Z-0409-03 | 4882 | [6 empty] | Ohmeda Medical Giraffe(tm) Omnibeds; and Ohmeda Medical Giraffe(tm) |  |  | October 21, 2002 | 2002 | OHMEDA MEDICAL | Electrical defects in neonatal incubators may lead to unintended movement | Other | Firm notified medical facilities' biomedical engineering department directors by certified letter 10/21/2002 that devices are usable with following precautions: (1) disable function of the elevating footswitch using the control panel and (2) place the 2 enclosed warning labels on incubators. Firm advised they will retrofit the incubators after evaluating possible solutions. Reply form enclosed.On 10/28/2002, firm notified consignees of a planned field correction to the OmniBed and Incubator devices by retrofitting isolation transformers to the units. The letter also advises that the firm has a retrofit kit available for correcting rechargeable battery back-up. The letter further advises that the units can continue to be used until corrective action is taken. |  | 1364 | 1364 units | Distribution to direct account medical facilities nationwide and international distributors.Foreign DistributorsCustomer NameAddress Line 1CityZipCountry DATEX-OHMEDA PTY LTD 149 ARTHUR STREET HOMEBUSH NSW 2140 AUSTRALIA SANITAS G M B H HOLUNDERSTRASSE 598 WALS B. SALZBURG A 5071 AUSTRIA MEDA N.V. OEYVAERSBOSCH 12 AARTSELAAR 2630 BELGIUM DYNAMED HEALTH CARE SYSTEMS 235 SHIELDS COURT MARKHAM ONTARIO L3R 8V2 CANADA CHINA HUA YANG TECHNOLOGY AND TRADE CORP (GROUP) BEIJING 100031 CHINA ESTIO SHANGHAI CO LTD RM 2202, ITALY BLDG SHANGHAI 200023 CHINA THE ORIENTAL SCIENTIFIC INSTR.I/E GRP,YINDU BLDG#67 BEIJING 100036 CHINA KIVEX LTD USSEROD KONGEVEJ 157 HORSHOLM 2970 DENMARK INSTRUMENTARIUM MEDICAL OU LIIVALAIA STR 22 TALLINN 10118 ESTONIA INSTRUMED INSTRUMENTARIUM OY ESPOO FIN-02630 FINLAND INSTRUMED PO BOX 300 HELSINKI FIN-00031 FINLAND INSTRUMENTARIUM CORP. INSTRUMED HELSINKI FIN-00510 FINLAND EUROCARE SRL RUE DES COMMERES COIGNIERES F-78310 FRANCE DATEX-OHMEDA GMBH DR ALRED-HERRHAUSEN DUISBURG 47228 GERMANY INTERMEDICA MANOS TZANIDAKIS ATHENS GREECE HONG KONG OXYGEN CHUN YAT ST, O INDUSTRIAL ES HONG KONG HONG KONG A KARLSSON HF BRAUTARHOLT 28 REYJAVIK 121 ICELAND PT MURTI INDAH SENTOSA JL SULTAN ISKANDAR JAKARTA 12240 INDONESIA OXYGENCARE LTD CORRIG ROAD DUBLIN 18 IRELAND OXYGENCARE LTD CORRIG ROAD DUBLIN 18 IRELAND DATEX-OHMEDA S.P.A. VIA CASSANESE, 100 MILANO ITALY AIR WATER INC 4019 OAZA YAMATO NAGANO 390-1701 JAPAN CHOONGWAE PHARMA CORPORATION 698, SHINDAEBANG-DONG SEOUL 156-757 KOREA, REPAL-ESSA MEDICAL &amp; SCIENTIFIC EQUIPMENT CO SAFAT 13036 KUWAIT INTERTRAME PUEBLA #14 COL. PE~NON V.E. MEXICO DF 15520 MEXICO SERVICIOS DE INGENIERIA EN MEDICINA SA DE CV,CUMBRES DE MEXICO D.F. 03630 MEXICO DATEX-OHMEDA B.V. KANTEMARSWEG 18 HOEVELAKEN 3871 AP NETHERLANDMEDISINK UTSTYR PO BOX 42 STOUNER OSLO 9 N-0913 NORWAY MUSCAT PHARMACY PO BOX 438 MUSCAT OMAN PRACTIKA PHILIPPINES, INC. 3RD FLR,DISPOPHIL BLDG, PASIG CITY 1605 PHILIPPINEPROMED SA UL KRAJEWSKIEGO 18 WARSAW PL 01-520 POLAND IBERDATA EQUIPAMENTOS SA RUA DR.RICARDO JORGE NO.4 AMADORA 2700-301 PORTUGAL CENTRE PERINATAL MEDICINE BUILDING 4 MOSCOW 109004 RUSSIAN FEABDULREHMAN ALGOSAIBI GTB ALGOSAIBI BLDG,AIRPORT RD RIYADH SAR-11411 Please see paper file for complete record. | Terminated | October 11, 2005 | 1086 | 1364 | Not_Computer
Z-0330-03 | 4896 | [6 empty] | Detection Reagent 1 of Digene's Hybrid Capture(TM) System Human |  |  | October 08, 2002 | 2002 | DIGENE CORP. | IVD test kit reagent may give false negative results | Other | Consignees were notified beginning 10/8/02 by phone and fax to discontinue use of the lot and reply to Digene representative with number of replacement reagents needed to retest patient specimens. Firm requests users to fax back the acknowledgement section of recall notification. Firm listed Digene's representative in Europe at 44 207 348 3500. |  | 217 | 217 kits | Kits were not distributed in the US. Kits were distributed to end users and distributors internationally.Genomed Ltd., United KingdomJK Trading, Czech RepublicInstitut fur Infektionskrankheiten, SwitzerlandAlpha Medical, SlovakiaHospital Universitario de Salamanca, SpainASL No. 18 Alba/Bra - Regione Piemonte, ItalyS.L. Durvis, SpainDr. Freischem &amp; Kollegen, GermanyDr. Med. Friedhelm Strube, GermanyLabor Dr. Peter Nae, GermanyUniversitatsklinikum Dusseldorf, GermanyDr. Med Eberhard Schneckenburg, GermanyDr. Med. Walter Franken, GermanyMedzinisches Labor Dr. Falke &amp; Partner, GermanyPathologisches Inst. Drs. Uhl &amp; Rosenkranz, GermanyArtzpraxis f. med. Mikrobiologie, GermanyZytologisches Labor Stocker, GermanyDr. Med. Ingo Scharmann &amp; Partner, GermanyDr. Med. Horparcy Gemeinschaftspraxis, GermanyDr. Med. Dietlinde Broer-Ahlers, GermanyGemeinschaftspraxis Leipzig, GermanyLandeskrankenhaus Feldkirch, AustriaApparategemeinschaft Allg auGmbH, GermanyDr. Menelaos Zaloumis, GermanyLabor 28, GermanyUniklinik Marburg Apotheke, GermanyHerm Dr. med. Reiss, GermanyFrauenklinik Hilden, GermanyDr. Med. Uwe Mundhenke, GermanyGem. Praxis. Drs. Loer, Winterhoff, GermanyDr. Thomas Weyerstahl, GermanyLabor Drs. Haubold Hofmann Jung Petihmann, GermanyFrau Dr. Med. Krompotic, GermanyDr. Med. Dietlinde Broer-Ahlers, GermanySteiermarkische Krankenanstaltenges, AustriaInstitut f. Medizinische Diagnostik, GermanyDr. T. Hermann / Dr. L. Holzer, GermanyDr. Josef Felixberger, GermanyZytologisches Labor Stocker, GermanyKlinikum der J. Gutenberg Uni Mainz, GermanyDr. Peter J. Stadler, Germany | Terminated | December 16, 2003 | 434 | 217 | Not_Computer
Z-0320-03 | 4890 | [6 empty] | Medtronic RF Enhancr SC Single-Curve Steerable Ablation Catheters, |  |  | October 29, 2002 | 2002 | MEDTRONIC, INC. | In some of the catheters, the catheter manipulator wire detached from the catheter insulating tip anchor during operation of the catheter deflection mechanism and after detachment penetrated the catheter wall and protruded from the shaft. | Other | A recall letter dated October 29, 2002 was delivered to the affected consignees in the United States by Medtronic representatives who retrieved the product. |  | 500 | 500 catheters | The catheters were distributed to 34 domestic consignees (hospitals) located nationwide in the United States. The product was also distributed to 2 Medtronic foreign subsidiaries and from there to foreign consignees located worldwide. | Terminated | January 01, 1990 | -4684 | 500 | Not_Computer
Z-0394-03 | 4898 | [6 empty] | 9 Fr. Percutaneous Sheath Introducer Kits, catheter introducer. |  |  | November 11, 2002 | 2002 | Arrow International | possibilty of leakage | Other | The recalling firm issued recall letters dated 11/11/02 to all their customers. The customers were instructed to examine their stock and cease using the affected lots. The letter also informed the customers that replacements would be shipped free of charge. |  | 25130 | 25130 kits | The product was shipped nationwide to 224 hospitals, 72 distributors and three sales represenatives. The product was also shipped to an international subsidiary account and an international distributor account. The government/military accounts are: VAMC, Houston, TX; VAMC, Portland, OR; VAMC, Birmingham, AL; VAMC, San Diego, CA; VA West Side Hospital, Chicago, IL; NIH, Bethesda, MD; VAMC, Decatur, GA; and VAMC, Minneapolis, MN.The international accounts are: Cecil E Mayo Ltd, 34 Morley Avenue, Rosenberg, NSW 2018, Sydney, Australia and Arrow Medical Products Ltd, 2300 Bristol Circle, Unit 1, Oakville, Ontario L6H 5S3, Canada. | Terminated | December 01, 2003 | 385 | 25130 | Not_Computer
Z-0389-04 | 28581 | [6 empty] | Lodox Statscan digital radiographic system. |  |  | May 10, 2004 | 2004 | Lodox Na Llc | The labels do not have proper identification and certification and the firm failed to provide adequate test and quality control procedures to establish certification by the manufacturer. | Other | The device in Illinois was re-exported. The device in Maryland was tested on 5/10/04 by a trained Lodox representative and proper labeling was reportedly attached. |  | 2 | 2 | Illinois and Maryland. | Terminated | October 06, 2004 | 149 | 2 | Not_Computer
Z-1350-02 | 4490 | [6 empty] | Disetronic D-TRON Insulin Infusion Pump, Models 8100001 D-TRON |  |  | June 17, 2002 | 2002 | Disetronic Medical Systems, Inc. | Some pumps started programming a bolus that was not initiated by the pump user, and the delivery of an unintended bolus was made unless interrupted by the device user. | Other | A notification letter dated June 17, 2002 was sent to patients who were using the pumps to direct them to check the buttons on their pumps to make sure they spring back when pushed and to make sure the acoustic alarms are turned on. All pump users then received instructions and mailing materials to return their pumps for inspection and replacement, if necessary. During February 2003 the firm changed their recall action from return for testing and possible replacement to recall and replace all of the pumps. They started sending a "Product Recall" letter about this to domestic patients on February 11, 2003. |  | 3357 | 3357 insulin infusion pumps in the U.S. | The pumps, imported into the United States, were distributed nationwide. The are an estimated 3,000 patients using the pumps in the U.S. | Terminated | June 01, 2004 | 715 | 3357 | Not_Computer
Z-1064-04 | 29393 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Osmetech OPTI-CCA Blood Gas Analyzer, Single Use, BP7564 OPTIÂżE-Glu Cassettes.The cassettes are labeled, Roche OPTICCA Cassettes E-Glu. There are 25 cassettes in each box. |  |  | June 22, 2004 | 2004 | Osmetech Inc | In some cases patient samples with high glucose concentrations, above the reportable range of the OPTI-CCA Analyzer, may be erroneously reported as 'Low'. | Other | Consignees were notified by letter on 6/22/04. A separate Recall Notice was mailed to distributors outside the United States with instructions on issuing the recall notice to their affected customers in the appropriate language. |  | 1087 | 1087 boxes | Nationwide and Argentina, Austria, Bolivia, France, Greece, Germany, Korea, India, Italy, Jordan, Netherlands, Sweden, Switzerland. | Terminated | January 12, 2005 | 204 | 1087 | Not_Computer
Z-0930-04 | 28972 | No | JCW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS Ambicor Penile Prosthesis, Penoscrotal Approach, Length: 14 cm, Diameter: 11 mm, Product Number 72401450 |  |  | April 29, 2004 | 2004 | American Medical Systems | The product is mislabeled as to length. | Other | Recall letters, dated April 29, 2004, state that it is important for the consignees to contact the firm's customer service for an exchange of the recalled products for new product. The letters recommend that the consignees send an attached letter to physicians who implanted the products. |  | 5 | 5 units | The affected products were distributed to consignees in the States of California and Texas, in Puerto Rico, and in the countries of Costa Rica and Taiwan. | Terminated | October 31, 2005 | 550 | 5 | Not_Computer
Z-0527-03 | 4679 | [6 empty] | Spinal 26 Gauge Quincke with Bupivacaine, Epinephrine, and |  |  | September 13, 2002 | 2002 | Abbott Laboratories | Mis-stenciling on tray. Tray stenciled Saddle Block instead of Spinal | Other | Letters to be sent out 09/12/2002. |  | 2980 | 2980 units | AR, CA, FL, MA, ME, NH, NM, PA, PR. There are no government nor foreign consignees. | Terminated | May 03, 2004 | 598 | 2980 | Not_Computer
Z-0874-05 | 31976 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nichols Advantage TSH - Third Generation |  |  | April 28, 2005 | 2005 | Nichols Institute Diagnostics | Does not meet performance claims in the labeling. | Other | Firm sent recall letter and response form April 28, 2005. Firm will follow-up to nonrespondants. |  | 1072 | 1072 | Nationwide, Germany, Greece and Spain | Terminated | March 16, 2012 | 2514 | 1072 | Not_Computer
Z-0847-05 | 31864 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Fresh Look * RadianceÂż * Moonlight, (phemfilcon) Contact Lenses * -2.75 median, Median Dia:14.5 * Rx only * Ciba Vision, Duluth, GA. The product has been distributed in twin packs containing six pairs of lenses. |  |  | April 21, 2005 | 2005 | Ciba Vision Corporation | The actual power of the lenses may be different from what is indicated on the package labeling. | Other | Consignees were notified via traceable mail on April 21, 2005. |  | 2250 | 2250 units | Nationwide, Europe | Terminated | October 21, 2005 | 183 | 2250 | Not_Computer
Z-1440-05 | 32931 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Aqueduct Hematology Workflow Manager, software |  |  | August 10, 2005 | 2005 | Beckman Coulter Inc | Potentail for misidentification of historical patient results within the Aqueduct Hematology Workflow Manager software. In instances where the workflow manager software receives the same patient identifier for more than one patient the Aqueduct Workflow Manager software merges all results into one Patient File. The may cause test results to be associated with the wrong patient. | Other | A Product corrective action letter to be mailed the week of August 8, 2005, to all customers directing that each patient is identified with a truly unique patient identifier. |  | 36 | 36 | AK, AZ, CA, DE, FL, IN, KS, MA, MI, MN, MO, NC, NY, OH, PA, TX, VA, WA | Terminated | April 03, 2012 | 2428 | 36 | Computer
Z-1065-04 | 29737 | [6 empty] | Sollux Systeme commercial tanning beds; Soleil Systems, Inc. |  |  | August 02, 2004 | 2004 | Soleil Systems, Inc. | The sunlamp products failed to incorporate a timer system and a control for termination of radiation emission in accordance with the Federal Sunlamp Product Performance Standard. | Other | Soleil Systems sent recall notification letters dated 6/28/04 to their dealers on 8/2/04, informing them of the noncompliances of failing to incorporate a permanently mounted timer and on/off switch in the sunbeds. Soleil will bring each sunbed into compliance and requested the dealers to review their sales records for any Sollux beds received between 4/28/04 and 11/01/03, and supply Soleil with a list of names and addresses of end users who received those units. Soleil sent out end user notifications dated 8/20/04 beginning on 8/23/04 and initiated the retrofit of the sunbeds on 10/25/04. |  | 70 | 70 beds | The sunbeds were distributed through dealers in Missouri, Texas, Virginia, Kentucky and Illinois | Terminated | March 28, 2006 | 603 | 70 | Not_Computer
Z-1168-2007 | 38273 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological Black Silicone Filiform Double Pigtail Ureteral Stent Set, 7.0Fr/24cm, sterile, Order Number (REF) 133724. |  |  | June 29, 2007 | 2007 | Vance Products Incorporated | The radiographic band may detach from the stent and remain in the patient after removal of the stent. | Other | Consignees were notified of the recall via letter dated 6//29/07, which instructed them to quarantine stock on hand and to fax number on hand back to Cook for return authorization. |  | 75 | 75 | Worldwide, including USA, Canada, and Japan. | Terminated | June 11, 2008 | 348 | 75 | Not_Computer
Z-1264-2007 | 44735 | No | LZK | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Spectrum Designs Medical-Nasal Dorsal Columella Implant, Catalog Number: S400-314S and Lot #31053; Serial Numbers: BAG612 ,BAG6I3, BAG6I4, BAG652, BAG654 and BAG655 |  |  | July 24, 2007 | 2007 | Spectrum Designs Inc | Mispackaging/Mislabeling: a portion of packaged Nasal LOT 31053 contained chin implants instead of the nasal implant. | Other | The firm initiated the recall on 07/24/2007 by contacting the customer (distributor) who filed the original complaint of mislabeling by email. The firm followed-up by sending recall letters (dated July 24, 2007) via FedEx or certified mail with signature confirmation to the one other direct account who received the recalled product on 08/03/2007. Customers were instructed to check their inventory for mislabeled products and call Spectrum at (800)239-6399 to report the number of this product being returned. They were given an address for product return. The firm indicated they will examine returned products, and scrap after examination. |  | 130 | 130 | Nationwide: California and China | Terminated | July 29, 2011 | 1466 | 230 | Not_Computer
Z-0289-2008 | 44752 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | A1CNOW+ with sampler 10-Test Kits (Product No: 3022), Metrika, Inc., Sunnyvale, CA 94085. (This test provides quantitative measurement of the percent of glycated hemoglobin (%A1C) levels in capillary (fingerstick) or venous whole blood samples. The test is used by professional healthcare providers to monitor long-term glycemic control in people with diabetes. The product is sold as a kit with a disposable monitor and reagents for ten A1C tests and is suitable for use at the point-of-care. |  |  | August 31, 2007 | 2007 | Metrika Inc | Incorrect Test Results: Due to manufacturing errors regarding temperature compensation factors, the device will give erroneously high results if the test is performed below 70 degrees Fahrenheit. | Process control | A September 14, 2007, Urgent Device Recall letter was issued to all Distribution Center Managers to return recalled product for replacement. Customer notification was made via a Dear Healthcare Professional letter dated Sept. 14, 2007, also titled Urgent Medical Device Recall. |  | 4286 | 4,286 | Nationwide including states of AK, AL, CA, CO, CT, DC, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MS, NC, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico. | Terminated | March 25, 2008 | 207 | 4286 | Not_Computer
Z-0288-2008 | 45333 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Clinical Chemistry Aeroset/Architect (c Systems)Bilirubin Calibrator, List Number (LN) 1E66-03 &amp; List Number (LN) 1E66-04, Abbott Laboratories, Abbott Park, IL 60064 |  |  | August 10, 2007 | 2007 | Abbott Laboratories Inc. | Incorrect Calibrator values: changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the standard. | Material/Component Contamination | An August 10, 2007, Product Correction letter was provided to all ARCHITECT cSystem and AEROSET System customers. The letter advises that changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the NIST standard 916a. Revised Bilirubin values for the calibrator lot numbers in this recall are attached to the recall notice. Six necessary actions were listed in the recall notice as: 1) Identify the lot numbers of Clinical Chemistry Bilirubin Calibrator(s) currently in use. 2) Discard the value sheet and replace with the appropriate values provided with the recall letter. 3) Enter the revised Bilirubin values. 4) Calibrate the assays. 5) Evaluate your quality control. 6) Evaluate your reference range and any shifts in patient results. |  | 7033 | 7,033 units total worldwide | Worldwide: USA, Canada, Mexico, Guatemala, Puerto Rico, Germany, Hong Kong, Singapore, Australia, New Zealand, Jamaica, Venezuela, Columbia, Ecuador, Chile, Uruguay, Argentina, Brazil, Thailand, Trinidad &amp; Tobago, Bahamas, Panama, Cayman Islands, and El Salvador | Terminated | April 19, 2012 | 1714 | 7033 | Not_Computer
Z-1653-2008 | 46686 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL700 Series analyzer - Human Blood Analyzer (Laboratory Device). |  |  | January 18, 2008 | 2008 | Radiometer America Inc | Reported calibration errors are not displayed on the analyzers parameter bar screen or on the patient results when the calibration number is between 32768 - 65535 and 98303 - 131070. This may cause the operator to perform measurements and obtain questionable results. | Software design | On February 28, 2008, a Field Action Notes Letter was issued advising them about a Mandatory Software Update. Customers were notified of the problem by fax and letter (Fed Ex). The Recall Letter instructed the user to disable the parameter bar feature and only rely on the traffic light system indicator which properly indicates system errors. The letter includes instructions to temporarily fix the problem and notice of a update of the software module that updates the parameter bar. If you have any questions contact Vincent Sigmund at (440)-871-8900 ext. 209. |  | 2481 | 2481 Units | Worldwide Distribution - USA and Canada - Products were sold in all fifty (50) US States. including twenty-nine 29 government accounts and 138 Canadian accounts. | Terminated | February 24, 2012 | 1498 | 2481 | Computer
Z-1363-2008 | 47340 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i LOCATOR ABUTMENT, Catalog/Item Number: LOA002, 2.0mm, BioMet 3i, Palm Beach, FL. |  |  | February 07, 2008 | 2008 | Biomet 3i, Inc. | Mispackaging: The Locator Abutment labeled as LOA002 may contain Catalog/Item # LOA003 product and visa versa | Other | Each USA customer was contacted by telephone and then by letter, Urgent Medical Device Recall, dated 2/15/2008, sent by fax. International customers were first contacted by E-mail then by fax letter also. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory for product return. They were also instructed to return the affected product to Biomet 3i. |  | 100 | 100 | Worldwide. | Terminated | July 10, 2008 | 154 | 197 | Not_Computer
Z-0627-2009 | 50139 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Excelsior SL-10 Microcatheter, placed in sterile pouch. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Packaging issue may result in failure to maintain sterile barrier. | Package design/selection | Consignees were notified by letter dated October 7, 2008. The letter identified the issue and affected product. Customers were instructed to: 1.) Discontinue use of and segregate recalled product, 2.) Complete and return the enclosed Reply Verification Tracking Form, and 3.) Package/ship the recalled product and send to the Boston Scientific Corporation (BSC).Return the completed Reply Verification Tracking Form via fax (800-836-4804) to the BSC Customer Service Call Center.A follow-up letter dated October 8, 2008 was also issued to customers. The letter notified customers of the correct date range of expiration for product within the scope of the recall which is September 2008 through October 2010. There were no other changes identified in the follow-up letter. |  | 21423 | 21,423 units | Worldwide Distribution. | Terminated | March 02, 2010 | 511 | 21423 | Not_Computer
Z-0906-2009 | 50737 | No | EMG | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Obwegeser Ramus Clamp w/ wide shield, 20 cm, 8", Right. ORAL-MAXILLO-CRANIO-FACIAL SURGERY, Original OBWEGESER Instruments. Walter Lorenz Surgical, Inc., Subsidiary of Biomet Microfixation, Inc. NOTE: None distributed domestically. |  |  | May 21, 2008 | 2008 | Biomet Microfixation, Inc. | The Biomet Microfixation Obwegeser Ramus Clamp, Wide Right 01-7996 were incorrectly etched as the left part number 01-7997. | Labeling mix-ups | Customers were contacted by e-mail notification with read receipt. Customers were notified in the e-mail to immediately discontinue use of the recalled product and return the affected items to Biomet Microfixation, Inc. The e-mail also included an Inventory Reconciliation form for customers to complete and return via fax (904-741-1801) to Biomet Microfixation, Inc. For questions related to this notice, please call 1-800-874-7711 or 904-741-4400 extension 448, Monday through Friday, 8 a.m.-4 p.m. Eastern Standard Time (EST). |  | 4 | 4 | International Distribution including countries of Ireland and Japan. | Terminated | September 11, 2009 | 478 | 4 | Not_Computer
Z-1620-2011 | 51208 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD FACS Sample Prep Assistant (SPA and SPA II) Instrument, |  |  | January 03, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | To prevent a needle stick while replacing the probe, the firm has initiated a Field Correction, and issued an improved Users Guide procedure (User Bulletin, Part Number 645050). | Nonconforming Material/Component | The firm, BD Biosciences, sent an information letter dated December 2008 with an enclosed New Probe Replacement Service/User Bulletin, to the customers. The letter described the product, problem and action to be taken. The customers were instructed to review the attached User Bulletin (Part Number 645050); ensure that their laboratory follows the updated procedure for removing and replacing probe; place the User Bulletin in their BD FACS SPA or SPA II User's Guide; always follow the instructions in User Bulletin to properly remove and replace probe, and complete and return the Effectivity Form via fax to 408-577-1409. Should you have any comments or questions, please contact BD Biosciences at 1-877-232-8995 in the US. Outside the US, contact your regional BD office or field service engineer for assistance. |  | 70 | 70 | Worldwide distribution: USA and countries including: Hong Kong, Australia, Belgium, Canada, El Salvador, Hong Kong, Japan, and Mexico. | Terminated | March 10, 2011 | 796 | 70 | Not_Computer
Z-1555-2009 | 52007 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Stratus CS Acute Care Troponin I CalPak. |  |  | April 28, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | The firm received and confirmed that some lots of Cardiac Troponin I (cTnl) CalPaks (CCTNI-CR) may exhibit calibration issues that cause lower than typical signal during calibration with the current population of cTnl TestPaks (CCTNI). This issue may result in failed calibrations due to "Slope Error" messages or falsely elevated patient sample and quality control (QC) results on calibrations that | Labeling False and Misleading | Siemens Healthcare Diagnostics, Inc. issued an Urgent Field Safety Notice dated April 29, 2008 to all affected customers instructing them to discontinue use of 3 CalPak lots and reduce the shelf-life for the remaining lots.For further information, contact Siemens Healthcare Diagnostics, Inc. at 1-302-631-7672. |  | 4415 | 4415 cartons | Worldwide Distribution - including United States, United Arab Emirates, Austria, Bermuda, Belize, Canada, Germany, Egypt, India, Japan, Saudi Arabia and Bengal. | Terminated | July 14, 2009 | 442 | 4415 | Not_Computer
Z-1625-2009 | 52189 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific EndoVive 20F (6.7 mm) Safety PEG Kit, Push Method for use with sterile Xylocaine ampule, latex free, sterile, Boston Scientific Corp., Natick, MA; Cat. no./REF 6647, inner pouch UPN M00566470 and outer carton UPN M00566471. |  |  | May 05, 2009 | 2009 | Boston Scientific Corp | The user may be unable to pass the device over the guidewire for placement in the patient, resulting in an inability to use the device. | Process control | Boston Scientific Corporation issued an "Urgent Medical Device Recall - Immediate Action Required" notice dated May 5, 2009 as follow-up to their "Recall of the Safety and Standard Push PEG Kits" dated December 30, 2008. The notice outlined the issue and Consignees were instructed to locate and return all affected product to the firm. In addition, the firm requests that users complete and return a Reply Verification Tracking Form via fax to 1-508-683-5578.For further questions, contact the Boston Scientific Corporation Field Action Team by phone at 1-508-683-4678. |  | 1490 | 1490 | Worldwide Distribution -- United States, India, Japan and Slovenia. | Terminated | October 13, 2010 | 526 | 10259 | Not_Computer
Z-1716-2009 | 52217 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | CONSERVE, TOTAL A-CLASS HEAD W/BFH TECH., REF 38AM-3635, O.D. Size: 36mm, Neck Length: +3.5mm, Taper: SLT, Sterile R, 1 each, Rx only, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | May 14, 2009 | 2009 | Wright Medical Technology Inc | Incorrect packaging, swapped sizes | Packaging process control | Wright Medical Technology, Inc. notified hospitals, surgeons and their distributors of the recall situation via an "Urgent: Medical Device Voluntary Recall" notice dated May 14, 2009 explaining the problem, the hazard involved and requesting confirmation that the product was removed from their inventory. The letter included a "Fax Back" response page to ensure the receipt of the notice.For further questions, contact Wright Medical Technologies, Inc. at 1-800-874-5630. |  | 15 | 15 units | Worldwide Distribution -- Including states of NC, TN, IA, IN, MD, WI, WV, OK, WA, PA, TN, TX, FL and CA and country of Canada. | Terminated | December 09, 2009 | 209 | 29 | Not_Computer
Z-0541-2011 | 52039 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | A-dec Decade Plus 1221 Dental Chair |  |  | April 01, 2009 | 2009 | A-Dec Inc | The design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. | Device Design | A-dec sent a letter dated March 23, 2009 to consignees/customers (doctors and network dealers). The consignees were informed that the design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. The potential for this occurs when returning the chair to the upright position. The consignees were advised to reprogram the Position 0/Entry-Exit preset to stop the back before it reaches a full upright position. This increases the gap between the armrest and the back slightly. To reprogram a chair, the firm included excerpts from a chair Instructions for Use (IfU) guide that detail the steps to reprogram a chair. Consignees can contact the firm's Customer Service at 503-538-9471, 1-800-547-1883, or via fax at 503-538-0276.The firm will replace the chair back assembly for all affected dental chairs. |  | 2723 | 2723 units | Worldwide distribution: US and Canada. | Terminated | December 16, 2010 | 624 | 2723 | Not_Computer
Z-1717-2009 | 52428 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | M-19 (Iridium 192): Source Production and Equipment Company, St. Rose, LA. |  |  | March 03, 2008 | 2008 | Source Production &amp; Equipment Inc | Radioactive source broke during use. | Other | Source Production and Equipment, Inc. notified the Distributor by phone and letter March 2008 and requested the return of all affected units. |  | 3 | 3 units | Nationwide Distribution -- GA, NV and NY. | Terminated | August 19, 2009 | 534 | 3 | Not_Computer
Z-1846-2015 | 71359 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems; |  |  | May 22, 2015 | 2015 | Siemens Medical Solutions USA, Inc | There exists a possible position sensor fault in the swivel base axis not being detected by the system software. When moving, the system could potentially exceed its usual speed, resulting in a collision. No injuries reported. | Software design | Siemens sent an Important Customer Safety Notice letter dated May 22, 2015, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter also informed customers that Siemens would contact them to schedule the installation of a software update to resolve the problem. |  | 475 | 475 | US Nationwide Distribution | Open, Classified |  |  | 475 | Computer
Z-1835-2015 | 71394 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optical Access System; 5mm x 100mm and 12mm x 100mm; |  |  | June 01, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling Trocars because they have received reports indicating that some may fracture at the distal tip resulting in fragments at the surgical site. | Under Investigation by firm | On 06/01/15 a customer notification letter was sent out by the firm.The notification letter instructs customers to stop using the affected lots of the recalled product, and should consider other trocar models reprocesses by Stryker as an alternative. The letter notes that attached is a Effectiveness Check Form and complete list of affected lots shipped to the facility.A Stryker Sustainability Solutions Sales Representative will be sent to check the inventory of the affected lots. Once the inventory is checked the affected lots will be removed. The firm will complete a Recall Effectiveness Check Form . The form is to be completed even if no affected product is found.The firm requests that the form be returned to the sales representative or fax to 480-809-4975. If the form indicates that affected devices remain in inventory a prepaid shipping label will be issued for the return of the product. Customers will receive a credit for all affected devices returned.The firm states that adverse reactions or quality problems experiences with the use of the product may be reported to the firm's hotline 888-888-3433 x5572, or at www.stryker.com/productexperience as well as through FDA's MedWatch Adverse Event Reporting program. |  | 147351 | 147,351 | Distributed Nationwide including Puerto Rico. | Open, Classified |  |  | 147351 | Not_Computer
Z-1839-2015 | 71373 | No | GAL | Absorbable surgical gut suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Fast Absorbing Plain Gut - MultiPass Needles, Sterile, Absorbable Surgical Suture, 5-0 (1.5 metric), 18" (45 cm), Conventional Cutting, Lot #JD6163, Exp: July 2019. |  |  | May 08, 2015 | 2015 | Ethicon, Inc. | Ethicon Fast Absorbing Surgical Gut (Plain) Suture is being recalled because the labeling on the box and foil are incomplete. | Labeling mix-ups | Ethicon initiated this voluntary recall by sending an Urgent: Medical Device Recall (Removal) Letter/Business Reply Form (BRF) dated May 8, 2015 to their affected customers via UPS. Stericycle will be handling returns. The letters identified the affected device and the reason for the recall. Customers were asked to examine their inventory immediately to determine if they had affected devices on hand. If so, they are to remove the affected product and communicate the issue to any necessary personnel. Customers with questions were advised to contact their local sales representative or call the Customer Support Center at 1-877-384-4266.For questions regarding this recall call 908-218-0707. |  | 1922 | 1,922 units | Worldwide Distribution -- US (nationwide) and Canada. | Terminated | June 06, 2017 | 760 | 1922 | Not_Computer
Z-1841-2015 | 71291 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | VersaTREK Instrument Series 240 REF 6240, Manufactured for Remel Inc. |  |  | May 19, 2015 | 2015 | Remel Inc | Use of the recalled product may result in false positive reports. | Release of Material/Component prior to receiving test results | Remel, Inc. a part of Thermo Scientific sent a Product Correction Notification dated May 19, 2015, to all affected customers. The notification states the problem with the device and that the recalling firm's Technical Services department will be contacting the customer to schedule a field service of the device to correct the issue. Customers were instructed to notify any personnel who need to be notified of the potential for a false positive resulting from rare inconsistent barometric pressure readings. Return the attached Medical Device Recall Return Response form acknowledging receipt of the notice of the affected product. If product was further distributed customers were instructed to contact those entities, advise them of the situation and provide them with a copy of the letter. Customers with questions were instructed to contact the Technical Services Department at 800-642-7029. For questions regarding this recall call 770-409-0713. |  | 143 | 143 devices | Worldwide distribution: US (nationwide) to states of AZ, CO, FL, GA, IA, ID, LA, MI, MO, PR, TN, TX, VA, and WA and the countries of Canada, People's Republic of China, Egypt, Finland, France, Indonesia, Israel, Jordan, Kenya, Republic of Korea, Latvia, Romania, Russia, Spain, Thailand, United Kingdom, Uruguay, and Vietnam. | Terminated | March 24, 2017 | 675 | 164 | Not_Computer
Z-1845-2015 | 71537 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Prestige Plus 185cm J-Tip: |  |  | June 04, 2015 | 2015 | Volcano Corporation | During a relabeling procedure at the manufacturing facility, the products were labeled with a shelf life in accordance with the date of the relabeling and not in accordance with the original shelf life date at the time of manufacture | Labeling mix-ups | Volcano sent an Urgent Medical Device Recalls lettersdated June 4, 2015 via Federal Express to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine product inventory, quarantine the affected product, complete the attached Customer Return form, and contact Volcano Customer Service at 800-228-4728, option 4 to arrange for the return of any product in inventory. |  | 162 | 162 total (both devices) | US Nationwide Distribution and one account in Canada. | Terminated | October 27, 2015 | 145 | 162 | Not_Computer
Z-1832-2015 | 71311 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EVAC 70 XTRA with Integrated Cable REF EICA5872-01 |  |  | May 15, 2015 | 2015 | ArthroCare Medical Corporation | Potential component failure causing device inoperability | Component design/selection | Smith &amp; Nephew sent an Urgent Product Recall notification letter dated May 15, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, quarantine immediately and return to the address listed on the inventory Return Certification Form. |  | 4111 | 4,111 units | US Nationwide Distribution | Open, Classified |  |  | 6466 | Not_Computer
Z-1837-2015 | 71294 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Surgical Kits containing Medtronic Covidien Devon(TM) Light Gloves: |  |  | April 27, 2015 | 2015 | DeRoyal Industries Inc | Surgical kits contain Medtronic Covidien Devon(TM) Light Gloves which were recalled for splits or holes | Nonconforming Material/Component | The firm, DeRoyal, sent an "URGENT! RECALL NOTICE" letter dated 4/27/2015 to their customers (i.e., end-user level, its direct accounts and hospitals) via overnight on 04/27/2015. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: a. send Recall Instructions and a listing of their accounts who were shipped affected product (also included were labels, Training forms, Record of Destruction Forms)b. notify Distributors and Hospitals with a list of products shipped to their account with instructions to place any remaining inventory in quarantine.c. document the lots/quantities labeled at their accounts and complete and return this information on the "Affected Product Listing" to DeRoyal via fax to 865-362-3716 or email to recalls@deroyal.com even if your distributor no longer has any inventory.d. conduct training at the hospitals and this training will be documented on the Recall Training form. e. leave Record of Destruction Forms at the hospital for completion/return to DeRoyal as products are discarded/destroyed.If you have questions or need assistance with the recall, please contact me at 865-362-6465 or tmarsee @deroyal.com. |  | 235785 | 235,785 units | Worldwide Distribution: US (nationwide) and country of Guatemala. | Terminated | December 29, 2015 | 246 | 235785 | Not_Computer
Z-1843-2015 | 71363 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | IMMULITE 2000/IMMULITE 2000 XPi Italian Cypress 3g Allergy, Catalog # T23L4 (40 test), Siemens Material Number (SMN) 10386079, IVD. |  |  | May 06, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | IMMULITE 2000/2000XPi Italian Cypress (T23L4), Lots 206 and 208, fail to meet specifications to support the shelf-life stability of 2 years (24 months) expiration date as indicated on the products label. | Under Investigation by firm | Siemens Healthcare Diagnostics issued an Urgent Medical Device Recall Letter (IMC-15-11A.US, dated May 6, 2015) in the US via FedEx, and an Urgent Field Safety Notice (IMC-15-11A.OUS, May 2015) outside the US on May 6, 2015. Customers were instructed to immediately discontinue use of and discard IMMULITE 2000/IMMULITE 2000 XPi Italian Cypress allergen, lot 206. The shelf-life of IMMULITE 2000/IMMULITE 2000 XPi Italian Cypress allergen, lot 208, has been shortened to June 30, 2015. Discontinue use and discard lot 208 after June 30, 2015. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within 30 days in the US and OUS and fax to Customer Care Center at 312-275-7795. |  | 1566 | Domestic: 17 kits, Foreign: 1549 kits | Worldwide Distribution-Distributed in the states of Florida, New Mexico and New York and the countries of Australia, Belgium, Chile, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, Guatemala, Israel, Italy, Japan, Kazakhstan, Latvia, Netherlands, Portugal, Romania, Russian Federation, Slovakia, and Spain. | Open, Classified |  |  | 1566 | Not_Computer
Z-1840-2015 | 71406 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The spare component Anti-HBs Cutoff Calibrator, 10 mLU/mL, Catalog number 26154. |  |  | May 13, 2015 | 2015 | Bio-Rad Laboratories, Inc. | The spare component Anti-HBs Cutoff Calibrator, Catalog number 26154, is recalled because the expiration date of 2016-04-30 printed on label in error. | Incorrect or no expiration date | Bio-Rad Laboratories Inc, sent an Urgent Product Correction letter dated May 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recalled calibrator is a spare component which is sold separately from the MONOLISA Anti-HBs EIA Kit and the kit is not affected by this recall.Customers were informed that testing was performed and the Anti-HBs 10 mLU/mL Calibrator was acceptable to use through the expiration date of the kit dated 2015-11-30.Customers with questions can call the firm at 1-800-224-6723, option #2, then #3. |  | 20 | 20 vials | US Distribution to the states of : AZ, CA, ID, MO, and NH. | Terminated | November 05, 2015 | 176 | 20 | Not_Computer
Z-2060-2010 | 55224 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System 120V, 50/60 Hz catalog number 809810, Terumo Cardiovascular Systems Corp., Ann Arbor MI. |  |  | June 11, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The HX2 Temperature Management System mixing valve may fail resulting in an unusable water channel and then the machine would go into standby mode. | Device Design | The firm, TERUMO Cardiovascular Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 11, 2010, to the consignees. The letter describes the product, reason for correction, what to do in the event of failure, potential hazard, correction, affected population and action to be taken by the customers. The customers are instructed to review the correction notice, assure that all users are aware of this notice, and complete, sign and fax the attached Customer Response Form to Terumo CVS Customer Service at 800-292-6551. The Terumo CVS Customer Service Dept. will contact you upon receipt of the Customer Response Form to schedule the replacement. The firm plans to inspect and correct all units in the field within 5 months.We encourage you to contact Terumo CVS Customer Service at 1-800-521-2818, Monday-Friday, 8 AM-6 PM EST, with any questions or concerns. |  | 100 | 100 | Nationwide Distribution: AR, CA, CO, CT, FL, HI, IA, IL, KS, LA, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA and WI. | Terminated | April 13, 2011 | 306 | 100 | Not_Computer
Z-2067-2010 | 55868 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | May 26, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity Laboratory software where results from interfaced external reference laboratories are not displayed or stored correctly. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the test result sent from the reference lab includes a comma in the result field. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated May 26, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue involving results from interfaced external reference laboratories that are not displayed or stored correctly if an unexpected alpha character is encountered in a result message that is expected to have a numeric value. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the result sent from the reference lab includes a comma in the result field. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to the their GE Service Representative or the GE Helpdesk at 888-778-3375. |  | 37 | 37 units | Worldwide distribution: California, Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee and Texas, and internationally to Australia, Canada, China, England, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 415 | 37 | Computer
Z-2093-2010 | 55578 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | The "RS-FBG Full Back Conductive Garment" with electrode attachments for use with interferential and muscle stimulation with a Transcutaneous Electrical Nerve Stimulation device (TENS) and/or muscle stimulator. |  |  | May 17, 2010 | 2010 | International Rehabilitative Sciences, Inc. | Device distributed with labeling and IFU outside of their 510(k) clearance. | Other | The firm, RS Medical, mailed "URGENT NOTICE OF VOLUNTARY CORRECTIVE ACTION" letters dated May 14, 2010, to customers on May 17, 2010 and letters to physicians/consignees on May 19, 2010. The letters included an updated RS-FBG Operation Manual and stated what the update covered; i.e. that the device is indicated for use on the "middle and upper back". The customers were ask to please read your new manual and replace your original RS-FBG Full Back Conductive Garment Operation Manual with this updated version and discard your old RS-FBG Garment Operation Manual. RS Medical stated that your RS-FBG Garment is not affected by this Corrective Action, only the Operation Manual. RS Medical informed the customers that you may continue to use the RS-FBG Full Back Garment with your RS Medical Stimulator as prescribed by your healthcare provider.If you have any questions or concerns, please call our Toll Free Customer Service number at 1-866-417-1293, email us at RSFBG@rsmedical.com or visit our special web page at www.rsmedical.com/rsfbginfo. |  | 190073 | 108134 units, 81,939 subject to recall | Nationwide distribution. | Terminated | August 15, 2012 | 821 | 190073 | Not_Computer
Z-2094-2010 | 56033 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Trapeze handles, part 148671, for use with Care Assist bed P1170 and Hill Rom 1000 bed P1160, Hill Rom Batesville, IN. |  |  | June 11, 2010 | 2010 | Hill-Rom, Inc. | The firm received a complaint regarding the plastic housing used to support the trapeze handle. The firm investigated and found that the manufacturer of the trapeze handle didn't correctly process the plastic component. Therefore the plastic had the potential to break prematurely. | Process control | The firm, Hill-Rom, sent an "Urgent Field Safety Notice" letter dated June 11, 2010, to its customers. The letter describes the product, problem and action to be taken by the customers. Hill-Rom provided the quantity of handles needed to replace handles customers received during the affected period. The customers were instructed to immediately remove and replace all Trapeze Handles that were delivered during the affected period; to replace the Trapeze Handle Assemblies (with a manufacturing date from January 2007 to January 2009) used on their Patient Helper with the enclosed unit(s); complete and return the Customer Response Form as soon as possible, or by July 26, 2010; advise the firm immediately if they have relocated the beds and/or the accessories to another location, and forward the letter to any other facility personnel they deem appropriate. If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 319 | 319 | Worldwide distribution: USA and Canada. | Terminated | February 18, 2011 | 252 | 319 | Not_Computer
Z-2104-2010 | 56161 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Merit Laureate Hydrophylic Guide Wire, Angled, Made In Ireland, CAT No: LWSTDA35260EX, 260cm, Sterile EO, Merit Medical Systems, Inc., South Jordan, UT 84095. |  |  | June 28, 2010 | 2010 | Merit Medical Systems, Inc. | Guidewires have decreased lubricity at the proximal end, inhibiting proper operation. | Process control | The firm, Merit Medical System, Inc., notified the Sales representatives by e-mail on June 28, 2010, and instructed them to contact all of their accounts and to advise them to discontinue use and return any unused products to Merit. A Product Recall Notice was to be completed by each facility under the direction of the rep and all activities were to be completed by July 9, 2010. Questions should be directed to Penny Goldman, Senior Product Manager at 801-208-4198 or Jackie Preece, Manager, Customer Service at 801-208-4365. |  | 199 | 199 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KS, MA, MD, MI, MN, NC, NH, NJ, NY, PA, and SD. | Terminated | October 19, 2010 | 113 | 964 | Not_Computer
Z-0492-2011 | 56405 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit Model 8000 only when used with Alaris Pump module or Alaris Syringe module |  |  | August 02, 2010 | 2010 | CareFusion Corporation | Than recall was initiated because CareFusion has identified identified a potential risk associated with the Alaris PC unit formerly known as Medley PC unit (PC Unit) model 8000 only when used with Alaris Pump module (Pump module) or Alaris Syringe module' (Syringe module), This recallnotification details the potential risk and recommended steps for users to take if they encounter a specific PC | Device Design | CareFusion will send by registered return receipt mail the customer notification letter to all customers usingthe affected Alaris PC Unit beginning August 24, 2010. The customer notification letter will be addressed tothe Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required toconfirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, selfaddressedmail, fax or email. The recall notice informed the customers of the reason for the notice, the unitsaffected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion is working to release a software update to address this display error and will contact your facility as soon as the software update is available. CareFusion does not require that you return your devices. Clinicians should weigh the risk/benefits to the patients beforecontinuing the use of the device.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems shouldalso be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 94146 | 94,146 units | Nationwide in the United States, Canada, Mexico, and the European Union | Terminated | July 25, 2013 | 1088 | 94146 | Computer
Z-2583-2010 | 56343 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SynFix-LR Spinal Implant System (Implant Holder for SynFix-LR and surgical technique brochure) |  |  | July 23, 2010 | 2010 | Synthes USA (HQ), Inc. | A potential for the implant holder tip to break off and lodge in the SynFix-LR implant, which pose a potential risk to the patient. | Error in labeling | Synthes Spine Hospital sent an ' URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated August 20, 2010 to all customers. The letter described the labeling issue and actions to be taken by the customers. The customers were instructed to review the supplementary labeling provided; contact their Synthes Spine Sales Consultant to arrange for provision of the Holding Forceps (p/n 388.401) prior to each SynFix-LR procedure; and complete and return the enclosed Verification Section form to Synthes via fax at 610-251-9005 or email: FieldAction@Synthes.com.If you have any questions, please call 1-800-620-7025 x 3500 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide Distribution: USA and the countries of Canada, Japan, Switzerland, Australia, Mexico, Malaysia, New Zealand, and Singapore. | Terminated | December 18, 2014 | 1609 |  | Not_Computer
Z-0350-2011 | 56753 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | September 01, 2010 | 2010 | Elekta, Inc. | It is possible to do couch corrections twice if both Couch Move Assistant in MOSAIQ and Remote Automatic Table in XVI are configured when in SYNERGISTIQ mode. | Other | Elekta sent an "Important Notice A340" dated September 1, 2010 to all customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to make a decision if XVI or MOSAIQ does the corrective table movement; configure and ensure that the system only uses RATM or CMA for the corrective table movement, when in SYNERGISTIQ mode, and complete and return the IMPORTANT NOTICE Confirmation of Receipt form via fax at +44(0)1293 654401. If you have any questions contact your local service representative at +44(0)1293 654200. |  | 23 | 23 units | Nationwide distribution: AZ, CA, DC, IL, KY, MA, MI, MS, NV, NY, OH, OR, PA, VT, and WV. | Terminated | December 14, 2010 | 104 | 23 | Not_Computer
Z-1650-2011 | 57404 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BIO-RAD BioPlex 2200 Syphilis IgG Control Set. Lot # 46636 |  |  | November 23, 2010 | 2010 | Bio-Rad Laboratories | An upward shift in the control values can result in a control value that falls outside of the current manufacturing ranges. | Other | The firm, Bio-Rad, sent a "Voluntary Field Correction: QC Ranges for Syphilis IgG Control Set Lot#46636" letter to Domestic customers via FedEx on November 24, 2010 and to the International customers via e-mail on November 23, 2010. The letter described the product, problem and actions to be taken by customers. The Customers were advised to temporarily establish their own means and acceptable ranges for Control Lot 46636 for the r15 and r47 markers. The firm will provide adjusted recommended means and acceptable ranges. Additionally, the customers were to complete and return the Customer Response Form for USA and/or the Subsidiary Response Form (Outside US) via fax to (510) 741-4846, Attn: Bio-Rad BPD-Product Support.If you have any questions about this notification, contact your local Bio-Rad Technical Support. |  | 202 | 202 | Worldwide distribution: USA and countries including: Canada, Czech Republic, France, Germany, and Italy. | Terminated | August 04, 2011 | 254 | 202 | Not_Computer
Z-1519-2011 | 57107 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pacific Hemostasis Âż Thromboplastin-DS, catalog #100354, packaged in 4 mL lypholyzed vials, 10 bottles/carton. Vials are labeled, in part ***Fisher Diagnostics a division of Fisher Scientific Company, LLC Middletown, VA 22645-1905 USA*** |  |  | October 06, 2010 | 2010 | Fisher Diagnostics, A Company of Fisher Scientific LLC | An in vitro diagnostic reagent used in coagulation testing may cause inaccurate clotting times in quality control and patient plasma samples. | Nonconforming Material/Component | The firm, ThermoFisher Scientific, sent two "Urgent Medical Device Removal" letters one via Federal Express on October 20, 2010 and certified mail for domestic customers on October 21, 2010. The letters described the product, problem and action to be taken by the customers. The customers were instructed to indicate the quantities of all affected product currently in stock at their facility; complete and return the attached Response Form/table to Fisher Diagnostics via fax to (540) 869-8129 or by mail to the attention of Technical Support; dispose of any remaining inventory of the affected lot numbers; provide a copy of this letter to their customers, and file this letter for future reference.If you have any questions regarding this notification, please contact our Technical Support staff at 800-528-0494, Option 2, or email techsupport.diagnostics.mtn@themofisher.com. |  | 1190 | 1190 kits | Worldwide distribution: USA including states of: Alaska, Florida, Kentucky, Ohio, New Jersey, New York, and Texas; and countries including: Argentina, Dubai, Italy, Mexico, Philippines, Romania, and Turkey. | Terminated | November 02, 2011 | 392 | 1190 | Not_Computer
Z-1741-2011 | 57705 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CR Tibial Insert, Size 5, 7mm, REF 163-2507, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | January 07, 2011 | 2011 | Ortho Development Corporation | Components of a replacement knee system could be mislabeled. | Error in labeling | The firm, Ortho Development, sent a "Field Notification Memo" date January 7, 2011 to its customers. The memo described the product, problem and action to be taken. The customers were instructed to promptly return any remaining parts of the listed product; locate all the parts in their inventory; identify any of the listed parts that may have been used in surgery; verify the quantity they have on hand and the quantity used in surgery and reconcile them with the quantity sent; if discrepancy exists, notify customer service resolution; prepare items for return and contact customer service to obtain RMA number. Once reconciled, arrangements would be made for sending new inventory and the return of unused parts.If you have any questions regarding this issue, please notify Regulatory Affairs Manager at (801) 619-3450. |  | 14 | 14 units (US), 2 (Intnl) | Worldwide distribution: USA including states of: AZ, CA, FL, MI, TX, and UT; and country of: Japan. | Terminated | December 27, 2011 | 354 | 72 | Not_Computer
Z-1728-2011 | 57814 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hc2 System Software Suite 4.0 version 3.0, available as component in Qiagen user guide, catalog #5050-00162, and as a component within a Qiagen PC, available as catalog # 5050-11140A. |  |  | January 12, 2011 | 2011 | QIAGEN Gaithersburg, Inc. | The software component of a diagnostic medical device may be defective and could cause incorrect patient values. | Software design | The firm, Qiagen, sent a "Notice of Product Field Action" letter dated January 12, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue the process of running multiple assay protocols on a single microplate to eliminate the potential for inaccurate patient result reporting, and to complete and return the acknowledgment section via fax to QIAGEN 661-702-3854 or scan the document and email to techservice-na@qiagen.com. The firm informed users that they will provide further detailed instructions for a workaround and software fix. A QIAGEN representative will contact the customers who indicated that they run multiple assays on a plate and are currently using the HC2 System Software Version 3.0 (Suite 4.0), to assist in the identification of those patients that may have been affected by the software defect.If you have any questions or concerns, please contact Qiagen Technical Services at 800-344-3631, Option 2. |  | 163 | 163 | The recalled product was sold to medical laboratories nationwide and in Canada. | Terminated | November 14, 2011 | 306 | 163 | Computer
Z-1895-2011 | 57948 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2 (part # 378600) is a microprocessor-controlled pipetting and diluting system. |  |  | January 18, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after erroneous, though potentially credible, results due to intermittent reagent level sense failures could be generated. During these failures, the instrument fails to produce a warning that a reagent vial is low or out of reagent. The low or absent levels of reagent used during these failures may lead to insufficient antibody being dispensed into a daughter tube there | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated January 18, 2011 to all their customers. The letter described the product, problem and actions to be taken. The customers were instructed to: 1) Follow PrepPlus 2 Instructions for Use (Part # 4277264-BA) for all reagent and specimen requirements prior to initiating a run on the PrepPlus 2 including:A) Ensuring all dead volume requirements for reagents as indicated in 7.11 and 7.12 of the Sample Processing sections and 12.2 of the Troubleshooting Guide. Appropriate volume of reagent in all vials per the planned number of tests to be run must contain at least the volume listed on the Worklist (plus an additional 10% of the full volume of the container to account for dead volume)".B) Ensuring specimen processing dead volume requirements are met as indicated in section B of Specimen Tube Size to ensure the minimum amount of specimen needed.C) Monitoring for any foaming/bubble accumulation in the reagent vial.2) A review of all clinically relevant results generated from samples prepared on the system be conducted. Abnormally dim or unexpected negative sample staining patterns should be reviewed to ensure correlation with clinical symptomatology and other diagnostic data. If the clinical information Is in accordance with the clinical data released, the validity of the results could be considered accurate and no retrospective action may be needed. Ifinconsistent data is observed, please adhere to good laboratory practices within your institution and consider data re-analysis and re-testing.The customers were also instructed to share this information with their laboratory staff and retain it as part of their laboratory Quality System documentation; if they have forwarded any of the affected product(s) listed above to another laboratory, provide a copy of the letter to them, and complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM via fax to: (786)639-7500/7501/750 |  | 369 | 369 units (140 in the U.S.) | Worldwide distribution: USA and countries including: Australia, Barbados, Belarus, Belgium, Canada, Croatia, France, Germany, Ireland, Italy, Japan, Malawi, Mozambique, Namibia, Poland, Puerto Rico, Russian Federation, Slovenia, South Africa, Spain, Swaziland, Sweden, Switzerland, Ukraine, United Kingdom, Zambia. | Terminated | September 02, 2014 | 1323 | 369 | Not_Computer
Z-1885-2011 | 57743 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | MDI(R), S.A.L.T. Airway Device, Supraglottic Airway Laryngopharyngeal Tube, Intended for use with 6.5mm -9.00mm ET tubes, Including SALT(TM) airway, ET tube clamp, Latex-free securing strap, Tongue blade, Lubricating jelly, NON STERILE, REF 79-580, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, Manufactured in the Dominican Republic of U.S. Materials |  |  | January 05, 2011 | 2011 | Microtek Medical Inc | Product contains Triad Lubricating Jelly which was recalled. | Other | The firm, Microtek Medical, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and place any S.A.L.T. Airway Device (79-580) in quarantine; if they have distributed any of the product, they should identify their customers and notify them at once of this recall; complete and return the enclosed URGENT: MEDICAL DEVICE RECALL-RETURN RESPONSE form via fax to: (662)244-3011 or email: tom.bonner@ecolab.com, and return any unused product inventory to Microtek Medical. The firm will provide the customers with a replacement product as soon as product becomes available.If you have any questions in regards to this recall, please feel free to contact the V.P.-Regulator Affairs at 1-888-844-0988 or email: Tom.bonner@ecolab.com. |  | 2230 | 223 cases of 10 units | Nationwide distribution: AZ, CA, CT, FL, IN, MS, NC, NH, PA, TX and WI. | Terminated | June 20, 2011 | 166 | 2230 | Not_Computer
Z-2046-2011 | 58240 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac High Energy Medical Linear Accelerator, a Trilogy Radiotherapy Delivery System; Model #s: H14, H18, H27, H29, HCX; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The wedge angle labeling on the wedge tray may not correctly match the wedge body. However, the wedge body is labeled with the correct wedge angle. | Labeling mix-ups | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 24, 2011, to all affected users. The letter described the product, problem and actions to be taken. The customers were instructed to IMMEDIATELY inspect all wedges to verify that the wedge angle label on the wedge body matches the wedge angle on the wedge tray; if any wedge is identified with mismatched labeling follow the written instructions provided in the letter; always measure wedge data at installation of new or replacement wedges; verify all treatment accessories of "matched" linear accelerators, and please advise the appropriate personnel working in your Radiotherapy Department of the content of this letter.Varian has enhanced their manufacturing quality control process to further ensure correct labeling of wedges. Also, Varian will replace any wedge that is identified as mislabeled, at the request of the customer. If you require further clarification, please feel free to contact your local Varian Customer Support Service Manager at: 1-888-VARIAN5 (888-827-4265) for USA and Canada; +14 41 749 8844 Europe; North American: support-americas@varian.com; Australia/New Zealand:support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America:soporte.al@varian.com. |  | 5541 | 5541 units | Worldwide distribution. | Terminated | November 05, 2012 | 620 | 5541 | Not_Computer
Z-2048-2011 | 58331 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Sheath Introducer Kit |  |  | March 25, 2011 | 2011 | Arrow International Inc | Labeling discrepancy - the outer corrugate label as well as a card inside the package state the product contains dry natural rubber - this is not duplicated on the unit packaging. | Other | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine any products with the listed lot numbers; complete and return the Immediate Attention Request form via fax to 800-343-2935 indicating how much product they have in stock and whether they need labels. In the event that an alternative approach is needed, the customers are to contact Arrow Customer Service for more information. Once the labels have been received and applied the product can be released from quarantine.An Arrow's Sales Representative will pay a personal visit to their customers and obtain the quantity and the amount of labels they need. If you have any other questions, feel free to contact your local sales representative or Customer Service at 866-396-2111. |  | 840 | 840 | Nationwide distribution: USA including states of: PA and NJ. | Terminated | January 29, 2014 | 1041 | 840 | Not_Computer
Z-2754-2011 | 59111 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 6F Zuma Guide Catheter with AL 1.5 Curve with side holes |  |  | June 08, 2011 | 2011 | Medtronic Vascular | Mislabeled; Product labeled as 6F Zuma Guide with AL 1.5 curve with side holes may contain a 5F Zuma Catheter with SAL 1.0 curve | Labeling Change Control | The firm, Medtronic, sent an "Urgent Medical Device Recall Notice(for US)" letter dated June 7, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to segregate the product so that it will not be used. A Medtronic Sales Representative will visit the customers to account for all units they received and initiate a credit as appropriate.If you have any questions, please call (978) 739-0371 or (978) 739-3259. |  | 5 | 5 units | Nationwide distribution: USA including states of: CA, FL, and MA. | Terminated | April 17, 2012 | 314 | 5 | Not_Computer
Z-0368-2012 | 60419 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer ZMR Femoral Spout &amp; Cone Bodies and Trabecular Metal Hip Prosthesis Femoral Stems: |  |  | November 16, 2011 | 2011 | Zimmer Inc. | Lack of assurance of proper metal fatigue strength due to metal grain structure anomaly. In 2006 a recall was initiated as a result of the grain anomaly; however, these 4 lots had been split off during normal production and had not been included in the 2006 recall. | Nonconforming Material/Component | On 11/16/11, Zimmer sent letters titled "URGENT MEDICAL DEVICE RECALL" to distributors via electronic mail and to risk managers and surgeons via FedEx indicating that the product must be immediately removed from use and quarantined." Stop using the affected lots and place affected product in quarantine." Review available product on the shelf" Identify those part/lot combinations listed in the recall communication" Immediately remove any implants found that match part/lot combinations listed in the recall communication and quarantine" Return affected lots and signature of acknowledgement to Zimmer |  | 1 | 1 unit (Distributed to Demark) | Worldwide distribution: USA (NY), AUSTRALIA, SWITZERLAND, and DENMARK | Terminated | August 28, 2013 | 651 | 4 | Not_Computer
Z-0870-2012 | 60786 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 73X Uni-Vent |  |  | November 04, 2011 | 2011 | Impact Instrumentation, Inc. | A piece of internal tubing in the 73X portable ventilator is of incorrect specifications for the device. | Component design/selection | Impact Instrumentation Inc. sent an URGENT Device Product Recall letter dated November 3, 2011 to all affected consignees. The letter identified the product, problem, instructions for product returns, and actions to be taken. The letter states that Impact will arrange for expedited return, at no charge to the customers of the affected units. The notice instructs customers to complete and return the Recall Tracking Form. If you need additional information, please call ImpactÂż Customer Service at (800) 969-0750 |  | 57 | 57 units | Worldwide Distribution - USA (nationwide) | Terminated | May 23, 2013 | 566 | 57 | Not_Computer
Z-0916-2012 | 60762 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Drystar Mammo. |  |  | March 02, 2010 | 2010 | AGFA Corp. | Labeling problem | Error in labeling | The firm, AGFA Corporation spoke to the customer directly via telephone concerning the mislabeled affected products. The firm also contacted the customer via e-mail on April 13, 2011 explaining the outcome of the customer complaint. Replacement product was sent to the customer. |  | 2 | 2 boxes | Product distributed in Florida. | Terminated | February 14, 2012 | 714 | 2 | Not_Computer
Z-0014-2012 | 59949 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Dynamic Multileaf Collimator (DMLC) |  |  | May 25, 2011 | 2011 | Elekta, Inc. | It is possible to load different patients in the record and verify (R&V;) system and in the multilead control software (MCS). | Device Design | Elekta, Inc. sent an "IMPORTANT NOTICE" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several precautions customers should use until Elekta finds a solution. Customers will be contacted when a solution is found. |  | 27 | 27 | Nationwide Distribution-USA including the state of CT, FL, GA, IA, MN, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA , and WA. | Terminated | January 14, 2014 | 965 | 27 | Computer
Z-3031-2011 | 59335 | No | JKD | C-peptides of proinsulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT C-Peptide Reagent Kit; a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of C-Peptide in human serum, plasma and urine on the ARCHITECT i System; Produced by Biokit S.A., 08186 Barcelona, Spain for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA and Abbott, 65205 Wiesbaden, Germany; list 3L53-25 |  |  | July 11, 2011 | 2011 | Abbott Laboratories | C-Peptide lots 03610K000 and 01711A000 have the potential to generate falsely elevated results with certain patient samples and with non-Abbott Controls. | Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated July 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the lots and destroy any remaining inventory of the lots according to laboratory procedures. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051.Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. |  | 3049 | 3,049 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Korea, Spain, Switzerland, Taiwan, Trinidad &amp; Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | February 10, 2012 | 214 | 3049 | Not_Computer
Z-0336-2011 | 56939 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | This recall has been initiated because Ethicon Endo-Surgery is aware of an increase in reports of leaking in the ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar, which potentially can lead to loss of pneumoperitoneum during a surgical procedure. Pneumoperitoneum leaks may lead to conversion from laparoscopic to open surgery. | Device Design | Ethicon Endo-Surgery sent an URGENT: DEVICE RECALL EVENT 2244 letter, dated October 8, 2010. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to stop using the product, examine all inventory to determine if they had the affected product on hand, and remove the affected products from inventory. Ethicon Endo-Surgery requested that each customer complete the Business Reply Form and fax it back to Stericycle at (877) 272-6141 with in 3 business days and also return any affected product to Stericycle.For question regarding this recall call (877) 598-5706 and reference 2244. |  | 671 | 671 lots totalling 2,479,134 instruments | Worldwide distribution: USA, ARGENTINA, AUSTRALIA,BELGIUM,BRAZIL,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, DOMINICAN REPUBLIC,ECUADOR, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, REPUBLIC OF, LEBANON, NETHERLANDS ANTILLES, NICARAGUA, PANAMASAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICASPAIN, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESURUGUAY, and VENEZUELA. | Terminated | January 25, 2012 | 485 | 671 | Computer
Z-2095-2010 | 55465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Evolis-Femoral Sizer Posterior Reference |  |  | September 11, 2009 | 2009 | Medacta Usa Inc | Medacta USA Inc is initiating a recall on the Femoral Sizer, an orthopedic manual surgical instrument, because the "L" and "R" markings are mislabeled which would cause incorrect usage. | Process design | The firm, Medacta USA Inc, sent an "Urgent Safety Notification" letter dated August 5, 2009, to the customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to check their inventory for devices listed, if found, to immediately sequester them and DO NOT use them in surgery. The customers were ask to please pack the affected femoral sizer(s) posterior reference for shipment to Medacta USA. Medacta informed the customers that Medacta USA personnel will contact them by telephone to provide them with an RMA number and the reworked femoral sizer(s) posterior reference will be returned as soon as they are available.If you have any questions, please feel free to contact the Medacta USA office at 805-437-7085 or contact Ms. Heather Neely directly at 805-886-1069. |  | 5 | 5 | Worldwide distribution: USA including states of Texas and Idaho and country of France. | Terminated | February 29, 2012 | 901 | 5 | Not_Computer
Z-3163-2011 | 59427 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoMobile, Hi-Art System, H-0000-0003, TomoTherapy 1240 Deming Way, Madison, WI 53717 |  |  | July 07, 2011 | 2011 | Accuray Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the TomoMobile Hi-Art System.TomoTherapy has discovered that when attempting to open the TomoMobile shielding doors, while the door hinges are in the locked position, the hinges may fail allowing the door to disengage from the shielding. | Device Design | The firm, TomoTherapy, sent an "Urgent Medical Device Correction" letter dated July 7, 2011 to its consignees/customers. The letter described the Issue, Product Affected, Recommended Action (advised consignees to not attempt to open the shielding doors when the hinges are in the locked position. The Hinges should always be in the unlocked position when opening the doors) and Resolution. A TomoTherapy field service engineer or an authorized service representative will contact the customers when a resolution is available.If you have any questions, please contact the TomoTherapy Technical Solutions Center by email at support@tomotherapy.com or phone at: North America: 1-866-368-4807; Belgium: 0800 38783; France: 0805 631 565; Germany: 0800 000 7541; Italy: 800 986 399; Netherlands: 0800 020 1364; Spain: 800 300049; Switzerland: 0800 001927; United Kingdom: 0808 238 6035; China/CNC (Northern): 10 800 712 1701; China (Southern): 10 800 120 1701; Hong Kong: 800 967912; Japan: 0044 22 132374; Singapore: 800 1204 683; South Korea: 0079 81 4800 7204, all other locations: +1 608 824 2900 or +32 2 400 44 44. |  | 4 | 4 | Worldwide distribution: USA including states of: AZ, CA and OK; and country of: United Arab Emirates | Terminated | June 29, 2012 | 358 | 4 | Not_Computer
Z-0036-2014 | 66335 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Mini CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply with an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 ÂżÂż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0419-2015 | 69584 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Disposable Falope-Ring Band Applicator Kits. REF 006889-901 and REF 005280-901 (with 8 mm Disposable Trocan). Sterile EO. Single Use Only. The disposable Falope-Ring Band Applicator is a sterile, two-band, dual-incision instrument. The two-band applicator enables the physician to occlude both Fallopian tubes without removing the applicator from the abdominal cavity. The Falope-Ring Band and Applicator Kits are indicated for use for female sterilization (permanent contraception). |  |  | October 21, 2014 | 2014 | Gyrus Acmi, Incorporated | All packages of Falope Ring Band Applicator kits are being recalled due to compromises in sterile packaging that could be associated with bacterial contamination that might lead to patient infection. | Process control | Olympus sent an Urgent Medical Device Recall letter dated October 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Olympus requires you to take the following action: I. Immediately cease any further use of any affected product you have, remove it from your inventory and quarantine it until it is shipped back to us. 2. Call your Olympus customer service representative at 1-888-524-7266 to obtain a Returned Goods Authorization so that you may return the product with no charge to you. Olympus will issue a credit t or replacement to your facility y for any returned product. 3. Please note on the enclosed questionnaire that you have received this information. 4. Fax the completed questionnaire to 484-896-7128 regardless of whether you have any affected inventory at your facility. ln addition, if you may have further distributed this product, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter. |  | 74792 | 74,792 units | Worldwide distribution. US Nationwide, Canada, Vietnam, and Europe | Completed |  |  | 74792 | Not_Computer
Z-1729-2011 | 57631 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) Creatinine Flex(R) Reagent Cartridges |  |  | December 27, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed the potential for falsely elevated creatinine results with specified CREA Flex(R) reagent cartridge lots when processing tests from a well that has been punctured. This issue may affect creatinine results for quality control (QC) and patient samples. | Material/Component Contamination | The firm, SIEMENS, sent an "Urgent Field Safety Notice" dated December 2010 to all customers. The letter described the product ( Dimension(R) CREA Flex(R) Reagent Cartridges DF33A lots BB1286, CA1260 and FB1251), problem and actions to be taken. The customers were instructed to please discontinue use and discard any remaining inventory of the product; complete and return the fax form to the Customer Service Support for replacement of the product; if they must continue use of the product until they receive a replacement, open a new well set every 8 hours to minimize any impact to CREA test results, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the Technical Solution Center at 302-631-8467.If you have any questions, contact the Technical Solution Center at 800-441-9250. |  | 20177 | 20,177 | Worldwide distribution: United States and the following foreign countries: Austria, Belgium, Bosnia and Herzegovina, Canary Islands, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Great Britain, Greece, Hungary, Italy, Latvia, Lithuania, Kuwait, Macedonia, Morocco, Norway, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey and the United Arab Emirates. | Terminated | December 21, 2013 | 1090 | 20177 | Not_Computer
Z-0041-2014 | 66344 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner ID ENT00001 |  |  | December 14, 2009 | 2009 | Xoran Technologies, Inc. | Failure of the gantry to rotate. The complaint alleged that the scanner would emit radiation but would not rotate. | Radiation Control for Health and Safety Act | The service report in response to this complaint stated that the Head select relay board, chip and fan were replaced. The machine (ENTOOOO I) was subsequently tested. All tests were successful. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1732 | 320 | Not_Computer
Z-2080-2012 | 62432 | No | IZH | Mammographic x-ray system. | Radiology | Radiology | 510(k) | Bella Blanket Protective Coverlet, for Mammography, 18cm x 25cm or 24cm x 30cm, Product Code: 311. The purpose of the product is to remove the cold by placing this sheet between the patients breast and the imaging receptor plate. |  |  | July 02, 2012 | 2012 | Beekley Corporation | Imaging artifact can occur on mammogram result | Nonconforming Material/Component | Beekley Medical issued "Urgent: Medical Device Recall" notifications dated June 29, 2012 that were sent via certified mail USPS on July 2, 2012. The product issue was was described and customers were provided the recommended actions. Customer Service number is 1-800-233-5539. |  | 86000 | 86,000 units | International Distribution including Nationwide(USA), and the countries of Japan and Canada. | Terminated | July 09, 2013 | 372 | 86000 | Not_Computer
Z-1948-2013 | 65472 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Colorimetric (CM)/ Rate (RT) incubator wear pads. Product Codes: 1) 6801375, 2) 6801890, 3) 6802445, 4) 6802413 |  |  | June 06, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling CM/RT wear pads which are a component in the CM/RT incubator due to becoming worn before their replacement. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 6, 2013 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that an OCD trained service personnel will contact them and schedule a replacement of the CM/RT wear pads. OCD requested customers to complete and return the enclosed Confirmation of Receipt form no later than June 17, 2013. Customers are asked to place a copy of the notification by each VITROS System in their facility. Foreign affiliates were notified by email on June 6, 2013 of the issue and instructed to notify their consignees of the issue and actions. If you have any questions or need additional information, please call Customer Technical Services at 1-800-421-3311. |  |  |  | Worldwide Distribution - USA Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Puerto Rico and the countries of Mexico, Singapore, Venezuela, England, Germany, Italy &amp; Spain. | Open, Classified |  |  |  | Not_Computer
Z-0039-2014 | 66338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Minicat00158 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evaluation of the hazards reasonably related to defect or the failure to comply with the Federal standard; An evaluation of the health ha | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  |  |  | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 |  | Not_Computer
Z-0040-2014 | 66340 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT00202 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xorans specifications. | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0756-2013 | 64010 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Buffered Sodium Citrate (9NC) Plus Blood Collection Tubes; Sterile; IVD; REF 363083; 13 X 75 mm, 2.7 mL |  |  | November 12, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer 2.7mL Plus Citrate Blood Collection Tube due to extended Partial Thromboplastin Time (aPTT) test results reported in two customer complaints. | Material/Component Contamination | BD sent "ATTENTION: PRODUCT RECALL" letters dated November 12, 2012 via UPS to distributors and customers instructing them to examine their inventory immediately and discontinue the shipment of the affected product. All customers were instructed to return all products from the lot affected, by following the instructions on the enclosed packing slip,and BD will send them replacement products at no charge. Contact the firm at 10201-847-4267 for assistance. |  | 2944700 | 2,944,700 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, India, Japan, Nepal, Malaysia, Singapore, Brazil, Costa Rica, and Virgin Islands. | Terminated | September 06, 2013 | 298 | 2944700 | Not_Computer
Z-0445-2013 | 63410 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | The Is-anti-Sm test kits****Catalog No: 720-240*** |  |  | August 06, 2012 | 2012 | Diamedix Corporation | Diamedix Corporation recalled their Is-anti-Sm Test Kit due to the Positive Control recovering high out of specification. | Nonconforming Material/Component | Diamedix sent a Customer Notification letter to all affected customers via Certified Mail. The letter identifed the affected product. problem and actions to be taken. The letter instructed customers to contact Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any remaining affected product. For questions call 305-324-2314. |  | 65 | 65 Test Kits | Worldwide Distribution - USA Nationwide including the states: CA, ID, NJ, NY, OH, PA, PR, RI, TN, and TX, FL, and the countries of Italy, Chile and Ecuador | Terminated | August 08, 2014 | 732 | 65 | Not_Computer
Z-2149-2014 | 68665 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SIEMENS ADVIA Centaur CP Immunoassay System BNP kit; SMN 10309044 - 100 test |  |  | June 23, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Complaints of failed calibration. | Other | Siemens sent an Urgent Medical Device Correction (UMDC) letter to all affected Siemens Healthcare Diagnostics customers in the United States on June 23, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 23, 2014 via hard copy, email, and/or fax per each country's local regulations. These notices inform the customer of the potential for low calibrator deviation failures when using ADVIA CentaurÂż Systems BNP kit lots ending in 172 on the ADVIA CentaurÂż CP due to an error on the master curve card. Siemens informs customers if a successful calibration is achieved, the results obtained are valid. If customers are unable to obtain a successful calibration on the ADVIA CentaurÂż CP, they should switch to an alternative lot of BNP assay. The notices should be reviewed with the facility's laboratory or medical director.Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 1191 | 1,191 kits | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada, Albania, Argentina, Australia, Bosnia Herzeg., Brazil, Chile, Colombia, Egypt, Georgia, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Japan, Kazakhstan, Lebanon, Malaysia, New Caledonia, Oman, P.R. China, Pakistan, Philippines, Qatar, Rep. of Yemen, Republic Korea, Russian Fed., Saudi Arabia, Serbia, Singapore, South Africa, Taiwan, Thailand, U.A.E., Ukraine, Venezuela, Vietnam, Austria, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece,Hungary,Iceland, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Romania,Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 31, 2016 | 647 | 2152 | Not_Computer
Z-1793-2012 | 61332 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Tibia Plateau (component of the Columbus REVISION Knee System) |  |  | February 21, 2012 | 2012 | Aesculap Implant Systems | During an internal quality inspection, the tibia plateau was found to be out of specification. There is a potential for misalignment between the tibia plateau and attachment of the extension stem. | Nonconforming Material/Component | Aesculap Implant Systems (AIS) initiated their recall of this product by first telephoning consignees and then on February 21, 2012 by sending an IMPORTANT RECALL NOTIFICATION letter dated February 21, 2012 to all consignees. The letter identified the affected product, problem and actions to be taken. The notification told consignees not to use the affected product and to return any stock to AIS (USA) promptly. For questions contact Scott Stephan at (610) 984-9239 or Christian Gabriel at (610) 984-9300. |  | 8 | 8 | Nationwide Distribution including PA and TX. | Terminated | January 15, 2014 | 694 | 8 | Not_Computer
Z-0694-2012 | 60851 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | MERZ AESTHETICS, 27G X 1.0 " (25MM) Cannula, Sterile R, QTY 20, Model 9015M1X20, Distributed by Merz Aesthetics, In., Manufactured by TSK Laboratory. |  |  | December 16, 2011 | 2011 | Merz Aesthetics, Inc. | Merz Aesthetics is initiating a medical device recall on TSK Closed Tip Cannula (27GA x 25mm) due to a labeling error on the secondary (outer) packaging. The outer package label states that the carton contains Cannulas that are 27GA x 25 mm in size instead of 27G x 38 mm. Length of the cannula is a matter of user preference and is unlikely to impact patient safety. | Packaging process control | Merz Aesthetics Inc, initially contacted customers by telephone on December 16, 2011. They were informed that the wrong size cannula were sent with their last order and would need to be returned and a replacement order would be sent. . Merz Aesthetics sent a "Urgent Medical Device Recall" letter dated January 6, 2012. The letter identified the product the problem and the action needed to be taken by the customer. In addition, if you have further distributed this product, please identify your customers and notify them of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.You may contact Merz Aesthetics Customer Service at (866) 862-1211 or (262) 835- 3333 ,for instructions regarding the return and immediate replacement of the product, or if you have any further questions or concerns.. |  | 5 | 5 boxes (20 units per box) | Nationwide Distribution including the states of CA, NY, OK, and TN. | Terminated | July 27, 2012 | 224 | 5 | Not_Computer
Z-0037-2014 | 66337 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply wvith an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0137-2014 | 66353 | No | JCA | Bleeding time device. | Hematology | Hematology | 510(K) Exempt | Surgicutt fully automated incision making device for bleeding time determination. ITC 8 Olsen Avenue, Edison, NJ 08820. 732.548.5700. www.itcmed.com. A subsidiary of Thoratec Corporation. |  |  | August 22, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that the blade in their Surgicutt with Blotting Paper (Product Code: SUB50I; Lot: AK301M) devices may be loose. | Nonconforming Material/Component | ITC sent a Medical Device Recall letter dated September 20, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove/quarantine affected product in their inventory and to return the affected product to ITC. Customers with questions were instructed to call 1-800-631-5945.Questions regarding this recall call 732-548-7000, ext 4476. |  | 300 | 6 boxes (Each box contains 50 Surgicutt Devices plus 55 sheets of Blotting Paper). | Nationwide Distribution including Arizona andTexas.. | Terminated | June 09, 2014 | 291 | 300 | Not_Computer
Z-1917-2015 | 71240 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Abnormal Power Source Switching. | Software design | HeartWare sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the enclosed Patient Communication and familiarize the contents, forward the notice to those individuals withing their organization who need to be aware of its contents, identify their patients currently supported by the HVAD System, distribute the Patient Communication to their patients via FedEX or some other reliable means of communication. Customers with questions were instructed to contact their HeartWare representative should they need any assistance with the process. Customers were asked to complete, sign, and return the Acknowledgement and Completion Form to HeartWare. Customers with quesitons were instructed to call 1-888-494-6365. |  | 3747 | 3,747 currently in use. | US including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1912-2015 | 71305 | No | MAL | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Gelseal Straight Diameter 8mm Length 15cm. |  |  | May 15, 2015 | 2015 | Vascutek, Ltd. | Product was labeled as a 30cm graft, when in fact the graft in the box was a 15cm graft. | Error in labeling | On 5/15/2015, Vascutek issued a URGENT FIELD SAFETY NOTICE to their consignee informing them of the corrective action of the affected product. The information included details of the affected devices, including potential risks and actions to be taken by medical staff and distributors. |  | 12 | 12 | Distributed in the state of Michigan (USA), and the countries of United Kingdom, Australia, Poland, and Germany. | Terminated | October 20, 2015 | 158 | 12 | Not_Computer
Z-1914-2015 | 71517 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DLP Femoral Arterial Cannula, models 57414, 57417, and 57421, and DLP Femoral Cannula Replacement Kit, Models 96017 and 96021. Sterilized using Ethylene Oxide. |  |  | June 12, 2015 | 2015 | Medtronic Perfusion Systems | Medtronic is recalling select lots of DLP Femoral Arterial Cannula product due to recent reports in which the user had difficulty or was unable to connect the cannula to the perfusion circuit tubing. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter, dated June 2015 to customers beginning June 12, 2015. The letter identified affected product, described the issue, requested that unused inventory be quarantined and returned to Medtronic. Customers should contact Customer Service at 800-854-3570 and reference RADAR #241958 to initiate a return and credit of product. Additionally a Customer Confirmation Certificate should be completed and faxed to 651-367-0612. |  | 1304 | 1304 devices | Worldwide Distribution - US Distribution including the states of : AK, AZ, CA, CO, IN, KS, LA, MN, MT, NM, NY, OH, OR, PA, TX, UT, and WI., and to the countries of : Belgium, Brazil, Canada, Colombia, Denmark, France, Germany, Greece, Hungary, India, Ireland, Italy, Korea, Morocco, Netherlands, Poland, Qatar, Russian Federation, Saudi Arabia, Slovakia, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | March 09, 2016 | 271 | 1304 | Not_Computer
Z-1849-2015 | 71451 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | The VariSource iX series afterloader systems are computer controlled remote electro/mechanical systems used for medical purposes, for placing a cable incorporating an irradiated iridium seed internally or close by a malignant tumor or tumor bed in a practice known as brachytherapy. |  |  | May 18, 2015 | 2015 | Varian Medical Systems Inc | After an application freeze and restart of the VariSource iX series by power cycling, the Partial Fraction generated by thesystem will not be correct. The application freeze issue affects only the systems equipped with magnetic Hard Disk Drives (HDD). | Software Design Change | All the consignees affected by this recall will be notified by:The Urgent Medical Device Correction Letter/ Field Safety Notification will bedistributed to all affected users, with a description of the problem.The Letter will also be distributed to the Varian Sales, Marketing, and Serviceorganizations, informing them of the issue.Additionally, the letter will be posted to the customer support site:http:/ /www.MyVarian.com. |  | 55 | 55 | Worldwide Distribution. | Open, Classified |  |  | 55 | Not_Computer
Z-1915-2015 | 71361 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | DePuy Synthes 6.5 mm Titanium Recon Screws w/T25 Stardrive (70mm, 75mm, 80mm, 85mm, 90mm, 95mm, 105mm); part of the Expert Lateral Femur Nail System and indicated to stabilize Subtrochanteric fractures, Ipsilateral neck/shaft fractures, Femoral shaft fractures, Impending pathologic fractures, and malunions/Non-unions. |  |  | May 27, 2015 | 2015 | Synthes, Inc. | certain lots of the Titanium Recon Screws (6.5mm ) were found to contain the incorrect insert GP0760-A (Synthes Titanium Buttress Locking Plate System), instead of the correct insert GP0706-E (Synthes Titanium Intramedullary Nails and Components). | Process control | DePuy Synthes sent an Urgent Medical Device Safety Notification, dated May 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions Required:Please take the following actions:Âż Review the information listed in this notification.Âż Forward this notice to anyone in your facility that needs to be informed.Âż Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided.Send a copy of the completed Verification Section by:Âż Fax: (888) 912-2182 orÂż Scan/email: Synthes4844@stericycle.comÂż If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification.Âż Maintain a copy of this notice.If you have any questions, please call (610)719-5450.Thank you for your attention to this issue. |  | 276 | 276 | Nationwide Distribution. | Terminated | December 18, 2015 | 205 | 276 | Not_Computer
Z-1847-2015 | 71342 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Navigated Cannulated Taps |  |  | May 26, 2015 | 2015 | Medtronic Navigation, Inc. | Medtronic Navigation is initiating a field correction due to the potential for injury which could occur as a result of cleaning the Medtronic Navigated Cannulated Taps. | Error in labeling | Consignees were notified via Urgent Medical Device Correction Notification on 5/26/15. The letter identified the affected product and described the issue. This field correction does not require the return of the subject products. Revised instructions for use are being provided to the identified consignees. |  | 15798 | 15,798 | Worldwide Distribution. | Terminated | October 23, 2015 | 150 | 15798 | Not_Computer
Z-1848-2015 | 71560 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Slipped Capital Femoral Epiphysis (SCFE) Drivers of the Free-Gliding SCFE Screw System with catalogue number SCF-MLD265 and SCF-MLD273. |  |  | June 15, 2015 | 2015 | Pega Medical Inc. | SCFE Drivers of the Free-Gliding SCFE Screw System are recalled because the handle of the instruments may detach from the shaft if high torques are applied in counter-clockwise rotation. | Reprocessing Controls | Pega Medical sent an Urgent Field Safety-Recall letter via email on June 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advise on action to be taken by the customer: Âż Identify the device with the same catalogue number and lot numberÂż Return the device to Pega Medical Inc. to the address shown below using the FedEx account.For further questions please call the firm at 450-688-5144 x. 242. |  | 4 | 4 units | US Distribution to the states of : FL and TN. | Terminated | December 29, 2015 | 197 | 4 | Not_Computer
Z-1916-2015 | 71424 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Bullet-Tip PEEK VBR/IBF System, 32mm(l) 13mm(H) Implant 32-13-32 UDI (01)008464680341 08(17)191218(10)200827; T-Plus PEEK VBR/IBF System, 10mm(W) x 27mm(l) x 13mm(H) Implant, 6degree lordotic 30-T-13-6 UDI (01) 00846468032920(17) 200116(10) 202726; and T-Plus PEEK VBR/IBF System, 10mm(W) x 36 mm(l) x 13mm(H) Implant, 6degree lordotic 30-T-1036-13-6 UDI (01)00846468032760(17)200115(10)202582. These products are part of the lnterbody Fusion Vertebral |  |  | May 21, 2015 | 2015 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | Three lots of the Bullet -Tip VBR 13mm x 32 mm implants are mislabeled with a light green colored label, used for 16 mm trial instead of a lime green label, used for a 13 mm trial. | Error in labeling | RTI Surgical sent an" Urgent Medical Device Voluntary Recall " Notification dated May 21, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate Actions: We are aware that one of this product was distributed to you. This one device has already been returned to RTI Surgical so no further action is needed. This notice is for your records.RTI takes this error seriously and has initiated a corrective and preventative action investigation to prevent recurrence. We apologize for any inconvenience this may have caused. For further questions please call (906) 226-4489. |  | 18 | 18 | US Distribution to the states of : FL and NY | Terminated | October 29, 2015 | 161 | 18 | Not_Computer
Z-1921-2015 | 71182 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholesterol Gen.2. |  |  | May 07, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Possible drug interference when using the Trinder reaction method on multiple Cobas Analyzers. Results of Trinder tests may be falsely decreased when the tests are run with urine samples (Creatinineplus and Uric Acid applications only) or blood samples collected before sufficient time has lapsed after the administration of the drugs Acetaminophen and N-acetylcysteine. | Under Investigation by firm | Roche Diagnostics initiated a voluntary recall and the Urgent Medical Device Correction notices were sent via UPS Ground on 05/07/2015 to customers. Customers were notified of the interference and the method sheets and package inserts will be updated with the new interference claims. The firm plans to monitor any non-responding accounts and address those customers with follow-up attempts to distribute recall notices. The firm plans to conduct level B effectiveness checks via fax and phone. Until the new method sheets are available, the Urgent Medical Device Correction serves as the updated labeling. Roche is not requesting return of the affected product. Customers may contact Todd SieskyCorporate Communications Director, Phone: 317.521.3911, Fax: 317.521.4696, E-mail: todd.siesky@roche.com. For technical questions, contact the Customer Support Center at 1-800-428-2336 |  | 6255 | 6,255 analytical units distributed | US (nationwide) including PR. | Terminated | March 24, 2017 | 687 | 6255 | Not_Computer
Z-1919-2015 | 71292 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeHealth, IRMA TRUPOINT model 443914, typically used in point-of-care settings. For the in vitro measurement of various critical care analytes in human whole blood. These analytes includes: pO2, pCO2, pH, iCa, Na, K and Cl. |  |  | April 29, 2015 | 2015 | LifeHealth, LLC | The barometric pressure measurement output of this particular IRMA TRUPoint SN 443914 device will drift over time. | Nonconforming Material/Component | Consignees were sent via e-mail a LifeHealth LLC " Urgent" letter dated April, 29th, 2015. The letter described the Purpose of this letter, Reason for the Voluntary Recall, Risk to Health, and Actions to be taken by the Customer. Advised consignees to immediately stop using the recalled product and return it to LifeHealth. Requested consignees to complete and return the Acknowledgement and Receipt Form. |  | 7 | 7 | Worldwide distribution. US in the state of MI; INDONESIA, UNITED ARAB EMIRATES, LITHUANIA, POLAND, and PERU. | Terminated | October 20, 2015 | 174 | 7 | Not_Computer
Z-1932-2015 | 71302 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ICT Serum Calibrator REF 146-03. |  |  | July 22, 2014 | 2014 | Abbott Laboratories | ICT Serum Calibrator may generate lower than expected Potassium Quality Control (QC) and patient results on the Abbott Architect cSystem. | Process control | Customers who received shipments of the affected lots were sent a Product Correction letter on July 22, 2014. |  | 11215 | 11,215 units | Worldwide Distribution-US (nationwide) including Puerto Rico, and the countries of Afghanistan, Algeria, Andorra, Angola, Argentina, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Bermuda, Bosnia &amp; Herzegovina, Brazil, Brunei, Bulgaria, Burkina Faso, Canada, Cayman Islands, Chile, China, Colombia, Croatia, Curacao, North Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Mali, Mauritania, Mauritius, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tajikistan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turks &amp; Calicos, UAE, Uganda, United Kingdom, Uruguay, Venezuela, Vietnam, and Zimbabwe. | Terminated | July 21, 2015 | 364 | 11215 | Not_Computer
Z-1918-2015 | 71346 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | StoneLight 30 Laser System (NS3000). |  |  | May 13, 2015 | 2015 | New Star Lasers, Inc. | During internal testing a situation was discovered where when changing the settings, a new setting could not be accepted and potentially result in a discrepancy in the energy delivered (higher or lower) from the value displayed on the Graphical User Interface display. | Software design (manufacturing process) | Medical Device Correction letters were sent to all affected consignees by Federal Express on May 13, 2015. The letter identified the affected product as well as the issue and risk involved. The software will be upgraded to address the issue; and, the firm will contact customers to coordinate installation of the upgraded software. Customers may continue to use their system safely by taking the precautions provided in the letter. Customers are to complete the attached response form. Questions should be directed to Nina Davis at 916-677-1909. |  | 52 | 52 units | Worldwide Distribution -- US, including the states of CA, IL, AZ, FL, OH, MA, MI, and AL; and the country of the Netherlands. | Terminated | October 28, 2016 | 534 | 52 | Computer
Z-1934-2015 | 71484 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 6; produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | June 04, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A potential safety issue involving two (2) broken screws and one (1) bowed, which is used to fix the Balance Weight of the DMS in the gantry, was found during a service call at the customer site. | Device Design | A safety advisory notice, dated June 4, 2015, was sent to direct accounts explaining the problem and potential risk to patients and operators. Siemens engineers will begin replacing the bolts in May 2015. |  | 40 | 40 total | Nationwide Distribution. | Terminated | March 24, 2017 | 659 | 40 | Not_Computer
Z-1935-2015 | 71023 | chest drainage kit | PAD | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | Enforcement Discretion | Ocean Water Seal Chest Drain, Single Collection W/AC, |  |  | May 07, 2015 | 2015 | Atrium Medical Corporation | Chest drains have a missing o-ring | Process control | Atrium Medical issued recall letter dated 5/7/15 to inform customers of the steps to take should they encounter affected product. If the O-ring is not present on the patient tubing:a. Do not use the chest drain.b. Replace the entire chest drain including the patient tubing. Open upanother chest drain package and repeat steps 1 to 3. c. Return any units with missing O-rings to Atrium.If the O-ring is present then reconnect in-line connector and proceed with set upas normal. Complete and return Field Action Reply Form to the following e-mail address: connector@atriummed.com or you may fax it to 1-603-386-6590. If you have any questions, please contact your local Atrium/Maquet chest drainage representative or Atrium Customer Service at 1-800-528-7486(603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 390 | 390 cases | NationwideCanada, Germany | Terminated | April 26, 2016 | 355 | 2080 | Not_Computer
Z-1940-2015 | 71383 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Accurus 2500 Standalone Vitreous Probe REF 8065741018; Accurus 23 GA Standalone Vitreous Probe REF 8065750821; Accurus 25+ 2500 CPM Standalone Vitreous Probe REF 8065751122; Anterior Accurus Probe with Infusion Needle REF 8065803650; Single Use Sterile Probes Rx only. |  |  | June 02, 2015 | 2015 | Alcon Research, Ltd. | Insufficient seal on the outside packaging, potentially affecting the sterility of the product. | Packaging | Recall notification letters were sent to consignees on June 2, 2015. |  | 35475 | 35,475 units | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, Germany, Hong Kong, Hungary, India, International APL, Italy, Japan, Kenya, Korea, Kosovo, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Poland, Russia, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | April 20, 2017 | 688 | 35475 | Not_Computer
Z-1942-2015 | 71204 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Data Management System; a clinical data management system that assists medical laboratory professionals with preanalytic and post-analytic functions in conjunction with multiple instruments, the laboratory information system (LIS) and Siemens StreamLAB Automation Solutions. |  |  | April 29, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | System Limitations and Software Issues related to the following features: Custom reference ranges, User-initiated sample and result actions, Auto-verification and quality control rules, and Sample reports. | Software design | Siemens sent an Urgent Field Safety Notice (UMDC # ISW-15-01.A.US) on May 5, 2015 to all affected customers. The letter instructed customers to follow specific work around instructions, review this letter with their Medical Director, complete and return the Field Correction Effectiveness Check attached to this letter within 30 days, retain the letter with laboratory records, and forward it to those who may have received this product.For questions regarding this recall call 800-441-9250. |  | 1355 | 1355 | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Canada, Switzerland, Germany, Denmark, Spain, France, Hong Kong, Italy, Japan, Korea, Malaysia, Netherlands, Norway, New Zealand, Philippines, Portugal, Saudi-Arabia, Slovenia, and Taiwan. | Terminated | May 11, 2017 | 743 | 1355 | Computer
Z-1850-2015 | 70985 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Palindrome SI Chronic Catheter Kit 14.5 Fr x 19cm |  |  | April 23, 2015 | 2015 | Covidien LLC | Catheters may not meet antimicrobial specification | Nonconforming Material/Component | Medtronic (Covidien) issued recalll letter dated April 24, 2015, via FedEx overnight delivery to all affected customers. Customers were requested to quarantine and discontinue use of the affected devices. Customers were instructed to complete A Recalled Product Return Form for product returns. Customer Service will respond with a Return Goods Authorization (RGA) number and shipping documentsCustomers with questions were instructed to contact their Medtronic representative or Medtronic Customer Service, Monday through Friday, 8am  6:30pm ET, at (800)-882-5878.For questions regarding this recall call 203-492-5000. |  | 96994 | 72,363 US - 24,631 Internationally | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Hong, Kong, India, Iran, Ireland, Israel, Italy, Libya, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, Norway, Oman, Panama, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, Thailand, UAE, and United Kingdom. | Terminated | July 26, 2017 | 825 | 96994 | Not_Computer
Z-1941-2015 | 71137 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Extractor Pro RX Retrieval Balloon Catheter 9-12 mm Above Retrieval Balloon Catheter; Universal Product Number (UPN) M00547000 |  |  | April 27, 2015 | 2015 | Boston Scientific Corporation | One lot of the Extractor" Pro RX Retrieval Balloon Catheter (9-12mm balloon) contains an incorrect syringe size (12-15mm), which is larger than the size identified on the label. | Employee error | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated April 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected product in their inventory, quarantine the affected product immediately and return to Boston Scientific in accordance with the enclosed recall instructions. Customers will receive replacements for all recalled product that is returned to Boston Scientific. Customers were instructed to forward the recall notification to their customers. Customers with questions were instructed to contact their local sales representative.For questons regarding this recall call 763-494-1133. |  | 39 | 39 units | Worldwide Distribution including the following countries: France, Great Britain, Spain, Ireland.. | Terminated | January 27, 2016 | 275 | 39 | Not_Computer
Z-1939-2015 | 71455 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Low Profile Wheeled Stretcher, Model No. 21762-110-33H, 21762-110-36H, 21762-110-41H, 21762-135-33HJ, 21762-010-36H, 21762-010-41H, |  |  | June 01, 2015 | 2015 | Sechrist Industries Inc | Sechrist Industries is recalling the Sechrist Low Profile Wheeled Stretcher because it may experience an unexpected incline angle during the height adjustment of the Low Profile Wheeled Stretcher when lifting the patient upward or when lowering the Wheeled Stretcher to a height at which the patient can be loaded/unloaded. | Device Design | Seachrist sent an Urgent: Medical Device Correction letter dated June 1. 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate the Sechrist Low Profile Wheeled Stretcher identified in the attached serial number list (Attachment A) are affected by the field correction and are located at your facility. Within the following weeks, a Sechrist trained Field Service Technician will schedule a visit to your facility to install a Hydraulic Wheeled Stretcher Alignment kit which will greatly minimize the degree of the incline angle that can be experienced after it is installed.Customers are informed that within the following weeks, a Sechrist trained Field Service Technician will schedule a visit to their facility to install a Hydraulic Wheeled Stretcher Alignment kit which will greatly minimize the degree of the incline angle that can be experienced after it is installed. For further questions, please call: (714) 579-8318, or (714) 579-8311.We apologize for any disruption that this correction may cause. Sechrist is dedicated to ensuring the safety and quality of our products. If you have any questions about what to do, please contact us. |  | 81 | 81 US Distribution | US Distribution and Internationally to Japan and Jordan. | Terminated | October 20, 2015 | 141 | 81 | Not_Computer
Z-1952-2015 | 70742 | No | EJI | Cartridge syringe. | Dental | Dental | 510(K) Exempt | Milestone Wand STA Handpiece for the delivery of local anesthesia to dental tissue. Model Number STA-5050-305. The Wand STA handpiece is a single use syringe used in conjunction with the CompuDent STA computer controlled dental injection system. |  |  | February 11, 2015 | 2015 | Milestone Scientific, Inc. | Wand Handpiece was mislabeled as STA handpiece. The Wand handpiece cannot be installed in the STA instrument. | Nonconforming Material/Component | Milestone Scientific notified their customers via e-mail and phone on 2/11/2015. |  | 165850 | 165,850 units | Worldwide distribution. US in FL, IL, NY WA and Puerto Rico, Israel, Croatia, Italy, Australia, Germany, Colombia, UK, Sweden, and UAE. | Open, Classified |  |  | 165850 | Not_Computer
Z-1949-2015 | 71448 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | INSTRUMENT ARM DRAPE, IS4000, 20 PACK; used on the da Vinci Xi Surgical System. |  |  | June 16, 2015 | 2015 | Intuitive Surgical, Inc. | Field Safety Notification to advise of an issue that may occur involving the combination of the da Vinci Xi Drapes and Xi instruments. A pinched drape may cause a message on the Vision Cart touchscreen - Missing reload error. | Device Design | Intuitive Surgical sent an Urgent Medical Device Notification letter on June 16 , 2015 to all affected consignees. Consignees were instructed to take the following actions:1. Ensure that all affected personnel are fully informed of this notice. Forward thisnotice to your Risk Manager, OR Director, Purchasing Manager, BiomedicalEngineering staff and members of your medical staff who perform da Vinci Surgeryprocedures.2. Complete and return the attached Acknowledgment Form to Intuitive Surgicalusing the instructions provided.3. Please retain a copy of this notice with your system.4. Refer to the da Vinci Xi System User Manual for more information on draping thesystem.Consignees with questions or needing spport should contact their Clinical Sales Representative or contact Intuitive Surgical Sustomer Service at the numbers listed below:Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail:customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 AM to6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST). |  | 310 | 310 | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Netherlands, Norway, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | October 05, 2015 | 111 | 310 | Not_Computer
Z-1947-2015 | 71475 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Etest Cefotaxime CT 256 WW and Etest Cefotaxime CT 256 US. in vitro diagnostic. |  |  | June 12, 2015 | 2015 | BioMerieux SA | QC results were out of range resulting in false resistant strains. | Process control | BioMerieux sent an Urgent Product Removal Notices dated June 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Required actions:ÂżÂż Please distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred our product.ÂżÂż Discontinue using and discard any remaining inventory of EtestÂżÂż CEFOTAXIME CT 256 lots reported in table 1.ÂżÂż Discuss any concerns you may have regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action.ÂżÂż Complete and return the Acknowledgement Form in Attachment A by Fax to confirm receipt of this notice and to request product replacement.bioMÂżÂżrieux is committed to providing our customers with the highest quality product possible. We sincerely apologize for any inconvenience that this may have caused you. If you require additional assistance or have any questions, please contact your local bioMÂżÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 674 | 674 | Worldwide Distribution - USA (nationwide) to the states of AL, AR, CA, CO, FL, GA, IL, IN, KY, MD, MI, MN, MO, NE, NY, NC, OH, OK, OR, PA, TN, TX, UT, VA, WA, WV and WI., and to the countries of: Austria, Australia, Canada, Switzerland, Germany, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Japan, Netherlands, Poland, Sweden, Thailand, Turkey, Saudi Arabia, Philippines, Slovenia, Kuwait, Brazil, Chile, Czech Republic, Algeria, South Africa, Kosovo, Costa Rica, Belarus, Guam, Honduras and Vietnam. | Open, Classified |  |  | 674 | Not_Computer
Z-1943-2015 | 71336 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Velocity" Biopsy Valve, 25 units per box, |  |  | February 17, 2015 | 2015 | US Endoscopy Group Inc | The company has received reports of structural non conformance at the connection between the valve body and the irrigation tube, resulting in separation of the components during use. | Under Investigation by firm | US Endoscopy sent an Urgent Recall Notification Letter dated to their customers. The letter identified the affected product, problem and actions to be taken. US Endoscopy requests that their customers:1. Remove all unused affected Velocity Biopsy Valves from your inventory.2. Complete the attached Customer Response Card and return to the attention of Kia Pennington, via fax 1-440-639-4495 or email kpenning@usendoscopy.com.3. Mark the shipping container with code: RMA xxxxxxxx QIF, and return the product to US Endoscopy via Fed Ex Account #293799547.For questions contact Kia Pennington in Customer Service at 1-800-769-8226. |  | 2173 | 2,173 units | Worldwide Distribution - US Nationwide in the states of: CA, CO, IA, MA, OH, TX &amp; WV and the countries of: Brazil, Costa Rica, Croatia, Czech Republic, Egypt, France, Germany, India, Indonesia, Italy, Latvia, Lithuania, Philippines, Poland, Romania, Russian Federation, Slovenia &amp; South Africa. | Terminated | September 21, 2015 | 216 | 2173 | Not_Computer
Z-1945-2015 | 71526 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Gentle Threads Interference Screw 9x35mm Full Thread. |  |  | June 03, 2015 | 2015 | Biomet, Inc. | Part Number: 905620 Lot: 824610 Gentle Threads Interference Screw 10x25mm Round Head is mixed with the Part Number: 905627 Lot: 753410 Gentle Threads Interference Screw 9x35mm Full Thread. If the incorrect screw is used various failures may occur which may ultimately lead to revision surgery or a delay in surgery. | Mixed-up of materials/components | On 6/3/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for locating and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 41 | 41 units | Distribution-Domestic: None. International: Australia, Mexico, Chile, Costa Rica, Netherlands, and Brazil. | Terminated | December 22, 2015 | 202 | 188 | Not_Computer
Z-1948-2015 | 71531 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural Knee II Revision Femoral Stem, 125mm, item 6215-12-185. packaged in configuration 803/791. Knee Prosthesis component. |  |  | October 02, 2014 | 2014 | Zimmer, Inc. | Failed to meet established requirements for sterile barrier integrity due to a breach of the inner cavity. If the outer tray is compromised in addition to the inner tray a periprosthetic infection may occur. The risk associated with a periprosthetic infection is a revision or potential multi-stage revision to treat the infection. | Package design/selection | On 10/2/2014, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, and consignee responsibilities. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759. Hours of operation are Monday through Friday, 8 a.m. through 8 p.m. ET. |  | 352 | 352 | US Nationwide Distribution, including DC and the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MD, MI, MN, MO, MT, NC, ND, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, and WY. | Terminated | July 01, 2015 | 272 | 352 | Not_Computer
Z-1946-2015 | 70992 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Axis-Shield Liquid Stable (LS) 2-Part HOMOCYSTEINE REAGENT; 100 test format, 1000 test format, and 200 test format; and Private Label: (Distributed by BECKMAN COULTER, for professional use only, on the BECKMAN COULTER AU platforms (AU400, AU480, AU680 and AU5800)) |  |  | March 20, 2015 | 2015 | Axis-Shield Diagnostics, Ltd. | Deterioration in performance of the Liquid Stable (LS) 2-Part HOMOCYSTEINE REAGENT resulting in under recovery. | Under Investigation by firm | Axis Shield sent an Urgent Field Safety Notice letter dated March 20, 2015 to their customers via email and written notice sent with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. Recommended Action to Customer is as follows: Confirm receipt of FSN to Axis-Shield. Do not use the affected Reagent lots referenced in the safety notice. In the event of holding one of the affected lots, immediately contact Axis-Shield Product Support Department to arrange for replacement material and return any defective product to Axis-Shield. Complete reconciliation form. Complete acknowledgement of the safety notice within 10 working days. Ensure safety notice has been distributed and reviewed by all appropriate personnel within your organization. If further distribution of affected lots to another organization has been made, inform them of this Product Recall and provide a copy of this safety notice to them. Retain a copy of the safety notice for their records. For any issues or concerns contact Axis-Shield via e-mail; MDL-AXD-Product-Support@alere.com. |  | 127 | 52 kits + 75 kits Beckman Coulter Brand OUS | Worldwide Distribution - US Nationwide in the states of: KY, WA, CA, and PA and the country of Germany. | Terminated | October 27, 2015 | 221 | 127 | Not_Computer
Z-1961-2015 | 71073 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Rapidlab 1260 system is a blood gas analyzer used for laboratory testing of blood gases, electrolytes, metabolites and CO-oximetry in arterial, venous and capillary whole |  |  | May 06, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | D50 and D51 Diagnostic error codes are not functional. | Software design | Siemens sent an Urgent Field Safety Notice April 2015, to all affected customers. Customers were notified of the recall beginning on May 6, 2015 via Federal Express. Foreign customers were notified starting on May 5, 2015 via method determined by each countries local regulations and procedures. The notification informs customers of the issue and asks customers to do the following: Ensure that the Glucose and Lactate sensor contacts are aligned and completely seated with the contacts on the analyzer, Analyze a minimum of 2 levels of quality control material after installing the sensors, as indicated in the RAPIDLab 1200 Operator's Guide, Review this letter with your Medical Director. Complete and return the Field Correction Effectiveness Check Form attached to this letter within 7 days. Retain this letter with your laboratory records, and forward this letter to those who may have received this product. For further questions please call (800) 242-3233. |  | 122 | 122 devices | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Argentina, Australia, Austria, Belgium, Bosnia, Herzog,Brazil, Brunei, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany,Georgia, Greece, Hong Kong, Iceland, India, Indonesia, Ireland, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, P.R. China, Peru, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom, Venezuela and Vietnam. | Terminated | August 02, 2017 | 819 | 2835 | Not_Computer
Z-1953-2015 | 71129 | No | MMS | Patient lubricant. | General Hospital | General Hospital | 510(k) | Dr. Sheffield's LUBRI GEL 3 oz (85 g) |  |  | April 27, 2015 | 2015 | Sheffield Pharmaceuticals, LLC | Recovery of high count of gram positive rods including single colonies of B. cepacia. | Process control | A letter was sent to inform customers of the voluntary recall of Sheffield Pharmaceuticals Lubri-Gel Personal Lubricant Jelly. The letter described the issue and asked customers to quarantine all related inventory immediately and provide their inventory counts on the specific lots. Questions or concerns related to the recall or letter can be directed to James Turner at James.Turner@sheffield-pharmaceuticals.com. |  | 37636 | 37,636 pieces | Worldwide Distribution -- CA, NJ, FL, PR, NJ, MI, NY, NC, FL, IL, GA, MN, KS, and NC; and, the countries of Jamaica and Georgia. | Open, Classified |  |  | 37636 | Not_Computer
Z-1957-2015 | 71362 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CentraLink" Data Management System; software is a network solution provider and multi-system data manager for the instruments and lab automation systems (LAS) within the lab. The CentraLink software consolidates data |  |  | May 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has determined that the sample query function that includes Instrument or Instrument Group as search criteria may not return all samples from the Dimension Vista systems if onboard aliquot support rules are configured in the CentraLink system. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated May 19, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be Taken by the CustomerWhen using any of the CentraLink sample queries, do not select Instrument or Instrument Group criteria for Dimension Vista 500 or Dimension Vista 1500. This would also include any preconfigured queries, commands or tool bar buttons that may use these criteria.Your local Siemens service provider will be contacting you to verify and/or update the configuration of your CentraLink system.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions please call (800) 441-9250. |  | 1941 | 1941 | Worldwide Distribution - Nationwide Distribution and to the countries of : Belgium, Canada, Czech Republic, Denmark, France, Germany, Italy, Netherlands, Saudi Arabia, Spain, Switzerland and United Kingdom. | Open, Classified |  |  | 1941 | Not_Computer
Z-1956-2015 | 71521 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | JustRight 5mm Reload, JR-REL25-2.0 Reload is packaged in cases of 12., placed in a tray and sealed with a Tyvek lid. |  |  | June 11, 2015 | 2015 | Justright Surgical, LLC | JustRight Surgical is recalling various lots of JustRight 5mm Stapler product because they may deliver incomplete staple formation and incomplete tissue closure. | Process design | JustRight Surgical sent an Urgent Medical Device Recall letter dated June 12, 2015 to aaffected customers and distributor via an email. The letter identified the affected product, problem and actions to be taken. Custers were instructed Review, complete, and return the enclosed Acknowledgement Form directly to ÂżRepÂż, your Sales Representative or email to orders@justrightsurgical.com or fax to (720) 287-7135.The following assistance is available:" For questions related to the issue or product replacement, please contact your local sales representative, ÂżRepÂż, or contact JustRight Surgical Customer Service at (866) 683-1743, Monday through Friday from 7:30 AM to 5:00 PM MST." For questions related to product pick-up by JustRight Surgical, please also contact your local sales representative or Customer Service (866) 683-1743. Return instructions are provided on the Acknowledgement Form." For questions related to the Acknowledgement Form and its return, please contact the JustRight Surgical Director of Quality Systems at (720) 287-7146. |  | 1452 | 1452 | Distribution - US Nationwide and Ireland. | Terminated | August 26, 2015 | 76 | 2274 | Not_Computer
Z-1965-2015 | 71434 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hamilton Medical Infant Flow Sensor, single use, (3.1m), Part number 260179. The infant flow sensor is an accessory for the Hamilton Medical Ventilators. |  |  | June 01, 2015 | 2015 | Hamilton Medical, Inc. | An issue has been discovered with the Hamilton Medical Infant Flow Sensor, single use. When a ventilation mode with adaptive volume control (APVcmv/CMV+) is used, the ventilator may adjust the ventilation on the inaccurate flow measurement and decrease the amount of ventilation delivered. | Device Design | Hamilton Medical Inc sent a Medical Device Safety Notices letter dated May 8, 2015 by traceable method to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Technical Support Manager at 800-426-6331, Ext 215. |  | 154 | 154 total, all part numbers | US Nationwide Distribution. | Terminated | August 12, 2015 | 72 | 154 | Not_Computer
Z-1972-2015 | 71507 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | SYNGO IMAGING (VERSION V30 and V31); Syngo Imaging is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. |  |  | June 11, 2015 | 2015 | Siemens Medical Solutions USA, Inc | During certain clinical workflows safety issues may occur involving monitor assignment order and incorrect display and measurements for images with non square pixel size. | Software design | Siemens sent a Customer Advisory Notice, dated June 11, 2015, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were also asked to check the software version and consider whether the issue impacts clinical routine and if a software update is necessary. |  | 5 | 5 | US in the states of NC and NE | Terminated | March 24, 2017 | 652 | 5 | Computer
Z-1959-2015 | 71469 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems between software versions VA16A and VA30A and with the Stress Echo Option. |  |  | May 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | When performing a 2D Stress Echo study, some following keystrokes result in a potential loss of data, where some selected clips may not be saved as part of the study. | Software design | Siemens sent an Customer Safety Advisory Notification on May 29,2015, to all affected customers. The letter identified the product the problems and the action needed to be taken by the customer.Customer Safety Advisory notifications inform users of the potential issue with the ACUSON SC2000 systems with software versions between VA16! and VA30A and the Stress Echo feature.Letters describe the steps that result in the potential loss of data and suggest that the following steps be taken to remedy the situation:1. Select Stress Echo Selection2. If applicable relabel the clips for all Stages and Views3. Do not return to Stress Echo Selection to select additional relabeled clips.The issue will be resolved in an upcoming software release for customers who did not receive the original software update. Until the necessary repairs have been completed, all personnel in the Organization who need to be aware of the issue should be notified. For further questions please call (650) 694-5398. |  | 573 | 573 devices | Worldwide Distribution - US Distribution and to the countries of : Albania, Angola, Armenia, Australia, Austria, Bangladesh, Belarus, Bolivia, Bosnia Herzog, Brazil, Brunei, Bulgaria, Canada, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Lithuania, Malaysia, Maldives, Mauritius, Mexico, Netherlands, Nigeria, Oman, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed, San Marino, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom and Venezuela | Open, Classified |  |  | 573 | Computer
Z-1958-2015 | 71270 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Cios Alpha; The Cios Alpha is a mobile X-ray system designed to provide X-ray imaging of the anatomical structures of patient during clinical applications. Clinical applications may include but are not limited to: interventional fluoroscopic, gastro-intestinal, endoscopic, urologic, pain management, orthopedic, neurologic, vascular, cardiac, critical care and emergency room procedures. The patient population may include pediatric patients |  |  | May 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | under certain circumstances the Cios Alpha system may freeze during a procedure. | Software design | Siemens sent an Safety Advisory Notice dated May 5 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were alerting them about the issue and risks involved. Customers were provided work-around instructions until a software fix is released. For further questions please call (610) 219-6300. |  | 37 | 37 | Nationwide Distribution | Terminated | January 13, 2017 | 619 | 37 | Computer
Z-1967-2015 | 71352 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite F500, in vitro diagnostic. |  |  | May 29, 2015 | 2015 | Tecan US, Inc. | A firmware bug causes incorrect values to be displayed for specific measured luminescence values. | Device Design | Tecan sent an Urgent Field Corrective Action letter dated May 29, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached Customer Response Form. For questions contact your local Tecan Helpdesk. |  | 868 | 868 in total | US Nationwide Distribution in the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, including DC | Terminated | August 18, 2017 | 812 | 868 | Not_Computer
Z-1954-2015 | 71505 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Treatment Planning and Delivery System Software version 3.6. ViewRay. |  |  | May 07, 2015 | 2015 | Viewray Incorporated | The software was failing to determine new patient locations if imaging is not enabled during treatment. | Software design | On May 7, 2015 the firm sent Customer Advisory Notification Letters to their customers. |  | 3 | 3 | Distributed in the states of CA, MO &amp; WI. | Terminated | February 04, 2016 | 273 | 3 | Computer
Z-1960-2015 | 71497 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Spine Wave Access Kit, Catalog Number 10-1530 |  |  | May 04, 2015 | 2015 | Spine Wave, Inc. | Mislabeled outer carton:Monopolar Probe is listed as 200 x 1.5mm when it should be listed as 280 x 1.5mm | Packaging change control | Spine Wave sent an Urgent: Medical Device Recall letter dated June 11, 2015 to affected customers. The letter identified the affect product, problem and actions to be taken. Customers were instructed to acknowledge receiptof notification by returning the enclosed Recall Notice Acknowledgement Form in the stamped envelope provided.For questions, contact Spine Wave directly at 203-944-9494and ask for the Recall Coordinator or by email at rsmith@spinewave.com. |  | 121 | 121 units | US Nationwide Distribution in the state of CA, FL, GA, LA, MA, MD, MI, MO, NC, NH, OH, OR, PA, TX, and VA. | Open, Classified |  |  | 121 | Not_Computer
Z-1966-2015 | 71572 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D and Advantx Legacy/Legacy-D Radiographic and Fluoroscopic Systems. |  |  | June 12, 2015 | 2015 | GE Healthcare | The SFD/IDD safety mechanism may not engage properly at the lower range of SFD/IDD travel. This mechanism is installed to hold and prevent the SFD/IDD from falling due to a counterweight cable failure. The lower range of travel is 1.1 in (28mm) from the lowest point of normal SFD/IDD vertical compression. A fall of the SFD/IDD could result in an injury to a patient or operator. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#10902 dated June 12, 2015 to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 4053 | 4,053 (3,804 US; 249 OUS) | Worldwide Distribution - US Nationwide including PR, GU and DC and the countries of: Afghanistan, Argentina, Australia, Austria, Belgium, Bermuda, Canada, Cayman Islands, Chile, Colombia, Czech Republic, Dominican Republic, Ecuador, Egypt, France, Germany, Ghana, Guatemala, Honduras, Indonesia, Ireland, Israel, Italy, Jamaica, Jordan, Korea (Republic of), Kuwait, Libyan Arab Jamahiriya, Malaysia, Malta, Mexico, New Zealand, Panama, Romania, Russia, Saudi Arabia, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | December 21, 2015 | 192 | 4053 | Not_Computer
Z-1973-2015 | 71508 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima IGS 320. |  |  | May 29, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue that may result in the loss of C-arm gantry motion capability on fluoroscopic Interventional imaging systems (Innova, Optima and Discovery). | Nonconforming Material/Component | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter on 5/29/2015- GEHC Ref#12231- dated May 29, 2015 to their Consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and contact Information. The customers were instructed that you may continue to use the system and avoid rapid direction changes with the gantry motions and prevent potential collisions with other equipment. If the failure occurs during interventional examination, please ensure that you have established procedures for handling patients in case of the loss of gantry motion as stated in your product labeling.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 3 | 3 | Worldwide Distribution: US (nationwide) including states of: CA, DE, FL, GA, IL, IN, KS, MS, NV, OH, OK, OR, PA, TN, TX, VA, WA and WI; and countries of: ARGENTINA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, INDONESIA, ITALY, JAPAN, RUSSIA, UNITED KINGDOM, and VIETNAM. | Terminated | October 28, 2015 | 152 | 83 | Not_Computer
Z-1986-2015 | 71443 | No | ISD | Measuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Primus (THE BTE WS30), model PrimusRS (PRRS) and Primus+ (PR30) a measuring exercise equipment device designed to provide active and passive torque resistance to patient motion during rehabilitation therapy. The Primus that includes the optional Chop / Lift Bar. |  |  | April 09, 2015 | 2015 | BTE Technologies, Inc. | On the Primus that includes the optional Chop / Lift Bar, there is a remote possibility that the Chop / Lift Bar could get accidently detached from the snap hook in specific use scenarios when the bar is rotated and twisted around the snap hook. The snap hook could open, resulting in possible injury to the user due to a fall. | Device Design | BTE sent an Medical Device Correction letter dated April 10, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.In order to correct this issue, please immediately install the provided retrofit kit according to BTE instructions 751W2519 included in the kit. From this point on, the ring on the Chop / Lift Bar shall not be used to connect the Chop / Lift Bar to the snap hook on the 191 Tool or any other device with a snap hook. Only the Eye part of the Jaw and Eye swivel will be used to connect to the snap hook.Please pass this Medical Device Recall notice to all those who need to be aware of it within the customer organization. Once the correction is completed, notify BTE using the provided Field Correction Confirmation card.We apologize for any inconvenience. Please do not hesitate to call BTE Customer Support at 410-850-0333 if you have any questions or concerns. |  | 510 | 510 Devices | Worldwide Distribution - Nationwide Distribution and to the countries of : Austria, Belgium, Canada, France, Germany, Greece, Hong Kong, Italy, India, Japan, South Korea, Mexico, P.R. China, Poland, Republic of South Africa, Romania, Saudi Arabia, Singapore, Spain, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 510 | Not_Computer
Z-1987-2015 | 71208 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips GoSafe Communicator 7000L |  |  | April 22, 2015 | 2015 | Lifeline Systems, Incorporated | Home communicator Model 7000L may not sense the phone line under certain unusual landline phone conditions. | Software design | Philips Lifeline notified all consignees of the recall via phone starting on April 22, 2015. The call informs consignees of the issue and actions to be taken. For questions, consignees can contact customer service at 1-800-451-0525 ext. 1608. Consignees were also sent a letter to all consignees instructing them that Lifeline would send out a qualified Installiation Representative for an equipment exchange. |  | 3189 | 3189 devices total, 38 are subject to recall | Nationwide Distribution including AK AZ CA CT FL IA IL KY MA MD MT NC NE NH NY OH OR SC VA WA and WI. | Completed |  |  | 3189 | Not_Computer
Z-1988-2015 | 71428 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | Natrelle Re-sterilizable Breast Implant Sizer, Size: 410 cc, Style No. MSZFX410. |  |  | June 08, 2015 | 2015 | Allergan Inc | Allergan is recalling the NATRELLE 410cc Re-Sterilizable Breast Implant Sizer because the expired product was shipped to the user level. | Error in labeling | A recall letter dated 6/8/15 was sent to all customers who received the NATRELLE 410 cc Re-Sterilizable Breast Implant Sizer. The letter provides the problem identified and the action to be taken. Customers are instructed to contact GENCO Pharmaceutical Services if the recall actions on the recall letter are unclear. For product returns, customers are instructed to contact GENCO at (877) 319-8966, 7am to 5pm CST. Customers with credit/reimbursements are instructed to contact Allergan at (800) 811-4148, 7am to 5pm PST. Customers with medical inquiries are instructed to contact (800) 433-8871, option 2, 8am-5pm, PST. Customers with adverse events/product complaints are instructed to call (800) 624-4261, option 3, 8am-5pm CST.On June 12, 2015, Actavis sent follow up letters to customers to include one additional serial number. |  | 9 | 9 units | CA, IN, MO, NJ, VA | Terminated | December 16, 2015 | 191 | 9 | Not_Computer
Z-1990-2015 | 71351 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The PowerLED / X-TEN / VOLISTA Surgical Light Systems are surgical light systems intended to be used to provide visible illumination of the surgical area or the patient. Ceiling surgical light systems typically include one or more arms that are able to rotate 360 degrees around a central axis. |  |  | May 29, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Maquet Suspension Arms used with PowerLED/X-TEN/Volista Surgical Light Systems may grind at the junction of the suspension arms which can cause paint chips or particles to fall into the surgical field, when the suspension arms are mounted above the surgical table. | Device Design | Maquet Medical Systems sent an Urgent Medical Device Field Correction letter dated May 29, 2015, to all affected customers via Fedex. The letter identified the product the problem and the action needed to be taken by the customer.The customer was informed that the correction will include inspection of PowerLED / X-TEN / VOLISTA Surgical Light Systems mounted above the surgical table and replacement of the suspension arms if found to be defective. As a remedial measure, pending replacement of the defective suspensions, Maquet personnel will install a protective plastic strip at each junction where the arms are grinding. The correction will be performed by an authorized Maquet Service Technician. The customer will be contacted by a representative of the Maquet Service team to schedule an on-site inspection and possible service of the PowerLED / X-TEN / VOLISTA Surgical Light Systems. If the customer has any questions or require additional information, they can contact their local Maquet representative, or Maquet Customer Service at 1-888-627-8383 (option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 646 | 646 units | Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Austria, Australia, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Croatia, Hungary, Indonesia, Ireland, Iraq, Jordan, Japan, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Greenland, Greece, Hong Kong, United Arab Emirate, United Kingdom, Kenya, Kuwait, Sri Lanka, Luxembourg, Libya, Mongolia, Martinique, Mexico, Malaysia, Namibia, Nigeria, Netherlands, Norway, Oman, Panama, Peru, Philippines, Pakistan, Poland, Saint Pierre and Miquelon, Portugal, Reunion, Romania, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Senegal, Thailand, Tunisia, Turkey, Trinidad and Tobago, Taiwan, Uruguay, Venezuela, Serbia, Yemen, Mayotte and South Africa. | Open, Classified |  |  | 646 | Not_Computer
Z-1989-2015 | 71457 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Nakoma-SL Anterior Cervical Plating System's Variable Single Barrel Drill Guide |  |  | June 09, 2015 | 2015 | Alliance Partners LLC | Inadequate interference between the screw and the plate when screw implanted into the plate at angulation greater than 7 degrees can result in the screw advancing through the plate during implantation. | Device Design | Alliance Spine sent an Urgent Medical Device Field Action Notification letter dated June 15, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the response form and fax to 210-314-2524. |  | 21 | 21 units | US Nationwide Distribution in the states LA, TX, and PR. | Terminated | October 09, 2015 | 122 | 21 | Not_Computer
Z-1996-2015 | 71130 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Endovenous Kit, Endo Pack and Endovenous Laser Kit - contains Devon Light Glove |  |  | April 20, 2015 | 2015 | Stradis Healthcare, LLC. | Convenience kits being recalled due to the recall by the manufacturer (Medtronic) of a component (Devon Light Glove) contained in the kits. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated April 16, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine and discontinue use of the affected device. Complete the Recalled Product Return Form and fax to 800-895-6140 or email to feedback.customerservice@Covidien.com. Customers with questions regarding return of the product should call 1-727-642--4124. |  | 15 | 15 | Distribution US nationwide to AL, AK, CA, DE, FL, GA, HI, ID, IL, IN, KS, KY, ME, MD, MA, MI, MS, MO, NM, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA; Canada and Dubai | Open, Classified |  |  | 204 | Not_Computer
Z-2043-2015 | 71404 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5x/CORE Mobile Intravascular Ultrasound Imaging and Pressure System; Product/Part numbers: s5 - 804200-002 and 804200-014; |  |  | September 23, 2013 | 2013 | Volcano Corp | A limitation in interoperability between the Philips Xper Flex Cardia Physiomonitoring System and Volcano's sS/Si, Core, and ComboMap systems. This limitation may impact the calculation of Fraction Flow Reserve (FFR) within the Volcano System when the system is connected to the Xper Flex Cardia system via its aortic pressure auxiliary output. | Nonconforming Material/Component | Volcano and Philips sent an Advisory Notice letter dated September 23, 2013 to all impacted customers. The letter identified the affected product, problem and actions to be taken. For question contact your Volcano and Philips representative. |  | 227 | 227 total units affected by the issue - all models | Worldwide Distribution - US Nationwide and the countries of Belgium, United Kingdom, Denmark, Netherlands, Sweden, Iceland, France, Saudi Arabia, Australia, Philippines and Thailand. | Terminated | July 09, 2015 | 654 | 227 | Not_Computer
Z-2047-2015 | 71420 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PowerPICC SOLO Catheter with Sherlock 3CG TPS Stylet 4F Maximal Barrier Tray with Microintroducer |  |  | May 27, 2015 | 2015 | Bard Access Systems | Bard Access Systems is recalling PowerPICC Solo Catheter Kits due to incorrect expiration dates on the primary labeling. | Labeling Change Control | BARD sent an Urgent Notification letter dated May 27, 2015, to all affected consignees via FedEx. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed that they may choose to return the affected product to Bard Access Systems or use with the understanding that the correct expiration date is November 2015 (2015-11). Customers with questions were instructed to call 1-800-290-1689. For questions regarding this recall call 801-522-5974. |  |  |  | Nationwide Distribution | Terminated | February 10, 2016 | 259 |  | Not_Computer
Z-2052-2015 | 71456 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Block Versions Used by Dental Offices often referred to as chairside milling: |  |  | June 15, 2015 | 2015 | 3M Company - Health Care Business | 3M is initiating a Field Correction to remove the crown indication from all Lava Ultimate products. These products will continue to be indicated for onlays, inlays and veneers. | Device Design | 3M ESPE Dental sent an URGENT: MEDICAL DEVICE CORRECTION letter dated June12, 2015 to affected customers. The letter described the affected product, problem and action to be taken. The letter instructed customers to complete and return the Acknowledgement Return/ Refund Form. For questions call 3M ESPE Dental Customer Care center at 1-800-634-2249. |  | 1800000 | 1.8 million (US 1,017,834) | Worldwide Distribution - US: Nationwide and the countries of: Australia, India, Japan, Korea, New Zealand, Singapore, Thailand, Vietnam, Canada, Bulgaria, Czech Republic, Estonia, Hungary, Israel, Latvia, Lithuania, Poland, Romania, Slovakia, Switzerland, Turkey, China Hong Kong, Taiwan, Brazil, Colombia, Costa Rica, Guatemala, Mexico, Egypt, Morocco, South Africa, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Italy, Netherlands Norway, Portugal, Spain, Sweden, UK and Ireland. | Terminated | April 26, 2016 | 316 | 1800000 | Not_Computer
Z-2051-2015 | 71440 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TITANIUM END CAP WITH T40 STARDRIVE RECESS, 0MM EXTENSION FOR TITANIUM TIBIAL NAILS-EX; Synthes Tibial Nail System EX is intended to stabilize fractures of the proximal and distal tibia and the tibial shaft; open and closed tibial shaft fractures; certain pre-and post-isthmic fractures; and tibial malunions and non-unions. |  |  | June 03, 2015 | 2015 | Synthes, Inc. | A specific part and lot number was packaged and shipped prior to the completion of a required internal inspection. | Employee error | The firm, Depuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" notification letter, dated June 3, 2015, to direct accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove the affected lots from your stock; If you DO have any of the identified devices, please take the following steps:ÂżÂż Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA)Complete and Return the Verification Section with the product to:ÂżÂż Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.ÂżÂż Send a copy of the completed Verification Section even if you do not have any product to DePuy Synthes by:ÂżÂż Fax: (610) 430-7083 or ÂżÂż Scan/email: Fieldaction@synthes.com Number. If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 3 | 3 | US Distribution to states of: MO and SC. | Terminated | September 15, 2015 | 104 | 3 | Not_Computer
Z-2046-2015 | 71511 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S. Foot Locking Calcaneus Plate,Small-Right. |  |  | June 04, 2015 | 2015 | Biomet, Inc. | The thread location on one of the thread holes is offset from the axis of the predrill hole. The threads are too deep on one side and too shallow on the other. The locking screw may back out of the plate if it does not achieve proper purchase. This may require a revision surgery to replace or remove the screw and/or plate. The threads, not uniformly loaded, may shear off, causing metal slivers. | Process change control | On June 4, 2015 an URGENT MEDICAL DEVICE REMOVAL letter was sent to all consignees. This action requires the immediate location and discontinued use of the product and its return to Biomet.The following actions are REQUIRED: Immediately locate and remove the identified device(s) listed below from circulation; Carefully follow the instructions on the enclosed "Response Form"; Email a copy of the Response Form to audrey.daenzer@biomet.com prior to return of product; Use priority carrier for your shipment; and If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear Risk/Recall Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. Please confirm receipt of this notice by sending back the response form within three (3) business days. Thank you in advance for your assistance and prompt attention. On behalf of Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 75 | 75 units | Distributed in the states of TX, MD, PA, NJ, NC, MA, AZ, GA, FL, and AR. and the countries of Costa Rica CP and The Netherlands. | Terminated | December 22, 2015 | 201 | 75 | Not_Computer
Z-1950-2015 | 71197 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube, Cuffless. Model Numbers: 2.5NEF,3.0NEF,3.5NEF,4.0NEF,4.5NEF,2.5PEF,3.0PEF,3.5PEF,4.0PEF,4.5PEF,5.0PEF,5.5PEF,5.0PELF,5.5PELF,6.0PELF,6.5PELF,2.5NCF,3.0NCF,3.5NCF,4.0NCF,4.5NCF,2.5PCF,3.0PCF,3.5PCF,4.0PCF,4.5PCF,5.0PCF,5.5PCF,5.0PLCF,5.5PLCF,6.0PLCF,6.5PLCF,3.0NEF-P,3.5NEF-P,4.0NEF-P,4.5NEF-P,3.0PEF-P,3.5PEF-P,4.0PEF-P,4.5PEF-P,5.0PEF-P,5.5PEF-P, |  |  | May 08, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports where patients who recently switched from the current Shiley Neonatal and Pediatric products to the affected products experienced discomfort and in some cases, breathing difficulties, corresponding to a negative effect on oxygen levels. | Device Design | Consignees were notified 5/8/15 via letter. Consignees were asked to return products. |  | 69461 | 69,461 units total | Worldwide Distribution. US nationwide, Canada, Australia, Belgium, Chile, Germany, Israel, Italy, Japan, Saudi Arabia, Singapore, South Africa, Turkey, and Uruguay. | Terminated | April 12, 2017 | 705 | 69461 | Not_Computer
Z-2056-2015 | 71147 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | IntegraÂż Long Epidural Steroid Injection Needle (Tuohy Needle), 20 gauge x 5 inch. Plastic Hub, Metal Stylet, calibrations (10BX). IntegraÂż Long Epidural Steroid Injection Needle (Catalogue #PITRW2050). Peel Pouch. The IntegraÂż Long Epidural Steroid Injection Needle is packaged in a labeled and sterilized sealed tyvek / plastic pouch. The sterilized pouches are subsequently packaged in a case of 10 that is labeled, sealed and distributed to customers. |  |  | May 01, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences Corp. d.b.a. Integra Pain Management Announces Voluntary Recall of Tuohy Needles (20 Gauge, 5 inch) due to incorrect packaging (labeled as 18 Gauge, 3. 5 inch). | Labeling design | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 1, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Integra is asking customers to do the following:1. Review your inventory and determine if you have Tuohy Needle, 20 Gauge X 5 inches (Cat. No. PITRW2050) with an affected Lot Number that is listed on page 2. If so, stop using them immediately and remove them from service.2. Complete the attached form. If you do not have affected product, check the box: I do not have any product(s) with a Lot Number from the list below.3. If you do have affected product, check the box:  I do have product(s) with a Lot Number from the list below.4. Complete the other information on the form and return it by email or fax as indicated on the form. Keep a copy of the form for your records.When your form is received, Customer Service will contact you and provide an RMA number and directions to return and replace the affected products. Should you have any questions regarding these instructions, please contact Customer Service at 1-800-241-2210. For questions regarding this recall call 609-936-6822. |  | 123 | 123 Tuohy needles | Nationwide Distribution | Terminated | December 04, 2015 | 217 | 123 | Not_Computer
Z-2055-2015 | 71048 | lens, fundus, hruby, diagnostic | HJI | Diagnostic Hruby fundus lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lombart 90D Clear with case |  |  | February 27, 2015 | 2015 | Volk Optical Inc | Mislabeling. The incorrect lot number was engraved on the anodized aluminum ring. The correct lot number is BE03356; however, BD03356 was engraved on the ring in error. | Other | Volk Optical sent a Medical Device Recall letter on March 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the affected product. If product was further distributed customers need to identify their customers and notify them at once. Customers were instructed to return the affected device for replacement. Customers with questions were instructed to contact Volk at 440-942-6161 or toll free at 800-345-8655. For questions regarding this recall call 440-510-0745. |  | 78 | 78 units | Distribution to VA only | Terminated | February 04, 2016 | 342 | 78 | Not_Computer
Z-2053-2015 | 71474 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Songer Spinal Cable System, Single Cable, Crimp, Bar, Stainless Steel, Sterile. |  |  | May 29, 2015 | 2015 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | The box end label is incorrectly labeled as Titanium instead of Stainless Steel on one lot of implantable Songer Spinal Cable System, Single Cable, Crimp, Bar, Stainless Steel, Sterile, distributed by DePuy Spine Inc. | Labeling Change Control | On 5/29/15, RTI Surgical sent DePuy/Synthes Spine Inc. with an URGENT: Medical Device Voluntary Recall Notification identifying the recalled product, reason for recall, potential risks, and immediate actions required. |  | 25 | 25 | Distributed in MA. | Terminated | March 17, 2016 | 293 | 25 | Not_Computer
Z-2054-2015 | 71315 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Dimension Vista Systems, used with the Dimension Vista Intelligent Lab System software version 3.6.1. In vitro diagnostic analyzer. |  |  | May 21, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Two software issues may occur in Vista software version 3.6.1. Issue #1: Samples may stop processing without notification. Issue #2: Timing issue in Vista 1500 may cause wrong reagent or no reagent delivery. This may lead to unflagged, unexpected low results and results flagged with assay errors. | Software design (manufacturing process) | Siemens sent an Urgent Medical Device Correction letter dated May 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be Taken by the Customer:For the corresponding issues listed above, please perform the following:Issue 1: If you experience this situation, sample processing can be resumed using the step(s) below:Step 1: From the Operation Icon select PAUSE, then immediately return to the Operation Icon and select Cancel PAUSE. Sample aliquotting should immediately resume.Step 2: If sample processing does not resume after Step 1, wait until the instrument displays System Ready, then restart the system software per the Vista Operator Guide,Section 2, Controlled Shutdown-Restart procedure.Issue 2:To prevent the Reagent Servers from temporarily losing synchronization, Siemens recommends setting the Auto Dispose option to No for methods assigned to Reagent Server 2 (see instructions below). These empty or expired reagent cartridges will now remain on the instrument and require the operator to unload them.Please consult your Vista Operators Guide or method Instructions for Use when results with flags and/or errors occur.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center-Technical Solutions at (800) 441-9250 or your local Siemens technical support representative. |  | 2315 | 2315 | Worldwide Distribution - Nationwide Distribution including Puerto Rico., and the countries of : Austria, Australia, Belgium, Canada, Denmark, France, Finland, Germany, Japan, Italy, New Zealand, Netherlands, Czech Republic, Portugal, Republic Korea, Norway, Spain, Switzerland, Slovakia, Saudi Arabia, Slovenia, and United Kingdom, | Open, Classified |  |  | 2315 | Computer
Z-2057-2015 | 71234 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilators with software versions V2.40/2.41 |  |  | March 24, 2015 | 2015 | Hamilton Medical, Inc. | Customer reports that the ventilator display can freeze. Ventilation continues but the information is no longer displayed on the screen and the user can no longer operate the device. | Device Design | *** Please note *** This is not a new recall. An error occurred, and this recall was originally classified as a class I on July 10, 2015. No further action is required.Hamilton Medical AG notified their customers with a letter dated February 19, 2015 and had a fix (SW 2.42) that was available as of February 20th 2015. Hamilton Medical proceeded to send out notification letters the 24th of March 2015 once they compiled the complete and validated list of customers that had the affected software. Customers with questions were instructed to call 800-426-6331, ext.215. |  | 248 | 248 | Nationwide Distribution | Terminated | December 23, 2015 | 274 | 248 | Computer
Z-2059-2015 | 71528 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Altiva / Exactech Spine, Large Set Screw, Catalog Number 28202, Lot Number AS05107. |  |  | March 06, 2009 | 2009 | Exactech, Inc. | Incorrect thread form on the screw. | Device Design | Consignees were notified with a Product Recall Notice on 03/06/2009 that all screws must be checked to ensure that any discrepant products to be removed from the field. |  | 749 | 749 devices | Distributed in the states of FL, GA, IL, NY, TX, and WA. | Terminated | July 14, 2015 | 2321 | 749 | Not_Computer
Z-2058-2015 | 71423 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S., 3.5mm x 60 mm Cortical Locking Screw. Bone fixation screw |  |  | June 04, 2015 | 2015 | Biomet, Inc. | Sterile packaging was not sealed. The implantation of an unsterile screw may lead to infection. | Packaging | On 6/4/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for removing the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday 8 a.m.to 5 p.m. |  | 50 | 50 units | Domestic: None.International:Netherlands and Japan | Terminated | March 22, 2016 | 292 | 50 | Not_Computer
Z-0811-2016 | 71008 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope Reusable Blades (GVL 3, 4, 5 and AVL 2, 3, 4, 5). |  |  | April 13, 2015 | 2015 | Verathon, Inc. | Verathon has revised the Operations and Maintenance Manual (OMM) to specifically state compatibility limitations of the reusable video laryngoscope Blade validated through simulateduse cycles, such that exceeding this number of cycles in the specific chemical agent may affect the potential life of the product. | Nonconforming Material/Component | Verathon sent an Urgent Medical Device Safety Alert letter dated April 13, 2015 and Safety Alert Reply Form to their consignees on April 17, 2015Users are informed that Verathon implemented a Safety Alert regarding updates to the following Operations and Maintenance Manuals (OMM):" GVL Operations and Maintenance Manual: 0900-1204-08-60" AVL Operations and Maintenance Manual: 0900-4200-02-60For a copy of the most recent GVL or AVL Operations and Maintenance Manual, please visit http://verathon.com/support.If you have any questions regarding the Operations and Maintenance Manuals revisions noted above, please contact your Verathon representative or Verathon Customer Care at 800.331.2313 or 425.867.1348. You may also email customerservice@verathon.com.PLEASE NOTE: NO ADDITIONAL ACTIONS NEED TO BE TAKEN. |  | 9088 | total 9088 units ( 5786 units in the US and 3302 units International) | Worldwide Distribution. | Completed |  |  | 9088 | Not_Computer
Z-2061-2015 | 71606 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Agility. |  |  | June 26, 2015 | 2015 | AGFA Healthcare Corp. | Firm received a complaint that the hospital had sent two accession numbers over the same morning and on the Agility side the patient name was incorrect. | Other | PRB0046372 was created to address the corrective actions related to this issue. On 6/26/2015, an email communication, as an effectiveness check was sent to the two customers. Acknowledgment, via email, that the communication was received was requested from the consignees. |  | 2 | 2 | Distributed in the state of MO. | Terminated | January 27, 2016 | 215 | 2 | Not_Computer
Z-2062-2015 | 71271 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | 0210-312-000-Woundcare Replacement kit with soft tip; 0210-318-000; Woundcare Replacement with fan tip; 0210-318-100- Woundcare Replacement Kit; 0210-318-200-Woundcare Replacement with retract coax tip. Accessory to Stryker InterPulse Irrigation System, used for wound debridement, soft tissue debridement, and cleansing of medical, clinical, or surgical sites. |  |  | May 08, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Potential sterility breach in the packaging. | Package design/selection | Stryker sent Customer Notification Letters on 05/20/2015 through Certified mail via USPS.Customers are asked to do the following: Immediately review this Recall Notification. Immediately check all stock areas and/or operating room storage and quarantine any affected product found. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you do not have any affected product on hand. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF. Fax (866-521-2762) or email (julie.forsyth@stryker.com) the completed Business Reply Form to Stryker Instruments Regulatory Department.Upon receipt of the Business Reply Form, if you have recalled product on hand, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled WoundCare Replacement kits to Stryker. |  | 5328 | 5,328 each (444 boxes) | Nationwide Distribution-including the states of FL, IA, IL, IN, LA, MD, MN, MO, NC, NM, NY, PA, TN, TX, VA, and WI. | Terminated | October 19, 2015 | 164 | 5328 | Not_Computer
Z-2060-2015 | 71530 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | ZNN Antegrade Femoral Nail (ZNN AF). Orthopedic internal fixation device. |  |  | June 17, 2015 | 2015 | Zimmer, Inc. | A single distributed ZNN Greater Trochanter Femoral Nail may have been dented. Potential for fatigue of implant prior to sufficient fracture healing, which may result in need for implant removal and revision of fracture fixation to assure proper healing and avoid fracture malunion or nonunion. | Process control | On 6/17/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 8 | 8 | Worldwide distribution. US states of NC and AK; Taiwan, Japan, and Germany. | Terminated | March 15, 2016 | 272 | 8 | Not_Computer
Z-2065-2015 | 71529 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Oxford Femoral Drill Guide Lateral IM Adapter |  |  | June 03, 2015 | 2015 | Biomet, Inc. | Oxford Femoral Drill Guide Lateral IM Adapter may have an undersized diameter that can cause an interference with PN: 32-422822, Oxford Microplasty IM Link. | Component design/selection | On 6/3/2015 an "URGENT MEDICAL DEVICE RECALL NOTICE" letter was sent to the affected distributors with instructions for locating and returning the affected product. The recall letter included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 18 | 18 | Worldwide Distribution - US Nationwide in the states of NC, IN, PA, OH and the country of the Netherlands. | Terminated | March 15, 2016 | 286 | 18 | Not_Computer
Z-2066-2015 | 71493 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G7 Variant Elution Buffer HiS no. 1 (M) component of the Tosoh Automated Glycohemoglobin Analyzer HLC 723G7. |  |  | February 10, 2015 | 2015 | Tosoh Bioscience, Inc. | The G7 Variant Elution Buffer HiS no. 1 (M) is incorrectly labeled with an expiration date of October 2015. The correct expiration date is October 2014. | Error in labeling | Tosoh Bioscience sent an Urgent Recall Notification letters dated February 10, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard any remaining pouches of the product, complete and return the attached response form. For questions call Tosoh Technical Support at 1-800-248-6764. |  | 45 | 9 boxes/5 pouches | Us Nationwide Distribution in the state ID including PR. | Terminated | October 01, 2015 | 233 | 45 | Not_Computer
Z-2064-2015 | 71407 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex ARC Hip Stem, HA Coated, Size 5; Product Code: HC-21005-HA |  |  | June 03, 2015 | 2015 | Omnilife Science Inc. | Product may breach the inner and outer sterile pouches during shipping or handling. | Packaging | All product is consigned to US Agents until implanted or returned to OMNI. All US Agents holding the potentially affected inventory were notified that the product is being recalled via letter on June 3, 2015. The letter instructs them to immediately discontinue use of the product and contact OMNI Customer Service at 800-448-6664 to request a return materials authorization, and to return the Acknowledgement and Receipt form by FAX to 508-819-3390 or email to cflores@omnils.com. |  | 146 | 146 devices | US Nationwide Distribution in the states of: UT, CA, FL, TX, MA, GA, OK, VA, CT, OK, and WY. | Terminated | September 08, 2015 | 97 | 146 | Not_Computer
Z-2067-2015 | 71527 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dual Delivery 4DD-1 of GYC-1000 Green Laser Photocoagulator system |  |  | March 25, 2015 | 2015 | Nidek Inc | Accessories to the GYC-1000 laser were missing Laser Aperture labels. | Error in labeling | Nidek sent Recall Notifications to doctors on March 31, 2015, by Federal Express. Letters were specific to the accessory involved.NIDEK detected that the laser aperture label has not been affixed on the delivery unit of GYC-1000, Green Laser Photocoagulator System and the device does not meet FDA regulation 21 CFR Subchapter J.In order to comply with FDA regulation, Nidek requests that the customer affix the enclosed laser aperture label to the Dual Delivery 4DD-1 of GYC-1000, according to the instructions attached.Please see the enclosed Laser Aperture label, instruction for affixing to the device and the response card.Nidek requests that you complete and return the response card as soon as possible. For additional information, please call Nidek at 1-800-223-9044 or visit http://usa.nidek.com |  | 22 | 22 units | US Distribution to the states of : CA, IL, MN, MO, OR, IL, UT RI, MA, WA, NE, NY, NJ, TX and WI. | Terminated | July 14, 2015 | 111 | 69 | Not_Computer
Z-2072-2015 | 71549 | No | HYA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | A.L.P.S. - Medium Bone Plate Forceps. Orthopedic manual surgical instrument. |  |  | June 17, 2015 | 2015 | Biomet, Inc. | Medium Bone Plate Forceps was assembled with the incorrect grip arm tip. Potential that the surgeon will be unable to provisionally clamp the plate to the patient's bone. | Vendor change control | On 6/17/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for locating and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 223 | 223 | Domestic: TX NJ, NC, KY, MA, SD, MO, AL, IN, GA, ARInternational: Netherlands, Japan, Costa Rica | Terminated | March 22, 2016 | 279 | 223 | Not_Computer
Z-2070-2015 | 71501 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Exactech Ambassador Cervical Plate System |  |  | May 21, 2015 | 2015 | Exactech, Inc. | Shipped to a single consignee prior to completion of final inspection. | Release of Material/Component prior to receiving test results | Exactech contacted their consignee by telephone on May 21, 2015, of this recall for their two products and instructed him to recover the devices from the hospital. Devices were recovered from the hospital May 22, 2015, and returned to the firm on May 26. All devices have been recovered. |  | 4 | 4 devices. | Nationwide Distribution to Florida only. | Terminated | July 16, 2015 | 56 | 6 | Not_Computer
Z-2063-2015 | 71573 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | July 01, 2015 | 2015 | Elekta, Inc. | A problem exists in MOSAIQ resulting in the incorrect field size being sent to the treatment machine for stereotactic plans using cones. | Software design | Elekta sent an Urgent Important Field Safety Notice 371-01-MSQ-003 on July 1, 2015, to all affected customers. The notice informed users of the specific product and version numbers affected and that the problem only exists if all the following criteria is met - (1) Running Mosaiq 2.41; (2) with at least one linac equipped with an Agility 160 MLC head (Versa HD, Infinity, Axesse (EOS) or Synergy); and (3) who are treating with Stereotactic Cones. The notice also advises users of the workaround to avoid the problem. The customers were also instructed to return the Important Field Safety Notice Acknowledgment form. For questions regarding this recall call 770-670-2422. |  | 74 | 74 Total (45 US) | Worldwide Distribution - US including AZ, CA, GA, ID, IL, MO, NJ, NY, NC, OR, SC, TN, WV, and Internationally to Austria, Brazil, Bulgaria, Canada, Colombia, Germany, Israel, Italy, Japan, Mexico, Russia, Taiwan, and Venezuela. | Open, Classified |  |  | 74 | Not_Computer
Z-2082-2015 | 70783 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Autopen Classic Automatic insulin delivery pen, 3.0ml x 2 unit 2-42 units, AN3800, Lot 7JH, NDC 08470-3800-01, |  |  | December 11, 2014 | 2014 | Owen Mumford USA, Inc. | The dials in the Autopens are not working - the dial does not stay dialed, but springs right back to 0. As a result, the patient could not dial up a dose of medicine to deliver via an injection. | Process control | Owen Mumford, Inc. sent an Urgent Medical Device Recall Notification letter dated December 19, 2014, to all affected customers via Fed Ex. The letter instructed customers to check their shelves for the product and specific lot number. The customer was instructed to contact Customer Service to obtain an RMA and a pre-paid, pre-addressed label for return of the devices. A form was included for the documentation of the inspection of their inventory, whether any are in stock and will be returned, and if they will provide a customer list of how they sold any devices to or that they will notify their customers of the issue with the product. They were instructed to fill it out and fax back to 770-977-2866. Customers with questions were instructed to contact Customer Service Department at 1-800-421-6936. For questions regarding this recall call 770-977-2226, ext 13. |  | 225 | 225 units | Nationwide Distribution including FL, IL, NV, NJ, OH, and TN. | Terminated | July 11, 2017 | 943 | 225 | Not_Computer
Z-2083-2015 | 70684 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | The FLOW-i Anesthesia System Axillary O2 and Suction Module, Device Part number: 66 79 847 |  |  | March 02, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Maquet has identified that in some FLOW-i Anesthesia System Auxiliary O2 and Suction Modules, the suctioning function cannot be activated or controlled by the suction unit on/off switch leading to a decrease or loss of suctioning capability of the modules. | Under Investigation by firm | Maquet sent an Urgent Medical Device Field Action Removal and Replacement Form/Device Removal Response Form dated March 2, 2015 to their affected customers via FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed read through the letter and inform all members of their staff who may use the affected device. Fill out the attached Device Removal Response Form with requested information, sign and date and fax/email the Response Form using the fax/email information provided. Customers with questions were instructed to contact their Maquet Service Representative or Maquet Technical Support at 1-888-627-8383 (press option 3, followed by option 1 and then optionn 1 again. |  | 104 | 104 units | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, China, Colombia, Croatia, Czech Republic, Ecuador, Finland, France, Germany, Hong Kong, India, Israel, Japan, Mexico, Netherlands, Peru, Poland, Portugal, Romania, Russia, Serbia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand and United Arab Emirates. | Open, Classified |  |  | 104 | Not_Computer
Z-2075-2015 | 71559 | No | HXY | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Allofix Insertion Kit, 2.0 mm, Single Drill, |  |  | May 27, 2015 | 2015 | Musculoskeletal Transplant Foundation, Inc. | Musculoskeletal Transplant Foundation is recalling a certain lot of Allofix Insertion Kits because they are labeled as having a 2.0mm diameter drill however it may contain a 2.4 mm diameter drill. | Process control | Musculoskeletal Transplant Foundation notified their affected customers of this recall by sending them a Voluntary Recall Notification Letter dated May 27, 2015 or via telephone. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Service at 1-800-433-6576. |  | 117 | 117 units | US Nationwide Distribution in the states of AZ, CA, CO, FL, IL, MI, MN, NC, NJ, NV, PA, TN, TX, UT, WA and WV. | Open, Classified |  |  | 117 | Not_Computer
Z-2076-2015 | 71593 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sorin Stockert Heater-Cooler 3T, 230 V |  |  | June 15, 2015 | 2015 | Sorin Group USA, Inc. | Potential colonization of organisms, including Mycobacteria, in Sorin Heater Cooler Devices, if proper disinfection and maintenance is not performed per Instructions for Use. | Error in labeling | Sorin Group issued a Field Safety Notice dated June 15, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and identify any affected devices. For each unit customers were instructed to determine if the device has been maintained according to the Instructions for Use. If yes, customers should strictly adhere to the new Instructions for Use. Customers were also provide with a Response form to confirm they received, read and understood the Field Notice. Customers were instructed to return the completed form to assist in monitoring the effectiveness of the communication. For technical support customers should call 1-800-221-7943, ext 6355. For questions regarding this recall call 303-467-6306. |  | 2837 | 2837 | Worldwide Distribution - US (nationwide) and Internationally to AL, AT, BE, BG, CH, CY, CZ, DE, DK, EE, ES, FI, FR, GB, GR, HR, HU, IE, IT, LT, LU, LV, MC, MK, MT, NL, NO, PL, PT, RO, RS, SE, SK, AE, AR, AU, AZ, BD, BH, BR, BY, CA, CL, CN, CO, CR, DZ, EC, EG, ET, GE, GY, HK, ID, IL, IN, IQ, IR, JO, JP, KR, KW, KZ, LB, LK, LY, MA, MN, MU, MX, MY, NG, NP, NZ, OM, PA, PE, PH, PK, PR, PS, QA, RE, RU, SA, SG, SV, SY, TH, TN, TR, TT, TW, UA, VN, ZA. | Open, Classified |  |  | 5055 | Not_Computer
Z-2073-2015 | 71515 | adaptor, hygiene | ILS | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Etac Supporter toilet armrest |  |  | June 15, 2015 | 2015 | Etac Supply Center Ab | The Supporter toilet armrest has two Fast Nuts fixing the product to the toilet porcelain. The firm has reports of the two Fast Nuts coming loose and causing the Supporter to become unsteady. | Nonconforming Material/Component | Etac sent an Urgent Notice of Field Correction beginning June 15, 2015, to all affected consignees. The letter directed the consignees to do an in field correction with replacement Fast Nuts. Defined replacement instructions were included with the replacement nuts. Customers with questions were instructed to contact their Etac distributor or the Etac coordinator, Snug Seal Inc. at 1-800-336-7683. For questions regarding this recall call 1-800-367-6844. |  | 25 | 25 units | Nationwide Distribution including GA, CT and PA. | Terminated | July 14, 2017 | 760 | 25 | Not_Computer
Z-2074-2015 | 71552 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Abutment Retrieval Instrument Zirconia CC RP/WP, Article No. 37882 and Abutment Retrieval Kit, Article No. 37508 (Article No. 37882 was assembled in this kit). Used to remove a zirconia abutment if the abutment is stuck in the implant connection due to a tight seat. Endosseous dental implant instrument. |  |  | June 22, 2015 | 2015 | Nobel Biocare Usa Llc | One dimension of the affected instrument is incorrect. Therefore this instrument cannot be used to remove the metal adapter as intended. | Under Investigation by firm | A customer notification letter dated 6/22/15 will be sent to all customers who purchased the Abutment Retrieval Instrument Zirconia CC RP/WP because an internal investigation showed that one dimension of the affected instrument is incorrect. Therefore this instrument cannot be used to remove the metal adapter as intended. The letter provides the problems identified and the actions to be taken. Customers are instructed to complete the acknowledgement form and return it to their local Nobel Biocare Subsidiary with the affected instrument. Customers with the affected instrument that is a part of a kit, Nobel Biocare Representative will contact them directly. Customers with any remaining questions, are instructed to contact their Nobel Biocare Representative. |  | 149 | 149 units total | Worldwide Distribution-US (nationwide) and the countries of Australia, Austria, Belgium, Canada, Finland, France, Germany, Greece, Hong Kong, Hungary, Italy, Luxembourg, Netherlands, New Zealand, Norway, Portugal, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | September 18, 2015 | 88 | 149 | Not_Computer
Z-2094-2015 | 69760 | Yes | NOU | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Breas Vivo 50 Home Care Ventilator |  |  | November 07, 2014 | 2014 | Breas Medical AB | Unintended treatment termination could result from a keypad malfunction in some situations. The device erroneously interprets this as a Stop Treatment Instruction. An alarm will not sound, or be registered. Accessories and monitoring equipment connected to the Vivo 50 will stop functioning as the device enters stand-by mode. | Device Design | Breas Medical sent Urgent Release: Field Safety Notice letters dated November 7, 2014 and Field Safety Notice dated November 21, 2014to their customers. The letters identified the affected product, problem and the actions to be taken. |  | 846 | 846 US | Worldwide - United States Nationwide Distribution and the countries: Argentina, Belgium, Chile, Cyprus, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Malaysia, Netherlands, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand, Turkey, UAE and United Kingdom. | Terminated | April 07, 2016 | 517 | 846 | Not_Computer
Z-2084-2015 | 69804 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NDI Disposable Reflective Marker Spheres for Brainlab IGS (Image Guided Surgical) system. |  |  | June 10, 2015 | 2015 | Northern Digital Inc. | Inadequate curing of adhesive, may cause the two halves of the spheres to separate. | Process control | NDI sent an Urgent Customer Notification letter dated June 10, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Acknowledgement Form attached to the notification letter and return using the enclosed pre-paid FedEx label. For questions, contact NDI at 1 (877) 634-6340 ext. 544 (if calling from North America) or +1 (800) 634 634 00 ext. 544 (if calling from outside North America) . |  | 6120 | 524, 700 spheres (41773 = 440,820 spheres; 41774 = 77,760 spheres and 41813 = 6,120 spheres) | US Nationwide Distribution | Open, Classified |  |  | 6120 | Not_Computer
Z-2090-2015 | 71502 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Flow Cytometers; used on all BD FACSCount, BD FACSCalibur, BD FACSCanto, BD FACSCanto II, BD LSR II, BD LSRFortessa,BD FACSAria, BD FACSAria II, BD FACSAria Ill, BD lnflux, BD FACSArray,BD Pathway 855, BD Pathway 435 and all SORP instruments. |  |  | June 15, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | Instruments installed between June 2007 and September 2011 do not contain labeling pursuant to the Food and Drug Administration's (FDA's) Guidance on Laser Products- Conformance with IEC 60825-1 and IEC 60601-2-22 (Laser Notice No. 50). | Labeling Change Control | BD Biosciences sent an Important Product Correction letter dated June 2015 to all affected customers. Customers were informed that a Field Service Engineer will schedule a time to perform the necessary correction. Customers with questions were instructed to contact BD Customer Support at 877-232-8995 (prompt 2). For questions regarding this recall call 408-954-6307. |  | 2188 | 2188 units | Nationwide Distribution | Open, Classified |  |  | 2188 | Not_Computer
Z-2092-2015 | 71366 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reamer / Irrigator / Aspirator (RIA): RIA Drive Shafts, RIA Tube Assemblies, sterile, RIA Reamer Head, sterile (12.0mm-19.0mm); intended to clear the medullary canal of bone marrow and debris and to effectively size the medullary canal for the acceptance of an intramedullary implant or prosthesis |  |  | May 28, 2015 | 2015 | Synthes, Inc. | It was reported that the RIA Drive Shaft, Tube Assembly, and Reamer Head have the potential to break when incorrectly assembled or used improperly. Precautionary statements are being added to the Reamer / Irrigator / Aspirator (RIA) Surgical Technique Guide . | Labeling False and Misleading | DePuy Synthes sent an Urgent Medical Device Safety Notification dated May 28, 2015, to direct accounts and sales consultants to advise them of the patient risk associated with misassembly or misuse of the device. Customers and sales force were instructed to remove the Surgical Technique Guide J4352 (all versions) dated 11/2011 or earlier and replace it with the updated Technique Guide (DSUS/TRM/0515/0620, Date 06/2015),and to return the completed verification response form.If you have any further questions, please call (610) 719-5450. Thank you for your attention to this matter. |  | 132564 | 132564 | Nationwide Distribution and Internationally to Canada and Switzerland. | Open, Classified |  |  | 132564 | Not_Computer
Z-2091-2015 | 71492 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Helical Blade Inserter for Trochanteric Fixation Nail-ADVANCED (TFNA); intended to secure the head element to the head element inserter to prevent disengagement during head element insertion. Manual orthopedic surgical instrument. |  |  | June 10, 2015 | 2015 | Synthes (USA) Products LLC | The height of one of the three guiding pins of the Helical Blade Inserter for the TFNA is oversized. The Helical Blade Inserter is therefore unable to pass through the Blade/Screw Guide Sleeve (part 03.037.017). | Process control | A recall notification letter, dated June 10, 2015, was sent to sales consultants and Synthes distribution sites requesting return of affected devices. |  | 32 | 32 | Nationwide Distribution. | Open, Classified |  |  | 32 | Not_Computer
Z-2085-2015 | 71555 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Patient Handling System (Motion Control Software), |  |  | April 01, 2015 | 2015 | Viewray Incorporated | ViewRay received a report that the couch moved unexpectedly into the bore after performing a RTCS reboot. | Software Design Change | ViewRay sent a Customer Advisory Notification Letters dated April 1, 2015 to thier customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns regarding this issue, contact ViewRaylM Customer Support at support@viewray.com or call the ViewRaylM Customer Support center at 855-286-8875. |  | 3 | 3 | US Nationwide Distribution in the states of CA, MO &amp; WI | Terminated | February 04, 2016 | 309 | 3 | Not_Computer
Z-2095-2015 | 71585 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Alere i flu, Alere i, Alere Influenza A &amp; B; Part Number (PN) 425-024 |  |  | June 26, 2015 | 2015 | Alere Scarborough, Inc. dba Binax, Inc. | Risk of false negative results due to microbial growth in the Sample Receiver of one lot of the Alere i Influenza A &amp; B kit | Process change control | Alere sent a Urgent Medical Device Recall letter dated June 26, 2015, to all affected consignees. Customers were instructed to cease use and discard any unused pouched Sample Receivers and Transfer Cartridges contained in Alere" i Influenza A &amp; B PN 425-024, Lot 0073853. Pouched Sample Receivers and Transfer Cartridges are identified in the kit as part number 425-431 lot number 074005 or 074108. Replacement Sample Receivers and Transfer Cartridges will be provided to the customer. Customers with questions were instructed to call Alere Technical Services at 855-731-2288 or email ts.scr@alere.com. |  | 257 | 257 Kits | Nationwide Distribution including AR CA CO CT FL GA IA IL IN KS KY LA MA MD ME MO MS NC NE NJ NY OH PA SC TX VA WA WY. | Terminated | September 15, 2015 | 81 | 257 | Not_Computer
Z-2093-2015 | 71350 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NDI Disposable Reflective Marker Spheres for Brainlab IGS (Image Guided Surgery) system. |  |  | May 20, 2015 | 2015 | Northern Digital Inc. | Inadequate curing of adhesive, which may cause the two halves of the spheres to separate. | Process control | NDI sent an Urgent Medical Device Recall letter dated May 20, 2015 to their customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. For question from Canada or the United States contact NDI at 1-877-634-6340 ext 544 and for outside North America call 1-800 634 634 00 ext 544. Customer Service can also be reached at spheres@ndigital.com. |  | 32400 | 32,400 spheres | Worldwide - US Nationwide Distribution and the countries of Australia and Canada. | Open, Classified |  |  | 32400 | Not_Computer
Z-2087-2015 | 71628 | No | PGY | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadiForce GX540, 5MP Monochrome LCD Monitor, Part Code No. 0FTD1826A5, 0FTD1826B5, 0FTD1826C5, UPC 690592034752. |  |  | July 07, 2015 | 2015 | Eizo Corporation | EIZO is recalling the RadiForce LCD Monitors could display abnormally while connecting to particular graphics boards with DisplayPort. | Device Design | EIZO sent An Urgent Field Safety Notice dated July 7, 2015, to all customers who purchased the RadiForce RX440, RX650, RX850 and GX540. The notice informs customers that the RadiForce monitors could display abornormally in 6-bit (64 gradations) instead of normal 8-bit (256 gradations) or 10-bit (1024 gradations) while connecting to particular graphics boards with DisplayPort. The notice provides the customers with the actions to be taken and contact information (1-800-800-5202, M-F, 8am-4pm, PST, us.feedback@eizo.com, 5710 Warland Drive, Cypress, CA 90630). |  | 19 | 19 units | Nationwide Distribution including OH, FL, IL, TX, OH, NY, AZ, and NJ. | Terminated | November 30, 2016 | 512 | 24 | Not_Computer
Z-2086-2015 | 71599 | No | FZO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 16mm Chisel Blade; used in the LCP Anterior Ankle Arthodesis Plate System and the Hip Preservation Surgery Set. Orthopedic manual surgical instrument. |  |  | June 24, 2015 | 2015 | Synthes (USA) Products LLC | It was discovered that a 10mm Chisel Blade was etched as a 16mm Chisel Blade. | Employee error | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated June 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Recipients were asked to check inventory and return affected devices along with the completed response form.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 100 | 100 | US Distribution to the states of : NJ, VA, OH, IN, MI, FL, WI, CA, TN, MD, NC, WA, UT, ID, MO, NV and KS. | Open, Classified |  |  | 100 | Not_Computer
Z-2097-2015 | 71446 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Magnus Hybrid OR Table Column (1180.01B2) used in conjunction with Siemens Artis Zeego Angiography System. Maquet GmbH &amp; Co. |  |  | June 10, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | There is a potential issue that could result to collision between the C-arm of the angiography system and the MAGNUS OR table system. After switching on the angiography-system, in rare cases, MAGNUS column provides wrong position data for the longitudinal movement leading to a collision. | Device Design | Maquet sent a field correction letter dated June 8, 2015, via Fed Ex. Customers are advised of the issue with the device and provided instructions. Questions or additional information can be obtained from the local Maquet representative or Maquet Technical Support at 1-888-627-8383 (option 3 followed by option 1) Monday through Friday between the hours of 8:00 am and 6:00 pm EST. |  | 12 | 12 units | Worldwide Distribution - US Nationwide Distribution and to the countries of : France, Australia, Belgium, Brazil, Switzerland, Chile, China, Germany, Spain, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Japan, Kuwait, Kazakhstan, Mexico, Netherlands, Norway, Russia, Sweden, Turkey, Taiwan and South Africa. | Terminated | April 18, 2016 | 313 | 12 | Not_Computer
Z-2098-2015 | 71582 | No | HDX | Vaginal stent. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Milex Vaginal-Hymenal Silicone Dilators Set of 4 |  |  | June 19, 2015 | 2015 | CooperSurgical, Inc. | Incorrect expiration date on outer carton kit label | Incorrect or no expiration date | CooperSurgical sent an Urgent Medicall Device Recall dated June 19, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected products and complete the attached Acknowledgement and Receipt Form for a replacement. For questions call 203.601.5200. |  | 40 | 40 kits | US Nationwide Distribution in the states of: CA, FL, GA, HI, IL, KY, MD, MN, NC, NY, OH, TX, VT, and W A. | Terminated | February 01, 2016 | 227 | 40 | Not_Computer
Z-2099-2015 | 71427 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 0702-001-350Q, Diverter Valve Assembly |  |  | May 22, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is initiating a recall of the Neptune 2 Waste Management System due to complaints related to installation failures of the Neptune 2 diverter valve assembly during docking. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Notification distributed customer notices on 06/09/2015 via certified mail. Customers were asked to review the notice, complete and return the response card, and forward the notice in the event the recalled device was further distributed. Customers were asked to fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com, and to please continue to rotate all Neptune2 Rovers. If Rovers are removed from circulation for more than 40 days, it is recommended that you contact Stryker Repair if any issue is experienced putting the Rover back into use. |  | 9245 | 9,245 | Worldwide Distribution in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN. KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NY, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the countries of: Canada, Europe, Asia Pacific, EEMEA. | Terminated | September 12, 2016 | 479 | 9245 | Not_Computer
Z-2096-2015 | 69762 | No | GCB | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety PICC Insertion Tray with Chloraprep OneStep - Model #5026CL-9 - 12 pr case. Packaged for Physician use only. |  |  | November 05, 2014 | 2014 | Custom Healthcare Systems, Inc. | An INCORRECT item number is packed within the OUTSIDE "PIGGYBACKED" PACKAGE. The interior sealed STERILE KIT is correct and NO adverse issues affect that part of the item. The NON-STERILE OUTSIDE "PIGGYBACKED" PACKAGE mistakenly contains ITEM #8881579121 - Syringe prefilled, Saline Flush, 09% Sodium Chloride and is a NONSTERILE PRODUCT that CANNOT BE DROPPED INTO THE STERILE FIELD. | Mixed-up of materials/components | Custom Healthcare System sent an Urgent Product Recall letter dated November 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.PLEASE CHECK YOUR INVENTORY AND QUARANTINE ANY AFFECTED PRODUCT MATCHING THE DESCRIPTIONS ABOVE. Call 804 (421-5959) with regards to returning the affected products. Return Authorizations will be issued and product pickup and return will be scheduled. |  | 293594 | 3 Cs-Lot#92414 , 25 Cs-Lot#100214, 24 Cs-Lot#100914 | US Distribution to CT. | Terminated | January 12, 2016 | 433 | 293594 | Not_Computer
Z-1838-2015 | 71056 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PROBEAT WITH DISCRETE SPOT SCANNING SYSTEM |  |  | April 21, 2015 | 2015 | Hitachi America, Ltd., Power Systems Division | At a PROBEAT III site installed in Japan, the operator of the system found that the radiation was delivered inconsistent with the treatment plan and stopped the irradiation. The irradiation control system may be potentially affected when rebooting a part of the control system. | Device Design | The recalling firm notified the one consignee about the incident that occurred in Japan on 4/7/15, then gave them an update on 4/13/15 and 4/21/15. The written notification had details for the consignee on how to prevent the device from malfunctioning. The recalling firm also had held conversations with the consignee to discuss the issue and revised work instructions via telephone conference and on site Hitachi personnel. |  | 3 | 3 | Distribution in US (Texas) and Japan | Terminated | March 24, 2016 | 338 | 3 | Not_Computer
Z-2105-2015 | 71468 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EVac 70 XTRA with Integrated Cable and PROcise EZ View with Integrated Cable |  |  | June 02, 2015 | 2015 | ArthroCare Corporation | Wands inspected with the adaptor fixture were released to the field not meeting the dielectric (Hipot) specifications. | Nonconforming Material/Component | Smith&Nephew; sent an Urgent Notification letter dated June 10, 2015 via mail/courier service with delivery tracking/confirmation to affected customers. A response form is included in the notification letter. A follow up letter will be sent to non-responding consignees. Customers were instructed to complete the acknowledgement form and return via fax or email. |  | 733 | 733 devices | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Germany, Spain, United Kingdom, Ireland, India, Italy and Singapore. | Open, Classified |  |  | 733 | Not_Computer
Z-2103-2015 | 71498 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | Natrelle 133 Series Tissue Expanders |  |  | June 18, 2015 | 2015 | Allergan Inc | Allergan is recalling the NATRELLE 133 and CUI Series Tissue Expanders because the product shipped was beyond the 47 month shelf life. | Incorrect or no expiration date | A recall letter will be sent on 6/18/15 to all customers who received the NATRELLE 133 and CUI series tissue expanders. The letter provides the problem identified and the action to be taken. Customers are instructed to contact GENCO Pharmaceutical Services if the recall actions on the recall letter are unclear. For product returns, customers are instructed to contact GENCO at (877) 319-8966, 7am to 5pm CST. Customers with credit/reimbursements are instructed to contact Allergan at (800) 811-4148, 7am to 5pm PST. Customers with medical inquiries are instructed to contact (800) 433-8871, option 2, 8am-5pm, PST. Customers with adverse events/product complaints are instructed to call (800) 624-4261, option 3, 8am-5pm CST. |  | 167 | 167 units total (114 units in US) | US in the state of Canada | Terminated | December 16, 2015 | 181 | 167 | Not_Computer
Z-2102-2015 | 71429 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker SelectCore Variable Throw Biopsy Device, 14G, 10CM |  |  | June 02, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is recalling the Stryker SelectCore Variable Throw Biopsy Device due to the potential for for the stylet to eject from the SelectCore device during actuation. | Employee error | Stryker distributed Customer Notification Letters via Fed Ex overnight to be delivered on 06/10/2015. Customers are asked to Immediately review this Recall Notification. 2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you do not have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in this Notification.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax (866-521-2762) or email (julie.forsyth@stryker.com) the completed Business Reply Forms; and return the recalled SelectCore devices(s) to Stryker. |  | 55 | 55 units | US Nationwide Distribution in the states of: TX, NJ and MS. | Terminated | September 03, 2015 | 93 | 55 | Not_Computer
Z-2100-2015 | 71430 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ring Transjugular Intrahepatic Access Set |  |  | May 28, 2015 | 2015 | Cook Inc. | Cook Inc. is initiating a voluntary recall of the Ring Transjugular Intrahepatic Access Set and the Flexor RTPS Introduce Sheath due to complaints of the dilators being too long. | Component change control | Written recall communication letters were shipped via FedEx on June 05, 2015 to all US Consignees. Customers were asked to complete and return response forms and recalled products. Recalled devices will be held for final disposition and destruction. Customers wree asked to forward notices if recalled devices were further distributed. Effectivenes checks will be conducted by mail, phone and visits. Customers may contact Cook Medical Relations (800) 457-4500 or 1 (812)- 339-2235 Monday through Friday from 7:30 AM to 5:00 PM or e-mail at CustomerrelationsNA@cook medical.com. |  | 661 | 661 | Worldwide Distribution - US Nationwide in the state of: AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SD, TN, TX, UT, VA, WA, WI, WV and the countries of Canada and EU. | Terminated | May 30, 2017 | 733 | 661 | Not_Computer
Z-2119-2015 | 71113 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hysto Pack, Non-sterile |  |  | April 16, 2015 | 2015 | Covidien LLC | Devon Light Gloves contain splits or holes compromising the sterility | Process control | Medtronic/Covidien initiated recall on April 16, 2015 via FedEx and certified mailCustomers were instructed to discontinue use of the product and return all inventory and kits to Medtronic (Covidien) -Attention: Devon" Light Glove 110 Kendall Park Lane. Atlanta, GA 30336.If you have any questions contact your Medtronic representative or Medtronic Customer Service, Monday through Friday, 8am to 6:30pm ET,at (800) 882-5878. |  | 23560 | 23560 | Worldwide Distribution-US (nationwide) including the Virgin Islands and the countries of Canada, Australia, New Zealand, Costa Rica, Austria, Belgium, Bulgaria, Czech Republic, Denmark, Finland , France, Germany, Netherlands, Sweden, Ireland, ISRAEL,Spain, Italy, Jordan, Kenya, Kuwait, United Kingdom, Lebanon, Liechtenstein, Luxembourg, Monaco, South Africa, Norway, Poland, Portugal, Qatar, Russia, Saudi Arabia, Slovenia, Switzerland, Turkey, and UAE. | Terminated | June 28, 2017 | 804 | 27407865 | Not_Computer
Z-2152-2015 | 70961 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Polaris 4.75 TI Spinal System, 4.0 x 20 MM, Closed MAS, Closed Head Multiaxial Screw, Non-Sterile, Rx Only, REF 14-583020, Biomet Spine and Bone Healing Technologies. Spinal fixation component. |  |  | March 12, 2015 | 2015 | Biomet Spine, LLC | Multiaxial screws may not meet internal requirements related to fatigue strength. Compression testing result and surface treatment are out of specification. | Nonconforming Material/Component | Biomet Spine issued an Urgent Medical Device Recall Notice/Response Form dated 3/12/2015 to their affected customers. |  | 4765 | 4,765 units | Distribution US Nationwide (AL, CA, CO, FL, GA, LA, MA, MD, MI, MO, NY, OH, PA, TX, WA, WI) and The Netherlands. | Terminated | December 21, 2015 | 284 | 4765 | Not_Computer
Z-2069-2015 | 70163 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Various OxyTOTE Portable Oxygen System regulators, Valve Integrated Pressure Regulator, one unit per package. |  |  | January 31, 2015 | 2015 | Western / Scott Fetzer Company | The firm has become aware of the possibility of ignition and rupture (explosion) of a compressed gas oxygen cylinder, which can cause serious injury and/or death. | Component design/selection | Western/Scott Fetzer sent an Urgent Medical Device Recall Notice notification dated January 31, 2015 to their customers via U.S. mail. The notification identified the affected product, problem and actions to be taken. The OxyTOTE/oxyQuik/AirTOTE Distributors and Oxygen Service Providers were instructed to:1) Notify and provide the letter to all accounts/customers to which any OxyTOTE/oxyQuik/AirTOTE products have been distributed. Alternately, please provide a list of customers who received affected product and Western/Scott Fetzer will notify them on your behalf. 2) Referring to the instructions for OxyTOTE/oxyQuik/AirTOTE Users and User Facilities, follow up with each account/customer to ensure that they locate and identify all affected product and establish the number of units in their possession subject to the recall. 3) Contact Western/Scott Fetzer to schedule remediation for any affected OxyTOTE/oxyQuik/AirTOTE product included under this notice. We will begin scheduling product for remediation on February 17, 2015. 4) Please complete the Recall Acknowledgement and Receipt Form and return it to Western/Scott Fetzer as soon as complete responses are available.The OxyTOTE/oxyQuik/AirTOTE Users and User Facilities (Hospitals, Health Care Facilities, Health Care Service providers, etc.) were instructed to: 1) Locate and identify any OxyTOTE/oxyQuik/AirTOTE product in your possession. 2) Review the instructions to determine if the product IS or IS NOT marked and if remediation is needed. 3) Contact your OxyTOTE/oxyQuik/AirTOTE supplier or Medical Gas Provider directly to schedule the remediation of product in your possession.For questions or concerns regarding this notification, please visit the website for additional details or contact Western/Scott Fetzer:Website: www.westernenterprises.com/recallBy Email: recall@WesternEnterprises.comBy Phone: (800) 783-7890, extension 2516By Fax: (440) 835-8283By Mail: Recall Coordinator, Wester |  | 162678 | 162,678 units | Worldwide Distribution -- US, including the states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Delaware, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, North Dakota, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, West Virginia, and Wyoming; and, the countries of CANADA, CHILE, GERMANY, ISRAEL, LEBANON, MACAU, MEXICO, PUERTO RICO, and SINGAPORE. | Open, Classified |  |  | 162678 | Not_Computer
Z-2155-2015 | 71509 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Cochlear Nucleus Sterile Silicone Template |  |  | June 05, 2015 | 2015 | Cochlear Americas Inc. | Cochlear Americas is recalling Nucleus Sterile Silicone Template CI24RE/CI422 part number Z421736 because expired product was distributed. | Process control | Cochlear sent an Urgent Medical Device Recall letter dated June 5, 2015 via Fed-Ex shipment. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return affected products for replacement. |  | 26 | 26 | US Nationwide Distribution to CO, OH, OR, MI, and CA. | Terminated | August 05, 2015 | 61 | 26 | Not_Computer
Z-2123-2015 | 71574 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision MPi fluoroscopic imaging systems. |  |  | June 12, 2015 | 2015 | GE Healthcare | A reported incident of a monitor that fell from the suspension. A fall of a monitor could result in an injury to a patient or operator. Recall affects the single and dual LCD monitor suspensions used on Precision 500D Classical R &amp; F system, Precision RXi 23A/32A Analog System, Precision RXi Digital system and Precision MPi fluoroscopic imaging systems. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated June 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Perform a visual check before and after each use for the presence of the four (4) mounting screws as shown in the pictures below. Check the secureness of these screws by hand. If these screws are present and secure, you may continue use of the system.If any of these screws are missing or can easily be turned by hand, follow the precautions below before you continue use with the monitor(s).1. Position the monitor suspension in the most frequently used position and limit further movement as much as possible.2. Clearly inform (e.g. through signage and verbal instructions) the users and other auxiliary personnel who may come in contact with the system to not move the monitor suspension once in position.3. Contact your GE Healthcare representative to inform them of any loose or missing screws.For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 788 | 788 (701 US; 87 OUS) | Worldwide Distribution - US Nationwide Distribution including GU and DC., and to the countries of : Australia, Bulgaria, Canada, Chile, Dominican Republic, Germany, Ireland, Italy, Kuwait, Lebanon, Portugal, Romania, Russia, Saudi Arabia, Slovakia, Spain, Taiwan,Turkey, United Kingdom, Venezuela and Yemen. | Terminated | August 03, 2017 | 783 | 788 | Not_Computer
Z-2153-2015 | 71411 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb SofPort Advanced Optics Aspheric Lens |  |  | May 26, 2015 | 2015 | Bausch &amp; Lomb Surgical, Inc. | Haptics breaking during lens loading and insertion. | Device Design | Bausch + Lomb sent an Urgent Medical Device Recall letter dated May 29, 2015 to their custometers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: segregate the identified IOLs, return the IOLs through your B&L; representative or by using the prepaid shipping label included in this package, order replacement product through customer service (800) 338-2020 or your B&L; representative and complete &amp; return the acknowledgement form included in this package. Customers were advised that if replacement product is needed for a scheduled surgery before your full return can be processed, please contact your B&L; representative to expedite shipment of the IOL. For questions contact Bausch + Lomb at (800) 338-2020. |  | 72773 | 72,773 units. 52,787 units (US) - 19,986 units (Foreign) | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MD, MN, MO, MS, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, Hawaii and Puerto Rico and countries of: Canada, Brazil, Argentina, Austria, Belgium, Switzerland, Germany, Spain, France, Great Britain, Georgia, Greece, Kyrgyzstan, Norway, Russia, Ukraine, United Arab Emirates, China, Australia, India Malaysia, Korea, and Taiwan. | Open, Classified |  |  | 72773 | Not_Computer
Z-2137-2015 | 71421 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ATTUNE FB PS ARTICULATION SURFACE SIZES 3-8 |  |  | June 12, 2015 | 2015 | DePuy Orthopaedics, Inc. | Potential for the Balseal, a small wire spring coil located on the post feature of the ATTUNE Knee Tibial Articulation Surface Instruments used in trialing to become damaged and disassociate (come off). If the Balseal is separated from the post, it has the potential to enter the surgical site and be left in the patient if the surgeon is unaware of the disassociation. | Device Design | URGENT INFORMATION: MEDICAL DEVICE CORRECTION was issued for all lots of the affected devices. The device correction is to inform US Distributors and Medical Professionals about the issue, how to prevent the issue, and any actions to take. US Distributors will provide Hospitals/User Facilities with the Medical Professional Notice and Acknowledgment Cards. US Distributors will also work with the sales consultants to ensure all consultants complete sales training regarding this issue. |  | 34676 | 34676 units | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Internationally to CANADA , AUSTRALIA, AUSTRIA, BELGIUM, ENGLAND/ UK, FRANCE, GERMANY, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JAPAN, MALAYSIA, NETHERLANDS, NEW ZEALAND, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH KOREA, SPAIN &amp; IBERIA, SWEDEN, SWITZERLAND, THAILAND, UAE, CHILE, COLOMBIA, MEXICO, PERU, and PUERTO RICO. | Terminated | November 12, 2015 | 153 | 280591 | Not_Computer
Z-2125-2015 | 71380 | No | JJJ | Beta or gamma counter for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | WIZARD2 2-detector, 550 samples |  |  | June 03, 2015 | 2015 | Perkinelmer | Error in WIZARD2 Barcode ID Label #023 Contentmay produce erroneous results | Nonconforming Material/Component | Perkin Elmer issued the recall letter on June 3, 2015. Action to be taken: Until the PerkinElmer correction has been completed, we request you immediately destroy all barcode ID labels #023 in your possession. The WIZARD Gamma Counter barcode label binder contains a total of 20 barcode labels identified as #023 (i.e. 2 labels on each of 10 pages). Please ensure you remove thebarcode labels ID #023 from all barcode label binders in your possession. Please be sure to inspect all sample racks to remove and destroy any barcode ID labels #023 already affixed to racks. A WIZARD2 Gamma Counter barcode label binder with corrected barcode labels ID #023 will be provided upon request and return of the completed Correction Response Form to PerkinElmer. Correction Response Form and return it by scan and e-mail to QMresponse.SG@perkinelmer.com |  | 53 | 53 units | Worldwide Distribution-US (nationwide) and the countries of ARGENTINA AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA CHILE CHINA COSTA RICA CROATIA CZECH REPUBLIC DENMARK ETHIOPIA FINLAND FRANCE GERMANY GREECE HONG KONG HUNGARY INDIA IRELAND ISRAEL ITALY JAPAN KOREA KUWAIT LUXEMBOURG MEXICO MYANMAR NETHERLANDS NORWAY PHILIPPINES POLAND REUNION SAUDIA ARABIA SINGAPORE SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UK. | Terminated | August 04, 2017 | 793 | 422 | Not_Computer
Z-2149-2015 | 71569 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | THOMAS by Gardner Denver, MEDI-PUMP(R) Aspirator, Model 1240 |  |  | June 04, 2015 | 2015 | Gardner Denver Thomas Inc | Vacuum pump may not perform to specification. | Nonconforming Material/Component | Gardner Denver Thomas sent an Important Medical Device Recall letter dated May 7, 2015 to affected distributors. The letter identified the affected products, problem and actions to be taken. For questions call your Customer Service Representative or Sales Representative or the Recall Coordinator at 920-451-6423. |  | 160 | 160 units | Internationally Distributed in the countries of :Mexico, Peru, El Salvador, Chile, Colombia, Vietnam, and Malaysia. | Open, Classified |  |  | 1038 | Not_Computer
Z-2147-2015 | 71597 | No | MPO | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(k) | Infant Heel Warmer and Infant Gel Warmer, sold under the brand names CooperSurgical and Fisher brand as follows : |  |  | July 06, 2015 | 2015 | CooperSurgical, Inc. | Product marketed without a 510(k) | No Marketing Application | CooperSurgical sent an " Urgent WarmGel Infant Heel Warmer" letter dated June 30, 2015, FEDEx, confirmed delivery receipt on July 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CooperSurgical is advising all WarmGel users to review your inventory of this product, remove it from your inventory and discontinue its use. informing them of the reason to cease using the product and either discard or return product to CooperSurgical. Accounts are provided with a response form to complete. Any questions contact CooperSurgical at 203 601 5200 Ext. 310 |  | 6344700 | 6,344,700 | Worldwide Distribution - Nationwide Distribution and to the countries of : AUSTRALIA, BERMUDA, ECUADOR , FRANCE, NEW ZEALAND, SAUDI ARABIA and UNITED KINGDOM. | Open, Classified |  |  | 6344700 | Not_Computer
Z-2133-2015 | 71654 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5i/CORE Intravascular Ultrasound Systems software versions 3.3 and 3.4 (with and without iFR Scout Technology) and v.3.4 software kits; |  |  | June 22, 2015 | 2015 | Volcano Corporation | A software issue was discovered where an inaccurate FFR/iFR value could be calculated under certain circumstances. | Software design | Urgent Medical Device Correction letters were sent on June 22, 2015 to all affected sites via FED Ex. |  | 4007 | A total of 4007 devices, all models | Worldwide Distribution-US (nationwide) and regions in Europe and the countries of Japan, Canada, South Africa, Taiwan, Thailand, Singapore, Philippines, Australia, New Zealand, El Salvador, Malaysia, Mexico, Lebanon, Kuwait, South Korea, Kenya, Jordan, Iran, India, Israel, Hong Kong, Egypt, Algeria, Dominican Republic, Costa Rica, Chile, Brazil, Bolivia, and Argentina | Terminated | April 14, 2017 | 662 | 4007 | Computer
Z-2148-2015 | 71635 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | The EKOS EkoSonic Control Unit is intended exclusively for use with the EkoSonic Endovascular Device and the EkoSonic MACH4 Endovascular Device. |  |  | June 24, 2015 | 2015 | EKOS Corporation | The Connector Interface Cable (CIC) was not recognized by the EkoSonic PT-3B control unit. | Vendor change control | Ekos sent an "Urgent Voluntary Medical Device and Replacement" letter dated June 24, 2015, to US customers. The letter identified the product the problem and the action needed to be taken by the customer. The replacement CIC and recall letter were sent by FEDERAL EXPRESS directly to the person responsible for maintaining the EkoSonic System. Tracking numbers and delivery notifications will be retained for each shipment. Letters to Risk Management and Purchasing were sent by REGISTER MAIL. The return cards will be retained for each letter.Customers with questions can contact EKOS Customer Service at 1-888-400-3567. |  | 46 | 46 units | Nationwide Distribution | Terminated | August 25, 2015 | 62 | 46 | Not_Computer
Z-2170-2015 | 71534 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Fiber Cleaver; |  |  | May 28, 2015 | 2015 | American Medical Systems Innovation Center - Silicon Valley | validation data related to cleaning instructions and sterilization methods and the methods defined in instructions for use (IFU) for the Sureflex Reusable Fibers and the Reusable Stripper and Cleaver Accessories do not meet the new guidance on qualifying reusable products | Reprocessing Controls | Urgent Medical Device Field notification letters were sent on June 9, 2015. Translation for countries out of the US will be conducted. Communications are being coordinated by the Prostate Health QA Director.Urgent Medical Device Field notification Letters advised customers of the issue and list all affected products with name and model numbers. Any SureFlex Reuseable Fibers which have not been used but are in inventory and with the 3 year shelf life should be returned as per instructions enclosed.Any of the SureFlex devices in inventory that have been used should be discarded or returned.All Reuseable Stripper and Cleaver Accessories that have not been used should be returned. Used devices should be destroyed or returned. Questions to 1-408-456-3320. The Acknowledgement form should be completed and returned. |  | 16710 | 16,710 total all affected devices | Worldwide Distribution - US Nationwide in the countries of: Argentina France Morocco, South Africa,Australia, Germany, Netherlands, Spain,Austria, Greece, New Zealand, Sri Lanka,Belgium, Guatemala, Norway, Switzerland,Bolivia, India, Panama, Thailand,Brazil, Ireland, Peru, Turkey,Canada, Israel, Poland, United Kingdom,Chile, Italy, Portugal, United States,Colombia, Korea, Qatar, Uruguay,Costa Rica, Lebanon, Romania,Finland, Malaysia and Saudi Arabia. | Terminated | January 19, 2016 | 236 | 16710 | Not_Computer
Z-2189-2015 | 71533 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | T-Lok Bone Marrow Biopsy Needle 8G x 4 Ext Can, Part Number DBMNJ0804TL; T-Lok Bone Marrow Biopsy Needle 11G x 4 Ext Can, Part Number DBMNJ1104TL; Bone Marrow Biopsy Needle 13G x 4, Part Number DBMNJ1304 |  |  | June 18, 2015 | 2015 | Argon Medical Devices, Inc | Potential defect in the forming of the primary package of the blister that may compromise the sterility of the product inside the tray. | Packaging process control | The recalling firm's sales force was notified of the recall on 6/17/15. The recalling firm will begin to send out Urgent Medical Device Recall letters by registered mail to affected customers on 6/18/15. Customers are instructed to inspect and quarantine and return any affected product and complete the Recal Response Form. |  | 10290 | 10290 units | US | Open, Classified |  |  | 16680 | Not_Computer
Z-2190-2015 | 71525 | No | DWO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mermaid Medical M-Biopsy Coaxial Introducer Needle, Coaxial Needle for M-Biopsy 18Gx10cm-PRO-MAG, RX Only. Single packed device in a tyvek pouches. 10 single packed devices in a |  |  | June 15, 2015 | 2015 | Mermaid Medical A/S | Mermaid Medical A/S is recalling some of their M-Biopsy Coaxial Introducer Needles because of the potential for small plastic pieces to be present in the luer of the cannula. | Process control | The firm, Mermaid Medical, notified both customers directly via email with written letters dated June 12 and June 15, 2015 followed up by sales and customer service for product return to MMI. The customers were instructed to "Please coordinate the return of the mentioned products with corresponding lot numbers with our Mermaid Medical's Customer Care the recalled product: either by phone at 1{720}420-9999 or by email: customercare@mermaidmedical.com".If you have any questions please do not hesitate to contact Vice President QA &amp; Regulatory Affairs toll free at (855) 427-6799 or (720) 420-9999 or email: customercare@memaidmedical.com. |  | 320 | 320 | US Distribution to FL including PR. | Terminated | December 07, 2015 | 175 | 320 | Not_Computer
Z-2197-2015 | 71699 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Ready CT/PET System. |  |  | July 01, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips has identified four (4) software defects in the Tumor LOC software application that may potentially result in irradiation of healthy tissue or non-irradiation of diseased tissue. | Software design | On 7/6/2015 the firm sent Field Safety Notices to their customers. |  | 1 | 1 unit | Worldwide Distribution - US Nationwide, Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, Israel, Italy, Japan, Lebanon, Netherlands, Portugal, Saudi Arabia, South Africa, Switzerland, Thailand, Venezuela &amp; Vietnam. | Open, Classified |  |  | 86 | Computer
Z-2200-2015 | 71689 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Pinnacle3 Software Version 9.0, 9.2, 9.4 and 9.6, Model Numbers 453560446041, 459800091001, 459800220161, 459800232931, 459800235871, 459800338451. |  |  | July 31, 2014 | 2014 | Philips Electronics North America Corporation | Philips, Pinnacle Radiation Treatment Planning System version 8.0h, 8.0k, 8.0m, 8.0n, 9 0, 9 2 9 4, 9 6 is being recalled because the dose may be inconsistent with the density of a density-overridden ROl. | Software design | Consignees were sent on 7/31/2014 a Philips "Urgent Field Safety Notice" dated 2014-July 20. The letter described the problem, product involved in the recall and the action that should be taken by the customers. For questions contact Customer Care Solutions Center at 1-800-722-9377.UPDATE: A second Philips letter dated 2 July 2015 was sent in August 6, 2015. This letter updated the recall to incorporate the additional products being recalled. |  | 1642 | 1642 | Worldwide Distribution-US: All states in the US including DC and PR. OUS: Australia, Austria, Bangladesh, Belgium, Brazil, China, Denmark, France, Germany, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands,New Zealand, Oman, Philippines, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, United Kingdom, Viet Nam.UPDATED: Canada, Bulgaria, Egypt, Hungary, Jordan, Luxembourg, Poland, Reunion, Romania, Russia, Sauth Arabia, Turkey. | Open, Classified |  |  | 1642 | Computer
Z-2202-2015 | 71473 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT c8000 System List Number 01G06-98 |  |  | June 11, 2015 | 2015 | Abbott Laboratories | The ARCHITECT c8000 instrument contains tubing that does not meet specifications and has the potential to leak. | Mixed-up of materials/components | Abbott notified thier affected consignees via phone or visit on 5/29/15. A follow up Product Recall notification letter dated June 4, 2015 was sent to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use of the product, Abbott representative will be make all necessary arrangements for replacing your impacted instruments(s), follow your laboratory protocol regarding the need to review previously reported patient results and retain this notification for laboratory records. For questions contact Customer Service at 1-877-4ABBOTT. |  | 5 | 5 units | US Nationwide Distribution in the states of CA, AL, NY and WI. | Terminated | February 09, 2016 | 243 | 5 | Not_Computer
Z-2201-2015 | 71650 | No | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | iovera 155 Smart Tip; |  |  | July 01, 2015 | 2015 | Myoscience Inc | The expiration date on the outer box label and the pouch label for the iovera Smart Tip product is incorrect. Instead of indicating the correct expiration or 2015-12. the labels were incorrectly labeled 2016-05. | Incorrect or no expiration date | Myoscience sent an Urgent Medical Device Correction letters dated July 1, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are requested to quarantine any affected product at their site. Myoscience will make arrangements to either correct the incorrect expiation date label or remove the product and replace it with correctly labeled product immediately.Question or concerns may be directed to Customers Service at 510-933-1500. |  | 36 | 12 boxes - 3 boxes of 10 count and 9 boxes of 5 count. | US Nationwide Distribution in the states of TX, NY, MD, CO and CA. | Terminated | September 08, 2015 | 69 | 36 | Not_Computer
Z-2205-2015 | 71227 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Head Holder shipped with Ingenuity CT Computed Tomography X-ray Systems |  |  | March 17, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | A head holder released for use with other CT systems but not yet qualified for use on the Ingenuity CT, which was shipped with this system. | Mixed-up of materials/components | On 6/4/2015, the firm mailed Customer Information Letters. |  | 66 | 66 units | Worldwide Distribution-Product was shipped to the following states: AR, AR, AZ, CA, CO, CT, FL, GA, IA, IN, KS, LA, MA, MD, MI, MO, MS, NC, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, VA, VT, WA, WV &amp; WY, and to the following countries: Argentina, Australia, Bangladesh, Belgium, Brazil, Bulgaria, China, Colombia, Denmark, Ecuador, France, Germany, India, Indonesia, Israel, Japan, Latvia, Libya, Myanmar, Norway, Oman, Poland, Republic of Korea, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom &amp; Viet Nam. | Open, Classified |  |  | 200 | Not_Computer
Z-2191-2015 | 71263 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | American Catheter / Cholangiogram Catheter 9100 Series. |  |  | March 17, 2015 | 2015 | American Catheter Corp | Catheters were not sealed before sterilization. and was not detected through subsequent acceptance activities, product release and distribution. | Process control | Consignees were contacted by telephone of the affected lot number and sent a Recall Script dated 3/17/2015 giving instructions how to respond to this recall. |  | 890 | 890 devices. | Distributed in the states of AL, AK, CA, CO, FL, GA, HI, IL, KS, LA, MI, MO, MT, NJ, OH, OK, SC, TN, TX, WA, and WI. | Completed |  |  | 890 | Not_Computer
Z-2192-2015 | 71577 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HardyCHROM VRE Agar-RSR |  |  | June 24, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling HardyCHROM VRE Agar due to lack of 510(k) clearance. | No Marketing Application | The firm initially called customers to inform them of the recall. The firm's phone script stated that they were discontinuing the product and as part of the process were recalling anything that still might be within expiration date. The phone script stated that the firm was going to be sending something in writing.On 06/24/15 the firm sent out a notification letters to customers to inform them that they were discontinuing the product, and were requested anything remaining within expiration to be discarded. The firm requests that customers complete, sign, and fax back the enclosed form stating their compliance with the above action. For any questions the firm states to contact Technical Services Department at 800-266-222, option 2 or via email at TechService@HardyDiagnostics.com. |  | 1500 | 1,500 | Distributed to one U.S. customer in SC. | Terminated | August 28, 2015 | 65 | 1500 | Not_Computer
Z-2208-2015 | 71615 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Silhouette Lift brochure |  |  | July 02, 2015 | 2015 | Silhouette Lift, Inc. | Silhouette Lift is recalling the Silhouette Suture brochure because it does not include a complete list of risks and side effects. Therefore, it can potentially mislead the patient regarding the safety of the product. | Labeling False and Misleading | Silhouette lift sent an "Correction of Information for Silhouette Suture" letter dated July 2, 2015, in response to FDA untitled letter. The letter identified the product the problem and the action needed to be taken by the customer.Action Requested:All brochures for Silhouette Suture should be destroyed and no longer be used for discussions with patients.Customers with questions are instructed to call (951) 734-7701. |  | 2260 | 2,260 units | US Distribution to the states of : CA, TX, IA, MA, RI, GA, WI, PA, VA, MO, FL, NE, MI, NY and CO., and Internationally to Canada. | Terminated | February 05, 2016 | 218 | 2260 | Not_Computer
Z-2207-2015 | 71337 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Tni-Ultra Assay; SMN 10317708 - 100 tests and SMN 10317709 -500 tests |  |  | May 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | System-to-system bias between the Tni-Ultra assay onthe ADVIA Centaur CP and ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT systems. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated via Federal Express to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Effectiveness Check attached to the letter.For questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 21159 | SMN10317708 = 20544 kits &amp; SMN 10317709 = 21159 kits | Worldwide distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MN, MS, MT, NC, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, WV, and Puerto Rico; and the countries of Algeria, Austria, Bahrain, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iceland, Indonesia, Ireland, Italy, Jordan, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Pakistan, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, AR, AU, BD, BR, CA, CL, CN , CO, CR, DO, EC, HK, ID, IE, IL, IN, JP, KR, LK, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, TH, TW, UY, VN, and ZA. | Open, Classified |  |  | 21159 | Not_Computer
Z-2254-2015 | 71124 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray Systems Brilliance CT 64-channel with Essence technology. |  |  | February 05, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Software change control | On 2/5/2015 the firm sent Customer Information Letters to their customers. |  | 4 | 4 Units | Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV &amp; WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 424 | Not_Computer
Z-2260-2015 | 70897 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MiniMed NGP 640G 1.8ml (mmol/L), Model No. MMT-1511; NGP 640G 1.8ml (mg/dL), Model No. 1512; NGP 640G PLGM 3ml (mmol/L), Model No. MMT-1711; NGP 640G PLGM 3ml (mg/dL), Model No. MMT-1712. |  |  | June 19, 2015 | 2015 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the MiniMed 620G and 640G insulin pumps because there are certain scenarios where the set Bolus screen will not timeout, which could cause confusion by showing a bolus amount that is no longer appropriate. | Under Investigation by firm | A customern notification letter sent to customers on 6/19/15 who received the MiniMed 640G insulin pumps. The purpose of the letter is to inform customers that when using Bolus Wizard with the MiniMed 640G insulin pumps, that there are certain scenarios where the set Bolus screen will not timeout, which could cause confusion by showing a bolus amount that is no longer appropriate. Customers are informed that Medtronic MiniMed has updated the user guide which instructs to respond promptly to set Bolus screens, and remind not to activate delivery on a bolus based on a blood glucose value that is more than 12 minutes old. The letter informs the customers of the summary of changes in the MiniMed 640G insulin pump user guides. |  | 1936 | 1936 units | Australia, Czech Republic, Denmark, Finland, France, Ireland, Netherlands, Norway, Sweden, United Kingdom. | Terminated | October 02, 2015 | 105 | 1936 | Not_Computer
Z-2261-2015 | 71715 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ottobock Pheon 3R62 Knee Joint |  |  | July 16, 2015 | 2015 | Otto Bock Healthcare GmbH | Otto Bock Healthcare GmbH is voluntarily recalling all 3R62 Pheon polycentric prosthetic knee joints distributed since February 2014 because the locking function may fail under certain conditions which may result in a fall. | Component design/selection | The firm, Ottobock, sent an "URGENT MEDICAL DEVICE RECALL 3R62 Pheon" letter dated 16 July 2015 to its customers. The letter described the product, problem, and actions to be taken. The customers instructed to read this notice in its entirety; examine your stock for 3R62 Pheon devices; quarantine 3R62 Pheon devices and prevent further distribution; contact Ottobock customer service at (800) 328-4058-option 1 or (512) 806-2000 -Option 1, to obtain a return authorization for those devices, and complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form via email to: USRecalls@ottobock.com "Pheon Recall"; fax (801) 974-6718-"Pheon Recall" or mail to: Otto Bock HealthCare LP, ATTN: PHEON RECALL, 3820 W. Great Lakes Drive, Salt Lake City, UT 84120-B7 is re-designing the locking function to prevent this failure from occurring. The re-designed devices will be available for replacement sometime within the next few months. If you have any questions contact Director, Regulatory &amp; Quality Compliance at (801) 974-6618 or email: Lex.Pearce@ottobock.com. |  | 518 | 518 (439 US; 79 OUS) | Worldwide Distribution: US (nationwide) and countries of: AU, BE, CN, CZ, IT, NL, and SK. | Terminated | June 21, 2017 | 706 | 518 | Not_Computer
Z-2209-2015 | 71459 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | InjectionPlasty 1.0 Bone Void Filler; an injectable and biocompatible calcium phosphate bone void filler. Sterilized using irradiation. The single-use InjectionPlasty 1.0 Kit contains the necessary components for mixing and delivery of the bone void filler. The InjectionPlasty 1.0 sterile kit contains: Calcium Phosphate Powder, Dilute Sodium Silicate Liquid, a Mixing System, Targeting K-Wires and a Delivery Needle (Cannula). |  |  | May 29, 2015 | 2015 | Skeletal Kinetics, Llc | Reports of venous uptake, DVT, pulmonary embolism in patients who were implanted with Injection Plasty Bone Void Filler. | No Marketing Application | Skeletal Kinetics sent letters to distributors by overnight mail with instructions to forward a Dear Doctor notification letter on May 14, 2015. Second letters dated May 29, 2015, advised that the firm is removing the product from the marker. Distributor letters asked that the Dear Doctor letter be forward within 24 hours of receipt of the letter. All Injection Plasty kits in their possession and those with customers should be retrieved and returned by Fed Ex. Dear Doctor letters advised that their distribution agent has been requested to return all unused kits. Questions may be directed to 408-366-5002 of duran@skeletalkinetics.com. |  | 67 | 67 kits | US Distribution including Puerto Rico. | Terminated | July 24, 2015 | 56 | 67 | Not_Computer
Z-2253-2015 | 71524 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | OxySure Model 615 disposable replacement cartridges, for portable oxygen generator. |  |  | June 12, 2015 | 2015 | Oxysure Systems, Inc. | Flow rate of the recalled replacement cartridges does not meet the draft FDA standard of a minimum 6 LPM for a minimum of 15 minutes. | Under Investigation by firm | The recalling firm will notify their affected customers via email beginning the week of 6/15/15. The recall notification will have a response for to be completed by the consignee. The recall notification calls for the customers to dispose of the affected product according to the MSDS. |  | 1005 | 1005 units | Nationwide throughout the US and Chile, Trinidad, China | Open, Classified |  |  | 1005 | Not_Computer
Z-2210-2015 | 71564 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Cortical Pin 2.0 mm,Sterile, FREEZE DRIED: Store at ambient temperature. Do Not freeze. |  |  | May 27, 2015 | 2015 | Musculoskeletal Transplant Foundation, Inc. | Musculoskeletal Transplant Foundation (MTF) is recalling one lot of MTF's Allograft Cortical Pins because the lot was labeled as a 2.0 mm diameter pin and it may contain a 2.4 mm diameter pin. | Labeling mix-ups | Musculoskeletal Transplant Foundation initiated this recall by either telephoning their customers or by sending a recall letter dated 5/27/15 to the customer. |  | 22 | 22 units | Distributed in the states of AZ, CO, MI, MN, NC, OH and TX. | Terminated | December 17, 2015 | 204 | 22 | Not_Computer
Z-2218-2015 | 71608 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Medical Device Exchange - Versaport Bladeless Trocar w/ Fixation |  |  | February 11, 2015 | 2015 | EXP Pharmaceutical Services Corp | EXP did not register as a medical device establishment, list the devices being recalled, orestablish its own quality system and instead relied on the fact that its vendors were registered and had their own quality systems. | Other | EXP first notified their customers to quarantine EXP devices February 11-17, 2015. Customers were contacted through telephone, e-mail, and written communications. Customers were instructed to recover and quarantine all EXP products from their inventories. EXP believes that they have contacted all customers to whom products were shipped between April 26, 2014, and December 9, 2014.A second letter is scheduled to be sent in July 2015 (date unknown at this time). This letter will advise users of the recall of due to lack of registration with the FDA. |  | 1 | 1 | Distributed in DC and the states of CO, CA, FL, TX, and AR. | Terminated | December 30, 2015 | 322 | 174 | Not_Computer
Z-2266-2015 | 71541 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Expression MRI Patient Monitoring System. |  |  | June 17, 2015 | 2015 | Philips Electronics North America Corporation | The device's labeling inaccurately indicates Endotracheal as a body temperature application site. The FlexTEMP reusable temperature sensor is not intended to be used for endotracheal placement. The application site should be listed as Esophageal. | Labeling False and Misleading | An URGENT Medical Device Correction was sent to all consignees by letter June 2015 regarding the Invivo Expression MRI Patient Monitoring System FlexTEMP Application Site explaining the reason for the affected labeling. |  | 697 | 697 devices. | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OR, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, and Alaska, Hawaii and the District of Columbia, and the countries of Argentina, Australia, Austria, Canada, Czech Republic, Denmark, Egypt, France, Germany, Hong Kong, Iceland, Israel, Italy, Japan, Netherlands, Norway, Oman, Poland, Portugal, Russian Federation, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 697 | Not_Computer
Z-2267-2015 | 71583 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | OxiMax N-65 Handheld Pulse Oximeter. N65, N65-1, N65P, and N65P-1. For continuous or spot check monitoring of functional arterial oxygen saturation (SpO2) and pulse rate. |  |  | June 29, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Potential missing segments on the display that can result in misinterpretation of data. | Device Design | Metronic sent an Urgent Medical Device Correction letter dated June 29, 2015 to affected customers. The letter identified the adfected product, problem and actions to be taked. Customers where instructed to return the attached verification form. For questions contact your Medtronic representative. |  | 211426 | 211,426 units | Worldwide Distribution - US Nationwide including Canada. | Open, Classified |  |  | 317257 | Not_Computer
Z-2269-2015 | 71376 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical P-440 Portable Ceiling Lift; PN 303070 and 303070R (refurbished); Patient lift with weight specification of 440 lbs (200 kg). |  |  | June 24, 2015 | 2015 | ErgoSafe Products, LLC (DBA) Prism Medical | There is a potential the sling loops may not stay attached to the carry bar hooks. | Other | Prism Medical sent a Field Safety Notice to affected customers via certified first class US mail. The notice identified the affected product, problem and actions to be taken. For questions contact your local Prism Medical dealer or Prism Medical Customer Service Representative 1-800-891-6502 |  | 7105 | 7105 lifts | Worldwide Distribution - US Nationwide and the countries of:Canada, Great Britain, Australia, South Korea, and Japan. | Terminated | April 03, 2017 | 649 | 7514 | Not_Computer
Z-2262-2015 | 71663 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | FOUNDATION KNEE SYSTEM NON STERILE Femoral A/P Cut Block - Type II REF 800-01-139 |  |  | June 24, 2015 | 2015 | Encore Medical, Lp | The markings on the distal face of the instruments are mis-oriented by 180Âż. This includes Anterior and Posterior, which the surgeon uses to correctly place the cut block on the resectioned face of the femur. The cut block is not symmetric, so correct orientation of the instrument in the A/P direction is critical to making proper anterior, posterior, and chamfer cuts to fit the femoral implant. | Component design/selection | DJO Surgical sent an " Urgent Field Safety Notice" dated June 30, 2015, to the affected customer. The recalling firm requested the customer return the affected product for a replacement. For further questions, please call (512) 834-6302. |  | 10 | 10 units | US Distribution to TX and GA. | Open, Classified |  |  | 10 | Not_Computer
Z-2265-2015 | 71636 | No | EIH | Porcelain powder for clinical use. | Dental | Dental | 510(k) | IPS Empress Universal Glaze Spray 270 mL, Product Code/Model/Catalog Number 609432AN, Rx only. Porcelain glaze spray used in the fabrication of dental restorations |  |  | July 09, 2015 | 2015 | Ivoclar Vivadent, Inc. | Complaints were received of coarse particles in the glazed surface of IPS e.max Ceram restorations. A basic component of the glass contains a higher number of large grain sizes than usual. | Nonconforming Material/Component | Ivoclar Vivadent sent an URGENT MEDICAL DEVICE RECALL Letter (dated 7/8/2015) and Response Form to affected customers via Certified Mail Return Receipt Request. The letter identified the affected product, problem and actions to be taken. Customers were instructed to notify dentists of the recall and request return of product. A notification letter was sent to dentists on 7/24/2015 requesting they destroy the product on hand (if any) and report back to Ivoclar Vivodent. For questions in the United States contact US Headquarters Ivoclar Vivadent Customer Services at 1-800-533-6825. |  | 27 | Domestic: 27 units | US Distribution in the states of: PA, IN, IA and SC. | Open, Classified |  |  | 27 | Not_Computer
Z-2271-2015 | 71592 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius MRI Anesthesia Machine; Fabius MRI is an inhalation anesthesia machine for use in MRI environments in operating, induction and recovery rooms. |  |  | June 22, 2015 | 2015 | Draeger Medical, Inc. | the Fabius MRI or parts of the system were attracted by the magnetic field of the MRI. If the Fabius MRI is positioned too close to the MRI, its parts can be loosened as a result of the magnetic force of the MRI. | Labeling False and Misleading | Draeger Medical sent an Urgent Medical Device Recall Letter, dated June 2015, and Instructions for Use Supplement (one supplement for each Fabius MRI Anesthesia Machine at the facility) to end users. The letter identified the affected product, problem and actions to be taken. For questions regarding the recall letter contact Michael Kelhart at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of the DrÂżger Fabius MRI anesthesia machine contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 373 | 373 | US Nationwide Distribution | Terminated | May 05, 2017 | 683 | 373 | Not_Computer
Z-2263-2015 | 71570 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Paragon Laboratory Management |  |  | June 15, 2014 | 2014 | McKesson Technologies, Inc. | For Paragon Laboratory Management 12.1 and 12.1.1 releases, if used with Microsoft SQL Server 2012, the Final Cumulative Report may incorrectly display the data. | Software design | McKesson issued a Clinical Alert to notify customers of the issue on June 15, 2015, and requested that customers disable Cumulative Reporting. McKesson is identifying customers currently in a live environment with Microsoft SQL Server 2012 and the affected configuration. McKesson will provide those customers with corrective software, and provide assistance in generating correct reports and updating those to the patient folder. For further questions, please call (404) 338-3556. |  | 128 | 128 | Nationwide Distribution including Puerto Rico and to the states of : CA, CO., CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OA, PA, SC, SD, TN, TX, VT, VA, WV and WY. | Open, Classified |  |  | 128 | Computer
Z-2272-2015 | 71595 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Synthes TI Matrix Pre-Bent Maxillary Plate / 2mm ADVMNT/0.8mm THK/L and Synthes TI Matrix Pre-Bent Maxillary Plate / 2mm ADVMNT/0.8mm THK/R. The Synthes Matrix ORTHOGNATHIC Fixation System is intended for use in selective trauma of the midface and craniofacial skeleton; craniofacial surgery; reconstructive procedures; and selective orthognathic surgery of the maxilla, mandible and chin in adults and children. |  |  | June 24, 2015 | 2015 | Synthes (USA) Products LLC | The affected part and lot numbers of the TI Matrix Pre-Bent Maxillary Plates may potentially have the incorrect laser etch denoting the orientation. | Process control | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated June 24, 2015, to all affected customers. Customers were asked to check their inventory and return affected devices with the completed response form. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. |  | 31 | 31 | Nationwide Distribution including CA, UT, CO, NJ, WA, VT, KY, WY, CT, NY, MS, and MD. | Open, Classified |  |  | 31 | Not_Computer
Z-2280-2015 | 71212 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ANSPACH - Fluted Drum, Side Cut only 10MM x 14MM , REF L-10DRC-G1 |  |  | April 13, 2015 | 2015 | The Anspach Effort, Inc. | G1 Dissection Tools are not compatible with all attachments listed on the labeling; G1 cutting tools are labeled with a Specified Shaft Exposure which may not be consistent when used with each attachment on the product label. Shaft exposure may be greater. | Device Design | Anspach sent an Urgent Notice: Medical Device Recall (Correction) letter of Anspach G1 Dissection Tools dated May 6, 2015, to all affected customers. Customers were not required to return their inventory of GI Dissection tools. Customers were instructed to complete, sign, and return the attached reply form confirming receipt of the letter by fax or email. Customers were also instructed to forward the letter as appropriate immediately. Customers with questions should contact the Complaint Handling Unit Manager at 561-494-3673 or contact their Anspach Sales Representative.For questions regarding this recall call 561-627-1080. |  | 65 | 65 tools. | Worldwide Distribution - US including AZ, CA, CO, FL, GA, IL, KY, LA, MN, MO, NC, NJ, NY, OH, OR, PA, TX, and Internationally to Canada and China. | Open, Classified |  |  | 5320 | Not_Computer
Z-2279-2015 | 71602 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Hardy Diagnostics |  |  | June 26, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling MacConkey Agar with Ciprofloxacin, Brain Heart Infusion Broth with Ciprofloxacin 1 ug, MDR Acinetobacter, BEA with Azide and Vancomycin, BEA Agar with Vancomycin, Gentamicin, and Amphotericin B, BEA Broth with Vancomycin, and VRE Broth due to lack of 510(k) clearance. | No Marketing Application | The firm initially called customers to inform them of the recall. The firm's phone script stated that they were discontinung the product and as part of the process were recalling anything that still might be within expiration date. The instructions provided via phone were to discard any remaining plates that are in stock in order for credit. The phone script stated that the firm was going to be sending something in writing. On 06/25-29/15 the firm sent out notification letters to customers to inform them that they were discontuing the product, and were requested anything remaining within expiration to be discarded. The firm requests that customers complete, sign, and fax back the endlosed form stating their compliance with the above action. The notification states if the laboratory was unable to be reached via telephone and would like a credit to contact the Customer Service Department at 800-266-2222, option 1. For any questions the firm states to contact Technical Services Department at 800-266-222, option 2 or via email at TechService@HardyDiagnostics.com.Follow-up calls will be made to customers who do not respond. |  | 900 | 900 each (90 pk/10) | U.S. Nationwide Distribution and the country of: Canada | Terminated | September 11, 2015 | 77 | 194840 | Not_Computer
Z-2314-2015 | 71687 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S 1000, ACUSON S 2000, or ACUSON S 3000 ultrasound systems with software version C3, C3, C3, or C1. |  |  | June 25, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system when using the Doppler manual trace measurement tool in full screen format that has a low probability of misdiagnosis. | Device Design | Customer Safety Advisory notifications were sent by certified mail to US customers on June 25, 2015. Siemens Service organization coordinate the implementation of the field correction with the distributors for all countries.Letters advised users of the potential miscalculation on ACUSON S1000, ACUSON S2000, or ACUSON S3000 ultrasound systems with software version VC30A, VC30B, VC30C, or VC31A. The manual trace tool may incorrectly calculate the Time Average Velocity and subsequently incorrectly calculate the Pulsatility index. This could lead to potential misdiagnosis.To mitigate the problem users should not perform measurement utilizing the manual Doppler trace tool in a full screen display format.This issue will be resolved in a future software release. If you have any questions, please contact your Siemens Service for more information.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue.Siemens is continuing to ship in forward production and will be including the customer safety advisory notice with the shipment. Siemens will update the software for those customers as well once it is released. For further questions please call (650) 694-5398. |  | 1789 | 1789 units | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Algeria, Angola, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Bulgaria, Canada, Chile, Colombia, Croatia, Cyprus, Czech Republic, Dem. Rep. Congo, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Malaysia, Mexico, Morocco, Nepal, Netherlands, New Zealand, Norway, P.R. China, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad,Tobago, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela and Vietnam. | Open, Classified |  |  | 1789 | Computer
Z-2315-2015 | 71605 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Aequalis Humeral Nail Targeting Jig, Product Code: 9020060. Orthopedic manual surgical instrument. |  |  | June 25, 2015 | 2015 | Tornier, Inc | Recall for the Aequalis IM Nail instrumentation set (Tray Number 9020000) due to several reports that the mast of the targeting jig (Part Number 9020060) is separating from the jig boom. Separation will impact alignment and can affect the ability to fixate Aequalis IM Nail screws. | Nonconforming Material/Component | Consignees were sent on 6/25/2015 a Tornier "Urgent Medical Device Recall" letter dated June 24, 2015. The letter described the problem and the product affected by the recall. The letter provided information on jig inspection and referred consignee to the attached "Perfect Circles' surgical technique that should be used when separation is observed. Requested consignees to complete and return the questionnaire to FieldAction@tornier.com. For questions they can contact Customer Service Department at 1(888)494-7950. |  | 99 | 99 | US: AL, AK, AZ, CA, CO, CT, FL, IL, IN, ME, MD, MI, MO, NC, OH, PA, TX.OUS: CANADA AND FRANCE. | Open, Classified |  |  | 99 | Not_Computer
Z-2322-2015 | 71550 | No | DJZ | Breath-alcohol test system. | Toxicology | Toxicology | 510(K) Exempt | Mission |  |  | June 17, 2015 | 2015 | Acon Laboratories, Inc. | Acon Laboratories, Inc. is recalling Mission Breath Alcohol Detector due to a labeling amendment. | Labeling False and Misleading | On 06/17/15 the firm sent out customer notification letters to their distributors. The letter identified the affected product, problem and actions to be taken. In the letter the firm requests that Mission Breath Alcohol Detectors be quarantined. The firm asks that distributor complete the "fax back" form to allow for the return of unexpired product for a replacement. Once the form is received the firm is going to issue a return authorization number as well as provide a return label.The contact noted in the letter is Qiyi Xie MD, MPH, Sr. Officer Regulatory and Clinical Affairs at 858-857-8011. |  | 1283 | 1,283 kits | U.S. Nationwide Distribution in the states of: NY, CA, FL, VA, LA, NJ, MN, PA, MI, OH, TX, VA, UT, AL, and GA | Terminated | December 08, 2015 | 174 | 1283 | Not_Computer
Z-2321-2015 | 71688 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 and ACUSON S3000 ultrasound systems with the Virtual Touch IQ option; |  |  | June 25, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system. When repositioning the Virtual Touch IQ region of interest from the original (default) location, the lateral position of the measurement tool may not align with the lateral position of the shear velocity data. | Software Design Change | On June 25, 2015 Customer Safety Advisory Notification letters were sent to all affected customers by Certified mail. The letter identified the affected product, problem and actions to be taken. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. For questions contact your Siemens Service for information. |  | 629 | 629 units | Worldwide Distribution - US Nationwide and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Kenya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. | Open, Classified |  |  | 629 | Not_Computer
Z-2316-2015 | 71300 | No | HXD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SYMMETRY SURGICAL DISPOSABLE TOWEL CLAMP, REF 19-4137, STERILE EO, Rx only. |  |  | June 16, 2015 | 2015 | Symmetry Surgical, Inc. | Lack of sterility assurance. | Package design/selection | The firm initiated their recall by via US certified mail, with return receipt requested, beginning on 06/18/2015. The Urgent Medical Device Recall letter, dated June 16, 2015, identified the affected product, as well as the reason for the recall. The letter discussed the potential harm associated with the issue, and requested that the consignees examine their inventory and discontinue use immediately. Customers should promptly call (800) 251-3000 for a return material authorization (RMA) number to support the return of the impacted lots. Symmetry Surgical Customer Service should be contacted with questions or any other needs. |  | 2708 | 2,708 units | Worldwide Distribution -- US, including the states of CA, VA, TX, FL, OH, NY, MS, KY, UT, GA, PA, MO, IN, NV, AL, MI, MA, and LA; and, the countries of Germany and Mexico. | Open, Classified |  |  | 2708 | Not_Computer
Z-1991-2015 | 71649 | No | RBS |  |  |  |  | Olympus Scientific Solutions Americas Analytical X-ray System. Delta, Delta Element, Delta Plus, DPO-4000-XX-U-EN-EN-PM-SP12, DPO-6000-XX-U-EN-EN-AP-SP10. |  |  | June 09, 2015 | 2015 | Olympus Scientific Solutions Americas | Higher than published scatter around the probe head when testing some samples. At the point the customer would interact directly with the instrument (Trigger position), most readings were lower. However, measurements around the probe head were higher particularly on samples know to be highly scattering (plastics, soils) with minimal increases on high density samples (steel, aluminum). | Component change control | Olympus Scientific Solutions planned action:1. Notify product managers as to potential issues 2. Distribute new table and notice to instrument owners 30 JUL 153. Post new table on website for access 15 JUN 15 |  | 1047 | 1047 | Nationwide Distribution. | Open, Classified |  |  | 1047 | Not_Computer
Z-2326-2015 | 71709 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S4C Occiput Torque Wrench F/Set Screw; used to tighten the S4 Cervical Spinal System clamping screws with a defined torque 2.8 Nm for tightening the clamping screw in the screw head/hook head/occipital plate/transverse connector and other connectors. |  |  | July 07, 2015 | 2015 | Aesculap, Inc. | A FW103R torque wrench that was improperly reassembled after reprocessing caused a screw to break during implantation due to too much torque. The firm did not include instructions for reassembly after reprocessing. | Labeling design | Aesculap sent Field Safety Notification Letters, dated July 7, 2015, to end users, sales consultants, and distributors to advise them about the issue. Enhancements to the instructions for use were made available on the firm's website. |  | 10 | 10 | Distributed US (nationwide) and to Canada. | Terminated | April 17, 2017 | 650 | 10 | Not_Computer
Z-2325-2015 | 71756 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Passive Biopsy Needle Kit, Part Number 9731754 - Passive Biopsy Fee Per Use (FPU) Kit. |  |  | July 17, 2015 | 2015 | Medtronic Navigation, Inc. | The adhesive between the sphere assembly and the biopsy needle shaft may be compromised, causing the sphere assembly to move on the needle shaft. This may result in inaccurate navigation and improper placement of the needles biopsy window in the brain. | Process change control | Medtronic sent an Urgent Medical Device Recall letter dated July 17, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for any of the affected lots and quarantine them if found. They will be asked to complete a confirmation form indicating they have completed this action and are to email or fax the form to Medtronic. Customers will be asked to contact Medtronic to obtain a return material authorization number and arrange for no-charge replacements of impacted products. Customers with questions were instructed to contact Medtronic Technical Services at 800-595-9709. For questions regarding this recall call 720-890-3302. |  | 1415 | 1415 total | Worldwide Distribution - US (nationwide) and Canada. | Terminated | April 07, 2016 | 265 | 1415 | Not_Computer
Z-2327-2015 | 71740 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Titan Spine Endoskeleton Locking Bone Screw Assembly, which is part of the Endoskeleton TCS Interbody Fusion Device, spinal implant. Part Numbers / Product Description: 5302-3514 / 03.5 x 14mm Locking Screws; |  |  | July 20, 2015 | 2015 | TITAN SPINE, LLC | Titan Spine is recalling specific Endoskeleton TCS Integrated Locking Bone Screw Assembly due to the locking bone screw collars not 'locking'. | Nonconforming Material/Component | Physicians and Distributors were sent on 7/20/2015 separate Titan "Urgent Medical Device Recall" letters dated July 20, 2015. The letters described the problem and the product involved in the recall. Advised consignees to complete and return the enclosed response form. For questions they contact Jane Rodd Senior AQ/RA Manager at 262-242-7801. Distributors were advised to examine their inventory and quarantine the product. If they have further distributed the product they should identify their customers and advise Titan Spine of their contact information. |  | 137 | 137 units | Distributed in the states of CA, TX, OK, MA, NY, IN, OR, and GA. | Terminated | October 20, 2015 | 92 | 137 | Not_Computer
Z-2328-2015 | 71584 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | June 05, 2015 | 2015 | Philips Electronics North America Corporation | Multiple software and hardware issues with device that can affect its function. | Device Design | On June 29, 2015, Philips sent an Urgent Medical Device Correction notification/Field Safety Notice (FSN) to inform customers of the issue. The FSN identifies details of the units affected, gives instructions on actions to be taken by the customer, and identifies what action Philips plans to take to remedy the issue. An upgrade consisting of both software and hardware will be provided free of charge for all units affected by these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades. Philips is asking customers to follow the "Action to be Taken by Customer/User" section of the Urgent Medical Device Correction notification/Field Safety Notice.For questions contact your local Philips representative. |  | 13168 | 13,168 devices | Worldwide Distribution - US Nationwide including PUERTO RICO, Canada, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICA, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GABON, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBERIA, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACAU, MALAYSIA, MALTA, MAURITIUS, MEXICO, MOROCCO, MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, RWANDA, SAUDI ARABIA, SENEGAL, SERBIA, SEYCHELLES, SINGAPORE, SLOVAKIA (Slovak Republic), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, United Republic of TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VIETNAM and ZIMBABWE. | Open, Classified |  |  | 13168 | Computer
Z-2331-2015 | 71712 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT S(40)-4R, Model Number 10248671 |  |  | July 14, 2015 | 2015 | Siemens Medical Solutions USA Inc. | Emission computed tomography systems and computed tomography xray systems with optional ceiling mounts may be missing a locking ring. | Packaging change control | The firm initiated their recall on 07/14/2015 by telephone to the sole US consignee. The update is scheduled for 07/15/2015. |  | 1 | 1 unit | Distributed in the state of Texas and the countries of Germany, France, Russia, Denmark, and Switzerland. | Terminated | March 03, 2016 | 233 | 12 | Not_Computer
Z-2340-2015 | 71711 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis zee systems; |  |  | June 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The possibility exists for a large amount of liquid to seep through the top cover of the table and penetrate the components inside the table. If this occurs, table movements may no longer be possible. No injuries reported. | Device Design | A customer safety notice, dated June 29, 2015, was sent to direct accounts to alert them about the issue and that Siemens' service organization will be contacting them to perform the corrective action. |  |  |  | US Nationwide | Open, Classified |  |  |  | Not_Computer
Z-2330-2015 | 71609 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | S2 Fluid Pack on the Cobas b 221 system- Catalog Number 03260925184 |  |  | June 24, 2015 | 2015 | Roche Diagnostics Operations, Inc. | The affected S2 fluid packs may generate a transponder error when the fluid pack is loaded on to the cobas b 221 system. The error prevents the fluid from being recognized by the system and the instrument cannot start measurements until the affected S2 Fluid Pack is replaced. | Process control | Roche Diagnostics sent an Urgent Medical Device Correction dated July 2, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to do the following: 1. Keep in the inventory a backup S2 Fluid Packs from an unaffected lot, or at least four additional S2 Fluid Packs from the affected lots when product from unaffected lots is not available.2. If the S2 Fluid Pack generates a Transponder error message, contact the Roche Support Network Customer Support Center at 1-800-526-2272 (provide the lot and serial number printed on the side of the affected S2 Fluid Pack).3. After contacting Roche Support Network Customer Support Center, dispose of the affected S2 Fluid Pack per local guidelines. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272, if you have questions about the information contained in this Urgent Medical Device Correction. |  | 407564839 | 19 affected devices of lot 21450781 were distributed in the US. | US Distribution: CT, TX, LA, ND, MI, NJ, OH, OK, PR, SC, and KY. | Terminated | January 12, 2016 | 202 | 407564839 | Not_Computer
Z-2345-2015 | 71659 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Battery Power Line II Oscillator Model 530.710, orthopedic surgical power tool. |  |  | July 01, 2015 | 2015 | The Anspach Effort, Inc. | Saw head separated from the drive unit or began to come loose. | Nonconforming Material/Component | DePuy Synthes sent an " Urgent Notice Medical Device Recall (Removal) letter dated July 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.ACTION REQUIRED:Our records show that your facility has the affected product(s) subject to this removal. The Anspach Effort, on behalf of DePuy Synthes asks that you review your inventory and immediately remove any affected BPLII Battery Oscillator devices from stock. In addition to this letter, your DePuy Synthes Sales Consultant has been provided with new product(s) to immediately replace the affected product(s).If you have any questions regarding this removal, please contact the Complaint Handling Unit Manager at 561-494-3673 or contact your DePuy Synthes Sales Consultant. |  | 360 | 360 handpieces (Currently there are 9 affected units remaining in the US and 17 units OUS all others were reworked) | Worldwide Distribution - US Distribution to the states of : AR, LA, OH, and WA., and the countries of : China, France, Italy, Russia, South Africa, Switzerland, and United Kingdom. | Open, Classified |  |  | 360 | Not_Computer
Z-2346-2015 | 71510 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium 36Fr Right Angled Firm PVC Thoracic Catheter, (Pkg of 10) |  |  | June 19, 2015 | 2015 | Atrium Medical Corporation | Inner pouches incorrectly labeled as part 8136, 36Fr Right Angled Thermosensitive PVC catheter, but are part 15236-36Fr Right Angled Firm PVC Catheter | Employee error | Atrium Medical notified accounts by letter dated 6/18/15, explaining the mislabeled inner units. If any product is discovered with incorrect labeling, the product may be used or product may be returned for credit. contact your local Atrium/Maquet chest drainage representative or Atrium Customer Service at 1-800-528-7486 (603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 62 | 62 cases (10/cs)=620 units | AZ, CA, LA. IN, NC, TN, TX, VA | Terminated | March 21, 2016 | 276 | 62 | Not_Computer
Z-2318-2015 | 71692 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Beacon Tip Royal Flush Plus High-Flow Catheter. Angiographic catheter. |  |  | July 02, 2015 | 2015 | Cook Inc. | Reports of catheter tip splits and or separation. Loss of device function, medical intervention to retrieve separated segment, or complications resulting from separated tip occluding blood flow to end organs may result. | Component design/selection | Cook Medical sent an Urgent Medical Device Recall letter dated July 2, 2015, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification and instructions for returning the affected product. All customers were notified via courier. Customers with questions were instructed to contact Cook Medical Customer Relations at (800) 457-4500 or (812) 339-2235.Firm issued press release 8/3/2015. |  | 268169 | 268169 | Worldwide Distribution - US (nationwide including Puerto Rico and US Virgin Islands), Algeria, Antigua, Armenia, Austria, Australia, Bahamas, Belgium, Bangladesh, Bulgaria, Bahrain, Brazil, Brunei Darussalam, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Ghana, Great Britain, Greece Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kuwait, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxemburg, Latvia, Libya, Macau, Malaysia, Malta, Martinique, Mexico, Mongolia, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Switzerland, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 967886 | Not_Computer
Z-2264-2015 | 71719 | Yes | NPT |  |  | Cardiovascular | PMA | EnVeo R Loading System, Model No. LS-EnVeoR-23, LS-EnVeoR-2629, LS-EnVeoR-23-C, LS-EnVeoR-2629-C. |  |  | July 15, 2015 | 2015 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic is recalling the EnVeo R Loading Systems because of the presence of particulate. | Process control | An Urgent Medical Device Recall Letter dated July 2015 will be be sent to customers who purchased the EnVeo R Loading System because Medtronic has received 8 reports related to this issue out of 7347 potentially affected units. The letter informs customers of the problems identified and the actions to be taken. Customers are instructed to complete the customer confirmation certificate and to email RS.CFQFCA@Medtronic.com or fax to Medtronic at (651) 367-0612 to the attention of Customer Focused Quality. |  | 6912 | 6,912 units total (540 units in US) | Worldwide Distribution - US Nationwide, Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Kuwait, Latvia, Lithuania, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Terminated | May 20, 2016 | 310 | 6912 | Not_Computer
Z-2344-2015 | 71697 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S Family ultrasound systems utilizing the 18L6 HD and/or 6C1 HD biopsy guidelines |  |  | June 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | When using the CIVCO biopsy attachments in combination with the ACUSON S Family on-screen guidelines, the biopsy needle may traverse outside the on-screenguidelines. | Under Investigation by firm | Siemens sent a Customer Safety Advisory Notification letters to affected customers via certified mail with return receipt. The letter identified the affected product, problem, potential risk to patients and actions to be taken. Seimens instructed that this notice be shared with all personnel within your organization who need to be aware of this issue. For questions contact your Siemens service representative. |  | 4140 | 4140 total | Worldwide Distribution - US Nationwide and the countries of Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Brazil, Brunei, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Dem. Rep. Congo, Denmark, Dominican Rep., Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Luxembourg, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, P.R. China, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Venezuela and Vietnam. | Open, Classified |  |  | 4140 | Not_Computer
Z-2339-2015 | 71788 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Signa Infinity MR System; Signa MR/i MR System; SmartSpeed; EchoSpeed Plus; HighSpeed Plus, Signa Infinity MR System with EXCITE Technology, Excite II, Signa 3T Infinity with Excite, Signa Excite 1.5T and Signa Excite 3T, Signa Excite HD, Signa HD 1.5T and Signa HD 3.0T, Signa HDx, Signa HDxt, Signa HDi. Whole body magnetic resonance diagnostic imaging scanners. |  |  | July 10, 2015 | 2015 | GE Healthcare | A gradient cable in the ACGD cabinet may be missing a spacer that prevents connecting the gradient cable terminals in the wrong polarity on X gradient amplifier. This could cause images to be flipped left-right with incorrect orientation annotation. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter dated July 10, 2015, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees with questions were instructed to contact GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 82 | 82 (39 USA; 43 OUS) | Worldwide Distribution - US (nationwide) and Internationally to UNITED ARAB EMIRATES, ARGENTINA, BRAZIL, INDIA, EGYPT, CHINA,FINLAND, FRANCE, UNITED KINGDOM, HUNGARY, JAPAN, INDONESIA, KOREA, MEXICO, NEW ZEALAND, CANADA, AUSTRALIA, ITALY, RUSSIAN FEDERATION, TAIWAN, and SOUTH AFRICA. | Terminated | November 19, 2015 | 132 | 82 | Not_Computer
Z-2348-2015 | 71575 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 18inch Dual Port/Single Bladder Disposable Tourniquet Cuff, Zimmer Surgical, Inc. Tourniquets are intended to be used by qualified medical professionals to exert enough pressure on the arterial blood flow in a limb to produce a bloodless operating field. Single bladder tourniquet cuffs are generally used for operations lasting less than 90 minutes. |  |  | June 24, 2015 | 2015 | Zimmer Surgical Inc | The inside of the folded Instructions For Use (IFU) pamphlet was missing the usage instructions and warnings. | Labeling False and Misleading | Zimmer sent an Urgent Medical Device Correction letter dated June 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification, identify the affected product and replace the IFU with the enclosed corrected IFU.Customers were also instructed to complete the attached Response Form and return it via email to CorporateQuality.PostMarket@zimmer.com. Customers with questions were instructed to call 330-364-0989. |  | 61 | 61 units | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV &amp; WY and Internationally to Canada. | Open, Classified |  |  | 391 | Not_Computer
Z-2347-2015 | 71814 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo, 9.10, 9.20.0, 9.20.1,9.20.2, 9.30, 9.40.0, 9.40.1, 9.40.2 with Massimo PHASEIN End Tidal CO2 (EtCO2) module. Merge Hemo (formerly named HeartSuite Hemodynamics) monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. |  |  | January 21, 2015 | 2015 | Merge Healthcare, Inc. | A system freeze-up of the Merge Hemo system, that included the PHASEIN End Tidal CO2 unit, resulted in the user needing to move the patient to another cath lab. | Software design | Merge Healthcare sent an Important Product Information letter dated January 21, 2015, to all affected consignees. The letter described the Issue, Affected Product, Instructions, Product Updates, and Contact Information. For questions contact Merge Customer Service at (877) 741-5369 or support@merge.comA second UPDATE letter was sent April 16, 2015. The letter included the Product Updates.For questions regarding this recall call 866-261-5761. |  | 451 | 451 | Nationwide Distribution including IL, SC, WI, AL, CT, CA, LA, NM, TX, OK, GA, IA, OH, FL, CO, TN, NY, NC, MA, AR, MO, HI, ND, AK, UT, WV, WA, AZ, MI, PA, ME. | Terminated | December 09, 2015 | 322 | 451 | Not_Computer
Z-2352-2015 | 71694 | No | GBQ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ReSolve Locking Drainage Catheters |  |  | June 29, 2015 | 2015 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily conducting a recall due to a potential discrepancy between the unit labeling of the Merit ReSolve Drainage Catheters and the actual device in thepackage. | Employee error | Merit Medical Systems sent an Urgent Product Recall Notice dated July 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions required of you:1. Please immediately determine if any of the devices are within your facility, quarantine them, and discontinue use. Due to the nature of this label discrepancy, you will need to physically verify whether product in inventory is within the scope of this recall. Your Merit Sales Representative can assist you with this determination.2. Ensure that all personnel to whom the devices were distributed are made aware of this field action.3. Please work with your Merit Sales Representative to arrange product return.4. If you have any questions concerning this communication, please don't hesitate to contact your Merit Sales Representative or call (801)208-4365. |  | 950 | 950 | Nationwide Distribution including the Virgin Islands. | Terminated | October 26, 2015 | 119 | 950 | Not_Computer
Z-2323-2015 | 71630 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | ReSTOR intraocular lens and ReSTOR Toric intraocular lens. |  |  | April 15, 2015 | 2015 | Alcon Research, Ltd. | There are reports of post - operative inflammation and/or toxic anterior segment syndrome (TASS) after cataract surgery. | Process design | Alcon Research sent recall notification letters dated April 16, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using the affected product and place on hold until their sales representatives visit their site for withdrawal. Follow up communication activities included personal contact via phone and site visits. For questions regarding this recall call 817-293-0450. |  | 45391 | 45,391 | Distribution Japan, No US distribution. | Terminated | June 30, 2017 | 807 | 45391 | Not_Computer
Z-2351-2015 | 71614 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | PathoDx Strep B Grouping Latex, REF 62031, 3 ml packaged in plastic dropper bottle with a pink cap. |  |  | June 30, 2015 | 2015 | Remel Inc | A reagent present may produce weak or slow reactions; continued use may result in a failure to identify target species. | Under Investigation by firm | On 4/28/2015, the recalling firm's European subsidiary notified the recalling firm of the recall. On 6/30/2015, the recalling firm notified their customers of the recall via letter, sent USPS standard mail. |  | 4420 | 4420 bottles | Nationwide Distribution, including Puerto Rico and the country of Canada. | Open, Classified |  |  | 6279 | Not_Computer
Z-2356-2015 | 71718 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Trumpf Ceiling Mounted Surgical Lighting Systems,TruVidia SD, User Manual #1400191. |  |  | July 27, 2015 | 2015 | Trumpf Medical Systems, Inc. | Complaints received of the front joint of the spring arm AC 2000 could break at the interface to the monitor. | Process control | Consignees were notified by letter on/about 7/27/2015. |  | 9287 | 9,287 units total | Worldwide Distribution-US (nationwide) including Puerto Rico, and the countries of Angola, Antilles, Argentina, Austria, Australia, Azerbaijan, Bosnia and Herzegovina, Belgium, Bulgaria, Bahrain, Brasilia, Belarus, Canada, Switzerland, Chile, China, Columbia, Cyprus, Czech Republic, Germany, Denmark, Algerian, Ecuador, Estonia, Egypt, Spain, Finland, France, Gabon, United Kingdom, Georgian Republic, Greece, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iraq, Iran, Iceland, Italy, Jordan, Japan, Kenya, South Korea, Kuwait, Kazakhstan, Lebanon, Liechtenstein, Lithuania, Luxembourg, Latvia, Libya, Morocco, Macedonia, Myanmar/Burma, Malta, Mauritius, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Pakistan, Poland, Palestine, Portugal, Qatar, Romania, Russia, Saudi Arabia, Sudan, Sweden, Singapore, Slovenia, Slovakia, San Marino, Syria, Thailand, Turkmenistan, Tunisia, Turkey, Taiwan, Ukraine, United Arabian Emirates, Uruguay, Venezuela, British Virgin Islands, Vietnam, Kosovo, Serbia and South Africa.* | Terminated | May 03, 2017 | 646 | 9287 | Not_Computer
Z-2358-2015 | 71613 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | InvacareÂż ProntoÂż Air Personal Transporter (power wheelchair), Invacare Corporation, One Invacare Way, Elyria, Ohio, 44035-4190, (800) 333 -6900. |  |  | February 02, 2015 | 2015 | Invacare Corporation | The batteries used in the InvacareÂż ProntoÂż Air Personal Transporter models power wheelchairs have the possibility to overheat during charging which could potentially result in leaking battery acid and/or swelling battery, loss of use, chemical burns and/or property damage. | Device Design | Invacare sent an " Urgent Medical Device Recall " letter on April 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.You should immediately contact your provider at the number below so a replacement batteries can be installed.If you have any questions concerning these instructions, please call Invacares recall support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. Visit www.invacare.com/recall for more information about this recall. |  | 1057 | 1,057 finished devices | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico, and to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA and WI., and to the countries of : AUSTRALIA, CANDA, FRANCE, NEW ZEALAND, SOUTH AFRICA, TRINIDAD AND TOBAGO | Terminated | September 14, 2015 | 224 | 1057 | Not_Computer
Z-2360-2015 | 71677 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon PS X3 Tibial Insert; No. 3 Triathlon TS Plus Tibial Insert X3 Poly 19 mm |  |  | May 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report that 5537-G-319 No 3. Triathlon TS Plus Tibial Insert X3 Poly 19 mm i insert was in a package labeled as a 5532-G-511 Triathlon PS Stibialism Insert. | Under Investigation by firm | Stryker Orthopaedics sent an Urgent Medical Device Recall Notification letter dated July 9, 2015 and Product Accountability Forms via return receipt. The letter identified the affected product, problem, potential hazard, risk mitigation and actions to be taken. Contact Stryker Orthopaedics Sales Representative to arrange for return of the affected product. For questions call 201-831-5826. |  | 26 | 26 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 26 | Not_Computer
Z-2366-2015 | 71431 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Oasis Coil Extension Cable used with Hitachi Oasis MRI System |  |  | April 25, 2014 | 2014 | Hitachi Medical Systems America Inc | The firm received a complaint stating the Coil Extension Cable accessory would overheat and be a burn hazard if touched. | Nonconforming Material/Component | Hitachi sent an Urgent: Device Removal Notice letter date April 24, 2014 to their customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers to discontinue using the Coil Extension Cable and that a Field Service Engineer will schedule a visit to remove the cable. For questions call 800-800-3106 ext 3720. |  | 173 | 173 units | US Nationwide in the states: AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NJ, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WY and the country of Mexico. | Open, Classified |  |  | 173 | Not_Computer
Z-2361-2015 | 71665 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-Stat PT/INR Cartridges |  |  | June 15, 2015 | 2015 | Abbott Point Of Care Inc. | Certain lots of PT/INR cartridges may generate a higher than typical number of PT/INR star outs potentially delaying the generation of patient results for PT/INR. | Under Investigation by firm | Abbott Point of Care sent a recall letter dated June 15, 2015 via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the enclosed Business Reply Card (BRC). A Proof of Destruction From (FG-4533 in included with letter and is required to be returned to APOC as part of the credit process. Customers with questions were instructed to contact APOC Technical Support at 800-366-8020, Option 1 or their APOC representative. |  | 423168 | 423,168 cartridges | Worldwide Distribution - US (nationwide) and Internationally to China, Lebanon, Costa Rica, Singapore, Brazil, Jordan, Russia, Greece, Austria, India, Nigeria, United Kingdom, Korea, Saudi Arabia, Morocco, Slovenia, Portugal, Tunisia, Kazakhstn, Qatar, U.A.E. Israel, Taiwan, Mexico, Argentina, Poland, Italy, Hong Kong, Kuwait, Switzerland, Turkey, Ireland, Chile, Germany, and Australia. | Terminated | February 19, 2016 | 249 | 423168 | Not_Computer
Z-2364-2015 | 71674 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Perfectum Tuberculin Glass Syringe, Non sterile |  |  | July 09, 2015 | 2015 | Cadence Inc. | Graduation marks of the syringe are not centered with the numeric position marks on the syringe barrel. | Device Design | Cadence issued the recall letter, specific to the customer dated 7/9/15 via registered mail on July 17, 2015. Accounts are requested to Recalled products must not be used.2. Locate and quarantine all affected products.3. Contact Cadence Science, Inc. using the Response Form provided.4. If this information is received by a dealer, wholesaler, or distributor/reseller that distributed any of these products to other facilities, then notify their customers of this action so that those customers can locate and remove all affected products. Any questions contact: Phone: 401-942-1031 |  | 766 | 766 units | NationwideForeign Canada, Denmark. Italy | Open, Classified |  |  | 2431 | Not_Computer
Z-2363-2015 | 71445 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ReUnion TSA Shoulder System, Standard SR Humeral Head Trial, 48X18, Catalog number 5901-S-4818. Orthopedic manual surgical instrument. |  |  | May 12, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | A package labeled as Part #5901-S-4818, Lot MAC7C14, standard humeral head trial, contained part # 5901-E-4818 Lot MAC7C14, offset humeral head trial. | Process control | Stryker Orthopaedics notified their branches/agencies via email on 5/12/2015 and sent notification letters and product accountability forms to hospital risk management/branches on 7/15/2015 and 7/16/2015 |  | 21 | 21 units | US Nationwide | Open, Classified |  |  | 21 | Not_Computer
Z-2359-2015 | 71729 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 8.5mm Reamer Head for Flexible Intramedullary Nailing System; An orthopedic manual surgical instrument is a nonpowered hand-held device intended for medical purposes to manipulate tissue or for use with other devices in orthopedic surgery. This generic type of device includes the cerclage applier, awl, bender, drill brace, broach, burr, corkscrew, countersink, pin crimper, wire cutter, prosthesis driver, extractor, file, fork, needle holder, impactor, bending or contouring instrument, compression instrument, passer, socket positioner, probe, femoral neck punch, socket pusher, reamer, rongeur, scissors, screwdriver, bone skid, staple driver, bone screw starter, surgical stripper, tamp, bone tap, trephine, wire twister, and wrench. |  |  | July 01, 2015 | 2015 | Synthes (USA) Products LLC | The possibility exists for intraoperative reamer head breakages which could also allow for un-retrieved fragments of non-implant grade material: additional lot numbers identified. No new adverse events reported. | Other | An urgent notice of medical device recall, dated July 1, 2015, was sent to end users and sales consultants to alert them about the device defect and potential risk to patients. Recipients were asked to complete the verification form and return it along with any affected devices their posession. |  | 4329 | 4329 | Worldwide Distribution. | Open, Classified |  |  | 4329 | Not_Computer
Z-2373-2015 | 71724 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte Liquid Bicarbonate Concentrate 6.4 liter bottle |  |  | July 16, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | Expansion of recall to lots manufactured in 2015 due to the potential for bacterial contamination. | Process control | Fresenius Medical NA issued an Urgent Expanded Medical Device Recall letter on 7/16/15 to clinics and patients. The letter identified the affected product, explained the reason for the expansion, and instructed users to immediately examine their stock to determine whether there was any Naturalyte Liquid Bicarbonate Concentrate beginning with 15AMLB, 15BMLB, 15CMLB.Users are to discontinue use immediately if any product of these lots is found, .All units are to be placed in a secure, segregated area.If affected product was on the machine prior to patient treatment, a [Heat Disinfect] program should be performed.The attached Reply Form should be completed and returned.Contact your FMCNA Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product. |  | 2264028 | 2,264,028 | Nationwide Distribution | Open, Classified |  |  | 2264028 | Not_Computer
Z-2320-2015 | 69849 | No | RER |  |  |  |  | GE Optional Laser Centering Device X-ray generator. |  |  | November 10, 2014 | 2014 | GE Inspection Technologies, LP | GE has identified a potential failure to comply with 21 C.F.R. Âż 1040.10 and 21 C.F.R. Âż 1010.2 for the alignment laser centering device. | Radiation Control for Health and Safety Act | GE's Planned Action for repair: 1. Replacement compliant laser centering devices will be shipped to purchasers, with prepaid shipping envelopes for purchasers to return the noncompliant units to GE.2.All mailings will be made at no cost to the purchaser; and3.The corrections will be completed by November 1, 2015Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. This office and the Food and Drug Administration (FDA) district office coordinator are to be included in the notification. For further questions please call (717) 447-1278. |  | 25 | 25 | US Distribution to the states of : NY, CA, WA, PA, MD, AL, FL and LA. | Open, Classified |  |  | 25 | Not_Computer
Z-2376-2015 | 71658 | No | RCE |  |  |  |  | PSW 1203.U/1, PSW 1301.U, PSW 1307.U, PSW 1403.U, and PSW 1501.U Pacemaker Programmer software. Software approved for use with BIOTRONIK device programmers (trade names: Renamic and ICS 3000). |  |  | June 30, 2015 | 2015 | BIOTRONIK, Inc. | Ventricular packing: LV software programming versions for BIOTRONIK CRT-P and CRT-D devices are recalled because they are not approved for use in the US. | Software Manufacturing/Software Deployment | BIOTRONIK sent the Important Product Information letter and the Acknowledgement of Important Product Information Form, dated June 20, 2015, to affected physicians with the following recommendations: -Review programming settings for your patients with CRT-P or CRT-D devices.-If you have a patient with a device programmed to Ventricular pacing: LV, arrange to have the patients device reprogrammed to an appropriate approved mode (i.e. "Ventricular pacing: RV" or "Ventricular pacing: BiV").-Physicians are requested to sign the Acknowledgement of Important Product Information Form and immediately fax to 800-913-6993 or email to UScompliance@biotronik.com.BIOTRONIK'S representative will update your programmer with approved 1502.U programmer software. Following this update, the Ventricular pacing: LV feature will no longer be available. However, a device previously programmed with"Ventricular pacing: LV will remain programmed in this mode unless and until the pacing mode setting is changed. If you have any questions concerning this Important Product Information, please contact the BIOTRONIK Technical Services Department:Telephone: 1-(800) 284-6689 (24-hour service) |  | 768 | 768 units (software) | Nationwide Distribution. | Terminated | December 24, 2015 | 177 | 768 | Computer
Z-2375-2015 | 71642 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Wall Chart for STERRAD 100NX System, P/N 10104. |  |  | July 17, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization PRoducts (ASP) is recalling the Wall Chart for STERRAD 100NX System because it incorrectly states "APTIMAX Instrument Trays can be enclosed in either polypropylene wrap or in Tyvek Pouches with STERRAD Chemical Indicators." | Under Investigation by firm | A Urgent recall letter will be sent to customers to inform them that Advanced Sterilization Products (ASP) is recalling the Wall Chart of the STERRAD 100NX System because it incorrectly states "APTIMAX Instrument Trays can be enclosed in either polypropylene wrap or in Tyvek Pouches with STERRAD Chemical Indicators. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete a business reply card for their acknowledgement of the customer letter and documentation of the removal and replacement of the wall chart. Customers are instructed to complete and return the attached Business Reply Card via mail or fax to Stericycle: 877-497-3125. Customers with questions are instructed to contact SteriCycle at 866-737-5389 and reference event #6677. |  | 1465 | 1465 units | Nationwide Distribution. | Open, Classified |  |  | 1465 | Not_Computer
Z-2381-2015 | 71707 | adaptor, hygiene | ILS | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Etac Relax wall mounted shower seat |  |  | July 12, 2015 | 2015 | Etac Supply Center Ab | The recalling firm has received reports of a malfunction of the Relax wall mounted shower seat where the wall bracket has cracked or broken. | Process design | Etac sent an Urgent Field Safety Notice dated July 2015 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that there was no reason for an immediate stop of using Relax, however, as an intermediate precautionary action awaiting replacement, consumers were requested to be observant and inspect the Relax to see if there are visible signs of a crack. Customers with questions were advised to contact their Etac distributor or the Etac coordinator at 1-800-336-7684 or fax 1-704-882-0751. |  | 13 | 13 units | Nationwide Distribution including FL, AZ, and CA. | Open, Classified |  |  | 13 | Not_Computer
Z-2382-2015 | 71666 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Celsite Implantable Access Port System, model ST301. The Celsite Implantable Access Port Systems (Celsite port systems) are implantable port and catheter systems that allow safe, repeated access to the patient's bloodstream. The port chamber and catheter design can be used for the administration of medication and fluids. The Celsite system consists of an access port with a silicone septum, which is connected to a catheter using a connection ring. The triangular shaped access port has a low profile nose, finger stops on the side of the housing, and a round base. Celsite access ports have suture holes or suture zones to secure placement during implantation. |  |  | June 16, 2015 | 2015 | B. Braun Interventional Systems | The manufacturer, B. Braun medical France, received endotoxin test results that are out of specification for the peelable sheath (A1537). | Unknown/Undetermined by firm | An Urgent Medical Device Recall letter, dated 6/16/2015 was sent to the 2 consignees via express mail. The letter explained the issue and requested the hospital review their inventory for the affected model and lot. If any quantities of the lot remained in inventory, the product was to be returned to BIS. A BIS sales representative will personally visit each account and complete an inventory sheet. Customers with questions can contact Paul O-Connell, President at 1-847-274-0097 |  | 11 | 11 units | CA and NY only. | Terminated | November 20, 2015 | 157 | 11 | Not_Computer
Z-2380-2015 | 71849 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioJet AVX Thrombectomy Set: UPN 105039-001, Catalog No.105039, Sterilized Using Ethylene Oxide. For breaking apart and removing thrombus from vascular system, used with the AngioJet Ultra System in breaking apart and removing thrombus from A-V access conduits. |  |  | July 31, 2015 | 2015 | Boston Scientific Corporation | AngioJet products were distributed with an incorrect use-by (expiry) date. The correct date is Nov-2015, but the units are labeled with a use-by date of Dec-2015. | Process control | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated July 31, 2015, to all affected consignees. The letter described the problem and the affected product. Consignees were advised to segregate the affected product and return it to Boston Scientific. They were also requested consignees to complete and return the Account Reply Verification Tracking Form. Contact your Sales Representative for any questions.For questions regarding this recall call 763-494-1556. |  | 9 | total of 9 units | Worldwide Distribution - US including AZ, MO, OH and Internationally to GERMANY and SPAIN. | Terminated | February 05, 2016 | 189 | 9 | Not_Computer
Z-2385-2015 | 71722 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 Hydrophilic-coated Female Intermittent Urinary Catheters (6"), 12 Fr, (30/box), Product Code 51612. |  |  | July 07, 2015 | 2015 | C.R. Bard, Inc. | Potential breach of the sterile barrier packaging. | Unknown/Undetermined by firm | A recall notice was sent to consignees on 7/7/2015 asking them to discontinue use of the product. The letter to the distributors requested a sub-recall. Both letters included a response form to be returned to Bard. |  | 91140 | 91,140 each | Nationwide Distribution-including AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI AND DC. | Terminated | July 21, 2017 | 745 | 694907 | Not_Computer
Z-2390-2015 | 71714 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | RadPro 40kW Digital X-Ray Panel and Mobile/Portable Diagnostic X-Ray. |  |  | July 13, 2015 | 2015 | Sedecal S.A. | Electrostatic energy may be stored in Systems that can cause a short-cicuit which, in turn, can generate an Electomagnetic Pulse which might disturb the electronics of the Digital Motion Control Board producing uncontrolled movements and such uncontrolled movements could cause the system to bump into the user or ppatient. | Device Design | FIELD CORRECTION NOTIFICATION letters dated July 13, 2015 were sent to all dealers. The letters instructed the dealers: to schedule the replacement of the current arrestor by a new design of resistive arrestor with a resistance of 1 M ohm between frame and floor (ground). Although the metallic arrestor is only implemented in the equipments manufactured from April 2013 till March 2015, the upgrade kit has to be implemented in all the installed base of Mobiles with the same configuration in order to ensure that even the arrestors installed as spare part are replaced. Sedecal will provide the necessary upgrade kits to all the customers or distributors. *** U.S. Distributor, Virtual Imaging (a Canon U.S.A. Company) will use letter entitled, FIELD CORRECTION NOTIFICATION, RADPRO 40kW - -DIGITAL MOBILE X-RAY SYSTEM and dated 7/13/2015 to communicate to end users. The letter states "Once Virtual Imaging receives the resistivearrestor, we will contact you to schedule the replacement on the Affected System that you service."Please contact Virtual Imaging, Inc.'s Technical Support for any questions at 561.893.8500 or technicalsupport@vifla.com. **** |  | 361 | 361 systems (U.S.); 5 systems (Foreign) | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 361 | Not_Computer
Z-2362-2015 | 71655 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Syringe Pump, Model No. 8110. Delivers fluids. |  |  | July 21, 2015 | 2015 | CareFusion 303, Inc. | Channel Error code is displayed on the PC unit with an audio and visual alarm, and on the syringe module. After the error is cleared on the PCU, the syringe pump is unresponsive to key presses until the next power cycle, or the module is detached and reattached. | Process control | An Urgent Medical Device Recall Notification letter was sent on 7/20/15 to customers who purchased the Alaris Syringe Module (Model 8110) to inform them that CareFusion has identified an issue with the Alaris Syringe module manufactured between March 2014 through September 2014. The letter informs customers of the problems identified and the actions to be taken. CareFusion will replace the syringe drive train assembly on the affected units and adjust, if required. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation. Customers are instructed to contact CareFusion Support Center (888-562-6018, 7am-4pm PT, M-F, supportcenter@carefusion.com) for recall related questions. Customers are instructed to contact Customer Advocacy (888-812-3266, 24hrs/day, 7days/wk, customerfeedback@carefusion.com) for adverse event reports. Customer are instructed to contact Technical Support (888-812-3229, 6am-5pm PT, M-F, dl-us-inf-techsupport@carefusion.com) for technical questions. Customers are instructed to complete and return the enclosed customer response card to expedite the corrective action process.A urgent recall letter for the expansion of the affected units was sent on 10/15/15 to inform customers of the expanded recall. The letter informs customers that the firm has identified an issue with the Alaris Syringe module model 8110 after receiving reports of customers experiencing an channel error on the Alaris Syringe module Model 8110. The letter informs customers that the recall is an expansion of affected units outlined in the CareFusion recall previously issued on 7/20/15. The letter provides the customers with the affected units, issue, potential risk, and actions to be taken.The corrective action for the expanded serial numbers will not change. The affected Alaris Syringe modules will be remediated with a hardware inspection and an adjustment, if required. CareFusion will contact all affected c |  | 68460 | 68,460 units | Worldwide Distribution-US (nationwide) and the countries of Australia, New Zealand, South Africa, Canada, Middle East, Malaysia, and Taiwan. | Open, Classified |  |  | 68460 | Not_Computer
Z-2206-2015 | 71696 | No | RCG |  |  |  |  | American Science &amp; Engineering Inc,Gemini Cabinet X-ray system. |  |  | May 22, 2015 | 2015 | American Science &amp; Engineering, Inc. | During factory testing, we discovered that Gemini systems could, in rare circumstances, experience a fault that leads to a state in which X-rays are being produced but the X-ray indicator lights (located on the operator's console and on both ends of the tunnel) are not illuminated. | Vendor change control | American Science Engineering, Inc (AS&E;) is developing a fix for this problem that includes updating a software configuration on the X-ray source. You will be notified by field service as soon as this update is available.Please inform all operators of the system about this communication, so they are aware of the need to shut down and re-start the system should this fault occur.We appreciate your understanding. if you have any questions, please feel free to contact AS&E; via e-mail at service@as-e.com. If you are under a service contract, you can also contact your customer service representative. For further questions please call (978) 262-8700. |  | 125 | 125 systems | Worldwide Distribution - US Distribution to the states of : NY, FL, TN, PA, NJ, DE, ME, TX, CT, MS, NC, MD, MO, VA, DC, NV and GA., and to the countries of : Saudi Arabia, Indonesia, Canada, Netherlands, Hong Kong, Lebanon and Kazakhstan. | Open, Classified |  |  | 125 | Computer
Z-2343-2015 | 71702 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | PS500 of an IACS Workstation Critical Care (Evita Infinity V500); intended for the ventilation of adult, pediatric and neonatal patients. |  |  | July 13, 2015 | 2015 | Draeger Medical, Inc. | The battery capacity of the optional PS500 of an IACS Workstation Critical Care (Evita Infinity V500) did not last as long as expected. | Nonconforming Material/Component | Draeger Medical, Inc. sent an Urgent Medical Device Recall Letter, dated July 2015, to end users to inform them of the issue and provide interim instructions to avoid the associated risk to patients. All PS500 batteries will be exchanged free of charge. |  | 2081 | 2081 | Nationwide Distribution. | Terminated | March 01, 2017 | 597 | 2081 | Not_Computer
Z-2329-2015 | 71675 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, Universal (with neonatal functionality enabled) and Neonatal Models. |  |  | July 16, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports in which tidal volumes reaching patients were lower than set tidal volumes in neonatal Volume Control Plus (VC+) Mode with active humidification. This situation may potentially lead to respiratory compromise if not recognized. | Software design | Consignees will be notified via letter on July 17, 2015. No products will be returned, the reconfiguration of affected ventilators will occur at customer sites. |  | 657 | 657 units | Worldwide Distribution-US (nationwide) and the countries of Canada, United Arab Emirates, Argentina, Australia, Bahrain, Brazil, Chile, Colombia, Costa Rica, Egypt, Ireland, Israel, Jordan, Japan, Kuwait, Lebanon, Morocco, Mexico, Panama, Peru, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Turkey, and Uruguay. | Open, Classified |  |  | 657 | Not_Computer
Z-2402-2015 | 71496 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Winco Inverness 24-Hour Treatment Recliner - with Nylon Casters |  |  | April 02, 2015 | 2015 | Winco Mfg., LLC | Shipment of foam for chairs was received classified as Fire Retardant but was in fact standard foam and would not meet Fire Retardant standards. | Nonconforming Material/Component | Winco sent a field correction letter dated April 2, 2015, to all affected customers The letter identified the product, the problem, and the action to be taken by the customer. Winco instructed customers that they will replace the recliners at no charge to the customer. Winco will provide detailed instructions for the replacement of the seats by their facility staff. Winco will extend the remaining warranty on the entire chair for a period of 6 months as compensation. If product was further distributed customers were instructed to notify those customers of the field correction. Customers were asked to complete and return the enclosed field correction response form as soon as possible. Customers with questions should contact Customer Care Monday-Friday 8:30 AM-5:00 PM EST at 800-237-3377 or email QA-RA@wincomfg.com.For questions regarding this recall call 352-854-2929. |  | 29 | 29 chairs | Nationwide Distribution including CA, FL, KS, KY, LA, MA, MD, MN, MS, NJ, NY, OK, TN, TX, and WA. | Open, Classified |  |  | 133 | Not_Computer
Z-2389-2015 | 71678 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon X3 Tibial Bearing Insert - CR |  |  | May 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker has received a report which identified that a Triathlon Tibial Bearing Insert CS X3 Size #2 9mm (5531-G-209) was reported to be in a pack that was labelled as a CR insert (5530-G-209). | Labeling mix-ups | Stryker Orthopaedics sent a recall notification via email on 5/29/2015 to their local branches. Notification letters and Product Accountability Forms were sent to Stryker branches/agencies with return receipts on 7/2/2015. |  | 10 | 10 units | OH, SC, RI, WV | Open, Classified |  |  | 10 | Not_Computer
Z-2405-2015 | 70741 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RenLane Renal Denervation Catheter; Cat No.: D135601. |  |  | April 15, 2014 | 2014 | Cordis Corporation | Potential Damage may occur to the helical tip of the RENLANE Renal Denervation Catheter, if the hemostasis valve of the arterial access device is not fully open before insertion or withdrawal of the catheter. | Labeling Change Control | Cordis sent an Urgent Field Safety Notice dated April 15, 2014 providing additional labeling to their consignees in Germany only, with a follow-up letter on August 4, 2014. Cordis explained When deploying the catheter via a compatible arterial access device, if a Tuohy-Borst type adjustable hemostasis valve is used, ensure the valve is fully open before insertion or withdrawal of the catheter. This is especially important, since the helical tip at the distal end of the RENLANE Renal Denervation Catheter has a larger diameter compared to the proximal catheter shaft. Although the event was not a Removal, all lots involved have expired. The product is no longer commercialized, and was never commercialized beyond the 10 customers in Germany. |  | 98 | 98 units | Distributed in Germany only. | Terminated | August 26, 2015 | 498 | 98 | Not_Computer
Z-2407-2015 | 71727 | Yes | BTS | Bronchial tube. | Anesthesiology | Anesthesiology | 510(k) | Bronchial Double Lumen Tube Set (Left), Sterile |  |  | November 21, 2014 | 2014 | Teleflex, Inc. | Customer complaints reporting that the cobb connector detached from the main connector prior to use. | Other | Recall letters were sent on 11/21/2014 to consignees requesting them to discontinue use and quarantine the products. The letter includes a reply form to be returned. The letter to the distributors requested a sub-recall, and included a recall letter to be sent to their accounts. |  | 143750 | 6,991 ea. - US &amp; 136,759 ea. - INTL | AL, AK, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV. WI, DC, Puerto Rico, Canada, Bolivia, Ecuador, Hong Kong, Malaysia, Philippines, Taiwan, Trinidad &amp; Tobago | Terminated | August 15, 2017 | 998 | 143750 | Not_Computer
Z-2404-2015 | 71762 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems; dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. |  |  | June 30, 2015 | 2015 | Siemens Medical Solutions USA, Inc | in case a system error occurs and the system enters the "Bypass Fluoro" mode while the X-ray locking function is active, the only way to exit the X-ray locking function, would be to either resolve the root cause of the system being in "Bypass Fluoro" or to restart the system. | Software design | A customer advisory notice, dated June 30, 2015, was sent to consignees informing them of the corrective action that would be performed to prevent the possible loss of the image generation function. Siemens plans to contact end users to install a software patch. For questions regarding this recall call 610-219-6300. |  | 586 | 586 | Nationwide Distribution | Open, Classified |  |  | 586 | Computer
Z-2406-2015 | 71673 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | Cell Search Circulating Tumor Kit. Intended for the enumeration of circulating tumor cells (CTC) of epithelial origin in whole blood. |  |  | May 15, 2015 | 2015 | Veridex, LLC | Complaints of the presence of artifacts that appears as small bead like structures in the image gallery when processing both control cells and patient samples as well as the failure of high control cells. | Under Investigation by firm | Janssen Diagnostics, LLC sent a urgent product correction notification/customer confirmation of receipt on 5/15/2015 via Fed Ex. |  | 1560 | 1560 kits | Nationwide Distribution. | Terminated | March 30, 2016 | 320 | 1560 | Not_Computer
Z-2422-2015 | 71911 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | IBC Suction Wand Catalog Number 1990S |  |  | August 07, 2015 | 2015 | International Biophysics Corp. | It was discovered that the heat seal placed onto sterile pouches may not meet strength acceptance criteria. | Packaging process control | The recalling firm began notifying their affacted customers on 7/2/2015 |  | 100 | 100 units | Worldwide - United States Nationwide and the countries of: Germany, Norway, Turkey, Chile, Pakistan, and Republic of South Africa | Terminated | March 01, 2016 | 207 | 11201 | Not_Computer
Z-2416-2015 | 71710 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | ADVIA Chemistry Acetaminophen reagents; for in vitro diagnostic use in the quantitative determination of acetaminophen in human serum and plasma (lithium heparin) on ADVIA Chemistry systems. |  |  | June 30, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | A change in the concentration of N-acetylcysteine (NAC) that may cause interference for the ADVIAÂż Chemistry Systems Acetaminophen assay for certain lots. | Unknown/Undetermined by firm | Siemens sent an Urgent Medical Device Correction letter dated June 30, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer. Siemens Healthcare Diagnostics is conducting a voluntary recall for the ADVIA Chemistry Acetaminophen reagent, for all in date and future kit lots. Customers were informed of the revised interference levels of N-acetyl cysteine (NAC) of 200 mg/dL, for the ADVIA Chemistry Acetaminophen assay.Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days.Field service personnel will be sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For further questions call 1 (800) 441-9250. |  | 1032 | 1032 | Worldwide Distribution - US (nationwide) and to the countries of : Canada, Netherlands, Great Britain, Norway, Sweden, Spain, Italy, Germany, France, Finland, Belgium, Portugal, Czech Republic, Israel, Poland, Turkey and Bahrain. | Open, Classified |  |  | 1032 | Not_Computer
Z-2421-2015 | 71783 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 2.0/2.4 Cannulated Screw Countersink Part # 210-24-002 NON-STERILE and 3.0/4.0 Cannulated Screw Countersink Part # 210-40-002 NON-STERILE |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | The cutting efficiency of the Countersinks was out of specification. | Nonconforming Material/Component | Trilliant sent an Important Medical Device Advisory Notice letter dated February 15, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers of the option to exchange affected instruments at no charge. To initiate the exchange customers were instructed to contact your local Trilliant Surgical Sales Representative, contact the Customer Service Department directly at 1-800-495-2919, or via email at orders@trilliantsurgical.com. |  | 102 | 102 units | US Nationwide Distribution in the states of: NY, CT, TN, TX, OH, IN, LA, NC, MO, MN, GA, FL, SC, WA, WI, IA, PA, CA, ND, OK, NJ, MI and MD. | Open, Classified |  |  | 102 | Not_Computer
Z-2418-2015 | 71869 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Sciatic Nerve Retractor and Sciatic Nerve Retractor Long; Synthes Sciatic Nerve Retractor and Sciatic Nerve Retractor Long are system specific instruments to the 3.5 mm Low Profile Pelvic System Retractor Set. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | Potential for micropores to form, allowing fluid to enter the hollow handle. | Process control | Urgent recall letters, dated July 29,2015, were sent to end users and sales consultants to inform them of the issue, potential risks, and to provide instructions for returning affected devices and the response form. |  | 1113 | 1113 | Nationwide and internationally, including Canada | Open, Classified |  |  | 1113 | Not_Computer
Z-2417-2015 | 71739 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | Radiometer Fixation Kit N20 for the TCM CombiM monitoring system, REF 905-873. The Fixation Kit N20 is used to apply tc Sensor 84 and tc Sensor 54 on the skin of a patient. TC sensor 84 and 54 are used with the TCM CombiM monitoring system and the Philips TcG10 module. The TCM CombiM monitoring system is intended for continuous transcutaneous monitoring of carbon dioxide (tcpCO2) and oxygen (tcpO2) partial pressures. |  |  | July 16, 2015 | 2015 | Radiometer America Inc | When the fixation ring is taken off the paper liner there is a potential for fractions of the paper liner's surface coating to stick to the surface of the ring. | Under Investigation by firm | Radiometer sent a Product Notification letter dated July 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.In the letter the firm requests that inventory be checked and removed for any the identified lot numbers for Fixation kit N20. The firm requests that the products be discarded. The letter states to complete p.2 and return it to their Radiometer representative. If you have any questions contact Radiometer America Technical Support at 1-800-736-0600 opt.4. |  | 4244 | 4244 kits | Worldwide Distribution - US Distribution and to the countries of : Canada, Spain, France, Great Britain, Philippines, Netherlands, Japan, China, Australia, Denmark, South Africa, Germany, Qatar, Turkey, Hong Kong, Costa Rica, Austria, Finland, Indonesia, Norway, Malaysia, Belgium, Czech Republic, Hungry, United Arab Emirates, Poland, Colombia, Belgium and Malta. | Terminated | January 29, 2016 | 197 | 4244 | Not_Computer
Z-2420-2015 | 71897 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic EOPA Arterial Cannula, Model # 77620. Lot: 2015060544 intended to deliver oxygenated blood to the heart during cardiopulmonary bypass procedures for up to 6 hours in duration. |  |  | July 24, 2015 | 2015 | Medtronic Perfusion Systems | EOPA Arterial Cannula Devices in this lot were shipped without the guidewire. | Packaging process control | On July 24, 2015 Medtronic sales representatives began verbally notifying the customers of the issue, requesting quarantine and return of un-used, affected productConsignees were sent a Medtronic "Urgent Medical device Recall" letter dated August 2015. The letter was addressed to Risk Manger and Health Care Professional. The letter described the problem and the product involved in the recall. Advised consignees to quarantine and remove the product from their inventory and return it to Medtronic. Requested consignees to complete and return the Customer Confirmation Certificate via email to RS.CFQFCA@Medtronic com or fax to Medtronic 651-367-0612 to the attention of Customer Focused Quality. For questions contact your Medtronic Sales Representative or Lifeline Technical Services at 1-877-526-7890. |  | 221 | 221 (111 US, 110 OUS) | Worldwide distribution. US nationwide including: FL, GA, IL, MI, NY, NC, PA, WV; Germany, Hong Kong, and Italy. | Terminated | April 04, 2016 | 255 | 221 | Not_Computer
Z-2425-2015 | 71713 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Tibial Nail Wrap Holds 11 Nails and Synthes Intramedullary Nail Wrap Holds 10 Nails; the Nail Wrap is intended to be used as a protective sheath specifically for Universal Femoral Nails, Universal Tibial Nails, Unreamed Tibial Nail, and the Stainless Steel Tibial Nail during steam sterilization by the hospital. |  |  | July 01, 2015 | 2015 | Synthes (USA) Products LLC | The wrap is for single use for nail sterilization but does not have sufficient testing for the device to be considered a multi-use item. | Under Investigation by firm | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated July 8, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.ACTIONS REQUIREDWe have on record that your facility has received the product(s) subject to this recall. DePuy Synthes asks that you review your inventory and immediately remove the affected lots from your stock.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note theReturn Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided.Return the Verification Section (page 3 of this letter) with the product to Credit/Returns, DePuy Synthes,Send a copy of the completed Verification Section to DePuy Synthes by: Fax: 888-670-4162 or Scan/email: Synthes5485@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information.Return the documents to DePuy Synthes by: Fax: 888-670-4162 orScan/email: Synthes5485@stericycle.comNote: If the Verification Section is answered on |  | 1487 | 1487 | Worldwide Distribution - US Distribution including Puerto Rico and to the countries of : Canada, Switzerland and New Zealand. | Open, Classified |  |  | 1487 | Not_Computer
Z-2433-2015 | 71720 | No | LWK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Medline Electrical Clipper with Charging base |  |  | July 08, 2015 | 2015 | Medline Industries Inc | Charging base of surgical clippers overheats and smokes. | Process control | Medline Industries sent a " Urgent Remedial Action Destruction Form " letter via first class dated July 9, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.***URGENT RECALL IMMEDIATE ACTION REQUIRED SURGICAL CLIPPER CHARGER BASE DYND70800 ***Medline Industries, Inc. is recalling specific lots of the Medline Surgical Clippers Charger Base (base only). This recall is being conducted due to a potential defective component on the circuit board causing the charger base to overheat which could result in a charger base malfunction. These lots were distributed October 2013 through March 2015. After further investigation we have determined the root cause to be a defective circuit board. This defect has since been corrected. REQUIRED ACTION: 1. Immediately check your stock for item number DYND70800 and the affected lots listed on the destruction form, dispose of the affected charger base only (do not dispose of the clipper). 2. Please return the completed enclosed destruction form listing the quantity of affected charger base(s) you destroyed. Even if you do not have any affected product please complete and return the form, as indicated on the destruction form. When we receive your completed destruction form we will send you replacement charger base(s). 3. If you are a distributor, promptly notify any of your customers that may have received affected product. If you have any questions, please contact 866-359-1704. As the distributor of this product, we sincerely regret the inconvenience. We, like you, place the health and safety of your patients first and foremost. This recall is being made with the knowledge of the U.S. Food and Drug Administration. |  | 7062 | 7062 units (U.S.); 540 units Foreign | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the countries of : United Arab Emirates, Barbados, Bahamas, Chile, Costa Rica, Italy, Lebanon, and Singapore. | Open, Classified |  |  | 7062 | Not_Computer
Z-2431-2015 | 71841 | No | LSL | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | Xpert CT /NG Vaginal/Endocervical/ Specimen Collection Kit: |  |  | July 27, 2015 | 2015 | Cepheid | Firm determined that some lots of CT/NG Swab 50 and CT/NG Urine collection kits contain a sub component with an expiry date that is earlier than the kit expiration date. | Incorrect or no expiration date | Notification letters dated July 28 and 29th to all affected customers via Federal Express.. |  | 776 | 729 kits US, 47 kits ROW. | Worldwide distribution. US nationwide, including North Mariana Islands, Aruba, Canada, Chile, Colombia, Costa Rica, Cyprus, Guam, Guatemala, Italy, Malawi, Micronesia Federated States, Panama, San Marino, South Africa, Spain, Switzerland, and Zimbabwe. | Terminated | September 28, 2015 | 63 | 2607 | Not_Computer
Z-2434-2015 | 71633 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer SST Plus Blood Collection Tubes, 5.0 mL, 13 x 100 mm, Sterile. |  |  | June 23, 2015 | 2015 | Becton Dickinson &amp; Company | Some of the tubes were manufactured with stoppers that did not meet current manufacturing specifications. Stoppers were manufactured with a reduced cycle cure time. | Process control | Becton Dickinson notified their affected customers of this recall by sending them an Urgent Product Recall Letter dated June 23, 2015 via UPS 2nd Day. |  | 238 | 238, 000 (Lot #5133959) and 243,000 (Lot #5133957) | Distributed in FL, IL, MA, MI, MO, PA and TN. | Terminated | April 06, 2016 | 288 | 238 | Not_Computer
Z-2371-2015 | 71543 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Medley LVP Frame Membrane. |  |  | June 03, 2015 | 2015 | Elite Biomedical Solutions LLC | Administration of inappropriate quantities of fluid can result, with the potential to cause injury or death. | Other | Elite Biomedical Solutions, sent a "Product Advisory Notices" letter dated May 21, 2015 to their customers. On June 3, 2015 the firm sent an Urgent Medical Device Recall letter to their customers. The letter identified the affected product , problem and actions to be taken. And on June 12, 2015 a press release was issued via ECRI (Emergency Care Research Institute) to all hospitals in the US. The notification letter requested customers to: 1) inspect and quarantine affected products, 2) identify your customers and notify them at once of this product recall, and 3) complete and return the enclosed response forms. For any questions, call Elite Biomedical Solutions, LLC, Quality Manager, at 1-855-291-6701. |  | 609 | 609 pieces | Nationwide Distribution in the states of AZ, CA, GA, IL, IN, KS, KY, MA, MD, MN, MO, MS, NC, NY, OH, PA, SC, TN, TX, VA &amp; WI. | Terminated | March 07, 2017 | 643 | 955 | Not_Computer
Z-2370-2015 | 71360 | Yes | CBI | Tracheal/bronchial differential ventilation tube. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCI, Sheridan, SHER-I-SWIV and SHER-I-SWIV/FO Double Swivel, Rx Only, Sterile. |  |  | June 01, 2015 | 2015 | Teleflex Medical | The double swivel connector may crack or separate on the endobronchial tube. | Unknown/Undetermined by firm | Teleflex Medical sent an Urgent Medical Device Recall Notification letter dated June 1, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were directed to contact any customers who had received the product included within the scope of this recall. Customers were directed to immediately discontinue use and quarantine any products with the referenced catalog numbers and lot numbers listed in the notification. Additionally, customers were asked to return any affected product together with a completed Recall Acknowledgement Form, for consolidation and return to Teleflex Medical. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 188195 | 188,195 total units | Worldwide Distribution - US (nationwide) and the countries of: The Bahamas, Belgium, Argentina, Canada, China, India, Japan, South Korea, Mexico, New Zealand and Vietnam. | Open, Classified |  |  | 188195 | Not_Computer
Z-2437-2015 | 71867 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Intevo Series system, a radiological imaging system that is a combination of a single photon emission computed tomography (SPECT) camera system for nuclear medicine images, and a computed tomography (CT) camera system for x-ray images. |  |  | July 31, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Fuses with the improper amperage rating were installed in the Symbia S, Symbia T Series and Symbia lntevo systems. These fuses, in conjunction with a specific failure mode of the Power Controller assembly, could lead to the overheating of the assembly. | Component change control | Siemens has sent an undated Customer Advisory Notice letter to all direct accounts (customers). The letters include instructions for customers to: 1) Continue to use the system(s) while waiting for the field correction to be scheduled and performed unless the user observes a burning smell or sees evidenve of smoke. the user should engage the emergency stop, power down and discontinue use of the system, and contact the local service representative. 2) Ensure that this advisory is placed in the system's instructions for use. Siemens reminds the user to observe the patient during studies. 3) If the product was sold to another party, forward the letter to the party. A local service engineer will contact the customer to schedule the servicing of the affected units by September 2015 which will consist of replacing the affected fuse with the correct amperage rating fuse. If you have any questions regarding the important notice, please contact your local service representative by calling 1-800-888-7436. Europe, Middle East and Africa: +49 9131 940 4000. Asia and Australia: +86 (21) 3811 2121. |  | 185 | 185 | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and Internationally to Algeria, Argentina , Australia , Austria , Bangladesh , Belgium , Brazil , Bulgaria , Canada , Chile , China, Colombia , Costa Rica , Croatia , Czech Republic , Denmark , Ecuador , Egypt , Estonia, Finland , France , Germany , Hong Kong , Hungary , Iceland , India , Indonesia , Iran , Ireland , Israel , Italy , Japan , Jordan , Kazakhstan , Korea, Kuwait , Latvia , Libya , Lithuania , Luxembourg , Malaysia , Mexico , Morocco , Myanmar , Netherlands , New Zealand , Norway , Oman , Pakistan , Peru , Philippines , Poland , Portugal , Qatar , Romania , Russia, Saudi Arabia , Singapore , Slovakia , Slovenia , South Africa , Spain , Sweden , Switzerland , Taiwan , Tanzania, Thailand , Tunisia , Turkey , United Arab Emirates , Ukraine , United Kingdom , Venezuela , Vietnam | Open, Classified |  |  | 3068 | Not_Computer
Z-2427-2015 | 71676 | No | KXE | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | uHead, Ulnar Implant, Standard Stem, Size 1, Sterile, Rx Only. Intended for replacement of the distal radioulnar joint following ulnar head resection arthroplasty. |  |  | June 24, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker is recalling rHead, uHead, Sigmoid Notch, Remotion, Radio Capitellum because packaging integrity (sterile barrier) of the packaging type KIT I may potentially be compromised by transportation. | Packaging change control | Branches/Agencies were notified by Stryker of this action by e-mails on June 24, 2015 and an Urgent Medical Device Removal Letter/Acknowledgement Response Form dated June 25, 2015 was sent to the attention of the Risk Manager. |  | 16992 | 16992 total | Worldwide Distribution. | Open, Classified |  |  | 16992 | Not_Computer
Z-2441-2015 | 71780 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChemVELOCITY Automated Urine Chemistry Catalog No. 800-3049, 800-3050, 800-3079, 800-3080, 800-3530, 800-7167, 800-7190, 800-7713,800-7714, 800-7719, 800-7720. An in vitro diagnostic device used to automate the urine chemistry analysis profile using iChemVELOCITY Urine Chemistry Strips. |  |  | July 15, 2015 | 2015 | Iris Diagnostics | Iris International is recalling the iChemVELOCITY systems because they do not maintain the on-board strip stability claim of 5 days at 18-28 degrees celsius at 20-80% relative humidity (RH) as stated in the iChemVELOCITY product labeling. | Nonconforming Material/Component | An Urgent Medical Device Recall letter dated 7/15/15 was sent to all customers who purchased the iChemVELOCITY Automated Urine Chemistry System using iChemVELOCITY Urine Chemistry Strips REF 800-7212 and 800-7212-001. As a result of internal investigations following customer complaints, Iris International has determined that the iChemVELOCITY systems do not maintain the on-board strip stability claim of 5 days at 18-28ÂżC at 20-80% relative humidity (RH) as stated in the iChemVELOCITY product labeling. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions or concerns regarding the recall notice, are instructed to contact Beckman Coulter Customer Support at http://www.beckmancoulter.com or call (800) 526-7694 in US and Canada. |  | 2448 | 2,448 units | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bahrain, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Jordan, Republic of Korea, Kuwait, Lebanon, Luxembourg, Macao, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 2448 | Not_Computer
Z-2448-2015 | 71941 | No | HRM | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LX3 Floor Stand for Endure Luxite Envision Optical System; Endure LuxOR Q-Vue. |  |  | August 11, 2015 | 2015 | Alcon Research, Ltd. | Alcon is initiating a voluntary recall of all LX3 Microscope Floor Stands after receiving reports concerning loss of functionality due to issues with the cabling between the power, host and illumination modules, potentially causing the system to display a system message, and/or resulting in illumination failure. | Device Design | Alcon sent an Urgent Medical Device Correction letter dated August 11, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Alcon at 800-862-5266. |  | 688 | 688 units | Worldwide Distribution - US Nationwide and the countries Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Republic, France, Germany, Greece, Guatemala, Hungary, India, Italy, Japan, Jordan, Korea, Lebanon, Malaysia, Mexico, Mongolia, New Zealand/Fiji, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine and United Kingdom. | Open, Classified |  |  | 688 | Not_Computer
Z-2439-2015 | 71395 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PTW UNIDOS webline: 1) REF T10023, 2) REF T10022 and 3) REF T10023 Firmware version 2.05. Dosimeter for radiation therapy, diagnostic radiology and health physics. |  |  | May 27, 2015 | 2015 | PTW-Freiburg | In some cases, especially with high dose-per-pulse-rates typical for particle (specifically proton) therapy, measuring errors can result. Individual measurement intervals can be lost that are not visible to the user when there is a high dose-per-pulse rate resulting in a too-low dose being displayed on the monitor. This could result in a radiation over-exposure. | Software design | PTW issues Urgent Safety Notice due to an error with firmware version 2.05 of the UNIDOS webline in some cases, especially with the high dose-per-pulse-rates typical for particle therapy, measuring errors can result. |  | 242 | 242 firmwares | Distributed in CA, MI, MA, NC, NJ, NY, PA, TN, TX, PA &amp; WI. | Terminated | March 31, 2017 | 674 | 242 | Not_Computer
Z-2445-2015 | 71461 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | APEX INTERFACE Hemi 54X and Flared 56X, 3 Hole Acetabular Shells; Model: H5-12354, H5-11356. |  |  | June 04, 2015 | 2015 | Omnilife Science Inc. | The device may have improper screw hole placement due to inaccurate location of the index line. | Process design | Omni sent an Urgent Medical Device Field Correction letter dated June 4, 2015 via email to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review inventory for affected product and complete and return the attached Acknowledgement and Receipt Form by FAX to 508-819-3390. |  | 112 | 112 | US Nationwide Distribution in the states of IL, VA, CO, UT, NY, FL, MA, OK, CA, TN and WY. | Terminated | September 28, 2015 | 116 | 112 | Not_Computer
Z-2442-2015 | 71816 | No | MDS |  |  | Clinical Chemistry | PMA | Guardian Real-Time Monitor, Model No. CSS7100. Indicated for continuous or periodic monitoring of glucose levels in the fluid under the skin, in adults (18 years and older) with diabetes mellitus, for the purpose of improving diabetes management. |  |  | July 22, 2015 | 2015 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Guardian Real-Time Continuous Glucose Monitoring System because it has a language translation error that impacts the Finnish language. Specifically, the Monitor has a Finnish translation error in Predictive Alerts setting: the on-screen Finnish translation for LOW/HIGH is reversed as HIGH/LOW. | Software design | A Urgent Medical Device recall letter was sent to customers on 7/22/15 to inform them that the Guardian Real-Time Continuous Glucose Monitoring System System has a language translation error that impacts the Finnish language. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contact Nordic Support Line 0-800-164-064. |  | 254 | 254 units | Distributed in Finland. | Terminated | September 18, 2015 | 58 | 254 | Not_Computer
Z-2438-2015 | 71950 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Otto Bock Healthcare Products GmbH Genium X3 |  |  | August 12, 2015 | 2015 | Otto Bock Healthcare Product | Otto Bock Healthcare Products GmbH is recalling the new Genium X3 knee joint prosthesis due to issues with the resistance-control of the hydraulics, which may fail under conditions. | Process control | The consignee was notified via e-mail and telephone on 12 August 2015. The affected product was identified, as well as the reason for the recall. The firm will replace the affected device with a new one. A return authorization and label will be provided for the recalled product to be returned as soon as possible. |  | 1 | 1 | Distributed to the state OH. | Terminated | October 28, 2015 | 77 | 1 | Not_Computer
Z-2440-2015 | 71898 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.0.x |  |  | August 05, 2015 | 2015 | Brainlab AG | ExacTrac 6.0 Patient Positioning System:Display of potentially incorrect Digitally Reconstructed Radiograph (DRR) for x-ray correction andverification. | Software in the Use Environment | Brainlab intends to distribute the Field Safety Notice / Product Notification letter CAPA-20150713-001446, dated July 20, 2015, to the customers. The letter provides instructions for the user on how to address this issue, and to avoid a patient and treatment target position potentially different than intended due to this issue.The Field Safety Notice / Product Notification letter states that as generally required, always verify that the correct DRRs for the current patient are displayed in ExacTrac and an accurate fusion to the X-Ray images is possible for patient positioning. In case incorrect DRRs are displayed, exit and re-start the ExacTrac software. To minimize the probability of this issue occurring for the affected ExacTrac versions: -Reduce the amount of contours exported to ExacTrac as far as possible.-Avoid exporting large contours, such as couch top models, to ExacTrac.-After every (re-)start of the ExacTrac application on any workstation, always first load a phantom plan not containing contours. The Field Safety Notice / Product Notification letter states that as a general reminder the user should always follow the instructions and warnings as described in the user guide. In the context of this notification please specifically consider the safety notes relevant for X-ray correction and verification:-The image overlay functions must be used to verify fusion accuracy in both image views, especially if the images contain a series of similar structures such as vertebrae.-Do not perform patient treatment unless accurate image fusion is possible.-Large correction shifts (e.g. greater than 10 mm) may indicate incorrect automatic fusion. In such cases, verify the correction shift a second time.Brainlab Corrective Action:1. Existing potentially affected ExacTrac 6.0.x customers receive the product notification information.2. Brainlab will provide a software revision (ExacTrac v. 6.0.6) with this issue solved to affe |  | 361 | 361 systems | Worldwide Distribution - USA including the states of AL, AZ, AK, CA CO, CT, FL, GA, ID, IL IN, KY, LA, MD, MA, MI, MN, NE, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, and WI; and, the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Japan, Netherlands, Panama, Qatar, Russia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | January 05, 2017 | 519 | 361 | Computer
Z-2443-2015 | 71784 | No | MPO | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(k) | Covidien Argyle Infant Heel Warmers (no tab) |  |  | July 31, 2015 | 2015 | Covidien LLC | Infant Heel warmers may break during activation resulting in spillage of the contents on patients or healthcare worker causing potential skin/eye irritations | Package design/selection | Medtronic/Covidien issued an Urgent Medical Device recall letter dated July 31, 2015 via mailed Certified. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to quarantine and discontinue use of the affected products.. For questions or concerns,contact Medtronic representative or Customer Service at (800)-882-5878. |  | 68800 | 68,800 units | Worldwide Distribution - US Nationwide and the countries: Canada, Mexico, Panama, and Japan. | Open, Classified |  |  | 233549 | Not_Computer
Z-2447-2015 | 71786 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | cobas b 123 Fluid Pack COOX REF 05170036001 400 |  |  | July 14, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Low PO2 results. QC failures of the PO2 parameter, affecting primarily Levels 1 and 2, caused by a calibration issue with the PO2 parameter. This issue may not be detected since QC results can be below mean values, but still within 2 standard deviations (SD) limits. Potential for erroneously low PO2 results in patient samples, especially in blood samples with PO2 values below 50 mmHg | Nonconforming Material/Component | Roche sent an Urgent Medical Device Correction letter dated July 14, 2015 and the Fax Back form via UPS ground to the customers list. The letter identified the affected product, problem and actions to be taken. Consignees are instructed to discontinue use of affected Fluid Packs and discard them per local guidlines; and complete and return the enclosed faxback form and fax it to 1-888-943-5171 to receive replacement product for Fluid Packs from the affected lot. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272. |  | 82 | 82 Fluid Packs | US Nationwide Distribution in the states of LA, SC, CA, TX, NE, VA and WV including: Puerto Rico. | Terminated | January 19, 2016 | 189 | 82 | Not_Computer
Z-2449-2015 | 71848 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX-EVOLUTION X-Ray System |  |  | August 07, 2015 | 2015 | Carestream Health Inc | Software defect that impacts image alignment when using the automatic stitching option for long length images. The defect could cause stitching inconsistencies which may go undetected on a radiograph. | Software design | Carestream sent an Urgent Medical Device Correction letters dated July 29, 2015 and Customer Notification Acknowledgement Forms to customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. Carestream will update the software on your device to correct this issue. For questions or concerns, contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910, available 7 days per week on a 24 hour basis. Outside of the U.S., please call your local Service support number. |  | 236 | Domestic: 76 units, Foreign: 160 units | Worldwide Distribution - US Nationwide and in the countries: Australia, Canada, Chile, China, India, Korea, Malaysia, New Zealand, Singapore, Switzerland, Taiwan and Thailand. | Open, Classified |  |  | 236 | Computer
Z-2450-2015 | 71831 | No | JZG | Beta-2-microglobulin immunological test system. | Immunology | Immunology | 510(K) Exempt | Dimension Vista BÂż2-Microglobulin Flex Reagent Cartridge (B2MIC); |  |  | July 23, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | customers are receiving a high number of Abnormal Assay Errors or are not able to calibrate B2MIC when using Flex reagent cartridge B2MIC various lots. In failure mode, there is the potential for falsely depressed beta2-Microglobulin results on the Dimension Vista system due to assay reagent contamination. | Environmental control | Siemens sent an Urgent Medical Device Correction letter dated July 2015 to all affected customers. The letter identified the affected product, problem, risk to health, and actions to be taken by customers. For questions contact your Customer Care Center - Technical Solutions at 1-800-441-9250 or your local Siemens technical support representative. |  | 1655 | Total 1655 units | US Nationwide Distribution in the states of TN, AZ, CT, Fl, CA, NY and WI | Terminated | March 24, 2017 | 610 | 1655 | Not_Computer
Z-2452-2015 | 71812 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.0/4.0 Cannulated Drill Bit; a component of the Tiger Cannulated Screw System |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | The cannulation of the 3.0/4.0 Cannulated Drill Bit inner diamter is too small to allow for the guide wire to pass completely through. | Nonconforming Material/Component | Trilliant Surgical sent an Important Medical Device Advisory Notice to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were instructed they have the option to exchange the affected drill bits at no charge. To initiate the exchange, customers were instructed to contact their local Trilliant Surgical Sales Representative or contact the Customer Service Department directly at 1-800-495-2919 or via email at orders@trilliantsurgical.com. |  | 121 | 121 units | Nationwide Distribution including NY, TX, OH, IN, FL, NH, MN, SC, OK, NJ, CA, NC, MO, GA, WI, IA, TN, WA, and Puerto Rico. | Open, Classified |  |  | 121 | Not_Computer
Z-2453-2015 | 71826 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD CD56 PE Cy 7: NCAM16.2 ASR |  |  | July 17, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Three lots of CD56 PE Cy 7, Catalog number 335809, contain a low amount of HLA-DR PE antibody which could result in an unexpected staining pattern. | Material/Component Contamination | BD Bioscience sent an Urgent Product Recall Letter dated July 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer Actions : 1. If you have CD56 PE-Cy7 integrated into a Laboratory Developed Test (LDT) or panel, evaluate the impact of a false positive in the PE channel the interpretation of on the results.2. Discard all remaining stock of CD56 PE-Cy7, Catalog Number 335809, lot number 4300708, 4304864, and/or 5013643, and request replacement product. BD will replace product at no cost.3. As a requirement of our Regulatory Tracking Process, we request that you complete the enclosed Tracking/Verification form, even if you no longer have the affected reagents listed in your inventory. Please return the Tracking/ Verification form to BD Biosciences, per the instructions (mail, fax, or email).We apologize for any inconvenience that this may have caused and are working toward a long-term solution to this issue. If you require further assistance, please contact BD Customer Support at 877.232.8995 (prompt 3) in the United States. For customers outside the US, contact your local BD Biosciences representative or distributor. |  | 362 | 362 units | US Distribution and Internationally to Taiwan. | Terminated | March 11, 2016 | 238 | 362 | Not_Computer
Z-2456-2015 | 71815 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Concave Reamer Head, 16 mm; Concave Reamer Head, 18 mm; Concave Reamer Head, 20 mm; Concave Reamer Head 22 mm; Convex Reamer Head, 16 mm; Convex Reamer Head, 18 mm; Convex Reamer Head 20 mm; Convex Reamer Head 22 mm, used with Gridlock Plating system, for use in trauma and reconstructive procedures in small bones. |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | The cannulation of the Concave and Convex Reamer Heads may be too small to allow the Guide Pin to freely pass through. | Component design/selection | The recalling firm sent out written notification of the recall on 7/3/14. The letter instructed the consignees to exchange affected instruments. |  | 638 | 638 units | Nationwide throughout the US | Open, Classified |  |  | 638 | Not_Computer
Z-2451-2015 | 71876 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Ex-Fix System (Small, Medium, Distraction Osteogenesis (DO) and Large) ; (S) intended to stabilize and provide treatment for fractures of the small bones, such as the hand, wrist, forearm, foot, and ankle; (M) for the construction of an external fixation frame for the treatment of pediatric and adult fractures; )L) for use to provide treatment for long bone and pelvic fractures that require external fixation; and (DO) for fracture fixation (open and closed),pseudoarthrosis or non-unions of long bones, limb lengthening by epiphyseal or metaphysical distraction, correction of bony or soft tissue deformities and correction of segmental bony or soft tissue defects in adult and pediatric patients. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | Labeling changes were made related to MR conditions that align with standard definitions for MR Safe, MR Conditional, and MR Unsafe. This recall identifies additional Trauma External Fixation System part numbers that should be considered MR Conditional. One serious injury reported. | Labeling design | Urgent recall letters, dated July 31, 2015, were sent to end users and sales consultants to notifiy them about the issue, the associated risks, and to provide instructions for returning the affected devices and response form. |  | 5306080 | 5306080 | Worldwide Distribution-including US (nationwide) and Canada. | Open, Classified |  |  | 5306080 | Not_Computer
Z-2455-2015 | 71470 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Offset Femoral Alignment Guide 4mm x 6Âż; Product Code: KS 67024 |  |  | June 09, 2015 | 2015 | Omnilife Science Inc. | The markings on both versions of the OFFSET FEMORAL ALIGNMENT GUIDES are incorrect. Specifically, the markings on the alignment guides are reversed from the intended position that should provide proper alignment for the offset revision stems when used with the femur. | Device Design | Omni sent an Urgent Medical Device -Field Safety Notice letters to their customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to Return the attached Acknowledgement and Receipt Form by FAX to 508-819-3390. |  | 15 | 15 | Worldwide Distribution - US Nationwide in the states of OK, CO, OH, MA and the countries of Spain and Australia. | Terminated | September 28, 2015 | 111 | 15 | Computer
Z-2461-2015 | 71556 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Wheel caster set screw for Ferno iNX Wheeled Stretcher. Ferno iNX Wheeled Stretcher for pre-hospital patient transport. |  |  | May 01, 2015 | 2015 | Ferno-Washington Inc | The firm discovered an improperly installed set screw in one caster arm assembly on a unit in production. The screw assists in preventing the bearing cup on the caster assembly from backing out of the caster arm during use. | Process control | Ferno's first action of May 1, 2015 was to telephone all customers to alert them to the potential issue and explain our planned inspection. Ferno's Technical Support Manager emailed a copy of the Service Notice and iNX Equipment Evaluation Form to the customers as well. |  | 26 | 26 units | Distributed in the states of AZ, CO, IA, MD, OH, PA &amp; WA, and the countries of Australia, Canada &amp; United Kingdom. | Terminated | November 06, 2015 | 189 | 26 | Not_Computer
Z-2459-2015 | 71708 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX Activated Dialdehyde Solution, 4 bottles/case, P/N: 2266 |  |  | July 27, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the CIDEX Activated Dialdehyde Solution because it may lack the adequate data to support the entire 24-month shelf life. | Device Design | A Urgent Product Notification Letter dated July 27, 2015, will be sent to customers and distributor who purchased the CIDEX Activated Dialdehyde Solution because it may lack the adequate data to support the entire 24 month shelf life. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the requested information on the enclosed postage paid Business Reply Card and return it to Stericycle. Customers with questions or need additional packing labels are instructed to contact Stericycle at (877) 550-0701, M-F, 8am-5pm ET. Customers are instructed to report complaints or suspected problems with CIDEX Activated Dialdehyde Solution by contacting (888) 783-7723. |  | 23426 | 23,426 cases | US (nationwide) Distribution | Open, Classified |  |  | 23426 | Not_Computer
Z-2462-2015 | 71854 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PICC Kits with Peel-Away Sheath Component; catheter, intravascular, therapeutic, ultrasound |  |  | July 30, 2015 | 2015 | Arrow International Inc | the catheter peel-away component hub tabs may prematurely detach. One injury reported. | Unknown/Undetermined by firm | Teleflex sent an Urgent Medical Device Recall Notificatin letters dated July 30, 2015 to their affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. immediately discontinue use and quarantine any products with the catalogand lot number provided in the notice. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507Attn: Customer Service or email it to recalls@teleflex.com. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product toArrow International. 3. Customers were also asked to complete the enclosed Recall Acknowledgment Form if there is no affected product, this will allow for documented receipt of this letter. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 52109 | 52109 | US Nationwide Distribution and internationally to Canada | Open, Classified |  |  | 52109 | Not_Computer
Z-2460-2015 | 71682 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Carepak Sheer Adhesive Bandages 50 Sterile Bandages Assorted Sizes |  |  | June 10, 2015 | 2015 | ASO, LLC | Contain 0.1% Benzalkonium Chloride in the 5/8 bandages inside the assortment which is not listed as an active ingredient in the product label. | Process control | Customers were contacted by phone and e-mail on 06/10/2015; First class letter was sent on 06/23/2015. |  | 655 | 655 cases. | US Distribution to Rhode Island only | Open, Classified |  |  | 655 | Not_Computer
Z-2457-2015 | 71414 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 4.7, version 4.7.1, used with Structure Definition or Automatic Breast Planning. Radiation Therapy Treatment Planning System, for treatment planning and analysis of radiation therapy. |  |  | May 27, 2015 | 2015 | RAYSEARCH LABORATORIES AB | Issue with the algorithm for ROI contraction in RayStation 4.7 when non-uniform contraction distances are used. The ROI contraction tool uses six distances as input: right/left, inferior/superior and posterior/anterior. These contraction distances are, in error, pairwise reversed, i.e. right is interpreted as left, inferior is interpreted as superior, and posterior is interpreted as anterior. | Software design | On May 27, 2015 RaySearch Laboratories distributed Medical Device Correction notices to their customers via courier service. Within the Medical Device Correction notice, RaySearch Laboratories included a workaround to the problem for their customers. Customers should not use structure templates containing derived ROI definitions with non-uniform contractions. They should avoid using non-uniform contractions. In the case non-uniform contractions must be used, enter switched distances manually and review the result carefully before using it for treatment planning. Lastly, do not use humerus shielding in Automatic Breast Planning. Customers are also expected to complete and return the lower portion of the correction notice via fax to 888-501-7195. **CAPA** Correct the software problem in patch version 4.7.2 of RayStation 4.7, scheduled for market release June 2015. |  | 126 | 126 programs | Distributed in CO, IL, MI, NC, OH,TN, and WA. | Open, Classified |  |  | 126 | Computer
Z-2419-2015 | 71777 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge 86 Series Washer/Disinfector with G1 control system software, Model 8666 and Model 8668. For use in healthcare facilities and laboratories where reusable items such as instruments, utensils, anesthesia sets, and glassware are handled for decontamination purposes. |  |  | June 26, 2015 | 2015 | Getinge Disinfection Ab | Getinge Disinfection AB has initiated a Field Correction due to the potential for unauthenticated remote access to the product over the network. | Software Manufacturing/Software Deployment | Getinge Disinfection AB has initiated a Field Correction due to unauthenticated remote access to the product over the network. |  | 20 | 20 devices | Distributed in CA, CT, GA, IA, ID, MD, MI, MN, NY, PA, TX, and WA. | Open, Classified |  |  | 20 | Not_Computer
Z-2458-2015 | 71889 | No | OFF | Venous blood pressure manometer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Central Line &gt;34 Weeks Kit . Cardiology department. |  |  | August 03, 2015 | 2015 | Medline Industries Inc | The print plate artwork for this lot was found to state: Central Line &lt;34 Weeks Kit displaying an incorrect sign in front of the 34. The description for this device should have stated Central Line &gt;34 Weeks Kit. The defect may lead to the product to be used on patient population that is &lt;34 weeks rather than intended population of &gt;34 weeks. | Error in labeling | Medline Industries contacted the account directly via phone on August 4, 2015, they confirmed no affected product is in their facility. The lot numbers are easily identified on the outside of each kit. Second and third follow up communications will not be sent as we have confirmed the facility does not have any affected product.For questions regarding this recall call 847-643-4129. |  | 12823 | 320 units for lot 14DB9277 &amp; 320 units for lot 14PB2878 | US including PA | Terminated | January 05, 2017 | 521 | 12823 | Not_Computer
Z-2464-2015 | 71813 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Tiger Cannulated Tray Lid Part # 210-00-001-1 NON-STERILE; Combined Tiger/Headless Screw Part # 210-99-101 NON-STERILE; These are components of the Tiger Cannulated Screw System and the Tiger Headless Cannulated Screw System |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | Sterilization tray lids for the Tiger and Tiger Headless Cannulated Screw Systems were reported to occasionally be difficult to remove from the sterilization tray. | Component design/selection | Trilliant sent an Important Medical Device Adivsory Notice dated November 6, 2013, to all affected customers. The letter instructed customers that they have the option of exchanging affected instruments for trays that meet specification at no charge. To initiate the exchange, customers were instructed to contact their local Trilliant Surgical sales representative or the Customer Service Department directly at 1-800-495-2919, or via email at orders@trilliantsurgical.com. |  | 71 | 71 units | Nationwide Distribution | Open, Classified |  |  | 71 | Not_Computer
Z-2465-2015 | 71913 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Salto Talaris Tibial Guide, Part # MJU333 ,Tray/Kit Code YKAL11, Non sterile. |  |  | August 06, 2015 | 2015 | Tornier, Inc | Tornier has initiated a voluntary recall for the Salto Talaris Tibial Guide (part number MJU333) due to acomplaint that occurred during physician training, that the guide cannot be assembled to the Salto Talaris Talar Pin Setting guide. This issue will cause the impossibility the time of putting the Talar pin in the guide. | Nonconforming Material/Component | Distributors were sent on 8/6/2015 a Tornier "Urgent - Medical Device Recall" letter, dated August 5, 2015. The letter described the problem and the product involved in the recall. Consignees were advised to review their inventory and the inventory of all reps within their area for the recalled device. The letter requested that consignees complete and return the questionnaire to FieldAction@tornier.com. A Tornier representative will organize the collection and replacement of the devices. For questions they can contact Customer Service Department at 1 (888) 494-7950. |  | 5 | 5 | Distributed to CA, ME, NC, PA, and TX. | Terminated | April 08, 2016 | 246 | 5 | Not_Computer
Z-2463-2015 | 71804 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HardyCHROM MRSA |  |  | July 09, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling HardyCHROm MRSA due to a report of the potential for breakthrough of the methicillin-sensitive Staphylococcus aureus (MSSA) QC strain, ATCC 29213 on the identified lot. | Unknown/Undetermined by firm | The firm initially notified customers via phone to inform them of a recall. The phone scripts asks customers if they have had any issued with the QC testing. The firm states that our of a precaution they are asking customers to discard the remaining plates of the identified lot, and that credit or replacement plates will be offered for any remaining.The firm followed-up phone calls with a written notification letter. The letter states to complete, sign, and fax or email back the enclosed form stating compliance with the above action. The letters states that is the customer could not be reached via phone and would like replacements or credit to contact the Customer Service Department at 800-266-2222, option 1.Any questions call the Technical Services Department at 800-266-2222, option 2 or email at techservice@hardy diagnostics.com. |  | 2820 | 2,820 (282 pk/10) | Distributed in the states of VA, PA, NY, NJ, IL, OH, and MD. | Terminated | November 10, 2015 | 124 | 2820 | Not_Computer
Z-2467-2015 | 71829 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie IRIS Wheelchair. model EIZ4-2 in combination with option code 183M02 - MONO Backrest system with Dynamic Backrest option. |  |  | July 27, 2015 | 2015 | Sunrise Medical (US) LLC | Quickie IRIS and Zippie IRIS tilt-in-space wheelchairs with MONO Backrest System with Dynamic Backrest may break over time resulting in a fall or injury to occupant. | Device Design | Sunrise Medical sent an Urgent Medical Device Field Correction letter dated July 27, 2015, to all affected dealers. The letter identified the product, the problem, and the action to be taken by the dealer. Dealers were instructed to immediately contact their customers to make arrangements to have the correction made with a replacement kit which will be supplied by Sunrise Medical to each dealer. Each dealer is requested to send back to Sunrise Medical the Acknowledgment and Response Form(s) by fax, email or regular mail once the work is completed. Customers with questions were instructed to contact Sunrise Medical Regulatory Affairs at (888) 208-4901. |  | 64 | 64 total wheelchairs, both models. | Worldwide Distribution - US (nationwide) and Internationally to Australia, Germany and Canada. | Terminated | December 10, 2015 | 136 | 64 | Not_Computer
Z-2374-2015 | 71551 | No | PDR |  |  | Anesthesiology | PMA | SEDASYS Computer-Assisted Personalized Sedation System, SEDPRU01. |  |  | May 13, 2015 | 2015 | Ethicon Endo-Surgery Inc | Ethicon has found that disinfecting methods not specified in the Instructions For Use (IFU) have led to the degradation of the plastic material of the SEDASYS System Control Unit. | Device Design | Sedasys, a Division of Ethicon Endo-Surgery, Inc. sent an Urgent Customer Notification letter dated May 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Please distribute this information to all staff within your facility who use the SEDASYS System2. Complete the Business Reply Form (BRF) (Attachment A) confirming receipt of this notice within three (3) business days. The BRF may be given to your sales or service representative or sent to Sedasys by email at SedasysTechSupport@its.jnj.com.3. Keep this notice visibly posted in your facility for awareness.4. If you have any questions or concerns regarding the condition of your SEDASYS System including the Control Units, please contact your Sales or Service Representative.If you need clinical or product support, please contact your local sales representative or call our Customer Support Center at 1-800-SEDASYS (1-800-733-2797). If you need additional communications packages, contact 1-800-SEDASYS and reference Event 052715SED01.For questions regarding this recall call 513-337-7439. |  | 14 | 14 units | Nationwide Distribution including OH, RI, TX &amp; WA. | Terminated | May 02, 2016 | 355 | 14 | Not_Computer
Z-2471-2015 | 64928 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MiniCap Extended Life PD Transfer Set Extra Short (4") with Twist Clamp. Product Code: 5C4483. |  |  | February 25, 2013 | 2013 | Baxter Healthcare Corp. | Labeling Revision: Additional instructions for the use of peritoneal dialysis transfer sets including the sanitization the connection, healthcare provider monitoring of thyroid function, caution against reuse or reprocessing of the devices, warning against use if tip protectors are not in place, and advisement that the product does not contain natural rubber latex. | Labeling design | Baxter sent an IMPORTANT PRODUCT INFORMATION letters dated February 25, 2013 to affected customers via USPS. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached reply form and return it to Baxter by faxing it to 224-270-5457 of scanning the form and e-mailing it to fca@baxter.com. For general questions about this communication can be answered by contacting The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). Clinical questions can be answered by calling Baxter's Renal Division Clinic Help line at 888-736-2543 x 2 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 260244 | 260,244 units | Worldwide Distribution - US: Nationwide + American Samoa + District of Columbia + Guam + Northern Mariana Islands + Puerto Rico; *** FOREIGN: Antigua, Argentina, Australia, Bahamas, Barbados, Bermuda, Brazil, Canada, Chile, Colombia, Cuba, Curacao, Dominican Republic, Ecuador, El Salvador, France (New Caledonia), Guatemala, Guyana, Honduras, Hong Kong, India, Indonesia, Jamaica, Japan, Macau, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Suriname, Taiwan, Thailand, Trinidad and Tobago, United Kingdom (Cayman Islands), Uruguay, Venezuela | Terminated | March 08, 2016 | 1107 | 1543011 | Not_Computer
Z-2468-2015 | 71805 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker |  |  | July 10, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions has received reports indicating that devices from one lot (lot# 3574704) were labeled as 100mm models (CB5LT) while the devices inside of the packaging were 75mm models (CB5ST). | Labeling mix-ups | On 07/10/15 the firm sent out notification letters to all customers. The letter states for the customer to discontinue use of the affected lot, and to use other trocar lots reporcessed by the firm. The firm requests that the Recall Effectiveness Check Form be completed. The firm is going to be sending a sales representative to the facilities to check inventory for the affected trocars. A Recall Effectiveness Check will also be completed. If the form indicates that affected devices remain in inventory a prepaid shipping label will be issued for the return of the product. Customers will receive credit for all affected devices returned. The form must be completed even if no affected devices are found in inventory. Stryker Sustainability Solutions Complaint Hotline: 1-888-888-3433 x5555 |  | 90 | 90 devices | U.S. Nationwide Distribution to the following states of:: NV, NM, TX, and MI. | Terminated | December 04, 2015 | 147 | 90 | Not_Computer
Z-2476-2015 | 71735 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, 10mL in 12mL., Single use, Rx only, Sterile Solution. |  |  | July 02, 2015 | 2015 | MRP, LLC dba AMUSA | Procedures for the acceptance and control of in-process product have not been adequately established. | Packaging process control | The firm notified their consignees of the recall by Fed Ex on 07/02/2015. The firm requested the return of the product. Distributors were instructed to notify their customers. |  | 551 | 551 cases | Distributed in the states of CA. PA, VA, TX, MA, MD, WV, UT, FL, and TN. | Terminated | April 03, 2017 | 641 | 551 | Not_Computer
Z-2475-2015 | 71834 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument for Sternal ZIPFIX; indications for use include primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion |  |  | July 22, 2015 | 2015 | Synthes (USA) Products LLC | The end cap may loosen and detach making the instrument non-functional. No injuries reported. | Device Design | Urgent recall notification letters, dated July 22, 2015 were sent to end users and consultants to inform them of the issue, the associated risks, and provided instructions to return affected devices along with the response form. |  | 451 | 451 | Nationwide and internationally to Canada. | Open, Classified |  |  | 451 | Not_Computer
Z-2470-2015 | 71964 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Universal Viewer |  |  | August 11, 2015 | 2015 | GE Healthcare | Images from the Centricity PACS-IW with Universal Viewer andCentricity Universal Viewer may be missing when a system parameter MapRoute is set to a value greater than 1. | Software design (manufacturing process) | The firm sent out an Urgent Medical Device Correction letter dated 8/11/2015 to all affected customers. Letter states that No User activity required. GE Healthcare has completed a remote inspection of your system and found that this system parameter is currently set to the default setting of 1. The firm will correct all affected systems by a providing software update.The letter also states that if the customer has any questions they should call GE Healthcare Service at 1-800-437-1171 or their local service representative. |  | 26 | 26 systems | Worldwide Distribution - U.S. nationwide in the states of: AK, CA, CO, DE, FL, HI, ID, NY, PA, SD, TN and TX; France, Germany, Greece, India, Israel, Malaysia, Mexico, Northern Ireland (UK), Panama, Poland, Russia, South Africa, Saudi Arabia, Switzerland, Turkey, UAE and Uruguay. | Terminated | January 05, 2017 | 513 | 108 | Computer
Z-2477-2015 | 71691 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | OPD-Scan III Refractive Power/Corneal Analyzer Opthalmic |  |  | June 15, 2015 | 2015 | Nidek Inc | Software bug was found where there was no difference in Total and Corneal high-order aberrations, but differences were found in Internal high-order aberrations. | Software design | Dear Doctor letters mailing started on July 15, 2015 - first 100. They will be mailed out as the correction is received from Nidek, Japan.Dear Doctor letter will include software upgrade kit including CD or USB, instruction for upgrade and pre-paid response card to be returned to Nidek Inc. |  | 3836 | 3836 | Nationwide Distribution. | Terminated | April 25, 2016 | 315 | 3836 | Computer
Z-2479-2015 | 71840 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RAPIDPoint 500 Blood Gas Analyzer, |  |  | July 28, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | When both ports (Serial and Ethernet) are configured to transmit data, the data stream from one port could potentially affect the data stream from the other port causing the data management system to include either: duplicate data, missing data, or data from a different patient. | Software design | Siemens Healthcare issued Urgent Field Safety Notice sent by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on July 28, 2015, both in the United States and outside of the United States, for communication with affected customers. This notice informs customers of the issue with the RAPIDPoint 500 dual port LIS transmission feature.Users are informed to: If the system is currently set up with dual LIS transmission enabled, one of the ports (either Serial or Ethernet) needs to be disabled. Instruction provided with the notification. Customer to complete Field Effectiveness Check Form to confirm receipt of the notice.The resolution of this issue will be implemented in the next software release for the RAPIDPoint 500 system.If you have any questions, contact Siemens Customer Care Center or your local Siemens technical support representative. Please fax this completed Customer Care form to (312) 275-7795. |  | 5241 | 5241 units | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Albania, Algeria, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg, Brazil, Bulgaria,Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany,Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands. New Zealand, Norway, P.R. China, Peru, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, U.A.E., United Kingdom and Vietnam. | Open, Classified |  |  | 5241 | Computer
Z-2478-2015 | 71970 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | FIRMap 60mm Catheter; |  |  | August 12, 2015 | 2015 | Abbott Electrophysiology | The expiration date was incorrectly printed on the pouch/box labels of 19 products that were shipped to customer sites. The manufacturing date was accidentally used (copied) for the expiration date. | Incorrect or no expiration date | The firm, Abbott Electrophysiology, sent an "URGENT PRODUCT RECALL" letter dated August 11, 2015. A customer recall notification letter was distributed to firm representatives who will visit each consignee and inspect inventory for affected product. Letters were provided to reps on August 12, 2015. The letter describes the product, problem and actions to be taken. The customers were informed that the product they have in stock will be return to the firm and will be provided a replacement product free of charge.If you have any questions, contact Quality Director at 650-681-1744. |  | 86 | 86 units | US Distribution in states of: KY, IN, MO, NY, OH, TX, LA, IL, PA, NJ, WA, VA, FL, GA, CT, CA, AZ, CO, and TN. | Terminated | September 23, 2015 | 42 | 86 | Not_Computer
Z-2481-2015 | 71835 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | COULTER ISOTON 4 Diluent, Catalog No. 8547148. The diluent provides the ability to analyze portions of the diluted blood sample for different blood cell types, such as red blood cells and platelets. |  |  | July 22, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the LH 750 and LH780 Hematology Systems because the use of the diluent lots may eventually result in a compromised Hgb cuvette. | Nonconforming Material/Component | An Urgent Medical Device Recall letter dated 7/22/15 was sent to the affectedcustomers to inform them of the issue and actions to take. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Beckman Coulter's Customer Support Center, http://www.beckmancoulter.com/customersupport/support, (800) 526-7694 in US and Canada. Customers outside of the US and Canada, are instructed to contact their local Beckman Coulter Representative. |  | 710 | 710 units | Distributed US (nationwide) including Puerto Rico and U.S. Virgin Islands and the countries of Bermuda, Canada, Guatemala, Hong Kong, Mexico, Panama, Taiwan, and Trinidad-Tobago. | Terminated | December 18, 2015 | 149 | 409752 | Not_Computer
Z-2482-2015 | 71611 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Contour Next Control Solution for Contour Next Blood Glucose Test System. |  |  | June 22, 2015 | 2015 | Bayer Healthcare, LLC | If the control solution bottle is not thoroughly shaken prior to use, the control test result may be higher than the control range and may not be recognized and marked as a control test by the blood glucose meter. Also, the control solution, used as a quality control test in blood glucose meter systems, may read outside the specified range too high or too low. | Device Design | On 6/22/2015, Bayer Healthcare began sending URGENT - MEDICAL DEVICE FIELD CORRECTIVE ACTION notifications to affected Health Care Providers, Pharmacies, and Users. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are instructed to always check the expiration, not to use and discard of expired control solutions. Also, only CONTOUR NEXT control solution should be used with the CONTOUR NEXT test strips. The control solution bottle should be shaken about 15 times before each use. The letter instructed customers to refer to their meter's user guide for further testing instructions. If test results outside the control range are received, another test should be performed following the instructions provided in the letter. If the problem persists, customers should contact Bayer Diabetes Support. Bayer will replace the control solution or provide new control ranges for the test strips, depending on the situation. The firm has implemented improvements to provide more consistent control solution product. |  | 5000000 | 5 million bottles of control solution | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Croatia, Denmark, Finland, France, Germany, Greece, Ireland, Italy, Luxembourg, Netherlands, Norway, Portugal, Slovenia, Spain, Sweden, Switzerland, UK, Albania, Azerbaijan, Bulgaria, Czech Republic, Estonia, Hungary, Korea, Latvia, Lithuania, Philippines, Romania, Serbia, Slovakia, and Thailand. | Open, Classified |  |  | 5000000 | Not_Computer
Z-2483-2015 | 71850 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vision RingLoc Shell 50 mm Size 23. Hip joint replacement prosthesis. |  |  | July 28, 2015 | 2015 | Biomet, Inc. | The device was assembled without screw hole plugs. | Under Investigation by firm | Zimmer Biomet issued notices on July 28, 2015, via FedEx overnight mail and e-mail to customers. Customers were asked to return response forms and recalled devices to Zimmer Biomet. Customers with questions may call Seth Pixton, Field Action Specialist, at 574-372-1570, Monday through Friday from 8:00 am to 5:00 pm, or e-mail at: seth.pixton@zimmerbiomet.com. |  | 5 | 5 | Distributed in Canada and Japan. | Open, Classified |  |  | 5 | Not_Computer
Z-2487-2015 | 71988 | No | JBF | Erythrocytic glucose-6-phosphate dehydrogenase assay. | Hematology | Hematology | 510(k) | Pointe Scientific G6PD Controls |  |  | May 19, 2015 | 2015 | Pointe Scientific, Inc. | A reduction in the reconstituted stability has been identified.Clinicians may notice a drop in recovery of the control after reconstitution, which could result in an out of specification result prior to the current reconstitution claim of 7 days. For this reason we suggest discontinuing use of the control set. | Nonconforming Material/Component | The firm, POINTE SCIENTIFIC, issued a "TECHNICAL BULLETIN" on 5/19/2015, notifications were sent to the affected distributors via electronic mail. The recall notification included a description of the reason for the recall, affected product, and actions to be taken. The customers were instructed to: discontinue use of the control set and to incorporate this bulletin into your Quality Control files.The firm noted that this product has been placed on indefinite back-order and a replacement product is not currently available. For questions contact Pointe Scientific's Technical Service Department at 1-800-757-5313. |  | 16632 | 2772 total vials (374 kits of 6 vials, 528 invidually sold vials) | Worldwide distribution: US (nationwide) including states of: PA, CA, KY, FL, NJ, DE, MI, MN,and IL; and countries of: United Arab Emirates, Austria, Italy, Israel, Greece, and Australia. | Open, Classified |  |  | 16632 | Not_Computer
Z-2488-2015 | 71862 | No | MDI |  |  | Orthopedic | 510(k) | VEPTR Vertical Expandable Prosthetic Titanium Rib, inferior cradles also known as caudal rib supports. The VEPTR/VEPTR II devices mechanically stabilize and/or correct thoracic deformities to allow the chest and lungs to grow. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | The last hole on certain VEPTR Inferior Cradles may be partially formed. Fully lengthening the construct to the partially formed last hole may result in mechanical failure of the construct due to rib sleeve breakage, and may require surgury to exchange components. Effects may also include pain, loss of deformity correction, and/or soft tissue/visceral injury. | Process control | Urgent recall letters, dated July 29, 2015, were sent to end users alerting them to the problem, the associated risk to patients, and provided instruction to return affected devices withthe response form. |  | 3262 | 3262 | Worldwide distribution including US nationwide, Canada and Germany. | Open, Classified |  |  | 3262 | Not_Computer
Z-2490-2015 | 71967 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 7 Hole VL Gridlock Fibula Plate, Part # 300-60-001 NON-STERILE SINGLE USE ONLY; 9 Hole VL Gridlock Fibula Plate, Part # 300-60-002 NON-STERILE SINGLE USE ONLY; 11 Hole VL Gridlock Fibula Plate, Part # 300-60-003 NON-STERILE SINGLE USE ONLY. For use in trauma and reconstructive procedures in small bones. |  |  | August 06, 2015 | 2015 | Trilliant Surgical Ltd. | The affected parts subject to the recall are out of specification, resulting in in the loss of the optimal locking screw functionality of the plates. The thread depth of the screw holes is insufficient, potentially limiting the ability to engage and lock with associated mating locking screws. | Nonconforming Material/Component | Trilliant Surgical sent an Urgent Medical Device Recall letter dated August 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed and confirmed consignees to discontinue use of the affected lots, quarantine, and return product (if not yet done so) in accordance with an included acknowledgement sheet. For further questions, please call 1(800) 495-2919 , 8:00am - 5:00pm CST |  | 28 | 28 units | US Distribution to the states of : FL, OH, NJ, MN, WI, NM and CA. | Open, Classified |  |  | 28 | Not_Computer
Z-2489-2015 | 71855 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow International, Inc. PICC Kits and VPS Access Kits with Touhy Borst Component Catheter, Intravascular, Therapeutic, Ultrasound, Long Term Greater than 30 Days |  |  | July 30, 2015 | 2015 | Arrow International Inc | saline leakage from unsecured connections between the Touhy Borst adaptor and the hub of the catheter. No injuries reported. | Under Investigation by firm | Arrow International sent an Urgent Medical Device Recall Notification letter dated July 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:If you have affected stock, immediately discontinue use and quarantine any products with the catalogand lot number listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507Attn: Customer Service or email it to recalls@teleflex.com. This will allow us to document the amountof product you have on hand for return. A customer service representative will contact you with aReturn Goods Authorization (RGA) Number and will provide instructions for the return of product toArrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. This will allow us to document your receipt of this letter.Customers with questions were instructed to call their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 92902 | 92902 | Worldwide Distribution - US nationwide and Internationally to Canada. | Open, Classified |  |  | 92902 | Not_Computer
Z-2492-2015 | 71966 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent* Closed Suction System for Adults, 12 F, T-Piece (Product Code 2155); KimVent* Closed Suction System for Adults, 14 F, T-Piece (Product Code 2205, 8169, 8248, 22015, 22025, 220125, 220135) |  |  | July 29, 2015 | 2015 | Halyard Health, Inc | The thumb valve of some KIMVENT Closed Suction Systems remained partially open after suctioning was completed, even though the valve appeared to be closed or in the "LOCKED" position. | Process control | Halyard sent a Product Advisory Notice on July 29, 2015, via FEDEX with a second letter dated August 12, 2015, being issued via FEDEX. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to report any occurrence of a valve leak to Halyard Customer Service at 1-844-425-9273, option 1, option 3. |  | 19250 | 19,250 cases | Worldwide Distribution - US (nationwide and Puerto Rico and Internationally to Canada | Terminated | January 12, 2016 | 167 | 19250 | Not_Computer
Z-2491-2015 | 71882 | No | OUO | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/MR, Diagnostic Imaging System , Philips Medical Systems |  |  | July 23, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has become aware of a problem in which the Ingenuity TF PET/MR PET Reconstruction Server (PRS) database may lock up after an MR acquisition is completed and before a PET acquisition is begun, even though no actual database lockup issue has been reported in the field-installed Ingenuity TF PET/MR systems. | Software design | On July 23, 2015 the firm sent Field Safety Notices to their customers. |  | 11 | 11 | Product was shipped to the following states: NY, OH &amp; TX. Product was also shipped to the following countries: Finland, Germany, Hong Kong, Japan, Netherlands, South Korea, Spain &amp; Switzerland. | Open, Classified |  |  | 11 | Not_Computer
Z-2486-2015 | 71803 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Bright Embrace, a Single Bili Light for newborn phototherapy |  |  | August 31, 2013 | 2013 | Physician Engineered Products, Inc. | Upper limit for ambient temperature for use of the device was printed incorrectly in the instructions | Device Design | Physician Engineered Products, Inc. sent an Bright Embrace Recall notification dated September 10, 2013, to notify customers of the recall of the Bright Embrace .PEP is issuing a Return Authorization to you. Upon receipt PEP will reimburse to you:1) The cost of the device(s) 2) Shipping costs-both from and to PEPFor further questions please call (207) 935-1256. |  | 97 | 97 units | Worldwide Distribution - US Distribution to the states of : CA, CO, FL, KY, MA, MD, NM, PA, TN, TX and VA., and to the countries of : Egypt, Honduras, Saudi Arabia and Singapore. | Terminated | September 02, 2015 | 732 | 97 | Not_Computer
Z-2529-2015 | 71893 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Mahurkar Elite Acute Dual Lumen Catheter Tray, High Flow Pre-Curved with IC* Safety Components, 13.5 Fr x 20 cm |  |  | August 06, 2015 | 2015 | Covidien LLC | Sterility may be compromised | Packaging process control | Medtronic/Covidien sent an Urgent Medical Device Recall letter August 7, 2015 to affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructs customers to quarantine, discontinue use and return the affected product as instructed in the letter. For questions contact your Medtronic representative or Customer Service at (800)-882-5878 |  | 4 | 4 kits | US Nationwide Distribution | Terminated | June 29, 2017 | 693 | 420002 | Not_Computer
Z-2538-2015 | 71653 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Panoramic X-ray Model PC-1000 |  |  | June 04, 2015 | 2015 | Panoramic Rental Corp. | Jackscrew nut part failure potentially leading to device malfunction and patient/operator injury. | Equipment maintenance | The firm, Panoramic Corporation, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 7/17/2015 via certified mail to customers on 7/20/2015, with response forms for customers to complete and return. The firm previously issued "IMPORTANT SERVICE BULLETIN" notices in 2003 and 2008 via FED-EX. The letter describes the product, problem and actions to be taken. The customers are instructed to select to use several options including replacing the failing part with one containing the safety feature, or trading in their machine for a newer one; discontinue using the devices immediately regardless the option you select, and complete and return the attached response card via email to recall@pancorp.com or fax to 260-489-5683. If the machine has been sold or donated to another dental outfit, the customer should notify Panoramic so that they can update their records. When the customer is reached, the firm will confirm whether or not the machine is still in service.Customers with questions may contact 800-654-2027 Monday through Thursday 8:00 AM to 6:00 PM and Friday 8:00 AM to 5:00 PM EST to discuss these options and determine which is best for their practice. |  | 7885 | 7885 | Worldwide distribution: US (nationwide) all 50 states, including Puerto Rico; and countries of Antigua, Armenia, Aruba, Austria, Canada, China, Colombia, Ecuador, Egypt, England, France, Germany, Greece, Haiti, Honduras, Indonesia, Japan, Kenya, Lithuania, Mexico, Panama, Saudi Arabia, Singapore, Taiwan, Turkey, UAE, Uganda, Venezuela, Vietnam, and Virgin Islands. | Terminated | May 30, 2017 | 726 | 7885 | Not_Computer
Z-2541-2015 | 71895 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CrossCHECK Locking Screw in the following sizes: |  |  | July 21, 2015 | 2015 | Wright Medical Technology, Inc. | These screws have a potential of passing through the plate in the case when torque is applied in excessive amounts, which may result in intraoperative malfunction. | Employee error | Wright sent an Urgent Medical Device Recall letter dated July 20, 2015, to all affected customers. Customers were asked to complete the attached response form confirming receipt of the notice and the status of any products on hand, and return a copy to Wright by email (complaints@wmt.com or fax (901)867-7401. Also, if customers have any product in their inventory the product should be sent back to Arlington, TN (ALM). For questions regarding this recall call 901-290-5662. |  | 107 | 107 units | Nationwide Distribution | Open, Classified |  |  | 11588 | Not_Computer
Z-2524-2015 | 71872 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 2 Waste Management System Docking Station (120V and 230V), New and Refurbished |  |  | July 20, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Fluid Coupling Sleeve, a component within the Neptune Waste Management System Docker, may be out of specification. This can cut or dislodge the PTFE seal and cause a leak between the Rover and Docker couplers during docking or leaking from the Rover coupler after docking. The potential hazard is waste fluid and/or fresh water dispensed from the Neptune 2 Rover at the incorrect time. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Notification letter dated August 4, 2015 via FedEx to all domestic customers, and all international cutsomers and sales reperesentatives were notified via e-mail. The letter identified the affected product, problem and actions to be taken. Customers are asked to complete the response form and return it to Stryker whether they have affected devices or not. Fax the completed, signed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com. For questions regarding this recall, contact Stryker Instruments at 269-389-2921. |  | 124188 | 524, 237 of these are non-conforming | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and countries of: Austria, Canada, Hong Kong, Netherlands, Finland and Spain. | Terminated | November 07, 2016 | 476 | 124188 | Not_Computer
Z-2383-2015 | 71838 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost Release 4.0.3, Release 4.1/4.1.1. excluding systems with SP1 |  |  | March 23, 2015 | 2015 | Philips Electronics North America Corporation | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal due to a defect in the Microsoft Windows 7 operating system. | Radiation Control for Health and Safety Act | Philips Healthcare sent an Urgent Field Safety Notice letter dated June 25, 2015 to customers. The letter identified the affected product, problem and actions to be taken. Philips plans to install a software upgrade in affected systems, which will eliminate this issue. Philips Healthcare Field Service Engineers will schedule an appointment with customers to install the software update. For questions contact your local philips representative. |  | 62 | 62 | US Nationwide Distribution to the states of : Arizona, Minnesota, State of Washington, Wyoming and Ohio. Internationally to Canada. | Open, Classified |  |  | 62 | Computer
Z-2546-2015 | 71963 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | MediTrace Cadence Adult Zoll, Preconnect |  |  | August 19, 2015 | 2015 | Covidien LLC | Multi-Function Defibrillation Electrodes, Pre connect-potential for damage to the wire insulation may result in a potentially increased risk for reduced or no patient therapy and patient and/or clinician burns | Process control | Medtonic/Covidien issued recall letter on August 19, 2015. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 2 Ludlow Park Drive, Chicopee, MA 01022, Attention: Field Action Returns.Contact your Medtronic representative or Customer Service at (800)-882-5878. |  | 114000 | 114,000 sets | NationwideForeign:CanadaAustraliaDenmarkGermanyNew Zealand | Terminated | May 24, 2017 | 644 | 177660 | Not_Computer
Z-2543-2015 | 71785 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobas Liat Analyzer, for in vitro diagnostic use. |  |  | July 24, 2015 | 2015 | Roche Molecular Systems, Inc. | Due to a supplier issue, the incorrect green and amber LED components were used in the manufacture of the LED printed circuit board assemblies (PCBA) in certain cobas Liat Analyzers. | Nonconforming Material/Component | Roche sent an Urgent Medical Device Recall letter dated July 24, 2015, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Actions Required Identify the affected systems in your inventory by referencing the serial numbers in the table on page 1. You will be contacted to return affected systems for immediate replacement.Complete the attached fax form and fax it to 1-317-521-4815.File this UMDR for future reference.Please contact the Roche Support Network Customer Support Center cobasÂż Liat hotline at 1-800-800-5973, if you have questions about the information contained in this letter. |  | 5 | 5 pieces | US Distribution to the states of :TX and NC. | Terminated | December 23, 2015 | 152 | 5 | Not_Computer
Z-2540-2015 | 71972 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SMR Shoulder/Spalla Trial Heads ECC .4 mm Adaptor, REF 9013.30.021; SMR Shoulder/Spalla Trial Heads ECC .8 mm Adaptor, REF 9013.30.031; for use in shoulder implant surgery. |  |  | August 14, 2015 | 2015 | Limacorporate S.p.A | The dimensions of the mating features of the recalled products are out of specification causing the device to be difficult to assemble. | Nonconforming Material/Component | Lima Corporate sent an Urgent Attention Recall letter on August 11, 2015, to US distributors and Sales Agents. The affected customers were instructed to examine their inventory, return the affected lots to the recalling firm, and complete a response form. |  | 93 | 93 units | US Distribution to the states of: OH, TX, NC, NY, CA, FL, KS, AL, MI, MT, CO, MO, NV, WA, MN and LA., and Internationally to Australia. | Terminated | February 26, 2016 | 196 | 203 | Not_Computer
Z-2544-2015 | 71619 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number/Product Code 6801375, Unique Device Identifier No. 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number/Product Code 6801890, Unique Device Identifier No. 10758750001644, IVD --- Ortho Clinical Diagnostics Inc. |  |  | June 30, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly may occur during the installation of an Assay Data Diskette (ADD) using Software Versions 2.2.1 through 2.8. If the user is utilizing User- Configured (Modified) parameters, ALL default settings are restored in the VITROS 5,1 FS Chemistry System from the ADD instead of retaining the user modified parameters. | Software design | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated 6/30/15 and response forms to customers via FedEx overnight courier, e-mail (ORTHO PLUS e-Communications) and/or via US Postal Service overnight letter on 6/30/2015. Foreign affiliates were informed by email on 6/30/2015 of the issue and instructed to notify their consignees of the issue and required actions. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the required actions and complete and return the Confirmation of Receipt-Response Required form, no later than 7/8/15, by fax at 1.888.557.3759 or 1.585.453.4110, or scan PDF and email to: ConfirmationAdmin@its.jnj.com. Customers with any additional questions, can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 210 | Catalog 6801375: Domestic 919; Foreign 1251; Catalog 6801890: Domestic 63, Foreign 147 | Worldwide Distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 210 | Computer
Z-2719-2015 | 71978 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza; picture archiving and communications system |  |  | July 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Potential issue leading to data loss and patient data mix-up | Software design | Siemens sent a Customer Information letter, dated July 6, 2015, to customers The letter identified the affected product, problem and action to be taken. The letter informed cusotmers about the Hotfix software update that was being released. |  | 37 | 37 | US Nationwide Distribution - in the states of CA, OR, FL, MO, NY, WA, MD, MA, NY, WI, KS, MN and PA. | Terminated | January 06, 2017 | 550 | 37 | Computer
Z-2718-2015 | 71881 | No | PDU | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano Pioneer Plus Re-Entry Catheter; |  |  | July 22, 2015 | 2015 | Volcano Corporation | Due to a manufacturing defect, there is the possibility that a small wire could extend through the catheter shaft of the Pioneer Plus intravascular Ultrasound Guided Re-Entry Catheters. | Under Investigation by firm | The Volcano Corporation sent an Urgent Medical Device Recall letter dated July 22, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately quarantine and return all affected product in their inventory. Customers were also asked to complete the attached form and fax back to 866-723-2482. Customers with questions were instructed to call 866-861-8923.For questions regarding this recall call 978-439-3586. |  | 5042 | 5042 devices | Worldwide Distribution - US (nationwide) Canada, Europe, Saudi Arabia, Iran, Egypt, and Qatar. | Terminated | April 28, 2016 | 281 | 5042 | Not_Computer
Z-0417-2015 | 69545 | No | JZH | Properdin factor B immunological test system. | Immunology | Immunology | 510(K) Exempt | MicroVue Bb Plus EIA is packaged as a kit which contains all reagent components to perform the test. In vitro diagnostic. |  |  | October 16, 2014 | 2014 | Diagnostic Hybrids Inc | Samples are quantitating incorrectly, with approximately 50% higher concentration values than previously reported and historically documented. | Device Design | Quidel sent an Urgent Medical Device Recall letter dated October 16, 2014, to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately review their inventory and determine if they have of the affected product. If so, they should remove it from their inventory and destroy it immediately by discarding it into their normal biomedical waste stream. Complete the enclosed Certificate of Destruction. Upon receipt of the Certificate of Destruction, Quidel, will then schedule their shipment of replacement product. If consignees further distributed the affected product, they should identify the customers and provide them with the Urgent Medical Device Recall. Consignees were instructed to monitor and reconcile the quarantine of product with the customers as this information may need to be provided to the authorities. For questions regarding this recall call 740-589-3300.For technical support consignees should contact technicalsupport@quidel.com or call 800-874-1517. |  | 171 | 171 | Worldwide Distribution - USA, including AL, CO, CT, FL, KY, MN, NC, NV, OH, PA, and WA; and, the countries of Canada, China, India, and Switzerland. | Terminated | September 03, 2015 | 322 | 171 | Not_Computer
Z-2548-2015 | 71721 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | ANGIODRAPE PK 4/CS |  |  | July 10, 2015 | 2015 | Customed, Inc | Recall initiated due to inadequate validation/qualification of sterilization processes, and package integrity for products released for distribution as sterile. | Packaging | Customed issued an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" notification letter dated July 10, 2015 to all affected customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 192 | 192 packs | Distributed Only in Puerto Rico. | Terminated | April 12, 2017 | 642 | 358405 | Not_Computer
Z-2732-2015 | 72024 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 Intermittent Catheter with Sure Grip Hydrophilic Coated Silicone Coude tip, Product Code 50614. The catheters are packaged in film-to-film pouches and packaged 30 per box. |  |  | August 24, 2015 | 2015 | C.R. Bard, Inc. | Potential breach of the sterile barrier packaging. | Unknown/Undetermined by firm | The firm, Bard, sent an "URGENT: MEDICAL PRODUCT RECALL" letter dated August 20, 2015 to customers via FedEx on 8/24/2015. The letter describes the product, problem and actions to be taken. The customers were instructed to not use or further distribute any affected product; check inventory and remove any identified affected product from your shelves; complete and return the Recall and Effectiveness Check Form by September 19, 2015, even if you have used or no longer have possession of the affected product, via fax to 1-770-784-6469 or email a scanned copy to Regan.Gealy@crbard.com ( if you cannot fax or email the form, call 1-770-784-6471 and report the required information verbally); and if you have further distributed any units of the affected lot, identify your customers and notify them at once of this product recall, and return product to Bard Medical Division, 8195 Industrial Blvd., Covington, GA 30014. BMD will provide replacement product or credit for your return product.If you have any questions, please call Vice President, Quality Assurance at 770-784-6119. |  | 8970 | 8,970 each | US Distribution to states of:CA, FL, IL, MA, MN, NH, NJ, NY, NC, OH, OK, PA, TX, and WA. | Open, Classified |  |  | 8970 | Not_Computer
Z-2734-2015 | 72026 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi Surgical System P4 software;with Xi Stapler; model number IS4000 : |  |  | August 17, 2015 | 2015 | Intuitive Surgical, Inc. | Complaints regarding the Exposed Knife Blade recoverable fault for the Xi Stapler with Xi System P4 version software. | Software design | Intuitive sent an Urgent Medical Device Correction letters dated August 21, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact Clinical Sales or ISI Customer service at 800-876-1310 Option 3. |  | 198 | 198 affected devices | Worldwide Distribution - US Nationwide and in the countries of Italy and Norway. | Terminated | February 08, 2016 | 175 | 198 | Computer
Z-2730-2015 | 71832 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Neuro Kinetics, Inc. (NKI) I-PortalÂż NOTC system with the OVAR research test; a rotary chair system used in vestibular and neuro-otologic diagnostic testing. This device provides no diagnoses nor does it provide diagnostic recommendations.The OVAR test has not yet been cleared by the FDA, and until then can only be used as an investigational device. |  |  | July 21, 2015 | 2015 | Neuro Kinetics, Inc. | The devices that included the OVAR research test were not cleared for marketing and mTBI research package was sold with the incorrect investigational label. | No Marketing Application | Neuro Kinetics sent notification letters to end users on/about July 24, 2015, to alert them to the reason for recall and inform them of the corrective action planned by Neurokinetics.Action to be taken by you the Customer/User:If you have already initiated OVAR research (an IRB approved study with non-significant risk designation is ongoing), please reply to this email and include a copy of the IRB approval letter or the IRB approval information (e.g., investigators name, IRB name, IRB location, and study reference number). In return, NKI will mail you a field correction package with the necessary and appropriate Investigational Device labeling.If you have not initiated OVAR research (an IRB approved study is not ongoing), please discontinue use of the OVAR test and reply to this email to confirm that you are not using the OVAR test as part of an IRB approved study. In return, NKI will mail you a field correction package to disable the OVAR test functionality, which can be restored at the time you have an IRB approved study. Disabling the OVAR research test will have no effects on the functionality of the clinical tests included with your NOTC system. There will be no perceivable changes to the system outside of not being able to run the OVAR test.For questions regarding this recall call 412-963-6649. |  | 45 | 45 | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 45 | Not_Computer
Z-2722-2015 | 71907 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Transonic Flow Probe, 6 mm, Catalog Number HQC6FMV -- |  |  | July 27, 2015 | 2015 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number. | Software Manufacturing/Software Deployment | Transonic sent an Urgent Medical Device Recall letters and Response Forms dated July 27, 2015 to customers via Certified Mail, email and/or fax. The letter identified the affected product, problem and actions to be taken. Customers are instructed to contact Transonic Systems' Customer Service at 1-800-353-3569 or 1-607-257-5300 to arrange for replacement (repair/exchange) immediately. |  | 1 | Domestic: 1 unit | Worldwide Distribution - US Nationwide and the countries of: Canada, Japan, Russia and The Netherlands. | Terminated | July 18, 2016 | 357 | 143 | Computer
Z-2733-2015 | 71878 | No | MSW | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access Thyroglobulin Assay; 100 determinations, 2 packs, 50 tests/pack. For the quantitative determination of thyroglobulin levels in human serum and plasma using the Access Immunoassay Systems. |  |  | July 28, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter, Inc. is recalling Access Thyroglobulin reagent packs because some may exhibit greater than 10% imprecision at concentrations &lt;10 ng/mL. | Under Investigation by firm | On 07/28/2015 the firm sent out customer notification letters via first class mail and e-mail. The firm's letter contained the following instructions:1. Discard all Access Thyroglobulin reagent pack lots listed in the letter.2. Evaluate Access Thyroglobulin results &lt;10 ng/mL when using the reagent pack lot numbers listed to determine if controls and/or patients were higher than expected or did not match the clinical status of the patient.3. At the discretion of the Laboratory Director, notify clinicians that it is possible the Access Thyroglobulin sample results reported by the laboratory were affected by this issue.The firm's notification letter requests that the information be shared with laboratory staff. If any affected products have been forwarded provide a copy of the letter to them. The firm requests that customers complete and return the Response Form within 10 days.Any questions contact the Customer Care Center:1. From website at http://www.beckmancoulter.com2. By phone: call 1-800-854-3633 in the United States and Canada.3. Outside the United States and Canada, contact local Beckman Coulter representative |  | 3035 | 3,035 (U.S.) and 7,123 (International) | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Algeria, Australia, Austria, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea, Republic of, Lebanon, Lithuania, Mexico, Morocco, New Zealand, Oman, Philippines, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Vietnam and Yemen. | Terminated | April 13, 2016 | 260 | 3035 | Not_Computer
Z-2484-2015 | 71651 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPodÂż Insulin Management System. Only the OmniPodÂż device is affected. Catalog Number(s): POD-ZXP420 (US). Intended for subcutaneous delivery of insulin at set and variable rates for the management of diabetes mellitus in persons requiring insulin. |  |  | July 13, 2015 | 2015 | Insulet Corporation | OmniPodsÂż (Pods) have a higher rate of failure causing:Cannula fails to deploy/retract causing insulin to be pumped or the audible alarm is displayed on the PDM and pod will not deliver insulin | Nonconforming Material/Component | The firm, Insulet Corporation, sent an "URGENT: Voluntary Device Recall" letter dated 7/13/2015 to each consignee/customer via Email notification (where possible) on 7/13/15 as well as Federal Express. Customers not responding to the email or Federal Express will receive follow up phone calls. The letter described the product, problem and actions to e taken. The customers were instructed to do the following: If you have Pods from any of these lots, set them aside, we will replace them at no charge; arrange for return and replacement by choosing one of the following methods:1. Complete the online form at https://www.myomnipod.com/podinfo (fastest method).2. Return the enclosed reply card by mail or fax to 1-855-407-3729.3. Contact Insulet Customer Care by calling 1-855-407-3729 (any time day or night).If you do not have any unused Pods form these lots, reply using one of the methods above.If you have any question, contact Acting VP RA/QA at 978-600-7000.EXPANDED RECALL:Insulet Corp. issued on 11/2/15 the URGENT: Field Safety Notification via Email notification and Federal Express. Letter describes problem of that certain lots of the pod's needle failed to deploy or there is a delay in the deployment of the needle. Customers not responding to the email or Federal Express will receive additional mailing and/or follow up phone calls. Call Customer Care at 1-855-407-3729 if you have any questions regarding this Field Safety Notification |  | 51059.9 | 24829 boxes Expanded: 26,230.9 boxes | Worldwide Distribution: US (Nationwide) and country of: Switzerland. | Terminated | January 26, 2017 | 563 | 51059.9 | Not_Computer
Z-2739-2015 | 71843 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CentaurÂż Folate Ascorbic Acid/Ascorbic Acid Diluent; SMN 10316877. |  |  | July 27, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | Low red blood cell control recoveries and low results for red blood cell patient samples when using the ADVIA Centaur Systems Folate assay. | Process design | An Urgent Medical Device Recall (UMDR) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on July 27, 2015 to be delivered to customers on July 28, 2015. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on July 27, 2015. These notices inform the customer to discontinue use of ADVIA Centaur Folate Ascorbic Acid/Ascorbic Acid Diluent kit lots 54942, 56230 and 70673and dispose of the remaining inventory per local regulations. The ADVIA Centaur Systems Folate red blood cell sample and control negative bias is minimal in the lower range of the assay and is more pronounced with increasing folate concentrations. The negative bias could lead to a misclassification of a patient as folate deficient. Siemens does not recommend a look back of previously tested patients. Fax-back forms will be used to confirm that the customers received the Urgent Medical Device Recall. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Recall and completed the required actions. |  | 466 | 466 | Worldwide Distribution-US (nationwide) including the states of CA CT DE FL GA IA IL IN KS KY MA MD MI MO NJ NM NV NY OH PA RI SC TN TX VA WA WI and OUS: Belgium Estonia Finland France Germany Great Britain Ireland Italy Pakistan Spain Sweden Switzerland Turkey. | Terminated | February 26, 2016 | 214 | 466 | Not_Computer
Z-2737-2015 | 71934 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DPM 7 Monitor, Multi Parameter Patient Monitor (with Arrhythmia Detection and Alarms) |  |  | July 27, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | The DPM 7 Monitors may display a black screen. | Under Investigation by firm | Mindray sent an Urgent Medical Device Correction / DPM 7 Monitor letter, dated July 27, 2015, to their affected customers via certified mail and return receipt. The letter identified the problem as well as the affected product, adverse effects on patients, and corrective action. The firm will repair the potentially affected monitor at no cost to the customer. Customers are to contact their Mindray Service Representative at 1-800-288-2121 to arrange for the repair. |  | 13 | 13 units | Distributed to the states of CT, IA, KY, MS, PA, UT and WA. | Open, Classified |  |  | 13 | Not_Computer
Z-2738-2015 | 71936 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.4, 3.0, and 4.0 mm Gridlock Screw Driver Bit, Model Number 310-30-003, Lot Number TSL002605, NON-STERILE. For use with the Gridlock Plating System for trauma and reconstructive procedures of small bones. |  |  | July 31, 2015 | 2015 | Trilliant Surgical Ltd. | Gridlock screw driver bits for the Gridlock Plating System were identified to be out of specification and have the potential for shallow engagement between the driver and the screw, resulting in an increased likelihood of stripping screw heads intraoperatively. Specifically, the minor diameter of the drive feature is above the specification, potentially reducing the engagement between the driver | Component design/selection | The firm sent out a written notification of the recall on 7/31/2015. |  | 108 | 108 units | NY, OH, CA, SC, NJ, AZ, VA, IN, FL, MN, OK, AR, MO, WI, GA, SD, IA, IL, TX | Open, Classified |  |  | 108 | Not_Computer
Z-2735-2015 | 72025 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MEDIVATORS HEMOCOR HPH700, HPH-1400TS, HPH-1000TS. |  |  | August 18, 2015 | 2015 | Medtronic Perfusion Systems | Hemocor High Performance Hemoconcentrators demonstrate low ultrafiltration performance that is below product specification. These potentially affected units have been distributed as stand-alone devices and in specific lots of Medtronic Perfusion Tubing packs. | Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent Medical Device Recall" letter dated August 2015. The letter was addressed to "Risk Manager". The letter described the problem and the product involved in the recall. Informed consignees that no action is needed for product that has been used, but patients should continue to be managed in accordance with their standard patient management protocol. For unused product they recommend to immediately quarantine and return the affected product to Medtronic. Contact Customer Service at 800-854-3570 and reference RADAR #253277 to initiated a return and credit of product. Requested consignees to complete the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. If the product has been forwarded to another facility, consignees are advised to notify that facility. For questions contact Medtronic Lifeline Technical Services at 877-526-7890 or your Medtronic representative. |  | 1594 | 1594 total (1306 US, 288 OUS) | Worldwide Distribution-US (nationwide) including the states of AL, AR, CA, CO, FL, HI, ID, IN, KS, LA, MA, MS, MO, NY, NJ, OH, PA, TN, TX, UT, and WI, and the countries of Bolivia, Denmark, Dominican Republic, Italy, Latvia, Lebanon, Mexico, Panama, Singapore, Venezuela, and Zambia. | Open, Classified |  |  | 1594 | Not_Computer
Z-2741-2015 | 71725 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | DimensionÂż CREA (Creatinine) and Dimension VistaÂż CREA (Creatinine); The CREA method used on the DimensionÂż clinical chemistry system is an in vitro diagnostic test intended for the quantitative determination of creatinine in human serum, plasma and urine. |  |  | July 15, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the Dimension CREA and Dimension Vista CREA may produce falsely depressed results up to - 12.7% and -22% respectively, for a creatinine concentration of 1.5 mg/dL when hemolysis exceeds 200 mg/dL and 300 mg/dL respectively. | Device Design | Siemens sent an Urgent Medical Device Correction letter, dated July 2015, was provided to all DimensionÂż CREA and Dimension VistaÂż CREA customers who received the products to notify them of the issue, the potential risk to health, and actions to be taken by customers.If you have any questions please contact your Siemens Customer Care Center Technical Solutions at (800) 441-9250, or your local Siemens Technical Support Representative. |  | 32165 | Total number Dimension Vista - 32,165, Dimension - 238,634 | Worldwide Distribution - US (nationwide) and to the countries of : Afghanistan, Algeria, Argentina, Armenia, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia Herzegovinia, Bulgaria, Brazil, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, France, Georgia, Germany, Ghana, Great Britain, Greece, Guatemala, Hungary, Iraq, Ireland, Honduras, Hong Kong, Iceland, Indonesia, India, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Lithuania, Libya, Macedonia, Malaysia, Mali, Mexico, Myanmar, Morocco, Netherlands, Nepal, Nicaragua, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tunisia, Turkmenistan, United Arab Emirates, Uruguay, Venezuela, Vietnam and Yemen. | Open, Classified |  |  | 32165 | Not_Computer
Z-2740-2015 | 71149 | [6 empty] | DermaPen Pro Model, Micro Needling System, a motorized AC powered surgical instrument for dermabrasive use. Each motorized handpiece is packaged within a small box. Each disposable needle cartridge is packaged individually and then placed into boxes of 30 quantities per box. |  |  | February 05, 2015 | 2015 | Derma Pen, LLC | Derma Pen initiated a field communication for Dermapen after receiving an FDA warning letter for distributing a device without FDA clearance or approval. | No Marketing Application | Derma Pen sent a notification letter to its customers on February 5, 2015. The purpose of the notification was to inform customers that the FDA issued a warning letter and the steps Derma Pen was taking to address the FDA' s concerns.For further questions please call (800) 981-5402. |  | 1466 | 1,466 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and The Virgin Islands and to the countries of : Canada, Brazil, Colombia, Ecuador, Costa Rica, Australia, Hong Kong, Netherlands, Mexico, Romania, Sweden, Peru, Greece, United Kingdom and United Arab Emirates. | Open, Classified |  |  | 1466 | Not_Computer
Z-2744-2015 | 71940 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Biomet Parsippany, NJ. 4 0 MM X 14 MM Variable Screw C-Tek Maxan Anterior Cervical Plate System. Lot 375440. |  |  | August 05, 2015 | 2015 | Biomet Spine, LLC | Some of P/N 14-521614B Lot 375440 was mislabeled as P/N 14-521614 Lot 375440. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall Notice dated August 5, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by customer. Customers were instructed to immediately locate and discontinue use of the affected product. This action requires the immediate location and discontinued use of the item and its immediate return to Zimmer Biomet Spine. Specifically, customers were REQUIRED to take the following steps:" Immediately locate and remove from circulation the items consigned/loaned to youraccount as identified on page two of this notice." Carefully follow the instructions on the enclosed "Response Form"." If you have further distributed these items to medical facilities, you MUST notify them of this action. This letter MUST be given to the person responsible for receiving recall notices. However, you are charged with the location and return of these items.Health care professionals and consumers may report serious adverse events or product quality problems with the use of this product to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail, fax or phone:Online:Fax:Phone:Regular Mail:www. fda .gov /MedWatch/report.htm(800) FDA-0178(800} FDA-1088Use postage-paid FDA form 3500 available at www.fda.gov/Medwatch/getforms.htmand mail to Medwatch, 5600 Fishers Lane, Rockville, MD 20852-9787.Questions related to this notice should be directed to 303-501-8400 Monday through Friday, 8 a.m.to 5 p.m. MST.Zimmer Blomet |  | 7 | 7 parts | Nationwide Distribution including Distributed nationwide to OH, LA, TX, and GA. | Terminated | February 10, 2016 | 189 | 7 | Not_Computer
Z-2742-2015 | 71937 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 500 Blood Parameter Monitoring System. |  |  | August 07, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Inaccuracies in SvO2, temperature, pH, pCO2, pO2, Hematocrit, and Potassium readings following a software upgrade to version 1.69. | Software design | A voluntary Urgent Medical Device Correction notice that clearly explains initial in-vivo calibration requirements, device operating ranges, and temperature measurements was sent on 08/17/2015, via express mail to consignees of CDI System 500 v1.69. Following the initial notice, Terumo will be updating the Operators Manual and will send the new manual to each consignee when available. One manual per unit at each facility will be provided. Customers with questions and return response forms may contact: Terumo Recall Email: tcvs.recall@terumomedical.comTerumo Recall Fax: 734-741-6149Terumo CVS Customer Service: 1.800.521.2818, Monday - Friday, 8 a.m.  6 p.m. ET. |  | 4638 | 4638 | Worldwide Distribution-US (nationwide) including DC and PR and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Australia, Belgium, Canada, Chile, Colombia, Hong Kong, Indonesia, Japan, Malaysia, Mexico, Philippines, Singapore, South Korea, Taiwan, Thailand, United Arab Emirates. | Open, Classified |  |  | 4638 | Computer
Z-2748-2015 | 71880 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Elcam Medical Minimal Residual Volume Luer activated Swabbable stopcock (MRVLS) for pressure monitoring and intravenous fluid administration; Manufactured by: Elcam Medical ACAL, Israel. Sterile, Quantity 20. |  |  | July 19, 2015 | 2015 | Elcam Medical, Inc. | Punctured blister packages were detected during the packaging process at the manufacturing site. Potential for device contamination. | Packaging | Elcam Medical sent "Urgent: Stopcock Recall" notices and "Recall Return Response Forms" dated 7/19/2015 to their customers. The recall notification informed the customers of the reason for the recall and the actions they must take. Customers were instructed to cease distribution and quarantine any affected products; return to Elcam Medical the enclosed Recall Return Response Form; and notify their customers as per their recall procedures. |  | 7721580 | 7,721,580 units | Texas, New Mexico and Pennsylvania | Open, Classified |  |  | 7969680 | Not_Computer
Z-2753-2015 | 72036 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Needle Driver; used with the da Vinci Si Surgical System (IS3000); |  |  | August 08, 2015 | 2015 | Intuitive Surgical, Inc. | In certain circumstances involving select da Vinci Si Single-Site instruments, the jaws of grip actuated instruments may become fixed in a closed position. | Device Design | Intuitive Surgical sent an Urgent Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Ensure that all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff and members of your medical staff who perform da Vinci Surgery procedures.2. Replace any instrument that shows signs of damage and contact Intuitive SurgicalCustomer Service.3. Complete and return the attached Acknowledgment Form to Intuitive Surgical usingthe instructions provided.4. Please retain a copy of this notice as well as the attached instructions with your Single-Site User Manual.If you need further information or support concerning this issue, please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail:customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 AM to 6PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST)For questions regarding this recall call 408-523-2100. |  | 89 | 89 total units | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, Canada, Chile, China, Cyprus, Denmark, France, Germany, Greece, Israel, Italy, Monaco, Netherlands, Panama, Romania, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | December 03, 2015 | 117 | 89 | Not_Computer
Z-2747-2015 | 71985 | No | NXT |  |  | Orthopedic | PMA | Smith &amp; Nephew BIRMINGHAM HIP (TM) RESURFACING DYSPLASIA CUP W/IMPACTOR, REF Numbers: 74120246 (46 MM), 74122248 (48 MM), 74120250 (50 MM), 74122252 (52 MM), 74120254 (54 MM) |  |  | June 03, 2015 | 2015 | Smith &amp; Nephew, Inc. | The firm observed that the BIRMINGHAM HIP(TM) Resurfacing (BHR) System had revision rates which were higher than established benchmarks. | Device Design | Smith &amp; Nephew sent an notification letter dated June 3, 2015 to customers via E-mail and Fed- Ex. The letter identified the affected product, problem and actions to be taken. Customers were asked to quarantine the product and return it. Distributors were asked to notified their customers of the recall and ask them to return the product. |  | 624 | 624 units | Worldwide Distribution - US Nationwide and countries of: Australia, Austria, Bahrain (UAE), Belgium, Canada, China, Denmark, Finland, Germany, Hong Kong, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand and United Arab Emirates (UAE). | Open, Classified |  |  | 16778 | Not_Computer
Z-2759-2015 | 71877 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Biomet Spine Lineum OCC Starter Flexible and Lineum Flexible Inner Shaft, Model Numbers 14-525050 and 14-525057, reusable devices distributed with bubble wrap inside a plastic pouch. |  |  | July 22, 2015 | 2015 | Biomet Spine, LLC | Biomet has initiated a recall on Biomet Spine Lineum OCC Starter Flexible and its inner shaft instrumentsafter identifying that these instruments were distributed prior to completion of all design control activities. | Device Design | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated 7/23/2015 to their Consignees/Customers via certified letter. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of device; remove from circulation; carefully follow instructions on enclosed "Response Form" and complete and return via fax to 303.501.8444; if you have further distributed these devices to medical facilities, you MUST notify them of this action; and return device and form to Zimmer Biomet Spine, 310 Interlocken Pkwy, Ste 120, Broomfield, CO 80021. Questions related to this notice should be directed to the Quality Engineer at 303-501-8457 Monday through Friday, 8 a.m. to 5 p.m. MST. |  | 49 | 49 | Worldwide Distribution: US (nationwide) in states of: NY, MO, WI, IN, CA, NC, and OH, and to countries of: Japan and Netherlands. | Terminated | December 10, 2015 | 141 | 49 | Not_Computer
Z-2761-2015 | 72122 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hemocor High Performance Hemoconcentrator, HPH700, sterilized with ethylene oxide |  |  | August 03, 2015 | 2015 | Medivators, Inc. | Hemoconcentrators may exhibit low ultrafiltration performance that is below product specification. | Process control | Medivators sent an "Urgent Medical Device" recall letter, dated August 4, 2015, via email. The letter described the issue, identified the affected lot, requested the product be quarantined, and to contact Medivators for return instructions and credit or no-charge replacement product. A response form was requested to be returned to Medivators. Customers with questions can contact Medivators Product at 763-553-3348. |  | 840 | 840 units | US Distribution to the states of CA and MI. | Terminated | November 04, 2015 | 93 | 840 | Not_Computer
Z-2762-2015 | 72019 | No | NAB | Nonresorbable gauze/sponge for external use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | XGAUZE Trauma Dressing, catalog number FIN-0002-01. |  |  | August 12, 2015 | 2015 | Revmedx Inc | The XGAUZE Trauma Dressing was recalled because the packages may acquire small tears or pinholes during shipping due to use of shipping containers that lacked adequate protective materials. If present, these small tears or pinholes may disrupt package integrity and render the device unsterile. | Package design/selection | RevMedx, Inc. sent the Urgent Medical Device Correction and Removal XGAUZE Trauma Dressing letter to US Customers on August 12, 2015.Customers are instructed to inspect packages for signs of damages or loss of vacuum. The affected packages will appear puffy and the dressing will be loose inside. The non-affected packages will have a good vacuum intact seal and the packaging material should contour snugly against the dressing.Customers are advised to return any packages appear to have damaged or loss of vacuum. Any undamaged packages may be used as indicated.Customers with question can call the firm at (503) 218-2172 and fax the Return Response Form via fax to (503) 218-2274 or send via email to info@revmedx.com. |  | 5004 | 5004 units | US Distribution to the states of : Arizona, Idaho, North Carolina, Tennessee and Virginia. | Terminated | September 17, 2015 | 36 | 5004 | Not_Computer
Z-2760-2015 | 72000 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | UCLA Abutment Hexed Castable Cylinder 3.4mm |  |  | July 29, 2015 | 2015 | Biomet 3i, LLC | MUCC2, non-hexed, castable abutment cylinders, has been labeled and distributed as a MUCC1C, hexed, castable abutment cylinders for this one lot only. | Process control | Biomet sent a Medical Device Recall Notice dated July 27, 2015, to all their clinicians informing them about the recall and to complete the attached recall return response form and any unused product to be replaced at no charge. Contact information 1-800-342-5454 |  | 61 | 61 pieces | Worldwide Distribution including Italy and Uruguay. | Open, Classified |  |  | 61 | Not_Computer
Z-2764-2015 | 71953 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific autoHDL Reagent H7545 |  |  | March 13, 2015 | 2015 | Pointe Scientific, Inc. | Crystals may be observed in the R1 reagent. Crystals may be a result of the reagent being frozen during transport. The crystals do not impact product performance. | Nonconforming Material/Component | Pointe Scientific sent a Field Correction Notice on March 13, 2015, to all affected customers via fax, email, and mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Any questions concerning the corrective action may be directed to Technical Service Department at 1-800-757-5313. |  | 40280 | H7545-R1, 8.75 L; 12-H7545-162, 4 kits; H7545-1000, 3 kits; HDL600, 6 kits | Worldwide Distribution - US including MI, FL, CA, TX, CT, PR and Internationally to Canada. | Terminated | October 29, 2015 | 230 | 40280 | Not_Computer
Z-2763-2015 | 71935 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | CIVCO Shoulder Retractors, Reusable non-sterile hand grip, REF 302070, packaged 1 set per box, RX. |  |  | August 19, 2015 | 2015 | Med Tec Inc | The wrist straps used for positioning the patient are too long | Vendor change control | Med Tec Inc sent an Urgent Field Correction Notice dated August 18, 2015, via e-mail on August 19, 2015, and made phone calls on August 19, 2015, to all affected customers requesting return of the product. Customers were instructed to forward a copy of the notification letter to all customers who have received the product. Customers with questions were instructed to reach their representative by calling CIVCO Medical Solutions, adiation Oncology, at 800-842-8688 or 319-248-8688 between 8:00AM and 5:00PM CST. For questions regarding this recall call319-248-6502. |  | 18 | 18 sets | Worldwide Distribution - US including NE, SC, and KS. Foreign distribution was made to Australia, China, Venezuela, Spain, Turkey, and Hong Kong | Open, Classified |  |  | 18 | Not_Computer
Z-2769-2015 | 71956 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Voyant Open Fusion Device |  |  | July 31, 2015 | 2015 | Applied Medical Resources Corp | Potential to improperly characterize tissue within the jaws, which may result in nonoptimal power delivery. | Device Design | The firm sent out customer notification letters to International customers the beginning of July 31st, 2015 and domestic customers the week of August 3rd, 2015. The firm's letter requested customers complete the following actions:- check inventory for recalled product- complete the attached Recall Notification Confirmation Form to acknowledge the recall and indicate if facility is returning or has already used any of the products- Return the confirmation form(s) to Applied Medical by emailing torecall60669166@appliedmedical.com or faxing to (949) 713-8936.- Return affected product and a copy of the confirmation form(s) to Applied MedicalFor product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by e-mail at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or bye-mail at lcontursi@appliedmedical.com. |  | 101 | 101 units (U.S) and 38 units (International) | Worldwide Distribution - U.S Nationwide in the states of:FL, GA, NV. and countries of: France, Netherlands, Spain, and United Kingdom. | Terminated | November 20, 2015 | 112 | 101 | Not_Computer
Z-2768-2015 | 71993 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MedTest DX Control Reconstitution Fluid |  |  | May 15, 2015 | 2015 | MedtestDx, Inc. | Vial to vial variation in the fill volume that could affect control recovery issues and cause a control result to fall outside the suggested recovery ranges. An under-fill of the vials was found ranging as low as 4.63 ml. When used to reconstitute HNC/HEC, control values were shown to be out of range running on the high end. | Process design | On 5/15/2015, MedTest DX Field Correction Notice notifications were sent to the affected customers via electronic mail and US mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Any questions concerning the corrective action may be directed to MedTest Technical Service Department at 1-800-757-5313. |  | 79 | 79 kits | US Nationwide Distribution. | Terminated | May 23, 2016 | 374 | 79 | Not_Computer
Z-2767-2015 | 72015 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | NICO Myriad Console |  |  | July 29, 2015 | 2015 | Nico Corp. | Two cables within the console were inadvertently twisted during manufacturing. The individual wires within the cables may break and/or contact one another where the cables terminate into two connectors. When these wires break or contact one another (i.e., short), various types of malfunction may occur with the console. One type of malfunction may result in unexpected aspiration increase. | Process control | The firm, Nico Corporation, sent an "URGENT: NICO Console Recall" notification dated 7/31/15 via fax to its customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The customers were instructed to return the console and complete and return the Recall Response Form via fax to: 317-683-0305 or email to: Jay.Dittman@niconeuro.com. The firm will provide customers with a Myriad Console loaner kit.If you have any questions, contact Director of Quality and Regulatory, at 317-660-7118 ext 104 or email to: jay.dittman@nicneuro.com. |  | 6 | 6 units | US Distribution to states of: MD, IN, CA, and NC. | Terminated | September 22, 2015 | 55 | 6 | Not_Computer
Z-2775-2015 | 71994 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiographic Injector. Medrad Mark 7 Arterion Injection System, which utilizes Software Version SW 005.006_SH; to be used specifically for the purposes of injecting contrast medium and common flushing solutions into humans for angiographic studies. |  |  | August 04, 2015 | 2015 | Bayer Healthcare | Software Version SW 005.006_SH, resulted in the removal of purge enforcement from traditional New-Case, Power Up and Syringe Change use cases while the injector head is in the upright position. | Software design | Direct accounts were notified by UPS Notification letter dated August 4, 2015 and advised when using the Arterion Injection System to ensure that purges are completed to avoid the injection of air. |  | 71 | 71 | Worldwide Distribution. US Nationwide, Germany, France, UK, Netherlands, Sweden, and Canada. | Open, Classified |  |  | 71 | Computer
Z-2772-2015 | 72054 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Aliquot Delivery Kit (Syringe and Plunger); for use in open and percutaneous procedures to facilitate placement of bone cement in weakened or diminished bone |  |  | August 20, 2015 | 2015 | Orthovita, Inc., dBA Stryker Orthobiologics. | Two isolated instances of holes in the outer header bag component of the double barrier sterile packaging configuration | Under Investigation by firm | Stryker Sent an Urgent Product Recall letter dated August 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken by the Customer/User:Our records indicate that you have received the above referenced instruments. It is Stryker Orthobiologics responsibility as the manufacturer to ensure that customers who may have received these affected instruments also receive this important communication. Please assist us in meeting our regulatory obligation by faxing back the attached Product Recall Acknowledgment Form as soon as possible to 215-253-5020.We regret any inconvenience this action may cause you and if you have any questions, please contact Stryker Orthobiologics at 610-640-1775 x 5299. |  | 9805 | 9805 | Worldwide Distribution - Nationwide Distribution and to the countries of : Australia, Belgium, Denmark, France, Germany, Italy, Malta, Netherlands, South Africa, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open, Classified |  |  | 9805 | Not_Computer
Z-2771-2015 | 71657 | No | FWB | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sony LMD-1951MD, Sony LMD1951MD/BS, Sony LMD-1951MC (Foreign) 19 inch medical grade liquid crystal display (LCD) Medical Monitors with light-emitting diode (LED) backlight technology. |  |  | January 13, 2015 | 2015 | Sony Medical Systems | Tthe medical monitor has either not turned on or has lost image during clinical use. | Device Design | Sony Corporation sent an URGENT: MEDICAL DEVICE FIELD ACTION field letter dated January 13, 2015. Customers are asked to affix the warning sticker to the rear of the monitor and place the adhesive seal over the AC inlet. If the firm does not possess an AC adaptor (AC-110MD) for each of the affected monitors, stop using the monitor and take the monitor out of service until the permanent corrective action described below is implemented.SEL will be replacing the power board of all affected units with a redesigned power board that does not exhibit the appropriate failure mode. Questions can be directed to 866-789-SONY (7669) between 8:00 am and 8:00 pm Monday through Friday. |  | 4952 | 4952 units (1016 US; 3936 OUS) | Worldwide Distribution - US Distribution to the states of : AZ, CA, NC, PA, IA, OH and FL., and to the countries of : China, Japan, Taiwan, South Korea and Indonesia. | Open, Classified |  |  | 4952 | Not_Computer
Z-2770-2015 | 71851 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Lumenis VersaCut+ Tissue Morcellator GA-0007500 |  |  | July 30, 2015 | 2015 | Lumenis | There is a probability the hand piece may operate inconsistently and unpredictably during the morcellation procedure resulting in damage to non-target tissue. | Device Design | Lumenis sent an Urgent Medical Device Recall letter on August 7, 2015, to all consignees. Consignees were instructed to destroy the operator manual, and return the affected device to Lumenis along with the completed Recall Verification form. Consignees with questions were instructed to call 801-656-2663. |  | 162 | 54 devices with 108 hand pieces | Worldwide Distribution - US including CA, PA, MA, FL, and NY, and Internationally to Italy, Germany, Australia, China, Great Britian, United Kingdom, Taiwan, France, Peru, Mexico, Turkey, Spain, and Russia. | Open, Classified |  |  | 162 | Not_Computer
Z-2776-2015 | 71971 | No | DXO | Catheter tip pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Edwards Pressure Monitoring Kit with TruWave disposable pressure transducer. A sterile, single-use kit that monitors pressures when attached to pressure monitoring catheters. For use on patients requiring intravascular, intracranial, or intrauterine pressure monitoring. |  |  | July 31, 2015 | 2015 | Edward Lifesciences, Llc | Flow restriction issues in a trifurcated IV set included in the pressure transducer kits. A restriction of the flow of IV fluid would result in either an inability to prime the system before use or difficulty in flushing the line after a blood draw. | Device Design | Edwards sent an Urgent Product Recall Letter on August 10, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm's letter requests that customers return any of the units that they currently have in inventory. The letter requests that customers verify the inventory that has been pre-populated on the acknowledgment form, and complete the remainder of the form. The firm states to fax back the acknowledgment form within three days of receipt of notification to Edwards Customer Service at 1-800-422-9329. The firm requests to call Edwards Customer Service at 1-800-424-3278 to arrange for return of any unused product from the implicated lots as well as to obtain information about product replacment. The notification states to call for an RGA prior to returning any unused product and return unused product to the firm at Edwards Lifesciences Return Goods. Any questions the firm states to contact Edwards Customer Service at 1-800-424-3278. |  | 13077 | 13,077 (U.S.) and 12,306 (International) | Worldwide Distribution - US nationwide, Australia, Austria, Canada, China, Colombia, Croatia, Denmark, France, Germany, Great Britain, Italy, Netherlands, New Zealand, Sweden, and Slovenia. | Open, Classified |  |  | 13077 | Not_Computer
Z-2774-2015 | 71887 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX OPA Solution, P/N 20390 for use as a high level disinfectant for reprocessing heat sensitive medical devices. |  |  | August 10, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is reminding all CIDEX OPA Solution customers to follow the warnings and rinsing procedure in their current labelling around the criticality of proper rinsing of medical device following disinfection with CIDEX OPA according to the CIDEX OPA Solution IFU (Instruction For Use). | Under Investigation by firm | A Urgent Medical Device Field Safety Notice will be sent to customers on 8/10/15 to inform them to follow the warnings and rinsing procedure in their current labelling around the criticality of proper rinsing of medical device following disinfection with CIDEX OPA according to the CIDEX OPA Solution IFU (Instruction For Use). The letter provides the customers with the description of the issues and the recommended actions to be taken. Customers are instructed to complete and return the attached Business Reply Card (double-sided) via mail or by fax to Stericycle. For questions regarding the business reply card, customers are instructed to contact Stericycle. Customers with any complaints or suspected problems with CIDEX OPA Solution related to the field safety letter, are instructed to contact 1-888-783-7723. |  | 318632 | 318,632 cases | US | Open, Classified |  |  | 318632 | Not_Computer
Z-2773-2015 | 71974 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Creatinine Reagent |  |  | December 12, 2014 | 2014 | Pointe Scientific, Inc. | Creatinine kit C7539-150 contains Creatinine reagent R1 that is labeled as Creatinine reagent R2 (30ml) and reagent R2 vial is labeled as R1 (120ml). With manual procedure, no result will be reported as the control will be out of specifications. | Process control | On 12/12/2014, Field Correction Notice notifications were sent to the affected distributors via electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Any questions concerning the corrective action may be directed to Pointe Scientifics Technical Service Department at 1-800-757-5313. |  | 93 | 93 kits | Worldwide distribution. US nationwide, Singapore, Kenya, Bolivia, Mexico, Greece, Guyana, and Philippines. | Terminated | October 29, 2015 | 321 | 93 | Not_Computer
Z-2779-2015 | 71910 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | The Leica ASP200S/ASP300S Tissue Processor. Pathology and Hematology Departments. |  |  | August 05, 2015 | 2015 | Leica Microsystems, Inc. | The Leica ASP200S/ASP300S Tissue Processor have an incorrect internal wiring to the remote alarm plug, which can cause the remote alarm to not function correctly in the case of an instrument failure. | Nonconforming Material/Component | Leica Biosystems sent an Urgent Medical Device Recall Notification letter dated August 4, 2015, to all affected customers to provide instructions on what the customer should do with the recalled product. These include 1) As an immediate countermeasure, the firm recommends that the customers do not use the remote alarm function on the Leica ASP200S/ASP300S in its current state. 2) The customers local Leica representative will contact them to arrange for a Leica service engineer to visit their facility to correct the internal wiring on their tissue processor and to check the remote alarm function. 3) The Medical Device Recall Notification needs to be passed on to all those who need to be aware within the custromers organisation or to any organisation where the potentially affected devices have been transferred. Please maintain awareness of this Medical Device Recall Notification and resulting action to ensure effectiveness of the corrective action. 4) The letter asks the customer use the enclosed Medical Device Recall Notification Response Form to confirm that they have received and understand the Medical Device Recall Notification. it states "Please confirm receipt of this notice by signing and dating the attached Acknowledgement Form and faxing it back to us at 1-847-236-3747, or scan and email it to LMGRA@leica-microsystems.com" |  | 901 | 901 | Worldwide Distribution - US including U.S. including AZ, CA, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MN, MO, MT, NC, NJ, NM, NY, OH, OR, PA, TX, VA, WA, WI, WV and District of Columbia. Internationall to AE, AT, AU , BD , BE, BF, BG,BR, CA, CL, CN, CO, CR, DE, DK, EC, EG, ES, ET, FI, FR, GB, HU, ID, IN, IT, JM, JP, KR, LK, MA, MT, MY, NL, NZ, PA, PE, PH, PL, PT, RO, RU, SA, SE, SG, TH, TR, TW, VN and ZA. | Terminated | January 05, 2017 | 519 | 901 | Not_Computer
Z-2781-2015 | 71909 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | NebuTech Nebulizer, Model number: Series 8660, 8960, 8961, 8966, 8984. Nebulizer used to generate aerosols delivered directly to the patient for breathing. Designed to be used in either a hospital or homecare environment. |  |  | August 04, 2015 | 2015 | Salter Labs | The Inspiratory valve cap on the nebulizer reservoir tower may be loose and may detach before or during treatment. | Component change control | Salter Labs sent a Medical Device Recall letter dated July 24 2015. The letter identified the affected product, problem and actions to be taken. The letter states that all affected product can be returned to Salter Labs for replacement at no charge with product specifically inspected for a tighter fit on the inspiratory valve cap. Customers were instructed to complete and return the Return Response Form as soon as possible. For any questions contact Salter Labs Customer Support at 800-421-0024 or the local Salter Labs representative. |  | 8975 | 8,975 | Worldwide Distribution - US Nationwide and in the countries of Greece, UK, Belgium, Australia, and Switerzland. | Terminated | May 06, 2016 | 276 | 8975 | Not_Computer
Z-2777-2015 | 72136 | No | LPB |  |  | Cardiovascular | PMA | Medtronic CryoConsole, models 106A3, 106E2, and 106A2-K |  |  | September 01, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | An issue with a USB memory component contained within a subset of CryoConsoles can result in extended procedure time. | Component design/selection | Medtronic began hand delivering an Urgent Medical Device Correction letter on September 1, 2015, to all affected customers. The letter described the issue. Customers were advised to continue using the device, per the operator's manual, and if a system notice appears, they are to carefully read and follow the instructions in the screen of the console. Medtronic personnel will track contact with the customer via Customer Contact Record. Customers with qeustions were instructed to contact their Medtronic AF Solutions representative.For questions regarding this recall call 1-877-464-2796 Technical Services. |  | 103 | 103 units | Worldwide Distribution - US (nationwide) and Internationally to Algeria, Australia, France, Germany, Greece, Iceland, Ireland, Israel, Italy, Lebanon, Malaysia, Martinique, Netherlands, Poland, Romania, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | April 08, 2016 | 220 | 103 | Not_Computer
Z-2782-2015 | 72158 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Fluorescence Imaging Procedure Kits used with the da Vinci Si and Xi Surgical Systems; |  |  | September 09, 2015 | 2015 | Intuitive Surgical, Inc. | B. Braun 2 gang 4-way stopcocks in the FluorescenceImaging Procedure Kit may experience damage to the sterile packaging during shipping. The damage canresult in tears to the packaging and has the potential to cause a breach in sterility of the supplied stopcocks. | Package design/selection | Intuitive sent an Urgent Medical Device Correction letters dated September 9, 2015 to all customers via traceable method. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to :1. Ensure that all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, BiomedicalEngineering staff, anesthesiologist and members of your medical staff who perform da Vinci Surgery procedures.2. Prior to use, open and inspect current supply of Fluorescence Imaging Procedure Kits. If kits include B. Braun stopcocks, discard stopcocks and replace with other stopcocks3. Continue to perform procedures using other stopcocks4. Continue to follow directions provided in Florescence Imaging Procedure Kit IFU (p/n 550976), If a sterile item in a Fluorescence Imaging Procedure Kit becomes compromised in any way, do not use the compromised item; replace it with a new item.5. Complete and return the attached Acknowledgment Form to Intuitive Surgicalusing the instructions provided.6. Please retain a copy of this notice near your Fluorescence Imaging Procedure Kit inventory.Questions or concerns should be directed to ClinicalSales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail: customersupportservicesupport@intusurg.com |  | 1200 | 1200 kits | US Nationwide Distribution | Terminated | November 02, 2015 | 54 | 1200 | Not_Computer
Z-2783-2015 | 72094 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System with software versions VB10B and lower; |  |  | August 19, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | The ACUSON SC2000 ultrasound system considers uppercase/lowercase differences in the same patient name as unique patient instances when registered on the same ultrasound system. If these differences are not corrected at the time of registration, the system does not capture images or clips. | Software design | Siemens sent a Customer Safety Advisory Notification letter on August 19, 2015 by certified mail. The letter identified the affected product, problem and actions to be taken. The letter states that this issue will be resolved in a future software release. If you have any questions,please contact your customer care center for information regarding timelines and status. Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. |  | 2039 | 2039 devices | Worldwide distribution. US Nationwide and the countries of Albania, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Canada, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Hungary,, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, P.R. China, Pakistan, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., San Marino, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2039 | Computer
Z-2731-2015 | 71905 | No | LOZ |  |  | Cardiovascular | PMA | Temporary Total Artificial Heart (TAH-t) with Freedom Driver System |  |  | August 06, 2015 | 2015 | SynCardia Systems Inc. | SynCardia Systems, Inc. is recalling Freedom Drivers due to a specific component of the drive mechanism may fail and cause the drive mechanism to stop pumping; resulting in the loss of life-sustaining function. | Component design/selection | SynCardia sent an Urgent Medical Device Recall letter dated August 6, 2015 to affected customers. The letter identified the purpose, reason for the voluntary recall, risk to health, how to recognize that the device may fail, actioms to be taken by the customer/user, product and distribution information and type of action by the company. Customer were instructed to completed and return Recall Acknowledgement and Receipt Form to SynCardia. A SynCardia Clinical Support Specialist or other SynCardia representative will assist customers with replacement product and returning affected product.Contact Information For QuestionsMary Pat SloanSenior Vice PresidentGlobal Certification and LogisticsSynCardia Systems, Inc1992 E Silverlake RdTucson, AZ 85713Direct 520-547-7470Toll-free 866-480-1122 Ext 1308Monday-Friday 7:30 AM to 5:00 PM Pacific/Arizona timeHotline 24/7 866-771-9437www.syncardia.com |  | 56 | 56 drivers (21 currently in use on patients) | Worldwide Distribution -- US Nationwide in the states of AR, NC, CA, OH, PA, TX, MI, and VA; and the countries of Germany, Turkey and France. | Terminated | November 04, 2015 | 90 | 56 | Not_Computer
Z-2778-2015 | 71695 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Medical Endoscope for the duodenum and upper G.I. tract. ED-530XT, ED-250XT5, ED-250XL5, ED-450XT5, and ED-450XL5 Endoscope Operation Manuals. |  |  | June 30, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | The suction button on the ED-530XT endoscope may become lodged (stuck) in the active position following the use of x-ray contrast agents during ERCP procedures. | Device Design | FujiFilm sent a Voluntary Field Correction/Verification Form dated June 30, 2015 to their affected customers. |  | 256 | 256 units | Distributed in the US and Canada. | Terminated | May 25, 2017 | 695 | 256 | Not_Computer
Z-2785-2015 | 71477 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System software |  |  | May 29, 2015 | 2015 | RAYSEARCH LABORATORIES AB | An error caused by a floating point precision problem tends to cause a problem with the display of isodose lines for dose and dose difference and color tablesfor dose, dose difference and PET images | Software design | The firm, RaySearch Laboratories, distributed a "Field Safety Notice, Medical Device Correction #7106" notice dated May 29, 2015,and reply forms to their customers via email. The notice describes the product, problem and actions to be taken. The Customers were provided a workaround as a temporary fix. Customers were also instructed to educate all users, inspect your product and identify all installed units, confirm you have read and understood the notice by contacting Quality and Regulatory Affairs Manager; and complete and return the reply form via email to freddie.cardel@raysearchlabs.com or by fax to 888-501-7195. The solution will be resolved in the next version of RayStation, scheduled for release in June 2015. Customers were advised to share the notice to all those who need to be aware within the organizations. Customers can contact Quality and Regulatory Affairs Manager at +46 8 510 530 12 or via email at david.hedfors@raysearchlabs.com. |  | 20 | 20 devices | Worldwide Distribution: US (nationwide) to state of WA; and countries of: France, Sweden, Austria, Spain, Germany, Italy, UK, New Zealand. | Terminated | March 13, 2017 | 654 | 20 | Not_Computer
Z-2780-2015 | 72046 | No | DPT | Extravascular blood flow probe. | Cardiovascular | Cardiovascular | 510(k) | ThruPort Knot Pusher. This device is sold individually (Model KP1) and as part of the ThruPort Valve Placement Pack (Model VPP). packaged sterile and non-pyrogenic in a sealed, peel-type pouch. |  |  | August 26, 2015 | 2015 | Edwards Lifesciences, LLC | The configuration of the slot at the tip of the knot pusher may inhibit intended ease of use. Fraying or splitting of the suture could occur, which could compromise the integrity of the suture knot. | Process control | Edwards sent an Urgent-Product Recall - Action Required letter dated August 26, 2015, to all affected customers. Letters were sent via Fed-Ex. Please check your inventory for the affected Lot Numbers listed above and return the enclosed confirmation form. If your inventory includes affected product, please contact Edwards Lifesciences to obtain an RGA number and return the affected product.We apologize for the inconvenience caused by this action and appreciate your attention to this matter. If you have questions that have not been answered by this letter, please call Edwards Customer Service at (800) 424-3278 from 6:00AM  4:30PM Pacific Time. |  | 759 | 759 units | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Canada, Japan and Europe. | Open, Classified |  |  | 759 | Not_Computer
Z-2796-2015 | 71933 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5mm Non Locking Cortical Screw, in the PediLoc Tibial Plate System |  |  | July 28, 2015 | 2015 | OrthoPediatrics Corp | The 00-0903-2510, 3.5mm x 10mm Non Locking Cortical Screw, Lot # 150192-J package may contain a different screw. The screw in the package is a 00-0903-2526, 3.5mm x 26mm, and is etched correctly. | Packaging process control | OrthoPediatrics has initiated a Voluntary RECALL of Part # 00-0903-2510, 3.5mm Non Locking Cortical Screw, Lot # 150192-J, initially by contacting customers via telephone, and then by e-mailing notification letters on 7/28/2015. Customers returned the 5 mislabled screws, which were repacked with correct labeling. All recall activities were complete as of 8/14/2015. |  | 5 | 5 | Domestic: TN, IAForeign: EnglandVA/DOD: None | Terminated | March 17, 2016 | 233 | 5 | Not_Computer
Z-2794-2015 | 71976 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Dynamic Locking Screw System (DLS) 3.7mm &amp; Dynamic Locking Screw (DLS) 5.0mm; intended for use in long bone fractures, the fixation of osteopenic bone, the fixation of osteotomies, and for the fixation of non-unions and malunions. |  |  | August 11, 2015 | 2015 | Synthes (USA) Products LLC | DLS pin breakage during planned implant removal, after uneventful and successful healing of the fracture. This notice to additional consignees identified after recall Z-0390-2014, related to complaints of breakage located at the bottom of the screws and under the locking screw head. | Labeling False and Misleading | Synthes sent an Urgent Notices of Medical Device Recall letter, dated August 11, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to return affected devices and response forms.For further questions please call (610) 719-6500. |  | 40038 | 40038 | Nationwide Distribution | Open, Classified |  |  | 40038 | Not_Computer
Z-2793-2015 | 71957 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | Microtek Medical, Inc. ORS Fluid Warming and Slush Drapes, Model #'s ORS-100, ORS-110, ORS-130, ORS-300, ORS-301, ORS-320, ORS-321, ORS-325, ORS-330, ORS-331 |  |  | August 18, 2015 | 2015 | Ecolab Inc | The sterile packaging may contain small (less than 250 microns wide) channels in the seal of the pouch. If channels are present, the package integrity is compromised and there is a remote chance that a non-sterile drape may be used during a surgical procedure. | Package design/selection | Consignees were sent on 8/18/2015 an Ecolab "Urgent Medical Device Recall" letter dated August 18, 2015. The letter was addressed to Recall Coordinator or Risk Manager. The letter described the problem and the product involved in the recall. Advised consignees to quarantine all affected lots, and notify their customers of the recall which is to be carried out to users level. Requested consignees to complete and return the enclosed Recall Response Form. For questions they can contact Customer Service at 1-800-824-3027 or e-mail customerservice@microtekmed.com. |  | 5172816 | 5,172,816 units | Worldwide Distribution -- US, including DC and PR; and, the countries of SAUDI ARABIA, LEBANON, GERMANY, HONG KONG, SINGAPORE, CANADA, and JAPAN. | Terminated | August 09, 2017 | 722 | 5172816 | Not_Computer
Z-2791-2015 | 72010 | [6 empty] | Terumo Custom Cardiovascular Procedure Kit X-Coated |  |  | August 24, 2015 | 2015 | Terumo Cardiovascular Systems Corp | Custom Cardiovascular Procedure Kits contain recalled component Medivators HPH700 High Performance Hemoconcentrators | Nonconforming Material/Component | Terumo CVS issued an Urgent Medical Device Recall Notice dated August 24, 2015, to consignees via express mail. The letter advised users to immediately quarantine affected Cardiovascular Procedure Kits and stop using the Hemocor HPH700 Hemoconcentrator within the kit.Replacement sterile, stand-alone Hemocor HPH700 Hemoconcentrators will accompany the letter so that users may replace the affected emoconcentrators at the point of use.The option of replacing the hemoconcentrator at the point of use is offered to facilitate imminently scheduled cardiopulmonary bypass procedures. Users are requested to return the unused hemoconcentrators to Terumo CVS. Alternatively, users may return the entire Cardiovascular Procedure Kit to Terumo CVS for replacement. Customers with questions were instructed to call 763-553-3348. |  | 19 | 19 kits | Nationwide Distribution including CA, CT , FL , WI. | Terminated | February 05, 2016 | 165 | 124 | Not_Computer
Z-2797-2015 | 72079 | No | EGS | Dental handpiece and accessories. | Dental | Dental | 510(k) | Contra-Angles CA1:1L EVO 15 |  |  | August 27, 2015 | 2015 | Bien Air Dental SA | Possible loosening of the dental tool, which may lead to potential swallowing of the tool. | Other | Bien Air Dental sent an Urgent Medical Device Safety Alert on August 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the mentioned device.The letter requests that customers inspect their inventory for the product those and affected within the scope, and immediately quarantine them. The firm states to fill out the attached REPLY FORM and returned it fax to +41 32 344 64 91.If customers have further distributed any of the subject devices to other organizations, please forward this note to them immediately, to stop using the affected devices and to place them in quarantine. All the devices at customer or end-user locations shall be also returned to Bien Air.Please return the affected products to Bien-Air Dental (the exact address for return will be communicated to you via Bien-Air Sales Manager). A Bien-Air Sales Manager will also be in contact with customers.Upon receipt of the completed REPLY FORM, a Bien-Air Dental Representative will contact customers to arrange for the return and/or replacement of any non-conforming units.Should customers require any further information or have any queries on the matter please contact Bien-Air Dental SAs Regulatory Affairs at: +41 32 344 6464 |  | 112 | 112 (U.S.) | Worldwide Distribution - US distribution to one distributor in FL. | Terminated | March 31, 2016 | 217 | 112 | Not_Computer
Z-2800-2015 | 72038 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PenBlade Safety Scalpel, for the incision of human skin and tissue, and to cut or trim sutures and catheters up to 6Fr. Models PB-M-10, PB-M-11P, PB-M-15 |  |  | August 18, 2015 | 2015 | Zien Medical Technologies, Inc | Cracks were discovered in the PET tray of a small population of sterile product. These cracks create a risk of loss of sterile barrier for the product. Sampling and root cause analysis indicated that the risk could extend throughout the lots manufactured from September 2014-July 2015. This led to the decision to recall the product. | Packaging process control | All consignees have been notified via email with an attached copy of the recall letter. Telephone calls with be made to all customers and distributors that do not respond directly to email notification to ensure compliance to the recall. |  | 28284 | 28,284 | Worldwide distribution to U.S. nationwide, Germany and England. | Terminated | December 09, 2015 | 113 | 28284 | Not_Computer
Z-2802-2015 | 71842 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | PET Discovery LS |  |  | July 17, 2015 | 2015 | GE Healthcare | Potential for expulsion of the Power Supply on Lightspeed CT scanner or Discovery ST PET CT or Discovery LS PET CT scanner. | Process control | GE Healthcare sent an "Urgent Medical Device Correction" letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Ref# 25463 |  | 6 | 6 | Worldwide Distribution - US (nationwide) and to the countries of : ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, HUNGRY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MEXICO, NORWAY, POLAND, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN and UNITED KINGDOM. | Terminated | March 08, 2016 | 235 | 440 | Not_Computer
Z-2798-2015 | 72008 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Regard Custom Surgical Pack, Item Number: 800628; Sterile, CV0805 - CABG B. Used in cardiovascular surgical procedures. |  |  | August 19, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | Surgical kits contain a surgical scrub product which may be chemically contaminated and lack sterility assurance. | Material/Component Contamination | On 8/19/2015, the recalling firm sent its two consignees an email with the recall notice attached. The notice notified their customers that they were affected by the recalling firm's supplier's recall. The notice instructed the customers to dispose of the recalled component and acknowledge receipt of the recall notice. |  | 24 | 24 kits | Distributed to one customer in Lake Charles, LA and one customer in Spartanburg, SC. | Terminated | April 05, 2017 | 595 | 91 | Not_Computer
Z-2807-2015 | 70858 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA II Duodenovideoscope OLYMPUS TJF TypeQ180V. |  |  | March 26, 2015 | 2015 | Olympus America Inc. | New and validated reprocessing procedures have been issued for the Olympus TJF-Q180V duodenoscope ("TJF-Q180V"), consisting of revised manual cleaning and high level disinfection procedures. | Labeling design | The firm, Olympus, sent an "Urgent Safety Notification" dated March 26, 2015, to direct consignees/customers. The notification described the product, problem and actions to be taken. The notification contained new reprocessing procedures. The customers were instructed to immediately implement new precleaning and high level disinfection procedure; ensure that all personnel involved in the reprocessing are completely knowledgeable and thoroughly trained on the new reprocessing instructions; complete and return the enclosed questionnaire via fax to (484) 896-7128 and implement the new cleaning instructions when the new brushes arrive on/about May 8, 2015.If you have question, contact our Technical Assistance Center (TAC) at 1-800-848-9024, option 1, Monday-Friday between 7 AM EST - 8 PM EST. For additional information on this matter, contact V.P., Regulatory/Clinical Affairs &amp; Quality Assurance at (484) 896-5688 or by email at laura.storms@olympus.com. |  | 6109 | 6109 | Worldwide Distribution: US (nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 6109 | Not_Computer
Z-2804-2015 | 71928 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | The Leica ASP6025 Tissue Processor. Pathology and Hematology Departments. |  |  | July 27, 2015 | 2015 | Leica Microsystems, Inc. | The ASP6025 Tissue Processor instrument is incorrectly getting into a state of bottle empty during retort filling, after passing the fill lever sensor 2 and prior to reaching levelsensor 3, thus not completing a 5 litre fill. | Nonconforming Material/Component | Leica sent an Urgent Medical Device Recall Notification letter dated August 11, 2015, to all affected customers via USPS 1st class mail ..Thel letter advised users to 1) do not use the instrument in 3 basket (5L) mode. Instead it advises the customer to change settings to use only the 2 basket (3.8L) mode only. 2)Prompt acknowledgement is requested by the customer 3) when a permanent solution is found a Leiica representative will contact the customer. 4) The Medical Device Recall Notification needs to be passed on to all those who need to beaware within the customers organisation or to any organisation where the potentially affected devices have been transferred.The customer is asked to sign the Medical Device Recall Notification Response Form to confirm that they have received and understand the Medical Device Recall Notification.The customer is requested to confirm receipt of the notice by signing and dating the Acknowledgement Form and faxing it back to Leica at 1-847-236-3747, or scan and email it to LMGRA@leicamicrosystems. com |  | 573 | U.S. Quantity Distributed: 163 systems; Foreign Quantity Diistributed: 410 systems. | Worldwide Distribution - US including AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH,PA, SC, TN, TX, VA, WA, WI, WV and District of Columbia. Foreign distribution to United Arab Emirates, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Colombia, Czech Republic , Germany, Spain, France, England, Georgia, Indonesia, Ireland, Israel, India, Italy, Jordan, Japan, Korea, Kuwait, Lithuania, Myanmar, Mexico, Malaysia, Netherland, Norway, Peru, Philippines, Poland, Qatar, Reunion, Romania, Serbia, Saudi Arabia, Singapore, Slovenia, Thailand, Turkey, Taiwan, and South Africa. | Terminated | January 05, 2017 | 528 | 573 | Not_Computer
Z-2806-2015 | 71706 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Processor EPX 2500 High Definition Endoscopy; (Sales Brochure). |  |  | June 30, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | The EPX-2500 Operation Manual and sales brochure incorrectly identify that the ED-530XT endoscope is meant to be used with the EPX-2500 Processor. The Operation Manual and sales brochure state that when the ED-530XT endoscope is used with the 2500 processor, a "super image" is displayed when it is actually a "standard image" that is displayed. | Device Design | Fujifilm Medical Systems USA, Inc. initiated this recall by sending a Voluntary Field Correction Letter and Tracking/Verification Form dated June 30, 2013 to their affected customers. |  |  | multiple | Nationwide US and Canada | Terminated | May 25, 2017 | 695 |  | Not_Computer
Z-2809-2015 | 71819 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Information Manager System, Version 1.0 to Version 4.11, Catalog No. 10714149, 10714150. Intended to manage both microbial identification (ID) and antimicrobial agent susceptibility testing (AST). |  |  | July 17, 2015 | 2015 | Beckman Coulter Inc. | The software incorrectly allows the operator to manually edit the carbohydrate substrates when manually reading dried overnight gram negative panels with an ID Hold status. | Unknown/Undetermined by firm | An Urgent Medical Device Recall Letter dated 7/17/2015 was sent to all customers who purchased the MicroScan LabPro Information Manager System, Version 1.0 to Version 4.11. Beckman Coulter is recalling the MicroScan LabPro Information Manager System because the software incorrectly allows the operator to manually edit the carbohydrate substrates when manually reading dried overnight gram negative panels with an ID Hold status. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Representative: http:www.beckmancoulter.com, (800) 677-7226 in US and Canada. Customers outside of US and Canada, are instructed to contact their local Beckman Coulter representative. |  | 2456 | 2,456 units total (1,034 units total) | Worldwide distribution: United States (including Puerto Rico), Bahamas, Belgium, Canada, Chile, China, Colombia, France, French Guiana, Germany, Greece, Guadeloupe, India, Israel, Italy, Japan, Kazakhstan, Lithuania, Mexico, Netherlands, Panama, Paraguay, Poland, Portugal, Romania, Russia, Singapore, South Africa, Spain, Turkey, , US Virgin Islands, and Northern Mariana Islands, Uruguay. | Open, Classified |  |  | 2456 | Computer
Z-2811-2015 | 72139 | No | GXN | Preformed nonalterable cranioplasty plate. | Neurology | Neurology | 510(k) | Synthes Cranial Flap Tube Clamp and Crimping Device |  |  | August 31, 2015 | 2015 | Synthes (USA) Products LLC | The Synthes Cranial Flap Tube Clamp and Crimping Device for Cranial Tube Clamp had been labeled MR Safe although they do not have the testing protocols currently required to designate them as MR Safe. | Labeling False and Misleading | DePuy Synthes sent an URGENT:FIELD SAFETY NOTIFICATION / MEDICAL DEVICE LABELING CORRECTION letter, dated August 31, 2015 to affected customers. The letter identified affected product, problem and actions to be taken. Customers were instructed to:"Update and review the package insert (GP1345-D)." Discard outdated package insert noted in the table above." Forward this Field Safety Notification to anyone in your facility that needs to be informed, especially personnel that conduct MR testing." If the package insert has been forwarded to another facility, contact that facility." Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided. Send a copy of the completed Verification Section by, Fax: 877-884-7821 or Scan/email: Synthes3911@stericycle.com" If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification." Maintain a copy of this notice.For questions call 610-719-5450. |  | 105882 | 105882 | US Nationwide and internationally, including Canada. | Open, Classified |  |  | 105882 | Not_Computer
Z-2808-2015 | 71757 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Marketing brochures for the PowerPICC and PowerPICC SOLO catheters |  |  | July 20, 2015 | 2015 | Bard Access Systems | Bard Access Systems is recalling marketing brochures for the PowerPICCÂż and PowerPICCÂż SOLOÂż catheters due to incorrect indication for use contained in the brochure. | Labeling False and Misleading | Customer letter notification with reply form will be sent via Federal Express with delivery confirmation for US consignees on 7/20/15 informing them of the incorrect indication for use contained in the brochure. |  | 1811 | 1811 | Distributed to domestic and foreign conference attendees. | Open, Classified |  |  | 1811 | Not_Computer
Z-2810-2015 | 71754 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident 10 degrees X3 Insert 36 mm ID |  |  | June 25, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | it was discovered that item 623-10-36I Lot 3Y0KHP, although meeting all device specifications, was determined to be unacceptable due to a documentation issue. However the parts were inadvertently shipped to the field. | Nonconforming Material/Component | Stryker Orthopaedics notified their branches via email on 6/25/2015. Notification letters and Product Accountability Forms were mailed to Stryker Branches/Agencies on 817/2015 and hospital risk management on 8/18/2015. |  | 16 | 16 units | Distributed in WI. | Open, Classified |  |  | 16 | Not_Computer
Z-2817-2015 | 72071 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CD4 (SK3), FITC; |  |  | August 27, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | CD4 FITC label contains an error in the Spanish small text - it states CD8 FITC. | Under Investigation by firm | Urgent Product Recall notification letters were sent to each customer on August 27, 2015 by certified mail. |  | 128 | 128 units | Worldwide Distribution. US nationwide (TX, CA, NC, OH, SD, PA, CT, UT, MD, TX, GA), Argentina, Taiwan, Australia, Japan, China, New Zealand, Chile and Colombia. | Open, Classified |  |  | 128 | Not_Computer
Z-2815-2015 | 72044 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDPoint 500 Measurement Cartridge Lactate (250) tests |  |  | August 24, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | RAPIDPoint 500 measurement cartridge serial numbers 2517102517 through 2519000012, negative sodium bias on both QC and patient results | Process control | Siemens Healthcare sent an Urgent Field Safety Notice letter dated August 2015 to all affected customers. The notice identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Effectiveness Check Form attached. For questions contact your Siemens Customer Care Center or ther local Siemens technical support representative. |  | 982 | 982 cartridges | Worldwide Distribution - US Nationwide and the countries of Canada, Albania, Austria, Australia, Bosnia Herzeg., Bangladesh, Belgium, Brazil, Switzerland, Chile, China, Colombia, Czech Republic, Germany, Denmark, Denmark, Estonia, Spain, Finland, France, Great Britain, Georgia, Greece, Croatia, Hungary, Ireland, India, Italy, Lithuania, Luxembourg, Latvia, Macedonia, Malaysia, Netherlands, Norway, New Zealand, Poland, Portugal, Paraguay, Romania, Sweden, Singapore, Slovenia, Slovakia, Turkey, Taiwan, South Africa, Unit.Arab Emirates | Terminated | May 23, 2016 | 273 | 4323 | Not_Computer
Z-2812-2015 | 71915 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE, ONCOR Avantgarde, ONCOR Impression Plus, ONCOR Expression, PRIMUS, PRIMUS Plus, Syngo RT Therapist; therapeutic treatment of cancer |  |  | August 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A software fix has been released to prevent automatic movement resulting in a collision safety risk for patients. | Software design (manufacturing process) | Siemens sent Notification letters to affected customers. The letter identified the affected product, problem and actions to be taken. The letter advise customers that the software fix was rolled out to address multiple safety issues. |  | 26 | 26 | Worldwide Distribution and US Nationwide | Open, Classified |  |  | 26 | Computer
Z-2819-2015 | 72035 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer, Reprocessed Zimmer Tourniquet Cuffs,34" Single Hose Single Bladder Brown (PLC), 34", Model (Reprocessor) Number: ZIM60-7075-106, Reprocessed by Sterilmed. |  |  | August 24, 2015 | 2015 | STERILMED, INC. | Sterilmed, Inc. is recalling its reprocessed Zimmer Tourniquet Cuff, Part #ZIM60-7075-106, lot number 1846885 due to mislabeling. The affected reprocessed product is labeled as single hose devices but the devices inside the packaging are dual hose devices. | Labeling False and Misleading | Sterilmed sent an "Urgent Medical Device Recall" letter dated August 24 2015. The letter was addressed to Materials Manger, Risk Manger and Operating Room Supervisor. The letter identified the affected product, problem and actions to be taken. Customers are requested to return all affected product, sign and return the enclosed Customer Response Questionnaire and the Customer Response Form per the instructions on the Forms. For questions contact Sterilmed's Customer Care Team at 888-541-0078. |  | 43 | 43 | US Nationwide Distribution in the states of CA, KY, MI, OH, UT and WI. | Terminated | March 11, 2016 | 200 | 43 | Not_Computer
Z-2818-2015 | 72051 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA II Gastrovideoscope OLYMPUS GF type UCT180; designed for endoscopic real-time ultrasound imaging, ultrasound guided needle aspiration and other endoscopic procedures within the upper gastrointestinal tract and surrounding organs. |  |  | August 21, 2015 | 2015 | Olympus Corporation of the Americas | Insufficient information for specific cleaning accessories used for the cleaning and reprocessing of the GF-UCT180 ultrasound endoscope | Device Design | Olympus sent an Urgent Device Safety Information letter dated August 2015, to all affected customers. Corrected instructions were provided with the letter and users were asked to begin implementing them immediately. Customers were asked to indicate on the enclosed questionnaire and fax the completed from to 484-896-7128. Customers with questions were instructed to call 484-896-5688. |  | 1014 | 1014 | Nationwide Distribution | Terminated | May 08, 2017 | 626 | 1014 | Not_Computer
Z-2784-2015 | 71995 | Yes | NVZ |  |  | Cardiovascular | PMA | Merlin@home RF Remote Monitoring Transmitter Model EX1150. |  |  | August 18, 2015 | 2015 | St Jude Medical Cardiac Rhythm Management Division | Some devices exhibit backup VVI operation due to a backup reset. | Software change control | The notification material is in the form of a Physician Letter was delivered to physicians by FedEx overnight service. The mailing went out on August 17, 2015 for receipt on August 18, 2015.The firm's notification letter identifies clinical implications.Clinical ImplicationsIn the event that an Ellipse, Fortify Assura, Unify Assura, or Quadra Assura ICD enters the back-up safety mode, the nominal operational settings will be VVI pacing mode, 67 ppm, 5.0v/0.6ms with bipolar pacing output and defibrillation settings of aVF detection rate of 146 bpm and 36J high voltage therapy. In the event an Assurity or Allure pacemaker enters backup mode, it will have output settings of VVI pacing mode, 67 ppm, 5.0v/0.6ms with unipolar pacing.This anomaly can only occur when the patient is being actively monitored by a Merlin@home bedside transmitter. If a device enters backup mode, the ICD will deliver a patient vibratory alert and the pacemaker will deliver a patient audible alert. If backup operation is encountered, St. Jude Medical Technical Services (1-800-722-3774) can assist with non-invasively restoring the device to normal operation.The firm's notification also idenfities recommendations and mitigations.Recommendations and MitigationsOnce available, a Merlin@home transmitter software update will be performed automatically over its telephone, broadband or cellular connection without requiring any action from you or your patients. No changes to the patients remote or in-clinic follow up schedules are required. In the event a patients device reverts to back-up mode, the firm recommends bringing the patient back in the clinic to clear the condition and return the device to full functionality.If you have any questions or concerns contact the local St. Jude Medical representative or St. Jude Medical's Technical Services Department. |  | 259214 | 259,214 transmitters (U.S. distribution) 23,604 (foreign distribution) | Worldwide Distribution-US (nationwide) and foreign distribution to the following: Japan, France, Austria, Canada, Italy, Spain, Netherlands, South Africa, Great Britain, Australia, Sweden, Belgium, Finland, Germany, Portugal, Singapore, Norway, Poland, Ireland, Czech Republic. Denmark, Hong Kong, Saudi Arabia, Hungary, Switzerland. | Open, Classified |  |  | 259214 | Computer
Z-2743-2015 | 71925 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROFEMUR PLUS CoCr MODULAR NECK, REF: PHAC1254, SIZE LONG, ORIENTATION 8 degree VAR/VAL. 12/14 SLT Taper. |  |  | August 04, 2015 | 2015 | MicroPort Orthopedics, Inc. | Unexpected rate of postoperative fractures resulting in the need for revision surgery. | Under Investigation by firm | Micro Port Orthopedics sent a Voluntary Device Product Recall letter dated August 7, 2015, via Fed Ex to all affected customers. The letter identified the product, the problem, and action to be taken by the customer. Customers were instructed to locate and return effected product immediately. Customers with questions were instructed to call 1-866-872-0211. Customers were also asked to return the enclosed form immediately even if they do not have any affected product to return. |  | 10489 | 10,489 units | Worldwide Distribution - US (nationwide) and Internationally to AT, BE, CA, CN, CZ, DE, ES, FI, FR, GB, IT, KR, LT, LV, NL, NO, PE, PL, PT, RU, and ST. | Open, Classified |  |  | 10489 | Not_Computer
Z-2820-2015 | 72132 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Guide Wires for orthopedic procedures. |  |  | August 27, 2015 | 2015 | Zimmer, Inc. | Product labeling of these sterile guide wires (external carton label and patient record label) states a 10-year expiration date. Package testing supports a 5-year expiration date for these same guide wires. | Package design/selection | Zimmer sent an "URGENT MEDICAL DEVICE RECALL" notifications dated August 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities:1.Review the notification and ensure affected personnel are aware of the contents.2.Locate part number and lot number combinations listed in attachment 1 and quarantine the devices immediately.3.Carry out a physical count of all affected product in your territory, including hospitals in your territory with directsales as well as consignment, and complete the Inventory Return Certification of Acknowledgement (Attachment 2). Email a completed copy of Attachment 2 to corporatequality.postmarket@zimmerbiomet.com.4.Return the recalled product along with the completed Attachment 2. Clearly mark the outside carton of each product return shipment made as Recall. 5.Please notify Zimmer Biomet of any hospitals and surgeons that have consumed the product in the past 12 months that are not included in either the distribution history or inventory system data information. Supply the information forany hospitals and surgeons that you have additionally identified by using the provided spreadsheet Additional Account template. The template, and the list of known consignees from the distribution history and inventory systems data, will be emailed to you for completion and return to: corporatequality.postmarket@zimmerbiomet.com.6.If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm ESTFor further questions about this recall, please call ( 1-800-447-5633). |  | 112568 | 112,568 units | Worldwide Distribution - US (nationwide) including Puerto Rico and Virgin Islands., and to the countries of Canada, Brazil, Algeria, Mexico, Austria, Taiwan, Belgium, Hong Kong, Croatia, Singapore, Denmark, Bahamas, Egypt, Panama, Finland, Dominican Republic, France, Guatemala, Germany, Australia, Ireland, Japan,Italy, Haiti, Jordan, Malaysia, Kenya, India, Lebanon, Luxembourg, Malta, Mauritius, Morocco, Netherlands, Norway, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, United Kingdom and United Arab Emirates. | Terminated | March 15, 2016 | 201 | 112568 | Not_Computer
Z-2825-2015 | 72086 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Kits with Medtronic Covidien 0.9% Sodium Chloride Flush Syringes |  |  | August 27, 2015 | 2015 | Arrow International Inc | Manufacturer of 0.9% Sodium Chloride Flush Syringes recalled due to compromised sterility of the outer barrel, Arrow purchased recalled items that were packaged with certain products. | Process control | An urgent recall notification letter, dated August 27, 2015, was sent to direct consignees to inform them of the issue and provided instructions for removing the syringes from the kit and returning product along with the response form. |  | 73714 | 73714 | Nationwide and internationally. | Open, Classified |  |  | 73714 | Not_Computer
Z-2823-2015 | 72183 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | P/N: 9735023, DRIVER 9735023 SOLERA 5.5/6.0 MAS; P/N: 9735024, DRIVER 9735024 SOLERA 5.5/6.0 MAS CAN; |  |  | September 14, 2015 | 2015 | Medtronic Navigation, Inc. | Specific use conditions can increase the possibility of a screwdriver tip deforming or breaking during surgery. | Labeling design | Medtronic sent an Urgent Medical Device Correction Notification letter dated September 14, 2015, to all affected consignees via FedEx. Revised Instructions for Use were provided to add additional warnings related to the careful inspection of the instruments and the importance of knowledge of operating procedures, patient selection, and product information. Consignees with questions were instructed to to call the local Medtronic Navigation Technical Services Representative or call 1-800-995-9709. |  | 2346 | 2,346 | Worldwide Distribution - US (nationwide) and Internationally to Canada; United Arab Emirates; Australia; Belgium; Switzerland; China; Germany; Denmark; Spain; Finland; France; Great Britain; Greece; India; Israel; Italy; Japan; Netherlands; New Zealand; Poland; Portugal; Sweden; Singapore; Thailand; Slovakia; Turkey; and South Africa. | Terminated | December 03, 2015 | 80 | 2346 | Not_Computer
Z-2822-2015 | 72106 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ASP-1021LipoTower with Aspiration Pump and Tumescent Pump. |  |  | June 12, 2015 | 2015 | MicroAire Surgical Instruments, LLC | There is potential risk of electrical shock if lead wires inside the device become detached and the ground connection fails. | Device Design | The firm has issued FDA recall letters to all consignees. The letter have been sent out by UPS to expedite the delivery. |  | 1189 | 168, ASP-1021 LipoTower with Aspiration Pump and Tumescent Pump | Nationwide Distribution. | Open, Classified |  |  | 1192 | Not_Computer
Z-0004-2016 | 71858 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips GoSafe Mobile Help Button; 7000MHB. |  |  | July 23, 2015 | 2015 | Lifeline Systems, Incorporated | Irregular battery disconnection resulted in issues with Mobile Health Button (MHB) Charging. | Process control | Customers were also notified of the recall via an Urgent: Medical Device Recall sent, by US Mail, between 8/19/2015 and 8/20/2015. The letter described the reason for the recall and how to recognize the device failure. The customer is asked to do the following: 1. Review and confirm the attached list of affected units for your program. Contact the affected subscribers listed and notify them that a service visit will be scheduled so that their units can be replaced within the next few weeks. 2. If any subscribers report any issues with their MHB, contact Philips for an expedited unit replacement and instruct them to discontinue use. 3. Philips Customer Care will follow-up with customers via telephone to ensure that they have received and understood the notification. During this follow-up call, we will also confirm arrangements for unit replacements. |  | 8616 | 8,616 devices | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, KY, MA, MD, MI, MN, MT, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TX, UT, VA, VT, WA, and WI. | Open, Classified |  |  | 8616 | Not_Computer
Z-0001-2016 | 71987 | No | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Histolock Resection Device, US Endoscopy. |  |  | July 14, 2015 | 2015 | US Endoscopy Group Inc | The company has determined the sterility cannot be assured for the affected lot. | Employee error | The firm, US Endoscopy, sent an "URGENT: HISTOLOCK" RESECTION DEVICE RECALL" letter and response form to their customers on July 14, 2015. The letter described the product, problem and action to be taken. The customers were instructed to complete the response form and return to your US Endoscopy Product Specialist and retain a copy of the letter for your records.If you have any questions, contact Sr. Manager, Reg. Affairs &amp; Quality Compliance at 440-358-6251 or email at: ccochara@usendoscopy.com. |  | 41 | 41 units | US Distribution to states of: AL, AR, CA, IL, IN, MS, OH, OK &amp; TN. | Open, Classified |  |  | 41 | Not_Computer
Z-0003-2016 | 72167 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SOUNDSTAR eco 8F and 10F Diagnostic Ultrasound Catheters. For intra-cardiac and intra-luminal visualization of cardiac and great vessel anatomy and physiology as well as visualization of other devices in the heart. |  |  | September 10, 2015 | 2015 | Biosense Webster, Inc. | Image disappeared from the cardiac ultrasound system when the CARTO 3 EP Navigation System needed restarting while the patient was experiencing pericardial effusion. Affects the CARTOSOUND Module of the CARTO 3 EP Navigation System when used with the SOUNDSTAR eco 8F and 10F Diagnostic Ultrasound Catheters. New precautions added. | Device Design | A Customer Notification Letter and Acknowledgement Form was distributed to customers on 09/10/15. One letter and form was sent via express mail, addressed to the Risk Management Office and one additional letter and form was sent by the firm's account representative to the Electrophysiology (EP/Cardiology Lab directly at the customer site.The firm's notification letter stated that they want to emphasize the following statement "The intra-cardiac ultrasound image will disappear if the CARTO 3 EP Navigation System power is disrupted and this may present a safety issue if the EP is using the ultrasound to monitor the patien during EP procedure. The ultrasound image will not reappear until the CARTO 3 System is restored."The firm will be updating the product labeling to further reinforce the precautionary statement.The firm requests that customers take the following actions:1. Read the Field Safety Notification carefully.2. Pass on this notification to anyone in the facility that needs to be informed of this issue, including appropriate clinical personnel involved in the use of SOUNDSTARÂż eco Catheters.3. Review, complete, sign and return the attached Acknowledgement Form in accordance with the instructions on the form. 4. Maintain a copy of this letter with the product.5. Maintain awareness of the Field Safety NotificationFor any questions the firm states to contact the BWI sales representative. |  | 1035 | 1035 (U.S.) | Worldwide Distribution. US nationwide, Italy, Canada, Japan, China, Singapore, Colombia, England, Russia, Spain, France, Germany, Finland, Ireland, Australia, Sweden, Hungary, India, Brazil, Slovenia, Netherlands, Mexico, and South Africa. | Open, Classified |  |  | 1035 | Not_Computer
Z-2826-2015 | 71983 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Ferno iNX Wheeled Stretcher, Integrated Patient Transport &amp; Loading System |  |  | August 03, 2015 | 2015 | Ferno-Washington Inc | Hole elongation in the telescoping load frame channel. A hole elongation could affect the ability of the retention pin to stay seated which is necessary to preclude movement of the telescoping load frame. | Device Design | Ferno began calling their customers on 8/3/2015. The firm informed their customers of the following: To provide optimal Customer Experience and minimal disruption to your service, the iNX field action involves the following steps:" Exchange of your current iNX with a new iNX;" Participation in a brief training session; and" Inspection of your Ferno Model 175 antler and rail fastening system as requested by the Dept. of Transportation.As you await delivery of your new iNX, it is important you ensure proper function of the fastening system and iNX telescoping load frame. |  | 72 | 72 units | Worldwide Distribution - US Nationwide in the states of: CO, IA, IL, IN, KY, MD, MN, MO, NC, NJ, NY, OH, OK, OR, TX, VA, WA &amp; WY and the countries of CANADA, AUSTRALIA, UNITED KINGDOM &amp; NEW ZEALAND. | Terminated | January 20, 2016 | 170 | 72 | Not_Computer
Z-2795-2015 | 71996 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Disposable Reflective Marker Spheres |  |  | August 18, 2015 | 2015 | Brainlab AG | Disposable Reflective Marker Spheres (DRMS) for Brainlab Image Guided Surgery (IGS) Systems: DRMS spheres may separate at the mid-point where the two halves of the sphere are sealed together. | Process control | The firm, Brainlab AG, sent out a" FIELD SAFETY NOTICE/ PRODUCT NOTIFICATION" dated 7/23/2015 to its customers. The firm sent this letter to mitigate the potential effects of the affected Disposable Reflective Marker Spheres (DRMS) when used with Brainlab Image Guided Surgery (IGS) systems. The firm advised the medical users that they should take the following steps prior to each procedure using Brainlab DRMS spheres for Brainlab IGS systems: -Prior to opening the package containing the spheres, visually inspect the spheres in the blister pack for any signs of separation between the two halves of the sphere. If there is any indication that the spheres may have separated, use a different blister pack for the procedure and repeat this step with the new pack. Discard any and all affected spheres and inform your Brainlab support representative about the occurrence of the issue. If there is no visual sign of sphere separation, proceed to the next step. -Prior to the start of the procedure, screw each DRMS sphere onto the tool and tighten. Upon tightening, inspect again to determine whether any separation between the two halves is evident. If separation is evident, remove and discard the affected sphere. Replace your surgical gloves, then replace the separated sphere with a new sphere, inspecting it as above. If no separation is evident, proceed with the clinical procedure. -Be sure to discard and replace your surgical gloves any time that a sphere with evident separation is found upon affixing the sphere to the tool, and every time after a separated sphere was touched.-Caution: Do not affix the sphere to the tool over the surgical area to minimize the risk of a sphere, or any portion thereof, falling onto the patient or into the surgical wound. Also avoid affixing the sphere over other sterile areas (such as surgical trays) to minimize potential contamination in case of sphere separation. -Reminder: In general, do not use defective or deformed marker |  | 369659 | For 28 US consignees, 9,900 Disposable reflective marker spheres ***For 451 non-US consignees, 359,280 Disposable reflective marker spheres | Worldwide Distribution: US (nationwide) including states of: TX, CA, TN, NY, LA, IL, MA, PA, OH, AZ, ID, CT, MI, and KY; and countries of: Algeria, Argentina, Australia, Bahrain, Belarus, Bolivia, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Germany, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Malaysia, Malta, Mexico, Moldova, Morocco, Myanmar, Netherlands, Philippines, Poland, Portugal, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, , Venezuela, and Vietnam. | Open, Classified |  |  | 369659 | Not_Computer
Z-2766-2015 | 72045 | No | NGV |  |  | Cardiovascular | PMA | Boston Scientific, WATCHMAN Access System Access Sheath with Dilator, 14F Single curve( M635TU10060, TU1006) , 14F Double curve (M635TU20060, TU2006) , &amp; 14F Anterior curve (M635TU40060, TU4006). For cardiac use. |  |  | August 25, 2015 | 2015 | Boston Scientific Corporation | Cross-threading of the hemostasis valve may occur if it is tightened with the dilator in place, possibly preventing subsequent sealing of the valve when desired. This Field Correction reinforces the Directions for Use (DFU) and gives further guidance on correct use of the hemostasis valve to avoid cross-threading and to securely seal the valve, minimizing potential for undesired blood leakage. | Use error | Consignees were sent a Boston Scientific "Urgent Field Correction" letter dated August 25, 2015. The letter was addressed to Materials manager / Field Action Contact. The letter described the problem and the product affected. The Field Correction reinforces existing Directions for USE (DFU) and provides further guidance regarding the correct us of the hemostatis valve in order to avoid cross-threading and to securely seal the valve, minimizing the poential for undesirable blood leakage. The consignees were informed of the action required which stated that no product is required to be returned and advised to consignees that they must complete the Customer Acknowledgement Form and return it to Boston Scientific (MapleGroveFieldAction Center@bsci.com or Fax to Field action Center 1-866-213-1806). The letter also provided a section of the Updated Directions for Use which were highlighted in blue. If the consignee is a distributor the Field Correction is to the hospital level. For questions, customers may contact their Local Sales Representative. |  | 29169 | 29,169 units total | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Iran, Ireland, Israel, Italy, Kazakhstan, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Oman, Panama, Philippines, Poland, Portugal, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, White Russia. | Open, Classified |  |  | 29169 | Not_Computer
Z-0022-2016 | 72245 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | Medtronic Medical Device Identification Card for SureScan pacemaker patients. The front of this card identifies the implanted pacing system components currently registered to the patient. The back of the card lists patient's following physician as well as specific messaging regarding the MR conditional status of the device components, where applicable. This card is provided to the patient by Medtronic after device registration. |  |  | September 15, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Some Medical Device Identification Cards provided to SureScan pacemaker patients indicate they have a complete Magnetic Resonance (MR) Conditional system, when in fact, not all of their implanted leads have been FDA approved as MR Conditional. | Process control | On September 15, 2015, Medtronic began mailing an Important Medical Device Identification Card Notification letter, including a corrected Medical Device Identification Card, to impacted patients. The letter described the issue, provided a new corrected Identification Card, asked that the old card be destroyed immediately, and to begin using the new card. Patients can contact Medtronic Patient Services at 1-800-551-5544, option 1, followed by option 3, Monday - Friday, 8am - 5pm Central time. |  | 777 | 777 cards | US nationwide distribution. | Terminated | April 14, 2016 | 212 | 777 | Not_Computer
Z-0024-2016 | 72238 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Periarticular Insertion Handle for 4.5mm LCP Proximal Tibia Plates and Periarticular Aiming Arm for 4.5mm LCP Proximal Tibia Plates |  |  | September 14, 2015 | 2015 | Synthes (USA) Products LLC | Affected lots of Periarticular Aiming Arms and Insertion Handles (Part Numbers 03.120.001, 03.120.002, 03.120.004 and 03.120.005) for the 4.5 mm LCP Proximal Tibia Plates are not able to be secured to the subsequent versions of these devices as intended. | Device Design | An urgent notice of medical device recall, dated September 14, 2015, was sent to end users and sales consultants to alert them to the issue and potential risks. Return of the completed response form and the affected devide was also requested. |  | 172 | 172 | Nationally | Open, Classified |  |  | 172 | Computer
Z-0023-2016 | 72212 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. |  |  | September 16, 2015 | 2015 | GE Healthcare | Images may be missing when a system parameter MapRoute is set to a value greater than 1. | Software Design Change | The firm, GE Healthcare, an "Urgent Medical Device Correction" by mail or hand delivered to all sites. The letter instructs the customer to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. Until GE Healthcare conducts remote inspection of the systems "MapRoute" parameter value. If the value is found to be greater than the recommended setting of 1, GE will reset this system parameter to a value of 1. The letter informs the user facilities not to make any changes to this parameter value.Additionally, until a product modification is available to correct this issue, the following actions are recommended. The firm is recommending user facilities to: 1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the PACS IW viewer.a. If a discrepancy is identified, attempt to retransmit the exam to PACS.b. If retransmission is unsuccessful, contact your GE Healthcare Service representative for assistance and instructions as to the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.GE Healthcare will correct all affected systems by providing a software update at no cost to the user facilities. If the user facilities have questions, they are encouraged to communicate questions or concerns regarding this notification to GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 921 | 921 systems are impacted. | Worldwide Distribution: US (nationwide) including states of: AL AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, ON, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and Puerto Rico; and countries of: Argentina , Austria, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile , China, Colombia, Costa Rica, Ecuador, Egypt, France, Germany, Great Britain, Hong Kong, Hungary, India, Israel, Italy, Japan,Korea, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Mexico, Netherlands, New Caledonia, Nigeria, PanamÂż, Peru, Philippines, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, Venezuela, and Vietnam. | Terminated | January 05, 2017 | 477 | 921 | Computer
Z-0020-2016 | 72091 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash; intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | August 14, 2015 | 2015 | Siemens Medical Solutions USA, Inc | software bug issues for SW-Version VA48A_SP0. The following safety issues were resolved: 1) Correction to improve visual warning and error indication son the gantry display. 2) Correction to improve acquisition data in order to optimize image quality. 3) Correction to improve robustness and general system behavior in some exception handling procedures. 4) Correction to improve auto post processin | Software design | A customer advisory notice, dated 8/14/15, was sent to direct accounts advising them of the issue to observe the safety notice and comply with the corresponding measure until the update has been fully completed. |  | 113 | 113 total | Nationwide Distribution. | Open, Classified |  |  | 113 | Computer
Z-0025-2016 | 72068 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Universal Compact Head Ring Adapter Plate (UCHRAP), a component of the Universal Compact Head Ring Adapter, UCHRA. The UCHRA is a component of the IntegraÂż CRWÂż Precision Arc Stereotactic System. |  |  | August 21, 2015 | 2015 | Integra LifeSciences Corp. | Integra identified that the UCHRAP component (Arc Adapter Plate) would not assemble properly to the UCHRAR component (ARC Adapter Ring) of the Compact Head Ring Adapter due to an error in the manufacturing drawing. | Process design | Integra sent an Urgent-Voluntary Medical Device Recall letter dated August 19, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm instructed their customers to:1. Review your inventory and determine if you have an UCHRAP Adapter Plate with an affected serial number listed on the next page. If you do not, it is not affected and you can continue to use it.2. If you do have product with a serial number listed on the next page, stop using it immediately and remove it from service. 3. Complete the attached form. If you do not have product with a serial number listed, check the box, I do not have any product with a serial number on the list below.4. If you do have product with a serial number listed, check the box, I do have product(s) with a serial number below. Circle the lot numbers and record the total quantity of all lot numbers you have.5. Complete the other information and return it by email or fax as indicated on the form or by providing it to your Integra Neuro Specialist.If an Integra Neuro Specialist has not already done so, Customer Service will contact you and provide an RMA number and directions for returning and replacing the product. Should you have any questions regarding these instructions, please call : (513) 533-7932. For any other questions, please call (609) 936-6822. |  | 6 | 6 units | Worldwide Distribution - US Distribution to the states of : NJ, TX and WY., and to the countries of : Algeria, Germany and India. | Open, Classified |  |  | 6 | Not_Computer
Z-0105-2016 | 72107 | No | HEE | Obstetric anesthesia set. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical LEEP REDIKIT-Used as a surgical preparatory kit for Loop Electrosurgical Excision Procedure |  |  | September 01, 2015 | 2015 | CooperSurgical, Inc. | Kit Label incorrect :Labeled with Lidocaine HCI drug cartridges Lidocaine HCl (2%) but contains vials Lidocaine HCl (1%) | Labeling Change Control | CooperSurgical issued an "Recall Notification" dated September 1, 2015, to all affected customers via FedEX with confirmed delivery receipt. The letter identified the product, the problem, and the action needed to be taken by the customer.Customers asked to discontinue use of affected products and complete the attached Acknowledgement and Receipt Form for a free replacement.Contact Product Surveillance at 203.601.5200 ext. 3300 |  | 305 | 305 units | US Distribution to the states of : AL, AZ, CA, GA, IA, IL, KS, LA, MA, MD, MN, MO, MS, NY, OH, PA, TN, TX, VA, and VT. | Terminated | January 29, 2016 | 150 | 305 | Not_Computer
Z-0107-2016 | 72137 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force; computed tomography x-ray system. |  |  | August 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Software and firmware bugs | Software design | Customers will be notified via the system interface and e-mail or U.S. postal mailing via a Customer Advisory Notice. If necessary the advisory letter may be hand delivered. The letter informed customers that the safety related issues identified were resolved by SW-Version VA50A_SP2. The update will be remotely pushed to affected systems. The letter also advised customers to notify anyone in their organization that should be aware of the information. In the event that the device had been sold, the notice should be forwarded to the new owner, and the firm should be informed of the new owner. |  | 21 | 21 | Nationwide Distribution | Terminated | May 08, 2017 | 629 | 21 | Computer
Z-0027-2016 | 72121 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Universal Locking System, Titanium 3.5mm. |  |  | August 11, 2015 | 2015 | Zimmer, Inc. | It was found through review of open Complaints and stock investigations that a single lot of Zimmer Universal Locking System 3.5 mm Stainless Steel Base and Trays has incorrect silk screening as applied by the supplier. The T-plate catalog numbers noted in the tray are listed as the 00-4946-xxx-xx series (titanium plates); however, the correct numbers should be 00-4936-xxx-xx (stainless steel plat | Process control | On 8/11/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 a.m. and 5:00 p.m. EST. |  | 18 | 18 units | Distributed in NC, AL, KS, TX, UT, CA, and SINGAPORE. | Terminated | March 15, 2016 | 217 | 18 | Not_Computer
Z-0106-2016 | 71979 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Oncologist, an optional accessory to the linear accelerator system and permits localization, contouring, segmentation, image review, and review and approval of treatment plan parameters. In addition, it includes tools and administrative functions to aid in the diagnosis, staging, and prescription of radiation therapy. |  |  | August 11, 2015 | 2015 | Siemens Medical Solutions USA, Inc | This update is intended to provide a software update for the Syngo RT Oncologist, which is currently running SW version 4.2 or 4.3. The safety-related issues, which were described in the Customer Safety Notice distributed as UFSN-RTT/RTO 4.3 Adaptive Targeting [Auto-registration]. The update contains several important safety and performance fixes. | Software design | The firm, Siemens, sent Recall letters, dated August 2015, that describe the product, the problem, and correction, will be hand-delivered to affected sites at the time the fix is implemented by a Siemens technician.If you have any questions, contact Regulatory Technical Specialist at 610-448-6471 or email: marlynne.galloway@siemens.com |  | 14 | 14 | US Distribution to states of: MA, MO, OH, PA, UT, and WI. | Open, Classified |  |  | 14 | Computer
Z-0026-2016 | 71899 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Ortho OS Solutions UltOS 3.5mm locking screw 12mm to 50mm Ortho Solutions LTD, Unit 5 West Station Business Park, Spital Road, Maldon, Essex CM96FF Rx Only a) Sterile b) Non-Sterile |  |  | July 14, 2015 | 2015 | Ortho Solutions Inc | Ortho Solutions Inc is initiating a recall on behalf of UltOS 3.5mm locking screw. The locking thread feature of the screw was not manufactured to specification. As a consequence there is a risk that a screw may not lock onto the plate at the time of primary insertion into the plate and the bone. In the event of this occurring the non functioning screw would need to be replaced with an alternativ | Nonconforming Material/Component | The firm, Ortho Solutions Group, sent an "Urgent Medical Device Recall" notifications and "Effectiveness or Recall Declaration Forms" dated 14 July 2015. The notification informed the consignee of the issue, clinical risks, how to identify the affected product, and the actions the consignees were requested to take. Requested actions included reviewing product on-hand, ceasing distribution of affected product and arrangement for product to be returned to Ortho Solutions, Inc. within 2 working days. Consignees were also requested to fill out the Recall Declaration Form and return it to Ortho Solutions, Inc. via fax to: 201-335-0759 by 16 July 2015.If you should have any queries relating to this matter or require any additional information, contact General Manager at 201-906-6175. |  | 875 | 875 units | US Distribution to states of: CA, CO, GA, IL, LA, NC, NY, TX and WA. | Open, Classified |  |  | 875 | Not_Computer
Z-0005-2016 | 71763 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Celesteion PCA-9000A/2 PET/CT System |  |  | May 08, 2015 | 2015 | Toshiba American Medical Systems Inc | It was found that if specific operations are performed in multi-phase helical scanning, the acquired raw data may not be saved due to a software problem. | Device Design | Planned Action by Toshiba Medical America Systems to address the defect:a. A notification will be sent to the customer with affected system in compliance with 21 CFR1 003.21 upon approval by the Secretary.b. To correct this issue revised software will be installed on the system to prevent this occurrence from happening. The consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.c. The customer will be advised to do the following until the corrective measures have been taken, i.e., revised software installed on their systems. If you have any questions please contact your local Toshiba Representative at (800) 521-1968. |  | 1 | 1 | US Distribution to the state of Nevada. | Terminated | December 03, 2015 | 209 | 1 | Computer
Z-0028-2016 | 71904 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | regard Item Number: 800091, Sterile, LD0692 - C Section - surgical kit containing cover light handle soft, RX. The responsible firm on the label is ROi, LLC, St. Louis, MO. |  |  | April 22, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | The kits contain an equipment glove component which may contain splits or holes compromising the sterility. | Process control | ROi initiated a recall of the custom kits via e-mail to their customers on April 22, 2015. The e-mail included their recall letter dated April 17, 2015. and the supplier's recall letter dated April 16, 2015. The customer was requested to discard the recalled component from their procedure tray at the time of set up and replace it with sterile stock. Customers were instructed to contact any ROi custom pack solutions manager with any concerns. The recalling firm issued a second e-mail notification on 10/19/2015 to the recall coordinator of the distributor. The notice informed them they had been receiving uncorrected product and reminded them to discard the defective component at the time the kit is set up for surgery. Any future distribution will be labeled with an alert to remove the defective component until the kits contain acceptable light handle covers.For questions regarding this recall call 314-364-6561. |  | 26 | 26 kits | Distribution was made to LA, MS, and TX. | Open, Classified |  |  | 38056 | Not_Computer
Z-0110-2016 | 72090 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Free T3 Reagent Kit, P/N A13422 |  |  | August 31, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter, Inc. is recalling Access Free T3 Reagent Kit because they were filled incorrectly and contain insufficient quantity of reagents in one or more of the pack wells. | Under Investigation by firm | Beckman Coulter sent an " Urgent Medical Device Recall" letter dated August 31, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informs the customers that some Access Free T3 reagent packs (p/n A13422) of lot 431433 were filled incorrectly and contain insufficient quantity of reagents in one or more of the pack wells. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the notice are instructed to contact Customer Support Center at http://ww.beckmancoulter.com; call (800) 854-3633 in US and Canada. Customers outside of US and Canada are instructed to contact their local Beckman Coulter representative. |  | 16648 | 16,648 units total (1874 units in US) | Worldwide Distribution - US (nationwide) and to the countries of : Australia, Brazil, Canada, China, Colombia, Costa Rica, Hong Kong, Israel, Japan, Kuwait, Macao, Malaysia, Mexico, New Zealand, Saudi Arabia, Singapore, Taiwan, Thailand and Viet Nam. | Open, Classified |  |  | 16648 | Not_Computer
Z-0112-2016 | 72221 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System. |  |  | September 15, 2015 | 2015 | Elekta, Inc. | Dose and MU are incorrect when CT images are viewed from the head, and, when using multiple prescriptions with forced densities. | Device Design | The firm issued an URGENT IMPORTANT FIELD SAFETY NOTICE on 09/15/2015. The letter identified the affected device, as well as the reason for the recall, and clinical impact. Customers were informed that both issues would be resolved in Monaco Patch Release 5.10.01. The attached acknowledgement form should be completed, signed, and returned to a local Elekta Office or representative within 30 days. Questions regarding the notice should be directed to a local Elekta office. |  | 154 | 154 units | Worldwide Distribution -- US, including the states of CA, IN, MI, MO, TX and WI; and, the countries of Australia, Canada, Germany, Greece, India, Netherlands, New Zealand, Turkey, and the UK. | Terminated | April 04, 2017 | 567 | 154 | Not_Computer
Z-0109-2016 | 72176 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MICroSTREP plus 2 Panel, Part No. B1027-202 |  |  | September 10, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling MICroSTREP plus 1 and MICroSTREP plus 2 panels because a manufacturing error may lead to antimicrobic degradation and the potential for elevated or false resistant results. | Equipment maintenance | Beckman Coulter Inc. sent an Urgent Medical Device Recall letter dated Setember 10, 2015 to all customers who purchased the MICroSTREP plus 1 and MICroSTREP plus 2. The letter identified the affected products, problem and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions regarding the notice, are instructed to contact their Customer Support Representative:website: http://www.beckmancoulter.com or by phone: 1-800-677-7226 in the United States and Canada. |  | 438 | 438 units total (433 units in US) | Worldwide Distribution - US Nationwide and the countries of Bahamas, Belgium, Canada, Ecuador, Germany, Italy, Republic of Korea, Mexico, Peru, Poland, Romania, Saudi Arabia, South Africa, Thailand, United Sates (including Puerto Rico), and Uruguay. | Terminated | January 22, 2016 | 134 | 438 | Not_Computer
Z-0111-2016 | 72113 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 Intramedullary (SP2 IM) Rod 400mm Instrument (PN 96-6120) |  |  | September 08, 2015 | 2015 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is initiating a voluntary recall of select lots of the Specialist 2 Intramedullary Rod 400mm Instrument (PN 96-6120) due to complaints of the rod fracturing and leaving a portion of the rod in patients. | Device Design | DePuy Synthes sent an Urgent Information Recall Notice to all affected customers. Instruments are not lot traceable in the system, so the company sent notices by e-mail on September 8, 2015, and US Distributors will hand deliver notices to all Medical Facilities that may have purchased the affected instrument. The US Distributor or US Sales Consultant will deliver the Medical Professional Notice and Reconciliation Form to the affected Medical Facilities. If assistance is needed, Sales Consultants will assist customers with inspection of all instruments, return of affected lots, receiving credit and returning the completed Medical Facility Reconciliation Forms to DePuy Orthopaedics, Inc. The affected instruments will be returned to the quarantine warehouse located at 700 Orthopaedic Drive, Warsaw, IN 46582. For questions regarding this recall call 574-371-4917. For questions regarding this recall call 574-267-8143. |  | 7488 | 7488 | Worldwide Distribution - US including AL; AZ; CA; CO; CT; DC; FL; GA; HI; IA; IL; IN; KY; LA; MA; MD; ME; MI; MN; MO; MS; MT; NC; NE; NH; NJ; NV; NY; OH; OR; PA; RI; SC; TN; TX; UT; VA; VT; WA; WI; and WV and Internationally to ARGENTINA; AUSTRALIA; AUSTRIA; BELGIUM; BRAZIL; CANADA; CHILE; COLOMBIA; CZECH REPUBLIC; EGYPT; FRANCE; GERMANY; GREECE; INDIA; ISRAEL; ITALY; KOREA; Mexico; NEW ZEALAND; NORWAY; P R CHINA; PANAMA; PERU; POLAND; RUSSIA; SINGAPORE; SLOVENIA; SOUTH AFRICA; SWITZERLAND; UAE; UK; URUGUAY; VENEZUELA; and VIETNAM. | Open, Classified |  |  | 7488 | Not_Computer
Z-0114-2016 | 72170 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Disposable Paper Filter for Sterilization Containers |  |  | September 17, 2015 | 2015 | Stryker Craniomaxillofacial Division | Stryker has become aware that the inhomogeneity of the filter paper might potentially compromise the ability of the filter to maintain a sterile barrier during post-sterilization shelf-life. | Device Design | An Urgent Medical Device Recall letter has been sent to the customers. The affected product was identified, as well as the reason for the recall and risk to health. The letter states that customers should inform users of the Medical Device Removal and pass the notice to all appropriate individuals within the organization; scrap all affected products available at the location or return it to Stryker Craniomaxillofacial, Attention: Recall Coordinator, at the address provided. If there are questions regarding the impacted product, don't hesitate to Contact Stryker Customer service and refer to PFA #2015-032. Complete and sign the enclosed Business Reply Form and fax or copy to: (877) 648-7114 or e-mail a copy to CMF-custserv@stryker.com. |  | 915 | 915 packs | Worldwide distribution -- US, Canada, France, Germany, Italy, Japan, Norway, Sweden, Switzerland, and the Netherlands. | Terminated | February 26, 2016 | 162 | 915 | Not_Computer
Z-0113-2016 | 72242 | No | FEQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MAJ-1606 Instrument Channel Adaptor; a single use adapter which enables a connection between the endoscope's instrument channel port and the O FP-2 pump's tubing. |  |  | September 11, 2015 | 2015 | Olympus Corporation of the Americas | small puncture marks in a sterile package for an accessory to an Olympus device | Process control | A letter of urgent device safety information, dated September 18, 2015, was sent to direct customers to inform them of the issue and instructions to cease use immediately. Customers were also asked to return the completed response form along with affected devices. |  | 31060 | 31060 | US (nationwide) and Canada. | Terminated | March 13, 2017 | 549 | 31060 | Not_Computer
Z-0017-2016 | 71923 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba Aquilion CT System TSX-302A |  |  | May 08, 2015 | 2015 | Toshiba American Medical Systems Inc | If two specific operations are performed in multi-phase helical scanning, the acquired raw data may not be saved. | Radiation Control for Health and Safety Act | Toshiba America Medical Planned Action to bring defect into Compliance:a. A notification will be sent to all consignees with affected systems in compliance with 21 CFR1 003.21 upon approval by the Secretary.b. To correct this issue revised software will be installed on each system to prevent this occurrence from happening. Each consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.c. Each consignee will be advised to do the following until the corrective measures have been taken, i.e., revised software installed on their systems.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Toshiba America Medical Systems, Inc. will implement this CAP byFebruary 1, 2016.If you have any questions please contact your local Toshiba Representative at (800) 521-1968. |  | 459 | 459 | US Distribution including Puerto Rico | Open, Classified |  |  | 459 | Computer
Z-0116-2016 | 72247 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Distal Access Catheters - 044 - 115 cm: |  |  | August 19, 2015 | 2015 | Concentric Medical Inc | One Distal Access Catheter intended for distribution in Korea, was found at a US location. | Employee error | Issue was discovered by a sales representative during a visit for a case in Kentucky, where multiple Distal Access Catheters (DAC's) were being used. It was noticed that the reference number on one DAC carton did not match, even through the product labeling appeared to be the same.The mismatched unit was removed and returned to Concentric (Stryker).A formal recall letter was not issued as the product was returned after discovery.For further questions please call (510) 413-2263. |  | 1 | 1 | US Distribution to one location in KY. | Terminated | October 09, 2015 | 51 | 1 | Not_Computer
Z-0115-2016 | 72166 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdianÂż ViewRayÂż Radiation Therapy System, ViewRayÂż |  |  | August 27, 2015 | 2015 | Viewray Incorporated | ViewRay discovered that in the event that an encoder breaks or fails on the Patient Handling System (PHS, or couch), when attempting to restart the system, the couch could move unexpectedly. | Device Design | ViewRay sent an Customer Advisory Notice dated August 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It is recommended that if you have a patient on the couch and you receive the fault ''The patient table has experienced an unintended movement. This may be due to the patient or operator leaning on the table.Clear the fault and retry. If the problem persists call Service" on the TPDS,first attempt to clear the fault.If it does not clear and requires you to restart the couch software (same switch as restarting the RTCS) you should manually remove the patient from the couch before restarting the system.A software patch to address this issue is planned for early 2016. In the meantime please practice our recommendation of removing the patient from the couch prior to restarting the RTCS. If you have any questions or concerns regarding this issue, please do not hesitate to contact ViewRay Customer Support at support@viewray.com or call the ViewRayÂżÂż Customer Support center at 855-286-8875. |  | 4 | 4 units | US Distribution to the states of : CA, MO and WI., and Internationally to Korea. | Terminated | June 09, 2017 | 652 | 4 | Computer
Z-0117-2016 | 71991 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne Software v5.94, Catalog/Part Number: 482-649600. |  |  | September 11, 2015 | 2015 | Natus Neurology Inc | When using the NicoletOne v5.94 software, after exiting the impedance check function and returning to the EEG screen, the impedance check signal remains active in waveform, obscuring the EEG signals. | Software change control | Natus Neurology sent an Ugent Medical Device Corrrection Notification letter dated September 9, 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were informed consignees that a Field Service Technician will contact them to install the new software. Once the software is installed they will ask the Recipient or Field service Technician to complete and return the Verification Form.For questions regarding this recall call 608-829-8600. |  | 101 | 101 (88 US, 13 OUS) | Worldwide distribution. US nationwide, INDIA, ROMANIA, DENMARK, JAPAN, AUSTRALIA, RUSSIA, BANGLADESH, SRI LANKA, BRUNEI, and UNITED ARAB EMIRATES. | Terminated | March 08, 2016 | 179 | 101 | Computer
Z-0120-2016 | 72165 | No | OUT |  |  | Neurology | PMA | -Pipeline Embolization Device |  |  | September 04, 2015 | 2015 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | The firm is recalling Pipeline and Pipeline Flex Embolization Devices from U.S. since the devices were shipped with an EU version of the Instructions for Use. | Process control | The firm, Medtronic, sent a "Notification of Recall" letter dated 09/04/15 to its customers. The letter described the product, problem and actions to be taken. The firm's letter instructed that customers return the Pipeline Embolization Devices to Medtronic for replacement with the US version of product. The firm asks customers to take the following steps:1) Stop using this product immediately.2) Segregate this product from other Inventory.3) Fill out the reply form and fax to the completed form directly to Customer Service at (800) 716-6700 or email to: andrew.schramm@medtronic.com. Customer Service will issue an RGA number for the return of the affected product back to Medtronic Neurovascular.4) Upon receipt of an RGA number, write the RGA number on the FEDEXslip and send the product back to: Medtronic Neurovascular, Attention: Product Recall, 9775 Toledo Way, Irvine, CA 92618. Prepaid and preaddressed FEDEX slips are enclosed for convenience.5. Once they receive the returned units, product replacement units will be shipped directly to customers. The local Neurovascular representative can also be available to assist in expediting the exchange of product.For any questions contact directly at 949-680-1340. |  | 23 | 23 devices | US Distributions to states of: TN, KY, VA, WI, MO, IL, UT, OR, and MD. No foreign consignees. | Terminated | May 18, 2017 | 622 | 23 | Computer
Z-0119-2016 | 72123 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail CM Long. Intramedullary Fixation Rod. |  |  | September 03, 2015 | 2015 | Zimmer Gmbh | In about 5% of tested samples, the outer tray was found to be compromised. Risks: A damaged outer tray detected during surgery may cause a delay in surgery to obtain another implant. If a compromised packaging is not detected prior or during a surgery by the user, it may lead to periprosthetic infection. This may result in revision or multi-stage revision to treat the infection. | Package design/selection | On 9/3/2015, URGENT MEDICAL DEVICE RECALL notifications were sent via courier to the affected distributors and Hospital Risk Managers with instructions for locating and inspecting the affected devices and contacting Zimmer if damage is observed. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have questions concerning this notification, please call 1-877-946-2761. |  | 57603 | 57,603 units | US Nationwide distribution. | Terminated | March 15, 2016 | 194 | 57603 | Not_Computer
Z-0126-2016 | 72243 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Screw Inserter; for use with Intramedullary Rod, orthopedic device. Used to secure the head element to the head element inserter to prevent disengagement during head element insertion. |  |  | September 21, 2015 | 2015 | Synthes (USA) Products LLC | Certain lots of the Screw Inserter may be etched with incorrect graphics. The orientation of the etched image of the oblique angle on the Screw Inserter may not correspond to the tip of the angle of the instrument. | Device Design | DePuy Synthes sent an urgent notice of medical device recall letter dated September 21, 2015, to all affected customers to notify them about the issue, the potential risks, and to provide instructions for return of the completed response form and affected devices. Customers with questions were instructed to call 610-719-5450. |  | 793 | 793 | US Nationwide Distribution | Open, Classified |  |  | 793 | Not_Computer
Z-0127-2016 | 72244 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | 4FR Single Lumen Pressure Injectable PICC; intravascular therapeutic catheter, for long-term use greater than 30 days. |  |  | September 18, 2015 | 2015 | Arrow International Inc | Labeling inaccuracy in the Instructions for Use (IFU) and Patient Card. The IFU and Patient Card with these products state the PICC can withstand up to ten (10) pressure injections, however, the PICC has only been tested to withstand five (5) pressure injections. | Under Investigation by firm | An urgent medical device notification, dated September 22, 2015, was sent to direct accounts to alert them about the issue, the potential risks, and provide instructions for returning the completed response form and affected devices. |  | 31590 | 31590 | Worldwide distribution. US Nationwide, Australia, Canada, and New Zealand. | Open, Classified |  |  | 31590 | Not_Computer
Z-0125-2016 | 72133 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Inserter for Titanium Elastic Nails (TEN), orthopedic device |  |  | September 03, 2015 | 2015 | Synthes (USA) Products LLC | potential for mechanical failures such as breakage. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated September 3, 2015, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer. ACTIONS REQUIRED:We have record that your facility has received the product(s) subject to this recall. It is important to note that the removal of the product would prevent the performance of emergency surgery, thus DePuy Synthes is not requiring an immediate removal of affected product(s). DePuy Synthes has developed a replacement &amp; recovery plan for indicated devices and replacement parts are becoming available. If you have questions on replacement timing and availability, please call DePuy Synthes at 1-800-479-6329.If you have questions regarding the return and recovery of the part, we ask that you contact your local DePuy Synthes Sales Consultant. |  | 1550 | 1550 | US (nationwide) distribution and one Internationally to Australia. | Open, Classified |  |  | 1550 | Not_Computer
Z-0122-2016 | 72120 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Multianalyte Controls, PN 88769 and 88773. |  |  | August 27, 2015 | 2015 | Alere San Diego, Inc. | Alere San Diego is recalling the Alere Cholestech LDX Multianalyte Control because the assigned control ranges for the Total Cholesterol (TC) and Triglyceride (TRG) analytes are incorrect. | Error in labeling | Customer notification letters dated 8/25/15 and 9/3/15 were sent to customers to inform them that Alere San Diego is recalling the Alere Cholestech LDX Multianalyte Control because the assigned control ranges for the Total Cholesterol (TC) and Triglyceride (TRG) analytes are incorrect. Customers are informed that the Alere Cholestech LDX Total Cholesterol and Triglyceride assays continue to perform as expected. However, if customers' facility has used the control lot to test Total Cholesterol or Triglycerides as part of their Quality Control program, Alere San Diego recommends that they verify previous results or perform control testing using the replacement Expected Value Card or replacement control materials. The letters also provide customers with actions to be taken. Customers with questions are instructed to contact: Alere San Diego, Inc. 9975 Summers Ridge Road San Diego, CA 92121 U.S.A. Phone: 877-308-8289, FAX: 866-333-9839, E-mail: Verifications.ts@alere.com. |  | 190 | 190 kits | Nationwide Distribution. | Terminated | December 22, 2015 | 117 | 190 | Not_Computer
Z-0123-2016 | 72018 | No | OQY | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | The Ebb Complete Tamponade System, model #CTS-1000 and Code #OQY. |  |  | January 05, 2015 | 2015 | Clinical Innovations, LLC | Clinical Innovations, LLC is voluntarily recalling the Ebb Complete Tamponade System (CTS-1000) Lot Number 1214-F-403.) to address the potential for a balloon leak which may affect the ability to deliver therapy, when needed. | Other | Consignees were notified via email on 08/21/15. All consignees were asked to acknowledge the recall and return devices. |  | 827 | 827 units | Nationwide to MS, MD, TN, NC, IN, FL, AL, and LA, and to France. | Terminated | April 21, 2016 | 472 | 827 | Not_Computer
Z-0121-2016 | 72168 | No | IMJ | Shortwave diathermy. | Physical Medicine | Physical Medicine | 510(k) | Monode, Shortwave Accessory |  |  | August 21, 2015 | 2015 | DJO, LLC | Overheating and burning of the wire insulation inside the applicator. | Device Design | The firm sent out two customer notification letters; one to distributors and one to direct users. Letters were sent September 1, 2015 via USPS certified return receipt.Distributors were asked to obtain a list of their customers with contact information that have purchase or received the monode accessory or Intelect Shortwave device that included the monode as an accessory item. Once the list has been assembled the firm requests that it be forwarded to DJO Global Chattanooga Product Support by phone 877-733-4888, or by e-mail at monode.support@djoglobal.com. The firm's notification states that customers will be sent an upgraded monode accessory along with a pre-paid shipper for the return of the current used unit.Any questions the letter states to contact Product Support per the contact information above or William Fisher, Global Regulatory Compliance at 760-734-3126.Letters sent to direct users received the same instruction of returning the product for a replacement. Any questions the letter states to call 877-733-4888 or email at monode.support@djoglobal.com. |  | 2887 | 2887 | Worldwide Distribution - US Nationwide in the states of WY, OH, WA, NC, FL, NE, GA, KS CA, MO, IN, MA, IL, OK, AZ, TX, AR, UT, NJ, WI, PA, SC, WV, NY, MI, MN, VA, AK, TN, OR, NC, NH, KY, CO, NV, ID and the countries of El Salvador, China, Ecuador, South Korea, Chile, Israel, Peru, Nicaragua, Bolivia, Indonesia | Terminated | October 24, 2016 | 430 | 2887 | Not_Computer
Z-0124-2016 | 72330 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.7MM/3.5MM VA-LCP Medial Distal Humerus Plates 10H/RT/186mm- Sterile and Synthes 2.7MM/3.5MM VA-LCP Extended Medial Distal Humerus Plates 10H/RT/189mm-Sterile |  |  | September 28, 2015 | 2015 | Synthes (USA) Products LLC | The 2.7MM/3.5MM VA-LCP Humerus Plates were found to be labeled incorrectly. Part number 02.117.410S was incorrectly etched with part number 02.117.610S | Labeling mix-ups | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 9/28/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to"Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s)."Complete the Verification Section (page 3 of this letter), (even if you do not have the product) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided."Return the Verification Section (page 3 of this letter) with the product to:"Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132."Send a copy of the completed Verification Section by:"Fax: 844-294-7175 or"Scan/email: Synthes4672@stericycle.com"Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450. |  | 36 | 36 | US Distribution to states of:OH, IL, PA, WV, FL, AZ, CA, AL, ND, KY, MN, NJ, MI, GA, IN, TN, VA, NC, and LA. | Open, Classified |  |  | 36 | Not_Computer
Z-0133-2016 | 72078 | No | LZA | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Ambitex NMD400 Nitrile Exam Gloves size medium. General Hospital and Personal use. |  |  | August 25, 2015 | 2015 | Cardinal Health | Ambitex Nitrile NMD400 Exam Gloves size medium on exam hold were inadvertently placed into commerce. The product was under Exam Hold because samples pulled from a shipment manufactured by the firm failed the water leak test performed by FDA as described at 21 CFR 800.20. | Nonconforming Material/Component | The firm sent an Urgent: Product Recall letter dated 8/25/2015. The firm is requesting the user discontinue using and return any remaining stock of the AMBITEX Nitrile Exam Gloves because the gloves were inadvertently shipped prior to U.S. FDA releasing them for sale.Additionally, the firm requests that the customer quarantine the affected examination gloves; and return a copy of the enclosed Acknowledgment Form confirming their receipt of the Urgent recall via fax to 216-651-9760. Customer Service should be contacted at 800-GLOVES-0 to arrange for the return and credit of any on hand product that customer may have. If the customer has further distributed the affected gloves they are advised to notify their customers of the recall. |  | 34 | 34 cases. | Nationwide Distribution. | Open, Classified |  |  | 34 | Not_Computer
Z-0130-2016 | 72223 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System. |  |  | August 31, 2015 | 2015 | Elekta, Inc. | Incorrect Treatment Delivery Using Third Party Fixed Wedges. | Software design | Consignees were notified via URGENT FIELD SAFETY ALERT, on 08/31/2015. |  | 1440 | 1440 units | Worldwide Distribution-Domestic: AK, AL, AZ, CA, CO, CT, DC, FL, IN, IL, KY, MA, MD, MI, MN, MO, NJ, NC, ND, NY, OH, OK, PA, PR, TN, TX, VA, WA and WI; International: Albania, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, France, Gabon, Germany, Greece, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Macedonia, Malaysia, Mali, Malta, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Suriname, Switzerland, Taiwan, Thailand, Turkey, The United Kingdom, Ukraine, Venezuela and Vietnam. | Open, Classified |  |  | 1440 | Not_Computer
Z-0131-2016 | 72208 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Humeral Socket Shell Trial, Part Number 804-06-052 |  |  | September 15, 2015 | 2015 | Encore Medical, Lp | Difficulty assembling the shell trial to the broach and the retaining ring disassembled from the screw. | Device Design | The firm, DJO Global, sent an "URGENT FIELD SAFETY NOTICE" letter dated 9/16/15 to affected consignees/customers. The letter described the product, problem, and actions to be taken. The consignees/customers were instructed quarantine the device, pass the notice on to all who need to be aware within your organization, contact customer service to place a replacement order and receive an RMA number; and to complete a response form and return the recalled device for a replacement.If you have any questions, please call (512) 834-6302 or by email at desiree.hubby@djoglobal.com. |  | 20 | 20 units | US Distribution to states of: PA, FL, ID, and CA. | Terminated | January 10, 2017 | 483 | 20 | Not_Computer
Z-0132-2016 | 71260 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ONE-LINK NON-DEHP Y-TYPE MICROBORE CATHETER EXTENSION SET NEEDLE-FREE IV CONNECTOR WITH NEUTRAL FLUID DISPLACEMENT |  |  | May 21, 2015 | 2015 | Baxter Healthcare Corp | Potential for a leak at the tubing to luer bond. | Packaging | Baxter sent their consignees a recall letter dated 05/21/2015 with the following instructions:Actions to betaken if product was purchased directly fromBaxter :1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001 betweenthe hours of 7:00 am and 6:00 pm Central Time, Monday through Friday. Pleasehave your Baxter 8-digit ship-to account number, product code, lot number(s) and quantity of product to be returned ready when calling.3. Complete the Âżenclosed Baxter customer reply form and return it to Baxter byfaxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com .Returning the Baxter customer reply form promptly will prevent you from receiving repeat notices.4. If you distribute this product to other facilities or departm ents within yourinstitution, please forward a copy of this communication to them.5. If you are a dealer, wholesaler, distributor/reseller, or Original EquipmentManufacturer (OEM) that distributed any affected product to other facilities,please conduct a consumer-level recall of the affected product that you distributed to customers.Action to be taken if product was purchased from a distributor:1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.The Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00pm central time, Monday through Friday.3. Pl ease note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to the |  | 23600 | 23,600 | Domestic Only | Terminated | May 10, 2017 | 720 | 23600 | Not_Computer
Z-0134-2016 | 72134 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 5.0mm Unit Rod 270mm, 5.0mm Unit Rod 290mm, 5.0mm Unit Rod 310mm, |  |  | September 04, 2015 | 2015 | Synthes (USA) Products LLC | This product was produced using a finishing process not identified as part of the manufacturing specification. The process used with the lots subject to this Recall was a bead blast process. (Bead Blasting vs. Shot Peened). | Process control | An Urgent Notice: Medical Device Recall, dated September 4, 2015, was sent to end users to alert them about the issue and possible risk to patients. Customers were requested to follow the actions to be taken for if they have affected product or not; complete the response form, and return affected product. Customers can call 610-719-5450 or a local Synthes Sales Consultant with any questions. |  | 492 | 492 | Nationwide Distribution | Open, Classified |  |  | 492 | Not_Computer
Z-0135-2016 | 68228 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | All Philips Ingenia, Intera, Achieva and Multiva MR systems using R5.1i and R5.1 .2 version of software. Indicated for use as a diagnostic device. It can produce cross-sectional images, spectroscopic images and/or spectra in any orientation of the internal structure of the head, body, or extremities. |  |  | May 05, 2014 | 2014 | Philips Medical Systems, Inc. | In spine clinical workflows, cross reference lines may be used to determine the position of slices. In cases, where MobiView fused Images are used to show the cross reference lines, the cross reference lines may be positioned incorrectly. | Software design | Consignees were notified of the recall by letter sent May 2014. They were instructed to do the following: Do not perform planning or review planning of transversal images on fused sagittal images which were generated through Mobi View postprocessing Instead, review the planning of transversal scans only on the unfused stations. A Field Change Order with reference FC0781 00431 is being released that PHILIPS requires Philips field service engineers to disable the MobiView post processing software. When possible this will be done remotely. Once it becomes available, a software update will be installed which will address the reference line positioning issue with MobiView Should you need to communicate with Philips with regard to this program, please reference FC0781 00431. If you need any further information or support concerning this issue, please contact your local Philips representative or the Technical Support Line: 1-800-722-9377. |  | 317 | 317 | Worldwide Distribution-US: AL AR AZ CA CO FL GA IL IN KS KY MA MI NY OH OR PA SC TX VA VT WA OUS: Argentina Australia Austria Bangladesh Belgium Brazil Canada Chile China Colombia Denmark Ecuador Estonia Ethiopia France Germany India Indonesia Iraq Ireland Italy Japan Jordan Korea, South Malaysia Mexico Netherlands New Zealand Norway Oman Palestine Poland Russian Federation South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates United Kingdom. | Open, Classified |  |  | 317 | Computer
Z-0137-2016 | 72366 | No | GWO | Preformed alterable cranioplasty plate. | Neurology | Neurology | 510(k) | Thinflap System Case |  |  | August 06, 2015 | 2015 | Biomet Microfixation, LLC | The lids on implant caddies within P/N 17-5000 and 19-5000 could not be opened by the user without application of undesirable/unexpected excessive force. | Component change control | Biomet Microfixation made phone calls on August 6, 2015 and an initial notice letter sent via FedEx on August 6, 2015, to all distributors. During the phone conversation the distributors confirmed that the inventory invoiced to them is currently in their possession and were asked to remove the lids from the trays and return them back to Biomet Microfixation. In conjunction to the Initial Notice, replacement lids were sent to their business address. FollowÂżup communication would be distributed until all units are returned to Biomet Microfixation.Distributors were instructed to do the following: 1. Immediately locate and remove from circulation the items invoiced to your account as identified on this notice.2. Remove the lids from the quarantined units and replace with the lids (15-5001-18 or 19-5001-18) received with this Initial Notice. The new lids are being distributed to address the issue with the trays.3. Return the lids removed from the quarantined units to Zimmer Biomet CMF and Thoracic team with the labeling information provided with this notice.RETURN THIS LID: lot number(101337 or 101~65) located on the back of the lidDO NOT RETURN NEW LID: Part numbert 15-5001-18 or 1~5001-18 respectively4. Use the FedEx label in this packet to return the lids; this will be Zimmer Biomet's method of tracking your return. If distributors have further distributed this product, they must notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field safety notices. Please understand it is your responsibility to locate and quarantine these products. Customers with questions were instructed to call at 1-800-874-7711 or 904-741-4400, extension 9574, Monday through Friday, 8am-4pm ET. |  | 5 | 5 units | Nationwide Distribution including CA, FL, OK, NY, NC, and IL. | Open, Classified |  |  | 42 | Not_Computer
Z-0152-2016 | 72112 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PORT-A-CATH Plastic Hub Needles, 0.9mm (20G) x 19mm (3/4 in.) 90-degree. REF 21-2287-24 |  |  | September 14, 2015 | 2015 | Smiths Medical ASD, Inc. | The label on the lid stock of the individual needle packages for a portion of Lot number 46262 contains the incorrect part number, part number description and product image. The part number should be 21-2287-24 Port-A-Cath Plastic Hub Bent Needle. | Labeling mix-ups | The firm, Smiths Medical, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 9/14/2015 to customers. The letter described the product, problem, and actions to be taken. The customers were instructed to Inspect your inventory for and quarantine the affected product, complete and return attached Confirmation form via fax to: 1-651-747-4959 or email to: recall.response@smiths-medical.com within 10 days; upon receipt of completed form, customer representative will contact you and arrange for exchange of your unused affected product; send a copy of letter and confirmation form to your customers identified, replacement product will be arranged for your customers; provide email confirmation of your customer notification to Smiths Medical at recall.response@smiths-medical.com, and notify downstream customers who may have received the product. If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at 1-800-258-5361. |  | 130 | 130 Cartons (1560 needles) | Worldwide Distribution: US (nationwide) to CA only and countries of: Australia, Belgium, Germany, Hungary, Italy, Sweden, and Switzerland. | Open, Classified |  |  | 130 | Not_Computer
Z-0150-2016 | 72237 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Perifix Continous Epidural Anesthesia Pain Control Trays; Injection of anesthetics to provide regional anesthesia for pain management. |  |  | August 27, 2015 | 2015 | B. Braun Medical, Inc. | Potential discoloration of the DuraPrep solution from unintended exposure to the sterilant during sterilization performed by the supplier. | Under Investigation by firm | BBMI provided written notification, dated August 27, 2015, to all customers in receipt of suspect product, along with acknowledgement forms detailing the suspect product that was distributed to each account. These customers include both BBMI direct end customers and BBMI direct distributors. Customers were notified by US Postal Service Certified Mail with registered return receipt mail or FedEx Priority. All customers were provided with instructions for disposition of suspect product in their inventory. Distributors shall be reminded of their responsibility in executing the recall. |  | 433 | 433 | Nationally | Open, Classified |  |  | 433 | Not_Computer
Z-0149-2016 | 72017 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE 3 Phototherapy system or Neonatal Phototherapy, Catalogue Part Number: 010066; Top Assembly Part Number: 001103 |  |  | August 05, 2015 | 2015 | Natus Medical Incorporated | The neoBLUE Instruction For Use (IFU) and Service Manuals for the neoBLUE 3 were revised to show the cleared intensity statement settings so the output intensity aligns precisely with the information listed in the cleared 510(k) submission. | No Marketing Application | Natus Medical starting sending the Product Labeling Change NeoBlue LED Phototherapy System letter, dated August 2015, to all affected customers. The firm also attached the electronic copy of the revised User Manual with the letter.Customers are informed that the user and service manuals related to intensity for the neoBLUE 3 system are revised. Also, while the neoBLUE system does allow the light output to be adjusted with the potentiometers to accommodate different treatment heights, the firm removed the phrase "adjustable to &gt;50 ÂżW/cmÂż/nm" from the specifications.If you would like a CD disk containing the updated User Manual and the Service Manual at no charge, please call Natus Technical Service at 888-496-2887. |  | 4560 | 4560 units | Nationwide Distribution. | Terminated | March 17, 2016 | 225 | 4560 | Not_Computer
Z-0143-2016 | 72108 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConForMIS iDuo Knee Bicompartmental Knee Replacement System , iDUO G2, Right Medial |  |  | August 31, 2015 | 2015 | ConforMIS, Inc. | May contain small amounts of ethylene glycol residue | Process design | ConForMIS issued a Press on August 31, 2015 to the general public The sales force was notified through email. The sales personnel for affected accounts that had upcoming surgeries scheduled were asked to confirm that the affected surgeries were cancelled and that the devices were not used in surgery (Formal notification letters will be distributed to the affected accounts by Tuesday, September 8, 2015. The user required to confirm their understanding of the recall and to return any unused product to ConforMIS. Distribution of the letters conducted in two phases as described below:Phase 1-Surgeons with product for which the surgery has not yet been completedPhase 2-Surgeons for whom product was confirmed to have beenconsumed/implanted. An updated Surgeon letter dated September 8, 2015 issued includes additional information for devices that have been used in surgery.Contact your local ConforMIS representative or Amita Shah, at 17813459164. |  | 12 | 12 units | Distributed Nationwide and in Germany, Switzerland, and UK. | Terminated | March 02, 2016 | 184 | 939 | Not_Computer
Z-0154-2016 | 72273 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TELETOM EQUIPMENT MANAGEMENT SYSTEM 4 and 6 series |  |  | September 23, 2015 | 2015 | Berchtold Corp. | During the use of the TELETOM Series 4 &amp; 6 there is potential for the controls and brakes not to work properly due to fluid intrusion within the control switches and/or the control board assembly. | Device Design | Stryker sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Stryker Representative would visit each account for product inspection. During the inspection the Stryker Representative will remove the control switches and replace them with new control switches. Customers were instructed if any malfunction prior to inspection were noticed customers should call local Technical Support immediately at 800-243-5135. |  | 2129 | 2129 units | Worldwide Distribution - US (nationwide), Canada, and Germany | Open, Classified |  |  | 2129 | Not_Computer
Z-0153-2016 | 72011 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | OneTray Sealed Sterilization Containers |  |  | February 03, 2014 | 2014 | Innovative Sterlization Technologies LLC | The wrong sterilization time was on the label. The label lists Steam Gravity 10 - 34 minutes at 270 degrees F instead of Steam Gravity 17 - 34 minutes at 270 degrees F. | Process control | The firm sent an Urgent Medical Device Recall letter dated September 14, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to fill out the attached response SASE mailer and return to or fax to the address/number identified below:Innovative Sterilization Technologies, LLC, 7625 Paragon Road, Suite A Dayton, OH 45459, 937-286-3571 (C), 937-630-4346 (F), 937-619-0138 x 202 (Office), dbillman@ISTSterilization.com |  | 4621 | M2104 - 25 units; M2408 - 84 units | US Nationwide Distribution in the states of IL, IN, LA, OH, PA, TX &amp; WV. | Terminated | February 04, 2016 | 731 | 4621 | Not_Computer
Z-0156-2016 | 72205 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Smith &amp; Nephew MOSAICPLASTY DP- Disposable Harvesting System 4.5 mm, Sterile |  |  | September 10, 2015 | 2015 | Smith &amp; Nephew, Inc., Endoscopy Div. | Sterility of device is compromised due to breach in the packaging | Package design/selection | Smith &amp; Nephew issued Recall Notification Letter dated September 10, 2015 via overnight Federal Express. Customers were requested to inspect inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately. Complete the Inventory Return Certification Form on the following page, indicating the lot numbers and quantities that need to be returned and include your phone and fax number in the spaces provided. If you do not have product to return, please place an x in the column No product to return. Send the completed form to the Smith &amp; Nephew's Field Action Department via email at FieldActions@smithnephew. com or fax +1-901-566-7975 to obtain a return authorization (RA) number.Return any affected product to the address listed on the Inventory Return Certification Form. Please indicate the RA number on your return shipment. |  | 1287 | 1287 units | Worldwide Distribution - US Nationwide, Canada, Austria, Australia, Belgium, China, Spain, Finland, France, Great Britain, Greece, Germany, Hungary, India, Italy, Japan, Korea, Lithuania, Latvia, Netherlands, Poland, Portugal, Sweden, Thailand, Turkey, Ukraine and Venezuela | Open, Classified |  |  | 4483 | Not_Computer
Z-0159-2016 | 72305 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | CC cartridges for IRMA TRUPOINT Blood Gas Analyzer, PN 039903. |  |  | September 16, 2015 | 2015 | LifeHealth, LLC | May report values outside the published performance specifications for pH. After a certain point in shelf life, the pH sensor may give results that are falsely elevated. The discrepancy is greater as shelf life progresses and can be seen on the aqueous controls before the blood. | Nonconforming Material/Component | Consignees were sent a LifeHealth "Urgent Medical Device Recall" letter dated September 16, 2015. The letter described the problem and the product involved in the recall. The letter described the Risk to Health, Action to be taken and Other Information. Consignees were advised to discontinue use and destroy all stock on hand. Requested consignees to complete and return the Verification Form to LifeHealth. For additional information contact Nancy Ring at 1-855-762-8978 or 651-638-1000, Monday through Friday, 8:30 AM to 4:30 PM, Central Time or see our website at www.lifehealthmed.com. |  | 116475 | 4659 boxes of 25. (116475 cartridges) | US: Nationwide (AL, AR, CA, CO, FL, IL, IN, IA, KS, KY, MD, MA,MI, MN, MS, MO, NE, NJ, NY,NC, OH, OK, PA, TN, TX, VA)OUS: ALBANIA, BELGIUM, BOLIVIA, CANADA, CHILE, CHINA, DENMARK. DOMINICAN REPUBLIC, ECUADOR, EGYPT, ENGLAND, ESTONIA, PHILIPPINES, FRANCE, GERMANY, GREECE, ICELAND, INDONESIA, IRELAND, ITALY, JAMAICA, JAPAN, MOROCCO, NETHERLANDS, NORWAY, PERU, PORTUGAL, SOUTH AFRICA, SOUTH KOREA, SPAIN, TAIWAN, TUNISIA, TURKEY, UNITED ARAB EMIRATES, VIETNAM. | Terminated | March 21, 2016 | 187 | 159850 | Not_Computer
Z-0161-2016 | 72365 | Yes | LLB | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Nipro Safe Touch TULIP Safety Fistula Needle, FT+152530TP, for use as a blood access device for blood purification and for other treatments requiring an extracorporeal circuit of larger volumes of blood. |  |  | August 13, 2015 | 2015 | Nipro Medical Corporation | Some 1 1/4" needles were packaged and labeled in the 1" blister package. | Mixed-up of materials/components | Customers were issued a RMA 60001203 on August 18, 2015. |  | 50000 | 50,000 pieces | Indiana | Open, Classified |  |  | 50000 | Not_Computer
Z-0118-2016 | 72138 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | CIOS ALPHA; image intensified fluoroscopic x-ray system |  |  | August 17, 2015 | 2015 | Siemens Medical Solutions USA, Inc | patient procedure interruption due to a potential system failure | Device Design | Customer Safety Advisory Notice XP038/15/S, dated August 17, 2015, was sent to direct accounts to alert them to the issue and provide work-around instructions to avoid potential risks. |  | 44 | 44 | Nationwide Distribution. | Open, Classified |  |  | 44 | Not_Computer
Z-0162-2016 | 72425 | No | OSF | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Signature Patient-Specific Surgical Guides; patient specific instruments used as a guide during surgical procedure for total and partial knee arthroplasty. Part of the Signature Personalized Patient Care System. |  |  | September 23, 2015 | 2015 | Materialise N.V. | Potential inaccuracies in the alignment of the guides. Surgery was inaccurately planned because the patient image data was incorrectly processed. | Use error | On September 16, 2015 Materialise NV identified four cases where not all available image sets were used to design the case. On September 18, 2015, prior to the surgery dates Materialise informed (phone call) the SignatureServices Manager of Zimmer Biomet that inaccuracies in the alignment of the knee might be observed with the use of the guides due to this issue. Biomet was asked to segregate the packages for these four cases and to return them to the Materialise US office. On September 23, 2015 this wasfollowed up by e-mail notification letter with response forms.Biomet was requested to contact Materialise customer service to return these four cases and arrange a replacement set of patient specific surgical guides Monday through Friday, 8:00 AM to 4:30 PM, Eastern Time. E-mail : ortho-quality@materialise. beCustomer ServiceMaterialise USA LLC44650 Helm CourtPlymouth, Michigan 48170734-259-7033 |  | 4 | 4 | Distributed in MI, ID, TX, and in France. | Terminated | February 12, 2016 | 142 | 4 | Not_Computer
Z-0165-2016 | 72216 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exprt Precision System: Revision Knee, Model Number 160-010-726/738 |  |  | September 16, 2015 | 2015 | Encore Medical, Lp | The labeling is missing the size/diameter information. | Labeling design | DJO Surgical sent an Urgent Field Safety Notice dated September 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter requested the consignees complete and return a response card and return affected devices for replacement. For further questions call (512) 832-6302. |  | 209 | 209 units | US Distribution to the states of : SC, WI, RI, NY, MO, AR, OK, GA, FL, AL, CA, AZ and CO. | Terminated | January 04, 2016 | 110 | 209 | Not_Computer
Z-0164-2016 | 72311 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sentinel MULTIGENT Ammonia Ultra Reagent, In-Vitro Diagnostic for ammonia measurement; List Number: LN 6K89-30. Intended for the in vitro quantitative determination of Ammonia (NH3) in human plasma. |  |  | September 22, 2015 | 2015 | Sentinel CH SpA | Results for ammonia are below the linear range of the assay for samples that have been collected in sodium or lithium heparin tubes. | Process control | Sentinel Urgent Field Safety Notice issued on 9/22/15 by Abbott to customers, including their US customers.Actions to be Taken1. Immediately discontinue use of lot 50083Y600 and destroy any remaining inventory.2. Order a replacement Ammonia Ultra reagent lot.3. Follow your individual laboratory protocol regarding reviewing previously reported patient results.4. Please retain this communication for future reference. If you have forwarded any MULTIGENT Ammonia Ultra reagent lot 50083Y600 to another laboratory, please provide a copy of this letter to them.ContactsIf you or any of the health care providers you serve have any questions regarding this information, U.S. Customers should contact Customer Service at 1-877-4 ABBOTT. |  | 765 | 765 kits | Nationwide Distribution. | Terminated | March 30, 2016 | 190 | 765 | Not_Computer
Z-0163-2016 | 72154 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000SR System |  |  | September 09, 2015 | 2015 | Abbott Laboratories, Inc | Incorrect tubing. | Nonconforming Material/Component | Abbott Laboratories executed a Technical Service Bulletin on August 14, 2015.. The firm sent out a Product Correction notice on September 9, 2015 to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed that an Abbott representative will make all necessary arrangements for replacing tubing on the impacted instruments. Customers were asked to retain a copy of the notice for their records. Customers with questions were instructed to contact Customer Service at 1-877-4ABBOTT. Customers outside the US should contact their local area Customer Service.For questions regarding this recall call 877-422-2688. |  | 65 | 65 units | Worldwide Distribution - US (nationwide) and Internationally to France, Netherlands, Greece, Italy, Belgium, India, Austria, Germany, and Japan | Open, Classified |  |  | 65 | Not_Computer
Z-0166-2016 | 72013 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer Specimen Gate, Screening Center. Data management of neonatal screening test results and demographics by qualified laboratory personnel in newborn screening programs. |  |  | August 24, 2015 | 2015 | Perkin Elmer Life Sciences Inc | It would be possible for an abnormal analytical test result value from an affected newborn to be reported with a default result interpretation of normal. The analytical test result value would be correct but the interpretation of the test result value would be incorrect. | Software design | Perkin Elmer issued the Recall Letter on 8/24/15 delivered to the customer by email .Software Services will be asked to make sure the customers receive the Recall Letter and implement the recommended corrective measures: Ensure all laboratory routine sample testing procedures are followed to avoid procedural anomalies. In the event a procedure anomaly occurs, review the final analytical result and the corresponding interpretation to confirm both are accurate before releasing the report. Once a PerkinElmer reporting error query is available, continue to run the query until the final corrective action is deployed to your site. The final corrective action will involve an installation update to the software.For further information, please contact your local PerkinElmer representative or specimen.gate.support@perkinelmer.com |  | 6 | 6 units | US distribution to FL, GA, and NV; and Canada | Open, Classified |  |  | 6 | Computer
Z-0168-2016 | 71207 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIMUS HI, MEVATRON M2/PRIMUS Mid-Energy, ARTISTE; intended use is to deliver x-ray radiation for therapeutic treatment of cancer. |  |  | May 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | potential safety risk of collision of the LINAC gantry with the patient or the treatment table in case of automatically sequenced treatment delivery techniques using SIMTEC" Auto Field Sequence Option. | Software design | Siemens sent Safety Advisory Notices, dated May 5th and 22nd, 2015 to all affected customers to inform them about an update of the Control Console software to correct defects affecting the Control Console with their digital Linear Accelerator. Customers were instructed to ensure everne at their organization is aware of the information in the Advisory Notice. If product was further distributed, please forward Safety Notice to all those affected.For questions regarding this recall call 610-219-6300. |  | 158 | 158 | Nationwide Distribution | Terminated | May 05, 2017 | 731 | 158 | Computer
Z-0183-2016 | 72104 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry XPT Chemistry Systems, Siemens Material Number (SMN) 10723034. |  |  | August 28, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | ISE module may produce discrepant, unflagged results for Sodium, Potassium or Chloride when the respective electrode is left unplugged after replacement, washing or maintenance activities. The calibration and 2-level QC will fail if the electrode is left unplugged. The IFU states to calibrate the ISEs after replacement and to run two levels of controls before running patient samples. | Device Design | Urgent Medical Device Correction letters (dated 8/27/2015) and Field Correction Effectiveness Check Forms were sent to customers via FedEx starting on 8/28/2015, informing them of the potential for discrepant ISE results. Customers must follow the instructions provided in the letter. The Actions to Be Taken section of the letter notes that customers must ensure that the electrodes are fully plugged in after replacement, washing or maintenance activities, followed by performing an ISE Calibration and running 2 levels of Quality Control material. Urgent Field Safety Notices were sent to customers outside the USA. The Field Correction Effectiveness Check form should have been completed and returned. A copy of the letter should be kept with laboratory records; and, the letter should be forwarded to anyone who may have received the affected product. Customers should contact their Siemens Customer Care Center or their local Siemens technical support representative with questions. |  | 74 | Domestic: 15 units; Foreign: 59 units | Worldwide Distribution -- US, Australia, Bahrain, Denmark, Egypt, Egypt, Finland, France, Germany, Italy, Netherlands, New Zealand, Norway, Republic Korea, Singapore, Slovakia, Spain, and the United Kingdom. | Open, Classified |  |  | 74 | Not_Computer
Z-0193-2016 | 72274 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | IntegraÂż Bone Marrow Biopsy Trays Catalogue No. 3404253 |  |  | September 22, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra received an Urgent Drug Recall Notice from Hospira Inc. for their 1% Lidocaine HCL Injection, USP, 10 mg/mL, Lot 44-359-DK. Hospira Inc. is recalling vials of Lot 44-359-DK, 1% Lidocaine HCL Injection, USP, 10 mg/mLm distributed by Hospira from February 2015 to March 2015 due to a confirmed complaint of visible, partially embedded particulate within a single-dose glass teartop vial. | Process control | Integra sent an Urgent Medical Device / Drug Recall Notice dated September 22, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed.to forward copies of the Integra and Hospira recall notification letters to their customers. Confirmation of recall notification delivery and instructed to stop use of devices and return to Integra.If you have questions of a customer service nature or beyond the information in the Hospira Inc. URGENT DRUG PRODUCT RECALL NOTICE, please contact either your local Integra Pain Management Sales Representative or customer service at 1-800-241-2210. |  | 470 | A total of 470 packages (47 cases) | US Distributed to the state of : MD. | Terminated | January 29, 2016 | 129 | 470 | Not_Computer
Z-0181-2016 | 72424 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco - Radiation Treatment Planning used to make treatment plans for patients with prescriptions for external beam radiation therapy. |  |  | October 16, 2015 | 2015 | Elekta, Inc. | Unintended update of Dose and MU and Incorrect Assignment of Bolus. | Other | Elekta sent Important Field Safety Notice 382-01-MON-003 to all customers on 10/16/2015. The notice informs users of the specific product and version numbers affected by both issues. The affected devices for both issues are those running Monaco 5.10.01. Customers are instructed to complete and return acknowledgement form. Solution/fix has been released in Monaco patch 5.10.02 and all affected users will be made aware of its availability by the end of October 2015. |  | 154 | 154 | CA, IN, MI,MO, TX, WI, Australia, Canada, Germany,m Greece, India, Netherlands, New Zealand, Turkey and United Kingdom | Terminated | April 04, 2017 | 536 | 154 | Not_Computer
Z-0170-2016 | 72374 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA Radiation Oncology, versions 10, 11 , 13.0 and 13.5 with Clinical Assessment License. |  |  | September 28, 2015 | 2015 | Varian Medical Systems, Inc. | An anomaly was identified with the ARIA for Radiation Oncology software with a Clinical Assessment license. Admin Instructions may not print on drug order prescriptions or be transmitted with e-Rx. | Software design (manufacturing process) | Varian sent an Urgent Medical Device Corrections letter dated September 29, 2015 via trackable method to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that Varian is developing a technical correction for this issue. Customers will be contacted by a Varian Field Service. A copy of the notification is to be kept with the most current product labeling and all appropriate personnel in the radiation oncology department should be aware of this issue. For questions contact your local Varian Medical Systems Customer Support District orRegional Manager |  | 1971 | 1971 units | Worldwide Distribution - US Nationwide and the countries of: Canada, Switzerland, South Korea, Germany, Spain, Ireland, Spain, France, Belgium, United Arab Emirates, United Kingdom, Norway, Italy, Trinidad and Tobago, Netherlands, Philippines, Portugal, Denmark, Sweden, Belgium, Saudi Arabia, Thailand, Australia, India, Taiwan, South Africa, Israel, Australia, Kuwait, Finland, Tunisia, Algeria, Argentina, Bangladesh, Brazil, Chile, China, Colombia, Costa Rica, Czech Republic, Guam, Guatemala, Hong Kong, Indonesia, Japan, Kazakhstan, Kuwait, Lebanon, Luxembourg, Madagascar, Moldova, Morocco, New Zealand, Panama, Poland, Qatar, Romania, Russia, Singapore, South Africa, Tunisia and Turkey. | Terminated | May 12, 2017 | 592 | 1971 | Computer
Z-0192-2016 | 72360 | No | EZX | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Synthes Threaded Plate Holder- Long for 3.5mm Locking Hold; Part Numbers 324.031. |  |  | October 07, 2015 | 2015 | Synthes (USA) Products LLC | The affected part number and lot listed above is incorrectly etched with the wrong part number and may not fit into a Small Fragment Plate. | Mixed-up of materials/components | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 10/5/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:" Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s)." Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided." Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Even if you do not have any product." Send a copy of the completed Verification Section by: Fax: 844-721-3048 orScan/email: Synthes4704@stericycle.com." Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 25 | 25 | Worldwide distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  | 25 | Not_Computer
Z-0189-2016 | 72052 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | Regard, Item Number 830500003, Sterile, CV0578C - CV PACK - Custom surgical kit for surgery. |  |  | August 25, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | Custom surgical kits contain a component (ORS fluid warming and slush drapes) that was recalled due to compromised package integrity potentially affecting sterility assurance. | Packaging | On August 25, 2015, the recalling firm sent an email to their direct customers informing them of the recall. The email instructed their customers to examine their inventory and determine whether they had any of the affected devices. The recalling firm attached the recall notice they received from their supplier.On September 4, 2015, the recalling firm sent a second email clarifying that they were instructing their customer to discard the supplier component . The notice also asked the customer to respond with how many affected devices they had in their inventory at the time they received the notice. |  | 17 | 17 kits | Distributed in Texas and Louisiana. | Open, Classified |  |  | 506 | Not_Computer
Z-0167-2016 | 72345 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.7 mm Locking Screw SLF-TPNG with T8 Stardrive Recess 12mm. The Synthes modular Mini Fragment LCP System is intended for fixation of fractures, osteotomies, nonunions, replantations, and fusions of small bones and small bone fragments. |  |  | October 06, 2015 | 2015 | Synthes (USA) Products LLC | Specific lots of the SS 2.7mm locking screw have been mislabeled as SS 2.4mm locking screw. | Packaging | Synthes sent an URGENT NOTICE: MEDICAL DEVICE NOTIFICATION UPDATED INFORMATION: CORRECTED LOT NUMBER SS 2.7mm Locking Screw, self-tapping letter dated January 13, 2014, to the affected customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number.Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Form (page 3 of this letter) with the product to:Credit/Returns, Synthes.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the documents to Synthes by: Fax: 1-866-792-5446 orScan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. For further questions please call 610-719-6500. |  | 98 | 98 | US Distribution to the state of : TN | Open, Classified |  |  | 98 | Not_Computer
Z-0182-2016 | 72372 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 ultrasound systems |  |  | October 07, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Siemens will provide all customers a new version of system software (VA35E or VB10C depending on the hardware configuration). Hardware fixes will be implemented to address situations where the metal component used to engage the wheel lock becomes disconnected, and where the system locks up due to misalignment between the bezel and MPI board, in all units that have not already been corrected. | Device Design | On August 19, 2015, Siemens Ultrasound began sending a customer notification to users, advising them to correct the uppercase/lowercase differences in the patient name at the time of registration until the issue is resolved with a future software release. Siemens believed that the other issues in this correction will begin approximately November 30, 2015 and December 31, 2015. Field service reps will visit sites to install software updates or hardware corrections. |  | 2099 | 2,099 units | Worldwide Distribution -- US and Albania, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Canada, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mexico, Netherlands, New Zealand, Nigeria, Norway, P.R. China, Pakistan, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2099 | Computer
Z-0178-2016 | 72282 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Flex Computed Tomography X-ray systems. Intended to produce cross-sectional images of the body. |  |  | September 14, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm discovered Ring/Dot artifact due to X-ray measurement error. | Software design | On 9-23-2015 the firm sent URGENT Field Safety Notices to their customers. The letter identified the products, problem, and actions to be taken by the customers. |  | 67 | 67 units | Worldwide Distribution-US (nationwide) including DC and the states of AL, AR, AS, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, SC, TN, TX, UT, VA, VT, WI &amp; WV, and the countries of Albania, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mexico, Mongolia, Myanmar, Netherlands, Nigeria, Norway, Oman, Panama, Philippines, Poland, Portugal, Romania, Russia, Rwanda, Samoa, Saudi Arabia, Singapore, South Korea, Spain, State of Palestine, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 643 | Not_Computer
Z-0172-2016 | 72211 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare CardioCall ECG Event Recorder, Models VS20, VS20x4, ST80, ST80x4, PCBAs PN # 670-1670-00 |  |  | September 03, 2015 | 2015 | Spacelabs Healthcare Inc | When the battery is inserted into the CardioCall ECG Event Recorder, the device beeps twice and is followed by a continuous ringing sound that only stops when battery is removed. The device has a self-test failure and will not operate normally. | Vendor change control | Spacelabs Healthcare sent an Urgent-Medical Device Correction letter, dated September 17, 2015 to US consignees (Hospital Administrator/Biomedical Manager/Office Manager). The firm plans to email a customer letter (translated as necessary) to international subsidiaries and distributors of record on 9/25/15. Customers were informed Spacelabs Healthcare will send a replacement CardioCall ECG Event Recorder(s) at no cost and a prepaid method to ease the return of the affected CardioCall device(s) back to Spacelabs.For additional information or technical assistance contact:Global Technical SupportSpacelabs Healthcare, Ltd.1 Harforde Court, John Tate RoadHertford SG13 7NWUnited Kingdom+ 44 (0) 1992 507700For USA Technical Support call 1-800-522-7025 and select 2 for Technical Support |  | 275 | 275 devices (20 in US and 255 international) and 113 PCBAs (international only) | Worldwide Distribution - US Distribution in the states of Nebraska, South Carolina, and Washington and in the following countries of AUSTRALIA, BELGIUM, BOLIVIA, CANADA, CHILE, DENMARK, GERMANY, GUATEMALA, HONG KONG, ITALY, NETHERLANDS, NEW ZEALAND, POLAND, PORTUGAL, QATAR, SINGAPORE, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 275 | Not_Computer
Z-0169-2016 | 72228 | No | HIB | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Welch Allyn KleenSpec Corded Illumination System, Vaginal Speculum Lighting System, Model Number 78x Series Illuminator, REF 901070, Catalog Numbers 78810, 78812, 78814 and 78816, Rx ONLY |  |  | October 02, 2015 | 2015 | Welch Allyn Inc | The firm became aware that the KleenSpecÂż 788 Corded Illuminator has the potential to be incorrectly connected to a commonly available AC power cord instead of the illuminator low voltage transformer power cord which is supplied with the illuminator. This may result in a shattered lamp bulb and a potential electric shock hazard. | Device Design | On 10/02/2015, Welch Allyn commenced mailing of the Urgent Medical Device Correction letters (dated 10/01/2015) via UPS Certified Next Day Air Saver Mail. The letters inform customers of the issue. To help customers avoid any such occurrences, Welch Allyn is providing supplementary User Precautions which re-emphasize safe use of this product and providing additional specificity to the warnings already on the labeling and packaging. |  | 194464 | Domestic: 179,401 units, Foreign: 15,063 units | Nationwide, Puerto Rico and foreign distribution to the following countries: Australia, Bahamas, Bahrain, Belgium, Bermuda, Cambodia, Canada, Czech Republic, China, Cyprus, Denmark, Egypt, France, Germany, Great Britain, Greece, Guam, Guatemala, Iceland, India, Israel, Japan, Kenya, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sri Lanka, Thailand, Turkey, United Arab Emirates. | Open, Classified |  |  | 194464 | Not_Computer
Z-0195-2016 | 72363 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Pak 9319-27. Single-use medical devices and accessories designed by medical professionals for a specific ophthalmic surgical procedure. |  |  | September 21, 2015 | 2015 | Alcon Research, Ltd. | The recalled lot of custom paks are labeled as latex free and contain latex gloves rather than the specified non-latex gloves. | Mixed-up of materials/components | The recalling firm notified the consignee on 9/21/15 via phone call, and with a follow up visit to the consignee site. The recalled units were recovered. |  | 5 | 5 units | Distributed in VT. | Terminated | March 24, 2016 | 185 | 5 | Not_Computer
Z-0194-2016 | 72362 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter 15L Cycler Drainage Bag Sterile, nonpyrogenic fluid path |  |  | October 06, 2015 | 2015 | Baxter Healthcare Corp. | Leakage from the inlet port of the 15L Cycler Drainage Bags. | Under Investigation by firm | Baxter sent an Urgent Product Recall letter dated October 8, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The recalling firm will ask customers to locate and separate all affected products, contact the recalling firm to arrange for a return, and complete the customer reply form. Further questions, please call (800) 422-9837. |  | 19440 | 19,440 units | US Distribution and Internationally to Canada. | Open, Classified |  |  | 19440 | Not_Computer
Z-0180-2016 | 72229 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reverse Shoulder Prosthesis Stem Impaction Fixture REF 804-03-053 |  |  | September 14, 2015 | 2015 | Encore Medical, Lp | During the assembly, the impaction forces caused the polymer, black acetal copolymer from the Impaction Fixture to wear off on the lateral surface of the Humeral Stem titanium plasma spray coating. | Device Design | DJO Global sent an Urgent Field Safety Notice dated September 15, 2015, to all affected consignees. The notice identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed become familiar with the new techniques and notify all of their applicable users of the RSP Monoblock and Turon stems of these revised methods. When each user has been notified of these new techniques, they should complete the form that was provided with the notification indicating that they have been informed of these changes and understand the new methods. Consignees were asked to complete and return forms. Consignees with questions should call 512-834-6255. |  | 634 | 634 units | Nationwide Distribution | Open, Classified |  |  | 911 | Not_Computer
Z-0196-2016 | 72448 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MR Surgical Suite II Table, Model Number(s): M8074SS, M8053SS, M0074SS. |  |  | September 23, 2015 | 2015 | GE Medical Systems, LLC | GE became aware of an issue with the table transfer release mechanism of the MR Surgical Suite Option. The issue may prevent patient transfer between the MR and surgical tables. | Process control | GE sent an "Urgent Medical Device Correction" letter dated September 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.You may continue to use the table. Monitor for any issues in releasing the table. If you see this issue, contact a GE Field Engineer.For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 23 | 23 (11 US; 12 OUS) Update 11-4-2015: 35 (14 US; 21 OUS) | Worldwide Distribution - US Distribution including the states of : CA, CO, FL, IL, MI, OH, TX and WI., and to the countries of : AUSTRIA, BRAZIL, FRANCE, GERMANY, JAPAN, POLAND and TURKEY. UPDATE 11-4-2015: Additional states: AZ, MN. Additional countries: Canada, China, Czech Republic, South Africa, and United Kingdom | Terminated | March 28, 2016 | 187 | 23 | Not_Computer
Z-0171-2016 | 72321 | No | LIZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD CD3 (SK7) FITC; IVD; Catalog # 349201 and 340542 |  |  | September 30, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Two lots of CD3 (SK7) FITC label contains an error in the Spanish Small text. | Under Investigation by firm | BD sent an "URGENT PRODUCT RECALL" letter dated September 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer Actions: 1. Discard all remaining stock of CD3 (SK7) FITC, catalog number 349201, lot number 5148598, and/or catalog number 340542, lot number 5113983, and request replacement product. BD will replace product at no cost.2.As a requirement of our Regulatory Tracking Process, we request that you complete the enclosed Tracking Verification form, even if you no longer have the affected reagents listed in your inventory. Please return the Tracking Verification form to BD Biosciences, per the instructions (mail, fax, or email).We apologize for any inconvenience that this may have caused and are working toward a long-term solution to this issue. If you require further assistance, please contact BD Customer Support at 877.232.8995 (prompt 3, 3) in the United States. For customers outside the US, contact your local BD Biosciences representative or distributor. |  | 120 | 120 units | US (nationwide) Distribution to the states of : TX, GA, IN, TN, CA, KY, WA, IL, MO, NY, CT, MN, NV, KS, WI, NH, FL, PA, OH, OR, SC. OK, NC and LA., and Internationally to Mexico. | Open, Classified |  |  | 120 | Not_Computer
Z-0197-2016 | 69082 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GEMINI PET/CT System, a diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. |  |  | May 01, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | The reason for this inspection is to determine if the Spindle Block Servo Controllers on the CT subsystem of potentially affected GEMINI PET/CT systems are susceptible to premature failure. | Unknown/Undetermined by firm | Philips field service personnel will conduct customer visits in order to inspect systems. For further questions please call (440) 483-7600. |  | 264 | 264 units | Worldwide Distribution - US Distribution to the states of : AL, AZ, CA, CO, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MO, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI, WV, and WY., and to the countries of : Canada, Argentina, Australia, Austria, Belgium, Brazil, China, Denmark, Finland, France, Germany, Hong Kong, Israel, Italy, Japan, Jordan, Luxembourg, Mexico, Netherlands, Poland, Portugal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey and UK. | Terminated | November 06, 2015 | 2745 | 264 | Not_Computer
Z-0200-2016 | 72115 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Achieve Cables, model 990066. The sterile, single use only Electrical Cable (Model 990066) provides the conduction elements from the proximal end of the Achieve Catheter handle to standard shielded ECG pins that connect into standard electrophysiology recording and pacing equipment. Medtronic Ablation Frontiers Achieve Mapping Catheter is indicated for multiple electrode electrophysiological mapping of the cardiac structures of the heart. |  |  | September 29, 2015 | 2015 | Medtronic Inc. | 64 units of Achieve Electrical Cables were shipped with a potential sterility breach. | Under Investigation by firm | Medtronic sales reps began notifying customers on 9/29/2015. Customers were told of the issue and were requested to quarantine and return un-used product. A follow-up Urgent Medical Device Recall letter was hand delivered beginning 10/6/2015. The letter again described the issue, identified affected product and asked that un-used product be returned to Medtronic. A response form was asked to be returned to Medtronic. Customer with questions can contact Medtronic AF Solutions at 888-843-8301 |  | 64 | 64 units | Distributed in DC and the states of AR, CA, CO, DE, GA, IA, KY, ME, MI, MN, NY, OH, PA, TN, TX, and VA. | Terminated | March 09, 2016 | 162 | 64 | Not_Computer
Z-0202-2016 | 72289 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Calibrator Kit 9 Calibrator, 2 mL, Product Code 8568040, Unique Device Identifier No. 10758750005031, IVD -- Ortho Clinical Diagnostics, Inc. --- VITROS Chemistry Products Calibrator Kit 9 is used to calibrate VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System for the quantitative measurement of ACET, CRBM, DGXN, and PHYT. |  |  | September 28, 2015 | 2015 | Ortho-Clinical Diagnostics | The firm is conducting a product correction due to an increase in U90-382 or 6LU condition codes (wash errors) when using VITROS Calibrator Kit 9, Lot 0954. | Under Investigation by firm | Customer letter (Ref. CL2015-191) was sent on 9/28/2015 via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) and /or ORTHO PLUS e-Communications to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected calibrator lot, discard the remaining inventory of the affected lot and recalibrate VITROS DGXN and PHYT Slides using Cal Kit 9 with an alternative lot of calibrator kit 9 upon availability. Distributor letter (Ref. DL2015-191) was sent on 9/28/2015 via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and the required actions. The foreign affiliates were notified via e-mail on 9/28/2015. |  | 3991 | Domestic: 2443 kits; Foreign: 1548 kits | Nationwide, Puerto Rico and foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela | Open, Classified |  |  | 3991 | Not_Computer
Z-0201-2016 | 72357 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Novation Element Press-Fit, Cementless, HA Coated, Collared, Standard Offset, Size 9, Femoral Hip Implant Stems, Intended for press-fit fixation |  |  | September 23, 2015 | 2015 | Exactech, Inc. | Eight (8) pieces did not receive the specified Hydroxyapatite (HA) coating. | Under Investigation by firm | The firm, Exactech, issued an "IMPORTANT NOTIFICATION-PRODUCT RECALL ADVISORY NOTICE" dated September 23, 2015, by phone and by mail to their consignees. The notice described the product, problem and actions taken. The consignees were instructed to 1) Immediately cease distribution or use of these products. 2) Extend this information to your accounts that may have this product in their possession. 3) Identify and quarantine any of the subject Novation Element Pres Fit Femoral stem in your inventory. 4) Complete and Return the attached Recall Inventory Response Form to Exactech,per the instructions provided on the notice itself, within 5 business days, via fax to: 352-337-3915 or email kaya@exac.com. Return products to: Exactech Distribution Center, 2411 NW 66th Court, Gainesville, Florida 32653, Attention: Exactech inventory representative.Any questions contact your Exactech inventory representative at kaya@exac.com or 1-800-392-2832. |  | 8 | 8 | US Distribution to states of: AZ, CO, FL, and MD. | Completed |  |  | 8 | Not_Computer
Z-0198-2016 | 72398 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Antibacterial Hydro Personal Catheter, Male 12 Fr, Product Code 63512, sold in boxes of 30 units |  |  | October 02, 2015 | 2015 | C.R. Bard, Inc. | Incorrect unit labeling. The case label and shelf box are correct, Product Code 63512 and Catheter Size 12 French. The unit label (catheter package) may state Product Code 63516 and Catheter Size 16 French. | Error in labeling | Bard sent a Customer Notification Letter dated September 29 to customers via FedEx with proof of delivery notice. The letter identified the product, the problem, and the action to be taken by the customer. The consignees were instructed to carry out the notification to the user level. Customers with questions were instructed to contact BMD at 1-770-784-6471. |  | 4170 | 4,170 each | Nationwide Distribution including CA, FL, IL, MO, NJ, OR, PA, TX &amp; WA. | Open, Classified |  |  | 4170 | Not_Computer
Z-0199-2016 | 72326 | No | IKK | Isokinetic testing and evaluation system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Robotic hand accessory to InMotion Arm, rehabilitation robot. |  |  | May 05, 2012 | 2012 | Interactive Motion Technologies, Inc. | Unsupervised patient could potentially mis-use the device and remove or rotate a protective cover and expose a potential pinch hazard. | Component design/selection | Interactive Motion Technology ( IMT) sent an Important Patient Safety Information letter dated March 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Summary of Actions to be taken For Hand Robot Cover B, firm recommends continued vigilance for both Adult and Pediatric Patients while in contact with this group of devices. For Hand Robot Covers A or C, at clinician discretion based on expected patient population, firm recommends continued use when appropriate of the IMT supplied Hypertonia Spring Assembly, or if not appropriate, insertion of the IMT Zero Force Spring Assembly. IMT will make this kit available for each Hand Robot upon request, or insert the Assembly when the Hand Robot is shipped to IMT using the special shipping container. If you have any questions, please contact your local IMT Representative or IMT Technical Services at support@interactive-motion.com. |  | 8 | 8 | Worldwide Distribution - US Distribution to the states of : Texas and Tennessee., and to the countries of : China, Brazil, Hong Kong, Austria and South Korea. | Terminated | July 13, 2017 | 1895 | 8 | Not_Computer
Z-0204-2016 | 69712 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillators |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | 1. Device will perform the weekly automated tests hourly, which could cause the therapy capacitors to degrade sooner than intended and 2. While connected to AC or DC power and with no battery installed or the battery installed has a charge level of less than 10%, the Ready for Use (RFU) indicator will not provide the expected low battery indication | Process control | Philips Healthcare sent an " Urgent Field Safety Notice " dated November 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A hardware upgrade will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:A Philips authorized service provider will service the device and ensure it is operating within specification. Prior to receiving service, you can continue to use your MRx provided that the following steps are followed. These steps will prevent further accelerated degradation of the therapy capacitor, and will eliminate the above described behavior for no/low battery conditions.To correct the problem, set the configuration of the MRx to factory defaults in configuration mode.Please contact your local Philips Service Representative for further questions, or call (978) 687-1501. |  | 1553 | 1553 units | Worldwide Distribution - USA (nationwide), Australia, Austria, Bahrain, Belgium, Brunei Darussalam, Canada, China, Finland, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Republic of Korea, Libyan Arab Jamahiriya, Lithuania, Mexico, Namibia, Netherlands, Palestine, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 1553 | Not_Computer
Z-0203-2016 | 72312 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | MaxAn 4.0mm and 4.5mm fixed screws, used in the MaxAn Anterior Cervical Plate System. The MaxAn Anterior Cervical Plate System is a decompression-based technique for cervical spine stabilization. |  |  | September 29, 2015 | 2015 | Biomet Spine, LLC | The packaging of 4.5 mm diameter MaxAn screws identifies/labels the screws as 4.0 mm diameter screws. | Labeling mix-ups | Zimmer Biomet Spine sent an Urgent Medical Device Recall Notice dated September 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the device has already been implanted, written confirmation of this will be required and the necessary parties will be contacted. They will store returned product in quarantine in an isolated location until they can undergo inspection to verify if they are conforming or not and then will be destroyed or retained. Necessary parts will be destroyed so that it is visually obvious they are no longer functional. For further questions call (303) 443-7500 ext.244. |  | 17 | 17 units | US Distribution to the states of : NY, MO, and FL | Open, Classified |  |  | 17 | Not_Computer
Z-0205-2016 | 71518 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Arthroscopic Energy 50Âż Probe with Suction, Catalog number: AES-50S |  |  | June 17, 2015 | 2015 | ConMed Corporation | Probe programming error. The incorrect program can generate more heat at the surgical site, resulting in an elevated probability of tissue damage. | Software design (manufacturing process) | On June 17, 2015 Con Med Corporation distributed Urgent Device Recall Notification Letters and Recall Reply forms to their customers via courier service. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory for any of the affected lot numbers referenced on the recall notice. If any of the recalled product are in stock, customers should discontinue the use or distribution of these items. If product was further distributed, customers should notify those individuals of this recall action. Customers with questions can call Gregory Connell at 727-399-5276 or email AES50S@conmed.com. Customers should also complete and return the Recall Reply Form by mail to the address listed on the recall notice. |  | 276 | 276 devices | Worldwide - US distribution in the states of FL, OH, and the country of Spain. | Terminated | April 11, 2017 | 664 | 276 | Not_Computer
Z-0208-2016 | 72296 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | e.cam Gamma Camera, an emission computed tomography system. It is intended to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588keV. |  |  | June 26, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential failure mode related to the radial brake, where the detector was able to drift, led to a re-design of the brake redundancy circuit. The redesigned Brake Redundancy Circuit will force even a defective brake to engage in the unpowered state by applying a reverse charge to the brake and ensuring the brake has demagnetized. | Component design/selection | The firm, Siemens, sent out an undated "CUSTOMER SAFETY ADVISORY NOTICE" on 04 April 2015 to its consignees informing them of the e.com system safety update. The notice described the product, problem and actions to be taken. The notice stated the following:In order for an injury to occur, the following series of events must be present;1. The detector configuration angle must be 180 degrees.2. The detector with the failed brake circuit must be above the patient.3. The detector with the faulty brake circuit must be under power (in motion).4. System power must cycle within a very short time period, preventing the brake circuit from charging.The customers were instructed to continue to use their e.com system and to ensure the safety advisory is placed in the systems instructions for use and to observe patients while system is in use. If they experience difficulty or failure while performing the daily touch pad test, they are advised to discontinue use and contact their local representative. If unintended movement is detected they are advised to remove the patient and discontinue using the system and contact their local service representative. If you have any questions regarding this important safety notice, please contact your local Service representative at the contact numbers provided below." America: 1-800-888-7436" Europe, Middle East, and Africa: +49 9131 940 4000" Asia and Australia: +86 (21) 3811 2121Local Service representative at the contact numbers provided below."America: 1-800-888-7436"Europe, Middle East, and Africa: +49 9131 940 4000"Asia and Australia: +86 (21) 3811 2121 |  | 32296 | 32296 e.cam Gamma Camera systems | Worldwide Distribution: US (nationwide) and countries of: United Arab Emirates, Albania, Argentina, Austria, Australia, Bosnia and Herzegovina, Bangladesh, Belgium, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, Cameroon, China, Colombia, Costa Rica, Czech Republic, Germany, Denmark, Dominica, Ecuador, Egypt, Spain, Finland, France, United Kingdom, Georgia, Greece, Croatia, Canary Islands, Ireland, Israel, India, Iraq, Iran, Islamic-Republic of, Italy, Jordan, Japan, Korea-Republic of, Kuwait, Lebanon, Sri Lanka, Luxembourg, Latvia, Morocco, Myanmar, Mauritius, Mexico, Malaysia, Netherlands, Norway, Nepal, New Zealand, Oman, Peru, Philippines, Pakistan, Poland, Portugal, Qatar, Serbia, Russian Federation, Saudi Arabia, Sweden, Slovakia, El Salvador, Syrian Arab Republic (Syria), Thailand, Tunisia, Turkey, Taiwan, Republic of China, Ukraine, Uganda, Uzbekistan, and South Africa. | Open, Classified |  |  | 32296 | Not_Computer
Z-0213-2016 | 72290 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | Helioseal F Sealant Refill 1 x 1.25 g, REF 558519AN, and Helioseal F Assortment (kit) 5 x 1.25 g, REF 558518AN. Dental sealant. -- ivoclar vivadent clinical --- Helioseal F Refill is provided in a syringe packed sealed foil bag. Helioseal F is a light-curing, white-shaded fissure sealant featuring fluoride release. Helioseal fissure sealants are suitable for the sealing of pits, fissures and foramina caeca. |  |  | October 07, 2015 | 2015 | Ivoclar A. G. | A defect occurred in the production process of Helioseal F Sealant; the particles in the faulty component seem to be larger than the size defined in the formula. This results in difficulties during extrusion of the material. | Under Investigation by firm | Ivoclar Vivadent sent an Urgent Medical Device Recall Letter dated October 7, 2015, with Response Forms via Certified Mail, Return Receipt Requested. Dealers were also contacted via telephone/e-mail and informed to expect the letter. Dealers are requested to check their stock and return any affected material for credit, plus notify their customers of the recall. For questions, please contact Ivoclar Vivadent Customer Service at 1-800-533-6825. |  | 337 | Domestic: 337 units | Nationwide Distribution including NJ, PA, SC, FL, IN, IA, WI, IL, MO, TX, CA and WA. | Open, Classified |  |  | 337 | Not_Computer
Z-0210-2016 | 72217 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO 400 and AMSCO C Small Steam Sterilizers |  |  | September 03, 2015 | 2015 | Steris Corporation | In AMSCO Small Steam Sterilizers equipped with the CS-iQ feature, the sterilizer software inadvertently prevents the capability to operate the foot pedal, which allows the Operator to open and close the sterilizer door. | Software design | On 9/21/2015 the firm sent Recall Correction Notices to their customers. |  | 32 | 32 units | US nationwide distribution. | Terminated | January 22, 2016 | 141 | 32 | Computer
Z-0151-2016 | 72319 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-G5 Ventilator with software versions between V2.00 and V2.31. |  |  | April 22, 2014 | 2014 | Hamilton Medical, Inc. | Ventilation and alarms of a HAMILTON-G5 ventilator can be suppressed unintentionally after the activation of a suctioning maneuver by the operator; this situation can occur regardless of the selected patient group (neonatal, pediatric and adult). | Device Design | Hamilton Medical issued a letter dated April 22, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that their Hamilton Medical Account Manager would contact them to coordinate the deactivation of the Ventilation Suppression feature. Customers were asked to review the Medical Device Field Safety Corrective Action.Customers with questions were instructed to call 800-426-6331, Ext 215. |  | 1128 | 1128 | Nationwide Distribution | Terminated | August 24, 2017 | 1220 | 1128 | Not_Computer
Z-0207-2016 | 72463 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW PICC Kit featuring Chlorag+ard Technology |  |  | October 16, 2015 | 2015 | Arrow International Inc | Possibility that the catheter peel-awaycomponent hub tabs may prematurely detach when the practitioner begins to peel apart the sheath body from thecatheter. Although the issue does not pose a significant risk to the patient, it may cause a delay of treatmentwhen placing the catheter. | Unknown/Undetermined by firm | Arrow International sent an Urgent Medical Device Recall Notification letter dated October 16, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed to immediately discontinue use and quarantine any effected products. Consignees were instructed to return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recalls@teleflex.com. A customer service representative will contact consignees with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Arrow International. For consignees with no affected product, should also complete the enclosed Recall Acknowledgement Form.For questions customers should call 1-866-246-6990. |  | 12157 | 12,157 | World Wide Distribution - US including AR, AL, TX, MA, CA, IA, MO, NY, AL, VA, RI, IN, TN, FL, KY, NH, SC, MD, WA, OK, IL, ME, OR, OH, PA, MI, KS, LA, TN, WV, NE, NJ, NV, GA, SD, CT, MS, CO, and Internationally to Canada and Belgium. | Open, Classified |  |  | 12157 | Not_Computer
Z-0209-2016 | 72488 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | iOS Galileo Application Version 1.1.1 or lower that programs the Triton Smart Ankle; 1C66 Triton Smart Ankle, an external prosthetic ankle/foot. |  |  | October 23, 2015 | 2015 | Otto Bock Healthcare GmbH | A software issue can make the foot move into dorsiflexion, unnoticed by the user, who could fall. When sitting, the prosthetic foot can go into the Relief Function to move the footplate towards the ground. When operating machines with foot pedals (e.g. a car) the foot could get stuck under the foot pedal and block it or, with the foot on the pedal, it could cause unintended vehicle acceleration. | Software design | OttoBock sent an Urgent Medical Device Correction letter dated October 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The US Agent will notify US customers via letter on October 23, 2015, of the correction on behalf of the Manufacturer. Customers will be requested to contact all users and advise them of the hazards when operating machines or vehicles with foot pedals. They may forward a copy of the insert if needed. Customers will also be requested to contact all users with TF-fittings, where the settings were done via iOS Galileo app., inform the user about this issue, and assure that settings in the device are as expected.The customer will be requested to complete a reply form and return to the US Agent. For further questions call (800) 328-4058. |  | 321 | 273 in US: 48 Foreign | Worldwide Distribution - US (nationwide) and to the countries of : Austria, Belgium, Germany, Israel, Luxembourg and Sweden. | Terminated | March 21, 2017 | 515 | 321 | Computer
Z-0206-2016 | 72161 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | BAUSCH + LOMB Biotrue ONEday lenses (nesofilcon A), Product Order Number Biotrue SVS - 7201, STERILE, Rx Only; The lenses are available in 5 pack, 30 pack and 90 pack sizes; Biotrue ONEday Foil - Part Number 8098001, Biotrue ONEday 5 pack Carton International - Part Number 8101403, and Biotrue ONEday 90 pack Carton International - Part Number 8098203. |  |  | September 03, 2015 | 2015 | Bausch &amp; Lomb Inc Irb | The firm determined that 25 lots had been manufactured with a minor surface defect occurring in the optical surface on each lens. Complaints of Halo or Poor Visual Acuity. | Process control | On 9/3/2015, the distributors received the Medical Device Recall Letter (dated 9/2/2015) via e-mail. Medical Device Recall Letters (dated 9/2/2015) and Customer Acknowledgement Cards were hand delivered to the direct customers via Bausch + Lomb Sales Representatives starting on 9/14/2015 or sent via FedEx if consignee is not a B+L account. The letter informs the accounts of the recall and instructs them to direct any consumers who would like to report a product complaint, arrange for a product exchange, or obtain more information on the product recall to contact the Bausch + Lombs Consumer Affairs Team at 1-800-553-5340 or via e-mail at BLCustomerCare@bausch.com. Customers may also contact Bausch + Lomb Customer Service at 1-800-828-9030 to arrange return of affected product. |  | 183920 | Domestic: 99,260 lenses; Foreign: 84,660 lenses | Distributed US (nationwide), Canada, Japan and Thailand. | Open, Classified |  |  | 183920 | Not_Computer
Z-0219-2016 | 70002 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx monitor/defibrillator |  |  | December 23, 2014 | 2014 | Philips Medical Systems, Inc. | MRx monitor/defibrillator could reboot at an indeterminate time, potentially causing therapy to be interrupted or delayed | Software design | Philips Healthcare issued on 12/23/14 The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the software upgrade. If the MRx reboots during clinical use, continue to treat the patient according to your organizations protocol, reinitiating therapy if necessary. |  | 75693 | 75,693 units | Worldwide Distribution - US Nationwide and the countries of AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GHANA, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LAOS, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACAU, MACEDONIA, MALAYSIA, MALTA, MEXICO, MONGOLIA, MOROCCO, MOZAMBIQUE, MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, UNITED REPUBLIC OF TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 75693 | Computer
Z-0218-2016 | 72236 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digital Lightbox, BrainLAB system, image processing, radiological |  |  | September 21, 2015 | 2015 | Brainlab AG | Potentially incorrectly displayed objects when actively deselecting a fused reference dataset. | Software design | Brainlab sent A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated 8/27/2015 to medical users regarding Content Manager, Patient Browser and the DICOM Viewer. The Product Notification letter decribes the corrective action information and to advise the medical users of the action Brainlab is taking to address the issue. If you require further clarification, the firm advises to contact the local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers)E-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax: Brainlab AG: + 49 89 99 15 68 33Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  | 897 | 897 | Worldwide Distribution - US Nationwide and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Cambodia, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Moldova, Morocco, Myanmar, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 897 | Not_Computer
Z-0217-2016 | 72322 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Envoy 500 system includes the following product reference numbers and names: |  |  | September 22, 2015 | 2015 | ELITech Clinical Systems SAS | ELITech clinical Systems reagents based on Trinder reaction used on the Envoy 500 instrument system are recalled due to potential drug interference. No complaints were received. | Under Investigation by firm | The firm sent recall letters or Technical Bulletins to US consignees on 9/22/15 via registered mail. Technical Bulletin #1022 (for EON 100 users)Technical Bulletin #2008 (for EON 300 users)Technical Bulletin #5064 (for ENVOY users)Technical Bulletin #3009 (for Selectra users)ELITechGroup is not requesting return of the affected product. The method sheets and package inserts will be updated with the new interference claims. When available, the revised labeling will be provided to the customer for customer's use until the new IFU is included in the product packaging. Consignees are requested to return the Confirmation of Receipt form back to the firm by FAX to 401-642-9001 or send email to customercare@elitechgroup.com or send via US Post: Vital Diagnostics, 27 Wellington Rd., Lincoln, Rl.If you have any questions, please contact Vital Diagnostics Technical Support at1-800-345-2822 or 855-ELITECH |  | 13455 | 13455 systems | US Nationwide Distribution | Terminated | March 04, 2016 | 164 | 15813 | Not_Computer
Z-0227-2016 | 72390 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Display, Model No. MDSC-2224 (24 inch), Model No. MDSC-2226 (26 inch) and Model No. MDSC-2242 (42 inch) |  |  | May 28, 2015 | 2015 | Barco Inc. | User settings were no longer on the monitor's memory after these systems were powered off and back on. | Process design | Barco sent an "Urgent Medical Device Correction Notice" dated May 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This notice contained "Interim Corrective Action" and information about a "Final Corrective Action". The notice included a reply notice to be sent back to Barco and also requested a sub-recall. For further questions, please call (678) 475-8156. |  | 462 | 462 devices in US | US Distribution to the states of : AZ, CA, FL, GA, IL, MA, NE, NJ, OH, OR and TX., and Internationally to Canada. | Open, Classified |  |  | 462 | Not_Computer
Z-0222-2016 | 72495 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker UniVise Spinous Process Fixation Plate System Inserter, Stryker Spine, Non-Sterile. Reference Number 48590100. Manual Surgical Instrument. |  |  | September 08, 2015 | 2015 | Stryker Spine | The Inserter Inner Shaft would not fit through the Inserter. The inner diameter of the cannulated shaft of the inserter was undersized. | Process control | Stryker Spine issued an "Urgent Product Removal Notifications /Customer Response Forms dated September 11, 2015. to all affected customers via Fed Ex priority overnight. The notification instructed customers on how to identify affected product; the related issue with the affected product; potential hazards; risk mitigations; and actions needed to be taken. Customers were asked to examine their inventory and related hospital location inventories; and reconcile any UniVise Inserters with Catalog numbers and Lot numbers (referenced in the notification). Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance (at 855-632-9049) or email it to Spine-RegulatoryActions@Stryker.com within 5 days of receipt of the notification.For further questions please call (201) 760-8346. |  | 41 | 41 units | Nationwide Distribution. | Open, Classified |  |  | 41 | Not_Computer
Z-0229-2016 | 72053 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Regard Custom Surgical Pack, Item Number 880239, Sterile IV Pack. Custom procedure trays contain surgical instruments and accessories intended for use during a surgical procedure. |  |  | August 25, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | The kits contain pressure monitoring kits or sets which are being recalled as a result of flow restriction issues in the trifurcated IV set. | Component design/selection | ROi sent a Notice of Correction/Removal letter dated August 25, 2015, to all affected customers. The letter notified customers that ROi recently received notice of a voluntary recall from Edwards Lifesciences regarding their Transducer Kit. ROi also sent an email to their customer on September 25, 2015, informing them of the recall. Attached to the email was the component manufacturer's recall notice and a recall notice from the recalling firm. The recalling firm's recall notice instructed their customer to respond to the notice indicating whether they had any of the recalled product remaining and to destroy any remaining product in their inventory.For questions regarding this recall call 417-820-2793. |  | 255 | 255 packs | Product was distributed to one customer in Oklahoma. | Terminated | April 06, 2017 | 590 | 255 | Not_Computer
Z-0221-2016 | 72384 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Extension Set, Model No. 20029E. Used to administer fluid and medications through a needle or catheter inserted into the patient's artery or vein. |  |  | October 09, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Extension Set due to disconnection and leakage issues. | Process control | An urgent medical device recall notification was sent to customers on 10/9/15 who purchased the SmartSite Extension Set. The letter informs the customers that CareFusion has identified potential risks with model code 20029E, lot numbers 1503541, 15015769 and 15045950 due to disconnection and leakage issues. The letter also provides the customers with the potential risk involved and the actions to be taken. Customers are instructed to complete and return to CareFusion the recall response card. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266 or email customerfeedback@carefusion.com. Customers with technical questions regarding technical questions are instructed to contact Technical Support at (888) 812-3229. |  | 21000 | 21,000 units | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 21000 | Not_Computer
Z-0223-2016 | 72459 | No | MIH |  |  | Cardiovascular | PMA | FLAIR Endovascular Stent Graft, Product Code: FAS08070, FAF08070, Implant Size: 8mmx70mm |  |  | October 19, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) is recalling the Bard Flair Endovascular Stent Graft because of the harm associated with failure to deploy or deployment related issues that could occur during use. | Device Design | Bard sent an Urgent Medical Device Recall Notification letter dated October 19, 2015, to customers who received the Bard Flair Endovascular Stent Grant. The letter informs customers that the harm associated with failure to deploy or deployment related issues that could occur during use may pose an incremental risk to the already inherit risk associated with stent graft procedures. The letter also provides customers with the actions to be taken. Customers were instructed to call Bard Peripheral Vascular's (BPV) Recall Coordinator at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST). Customers were instructed to fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. |  | 292 | 292 units | Nationwide in US | Terminated | April 12, 2016 | 176 | 292 | Not_Computer
Z-0224-2016 | 72224 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Robotic Arm Interactive Orthopedic (RIO) system component PN#205143 (locking pin) and PN# 112220 (connector). Indicated for use in surgical knee resurfacing and total hip Arthroplasty joint replacement procedures. |  |  | July 16, 2015 | 2015 | Mako Surgical Corporation | The locking pin on the RIO Base Array is damaged during assembly onto the RIO Base Array Connector. | Device Design | As of July 16, 2015, Mako communicated by mail the nature of the recall to their Product Specialists and the OR Administrators describing the specifics regarding the issue for correction. |  | 1322 | 1,322 units | Worldwide Distribution-US (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY and HI, and the countries of Italy, Japan, Hong Kong, Scotland, Germany, Turkey, Greece, Australia, Singapore, Taiwan, and South Korea. | Open, Classified |  |  | 1322 | Not_Computer
Z-0220-2016 | 72097 | [6 empty] | EnVision FLEX/HRP visualization reagent found in the following kits: |  |  | September 04, 2015 | 2015 | Dako North America Inc. | There is a defect in one lot of a buffer, non-specific background staining that has, in some instances, interfered with interpretation of the staining results. | Other | Dako sent an customer notification letter on September 4, 2015, to all affected customers via mail. The letter identified the product the problem and the action needed to be taken by the customer.The customers/distributors are instructed in the Recall Notification Letters to 1) remove any unused kits bearing the affected lot numbers listed in their letter from diagnostic use 2) destroy any unused kits bearing the affect lot numbers listed in their letter, and 3) determine any possible impact of this change in performance on patient results from the affected lots.The letter states actions should be taken within 10 days and that customers/distributors should confirm they have received the information by returning the attached, fully completed Device Recall Form and return it to Dako QA Vigilance by email at dako.dkvigilance@dako.com.The firm lists on their letter for customers to contact their Dako representative in there are any questions. For further questions, call (805) 566-6655. |  | 300 | 300 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the states of : TX, SD, CA, FL, MA, MD, WV, NC, OH, DE, MI, NJ, CO, HI, AR, NY, LA, GA, NV, RI, IL, OK, CT, MO, VA, and DC., and to the countries of : Canada, Brazil and Costa Rica. | Terminated | April 25, 2016 | 234 | 300 | Not_Computer
Z-0228-2016 | 72096 | No | GAX | Nonpneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MATResponder Tourniquet; Catalog number MATR: UPC 78353 00008; |  |  | August 31, 2015 | 2015 | Pyng Medical Corporation | Once tested, it is required that the tourniquet be returned to its Ready position then placed in its packaging. Eachtourniquet was tested for functionality, but in some cases, the tourniquet was not returned to its Ready position prior to packaging. | Process control | PYNG Medical sent Recall Notification letters dated August 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:Please immediately cease use of the MATResponder and MATCombat tourniquet lots specified above. Inform all of your teams to cease use of or distribution of these lots.Please immediately place all MATResponder Tourniquets and MATCombat Tourniquets that you have in inventory into a quarantined location. For further questions please call (604) 303-7964.. |  | 6728 | 6728 | Worldwide Distribution - US Distribution and to the countries of : Colombia, France, Canada, Italy, Germany, Finland, United Kingdom, Australia, The Netherlands, Japan, Hungary, Indonesia, Israel, Mexico, Poland and Spain. | Open, Classified |  |  | 6728 | Not_Computer
Z-0231-2016 | 72220 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number/Product Code 6801375, Unique Device Identifier No. 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number/Product Code 6801890, Unique Device Identifier No. 10758750001644. Intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest. |  |  | September 10, 2015 | 2015 | Ortho-Clinical Diagnostics | Software Anomaly during ADD Installation on VITROS 5,1 FS Chemistry Systems using Software Versions 2.2.1 through 2.8. The user modified parameters would not be saved as expected, but remained changed to the default values obtained from the ADD disk. | Software design | The firm, Ortho Clinical Diagnostics, sent a "FOLLOW UP to URGENT PRODUCT CORRECTION NOTIFICATION" letter dated 9/10/2015 and response forms to customers via FedEx overnight courier, e-mail (ORTHO PLUS e-Communications) and/or via US Postal Service overnight letter on 9/10/2015. Foreign affiliates were informed by email on 9/10/2015 of the issue and instructed to notify their consignees of the issue and required actions.The letter described the product, problem, and actions to be taken. The customers were instructed to follow the required actions and complete and return the Confirmation of Receipt-Response Required form, no laterthan 9/18/15, by fax at 1.888.557.3759 or 1.585.453.4110, or scan PDF and email to: ConfirmationAdmin@its.jnj.com. Customers with any additional questions, can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 210 | Catalog # 6801375: Domestic 919; Foreign 1251; Catalog 6801890: Domestic 63, Foreign 147 | Worldwide Distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 210 | Computer
Z-0225-2016 | 72472 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS-IW with Universal Viewer is a device that displays medical images (including mammograms) and data from various imaging sources. Images and data can be viewed, communicated, processed and displayed within the system or across computer networks at distributed locations. Typical users of this system are trained professionals, including but not limited to radiologists, physicians, nurses, medical technicians, and assistants. |  |  | September 28, 2015 | 2015 | GE Healthcare | Inaccurate distance measurements with magnified projection X-ray images. | Software design | An URGENT MEDICAL DEVICE CORRECTION letter was sent to all medical users on 9/28/2015 regarding Centricity PACS-IW with Universal Viewer and Centricity Universal Viewer- Inaccurate distance measurements with magnified projection X-ray images.GE Healthcare explains that for all projection X-rays, GE Healthcare recommends that the user perform the necessary measurements on the standard contact images, and not rely on measurements made on the magnified views unless the measurements have been independently verified for each X-Ray machine installed or manually calibrated using a radio-opaque fiducial of a known size. Attention should be given to prior cases acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue. GE Healthcare will correct all affected systems by providing a software update at no cost to the customer. GE Healthcare advises customers if they have any questions or concerns regarding the notification, to please contact GE Healthcare Service at1-800-437-1171 or their local Service Representative. |  | 5 | Centricity PACS-IW with Universal Viewer Versions 5.0 SP2 and higher 1052 units | Worldwide Distribution-US (nationwide) including the states of AZ, CA, CO, DE, FL, GA, IA, IL, IN, LA, MA, MI, MN, MO, NJ, NY, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV, and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Ecuador, France, Germany, Ghana, Hong Kong, India, Israel, Italy, Japan, Kenya, Korea, Malta, Malaysia, Namibia, Netherlands, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, Taiwan, United Arab Emirates, United Kingdom and Venezuela. | Open, Classified |  |  | 11 | Computer
Z-0241-2016 | 72098 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal Sterile Custom Sterile Custom Surgical Kits containing 3M DuraPrep (TM) Surgical Solution, Catalog No. 8630, Lot Nos: 2016-09EE or 2016-09EF, or 2016-09EG, or Catalog No. 863, Lot No. 2016-08DC: |  |  | August 31, 2015 | 2015 | DeRoyal Industries Inc | Custom Sterile Surgical kits contain 3M DuraPrep (TM) Surgical Solution, Catalog Number 8630, Lot Number 2016-09EE which was recalled due to unintended exposure to the sterilant during package sterilization. | Other | DeRoyal initiated their recall on 08/31/2015 to the end-user level by sending overnight letters to its direct accounts and hospitals. Distributors were asked to put the product in quarantine, place the provided labels on the products and complete and return the spreadsheet of how many they labeled. They were asked to return the spreadsheet even if they did not have inventory. Hospitals were provided with labels to affix to the kits with direction to discard the recalled component and notice of destruction forms to be completed and returned. |  |  |  | SC, FL, OH, NJ, OK, MN, AR | Terminated | April 12, 2016 | 225 |  | Not_Computer
Z-0237-2016 | 72434 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | TEKIA Single Piece Hydrophilic Acrylic Intraocular Lenses, Model 811 |  |  | September 29, 2015 | 2015 | Tekia, Inc. | Tekia is recalling the Single Piece Hydrophillic Acrylic Intraocular Lenses (IOLs) because they may be labeled with the incorrect diopter power. | Error in labeling | Tekia sent an Field Safety Notice dated September 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were advised that Tekia is recalling a limited a number of Single Piece Hydrophilic Acrylic IOLs. If the customer has any questions, the customer is instructed to contact a TEKIA Customer Service or QA Representative at (949) 699-1300. Customers with product complaints or adverse events regarding the use of Single Piece Hydrophilic Acrylic IOLs are instructed to inform TEKIA by phone, email, or fax. |  | 50 | 50 units | Internationally to Vietnam. | Terminated | December 03, 2015 | 65 | 50 | Not_Computer
Z-0239-2016 | 72408 | No | DHX | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access CEA Assay, Part No. 33200. For the quantitative determination of Carcinoembryonic Antigen (CEA) levels in human serum. |  |  | October 13, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access CEA reagent packs because they were filled incorrectly. | Under Investigation by firm | An Urgent Medical Device Recall letter dated 10/13/15 was sent to all customers who purchased the Access CEA Reagent Kit to inform them that Beckman Coulter has determined that the Access CEA reagent packs (P/N 33200) were filled incorrectly. The packs contain insufficient quantity of reagents in one of the pack wells. The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Technical Support at http://www.beckmancoulter.com, (800) 854-3633 in the US and Canada. Customers outside of the US and Canada are instructed to contact their local Beckman Coulter representative. |  | 11406 | 11,406 units total (1,660 units) | Worldwide Distribution-US (nationwide) and the countries of Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Croatia, Czech Republic, Egypt, France, Germany, Ghana, Gibraltar, Greece, Guam, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Republic of Korea, Lebanon, Mexico, Morocco, Myanmar, Netherlands, Pakistan, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Vietnam. | Terminated | April 13, 2016 | 183 | 11406 | Not_Computer
Z-0232-2016 | 72225 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific SQ-RX Model 1010 subcutaneous pulse generator, (S-ICD). |  |  | September 17, 2015 | 2015 | Boston Scientific Corporation | Devices shipped to US in dual-channel rather than US approved single-channel RF Telemetry communication mode. While this does not impact therapy delivery of device; this can impact likelihood of successfully establishing telemetry during first scan of device when using a single-channel programmer. Telemetry can be established with additional scans; once established no further difficulty is expecte | Process design | One consignee was contacted via a Boston Scientific "Medical Device Retrieval" letter dated September 17, 2015. The letter was addressed to Hospital Administrator. The letter described the problem and the product involved in the recall. Informed consignee that their local Boston Scientific sales representative will be retrieving the product form their inventory. For questions they can contact their local sales representative or Boston Scientific Technical service at 1-800-227-3422.A second Boston Scientific "Important Medical Device Information" letter dated October 2015 was hand delivered to physicians starting on 10/15/2015. The letter described the problem and provided information on Clinical Considerations, Affected Population, Recommendations and Further Information. For questions they can contact their Boston scientific representative or the Americas Technical service at 1-800-227-3422. |  | 3 | 3 | Distributed in the US to MA and PR. | Terminated | March 08, 2016 | 173 | 3 | Not_Computer
Z-0128-2016 | 71958 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Villa Sistemi Rotograph EVO 3D X-ray |  |  | August 05, 2015 | 2015 | Villa Radiology Systems LLC | It has been discovered that due to lack of information accompanying documents the product fails to comply with 21CfR 1020.33(c),1020.33(d),1020.33(g),and 1020.330). | Other | Villa Radiology Systems Planned Action:1. The current users manuals have been updated and addendums to the user manuals have been developed for currently installed Rotograph Evo3D and Imax Touch 3D. The revisions and addendums contain all information required under 21 CFR 1020.33(c), 1020.33(d), 1020.33(g), and1020.33(j).2. New Rotograph Evo 3D and Imax Touch 3D systems will ship with the updated user manual.3. Addendums will be provided to all end usersCDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. |  | 8 | Rotograph EVO 3D - 8 installed | US Distribution | Terminated | March 22, 2017 | 595 | 8 | Not_Computer
Z-0238-2016 | 72395 | research use only/microbiology | OTT |  |  | Microbiology | Enforcement Discretion | Focus Diagnostics Anaplasma Phagocytophilum IFA IgM Test Kit, Model No. IF1450M in vitro diagnostic. |  |  | October 02, 2015 | 2015 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Anaplasma phagocytophilum IFA IgM kit because of the potential increase in invalid results. | Device Design | A recall letter dated 10/2/15 was sent to distributors who distributed the Anaplasma phagocytophilum IFA IgM (IF1450M) kit lots 26431 and 27011. The letter informs the customers that Focus Diagnostics received customer complaints of IF1450M with low reactivity, which may result in higher invalid rate. The letter also provides the customers with Focus Diagnostics' recommendations and actions to be taken by the distributors. Customers with questions or require additional information are instructed to contact Focus Diagnostics' Technical Services department at (562) 240-6550 from 7am to 5pm (PST) or send an email to DxTS@focusdx.com. |  | 38 | 38 kits | Non US (worldwide): Czech Republic, Denmark, Germany, Italy, Spain and Colombia | Terminated | January 29, 2016 | 119 | 38 | Not_Computer
Z-0234-2016 | 72386 | No | OUB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | miraDry System consisting of the MD4000-MC Console, the MD4000-HP miraDry Handpiece, and the MD4000-BT miraDry bioTip and accessories. |  |  | October 09, 2015 | 2015 | Miramar Labs, Inc. | Update to labeling for the miraDry Console to remind users to only use approved lubricants in a thin layer to prevent potential burns and blisters. | Labeling design | Miramar Labs sent Medical Device Correction letters to all users on October 9, 2015 to inform them of the labeling clarification by UAPS First Class Certified mail or courier such as Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the firm added a new warning to their labeling. The firm included an updated copy of the MD4000-MC User Manual with the new warning included and asked that customers destroy their older copy. Customers with questions were instructed to call Customer SService at 1-855-647-2379 or email info@miraDry.com |  | 277 | 277 units | Nationwide Distribution | Terminated | December 23, 2015 | 75 | 277 | Not_Computer
Z-0236-2016 | 72407 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Small Notch Titanium Reconstructive Plate set (SNTRP). Ti PLATE/ROD/PLATE 231mm 6HOLES/5.0mm ROD/17HOLES. Orthopedic implant. |  |  | October 08, 2015 | 2015 | Synthes (USA) Products LLC | The Synthes SNTRP set was released for sale as a Trauma product in 1996 with specific trauma indications and contraindicated for certain spine applications. In 1999, a module was released containing a small offering of the SNTRP plates. Synthes determined that the product in the SNTRP module is being used in spine surgeries, which are a contraindicated use. | No Marketing Application | Synthes sent an Urgent Notice Medical Device Letter dated October 8, 2015. to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number, and signature in the spaces provided."Return the Verification Form (page 3 of this letter)If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number, and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the documents to Synthes by:"Fax: (866) 240-5364 or"Scan/email: Synthes4642@stericycle.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 20494 | 20,494 | US Distribution to the states of :MO, NJ, MN, MS, AZ, CA, LA, OK, VA, OH, AR, FL, PA, VT, IN, AL, WA and NY. | Open, Classified |  |  | 20494 | Not_Computer
Z-0245-2016 | 72037 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | RÂżsch Flexi-Slip Endotracheal Tube Stylet with Soft Distal Tip, Sterile, Rx only, Teleflex Medical. |  |  | August 24, 2015 | 2015 | Teleflex Medical | Due to reports of the plastic coating of the stylet splitting and/or breaking off of the stylet. | Unknown/Undetermined by firm | Consignees were notified by letter on 08/24/2015. They were instructed to immediately discontinue use and quarantine any products with the listed catalog numbers. They were to return the product, complete the enclosed Recall Acknowledgement Form and fax or e-mail to recalls@teleflex.com, even if they do not have any of the affected stock. Distributors were instructed to communicate the recall to any of their customers who have received the affected product and to have the customers return any affected product to them with the completed Recall Acknowledgement Form. |  | 233641 | 233,641 units | Worldwide Distribution -- U.S. Nationwide; International Distribution: Argentina, Austria, Belgium, Belgrade, Cameroon, Canada, Chile, Costa Rica, Croatia, Czech Republic, Ecuador, Denmark, Estonia, Finland, France, French Polynesia, Great Britain, Greece, Guyana, Hong Kong, Ireland, Italy, Jordan, Kuwait, Latvia, Lebanon, Martinique, Netherlands, Panama, Peru, Poland, Portugal, Reunion, Romania, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Venezuela, and the United Arab Emirates. | Open, Classified |  |  | 233641 | Not_Computer
Z-0242-2016 | 72331 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Triton Infusion Pump (model 300000) and Triton fp Infusion Pump (model 400000) (proprietary name: AFF (Mark I) Volumetric Infusion Pump). Used to pump fluids into a patient in a controlled manner. |  |  | October 01, 2015 | 2015 | WalkMed Infusion, LLC | The device might fail to detect air in line while infusing the medication Venofer. This may occur when a drop of Venofer gets stuck in the tubing adjacent to the bubble detector and the Total Volume Limit is programmed higher than the actual bag volume. | Device Design | A Field Notification (letter) was provided to users on 10/01/15 to remind them of the following:use air elimination filters, as appropriate, not to over program into the device the volume to be infused, remove air from the IV bag prior to starting the infusion, issues associated with the use of Venofer or similar drugs. No products are being returned. |  | 4695 | 4695 units | Distributed in the states of AL, CA, CO, DE, FL, GA, IL, MA, MD, MI, MN, MS, NC, NJ, NM, NV, NY, OH, PA, TN, TX, and WA. | Open, Classified |  |  | 4695 | Not_Computer
Z-0244-2016 | 72518 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Cook Medical CXI Support Catheter |  |  | October 21, 2015 | 2015 | Cook Inc. | Cook Medical has received six product complaints associated with reports of the 2.6FR curved catheter being mixed with the 4.0FR straight catheter and vice versa. Investigation revealed the products were mixedduring processing. | Labeling mix-ups | Cook Medical sent an Urgent: Medical Device Recall letter dated October 20, 2015, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 38 | 38 | Worldwide Distribution - US Distribution to the states of : CA, FL, ID, IL, MA, NY, OK, TX, UT and VA., and to the countries of Canada, Austria, Germany, Spain, France, Hungary, Sweden and UK. | Open, Classified |  |  | 38 | Not_Computer
Z-0235-2016 | 72456 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SSS Reprocessed Zimmer, Synthes, Stryker and Linvatec Drill Bits and Blades |  |  | October 14, 2015 | 2015 | Stryker Sustainability Solutions | It was reported that packaging was not properly sealed. | Employee error | Stryker sent an Urgent Medical Device Recall letter dated October 14, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm's letter states that the use of the affected lots of the recalled product should be discontinued. Other bit and/or blade lots reprocessed by Stryker Sustainability Solutions should be considered as alternatives. A Recall Effectiveness Check Form including the quantity of affected devices shipped is enclosed.The letter states to complete the Recall Effectiveness Check Form and indicate if any devices from the affected lots remain in inventory. The form needs to be completed even if no affected product is found.Return the completed and signed form to the local Stryker Sustainability Sales Representative or email to to SSSPFA@stryker.comIf the form indicates that affected devices remain in inventory, a prepaid shippinglabel will be issued for the return of the product. Customers will receive a credit forall affected devices returned. For further questions, please call (888) 888-3433. |  | 1202 | 1202 | US ( nationwide) Distribution to the states of : MN, OH, CA, IN, WI, FL, GA, MO, IA, MT, LA, KS, TX, TN, WY, MD, NJ, WA, PA, SC, NY, MI, AT, HI, MN, AR, AZ, IL, IA, KY, MS, UT, NC, VA, CO, ME and ID., Internationally to Canada. | Open, Classified |  |  | 1202 | Not_Computer
Z-0243-2016 | 72502 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Amphirion Deep OTW PTA Balloon Catheter: |  |  | October 23, 2015 | 2015 | Medtronic Vascular, Inc. | Medtronic has discovered that select lots of the Amphirion Deep PTA Balloon Dilation Catheter have a potential breach in the pouch seal enclosing the sterilized product. | Equipment maintenance | Beginning October 23, 2015, Medtronic will send the customer letter andconfirmation form to all identified US customers via 2-day UPS mail delivery. Internationally, communications will be sent per locally approved methods. |  | 5895 | 1652 units in US; 4243 units OUS. | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Australia Austria Belgium Brazil China Colombia Costa Rica Czech Republic Denmark Dominican Republic Egypt France Georgia Germany Hong Kong Hungary India Iran Ireland Israel Italy Japan Jordan Kazakhstan Korea Kuwait Latvia Lebanon Lithuania Luxembourg Mexico Netherlands Norway Oman Pakistan Panama Peru Philippines Poland Portugal Russian Federation Saudi Arabia Singapore Slovakia Slovenia South Africa Spain Sri Lanka Sweden Switzerland Syria Taiwan Thailand Trinidad and Tobago Turkey United Arab Emirates United Kingdom and Uruguay.. | Terminated | July 31, 2017 | 647 | 5895 | Not_Computer
Z-0233-2016 | 72462 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Myelo-Nate CSF Sampling Kit, Neonatal/Pediatric cerebralspinal fluid sampling sterile tray, including a lumbar puncture needle, four collection vials, an iodine swabstick package, a fenestrated drape, a polylined drape, three gauze pads, and one spot bandage. |  |  | October 16, 2015 | 2015 | Utah Medical Products, Inc | Utah Medical is conducting a recall of Myelo-Nate CSF Sampling Kit due to an internal post-sterilization inspection that found a sterile finished tray where the seal between the Tyvek lid and the PETG tray on oneside of the tray had been compromised. | Package design/selection | Consignees notified on 10/15/15 via letter and email, instructed to return product to distributor. |  | 21403 | 21,403 trays | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of CA, CO, MD, WI, FL, GA, IL, IA, ID, IN, KS, KY, LA, MI, MN, MO, MS, NC, NE, NJ, NV, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI, and the countries of Canada, United Kingdom, Saudi Arabia, Jordan, United Arab Emirates, Kuwait, and Costa Rica. | Open, Classified |  |  | 21403 | Not_Computer
Z-0250-2016 | 72500 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | NovoCut Suture Manager; Ceterix catalog number CTX-001. |  |  | October 13, 2015 | 2015 | Ceterix Orthopedics, Inc. | The NovoCut Suture Manager device may cut the suture during knot tensioning and this may cause the suture knot to unravel. | Device Design | Ceterix Orthopaedics sent an Urgent Field Safety notification letter dated October 13, 2015, to all affected sites. The product was identified, reason for recall explained and letter contains list of product lot numbers affected. Customers were instructed to identify and quarantine the affected product, contact Ceterix Orthopaedics to arrange a no cost return, and complete and return the attached acknowledgement form. Customers with questions were instructed to call 1-888-305-0615 or email: docs@ceterix.com. For questions regarding this recall call 650-396-3478. |  | 1364 | 1364 units | Nationwide Distribution | Terminated | February 12, 2016 | 122 | 1364 | Not_Computer
Z-0251-2016 | 72376 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu C-arm Table, Catalog No. IVS-110, Medical Imaging Table. |  |  | September 30, 2015 | 2015 | Shimadzu Medical Systems | The C-arm may interfere with the table top under certain circumstances. | Software design | An Urgent Voluntary Medical Device Recall notice was sent on 9/30/2015 to all the customers who purchased the Shimadzu C-arm Table (Catalog No. IVS-110). The notification informs the customers of the description of the issue, the actions to be taken by Shimadzu to correct the issue, and the actions to be taken by the customers. Customers with questions or concerns about the device correction, are instructed to contact their local Shimadzu representative or contact ShimadzuNational Technical Support at (800) 228-1429. |  | 26 | 26 units total (4 units in US) | Worldwide Distribution. US (Texas, Wisconsin, Ohio, and Connecticut), Iran, India, Saudi Arabia, Serbia, Thailand, Germany, Croatia, Turkey, Lebanon, Korea, and China. | Open, Classified |  |  | 26 | Not_Computer
Z-0249-2016 | 72381 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, MIS MODULAR DISTAL CAPTURE TRIATHLON MIS INSTRUMENTS, REF 6541-5-723, Lot AdminECR19205, Non-Sterile, The Triathlon MIS Distal Capture is used when a surgeon elects to use a capture for the distal femoral resection in a Triathlon primary MIS TKA. Once the Triathlon MIS distal Resection Guide (cat. no. 6541-5-721/722, for left or right preparation, respectively) has been secured to the femoral, the MIS distal capture can optionally be attached to the distal resection guide to complete the cut, per Trithion MIS surgical protocol TRI ATH-SP-5. Surgical instrument for knee prosthesis implantation. |  |  | September 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker is recalling various lots of the Trithion MIS Modular Distal Capture because complaints were received that specific lots of this product may disassociate during use. | Nonconforming Material/Component | Stryker notified their Branches/Agencies of this recall by email on September 29, 2015 and notifiction letters and product accountability forms were sent via UPS (with return receipt) on September 29, 2015 to branches and on September 30, 2015 to hospitals. Stericycle wil be handling the returns. Customers with questions were instructed to call 201-831-5826.For questions regarding this recall call 201-831-5272. |  | 3444 | 3,444 units | Nationwide Distribution including AZ, DE, ID, IL, KS, LA, MI, NH, NM, NJ, NY, OH, PA,TX, VA, VT, WA and WI. | Open, Classified |  |  | 3444 | Not_Computer
Z-0247-2016 | 71998 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Getinge Group, Adjustable Positioning Elements (1007.28A0) Manufacturer: Maquet GmbH, Germany. Patient positioners. |  |  | August 17, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Two serious thromboembolic events were reported involving the Adjustable Positioning Elements (Part Number 1007.28A0) in conjunction with Universal Frame (Part Number 1007.24F0) for lumbar spine surgery. The hip elements may have a negative effect on patients- venous backflow under certain conditions and may have been a contributing factor to thromboembolic events. | Under Investigation by firm | MAQUET sent an "Urgent Product Recall Medical Device Removal" notice and "Medical Device Removal Response Form" dated August 17, 2015 via UPS to affected customers. The notification identified the issue and the affected product. The notification instructed customers to read the notification and disseminate the information provided to all the users of the device; examine their inventory to determine if they have the Adjustable Positioning Elements; if so, remove the affected products, quarantine them and keep them in a secure location for shipment back to Maquet; and fill out and sign the accompany response form and send via email or fax to Maquet. If the customers have any questions regarding this field action, they were instructed to please contact their local MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383 (Press option 2, then followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 5:00 p.m. EDT. |  |  | United States (3 sets); Rest of World (47 sets) | Worldwide Distribution - US including AL, CA, MO and Internationally to Australia, Brazil, Canada, China, Croatia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Russia, Singapore, South Africa and United Arab Emirates. | Terminated | October 19, 2016 | 429 |  | Not_Computer
Z-0252-2016 | 72160 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Smooth Muscle Action (SMA) Immunohistochemistry reagent. Catalog numbers: CME305A-0.1ml; CME305B - 0.5 ml, PME305AA - 6.0 ml prediluted and ready to use. They are the same product with the same intended use. |  |  | September 08, 2015 | 2015 | Biocare Medical, LLC | Certain catalog numbers and lots of Smooth Muscle Actin may stain some cell types that are typically expected to be negative for SMA. | Device Design | Biocare sent a Medical Device Recall Notification letter dated September 8, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue and dispose of all affected product. Customers were asked to complete an return the enclosed Return Response Form. Customer's accounts would be credited for all affected product.For questions regarding this recall call 925-603-8033. |  | 151 | 151 units | Worldwide Distribution - US including CO, FL, OH, SC, AZ, AL, FL, CT, NJ, CA, DC, KY, NC, TX, ID, MA, MD, OR, LA, PA, KS, NY, UT, GA and Internationally to Spain, Sweden, Switzerland, Hungary, Italy, Germany, Turkey, Canada, India, Bolivia, Mexico, Indonesia and Australia. | Terminated | January 11, 2016 | 125 | 151 | Not_Computer
Z-0253-2016 | 72204 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MILEX PESSARY KIT; Model Numbers: MXPES00; MXKPES00; MXKPREG03; MXKPGSK07 |  |  | September 14, 2015 | 2015 | CooperSurgical, Inc. | The products have been identified to have incorrect size prints on the MILEX PESSARY KIT packaging. While the Pessary pouch is marked with correct size, an incorrect label onthe packaging does not meet the required release specification. | Labeling mix-ups | Cooper Surgical sent an Urgent Medical Device Recall letter dated September 14, 2015 via FedEx with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. Customers were asked to discontinue use of the affected products and complete the Acknowledgement and Receipt form for a free replacement. For questions contact Product Surveillance at 203.601.5200 ext. 3300. |  | 144 | 144 units | Worldwide Distribution - US Nationwide in the states of AR, CA, CO, GA, IL, KY, LA, MA, MD, MI, MT, NC, NY, OH, OK, PA, RI, SD, VA, WA, WI,WV and in the countries of FRANCE and VIETNAM. | Terminated | January 08, 2016 | 116 | 144 | Not_Computer
Z-0255-2016 | 72301 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, SIGNA 3.0T. |  |  | August 31, 2015 | 2015 | GE Healthcare | When performing head or neck scans, the currently displayed SAR values could be lower than the actual SAR in the head as predicted by SAR modeling. The predicted HEAD SAR value from the modeling suggests the actual SAR in the head could exceed IEC60601-2-33 limit of 3.2 W/kg for some scans. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated August 31, 2015. The letter was addressed to Hospital Administrators/Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 19 | 19 | Worldwide Distribution - US: Nationwide including DC, PR and GU and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BERMUDA, BOSNIA HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GREECE, HONG Kong, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LUXEMBOURG, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VENEZUELA. | Open, Classified |  |  | 1509 | Not_Computer
Z-0262-2016 | 69714 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | MRx Defib can be susceptible to one or both issues: 1. The C02 Inlet Port associated with end-tidal carbon dioxide (EtCO2) monitoring on MRx Monitor/Defibrillators can be pushed into the MRx housing, making it inaccessible. 2. The handle can separate from the MRx housing due to breakage of mounts on the rear case. | Device Design | Philips Heatlhcare issued on November 19, 2014, an Urgent Medical Device Correction notification/Field Safety Notice. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade provided you follow the steps given.For questions regarding this recall call 978-687-1501Customers with questions were instructed to contact their local Philips representative. |  | 27918 | 27,918 units | Worldwide Distribution - US (nationwide) and Internationally to ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CROATIA (local name: Hrvatska), CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NETHERLAND ANTILLES, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, AND YEMEN. | Open, Classified |  |  | 27918 | Computer
Z-0263-2016 | 72171 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Samsung GM60A-40S |  |  | September 14, 2015 | 2015 | NeuroLogica Corporation | When the system is driving over a non-conductive floor and then over a conductive, grounded part, proper control of the drive behavior may not be possible. The system can be stopped by releasing the drive handle or by pressing the emergency switch. Proper behavior is re-enabled when the system is rebooted. | Device Design | An Urgent - Field Safety Notice identifying the affected products, problem involved, root cause, and potential hazards, was sent to customers. In the event the issue occurs, users can release the drive handle or press the emergency switch and the unit will stop. Then, reboot the system. If unintended movements occur, please contact local representative. To correct the issue, service engineers will visit to apply the corrective action. For further information or support concerning this problem, contact local representative. |  | 7 | 7 devices | Distributed to the states of NC, OH, MA, and MO. | Terminated | January 19, 2017 | 493 | 7 | Not_Computer
Z-0261-2016 | 72297 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Tapers LFIT V40 Vitallium femoral heads. hip prosthesis component. |  |  | September 08, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Complaints that the femoral head could not be assembled with its corresponding V40 stem trunnion at the time of surgery. | Process design | Stryker sent an Urgent Medical Device Recall Notifiications dated September 9 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to complete the attached form and return within 5 days by using email strykerortho6632@stericycle.com or FAX 1-888-965-5803.For further questions please call ( 201) 831-5000. |  | 201 | 201 | Nationwide Distribution | Terminated | July 25, 2017 | 686 | 201 | Not_Computer
Z-0264-2016 | 72475 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE VC150 Vital Signs Monitor |  |  | October 22, 2015 | 2015 | Innokas Medical Oy | Monitor may shut down unintentionally without restarting. | Software design | The firm, Innokas Medical, sent a "Urgent Field Safety Notice(s)" dated October 22, 2015 and October 23, 2015 by e-mail to their consignees. The notice informed the distributor of the device, about the deficiency, and correction of the device. The notice included the instructions regarding what to do with the product and information that the devices will be updated. The consignees were instructed to complete and return the Medical Device FSCA/Recall Return Response Acknowledgement and Receipt Form via fax to +358 8 562 3151, mail to: Innokas Medical Oy, Vihikari 10, FI-90440-Kempele, Finland or email to: feedback@innokasmedical.fi.If you have any questions, contact the QA/RA Director at +358 44 562 3131. |  | 424 | 424 units (38 in US) | Worldwide Distribution: US (nationwide) and Europe. | Open, Classified |  |  | 424 | Not_Computer
Z-0270-2016 | 72413 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Radius 5.5 to 3.5 Rod to Red Connector, REF 48664530. Intended to provide additional support during fusion using autograft or allograft in skeletally mature patients. |  |  | September 25, 2015 | 2015 | Stryker Spine | Stryker Spine is recalling Radius Rod to Rod Connector 5.5 to 3.5 mm (Catalog #486614530) because it does not have an associated 510(k) submisstion and subsequent clearance, which is required as this is a class 2 device. | No Marketing Application | Stryker Spine initiated this recall by sending an "Urgent Product Removal Letter and Product Accountability Form dated September 25, 2015 to the Branch Manager, Agency Manager, Quality Contact via FedEx priority overnight. |  | 42 | 42 units | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MT, NC, NM, NY, OH, OK, OR, PA, TN, TX, UT, WA and WI. | Open, Classified |  |  | 42 | Not_Computer
Z-0268-2016 | 72550 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Perfusion Tubing Packs, Models TL7M64R3, TL7P17R1 &amp; HY2W85R9, Sterilized using Ethylene Oxide. |  |  | October 28, 2015 | 2015 | Medtronic Perfusion Systems | This recall is being initiated due to a potential breach of the Tubing Pack Header Bags manufactured in these select lots. The seal in question is between the Tyvek and the Poly side of the header bag. | Packaging process control | Medtronic sent an Urgent Medical Device Recall letter dated October 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were advised to quarantine and return all affected product and to complete and return the Customer Confirmation certificate. For questions contact your Medtronic representative or contact Medtronic's national Answering service at 800-633-8766. |  | 496 | 496 | Nationwide Distribution including US: AL, IL, IN, and WI. | Open, Classified |  |  | 496 | Not_Computer
Z-0265-2016 | 72406 | No | LNS | Extracorporeal shock wave lithotripter. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MODULARIS VARIOSTAR; Lithotripter device designed to treat urolithiasis. |  |  | September 24, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Display freeze of MODULARIS hand control results in information not being updated on the display. Current treatment data is not shown to the user. | Software design | Siemens sent a Field Safety Notice (XP044/15/S) dated September 24, 2015 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 15 | 15 | Distributed in PR and the states of MO, NC, GA, MS, LA, and KY. | Open, Classified |  |  | 15 | Not_Computer
Z-0267-2016 | 72368 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Biomet 3i Narrow Right Angle Square Driver Tip. Driver tips are attached to a surgical drilling unit to carry components such as implant mounts, healing abutments and screws for placement and removal. For dental and restorative procedures. |  |  | August 10, 2015 | 2015 | Biomet 3i, LLC | RASQ3N, a square, abutment screw driver tip has been labeled and distributed as RASH3N, a hexed, abutment screw driver tip. Additionally, the RASQ3N received the RASH3N's laser marking "RASH3N" during manufacturing. | Labeling mix-ups | Beginning on or about 08/10/2015 customers will be notified in writing via FEDEX and instructed to to check their inventory for the identified product and to return any unused quantiles along with the recall response. Biomet 3i will promptly issue replacement product. |  | 103 | 103 pieces | AK, AZ, CA, CO, CT, FL, GA, IA, ID, IL, KY, LA, MA, MD, MI, MS, NC, NE, NJ, NY, OH, OR, PA, SD, TN, TX, UT, VA, WA, and District of ColumbiaArgentina, Australia, Belgium, Canada, Colombia, Greece, Hong Kong, Ireland, Italy, South Korea, and Kuwait. | Open, Classified |  |  | 199 | Not_Computer
Z-0269-2016 | 72396 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Ventilator Tubing |  |  | September 23, 2015 | 2015 | Instrumentation Industries Inc | As a distributor of tubing manufactured by Smooth Bore Plastics, we are initiating a product recall for specific lots of KC 036, KC 060 and KC 072 reusable ventilator tubes due to a possible manufacturing defect: A few tubes may have the potential for delamination to occur between the layers of the tube wall which, should it occur, would result in leakage | Nonconforming Material/Component | Instrumentation Industries notified consignees by e-mail and telephone of the recall. The firm also sent a Notice of Product Recall letter dated September 23, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were asked to determine if they have any of the affected product and contact the firm to make arrangements for return to Instrumentation Industries, Inc. Replacement product will be sent as soon as possible after all quality inspections have been completed. If product was further distributed consignees were asked to provide recall information to their customers. Customers with questions should call 412-854-1133, ext 433. |  | 567 | 567 | Worldwide Distribution - US including PA, FL, GA, WA, OH, CA, KA and Internationally to Canada. | Open, Classified |  |  | 567 | Not_Computer
Z-0272-2016 | 72449 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Easy Dial Oxygen Regulator, 168708D and 168715D. |  |  | October 08, 2015 | 2015 | Precision Medical, Inc. | These lots of Oxygen Dial Regulators could have the settings label off by one position. When the indicator arrow points to OFF position, the regulator is actually on the half liter setting (0.50), for the 15 liter. For the 8 liter, instead of being in the OFF position, the regulator will be in the quarter liter setting (0.25). | Error in labeling | The firm sent "Attention Field Corrective Action" letters, dated Oct 8, 2015, to customers. The letter identified the affected device and explained the reason for the recall. The letter provided instructions on how to check the regulator; and asked customers to call the number provided for a replacement if a regulator is labeled incorrectly. |  | 315 | 315 | Distributed in the US to the states of PA, MA, MO, MT, ID, MI, MN, ND, KY, NY, NC, CA, AL, AZ, LA, OH, TX, GA, IN, and OK. | Open, Classified |  |  | 315 | Not_Computer
Z-0273-2016 | 72477 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ECG Out Cable,for use with a Philips ALS monitor/ defibrillator and bedside monitor for Synchronized cardioversion and Transcutaneous pacing. M1783A/M5526A 12-pin Sync cables, used with HeartStart MRx and HeartStart XL and HeartStart XL+ |  |  | November 19, 2014 | 2014 | Philips Electronics North America Corporation | When a Philips monitor/defibrillator is receiving an ECG signal from an auxiliary bedside monitor via a sync cable, the following can occur if the monitor/defibrillator experiences interference from electrical fast transients (EFTs) while connected to AC power: On the HeartStart MRx and HeartStart XL, EFT noise can be misinterpreted as an R-wave. On the HeartStart XL+, EFT noise can disable ECG mo | Device Design | Philips sent an Urgent - Field Safety Notice dated November, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips plans to take to remedy the issue. A replacement cable will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the: Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: Prior to receiving a replacement sync cable, you can continue to use affected sync cables, provided that they are only used to connect the ECG Out jack on the monitor/defibrillator to the ECG (input) port on a bedside monitor. The issues described in this letter do not impact this application of the sync cable.Affected sync cables should not be used to connect the ECG Out jack of a bedside monitor to the ECG (input) port on the monitor/defibrillator.For further questions please call (800) 722-9377. |  | 18604 | 18,604 | Worldwide Distribution - US (nationwide) and to the countries of : Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Ethiopia, Finland, France, Gabon, Germany, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Republic of Korea, Latvia, Lebanon, Macau, Malaysia, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela and Yemen. | Open, Classified |  |  | 18604 | Not_Computer
Z-0212-2016 | 72409 | No | MCX |  |  | Cardiovascular | PMA | RotaWire Elite and wireClip Torquer, Guidewire and Guidewire Manipulation Device, Extra Support. REF H802233301 (5-pack outer package), REF H802233300 (single unit). |  |  | October 09, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific is recalling its recently released RotaWire Elite Guidewire and wireClip Torquer Guidewire Manipulation device that is used in conjunction with the Rotablator Rotational Atherectomy System because of complaints of wire fracture. | Component change control | The firm, Boston Scientific, sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated 10/9/2015. Letters were sent via overnight delivery (FedEx or other equivalent service). The letter stated the issue, identified affected devices, stated that further distribution or use of the devices should cease immediately, and affected product should be returned to Boston Scientific in accordance with the enclosed recall instructions. Complete and return the Recall Instructions via email: MapleGroveFieldActionCenter@bsci.com or Fax to: Field Action Center 1-866-213-1806. Healthcare professionals and consumer may report serious adverse events (side effects ) or product quality problems with the use of this product to Boston Scientific by calling 1-800-811-3211.If you have any questions, please contact Boston Scientific Quality Systems at 763-494-1133, email: MapleGroveFieldActionCenter@bsci.com or your local sales representative can answer any questions that you may have regarding this recall removal. |  | 955 | 955 single units (191 5-packs) | Worldwide Distribution: US Distribution in the states of AZ, CA, DE, FL, GA, IL, LA, MD, MI, MN, NE, NY, PA, SC, TX, WA, and WI, and the countries of Belgium, Germany, Great Britain, Italy, Japan, and South Africa. | Open, Classified |  |  | 955 | Not_Computer
Z-0274-2016 | 72423 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 100 Low Bed (GPAC) AC Powered adjustable hospital bed |  |  | October 07, 2015 | 2015 | Hill-Rom, Inc. | Hill-Rom is voluntarily recalling the Hill-Rom 100 Low Bed due to complaints of patient injury when exiting the beds set at the lowest position setting. | Device Design | Hill-Rom initiated Phase 1 of the recall on October 2, 2015, which included notices of the recall being issued to customers via certified mail, which included safety information on how to properly use the bed, and instructions to forward the information on if devices have been transferred. Phase 2 for the recall will include the distribution of caps for the edges of the bed to be placed on by the customers. Customers may contact Hill-Rom Technical Support at 800-445-3720. |  | 12251 | 12251 units | Worldwide Distribution - USA (nationwide) to the states of : AL, AR, AZ. CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV., and to the countries of : Canada and France. | Open, Classified |  |  | 12251 | Not_Computer
Z-0275-2016 | 71121 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GMK Hinge knee prosthesis; Model Number: 02.09.TA410; Orthopedics and Orthopedic Surgery |  |  | April 17, 2015 | 2015 | Medacta Usa | Incorrect parts: Packages of the GMK Hinge Tibial Augmentation size 4-10 mm have 5 mm screws instead of 10 mm screws. | Employee error | On 4/16/15 all distributors with recalled product were immediately notified of the recall by phone. On 4/20/15, all distributors were further notified of the recall via email. The URGENT MEDICAL DEVICE RECALL letter dated 4/17/2015 the firm described the initiation of a recall for regarding the GMK HINGE Tibial Augmentation Size 4-10mm Reference 02.09TA410 / Lot 145129.The firm provided the following recall actions to be performed:1.Check your inventory for the device listed above. If you find it, IMMEDIATELY place it in quarantine and DO NOT use it in surgery.2.If you have further distributed this device to other customers please notify me and initiate their return to Medacta USA.3.We will provide you with an update on a replacement within the next few days.The firm also provided a contact for customers requiring additional information.Mike Loiterman Director of Regulatory, Quality and Compliance Medacta USA Phone: 312548-9971Fax: 312-546-6881MLoiterman@medacta.us.com |  | 4 | 4 units (US) | Distributed in the states of IL, OH and UT, and the countries of in Austria, Spain, Australia, France and Switzerland. | Terminated | January 05, 2017 | 629 | 4 | Not_Computer
Z-0276-2016 | 72469 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument for Sternal ZIPFIX. Indications for use include primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion. |  |  | October 19, 2015 | 2015 | Synthes (USA) Products LLC | Certain lots of the Application Instrument for Sternal ZIPFIX that were previously recalled on August 19, 2015 and subsequently repaired and returned to customers between August 26, 2015 and September 23, 2015. The screws on the repaired instruments have the potential to break or loosen. DePuy Synthes is requesting the return of the repaired device in exchange for a new device. | Device Design | Communication to the accounts having record of receiving this product was sent on October 19, 2015. |  | 267 | 267 | Distributed in the states of KY, UT, CA, WI, OH, CO, IL, TX, OK, WA, MI, LA, MA, PA and NE. | Open, Classified |  |  | 267 | Not_Computer
Z-0278-2016 | 72285 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE+ 7 Laparoscopic Shears with Advanced Hemostasis (23cm, 36cm, 45cm length), Ethicon Endo-Surgery. Indicated for soft tissue incisions when bleeding control and minimal thermal injury are desired. |  |  | September 23, 2015 | 2015 | Ethicon Endo-Surgery Inc | Complaints of holes in Tyvek which can compromise the sterile barrier provided by the packaging. | Device Design | On October 5, 2015, the firm sent Urgent Medical Device Recall Removal Letters to their customers. |  | 441490 | 441,490 pieces | Distributed US (nationwide) including the District of Columbia &amp; Puerto Rico and in the countries of ARUBA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, ECUADOR, HONG KONG, INDIA, JAPAN, KOREA, MALAYSIA, MEXICO, PANAMA, PERU, PHILIPPINES, SINGAPORE, SINGAPORE, THAILAND, TAIWAN &amp; VIETNAM. | Open, Classified |  |  | 441490 | Not_Computer
Z-0279-2016 | 72552 | No | PAH | Surgical mesh. | General &amp; Plastic Surgery | Obstetrics/Gynecology | 510(k) | MiniArc Pro Single incision Sling System, Model number / Product Reference 9000261, Lot 936405xxx. Designed to treat female stress urinary incontinence. |  |  | October 26, 2015 | 2015 | Astora | The lot number on the outer carton label did not match the lot number printed on the tray label or identification labels inside the box. The identification labels inside the box are placed on medical records and other documents associated with the procedure. Use of incorrectly identified labels may prevent direct traceability of the implanted unit in the future. | Labeling mix-ups | Consignees were hand delivered by sales representatives a Astora Women's Health "Urgent Recall Notice". The letter was addressed to Physician, Health Care Professional, Nurse &amp; Risk Manager. The letter described the problem and the product involved in the recall. Consignees were advised to check their inventory and return all affected product, notify those that might be affected and to complete and return the Acknowledgement Form. For questions contact ASTORA Women's Health Customer Service at 1-844-879-0787. |  | 52 | 52 | Distributed in the states of AZ, CA, DE, IL, KS, NE, NJ, NY, OH, PA, TX | Terminated | March 09, 2016 | 135 | 52 | Not_Computer
Z-0277-2016 | 72418 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Taut Intraducers 10/BX7.5 FR 3.5, Product Code PI-93. To provide an access port into the peritoneal cavity. |  |  | October 14, 2015 | 2015 | Teleflex Medical | The defective component can cause gas leakage through the device. This can result in loss of pneumoperitoneum, potentially necessitating removal of the defective device and insertion of another device. | Nonconforming Material/Component | Recall letters were sent to customers and distributors. The letter requested that the accounts discontinue use and quarantine for return any product. The distributors were requested to conduct a sub-recall. The letters included a Recall Acknowledgement Form which was to be returned to Teleflex. |  | 2050 | 2050 ea. | Nationwide Distribution including the states of AK, AZ, AR, CA, CT, FL, GA, ID, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MO, MT, NE, NJ, NY, NC, OH, OK, PA, TN, TX, VA, VT, and WI. | Open, Classified |  |  | 2050 | Not_Computer
Z-0287-2016 | 72378 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Needle Guide Starter Kit, Reusable non-sterile bracket with Infiniti Plus (12, 14, 16, 18GA) needle guides, Rx ONLY. REFs 674-030, 674-030-T, 674-031, 674-031-T, 674-035, 674-036, 674-044, 674-052, 674-059, 674-061. Large Gauge Infiniti Plus Needle Guides, used in needle-guided (or catheter) procedures with diagnostic ultrasound transducers. |  |  | October 21, 2015 | 2015 | Civco Medical Instruments Co. Inc. | Sterility of the product cannot be assured. | Process design | Civco Medical Solutions sent an Urgent Medical Device Recall letter dated October 19, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected needle guides. If the needle guide is contained within a kit, only the needle guide from the kit needs to be returned. The communication contains a response form to be completed and returned by the customer. The customer filling out the response form is asked to acknowledge receipt of the recall notice and indicate the number and type of affected product they are returning. Customers withh questions were instructed to call 800-445-6741 or 319-248-6757 or email order@civco.com.For questions regarding this recall call 319-248-6502. |  | 508 | 508 boxes | Worldwide Distribution - US (nationwide) and Internationally to AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHILE, CHINA (MAINLAND), CHINA (TAIWAN), CZECH REPUBLIC, DENMARK, ECUADOR, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, ISRAEL, ITALY, JAPAN, KUWAIT, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PUERTO RICO, QATAR, SOUTH AFRICA, ROMANIA, SAUDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, SWITZERLAND, TRINIDAD AND TOBAGO, TURKEY, UGANDA, UKRAINE, and UNITED KINGDOM. | Open, Classified |  |  | 2624 | Not_Computer
Z-0281-2016 | 72506 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker PCA 10 Degree Hooded Acetabular Insert ID 22mm, |  |  | October 15, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Six (6) PCA 10 Degree Hooded Acetabular Insert-Duration 22 mm (ID) have been distributed in the US, although the 22 mm size was inadvertently omitted from 510(k) #K963612. | Under Investigation by firm | Stryker sent an Urgent Medical Device Recall Notification letter dated October 15, 2015, to all affected customers. The notification instructed consignees on how to identify affected product; the related issue with the affected product; potential hazards; risk mitigations; and actions needed to be taken. Customers were instructed to quarantine the affected product. Consignees were asked to complete and return the attached Product Recall Acknowledgement Form within 5 days and either email (strykerortho7984@stericycle.com or Fax (1-866-672-0627 ) the response form back. Customers with questions were instructed to call 201.831.5272. |  | 1 | 1 unit | Distribution to Virginia only. | Terminated | June 08, 2017 | 602 | 1 | Not_Computer
Z-0283-2016 | 72411 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX, PACS, Picture Archive and Communications System, IMPAX CV 7.8.x and IMPAX CV 12.x. Corrects certain demographic information (Patient Name, Patient Medical Record Number and/or Accession Number in previously signed reports that have since changed. |  |  | September 01, 2015 | 2015 | AGFA Healthcare Corp. | Customers have experienced IMPAX CV software, specifically, Demographics Manager (DM), not producing correct pdf reports after the reports, at the time of physician signature, were correct. | Software design | An Urgent Field Safety Notice was sent to accounts via email on 9/1/2015. The letter described the safety alert. Customers were informed of the problem and that Service would be contacting them to arrange a date to deploy the corrective action and install the Service Update at their facility. An acknowledgement was included which was to be faxed back or emailed back to Agfa. |  | 19 | 19 | Distributed in the states of CA, KY, NH, NY, NC, OH, OR, PA, SC, SD, TN, TX, WA and the country of Canada. | Open, Classified |  |  | 19 | Computer
Z-0291-2016 | 71818 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter HomeChoice and HomeChoice Pro APD systems, designed to provide Automated Peritoneal Dialysis (APD) therapy for pediatric and adult renal patients. |  |  | June 29, 2015 | 2015 | Baxter Healthcare Corp. | Loud operating sounds, which was unacceptable to the end users when the device was powered on. The HomeChoice devices continue to perform as intended, as the issue is only associated with the noise coming from the device and not the function of it. May delay treatment if changing to a different system. | Device Design | Baxter sent an Important Product Information communication to affected customers (patients and clinicians) via U.S.P.S., first-class mail on June 29, 2015. The letter states that Baxter Healthcare Corporation is updating the HomeChoice Patient At-Home Guide to include a discussion of the expected operating sounds associated with the HomeChoice and HomeChoice PRO devices. The new labeling will address the following:The HomeChoice cycler utilizes an air (pneumatic) pump; some sounds related to the operation of this pump are expected as part of normal operation of the HomeChoice cycler. There will be sounds like humming, swishing, clicking, and venting (air being released) that are normal for the operation of the cycler. In certain portions of the therapy, the sound level is expected to increase which is also part of normal operation. If there is a significant change in the sound level or a new, previously unheard sound when using the HomeChoice cycler, please contact your doctor and/or nurse or Baxter Technical Services. The firm advises customers to keep the letter with the Patient At-Home Guide.The firm also requests that customers complete the enclosed home patient reply form (enclosure) and return it to Baxter by faxing the reply form to 224-270-5457, scanning and e-mailing it to fca@baxter.com or by mail in the enclosed, preaddressed, stamped envelope. Returning the home patient reply form promptly will prevent you from receiving repeat notices. If the customer has any questions about their Peritoneal Dialysis (PD) therapy, they are advised to please contact their doctor and/or nurse.Additionally the letter states that any adverse reactions or quality problems experienced with the use of this product may be reported using one of the following options:" Calling Baxter Product Surveillance at 800-437-5176, Monday through Friday, between 8:00 am and 5:30 pm Central Time Baxter will advise customers of the additional product information regarding |  | 29355 | 29355 | Nationwide Distribution, District of Columbia, and Puerto Rico. | Open, Classified |  |  | 29355 | Not_Computer
Z-0282-2016 | 71248 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BrainLAB Image Guided Surgery (IGS) System, Stereotaxic Instrument navigation software, Cranial ENT version 2.1, and Spine and Trauma 3D versions 2.0 and 2.1 An intraoperative image-guided localization system to enable minimally invasive surgery. |  |  | May 08, 2015 | 2015 | Brainlab AG | Instances of data sets not being accurately registered to the patient anatomy were observed. | Nonconforming Material/Component | A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated April 30, 2015 was sent to all direct accounts (customers). The letters included instructions for customers to: To reduce the risk of an inappropriate automatic registration, Brainlab, in cooperation with NeuroLogica, has developed an improved integration- and verification-workflow to be followed, and additionally wants to re-iterate important use instructions, as described below and specified in detail in the Appendix of this document. When using Brainlab's navigation software to perform automatic registration of CT image data sets in combination with NeuroLogica BodyTom, please ensure that the following conditions are always fulfilled: 1) Ensure suitability of OR Ensure that the operating room where the BodyTom is used in combination with the Brainlab navigation has a continuously even floor. The flatness must not exceed 1mm (0.04 in) for 1m (39.4 in). Please take into account that the flatness may change over time due to wear of the floor coating material and/or negative impacts under the load of equipment used in the room. Ensure that the OR still fulfills the initial specifications before use. 2) Ensure a valid calibration of the BodyTom Automatic registration with the BodyTom may only be used if a valid calibration and verification of the BodyToms mechanical position is given. Both BodyTom and Brainlab calibration must be performed on the same position and movement path that is used for the navigated clinical procedure. The use of Brainlab navigation system is restricted to the one room and movement path where the calibration has been performed. If multiple calibrated rooms are listed on the BodyTom, please ensure that the correct room is selected on the BodyTom workstation when preparing the scan. As part of the Brainlab Corrective Action (see below) you will be contacted by a Brainlab employee to assist in the selection and marking of a suitable movement path and to perform th |  | 12 | 12 systems total (10 in U.S., 2 in Australia) | Distributed in the US and Australia. | Open, Classified |  |  | 12 | Computer
Z-0280-2016 | 72385 | No | LYJ |  |  | Neurology | PMA | Cyberonics Vagus Nerve Stimulator (VNS) Therapy AspireSR Generator, Model 106 |  |  | October 07, 2015 | 2015 | Cyberonics, Inc | Certain Model 106 Pulse Generators demonstrate delays in sensing during use of the 'Verify Heartbeat Detection' feature and exhibit potential for decreased battery longevity. | Manufacturing material removal | Cyberonics sent written recall notifications beginning October 14, 2015, to all affected customers. The letter identified the product the problem and the action to be taken by the customer. The notice also directs customers to contact Cyberonics to schedule additional on-site troubleshooting. The on-site troubleshooting proposed by Cyberonics will utilize a feature included within the Model 250 Version 9.0 Programming Software as an alternate method of verifying heart rate and determining that the appropriate heart beat sensitivity setting is programmed. The Version 9.0 software displays an instantaneous heart rate when the user selects the 'Refresh' feature. This heart rate is measured directly by the generator and displayed on the programmer.For further questions, please call (866) 882-8804. |  | 27 | 27 units | US Distribution to the states of : OR, NY, GA, PA, CA, NC, KS, FL, KY, MI, TX, IL, OH, IN, UT, CO and WV. | Open, Classified |  |  | 27 | Computer
Z-0284-2016 | 71680 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical C-300 Consumer Ceiling Lift; PN 330050, 330050R, 330053, 330053R, and 330057; weight specification 300 lbs. /136 kg. |  |  | September 08, 2015 | 2015 | ErgoSafe Products, LLC (DBA) Prism Medical | Drive shaft and over speed arm may simultaneously fail potentially resulting in a patient falling during use | Device Design | The firm, Prism Medical, sent a "Field Safety Notice (FSN 0008)" letter dated 9/8/2015 to all consignees by email via their sales portal. The letter describes the product, problem and actions to be taken. The consignees were instructed to call Prism Medical directly at 1-314-219-8655 and indicate that you are participating in the 300 series lift recall; provide your existing 300 series lift model number and serial number; follow the instructions provided; complete and return the Acknowledgement Form via scanned &amp; signed to bschoenman@prismmedicalltd.com or fax to: 314-692-7858, Attention: 300 series recall.For further information please contact: Prism Medical Quality department at: 1-314-219-8655 or North American Director for Quality &amp; Regulatory Affairs at 1-314-219-8614. |  | 561 | 561 lifts | Worldwide Distribution: US (nationwide) and countries of: Australia, Canada, South Korea, Taiwan, Mexico, Japan, and Great Britain. | Terminated | May 04, 2017 | 604 | 2282 | Not_Computer
Z-0298-2016 | 72471 | No | JXM | Benzodiazepine test system. | Toxicology | Toxicology | 510(k) | 1 Step 12 Panel Cup, Item No. NBCA-12M-W |  |  | October 20, 2015 | 2015 | Ameditech Inc | Ameditech is recalling drug abuse tests because they have shown reduced reactivity. | Nonconforming Material/Component | Ameditech Inc. sent an Urgent Medical Device Recall Letter to al affected customers on October 20, 2015, to inform them that Ameditech is recalling affected lots of Ameditech's Drugs of Abuse Tests because they have shown reduced reactivity for up to three of the nineteen Benzodiazepine (BZO) compounds. The letter informs customers of the actions to be taken. Customers were instructed to complete and return the customer verification form to (858) 677-0243 or scan and email to amt.fieldaction@alere.com. For questions regarding this recall call 858-535-1968. |  | 787 | 787 kits | Worldwide Distribution - US (nationwide) and Internationally to Australia. | Open, Classified |  |  | 149975 | Not_Computer
Z-0294-2016 | 72484 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 132 Gallon GranuFlo Dissolution Unit - P/N G047-80101. |  |  | October 23, 2015 | 2015 | Fresenius Medical Care Renal Therapies Group, LLC | Dissolution units and replacement motors units do not provide full thermal protection for both the Main and Starter coils resulting in the generation of smoke and/or potentially lead to fire. | Device Design | Fresenius issued an Urgent - Medical Device Correction notification, dated October 21, 2015, on 10/27/15. The letter identified the affected device and the reason for the recall. It also contained instructions (on how to perform the required wiring modification) along with a Reply Form. Customers are to return the form after completing the modification. Inquiries regarding the correction should be directed to Fresenius Technical Services at 800-227-2572, referencing the Field Action number FA-2015-01-W. |  | 918 | 918 units | Nationwide Distribution. | Open, Classified |  |  | 2258 | Not_Computer
Z-0293-2016 | 72451 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Neptune Plus, HEMOSTASIS PAD WITH ANTIMICROBIAL BARRIER, For hemostasis following needle/sheath removal, Part number 8870-03 |  |  | October 14, 2015 | 2015 | TZ Medical Inc. | The IFU of the Neptune Plus Pads, does not match that submitted in the original 510k. Need to remove: for haemostasis following needle/sheath removal | Labeling False and Misleading | TZ Medical Inc. began sending the Urgent: Medical Device Correction 8870-03 Neptune Plus letter, dated 10-7-2015, to consignees in October 2015. The firm is notifying all users of Neptune Plus wound dressing to a correction to the Instructions for Use (IFU). Consignees should discard the IFU in your current inventory and replace with new IFU. Consignees with questions should call TZ Medical Inc. at (503) 639-0282. |  | 32170 | 32,170 units | US nationwide distribution. | Open, Classified |  |  | 32170 | Not_Computer
Z-0297-2016 | 72427 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Bulk Loader Module for cobas p 512/612 |  |  | September 22, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Possible for sample tubes in the loader to be opened and contaminate the system as well as adjacent samples during processing. | Device Design | Roche Diagnostics Operations, Inc. sent an Urgent Medical Device Correction letter to affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were asked to retain the Urgent Medical Device Correction letter until a Roche Field Service Representative has installed the BLM modification kit. Consignees were also asked to complete the attached form and fax to 1-877-906-8982, ext 2015. Consignees with questions were instructed to call 1-800-428-2336.For questions regarding this recall call 317-521-3911. |  | 1 | 1 | Distributed to TX only. | Terminated | February 03, 2016 | 134 | 1 | Not_Computer
Z-0292-2016 | 72575 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | ECG Electrodes: |  |  | November 02, 2015 | 2015 | Leonhard Lang Medizintechnik GmbH | There is a risk ECG electrodes of specific lots fail to transmit signals. There is a risk that the affected electrodes will transmit a poor signal or no signal at all before reaching the indicated expiration date. | Nonconforming Material/Component | Leonhard Lang emailed an "Important Safety Notice" letter dated November 2, 2015, to all affected customers on November 2, 2015.. The letter identified the product the problem and the action needed to be taken by the customer. The letter provided information on: 1. Description of the defect; 2. Actions and time frame of the recall; 3. Replacement of the electrodes &amp; contact information. For questions call : +43 512 33425-39. |  | 16410800 | 16,410,800 (10,226,600 US:, 184,200 OUS) | Worldwide Distribution - US Distribution to the state of : FL., and to the countries of : Australia, Austria, Croatia, Czech Republic, Estonia, Finland, France, Germany, Great Britain, Greece, Ireland, Italy, Japan, Malta, Norway,Poland, Portugal, Romania, Serbia, Spain, Switzerland, Thailand, Turkey and United Arab Emirates. | Open, Classified |  |  | 16410800 | Not_Computer
Z-0230-2016 | 72491 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Inspiration Ventilator System, i Series and LS 12.1 LCD. The Inspiration 7i / 5i Ventilator System is to be used by healthcare professionals in hospitals or healthcare facilities and intra-hospital transport. |  |  | October 13, 2015 | 2015 | Event Medical LTD | A component on the systems power board can fail in a manner which can cause the ventilator to shut down without alarming. | Component design/selection | eVent Medical sent an Urgent Field Safety Notice dated October 12, 2015, to all affected customers on October 13, 2015, to inform them that eVent Medical is recalling Inspiration Ventilators i Series and LS 12.1 LCD because a component on the systems power board can fail in a manner which can cause the ventilator to shut down without alarming. The letter provides the customers with the problems identified and the actions to be taken. Customers with questions are instructed to call (949) 900-1917, x232. |  | 6441 | 6441 units total (251 units in US) | Worldwide Distribution - US (Nationwide) and Internationally to:AlgeriaArgentinaAustraliaAustriaBahrainBangladeshBarbadosBelarusBelgiumBoliviaCanadaChileChinaColombiaCzech RepEcuadorEgyptEnglandEstoniaFinlandGermanyGhanaGibraltarGreeceGuatemalaHollandHong KongHungaryIndiaIndonesiaIranIraqIrelandIsraelItalyJapanJordanKenyaKoreaKuwaitLaosLatin AmericaLebanonLibyaLithuaniaMalaysiaMexicoMoroccoNepalNetherlandsNew ZealandNorwayPakistanPalestinePanamaPeruPhilippinesPolandPortugalRussiaSaudi ArabiaSerbiaSingaporeSouth AfricaSpainSri LankaSudanSwitzerlandSyriaTaiwanThailandTurkeyUAEUkraineUSAVenezuelaVietnamYemenZimbabwe | Terminated | June 22, 2017 | 618 | 6441 | Not_Computer
Z-0318-2016 | 72450 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | SilverStat, sterile 2" x 2 " calcium alginate hemostasis pad with sliver nitrate, Part number 8870-07. |  |  | October 14, 2015 | 2015 | TZ Medical Inc. | During a review of the IFU of the SilverStat Pads, it was found that the IFU did not match that submitted in the original 510k | Other | TZ Medical Inc. sent an Urgent: Medical Device Correction 887007 SilverStat letter, dated October 13, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard the Instructions for Use (IFU) in their current inventory and replace with new (IFU). Consignees with questions should call TZ Medical Inc. at (503) 639-0282. |  | 10180 | 10,180 units | US in the states of AL, AR, CA, DC, FL, GA, IA, MD, NC, NE, NH, PA, RI, SC, TN, TX, UT, and WV. | Open, Classified |  |  | 10180 | Not_Computer
Z-0322-2016 | 72537 | No | NTR | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Thermo Scientific p21 WF1 Ab-3 (DCS-60.2) |  |  | October 27, 2015 | 2015 | Lab Vision Corporation | It was discovered that the labeling of MS-230 antibody incorrectly describes the concentration - label reads 0.2mg/ml but it is actually 0.4 mg/ml. | Labeling design | The firm, ThermoFisher Scientific, sent an "URGENT MEDICAL DEVICE RECALL-Lab Vision Corporation" letter dated 10/29/2015 to its customers on 10/30/2015. The letter describes the product, problem and actions to be taken. The customers were instructed to Upon identification of the affected lot, notify Lab Vision of the quantity previously used and the quantity still in inventory; review any associated test results to ensure an accurate diagnosis was able to be made; If routine laboratory use of this product follows the Lab Vision Data Sheet, the antibody may safely be used as is; hold any affected material in quarantine until the new labels are received-affected product may also be returned to Lab Vision, and complete and return the attached Recall Return Response Acknowledgement &amp; Receipt Form via email to: sarah.rickert@thermofisher.com or Fax to: 269-372-2674, Attn: Quality Assurance &amp; Regulatory Affairs Manager. A new label with the correct concentration can be sent upon request, for relabeling of the product at the customers facility.If you have any questions, please contact Quality Assurance &amp; Regulatory Affairs Manager at 269-544-5628 or sarah.rickert@thermofisher.com. |  | 5 | 5 vials | Worldwide Distribution: US distribution to states of: CA and WA, and countries of: Serbia, Germany and Netherlands. | Open, Classified |  |  | 5 | Not_Computer
Z-0324-2016 | 72529 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SCFE Driver (SCF-MLD273 &amp; SCF-MLD265) instrument used for the insertion of the Free Gliding SCFE Screw System. |  |  | October 20, 2015 | 2015 | Pega Medical Inc. | SCFE Driver's threaded shaft broke close to the knob when a manual force was applied during the fixation of the screw. | Unknown/Undetermined by firm | The firm sent the Urgent Field Safety Notice-Recalls to US consignees via email. Consignees are advised to return affected devices back to the firm. Customer with questions can contact Enrique Garcia at 450-688-5144 ext. 242 or send email to egarcia@pegamedical.com |  | 4 | 4 units distributed in the US | Distributed in the states of FL and TN. | Open, Classified |  |  | 4 | Not_Computer
Z-0320-2016 | 72344 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator; M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6. |  |  | October 14, 2015 | 2015 | Philips Electronics North America Corporation | The following MRx software issue has been identified: MRx model M3535A with software version F.03.06 and earlier, and model M3536A with version T.00.05 and earlier may stop the automated Ready-For-Use (RFU) test in an abnormal state when the device is turned off. | Software design | On 10/14/2015, Philips sent an Urgent Field Safety Notice to each known affected customer. The Urgent Medical Device Correction notification/ Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will provide a software upgrade for M3535A (software version F.03.07 or higher) and M3536A (software version T.00.06 or higher) to correct this issue. A software upgrade will be provided free of charge for all affected products. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: All MRx users should be informed that if the pattern appears, they can continue to use the device. Once the MRx is turned on, the pattern will disappear and the device will function normally in all monitoring and therapy modes. |  | 81161 | 81,161 | Worldwide Distribution - US (nationwide) including Puerto Rico, and the countries of AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GHANA, GREECE, GUADELOUPE, HONG KONG, Hungary, ICELAND, India, INDONESIA, ISLAMIC REPUBLIC OF IRAN, Iraq, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LAOS, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACAU, MACEDONIA, MALAYSIA, MALTA, Mauritius, Mexico, MONGOLIA, MOROCCO, MOZAMBIQUE, Myanmar (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, State of Palestine, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH, AFRICA, SPAIN, SRI, LANKA, SWEDEN, SWITZERLAND, TAIWAN, UNITED REPUBLIC OF TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED, ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN | Open, Classified |  |  | 81161 | Computer
Z-0321-2016 | 72541 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Tool Legend 16cm 2.4 mm wire pass for Midas Rex Legend Dissecting Tool. Used by surgeons to drill holes. |  |  | October 21, 2015 | 2015 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | The wire pass hole is not present on the tools. | Process control | A recall letter was mailed to consignees on October 21, 2015. |  | 9 | 9 pieces | Distributed in the states of California, New York, and Texas and the country of Belgium. | Open, Classified |  |  | 9 | Not_Computer
Z-0319-2016 | 72576 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging VB36D_HF02. Radiological image processing system. |  |  | October 28, 2015 | 2015 | Siemens Medical Solutions USA, Inc | To provide supplementary information regarding the release of the syngo Imaging Software Version VB36D_HF02. The software provides improvements for all syngo Imaging installations running on the software Version VB36D. | Software change control | Siemens sent an Customer Information Letter dated October 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call: (610) 219-6300. |  | 3 | 3 | US Distribution to the states of : NC, NE and OH. | Open, Classified |  |  | 3 | Computer
Z-0323-2016 | 72553 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 2.4MM locking screw SLF-TPNG with Stardrive Recess 18 mm. Intended for fixation of fractures, osteotomies, nonunions, replantations, and fusions of small bones and small bone fragments, particularly in osteopenic bone. |  |  | November 02, 2015 | 2015 | Synthes (USA) Products LLC | It was discovered that the above part number and lot of 2.4mm Locking Screws may contain the incorrect package insert (GP2613, VA-LCP Distal Radius System). The correct package insert for the above part is GP2615 (Modular Mini Fragment LCP System). | Process control | Synthes sent an "URGENT FIELD SAFETY NOTIFICATION" letter dated November 2, 2015. Customers were asked to take the following actions:-Review the correct package insert (GP2615, attached).-If found, discard incorrect package insert associated with the locking screw noted in the table above.-Forward this Field Safety Notification to anyone in your facility that needs to be informed.-If the package insert has been forwarded to another facility, contact that facility.-Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided. ---Send a copy of the completed Verification Section by:-Fax: 877-907-7514 or Scan/email: Synthes7766@stericycle.com-If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification.-Maintain a copy of this Notification. |  | 229 | 229 | Distributed in the states of IA, OH, CO, FL, TX, MA, NY, OR, SC, VA, WV, TN, MO, PA, CT, GA, WA, CA, AL, , IN, MI, NJ, VT, ID, KS, MN, and MD, and the country of Canada. | Open, Classified |  |  | 229 | Not_Computer
Z-0325-2016 | 69127 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia, Intera, Achieva and Achieve dStream MR systems on Software versions R5.1.7 and R5.1.8 with the Mobi View software option. Diagnostic imaging system. |  |  | August 21, 2014 | 2014 | Philips Medical Systems, Inc. | When a fused series of a sagittal, coronal or radial multi station scan is generated in MobiView, the resultant image order is reversed. This occurs with software release R5.1.7/R5.1.8. | Software Design Change | Philips Healthcare sent an Customer Information Letter dated July 2014, on August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips will issue a Field Change Order to install updated software in a Service Pack that makes the slice numbering consistent for fused and unfused images.If you need any further information or support concerning this issue, please contact your local Philips representative: 1-800-722-9377 |  | 307 | 307 units | Worldwide Distribution - USA (nationwide) and to the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Denmark Estonia, Ethiopia. France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Italy,Japan, Jordan, Korea, Republic of, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Oman, Palestinian territory, Peru, Poland, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 307 | Computer
Z-0326-2016 | 72543 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage System One S Cycler (High Permeability Hemodialysis System) |  |  | October 29, 2015 | 2015 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal | Software Design Change | NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). |  | 323 | 323 units | Worldwide Distribution. US Nationwide, Netherlands, and Sweden. | Open, Classified |  |  | 339 | Computer
Z-0328-2016 | 72197 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CONSULT(TM) diagnostics hCG CONTROLS Kit, Part/Reorder Number 5011, IVD |  |  | September 14, 2015 | 2015 | Biochemical Diagnostics Inc | CONSULT diagnostics hCG Controls have discrepant storage temperatures listed on the kit labels when compared with the individual vial labels and the package insert. | Labeling False and Misleading | Biochemical Diagnostics sent a Notice of Correction Letter (dated August 26, 2015) to customers via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return the verification form, per the notification. For questions contact Alere Technical Service at 866-216-0094 or by email at verifications.at verifications.ts@alere.com. |  | 4473 | Domestic: 4,473 kits | US Nationwide Distribution | Terminated | January 26, 2016 | 134 | 4473 | Not_Computer
Z-0332-2016 | 72416 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Palindrome Precision Chronic Catheter Kit Symmetrical Tip, Tal VenaTrac" Stylet 14.5 Fr/Ch (4.8mm) X 33 cm |  |  | October 14, 2015 | 2015 | Covidien LLC | Incorrectly packaged with a 90 degree bend at the distal tip | Packaging process control | The firm, Medtronic (Covidien), sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/16/2015 to their customers. The letter described the product, problem, and actions to be taken. Customers were instructed to quarantine and discontinue use of the product, return affected product and complete and return the RECALLED PRODUCT RETURN FORM via Fax to Medtronic to (800)-895-6140 or email to: feedback.customerservice@Covidien.com/Distributor to (203) 492-7719 or email to: PalCathFCA@Covidien.com.Questions or concerns contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 150 | 150 units | Worldwide distribution: US (nationwide) and countries of: Canada, Belgium, Chile, Panama, Denmark, France, Germany , Netherlands,Spain, Switzerland, and United Kingdom. | Terminated | June 02, 2017 | 597 | 965 | Not_Computer
Z-0334-2016 | 72548 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD CD64 APC-R700; |  |  | October 30, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Three lots of CD64 (MD22) are contaminated with CD4 antibody. | Process control | Urgent Product Recall letters were sent out on October 30, 2015 via certified mail to notify customers of the issue. |  | 23 | 23 vials | Distributed in the states of FL, WA, PA, IL, GA, RI, TX, ND, and IA. | Open, Classified |  |  | 23 | Not_Computer
Z-0341-2016 | 72288 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Lyophilized Mouse Monoclonal Antibody CD10 (CD10). Immunohistochemistry. in vitro diagnostic. 0.1ml NCL-CD10-270 and 1mL NCL-CD10-270 are intended for the qualitative identification by light microscopy of CD10 molecules in paraffin sections. |  |  | September 03, 2015 | 2015 | Leica Microsystems, Inc. | These Lot numbers are not stable up to the expiry date on the product labeling. | Nonconforming Material/Component | Leica sent a MEDICAL DEVICE RECALL NOTIFICATION letter dated September 3, 2015 via mail to all of the consignees. The letter identified the affected, product, problem and actions to be taken. Customers are instructed to discontinue use of the reagent lots listed in the recall notification. It is requested that customers appropriately destroy any unused or partially used affected lots of the reagent, and indicate on the attached Medical Device Recall Notification Acknowledgement Form that this action has been undertaken. Customers were asked to confirm receipt of the notice as soon as possible by signing and dating the attached Medical Device Recall Notification Acknowledgement Form and fax it back to1-847-236-3747, to receive youralternative replacement for applicable orders. For questions contact your local Leica representative. |  | 66 | U.S: 1 unit of CD10-270-CE-S &amp; 66 units of CD10-270-CE. Foreign: 83 units of CD10-270-CE-S &amp; 219 units of CD10-270-CE | Worldwide Distribution - US Nationwide, Australia, Austria, Brazil, Canada, China, Czech Republic, Finland, France, Germany, Hungary, Indonesia, Ireland, Israel, Italy, Japan, Netherlands, Norway, Romania, Singapore, Slovenia, South Korea, Spain, Sweden, Switzerland and UK. | Open, Classified |  |  | 66 | Not_Computer
Z-0271-2016 | 72377 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Expansion of previous recall of AcrySof IQ ReSTOR and AcrySof IQ ReSTOR Toric IOL; now to include AcrySof IQ Toric models SN6AT6, SN6AT7, SN6AT8, and SN6AT9. |  |  | September 29, 2015 | 2015 | Alcon Research, Ltd. | Continued increase in reports of post-operative inflammation in patients who received AcrySof IQ Toric IOL. | Process design | The firm, Alcon, sent an "FDA Notification-Update to Recall Z-2323-2015" letter, dated October 1, 2015, to its Consignee in Japan who have received AcrySof IQ Toric IOL models SN6AT6 to SN6AT9 in order to initiate the recovery of product. The letter described the product, problem and actions to be taken. The consignee was instructed to stop using these immediately and put on hold until sales representatives visit their site for withdrawal. Recall does not affect AcrySof IQ Monofocal IOLs or AcrySof IQ Toric IOL models SN6AT3, SN6AT4 and SN6AT5. |  | 43651 | 43,651 units | International Only Distribution -- Japan. | Terminated | June 30, 2017 | 640 | 43651 | Not_Computer
Z-0340-2016 | 72392 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry Hemoglobin A1c_3, A1c_3M Reagent kits. Used on the ADVIA 1200, 1650, 1800, 2400, and XPT Chemistry Systems. For in vitro diagnostic use in the quantitative determination of Hemoglobin A1c, a diabetes marker, in whole blood on the ADVIA Chemistry systems. |  |  | September 24, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Reagent in these lots may demonstrate an increased occurrence of high %HbA1c bias, a positive bias averaging 0.6% HbA1c units, ranging from -0.1% to 1.1% HbA1c units. The maximum bias was observed at higher %HbA1c concentrations. QC samples may exhibit a similar bias. | Under Investigation by firm | Siemens has issued an Urgent Medical Device Recall in the US and an Urgent Field Safety Notice Outside the US on September 24, 2015, instructing customers to discontinue use, discard reagent and request no charge replacement kits for these products. |  | 3031 | 3031 | Distributed in the states of AZ, NY, CA, IN, LA, MN, OH, MI, NM, and GA, and the countries of Mexico and Canada. | Terminated | January 13, 2017 | 477 | 3031 | Not_Computer
Z-0338-2016 | 72626 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | rainbow Reusable Sensors. RAINBOW DCI-DC3, SpO2/SpCO/SpMet, Part No. 2201; RAINBOW DCI-dc8, SpO2/SpCO/SpMet, Part No. 2407; RAINBOW DCI, SpO2/SpCO/SpMet, Part No. 2696; RAINBOW DCIP, PED, SpO2/SpCO/SpMet, Part No. 2697. rainbow reusable sensors are indicated for spot check or continuous non-invasive monitoring of arterial saturation (SpO2), arterial carboxyhemoglobin saturation (SpCO), arterial methemoglobin (SpMet), and pulse rate. |  |  | October 28, 2015 | 2015 | Masimo Corporation | These sensors were manufactured with incompatible configurations. This could result in sensors that will provide either no readings or inaccurate readings for the SpCO and SpMet parameters. | Finished device change control | Masimo sent an Urgent Product Recall letter on October 28, 2015, to all affected customers to inform them that Masimo has identified specific lots of sensors that were manufactured with incompatiable configurations. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to complete the tracking/verification form and email to Masimo. Customers with questions were instructed to contact Masimo Technical Services at http://Service.Masimo.com. For questions regarding this recall call 949-297-7000. |  | 3476 | 3476 units total (3322 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Costa Rica, France, Germany, Kuwait, Martinique, Netherlands, Norway, South Korea, and Turkey. | Terminated | May 18, 2016 | 203 | 3476 | Not_Computer
Z-0339-2016 | 72399 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista System (MG) ; Flex reagent cartridge Catalog number K3057, SMN 10445158. For the quantitative measurement of magnesium in human serum, plasma and urine on the Dimension Vista System. |  |  | September 30, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Erroneous low results on a small number of reagent wells. There is the potential for under-recovery on certain wells for quality control (QC) and patient sample results. The under-recovery ranged from -0.3 mg/dL [0.12 mmol/L] to -1.6 mg/dL [0.66 mmol/L]; equivalent to a maximum negative bias of -88%. | Nonconforming Material/Component | An Urgent Medical Device Recall dated September 2015 was provided to all Dimension Vista MG customers who received the affected lot to notify them of the issue, the potential risk to health, and actions to be taken by customers.1. An Urgent Medical Device Recall was provided to all customers who received the affected lot to notify them of the following:Siemens has received customer complaints for Dimension VistaÂż MG lot 15063BA, regarding sporadic depressed serum/plasma and QC results.Discontinue use and discard their remaining inventory of the affected lot.Review the letter with their Medical Director. Siemens will replace any unused inventory of the affected lot. |  | 3601 | 3601 | Distributed US (nationwide) including DC and the states of IL, CA, GA, LA, TX, TN, FL, SC, IA, TN, DE, NJ, WA, MS, KY, NY, WI, OH, OH, MO, UT, KS, VA, AR, WV, MA, PA, AL, CO, IN, MT, NM, ME, AZ, NM, MO, and NC, and the country of Canada. | Open, Classified |  |  | 3601 | Not_Computer
Z-0335-2016 | 72554 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage System One S Cycler - Model no. NX1000-3. |  |  | October 29, 2015 | 2015 | NxStage Medical, Inc. | Ultrafiltration Volume (UFV) may not decrease during treatment-software error. The UF pump may continue to run and remove fluid even after the target UF volume has been removed | Software design | NxStage issued a Field Safety Notice (FSN), dated October 29, 2015, to clinic and home patients with a NX1000-3 or NX1000-4 cycler. For patients performing short daily hemodialysis, the FSN included instructions for steps to be taken so that the cycler and treatments are not affected by the software error. Patients performing nocturnal hemodialysis were instructed to contact their Center for a replacement NX1000-1 Cycler. A software update will be released to correct the error. Customers should complete and return the reply form. Contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243) with questions or comments. |  | 2134 | 2,134 units | Worldwide Distribution. | Open, Classified |  |  | 2274 | Computer
Z-0342-2016 | 72618 | Yes | NKE |  |  | Cardiovascular | PMA | Medtronic InSync III, model 8042, 8042B and 8042U. Implantable pulse generator pacemaker. |  |  | November 09, 2015 | 2015 | Medtronic Inc. | Medtronic has confirmed an issue with InSync III CRT-pacemaker devices related to long-term battery performance. The root cause is unexpected high battery impedance. | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated November 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Medtronic records indicate you are following one or more patients with this device. Please see the enclosed Physician/Patient Detail Report.We regret any difficulties this may cause you and your patients. We will continue to monitor device performance and provide regular updates in our product performance report available at wwwp.medtronic.com/productperformance/. Medtronic Patient Services is available to assist patients at 800-551-5544.Medtronic will offer a supplemental device warranty. Contact your sales representative for terms and conditions. If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 96787 | 96,787 | Worldwide Distribution - USA (nationwide) including Puerto Rico and to Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belarus, Belgium, Bermuda, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary , India, Indonesia, Iran, Ireland, Israel, Italy, Japan Jordan, Kazakhstan, Kenya, Kuwait, Korea, Latvia, Luxembourg, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius , Mexico, Morocco, Netherlands, New Zealand, Norway,Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Trinidad, Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Open, Classified |  |  | 96787 | Not_Computer
Z-0343-2016 | 72590 | No | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Copper NiTi Preformed Archwires; Copper NiTi 40 BAF .016X.022 LO Pk10, Part No. 211-0952. Intended to provide tooth movement. |  |  | November 04, 2015 | 2015 | Ormco/Sybronendo | Ormco Corporation is recalling Copper NiTi Preformed Archwires because the product transformation temperature as listed on the labeling is incorrect. | Employee error | A recall letter dated 11/5/15 was sent to all their customers who purchased the Copper NiTi Preformed Archwires because it has been brought to Ormco's attention that the product transformation temperature as listed on the labeling is incorrect. The wire is labeled as having 40 degrees Celsius transformation temperature, but the actual wire has a transformation temperature of 29 degrees Celsius. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to complete and fax back the enclosed acknowledgement form. |  | 161 | 161 units | Distributed US (nationwide) and the countries of Canada, Japan, Mexico, Russia, and Peru. | Terminated | January 19, 2016 | 76 | 161 | Not_Computer
Z-0345-2016 | 72454 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | HYDRO-TEMP(R) Neonatal Temperature Skin Sensor, REF HNICU-32, 50 units per box, NON STERILE, Rx only. |  |  | October 14, 2015 | 2015 | DeRoyal Industries Inc | DeRoyal manufactured four lots of HNICU-32, Hydro-Temp(R) Neonatal Temperature Skin Sensor incorrectly. The pins contained in the probe were soldered incorrectly. When probe is connected to the proper port, this defect results in the inability to read or display temperatures on the infant warming system. | Process control | DeRoyal issued their recall on 10/14/2015 to the end-user level by sending overnight letters to its direct accounts and hospitals. The consignees were instructed to destroy and discard the units. In the case of further distribution, the firm requested that the consignee either notify its customers of the recall or to submit a customer list for DeRoyal to contact them. |  | 800 | 16 cases of 50 units (800 eaches) | Distributed in the states of VA, AL, FL, and CA. | Terminated | July 01, 2017 | 626 | 800 | Not_Computer
Z-0344-2016 | 72466 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | Two-stage and three-stage I.V. poles are shipped as an optional accessory to a bed or stretcher. The I.V. poles are stationary, height adjustable stands intended to hold infusion liquids, infusion accessories, and/or other medical devices. They may be used on the following parent devices: Model 2030 (EPIC II), model 3005 S3 beds and stretcher. |  |  | October 13, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | It was identified by a customer complaint that some I.V. poles wereretracting/lowering from an extended position to the low height unexpectedly. Aninvestigation was initiated which identified that the I.V. Latch Housing wasnonconforming. potential hazards:" Suspended I.V. pole retracts unexpectedly" Suspended I.V. pole retracts unexpectedly with pump/bags attached | Process control | On 10/13/2015, URGENT MEDICAL DEVICE NOTIFICATION letters were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-5 p.m. (ET). |  | 19876 | 19,876 | Worldwide Distribution-US (nationwide) including DC and Guam, and the countries of New Zealand Canada Chile Hong Kong Colombia Argentina Brazil France Germany Italy Netherlands Spain Switzerland United Kingdom India Japan South Korea Mexico South Africa Poland Singapore. | Open, Classified |  |  | 19876 | Not_Computer
Z-0354-2016 | 70370 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Xper vascular systems R7.6 |  |  | February 03, 2015 | 2015 | Philips Medical Systems, Inc. | Monitor Ceiling Suspension system may fall | Process control | Philips sent an Urgent Field Safety Notice letter dated January 8, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. The actuator of all affected systems will be replaced free of charge. All affected customers will be contacted in order to schedule the implementation of this action. For questions contact your local Philips representatice: Technical Support Line: 1-800-722-9377. |  | 22 | 22 units | Worldwide Distribution - US Nationwide and the countries of: Canada,Australia, Austria, Brazil, Chile, China, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, South Africa, Spain, Switzerland and United Kingdom. | Open, Classified |  |  | 176 | Not_Computer
Z-0359-2016 | 72586 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion Energy Sources, Catalog Numbers: 1) 60-9500-120 &amp; 2) 60-9500-INT. Thermal Ligation and Cutting Energy Source. |  |  | November 05, 2015 | 2015 | ConMed Corporation | Miscalibration of the Energy Source monitoring function, which may incorrectly sense the handpiece temperature as higher than the actual handpiece's temperature. This causes some devices to generate a "Check Seal" message with an audible alarm and to stop the heating cycle. | Process control | On November 4, 2015, ConMed Corporation distributed Urgent Device Recall Notification letters dated November 5, 2015 to their customers via overnight courier service. Customers should stop the use of these devices immediately. The recall notice recipient should contact all organizations within their facility and any other facilities that may have received the affected products. It is imperative that all end users of these devices receive this notice and respond immediately. Customers with questions can call Patricia Cotter, ConMed Recall Coordinator 315-624-3237 (8AM - 5PM EST, Monday through Friday), or fax to 315-624-3225 or email altrus@conmed.com. Customers who have any of the devices subject to recall (listed on the recall notice-attachment I) should immediately contact their sales representative. The sales representative will deliver a replacement unit to your facility and will pick up and return the recalled devices to ConMed Corporation. If you are unable to reach your sales representative, please contact Patricia Cotter at the numbers above. If you do not have any devices to return, customers should complete the Business Reply Form (attachment II), indicating you have no devivces and fax it to 315-624-3225, Attn: Patricia Cotter. |  | 9 | 9 devices | Distributed in the states of CA, and TX and in the country of Mexico. | Open, Classified |  |  | 9 | Not_Computer
Z-0355-2016 | 72596 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | U-SDN 2 Part Ultra Modified Seldinger Needles, NS/37-0387. |  |  | October 27, 2015 | 2015 | Procedure Products, Inc. | The firm had received a complaint alleging that a medical professional had difficulty inserting a guide wire into a needle | Nonconforming Material/Component | The firm sent the Urgent Medical Device Recall letter to their consignees via email on 10/27/15.Customers with questions can call Doug Rowley at 360-693-1832 or send email to doug@procedureproducts.com |  | 1100 | 1100 pieces | Distributed in the states of MT and NY. | Terminated | April 01, 2016 | 157 | 1100 | Not_Computer
Z-0357-2016 | 72564 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | IOL Master 5.5: Software versions 7.5 and 7.7; |  |  | October 27, 2015 | 2015 | Carl Zeiss Meditec AG | IOL Master software versions 7.5 and 7.7 calculation printouts and exported reports can contain the wrong IOL power data. | Software design | Urgent Medical Device Correction letters will be mailed on Friday, November 13, 2015 as the IOL SW update material is now available. Letter informs customers of the issue and provide instructions for identifying the software version and for installing the software update with illustrations. All customers are requested to complete and return the confirmation form to Stericycle as soon as the upgrade is complete or it has been determined that your unit is not running the affected version of the software. |  | 49 | 49 | Nationwide Distribution. | Open, Classified |  |  | 3862 | Computer
Z-0356-2016 | 72539 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ProTrack Microcatheter. Model/Catalog Numbers CIC38-145, CIC35-145. Sterile, single use device intended to be used as a standard infusion wire or as an exchange guide for exchanging guidewires. Coaxial Injectable catheter guide wire. Cardiovascular. |  |  | October 28, 2015 | 2015 | Baylis Medical Corp * | Microcatheter may have circumferential defects (cracks) along its shaft. | Process control | Recall notification letter was sent to affected consignees on 10/28/2015 via courier. The letter contained a fax-back form to be completed by the consignee and sent to the quality department of the recalling firm. |  | 433 | 433 units | Nationwide Distribution. | Open, Classified |  |  | 433 | Not_Computer
Z-0362-2016 | 72498 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Duracon Flat Tibial Wedge Cemented use only Sterile. Tibial Wedge is a component of Total Stabilizer Knee which consists of femoral component, tibial insert, modular offset adaptor, femoral/tibial stem extenders and tibial wedges. Tibial wedges are intended to be used with the Universal tibial baseplate to augment bone loss on the surface of the tibia. |  |  | September 21, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received two customer complaints regarding Duracon Tibial Wedge implants puncturing the packaging's Tyvek lid(s). | Under Investigation by firm | On September 18, 2015, Stryker notified their Branches/Agencies of their action via email. The Branches/Agencies were instructed to quarantine affected devices. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated September 21, 2015 were sent to Branches/Agencies via UPS on 9/21/2015. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated September 21, 2015 were sent to Hospital Risk Managers via UPS on 9/23/2015. The notification instructed consignees on how to identify affected product; the related issue with the affected product; potential hazards; risk mitigations; and actions needed to be taken. Consignees were asked to contact their Stryker Orthopaedics Sales Representative to arrange for return of any recalled product. Customers were asked to complete and return the attached Product Recall Acknowledgement Form within 5 days and either email (strykerortho7984@stericycle.com or Fax (1-866-672-0627) the response form back. A point of contact was provided in case the customer had any questions 201.831.5826. |  | 2577 | 2,577 units | Distributed US (nationwide) and the countries of Australia, Canada, Germany, Hong Kong, Italy, Japan, Malaysia, New Zealand, Netherlands, Poland, Romania, Spain, Sweden and Switzerland. | Terminated | June 08, 2017 | 626 | 2577 | Not_Computer
Z-0363-2016 | 72574 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal sterile custom surgical kits: |  |  | September 10, 2015 | 2015 | DeRoyal Industries Inc | These custom sterile surgical kits contain the component, ECOLAB ORS Fluid Warming and Slush Drape, which was recalled due to the possibility that the sterile packaging may contain small channels in the seal of the pouch. | Other | DeRoyal issued their recall on 09/10/2015 to the end-user level by sending overnight letters to its direct accounts and hospitals. Distributors were asked to put the product in quarantine, place the provided labels on the products and complete. Hospitals were provided with labels to affix to the kits with the direction to discard the recalled component and notice of destruction forms to be completed and returned. |  | 836 | 836 units | Distributed in MO, MA, NC, CT, Kuwait. | Open, Classified |  |  | 836 | Not_Computer
Z-0361-2016 | 72524 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Model 5001102 MPS Delivery Set w/ arrest agent and additive cassettes, heat exchanger &amp;10 ft. delivery tubing; Model 5001102-AS MPS Delivery Set with 6 ft. delivery tubing; Model 7001102 MPS Low Volume Delivery Set for use with cardiopulmonary bypass equipment. |  |  | October 28, 2015 | 2015 | Quest Medical, Inc. | The products have been found to intermittently exhibit a seal failure during use. | Equipment maintenance | The recalling firm issued a press release on 10/28/2015. The press release was posted on FDA website on 10/29/2015. The recalling firm emailed affected consignees on 10/28/15 and sent out recall letters via USPS beginning 10/30/15. |  | 18761 | 18,761 units | Distributed US (nationwide) and the countries of Canada, Japan, Saudi Arabia, Kuwait, South Africa, Switzerland, France, and Italy. | Open, Classified |  |  | 18761 | Not_Computer
Z-0364-2016 | 69090 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Patient Support Material Number: 9890 010 87431 used with: BuckyDiagnost ; Model: 704060/706037 |  |  | August 19, 2014 | 2014 | Philips Medical Systems, Inc. | The hook does not securely hold the footplate in vertical position. Therefore the footplate can fall or swing down and may hurt the operator or patient. | Device Design | The Field Safety Notice FCO70600086 will be mailed on August 19, 2014. Philips Medical Systems will send out certified letters and will have the receipts returned for tracking purposes as well as have a Field Service Engineer visit each site to provide the fix when the Field Change Order is released. |  | 1517 | 1517 systems, not all units have the patient support accessory | Distributed in the states of MN, KY, MA, VA, TX, NC, HI, FL, AZ, CT, and WY and the countries of Argentina, Australia, Austria, Canada, Denmark, Germany, Netherlands, Russia, Saudi Arabia, Singapore, Spain, and Switzerland. | Open, Classified |  |  | 1517 | Not_Computer
Z-0368-2016 | 72591 | No | KHJ | Ear, nose, and throat synthetic polymer material. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | PROLARYN PLUS Injectable Implant 1.0 cc Plus Calcium Hydroxlyapatite (CaHA) Injectable Implant, Model No. 8044M0K5. An injectable implant used to treat vocal fold insufficiency. |  |  | October 26, 2015 | 2015 | Merz North America, Inc. | The expiration date encoded in the UDI barcode on the label on the product carton and the peelable patient identification label is incorrect. | Other | Letters were issued 11/4/2015 to consignees. The letter informs the consignees that they can continue to safely use the affected products by referring to the USE BEFORE date. They letter does not require for any product to be returned. Also, there is no sub-recall required as the product was not distributed to distributors or wholesalers. |  | 9439 | 9,439 units | Distributed US (nationwide) including the Virgin Islands and South Africa. | Terminated | May 08, 2017 | 560 | 17154 | Not_Computer
Z-0366-2016 | 72585 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Radiesse (+) Lidocaine Injectable Implant 1.5cc, P/N: 8063M0K1. Indicated for subdermal implantation for the correction of moderate to severe facial wrinkles and folds. |  |  | October 26, 2015 | 2015 | Merz North America, Inc. | The expiration date encoded in the UDI barcode on the label on the product carton and the peelable patient identification label is incorrect. | Other | Letters dated 11/4/2015 were sent to customers with tracking and delivery confirmation. The letter informs the customer that they may continue to safely use the affected product by referring to the USE BEFORE date. |  | 1616 | 1,616 units | Nationwide Distribution including the states of AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, LA, MD, MA, MI, MN, MO, NV, NH, NJ, NY, NC, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI. | Terminated | May 08, 2017 | 560 | 1616 | Not_Computer
Z-0374-2016 | 70000 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) are impacted by this issue: |  |  | December 19, 2014 | 2014 | Philips Medical Systems, Inc. | Sp02 and/or Non Invasive Blood Pressure (NBP) alarms may become disabled without visual notification | Software Design Change | The firm, Philips Healthcare, issued a "URGENT-Field Safety Notice" on 12/19/14 to inform customers of the issue, identify details of the units affected, instructs on actions to be taken by the customer and identify what action Philips plans to take to remedy the issue. Philips will provide a software correction free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice-To prevent this issue from occurring, customers/users should:1) Disable the Change Bed Label field in Patient Management. This is accomplished by changing the feature from Optional to Hidden2) All patient transfers must be performed using the Transfer key in the Manage Patient application.Remember that the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.If you need any further information or support concerning this issue, please contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. |  | 10013 | 10,013 units | Worldwide Distribution: US (nationwide) and countries: AUSTRALIA, AUSTRIA, BELGIUM, BULGARIA, CANADA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, HUNGARY, ICELAND, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, LUXEMBOURG, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 10013 | Computer
Z-0389-2016 | 72534 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre NovaSil Silicone Single Lumen Embolectomy Catheter 6F, 80 cm, Model 1601-68. Indicated for the removal of arterial emboli and thrombi. |  |  | October 21, 2015 | 2015 | LeMaitre Vascular, Inc. | Manufacturing issue that could cause a pinhole in the formed end of the packaging, which may lead to compromised sterility. | Packaging process control | Consignees were notified of the recall on October 21, 2015 via letter. Consignees were asked to do the following: Identify all catheters with the affected LOT and model numbers in inventory; quarantine all unused catheters from the affected LOTs and record this action in the attached form; send the form via mail, email or fax to Customer Service who will then issue an RMA-number (Return Merchandise Authorization number) for the return shipment and replacement of the products; pass this notice on to all those who need to be aware within your organization and to any organization to which the potentially affected devices have been transferred. For questions concerning this safety notice, please contact Salva Kozin, Quality Assurance Engineer, at 781-221-2266 ext. 183. |  | 4315 | 4315 total | Worldwide Distribution. US nationwide including PUERTO RICO, and Australia, AUSTRIA, BELGIUM, COLOMBIA, COSTA RICA, Czech Republic, DENMARK, EGYPT, FRANCE, GERMANY, ICELAND, Italy, JAPAN, JORDAN, Latvia, LUXEMBOURG, MOROCCO, NETHERLANDS, Norway, Oman, PERU, PORTUGAL, Poland, Slovenia, Spain, Switzerland, Sweden, THAILAND, United Kingdom, United Arab Emirates, and VIETNAM. | Open, Classified |  |  | 4315 | Not_Computer
Z-0372-2016 | 72655 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew PERI-LOC VOLAR DISTAL RADIUS PLATE STANDARD 3 HOLE SHORT, LEFT - 62 MM, STERILE R, Rx only, REF Numbers: 71863110 Bone fixation plate |  |  | October 20, 2015 | 2015 | Smith &amp; Nephew, Inc. | The devices were packaged and laser marked anatomically incorrect. The Left plates were laser marked and packaged as Right plates and the Right plates were laser marked and packaged as Left plates. | Employee error | Distributors, notified by Fed- Ex on 10/20/2015, were asked to quarantine the product and return it. |  | 9 | 9 units | International onlyGermany, Japan, Russia | Open, Classified |  |  | 9 | Not_Computer
Z-0371-2016 | 72198 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Valeant Pharmaceuticals North America Model No. Spag-2 |  |  | September 14, 2015 | 2015 | Valeant Pharmaceuticals North America LLC | SPAG-2 6000 Lot J3522070 was released to commerce without being tested for particle size per quality control procedure SPAG-2515. | Under Investigation by firm | Valeant Pharmaceuticals International issued recall letter/return response form dated 9/11/2015 via 2nd day air service Customers were requested to check their inventory for the SPAG-2 device and sgregate it for return shipment. |  | 61 | 61 units | AL, AZ, CA, CO, IL, OH, MO, NC, NY, PA., SC, UT and TX | Terminated | June 27, 2017 | 652 | 61 | Not_Computer
Z-0376-2016 | 72604 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AMS 4674 Custom Pack. For use in a general clinical procedure. |  |  | November 02, 2015 | 2015 | Windstone Medical Packaging, Inc. | AMS 4674 Custom Pack is recalled because this kit containing U-SDN 2 Part Ultra Modified Seldinger Needle which is recalled by supplier. | Process control | Aligned Medical Solutions (AMS) sent the Procedure Products -Urgent Medical Device Recall Letter, dated November 02, 2015 to consignees.This letter listed the items affected by recall as: AMS 4674 Custom Pack and AMS 4675 Angio Pack. Both direct account and sub account are informed that the part # NS/37 -0387 01 U-SDN 2 Part Ultra Modified Seldinger Needle (introducer needle) is recalled because some was unable to accept guide wires up to .038" as intended. The inner diameter was found to be out of specification in some cases. All other components in these two AMS kits are not affected by this recall.Customers with questions about this recall can call Vicki Davis at 321-527-7714 or send email to vdavis@alignedmedicalsolutions.com. |  | 40 | 40 kits | Distributed in AL and VA. | Open, Classified |  |  | 40 | Not_Computer
Z-0375-2016 | 72555 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Trochanteric Fixation Nail- Advanced System-11MM/130 DEG TI CANN TFNA 420MM/Right Sterile (The TFN-ADVANCED Proximal Femoral Nailing System) The TFN-ADVANCED Proximal Femoral Nailing System is intended for treatment of fractures. Intramedullary fixation rod, orthopedic. |  |  | November 02, 2015 | 2015 | Synthes (USA) Products LLC | TFNA nails may have the locking mechanism too close to the top of the nail. This may 1. Prevent the cannulated connecting screw from fully tightening the insertion handle to the nail resulting in a loose or toggling nail in the handle. Or 2. Cause the bottom of the cannulated connecting screw to tighten against the top of the locking mechanism preventing advancement of the locking mechanism. | Process control | The firm, DePuy Synthes, sent an "URGENT NOTICE:MEDICAL DEVICE FIELD SAFETY NOTIFICATION" letter dated 11/2/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Review the instructions listed in this notification; forward this notice to anyone in your customer's facility that needs to be informed; complete the attached Verification Section and send a copy of the completed Verification Section by: Fax: (866) 629-6186 or Scan/email: Synthes3826@stericycle.com, if the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual, and maintain a copy of this notice.If you have any questions, please call (610)719-5450. |  | 5 | 5 | US Distribution to states of: CA, FL, MI and WI. | Open, Classified |  |  | 5 | Not_Computer
Z-0391-2016 | 72598 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | RBT, Internal Implant, Laser-Lok, 5.0 x 15mm, 5.7 Platform, REF LPBR5015, Rx only, Sterile R, BIOHORIZONS(R), |  |  | October 30, 2015 | 2015 | BioHorizons Implant Systems Inc | An incorrect label reading 10.5mm and not the specified 15mm was placed on the implant inner vial packaging. | Employee error | BioHorizons sent an Urgent Medical Device Field Correcton letter dated Friday, October 30, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to complete the enclosed Customer Reply Form and fax to 205-484-2195 or scan to PDF and e-mail to lpbr5015@biohorizons.com. The sole international consignee was notified Monday, 02 November 2015. Customers with questions were instructed to contact the Customer Care Department at 888-246-8338. For questionsn regarding this recall call 205-986-1205. |  | 32 | 32 units | Worldwide Distribution - US including CA, CO, NY, OH, WA, AZ MI, MD, GA, NC, and Internationally to Iran. | Terminated | February 04, 2016 | 97 | 32 | Not_Computer
Z-0395-2016 | 72436 | No | HCC | Biofeedback device. | Neurology | Neurology | 510(K) Exempt | The "Pacifier Activated Lullaby" (PAL) System composed of three major components: 1) the PAL Player Module (aka Player Unit), 2) the PAL Pacifier Sensor, and 3) power supply. The system is designed to encourage and reinforce effective non-nutritive sucking of premature infants. |  |  | October 01, 2015 | 2015 | Powers Medical Devices, LLC | PAL unit has undergone significant technological changes since the previous 510(k) and labeling is not consistent with the cleared indications of the 510(k). | No Marketing Application | Powers Medical Devices, LLC intends to notify customers via certified U.S. Mail and supply the updated Instruction Manual via U.S. Mail or other tracked service. Instructions to customers include: 1) destroy the previous Instruction Manual, and replace it with the new one provided, 2) attach the new product label over the existing labeling without covering up the Model # and serial # of the device, 3) certify the destruction and replacement of the Instruction Manual nd replacement of the product label to the company using the form and self-addressed stamped envelopw provided. |  | 128 | Total of 128 units | Nationwide Distribution. | Open, Classified |  |  | 128 | Not_Computer
Z-0396-2016 | 72549 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | DGDE, DirectGuide Drill Extender. Instrument for use with dental implants. A manufactured device that adds a 6mm extension for a 25mm and 28mm tissue height capability suitable for individual patient needs. The thru hole in the DGDE allows existing surgical drills to be inserted and locked in place. |  |  | October 28, 2015 | 2015 | Implant Direct Sybron Manufacturing, LLC | Implant Direct Sybron Manufacturing, LLC is recalling 26 lots of DGDE DirectGuide Drill Extender because of a possible dropping of the drill. | Unknown/Undetermined by firm | On 10/28/15 customer notification letters were sent.Within the letter the firm states that if customers have any of the affected product listed to contact the firm's Customer Care at 1-888-649-6425. A return label is provided with notification.The firm states that if customers are an authorized Implant Direct Sybron Manufacturing distributor they request that they identify their customers that have been shipped the affected product, and contact them within 48 hours of receiving notification in order to obtain the product.The letter states if any customers or patients have experienced any issues as a result of the affected products to report the incident to the FDA through he MEDWATCH reporting system. |  | 478 | 478 | Worldwide Distribution- US (nationwide) including the states of CA, FL, TX, MO, MI, MA, NV, OR, PA, AZ, IL, VA, OH, GA, and SC, and the country of Canada, Poland, Serbia, Italy, Germany, Cairo, Jordan, El Salvador, Kuwait, Costa Rica, Switzerland, Colombia, United Kingdom, Austria, Netherlands, Chile, and Egypt. | Terminated | April 08, 2016 | 163 | 478 | Not_Computer
Z-0397-2016 | 72711 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | IBS - Osteosynthesis screws for bone fixation Model S25 ST014. |  |  | November 18, 2015 | 2015 | In2Bones, SAS | One batch of osteosynthesis screws were manufactured with the incorrect Torx imprint on the screw head. | Nonconforming Material/Component | IN2Bones - USA sent an Field Safety Corrective Action Medical Device Batch Recall letter dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Recommended actions :We therefore you to follow the instructions here below:1. Identify all I.B.screws with the batches subject of this recall that might still be in your inventory and quarantine them2. Inform and distribute this Recall Notification to all relevant persons within your organization3. To distributors only : Identify all I.B.screws with the batches subject of this recall that were delivered to your customers, and if relevant instruct them to also follow these instructions (identification and quarantine)4. Fill in and return the fax back form enclosed. With this form, you will certify that you have received this Recall Notification and intend to, comply with the recommendations listed.This fax back form will enable ln2Bones to conduct effectiveness checks.For any question,plea1se contact our Quality and Regulatory Affairs team at: +33 4 72 29 26 26/ +33 7 6188 44 30 or by email:qualite@in2bones.com. |  | 20 | 20 units | US Distribution to the state of :TN | Completed |  |  | 20 | Not_Computer
Z-0398-2016 | 72667 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Thin-Flex Venous Cannula, Model Number:TF292902A |  |  | November 04, 2015 | 2015 | Edwards Lifesciences, LLC | Edwards Lifesciences initiated a field action for the Thin-Flex Venous Cannula (Model TF292902A) because the connector size is incorrectly described in the Edwards product guide as 3/8 inch. The actual and correct connector size is 1/2 inch. | Under Investigation by firm | Edwards Lifesciences sent an Urgent Medical Device Safety Notice dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer Instructions1. Review this field safety notice to understand the potential hazard.2. Meet and review with appropriate clinical staff at your hospital.3. Complete and return the acknowledgement form attached to this letter via fax to (800) 422-9329 within 5 business days of receiving this notice.4. Distribute this notice within your organization or to any organization where the potentially affected devices have been transferred.5. Product return is not required. Actual product is acceptable for use and is not being recalled as part of this field action.Your assistance is appreciated and necessary to ensure that this notice is reviewed and understood.We appreciate your attention, and apologize for the inconvenience caused by this matter. If you have questions that have not been answered by this letter, please call Edwards Customer Service Monday through Friday at (800) 424-3278 from 8:00AM - 4:00PM Pacific Time. |  | 8915 | 8,915 units | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : United Arab Emirates , Belgium, Czech Republic, Spain, France, United Kingdom, Italy, Kuwait, Netherlands, Romania, Saudi Arabia, Sweden, Singapore, Thailand and South Africa. | Open, Classified |  |  | 8915 | Not_Computer
Z-0400-2016 | 69121 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT (computed tomography) scanners, Big Bore configuration, Model number 728243. software version 2.0 The products are intended to provide cross sectional images of the human body and visualization of the internal organs of the body. |  |  | April 25, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | Brilliance Big Bore- An anomaly was identified with software version 2.0 tumor localization application when printing work sheets. this causes information regarding relative locations of marked regions of interest on the patient work sheets to be incorrect and may cause mistreatment. | Software design | Philips Medical Systems sent a product safety notice to customers and corrective software for the issue was released to the field on July 25, 2006. Service Reps will contact each customer to install the new version of software. For questions regarding this recall call 440-483-7600. |  | 54 | 54 units | Worldwide distribution, including US states of CA and NV. | Terminated | December 15, 2015 | 3521 | 54 | Computer
Z-0399-2016 | 70931 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The eCareManager system. Software intended for use in data collection, storage and clinical information management with independent bedside devices, and ancillary systems that are connected either directly or through networks. For use in a hospital environment. eCareManager allows clinicians to generate a PDF display of an order (new, change, discontinue), which is printed at the bedside and/or hospital pharmacy. This printed PDF should be reviewed by the clinician and verified by the hospital pharmacist and entered into the hospital pharmacy system. |  |  | March 27, 2015 | 2015 | Visicu, Inc. | A software defect may cause incorrect medication order change. If the user decides to edit the order prior to electronic signature AND selects the return to Previous Screen function, the previously selected medication is replaced with whatever medication is in the number one spot on the complete medication list. At the stage where the order is signed, the medication displayed may be incorrect. | Software design | Philips Visicu sent an "Field Safety Notification" to all affected customers still using an affected version of eCareManager. (4 Customers) 03/27/2015 The letter identified the problem and circumstances in which the defect can occur; recommend actions the customer should take to prevent risks for patients; and list the actions planned by Philips to correct the problem."Philips Hospital to Home has taken corrective action to address the software defect in all supported versions of eCareManager software. Philips Hospital to Home is working with affected customers to complete the necessary upgrade.Please contact your local Philips Representative for further questions. |  | 43 | 43 | US distribution only. | Open, Classified |  |  | 43 | Computer
Z-0409-2016 | 69055 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, Model number 728231, running version 2.6.0. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | March 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Multiple software defects which could impact the performance or reliability of the system. These impacts include, but are not limited to: error messages which prevent operation, incorrect scan region, couch movement failure, inaccurate CTDI reporting, image artifacts, lost images. There is a risk of misdiagnosis. | Software design | On March 3, 2011, the firm sent Field Service Notices to their customers. |  | 1765 | 1,765 units | Worldwide Distribution. US nationwide, including Puerto Rico, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Equatorial Guinea, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia &amp; Montenegro, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Vietnam &amp; Yemen. | Terminated | December 15, 2015 | 1748 | 1765 | Computer
Z-0404-2016 | 72659 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer, Cuff Tourniquet Dual Hose Single Bladder Maroon Reprocessed by Sterilmed, Sterile EO, Rx only. |  |  | November 13, 2015 | 2015 | STERILMED, INC. | Sterilmed, Inc. is voluntarily recalling all unexpired Reprocessed Tourniquet Cuffs. Sterilmed cannot assure the sterility of all units in the recalled lots. The introduction of non-sterile product into a sterile operative field raises the risk of potential infection. Sterilmed has not received any reports of Adverse Events for this issue. | Process control | Consignees were sent a Sterilmed "Urgent Medical Device Recall: letter dated November 13, 2015. The letter was addressed to Materials Manager, Risk Manager and Operating Room Director. The letter described the product affected, the description of the problem and the Actions Required. Requested consignees to quarantine and return affected product and to complete and return the Business Reply Form. For questions contact Sterilmed's Customer Care Team at 888-541-0078. |  | 174960 | 174,960 | US: Nationwide | Open, Classified |  |  | 174960 | Not_Computer
Z-0403-2016 | 72713 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Olympus ENDOEYE HD II Video Telescopes. Models WA50040A ENDOEYE HD II, 10 mm, 0 degree, WA50042A ENDOEYE HD II 10 mm, 30 degree, W7100603 Repair cable 10 mm 0/30 degree. The ENDOEYE video Telescopes are used with other supporting equipment for endoscopy and endoscopic surgery within the thoracic and abdominal cavities including the female reproductive organs. |  |  | November 17, 2015 | 2015 | Olympus Corporation of the Americas | A damaged temperature sensor at the ENDOEYE tip caused the distal end to become abnormally hot. Excessive heating of the ENDOEYE distal end could result in patient or user injury. | Unknown/Undetermined by firm | Olympus Corporation of Americas sent a letter dated November 25, 2015, to all affected customers to inspect the serial #'s to identify if the affected lots were in inventory and discontinue of the affected lots. Then return all aftected lots for repair and service. Customers were instructed to contact the Customer Care Center at 800-537-5739, Option 2 for instructions on returninng the ENDOEYE for repair and service. The form should be faxed to 484-896-7128. Customers with questions were instructed to call 484-896-5688. |  | 182 | 182 | US Nationwide Distribution including these states: NE, IL, FL, TX, OH, PA, IN, CA, NY, AL, ND, GA, AK, MI, WA, MS, NJ, IA, AZ, and RI. | Open, Classified |  |  | 182 | Not_Computer
Z-0402-2016 | 72364 | No | KGC | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | One Step Button, Low Profile Initial Placement PEG Kit, gastro-enterostomy tube. |  |  | September 11, 2015 | 2015 | Boston Scientific Corp | Boston Scientific Corp. initiated a voluntary recall of specific lots of the EndoVive One Step Button, including the Decompression Tube and Low Profile Replacement Button devices because they may contain decompression tube adapters with a protrusion that could impede its ability to be introduced or removed from the PEG Button. | Under Investigation by firm | Boston Scientific Corp. sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated September 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you identify any product from the affected lot within your inventory, please segregate the affected product immediately and return it to Boston Scientific in accordance with the enclosed recall instructions. You will receive replacements for all recalled product that is returned to Boston Scientific.If you are a distributor, or hospital please note that this recall is to the customer level. Please notify any customer who has received affected product.We are notifying affected worldwide regulatory authorities of this Recall Removal as required.Please read carefully through the enclosed Recall Removal Instructions. Your local Sales Representative can answer any questions that you may have regarding this Recall Removal.We regret any inconvenience that this action may cause, and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. We are committed to continuing to offer products that meet the highest quality standards that you expect from Boston Scientific.Health care professionals and consumers may report serious adverse events (side effects) or product quality problems with the use of this product to Boston Scientific by calling 1-866-868-4004. |  | 22 | 22 units | US Distribution to the states of : IN, NV, CA, PA, CT, TN, NY, WI, and WA, and Internationally to Ireland. | Open, Classified |  |  | 22 | Not_Computer
Z-0401-2016 | 72617 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Revolution CT Scanners. |  |  | October 23, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential issue with routine head scans on your Revolution CT scanner. A potential hazardous situation can occur during a routine head scan with possible artifacts that may emulate pathology between the brain tissue and bone in the head images. No injuries have been reported to date related to this issue. | Software design | The firm, GE Healthcare, sent an"Urgent Medical Device Correction" letter GEHC Ref#25465 dated October 23, 2015 to its Consignees. The letter described the Safety Issues, Safety Instructions, Product Correction, Affected Product Details and Contact Information. The consignees were instructed to follow the instructions in existing CT User Manual and the safety instructions listed in the letter. GE will be delivering an addendum to the User Manual that reiterates these mitigations. If you have any questions or concerns regarding this notification, please contact GEHealthcare Service at 1-800-437-1171 or your local Service Representative. |  | 125 | 125 (US = 37; OUS = 88) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CT, FL, GA, HI, IL, NY, NC, OH, OR, PA, TX, UT, VA, WA, WI; and countries of: AUSTRALIA, BELGIUM BRAZIL, CANADA, CHINA,DENMARK, FRANCE, GERMANY, HONG KONG, INDONESIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, NORWAY, PERU, POLAND, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 125 | Not_Computer
Z-0411-2016 | 72714 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Persona EM Distal Rod, orthopedic manual surgical instrument for use with Persona Knee Prosthesis, item 42-5399-002-00, part of the Persona Tibial Resection System. |  |  | November 24, 2015 | 2015 | Zimmer, Inc. | Complaints that the ratchet teeth of the tube are not latching onto the rod. | Under Investigation by firm | Zimmer Biomet initiated a voluntary recall of the PersonaÂż EM Proximal Tube and PersonaÂż EM Distal Rod due to complaints that the ratchet teeth of the tube are not latching onto the rod via certified mail on 11/24/2015. The firm requests customers quarantine recalled devices, return the customer response forms, and assist in returning the devices to the Zimmer Biomet sales reps. Customers with questions may contact 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 156 | 156 | Distributed in the states of AZ, CA, ID, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, PA, TN, TX, VA, and WA, and the country of Germany. | Terminated | March 24, 2017 | 486 | 156 | Not_Computer
Z-0405-2016 | 71926 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Brilliance 64 Computed Tomography X-ray system |  |  | April 14, 2015 | 2015 | Philips Electronics North America Corporation | During scans with specific protocol steps, the software unexpectedly sets the Axial ( 2 axis ) scan length to a different length than the user set in the plan. | Radiation Control for Health and Safety Act | Philips Healthcare Planned Action: 1. Philips Healthcare will contact customers and initiate a software update to correct the defect.2. Philips Healthcare engineers will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that Philips Healthcare will without charge, remedy the defect or bring the product into compliance.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Philips Healthcare will implement this CAP by April 1, 2016.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. |  | 7 | 7 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Canada, China, Colombia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Latvia, Lebanon, Libya, Lithuania, Malaysia, Martinique, Mayanmar, Mexico, Mozambique, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Republic of Korea, Reunion, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, State of Palestine, Sweden, Switzerland, Taiwan, Thailand, Tanzania, Tunisia, Turkey, United Arab Emirates, Uzbekstan and Vietnam. | Open, Classified |  |  | 322 | Computer
Z-0414-2016 | 72740 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima IGS 320. |  |  | November 13, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue related to a non-recoverable loss of displayed imaging (loss of monitor video) involving Optima CL323i &amp; Optima IGS 320 systems. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#12237 dated November 13, 2015 to customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 45 | 45 | Worldwide Distribution - US Nationwide in the states of: AR, AZ, DC, FL, GA, IL, IN, LA, MI, NC, NE, NJ, NV, NY, OH, SC, TN, TX, WA. and the countries of Algeria, Bangladesh, Bosnia and Herzegovina, Brazil, China, Dominican Republic, Ethiopia, Georgia, Germany, India, Indonesia, Italy, LebanonMartinique , Morocco, Poland, Republic of Moldova, Russia, SpainThailand, Trinidad, Turkey, and Uzbekistan. | Open, Classified |  |  | 153 | Not_Computer
Z-0360-2016 | 72649 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Perseus A500 Anesthesia Workstation, MK06000. |  |  | November 12, 2015 | 2015 | Draeger Medical, Inc. | There is a potential problem with the device main power switch of the Perseus A500 Anesthesia Workstation. The device power switch may spontaneously fail during use, causing one of the following to occur: 1. The workstation has a fresh gas delivery failure and a ventilator failure with respective alarms. 2. The workstation shuts down completely with an audible alarm. | Nonconforming Material/Component | Draeger Medical, Inc. sent Urgent Medical Device Recall Letters, dated November 2015, to consignees on November 10, 2015. The letter identified the affected device and the reason for the recall. The device power switch of the affected workstations must be replace. Customers should refer to the list of affected serial numbers provided. A Drager Service Representative will contact customers to schedule the replacement free of charge, as soon as new switches are available. Until the replacement takes place, the workstation may only be operated under continuous supervision. Increased attention is necessary to notice the potential failure case early. All affected users should be informed. Customers can contact Michael Kelhart at 1-800-543-5047, (press 1 at the prompt, then 2, then 32349), with questions pertaining to the letter. Contact Drager Service Technical Support at 1-800-543-5047 for questions regarding the operation or servicing of the Drager Perseus A500. |  | 34 | 34 | US Distribution -- To the states of WI, PA, OH, ME, IL, and NY. | Terminated | April 05, 2017 | 510 | 34 | Not_Computer
Z-0416-2016 | 72718 | No | NHC | Ventricular catheter. | Neurology | Neurology | 510(k) | VentriClearÂż II Ventricular Drainage Catheter Set. Allows external access and drainage of cerebrospinal fluid (CSF) from the ventricles of the brain. |  |  | November 19, 2015 | 2015 | Cook Inc. | Cook Medical initiated a voluntary recall of specific lot numbers of the VentriClear II Ventricular Drainage Catheter Set due to the device being sterilized twice, which may result in the minocycline and rifampin catheter coating to be less than what is indicated on the labeling. | Under Investigation by firm | Cook Medical issued letter via certified mail on 11/19/2015, for their voluntary recall of the VentriClearÂż II Ventricular Drainage Catheter Set due to being sterilized twice. The Distributor was instructed to forward the recall letter on to its customers, and quarantine any unused recalled devices, and return them to Cook Medical. Customers with questions may contact: Cook Medical Customer Relations 1-800-457-4500 or 1-812-339-2235. Mondaythrough Friday between 7:30a.m. and 5:00p.m. Eastern Daylight Time or email atCustomerRelationsNA@cookmedical.com. |  | 35 | 35 | Distributed in Tennessee, Florida, Michigan, and Pennsylvania and in Saudi Arabia. | Open, Classified |  |  | 35 | Not_Computer
Z-0417-2016 | 72775 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cournand Modified 2-PC 19G, 5.4cm or (Guidewire Introducer Needle) |  |  | November 23, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Guidewire Introducer Needle (0910-192-000) was shipped after the documented expiration date or with minimal time until expiration.There is a potential risk of a soft tissue infection to the patient if device is used after expiration date. | Process control | Stryker sent an Urgent Medical Device Recall Notification letter dated November 24, 2015, to all affected customers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Monday-Friday 8am-5pm (EST)269-389-2921 |  | 120 | 120 units | Nationwide Distribution including PA and SC. | Terminated | September 06, 2016 | 288 | 120 | Not_Computer
Z-0415-2016 | 72682 | No | PIF | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Catalog No. INF0020-A  EnteraLite Infinity Enteral Pump Delivery Set with ENFit Connector System and Transitional Stepped Connector (spiked set); Catalog No. INF0500-A - 500ml EnteraLite Infinity Enteral Pump Delivery Set with ENFit Connector System and Transitional Stepped Connector; Catalog No INF1200-A - 1200ml EnteraLite Infinity Enteral Pump Delivery Set with ENFit Connector System and Transitional Stepped Connector; Catalog No GR1200-A  1200ml Enteral Feeding Delivery Set with ENFit Connector System and Transitional Stepped Connector (gravity set). |  |  | November 20, 2015 | 2015 | MOOG Medical Devices Group | Potential for leaking between the purple ENFit connector and the white transitional stepped connector due to cracking or breaking of the purple connector. | Unknown/Undetermined by firm | MOOG Medical sent an URGENT Field Safety Notice, dated November 20, 2015, to its customer. The letter identified the affected device, provided a description of the problem, and discussed the risk to health. The customer was asked to contact all known distributors and inform them that MOOG Medical will cease production of all its enteral administration sets using ENFit connectors and transition back to the previous revision of the product codes that do not include the ENFit connector (INF0020, INF0500, INF1200, and GR1200). MOOG Medical will produce the previous revision until a solution to the problem is found. All necessary users should be made aware of the notice. Christopher Dodge, Manager, Regulatory Affairs, should be contacted with questions or for additional info at 801-264-1001, ext. 112 or cdodge@moog.com. |  | 12014130 | 12,014,130 sets | Distributed to the state of MN. | Open, Classified |  |  | 12014130 | Not_Computer
Z-0419-2016 | 72731 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI Adult Universal Ventilator Circuit, Duel Heated Limb and Hudson RCI Adult Heated Wire Circuit Kit with Column |  |  | April 17, 2014 | 2014 | Teleflex Medical | Product is being recalled due to complaints that the inspiratory and expiratory tubes are transposed. | Process control | Teleflex sent an "Urgent Medical Device Recall Notification" dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate that you have received products that are subject to this voluntary recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2600 | 2,600 ea | US Distribution to the states of : AL, AZ, CA, GA, IL, KY, LA, MA, MI, MN, NJ, NY, NC, OH, OK, OR, PA, TN, TX and WA. | Terminated | December 16, 2015 | 608 | 2600 | Not_Computer
Z-0393-2016 | 72535 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPodÂż, Insulin Management System (US) |  |  | November 02, 2015 | 2015 | Insulet Corporation | Pod's needle mechanism fails to deploy or there is a delay in the deployment of the needle mechanism. | Process change control | Insulet issued on 11/2/15 an URGENT: Field Safety Notification via Email notification and Federal Express. Letter describes the problem that certain lots of the pod's needle failed to deploy or there is a delay in the deployment of the needle. Customers not responding to the email or Federal Express will receive additional mailing and/or follow up phone calls.Call Customer Care at 1-855-407-3729 if you have any questions regarding this Field Safety Notification. |  |  | 26,230.9 boxes | Worldwide Distribution - US Nationwide and countries of Switzerland, Germany, and Israel. | Terminated | March 07, 2017 | 491 |  | Not_Computer
Z-0418-2016 | 72258 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Medical Systems Duodenoscope, Model Numbers ED-200XU, ED-200XT, ED-250XT, ED-310XU, ED-420XL, ED-410XT, ED-410XU, ED-450XL, and ED-450XT. Medical endoscope for the duodenum and upper G.I. tract. Open elevator wire channel duodenoscopes. |  |  | August 31, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | A precautionary measure because the Duodenoscopes have been redesigned to incorporate a closed elevator wire channel that facilitates reprocessing. | Device Design | FujiFilm mailed a Medical Device Recall Letter and Tracking/Verification Form dated August 31, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to check their inventory for affected product, if found, affected product should be removed from circulation and returned to Fugifilm. Customers were also asked to complete and return the enclosed tracking/verification form to Fujifilm Medical Systems USA, Inc., Endoscopy Division (FMSU-ESD), 10 High Point Drive, Wayne, NJ 07470 or via fax to 973-633-8818. Customers with questions were instructed to contact their Fujifilm sales representative. For questions regarding this recall call 973-686-2482 |  | 79 | 79 units (US) and 985 units (Internationally) | US including CA, CO, FL, IA, IL, KS, LA, MD, MN, MO, NC, TN and WA and Internationally to Asian Pacific, EU, Latin America, and Japan. | Open, Classified |  |  | 79 | Not_Computer
Z-0421-2016 | 72642 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope AVL Video Laryngoscope (AVL), Part Number AVL 3- 0574-0115; AVL 4-0574-0116, AVL 5- 0574-0117. |  |  | November 30, 2015 | 2015 | Verathon, Inc. | Certain models and serial numbers of GlideScope GVL and AVL laryngoscope blades may separate, due to inadequate bonding of the silicone adhesive affixing the halves of the blade. The potential for blade separation may not be readily visible during routine inspection before or after intubation. | Process control | The Urgent Medical Device Recall letter, dated November 30, 2015, had an error on the dates of manufacture for the AVL devices. Once Verathon discovered the incorrect information, the notification letter was revised and it was sent on December 7, 2015 to any customers that received the previous letter between 11/30/15 and 12/2/15. ****************************************************************************************Verathon sent an Urgent Medical Device Recall letter dated November 30, 2015, to all affected consignees. Verathon will, at no cost to you, test your GlideScope GVL and/or AVL blades that fall within the serial number ranges identified to determine whether the blades exhibit the defect that may result in separation of the blade. For all blades that we determine to exhibit this defect, Verathon will replace the blade, free of charge, with a blade of the same model and size. If the blade passes the testing without any evidence of the defect, we will return it to you in its current condition. Please note that this testing will not return blades to their original, factory condition. Verathon will provide you with a loaner blade for each blade in your possession that is subject to this recall, so you will not experience any interruption in use of your GlideScope system.To comply with this recall notice for the affected models and serial numbers of GlideScope GVL and A VL blades, please take the following actions:" Fill out the attached Recall Response Form and return it to Verathon by fax, (425) 883-2896, or email, CSNotifications@verathon.com . Please return the form even if you do not have any blades subject to the recall." Contact Customer Care to arrange for delivery of a loaner blade and return of your recalled blade. Customer Care will issue you an "RMA" number and arrange for shipment of a loaner blade. Do not send us your current blade(s) until you receive the loaner blade(s).Should you have any questions about th |  | 448 | 448 units (189 units in US and 259 units outside US) | Worldwide Distribution - US (nationwide including Puerto Rico) and Internationally to Bahrain, Belgium, Bermuda, Brunei Darussalam, Bulgaria, Canada, China , Czech Republic, Denmark, France, Georgia, Germany, Greece, India , Indonesia, Iran , Israel, Italy , Japan, Kuwait, Lebanon, Malaysia, Morocco, Palestinian Territories, Panama, Peru, Philippines, Republic of Serbia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, and U.A.E. | Open, Classified |  |  | 448 | Not_Computer
Z-0412-2016 | 72686 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Digital Temple Thermometer labeled under: Bestmed, Good Neighbor (Amerisource Bergen), Kroger, Medline, Meijer, Premier Value, Safeway, Life Brand, Target, Top Care, Best Choice, Western Family. Device is packaged in plastic blister with cardboard insert, 2 AAA batteries installed, a Quick Start Guide, and an Instruction Manual booklet. |  |  | November 12, 2015 | 2015 | Bestmed, LLC | An incorrect calibration of the thermometer that causes the device to display inaccurate and frequently lower temperatures than the actual body temperature of the user. | Process control | Consignees were notified verbally about 11/12/15. Written notification is planed and will include instructions on the identification of the Affected Device, potential health risk, method for return, and manner of restitution for the cost of the Affected Device. |  | 305548 | 305,548 | US Nationwide Distribution and the country of Canada. | Open, Classified |  |  | 305548 | Not_Computer
Z-0426-2016 | 72643 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MYLA MASTER DVD 3.2 CLI DL380 |  |  | November 03, 2015 | 2015 | bioMerieux, Inc. | MYLAÂż software connected to a Laboratory Information System (LIS) and a VITEKÂż 2 system has the potential to link a test result to an incorrect patient record with the same specimen ID. | Software design | bioMerieux sent an Urgent Product Correction Notice letter dated November10, 2015 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local bioMÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 626 | 626 systems in total | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and countries of: Foreign accounts: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, Germany, Greece, Hungary, India, Italy, Japan, Korea, Netherlands, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, Angola, Puerto Rico, Republic of Belarus, Republic of Belarus, Benin, Bosnia, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Chad, Democratic Republic of Congo, Costa Rica, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Gabon, Guam, Guatemala, Guatemala, Guinea, Honduras, Indonesia, Iraq, Israel, Ivory Coast, Kazakhstan, Kenya, Kenya, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Malaysia, Malaysia, Mali, Portugal, Myanmar, Managua, Niger, Nigeria, Palestine, Philippines, Philippines, Qatar, Kingdom of Saudi Arabia, Senegal, Serbia, Slovakia, Togo, Tunisia, Ukraine, UAE, Vietnam, Vietnam, Ivory Coast and Zimbabwe. | Open, Classified |  |  | 626 | Computer
Z-0448-2016 | 72646 | No | LRZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tray Set containing multiple instruments and may contain the flexible bone awl, flexible drill and/or flexible screw driver in varying amounts depending on order. Tray set is not sealed. |  |  | November 10, 2015 | 2015 | Aesculap, Inc. | The flexible shaft of the instruments may break during orthopedic spinal surgery. There is a risk of a delay in surgery or that a fragment from the shaft could be left in a patient. | Component design/selection | The firm, Aesculap Implant Systems, sent a "URGENT MEDICAL DEVICE RECALL" letter dated Nov. 23, 2015 to all consignees via certified letter with return receipt. Distributors and sales reps will be notified via e-mail with return receipt. The letter described the product, problem and actions to be taken. The consignees were instructed to review the product recall notification in its entirety and ensure all users your organization and other concerned persons are informed; determine your current inventory; do not destroy any affected product; immediately remove and return the affected instruments-these instruments may be located within Set ST0485; use the U joint screwdriver, ME014R in place of the SJ706R of which is already part of set ST0485; If you require Customer Service assistance please call: Customer Service: 866-229-3002 or Email: SetWorld_Orders.BBMUS_Service@aesculap.com., and complete and return the Product Removal Acknowledgement Form to Aesculap Quality Assurance Department via mail to: Aesculap Inc., 615 Lambert Pointe Drive, Hazelwood, MO 63042; fax to: 610-791-6882 or email to: val.strawn@aesculap.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). A replacement for the SJ607R will be sent to you when one becomes available (ME190R). If you have any questions, contact via email: val.strawn@aesculap.com/Tel: 610-984-9414 or email: Gordon.young@aesculap.com/Tel: 610-984-9043. |  |  | unknown | Worldwide Distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  |  | Not_Computer
Z-0434-2016 | 72773 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 600 Coulter Cellular Analysis System, Software Version 1.1.1.0, Part No. B23858 |  |  | December 01, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the DxH 800 Slidemaker Stainer Software version 3.0.2.0 and DxH 600 software version 1.1.1.0 because the software for the DxH systems allows the creation of multiple orders with the same Specimen identification (ID) but different Patient identification when manually editing pendingorders at the System Manager. | Software design | Beckman Coulter sent an Urgent Medical Device Recall letter dated December 1, 2015, to all affected customers to inform them that Beckman Coulter is recalling the the DxH 800 Slidemaker Stainer Software version 3.0.2.0 and DxH 600 software version 1.1.1.0 because of the possibility of sample misidentificaton. Customers were informed that during internal testing, Beckman Coulter determined that the software for the DxH systems noted above allows the creation of multiple orders with the same Specimen identification (ID) but different Patient identification when manually editing pending orders at the System Manager. The issue does not occur for edits of released results or for test orders requested through host transmission. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions were instructed to contact Customer Support Center at www.beckmancoulter.com/customersupport/support or via phone (800) 526-7694 in US and Canada. Customers outside of US and Canada are instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-961-5321. |  | 666 | 666 units total (327 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Oman, Panama Philippines, Poland, Portugal, Puerto Rico , Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay, and Vietnam. | Open, Classified |  |  | 666 | Computer
Z-0447-2016 | 72725 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System. |  |  | November 24, 2015 | 2015 | Elekta, Inc. | First Databank (FDB) is a third party product incorporated into MOSAIQ. FDB routinely modifies drugs in their database to reflect market and formulary changes. In FDB Framework v2.0, those drugs are not being included in drug-drug, drug-allergy and duplicate drug therapy interaction checking. | Software design | Elekta sent an Urgent Important Field Safety Notice 371-01-MSQ-004 to all affected customers on 11/24/2015. The notice informs users to the issue, the work around and the solution. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta Office or Representative, as soon as possible but no longer than 30 days from receipt. |  | 211 | 211 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WY, &amp; Puerto Rico. | Open, Classified |  |  | 211 | Not_Computer
Z-0437-2016 | 72352 | No | OYC |  |  | Clinical Chemistry | PMA | MiniMed 620G, 3.0mL mg/dL, Model No. MMT-1710 |  |  | September 25, 2015 | 2015 | Medtronic MiniMed Inc. | The pump drive motors may experience a malfunction which would result in a pump error message alarm notifying you that insulin is no longer being delivered. Possible cause is either lack of greasing or deficient greasing in a ball bearing assembly of the motor. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall Notification dated September 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letters stated that the pump drive motors may experience a malfunction which would result in a pump error message alarm notifying the patient that insulin is no longer being delivered. The letters also inform the customers of what actions are to be taken. Customers with questions are instructed to contact Medtronic MiniMed's customer support team. For further questions please call (818) 576-4700. |  | 53 | 53 units | Internationally Distribution only to the countries of : Australia, Austria, Belgium, Chile, Czech Republic, Denmark, Finland, France, Germany, Greece, Iceland, India, Ireland, Italy, Japan, Kenya, Luxembourg, Namibia, Netherlands, Norway, Slovenia, South Africa, Spain, Sweden, Switzerland, United Kingdom. | Terminated | February 05, 2016 | 133 | 1105 | Not_Computer
Z-0446-2016 | 72690 | No | DXO | Catheter tip pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Intravascular (IV) Administration Sets, Model No. 54IV12, 54IV22, 54IV32, MK2450, 54IV42, MK132, MK2156, MK2270, MK2292, MK2300, MK2374, MK2394, MK2508, MK2227 |  |  | November 13, 2015 | 2015 | Edwards Lifesciences, LLC | Edward Lifesciences is recalling Intravascular (IV)Administration Sets due to the requirement of an additional FDA submission to support commercialization (clearance) of the devices. | Under Investigation by firm | An urgent product notification letter dated 11/13/15 was to customers to inform them that during an internal documentation review, the requirement of an additional FDA submission to support commercialization (clearance) of the devices was discovered. Customers are informed that the Edwards Lifesciences Intravascular (IV) Administration sets are temporarily unavailable for sale. Customers are informed that there is no increased risk to the patients' health or safety by using products in their current inventory and there is no impact to the IV sets included in pressure monitoring kits (with TruWave pressure transducer or VAMP system). Customers are informed that if they are in need of a substitute IV set, then to call customer service at (800) 424-3278. Customers instructed to respond to the customer response form by faxing it to (800) 422-9329. |  | 147350 | 147,350 units | US | Open, Classified |  |  | 147350 | Not_Computer
Z-0435-2016 | 72276 | applicator, ent drug | LRD | Ear, nose, and throat drug administration device. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | AbsorbENT, Sterile Compressed Absorbent Sponge, Nasal Pack, 8.9cm(L) x 3.0cm (W) x 1.5cm (H), RH-7413-10, Sterilized by Gamma Radiation, Rx only. |  |  | September 03, 2015 | 2015 | Summit Medical, Inc. | Summit Medical is recalling Doyle P/F Nasal Pack w/Airway, string 8 cm x 3 cm x 1.5 cm, Catalog # RH-7413-10, Lot # 160889 &amp; 161354 because the product does not have the split in the packing. | Process control | Consignees were sent a Summit Medical "Medical Device Recall" letter, dated September 2, 2015. The letter described the problem and the product involved in the recall. It also advised consignees to immediately examine their inventory, quarantine the product, and to complete the attached response form, as well as follow the instructions provided regarding the affected product. For questions, contact Nicole Dove at 651-789-3921 or ndove@summitmedicalusa.com. |  | 9 | 9 boxes (10 /box) | Distributed to CA. | Open, Classified |  |  | 9 | Not_Computer
Z-0444-2016 | 72616 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Unolok Blunt Fill Needle 18G x 1.5" Red, Sterile, Single use, Non-injectable needle, |  |  | November 06, 2015 | 2015 | Myco Medical Supplies Inc | Particulates: Complaints of rubber coring from the closure of injection vials. | Under Investigation by firm | MYCO Medical sent an Advisory Notice dated November 6, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory, if any affected product found MYCO request that the customer return the product to MYCO Medical following the instructions provided on the Materials for Return Form. Customers with questions were instructed to call Customers Service at 919-460-2535/ customerservice&amp;mycomedical.com; or cal the Quality Department at 919-460-2535/ QA@mycomedical.com. |  | 2401419 | 2,401,419 units (100ea/bx, 12bx/cs, 1200ea/cs) | Nationwide Distribution including AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MI, NJ, NY, NV, OK, PA, TX and UT. | Terminated | May 03, 2017 | 544 | 2401419 | Not_Computer
Z-0445-2016 | 72583 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Alaris Pump SmartSite Infusion Set, Model No. 2420-0500. |  |  | November 05, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the Alaris Pump Module SmartSite Infusion Set because the disconnection and leakage may occur during infusion. | Under Investigation by firm | An urgent recall letter dated 11/6/15 will be sent to all customers who purchased the Alaris Pump Module SmartSite Infusion Set because CareFusion has received reports of breakage of the spin collar on the male luer. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed recall response card. Customers with questions related to recall are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PST. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266, customer feedback@carefusion, 24 hrs/day, 7days/wk. |  | 1445180 | 1,445,180 units | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia, and New Zealand. | Open, Classified |  |  | 1445180 | Not_Computer
Z-0443-2016 | 72525 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Gentamicin Antimicrobial Susceptibility Testing Disc (CN10), REF# CT0024B packaged in packs of 5 cartridges - 50 disc/cartridge. |  |  | November 06, 2015 | 2015 | Remel Inc | Antibiotic concentration may be insufficient which may result in incorrect test results. | Nonconforming Material/Component | ThermoFisher Scientific sent an Urgent Medical Device Recall letter dated November 6, 2015, to all affected customers via first class mail. The recall notice identified the problem with the device and asked customers to inspect their stocks for the affected lots. The notice instructed customers to destroy any remaining inventory from the affected lots, contact the recalling firm for replacement, and complete the accompanying Acknowledgement form. Customers with questions were instructed to contact the Technical Services Departmentn at 800-255-6730 (US) or 913-888-0939 (International). For questions regarding this recall call 913-895-4077. |  | 290 | 58/5/50 discs/catridges/packs | Nationwide Distribution including IL, GA, FL, NC, CA, AL, MA, IN, MO, NY, TX, VA, KS, and IA. Product was also distributed to Puerto Rico. | Open, Classified |  |  | 290 | Not_Computer
Z-0449-2016 | 72729 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Label reading in part as: "NexGen System, Complete Knee Solution; Reference Number 5980-37-01; Tibial Component, Precoat, Stemmed Size 3; |  |  | December 02, 2015 | 2015 | Zimmer Manufacturing B.V. | Complaints were received from Thailand reporting that there was no white paper (Tyvek) seal on the inner package of P/N:00-5980-037-01 unit from lot #62460264 | Packaging | An Urgent Medical Device Recall Letter, dated 12/2/15, was sent to the firm's consignees. All distributors will be notified via electronic mail. Hospital risk managers, surgeons, and distributors with product will be notified via courier. |  | 179 | 179 | US Distribution and the country of Thailand. | Open, Classified |  |  | 179 | Not_Computer
Z-0450-2016 | 72741 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S. 4.0mm Cancellous Locking Screws |  |  | November 20, 2015 | 2015 | Biomet, Inc. | These screws have a single lead (instead of a triple lead) thread. The screw may sit proud in the plate, If the screw does not lock into the plate; the screw may back out. If the screw sits proud, soft tissue irritation may occur. If the screw does not lock into the plate and backs out, a revision surgery may be necessary. | Process change control | On 11/20/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors, hospital risk managers, and implanting surgeons via courier. Distributors were instructed to locate and return the affected product, hospitals to locate and discontinue use, and implanting surgeons to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 182965 | 182, 965 | AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, PRCHILE, COSTA RICA, NETHERLANDS, CANADA | Open, Classified |  |  | 182965 | Not_Computer
Z-0451-2016 | 72683 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile X-ray system MobileDaRt Evolution/FDR Go Software |  |  | November 10, 2015 | 2015 | Shimadzu Medical Systems | The firm is recalling the Shimadzu Mobile X-ray system because an image may not transfer to image server properly. | Software design | Shimadzu sent an Urgent Voluntary Medical Device Recall Notice dated November 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter was sent to customers to inform them that Shimadzu is recalling the Mobile X-ray system MobileDaRt Evolution because of reports that an image was not transferred to image server properly. Customers with questions are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 297 | 297 units | US Distribution and Internationally to Canada | Open, Classified |  |  | 297 | Not_Computer
Z-0392-2016 | 72702 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Chariot ST-CCV Guiding Sheath, Sterilized using ethylene oxide: |  |  | November 19, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific is voluntarily recalling its Chariot" Guiding Sheath. To date, Boston Scientific has received fourteen complaints for shaft separation, four for distal shaft separation.The most severe outcome of this failure is embolism of device fragments. | Nonconforming Material/Component | Consignees were sent a Boston Scientific "Urgent Medical Device Recall Removal" letter dated November 19, 2015. The letter was addressed to Materials Manager /Field Action Contact. The letter described the problem and the product involved in the recall. Advised consignees to segregate and return affected product to Boston Scientific. Requested consignees to complete and return the "Recall Removal Reply Verification Tracking Form" to MapleGroveFieldActionCenter@bsci.com or fax to 1-866-213-1806. |  | 7985 | 7985 | Worldwide Distribution - US (nationwide) including DC and Puerto Rico, and in the states of AK, AL, AR. AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV, and WY and the countries of Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hong Kong, Hungary, Ireland, Italy, Luxembourg, Macedonia, Norway, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Switzerland, and Ukraine. | Terminated | February 23, 2017 | 462 | 7985 | Not_Computer
Z-0455-2016 | 72632 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Alaris EtCO2 Module, Model 8300. A capnograph that continuously monitors end tidal carbon dioxide (EtCO2), fractional inspired carbon dioxide (FiCO2), and respiratory rate (RR). |  |  | November 12, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the Alaris EtCO2 module model 8300 because it may have been tested during service using a higher than specified concentration of CO2. The effect to the EtCO2 module tested with a higher concentration of CO2 could lead to an incorrect measurement of CO2. | Other | CareFusion sent an Urgent: Medical Device Recall Notification Important letter dated November 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The Alaris EtCO2 Module ( model 8300) is being recalled because it may have been tested during service using a higher than specified concentration of CO2Customers with questions regarding clinical inquiries, product complaints or clinical troubleshooting are instructed to contact Customer Advocacy at (888) 812-3266 or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris are instructed to contact Technical Support at (888) 812-3229 or by email at DL-US-INF-TechSupport@carefusion.com. |  | 157 | 157 units | US Distribution. | Open, Classified |  |  | 157 | Not_Computer
Z-0452-2016 | 72688 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer UNIVERSAL Power System Loaner and Modular Electric/Battery Double Trigger Handpiece. Rx only Made in Switzerland |  |  | November 10, 2015 | 2015 | Zimmer, Inc. | The firm identified a malfunction which made it possible for a handpiece to start by itself when the power source is connected. | Device Design | Zimmer sent an Urgent Medical Device Correction letter dated November 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Review this notification.2. You will receive copies of notification sent directly to hospital risk managers and surgeons in your territory. As necessary, review and facilitate understanding of this notification by those entities.3. Further identify hospital risk managers and/or surgeons in your territory who should also receive notification of this update, and supply the information to the entities you have identified.4. If after reviewing this notification you have further questions or concerns, please call 330-364-0989 between the hours of 8 a.m. and 5 p.m. EST, Monday through Friday.Zimmer instructed their direct customers that they will upgrade the device with the new electronic during their annual preventative maintenance activities. Until the new upgrade is installed, the firm instructed their customers to follow the Instructions for Use. |  | 87 | 87 | US Distribution to the states of : AR, CA, CO, CT, FL, GA, HI, IL, KS, KY, LA, ME, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, PA, RI, SC, TX, UT, VA, WA, WI, WV and WY. | Terminated | February 19, 2016 | 101 | 176 | Not_Computer
Z-0456-2016 | 72700 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer K2 EDTA (K2E) 5.4mg Plus Blood Collection Tubes. |  |  | November 17, 2015 | 2015 | Becton Dickinson &amp; Company | Some tubes were manufactured with the stopper not fully inserted into the tube. This may result in the stopper popping-off in a small number of instances which could lead to blood exposure and in some cases to blood splatter or spraying. | Under Investigation by firm | BD sent an "Urgent Product Recall notifications/Recall Response Forms dated November 17, 2015 via email and/or UPS Next Day Mail on November 17, 2015 to their customers. The notification identifies the affected product and problem and actions to be taken. Customers are requested to take the following actions (1) Immediately review their inventory and quarantine the product subject to the recall and immediately discontinue the shipment of the affected product (2) complete the enclosed Recall Response Firm and fax it bat to BD at 1.866.873.0312 or emai to bd8554@stericycle.com (3) Return all affected products with the completed Recall Response Form following the instruction on the enclosed packing instruction and (5) if distributed, identify your customers and notify them immediately of this product recall. If the customers had any questions, they were instructed to contact BD at 1.888.628.0732 between 8 AM and 5 PM Monday through Friday. |  | 4835000 | 4,835,000 units | Worldwide Distribution - US Nationwide and the countries of: Canada, Indonesia, Malaysia and South Korea. | Open, Classified |  |  | 4835000 | Not_Computer
Z-0454-2016 | 72654 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Aegis Torque Handle CATALOG NO. : 2871-10-800. A spinal surgical instrument used on the AEGIS Anterior Lumbar Plate System. |  |  | November 20, 2015 | 2015 | DePuy Spine, Inc. | Manufacturing error within the Torque Handle may prohibit some handles from being able to fully assemble with the tightener shaft. The design of the inspection gage does not account for the additional shift length | Device Design | DePuy sent an Urgent Voluntary Product Recall Notification dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were asked to do the following: 1. Review inventory located within your facility and quarantine any products listed in this recall notice. Please contact your DePuy Synthes Spine* Sales Consultant to return any inventory located within your facility. 2. Review, complete, sign and return the attached business reply form to in accordance with the directions on the business reply form within 5 business days of receipt of this notification.3. Forward this notice to anyone in your facility that needs to be informed.4. If any product listed below has been forwarded to another facility, contact that facility and provide them with this letter. 5. Keep a copy of this notice.For further questions please call (508) 828-3774 |  | 16 | 16 devices | US Distribution to the states of : NC, MI, CO, MD, FL, ID, TX, CA and NY. | Terminated | July 03, 2017 | 591 | 16 | Not_Computer
Z-0457-2016 | 72619 | No | DRL | Defibrillator tester. | Cardiovascular | Cardiovascular | 510(k) | DA-2006P - Defibrillator / Pacer Analyzer; |  |  | October 28, 2015 | 2015 | BC Group International Inc | Device is not functioning as intended: Two wires running to the ECG distribution circuit board are reversed resulting in inverted waveforms. | Employee error | The recalling firm issued return authorization numbers to the two customers who submitted complaints regarding the malfunctioning device. The customers sent back the units and the recalling firm replaced their devices with new units. On 10/29/2015, the recalling firm contacted their third customer by phone and informed them that the three devices they had in their possession were miswired which would cause the inverted ECG waveform. The recalling firm issued a return authorization number to their customer for the three devices. The recalling firm told their customer they would repair the devices and return them. |  | 5 | 5 units | US Nationwide Distribution in the states of South Carolina, Florida, and Canada. | Open, Classified |  |  | 5 | Not_Computer
Z-0467-2016 | 72861 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM systems, magnetic resonance diagnostic devices (MRDD). |  |  | December 07, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possible installation error of the quench lines of superconducting magnets, used to vent helium gas in the rare event of a quench. Any restriction, obstruction, or improper installation of the quench line may cause the helium gas to be blown directly into the magnet room or other areas. This could lead to displacement of oxygen, cause cold burns, and/or pose a risk to users, patient or others. | Other | Siemens sent an Customer Advisory letter dated December 7, 2015, certified mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter notifies customers of the potential injuries, which may result of improperly installed quench lines. The letter indicates it is the system's owner who is responsible for installing the quench line per the requirements. If Siemens installed the quench lines, the consignee is directed to contact Siemens, who will inspect the quench lines during next maintenance visit. Siemens will monitor service reports for affected sites. For further questions please call ( 610) 219-6300. |  | 3752 | in US 3752 units | Nationwide including : DC, PR, Virgin Islands and Bermuda | Terminated | June 21, 2017 | 562 | 3752 | Not_Computer
Z-0468-2016 | 72752 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Factor VIII Chromogenic Assay, REF/Catalog No. B4238-40, Siemens Material Number (SMN) 10445729, IVD. --- CLASSIFICATION NAME: Test, Qualitative and Quantitative Factor Deficiency |  |  | November 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Factor VIII Chromogenic Assay Lot 44637 is showing an increased occurrence of irregular curves when used for measurement on Sysmex CA-1500, CA-7000, CS-2000i, CS-2100i, CS-5100 Systems. This could lead, under certain circumstances, to an erroneously increased result with a Factor VIII Chromogenic Assay. | Employee error | Siemens sent an Urgent Medical Device Correction Letter dated November 19, 2015, to all affected customers . The letter identified the product the problem and the action needed to be taken by the customer.Actions to be Taken by Customers: If you are using the affected lot of Factor VIII Chromogenic Assay on a Sysmex CA-1500 or CA-7000 System, please discontinue use. - Due to a different detection principle, the use on BCS and BCS XP is not affected and you can continue using the product on BCS / BCS XP. - Complete and return the Field Correction Effectiveness Check attached to this letter within 30 days. - Review your inventory of this product to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities. --- Corrective Action: The Factor VIII Chromogenic Assay Lot Number 44637 kits were placed on product hold on August 19, 2015, at the Siemens Distribution Centers. For further questions, please call ( 914) 524 - 2495. |  | 675 | Domestic: 52 kits, Foreign: 623 kits | Worldwide Distribution - US Distribution to the state of : Massachusetts and to the countries of : Austria, Czech Republic, France, Germany, Hungary, Italy, Poland, Serbia, Spain and Sweden. | Terminated | August 15, 2017 | 635 | 675 | Not_Computer
Z-0461-2016 | 72738 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Reagent Management System (RMS) |  |  | November 10, 2015 | 2015 | Siemens Healthcare Diagnostics Inc. | Siemens hasidentified RMS refrigeration compressors that are missing the safety cover on the electricaltermination block. The exposed electrical hazard on the electrical termination block is locateddirectly behind the RMS Waste Container. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction (UMDC # DI-16-01.A.US) and Urgent Field Safety Notice (UFSN # DI-16-01.A.OUS) letters dated November 2015 to all affected customers. The letters identified the affected product, problem and actions to be taken. For questions contact your local Siemens technical support representative. |  | 151 | 151 | US Nationwide Distribution: FL, NY, SC, TX, LA, AZ, WI, IA, IL, OH, AL, AK, NH, OK, MD, CA, NM, MS, KY, NJ, MN, GA, CO, MA, SD, TN, MO, PA, MT, SC, NE, KS, VA, AL, NV, WY, OR, PR, MO, OK, GU, HI, AR, ME, VT, CT, NC and the countries of Canada and Mexico. | Open, Classified |  |  | 264 | Not_Computer
Z-0466-2016 | 72811 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Oscillating Flat - TS ST Hub; |  |  | December 08, 2015 | 2015 | Synvasive Technology Inc | Two lots of oscillating saw blades incorrectly identify the blade cutting width as 25mm instead of the correct specification of 19mm. | Labeling False and Misleading | Zimmer sent an Urgent Medical Device Removal letter on December 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed: I. Review this notification.2. You will receive copies of notification sent directly to hospital risk managers and surgeons in your territory. As necessary, review and facilitate understanding of this notification by those entities.3. Further identify hospital risk managers and/or surgeons in your territory who should also receive notification of this update, and supply the information to the entities you have identified.4. If after reviewing this notification you have further questions or concerns, please call 330-364-0989 between the hours of 8 a.m. and 5 p.m. EST, Monday through Friday.Letters to Rick Managers are similar and the actions requested by them are:1. Reviewing this notification and forwarding to all people within your organization that may utilize this product.2. Identifying and quarantining the affected product to prevent further distribution or use.3. Completing the attached Response Form and returning it via e-mail toCorporateQua1ity.PostMarket@zimmerbiomet.com. Returning affected product to:Zimmer SurgicalAttn: QA!RA Dept. - RecallPlease include a copy of the Response Form with the shipment.4. Upon receipt of affected product, Zimmer will credit your account for unused Zimmer Oscillating Flat Blade - TS ST Hub, 75mm x 19mm, xl.OO/l.27mm, lots 43054 and 45104 only. Please return a copy of the completed response form along with your returned product to ensure proper credit.Important: Please distribute this notification to all personnel within your organization who need to be aware. If you have further distributed affected product(s), please provide the customer 's information on the Business Response Form to Zimmer.Letters also included a FAQ sheet regarding the labeling error. |  | 14 | 14 | US Distribution to one location each in CT, MA and CA. | Open, Classified |  |  | 14 | Not_Computer
Z-0462-2016 | 72721 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi XP Asymmetric (N4) Curve, |  |  | November 20, 2015 | 2015 | Boston Scientific Corp | During routine monitoring of complaints, BSC noted that the IntellaTip MiFi XP N4 complaint rate for distal tip bends/kinks was above the anticipated rate in BSC's product risk documentation. | Device Design | Boston Scientific sent an Urgent Medical Device Removal letters to all affected customers dated November 20, 2015 via Fed Ex or other equivalent service. The letter identified the affected products, problem and actions to be taken. Customers were instructed to contact their local Sales Representative to answer any questions. |  | 1556 | 1556 - all models | Worldwide Distribution - US Nationwide in the state of: PA, FL, IN, MI, TN, AL, IL, KY, MI, TX, OH, MO, MA, MN, ID, AL, TN, VA, CA and internationally to Austria, Australia, Belgium, Switzerland, Germany, Great Britain, Greece, Hong Kong, Israel, Iran, Italy, Japan, Kuwait, Latvia, Malaysia, Netherlands, Saudi Arabia, Singapore, Thailand, and South Africa. | Open, Classified |  |  | 1556 | Not_Computer
Z-0469-2016 | 72748 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Hook Blade, Angled. Part number 7058A. Sterile disposable blade for use with Endoscopic Tarsal Tunnel Decompression (ETTD) System. |  |  | November 24, 2015 | 2015 | Instratek, Incorporated | The sterile blade will not fit into the handle provided with the ETTD system. | Employee error | Instratek contacted customers by phone call either by direct contact or voice mail message on November 23, 2015. that a recall was being initiated. Formal letters were sent via UPS to the Customers annoucing the recall on November 24, 2015.Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product,please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.You should contact lnstratek's Customer service for a return authorization number and further instructions for return of unused product. |  | 39 | 39 units | US distribution to the states: TX, UT, NY and CO. | Open, Classified |  |  | 39 | Not_Computer
Z-0442-2016 | 72755 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | EtestÂż PIP/TAZO/CON-4 PTC 256 - Foam of 100 tests US. |  |  | November 19, 2015 | 2015 | Biomerieux | Due to the potential for reporting erroneous results for Piperacillin/Tazobactam; specifically false susceptible results. | Process change control | Consignees (Subsidiaries and Distributors) of bioMÂżrieux, impacted by the identified issue, were notified of the FSCA and corrective action (FSCA notification) electronically on November 24th, 2015. Additionally, the bioMÂżrieux United States subsidiary notified US consignees (distributors/end-users) of the FSCA notification via FedEx on December 2nd, 2015. |  | 2422 | 2422 units | Worldwide Distribution- US Nationwide including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GS, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV and WY, and the countries of Argentina, Australia, Austria, Brazil, Canada, Chile, Colombia, Czech Republic, Estonia, France, Germany, Hungary, India, Israel, Italy, Japan, Kazakhstan, Kenya, Kingdom of Saudi Arabia, Kuwait, Lebanon, Lithuania, Malaysia, Macedonia, Oman, Netherlands, Pakistan, Paraguay, Philippines, Poland, Portugal, Qatar, Romania, Russia, Serbia, Singapore, Slavonia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Uruguay, the United Kingdom, Vietnam and Venezuela. | Terminated | June 19, 2017 | 578 | 11591 | Not_Computer
Z-0479-2016 | 72703 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VAPRÂż TRIPOLAR 90" Degree Suction Electrodes (Model/Catalog No: 225028) |  |  | November 23, 2015 | 2015 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Mitek identified that on VAPR Tripolar 90 Degree Suction Electrode the ablation and coagulation buttons are colored incorrectly. The wiring for each button function was confirmed as correct and the non-conformance impacted button color only | Process control | DePuy sent an Urgent Field Safety Notice dated November 23, 2015, mailed FEDEx to all affected Customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were asked to take the following steps: 1. Review inventory located within your facility and quarantine any products listed in this product recall notification. Please follow the instructions on the Business Reply Form to return any affected inventory located in your facility. 2. Review, complete, sign and return the attached business reply form in accordance with the directions on the business reply form within 5 business days of receipt of this notification. 3. Forward this notice to anyone in your facility that needs to be informed. 4. If any potentially affected product has been forwarded to another facility, contact the firm. 5. Keep a copy of this notice. For further questions please call (508) 880-8100. |  | 127 | 127 | Worldwide Distribution - US Distribution to the states of : AL, CA, CT, FL, KY, MN, MS, ND, NJ, NM and TX., and to the countries of: France, Netherlands and Norway. | Terminated | April 12, 2017 | 506 | 127 | Not_Computer
Z-0475-2016 | 72695 | No | LZJ |  |  | Orthopedic | 510(k) | Total Toe Instrument Kit, Labeled in part: "Instrument Set, 15mm Toe, Dorsal Flange Combi" Part Number: 9000-3002 |  |  | June 09, 2014 | 2014 | Arthrosurface, Inc. | An oversight in the implementation of the design change record resulted in outdated components being shipped to customers. The component did not pass a half-cycle sterilization. | Process change control | The distributors were verbally notified, and replacement Rev F Hex Drivers sent to the locations of the kits. The kits had been moved in preparation for upcoming surgeries, but the Drivers were replaced prior to any additional surgery taking place. No affected product remains in the field. For further questions, please call (508) 520-3003. |  | 3 | 3 units | US Distribution to the states of : PA, WV and CA. | Terminated | February 05, 2016 | 606 | 3 | Not_Computer
Z-0476-2016 | 72726 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Dressol-X, Radiopaque D-502, NDC 11004-502-40, Dental Use Only, Active Ingredients: A.S.A. (Acetylsalicylic Acid) + Eugenol |  |  | November 18, 2015 | 2015 | Rainbow Specialty &amp; Health Products | Rainbow Specialty &amp; Health Products initiated a field action for Dressol-X and Dressit-X because the labeling does not currently list the percentage of drug components. | Labeling design | Consignees have been made aware of the recall by FAX, e-mail and via telephone calls on 11/18/2015. Requested that distributors contact their clients, by telephone initially and then by email with written correspondence requesting them to recall the product from their clients. |  | 2775 | 2775 | Nationwide. No foreign, govt/VA/military consignees. | Open, Classified |  |  | 2925 | Not_Computer
Z-0470-2016 | 72786 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | FB Broth, 10ml |  |  | November 18, 2015 | 2015 | Hardy Diagnostics | The firm is recalling FB Broth Cat. no. K31, Lot no. 15231 because of potential Burkholderia fungorum contamination. | Under Investigation by firm | Hardy Diagnostic notified Customers via phone on November 18, 2015, followed up with written notifications being sent via first class, email, or facsimile.Customers were instructed to destroy the recalled product according to local regulations for nonhazardous waste if uninoculated or as medical waste if inoculated. and to discard any remaining tubes of the recalled lot. Any remaining product on hand the firm stated would be replaced. The firm stated to customers that they would follow-up the phone conversation with written communication and requested a full name, and email or fax number.We sincerely apologize for the inconvenience. Thank you in advance for completing the attached form. Please feel free to contact our Technical Services Department if you have not received your replacements or if you have additional questions about this notification. Technical Services can be reached at (800) 266-2222, option 2 or via email at TechService@HardyDiagnostics.com. |  | 500 | 500 25pk/20) | US. Distribution to the states of : CA, WI, MD and UT.. | Terminated | March 28, 2016 | 131 | 500 | Not_Computer
Z-0436-2016 | 72754 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Optional PS500 Power Supply Unit for the Evita V500 Ventilator and Babylog VN500 Ventilator. |  |  | December 01, 2015 | 2015 | Draeger Medical, Inc. | The firm became aware of cases in which the battery run times of the optional PS 500 power supply unit with the Infinity Workstation Critical Care (Evita Infinity V 500) were unexpectedly short due to the design of the charging algorithm in the current software. Devices used for patient transport will be a priority. | Software design | Draeger Medical, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated November 2015 to all affected consignees. The letter identified the product, problem, and actions to be taken by the customers. The Battery Charging Software will be updated and, if necessary, the batteries replaced in all affected devices Free of Charge. |  | 2422 | 2422 | US Nationwide-including DC and PR, and the states of AK; AR; AZ; CA; CO; CT; DE; FL; GA; HI; IA; ID; IL; IN; KS; LA; MA; MD; ME; MI; MN; MO; MS; MT; NC; ND; NE; NH; NJ; NM; NV; NY; OH; OK; OR; PA; SC; SD; TN; TX; UT; VA; VT; WA; WI; and WV. | Open, Classified |  |  | 2422 | Computer
Z-0471-2016 | 72829 | No | HWQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Black Mamba Suture Passer |  |  | December 02, 2015 | 2015 | Biomet, Inc. | There may be a separationand an insufficient weld between the spring and washer. The weld between the spring and washer may fracture,causing the Mamba Suture Passer instruments jaw to not actuate properly.6 complaints have been received in which 3 have been reported with a delay up to 10 minutes. | Nonconforming Material/Component | On 12/2/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors and hospital risk managers via courier with instructions to locate and return the affected products. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 50 | 50 | FL, NJ, WI, UT, TX, KY, KS, OH, GA, LA, AL, NY, NC, AR, MO, VA, CA, MN, IN, HI | Open, Classified |  |  | 63 | Not_Computer
Z-0480-2016 | 72747 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Synaptive Medical BrightMatter Navigator System (Guide System). Part number SYN-0026. The Guide system is comprised of a cart labelled SYN-0024 and a cart labelled SYN-0025, Operator Cart. A planning and intraoperative guidance system to enable open and percutaneous computer assisted surgery. The system is indicated for medical conditions requiring neurosurgical cranial procedures. It can be used for intra-operative guidance where a reference to a rigid anatomical structure can be identified. |  |  | December 04, 2015 | 2015 | Synaptive Medical, Inc. | Out of tolerance for radio frequency emissions. At the 150-1000MHz frequency, the testing indicated the BrightMatter Navigation system was up to 20dB uV/meter higher than the applicable IEC 60601-1-2:2007 (Ed3.0) standard specification. | Component design/selection | Consignees were sent via courier on 12/4/2015 a Synaptive "Urgent Medical Device Recall" letter dated 30 November 2015. The letter described the problem and the device involved in the recall. Informed consignees that a technical support team member from Synaptive Medical will conduct the field correction. They also requested consignees to complete and return the Recall Response Form. For questions contact Thas Yuwaraj, PhD, .PEng Monday through Friday, 9 AM to 5 PM, Eastern Time at 1-844-462-7246. |  | 8 | 8 | US: FL, WI, IL, LA, WA. | Terminated | March 04, 2016 | 91 | 8 | Not_Computer
Z-0473-2016 | 71883 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vancomycin Screen Agar |  |  | June 29, 2015 | 2015 | Becton Dickinson &amp; Co. | The affected expired lots of product may have exhibited breakthrough growth of vancomycin susceptible Enterococcus facaelis ATCC 29212 near the end of shelf life. The in-date product may also exhibit breakthrough growth of vancomycin susceptible Enterococcus facaelis ATCC 29212 near the end of shelf life. | Under Investigation by firm | Becton Dickinson (BD) sent US customers a Urgent Product Recall letter dated June 2015, via UPS.Customers were requested to return response forms indicating acknowledgement of the recall and quantity of remaining product that was at their facility. The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 10% of non-respondents to verify receipt of the customer letter.Any product returned to BD will be quarantined and discarded.If further assistance is needed, please contact BD Technical Services at 1-800-638-8663. |  | 60840 | 60,840 | Worldwide Distribution - US (nationwide) Distribution to the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI and WV., and to the countries of : Canada, Colombia, India, Mexico and Europe. | Terminated | August 12, 2016 | 410 | 61610 | Not_Computer
Z-0552-2016 | 72848 | Yes | BYE | Positive end expiratory pressure breathing attachment. | Anesthesiology | Anesthesiology | 510(k) | Airways Development LLC WaterPAP Positive Airway Pressure Device Reference #WP-7700. |  |  | September 24, 2015 | 2015 | Airways Development LLC | Airways Development LLC has received a complaint about a canister leaking from WaterPAP Lot 8605A. During a review with their supplier it could not be determined if this was an isolated incident or if the whole lot is impacted. | Nonconforming Material/Component | Airways Development Inc. sent Product Alert notices dated 9/24/2015 to their customers via email. The alert informed customers about the related product issue and how to identify affected product. Airways Development is replacing any affected product left in their customers stock. Customers were also instructed to notify their customers if the product was further distributed. A point of contact was provided if the customers had any related questions. |  | 1464 | 1,464 units | Nationwide | Open, Classified |  |  | 1464 | Not_Computer
Z-0481-2016 | 72742 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis systems with large display from Siemens Medical Solutions USA, Inc. Interventional Fluoroscopic X-ray for angiographic procedures. |  |  | October 30, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Due to improper soldering in a limited number of large display bypass module components of a specific production lot, a loss of video signal could potentially occur. This has not been observed in the field and only sporadic cases have been observed in factory screening. | Software design | Siemens sent an Important Customer Safety Notice dated October 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The issue will be remedied with a hardware update AX025/15/S. The affected bypass module hardware will be exchanged with unaffected production lots to eliminate the problem. For further questions call: (610) 219-6300. |  | 17 | 17 | US Distribution to the states of : IL, WI, MA, FL, AL, TN, ND, OH,TX, CA, NJ and MI. | Terminated | March 28, 2017 | 515 | 17 | Not_Computer
Z-0543-2016 | 72681 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cardinal Health Tiny Toes; 11470-010T Infant heel warmer |  |  | November 20, 2015 | 2015 | Philips Electronics North America Corporation | Heel warmers may cause first and second degree burns. | Labeling design | Cardial Health was notified of the recall via letter on 11/20/2015. Cardinal Health notified all of their customers as of 12/1/2015. Philips instructed Cardinal Health do take the following steps: 1.) Check for and destroy any inventory of Cardinal Health Tiny Toes. 2.) Complete the Distributor Reply Form included in this letter to document this action. Fax to Philips Healthcare at +1 404-855-4900. Please do not return any product to Philips. 3.) Determine which of your customers may have affected product. 4.) Ask your customers to follow the instructions in the section entitled: ACTIONS FOR CARDINAL HEALTH CUSTOMERS/USERS TO TAKE. 5.) As you receive signed Customer Reply Forms from your customers, please forward a running monthly summary to us. Please email the summary to Jennifer Maglio at jennifer.maglio@philips.com. To maintain complete documentation, please retain your customers completed reply forms. Philips may request access to particular reply forms to respond to an audit inquiry.Actions for Cardinal Health Customers to Take: 1.) Discontinue use of Cardinal Health Tiny Toes. 2.) Check for any inventory of Cardinal Health Tiny Toes. 3.) Destroy all products on hand (eg. by activating or puncturing each unit). 4.) Complete and sign the Customer Reply Form provided in this letter or Cardinal Healths version thereof, to notify Cardinal Health of how many Cardinal Health Tiny Toes were destroyed. Submit your signed reply form only after you have completed all steps |  | 7327100 | 7,327,100 individual units | Moberly, MO | Terminated | August 07, 2017 | 626 | 7327100 | Not_Computer
Z-0545-2016 | 72684 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-ray system |  |  | October 29, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Perfusion scan feature may not be available on machines running software versions 4.1.2, 4.1.3, and 4.1.4. Customers previously using the basic axial perfusion software (versions preceding 4.1.2) without buying the helical perfusion license key could not access the basic axial perfusion feature after upgrading to software version 4.1.2. | Software design | On 10/30/2015 the firm sent Customer Information letters to their Consignees. |  | 51 | 51 Units | Worldwide distribution. US nationwide, Australia, Austria, Bangladesh, Belgium, Brazil, Cambodia, Canada, China, Egypt, France, Germany, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Lebanon, Libya, Lithuania, Malaysia, Martinique, Myanmar, Mexico, Mozambique, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 436 | Computer
Z-0482-2016 | 72687 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroCutter Xchange 30 Blue Cartridge, FG-025320; |  |  | November 06, 2015 | 2015 | Cardica, Inc. | Cardica, Inc. has received reports that use of the Blue Cartridges with the MicroCutterXCHANGE 30 Stapler was associated with incomplete firing of staples, potentially resulting in anincomplete transection or anastomosis of tissue during the surgical procedure. | Process design | An Urgent: Medical Device Removal letter, dated November 17, 2015, was sent via Fed Ex. The letter identified the affected device, described the reason for the recall, and provided the instructions to be followed. Customers were asked to check their inventory, and remove and quarantine any affected product. The notice should be forwarded to relevant personnel, and the affected product should be returned along with the completed Reply Form. A credit will be issued for any returned product. A sales representative can be contacted with questions. |  | 7384 | 568 boxes or 6816 cartridges | Worldwide Distribution -- USA, distributed to the states of Alabama, California, Colorado, Florida, Georgia, Iowa, Massachusetts, Maryland, Minnesota, New York, Ohio, Pennsylvania, Rhode Island, Texas, Washington, and Wisconsin ; and, to the countries of Austria, Belgium, France, Germany, Ireland, Italy, Luxembourg, Netherlands and United Kingdom, Switzerland, and Qatar. | Terminated | February 12, 2016 | 98 | 7384 | Not_Computer
Z-0512-2016 | 72651 | No | JXM | Benzodiazepine test system. | Toxicology | Toxicology | 510(k) | ImmuTest Cup 7-panel w/ Adulteration, Item No. IMCA-7P |  |  | November 16, 2015 | 2015 | Ameditech Inc | Ameditechs Drugs of Abuse Tests listed above have shown reduced reactivity for up to two (2) of the nineteen (19) Benzodiazepine (BZO) compounds for which specificity information is provided in the product insert. The particular BZO compounds for which there may be reduced reactivity are Clonazepam and Chlordiazepoxide. | Nonconforming Material/Component | A Medical Device Labeling Correction letter was sent on 11/16/15 to distributors and end users to inform them that Ameditechs Drugs of Abuse Tests have shown reduced reactivity for up to two (2) of the nineteen (19) BZO compounds for which specificity information is provided in the product insert. The letter informs the customers that the particular BZO compounds for which there may be reduced reactivity are Clonazepam and Chlordiazepoxide. Customers with questions are instructed to contact Ameditech at (844) 616-4453 or email at amt.fieldaction@alere.com. |  | 60 | 60 kits | Nationwide Distribution. | Open, Classified |  |  | 161537 | Not_Computer
Z-0551-2016 | 72828 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | Terumo Surflo Winged Infusion Set with Filter and Needle Protection (Surshield), Product Code: SV-S25FL35. |  |  | November 23, 2015 | 2015 | Terumo Medical Corp | Potential improper bonding between the tube and luer for certain lots of Surflo Winged Infusion Set with Filter and Needle Protection (Surshield). | Process control | Terumo N.V. Europe notified their US Importer (Termumo Medical Corporation, Somerset, NJ) of the voluntary recall on 11/23/2015 via phone and email. Terumo Medical Corporation in turn notified their sole consignee on 11/24/2015 by email that included the manufacturers Recall Notification. The notification informed the customer of the reason for the recall; how to identify affected product; to immediately identify and quarantine any available stock; and complete and return the attached Reply Form. |  | 80500 | 80,500 units | California | Terminated | August 01, 2017 | 617 | 80500 | Not_Computer
Z-0550-2016 | 72838 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS is a Picture Archiving and Communication System (PACS) |  |  | November 17, 2015 | 2015 | Siemens Medical Solutions USA, Inc | For the Syngo Workflow SLR System with Software Ver: VB10C: Printouts may be printed in incorrect anatomical size when using syngo Imaging XS filming application in conjunction with a printer not released for anatomical print usage. This may happen due to an invalid combination of printer and the syngo Imaging XS printer configuration file. | Software in the Use Environment | 11/17/2015, CSAN sent to customers with a list of suitable printers for the anatomical printing. |  | 42 | 42 | AL; CA; CO; CT; FL; GA; IA; IL; IN; KS; MD; MI; MO; NC; NE; NJ; NV; NY; OH; PA; TN; TX; WI; | Open, Classified |  |  | 42 | Computer
Z-0554-2016 | 72661 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo with software versions 10.0 LA. |  |  | October 26, 2015 | 2015 | Merge Healthcare, Inc. | The firm, Merge Healthcare, sent out Merge HEMO V10.0 &amp; Merge CARDIO V10.1 software to 3 customers before the product was completely validated in house. | Process control | Consignees were either contacted via phone 9/2015 or in person 7/31/2015. The customer communications involved the need to upgrade the consignees software. Merge Healthcare did not script the conversations because they opted to hold one-on-one management level discussions to carefully explain the importance of moving to the Generally Available (GA) release. |  | 2 | 2 | Distributed in the states of IL, NC, and VT. | Open, Classified |  |  | 3 | Computer
Z-0553-2016 | 72892 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, 3.0T 6 Channel Phased Array Flex Coil, M0050SS. |  |  | December 07, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue with the 3.0T 6 Channel Flex Coil used with the 3T MR750W Surgical Suite Scanners.Coil overheating can occur when the device is used in Mode 2 setup. This could lead to a serious patient thermal injury. There have been no injuries reported as a result of this issue. | Nonconforming Material/Component | Consignees were sent on 12/7/2015 a GE Healthcare "Urgent Medical Device Correction" GEHC Ref#60888 letter dated December 7, 2015. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 90 | 90 (54 US; 36 OUS) | Worldwide Distribution- US Distribution including the states of AZ, CA, CO, IL IN, MI, MN, NY, OH, OK, SD, TX, WI, and the countries of CANADA, CHINA, CZECH REPUBLIC, ITALY, JAPAN, NETHERLANDS, NEW ZEALAND, PORTUGAL, ROMANIA, SOUTH AFRICA, SWITZERLAND, TURKEY. | Open, Classified |  |  | 90 | Not_Computer
Z-0556-2016 | 72760 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Granada Biphasic Broth |  |  | November 23, 2015 | 2015 | Biomerieux Inc | FDA has determined that the product requires 510(k) clearance to continue marketing/distribution in the United States. | No Marketing Application | The firm, Biomerieux, sent an "Urgent Product Removal Notice" dated November 23, 2015 to its consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to confirm that the letter has been distributed and reviewed by all appropriate personnel, discontinue and discard any remaining inventory, store letter with your bioMerieux documentation and complete the attached Acknowledgement Form via fax to: (919) 620-3306 attention of: bioMerieux NA ComOps QA/RA.If you have any questions or concerns, please contact your local bioMÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 24 | 24 cartons | US Distribution to states of: CA, GA, IN, MA, MN, MO, NY, OH and PA. | Terminated | February 24, 2016 | 93 | 24 | Not_Computer
Z-0557-2016 | 72801 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | December 09, 2015 | 2015 | Elekta, Inc. | Incorrect drug dosage due to "Age Limit" and patient weight data item issue. | Software design | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTICE 371-01-MSQ-005" to all customers on 12/9/2015. The notice informs users of the specific product and version numbers affected by both issues, also includes a workaround for both issues. A copy of the Important Field Safety Notice is to be kept with the most current labeling and all personnel working with the product should be made aware of the content of the letter. The customers are instructed to return the acknowledgement form to their local Elekta as soon as possible-Fax: 408-830-8003.If you have any queries about this Notice, please contact your local Elekta office, call 408-830-8000 or toll free: 855-693-5358 or www.elekta.com. |  | 399 | 399 | Worldwide Distribution: US (nationwide) including states of: AK, AZ, AR, CA, CO, CT, DE DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, ,MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, and WY; and countries of: Antiqua &amp; Barbuda, Australia, Austria, Bahama, Bahrain, Canada, Czech Republic, Germany, Hungary, Saudi Arabia, Sweden, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 399 | Not_Computer
Z-0558-2016 | 72779 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 21, Part No. B1017-201 |  |  | December 01, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter has confirmed falsely negative results for Voges-Proskauer (VP) (pale-pink/brown/colorless) with Quality Control (QC) American Type Culture Collection (ATCC) organism Staphylococcus aureus ATCC 29213 in a portion of affected lots of MicroScan Pos Combo and MicroScan Pos Breakpoint Combo panels. The expected result is positive. | Device Design | The firm, Beckman Coulter, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to discontinue use and discard any remaining inventory of lots listed; return the attached form for product replacement; retain your inventory of the other lot numbers of the panel types, as they are not impacted by this issue; discuss content of the letter with your laboratory director/staff, retain the notification as part of your laboratory QS documentation; if you have forwarded any of the affected products to another laboratory provide them a copy of the letter, and complete and return the enclosed Response Form within 10 days via mail to: Beckman Coulter, Inc., 1584 Enterprise Blvd., West Sacramento, CA 95691, Attn: Quality Systems &amp; Compliance or Fax: 916-374-2119 or email: MicrobiologyCustomersHCUS@beckman.com.Customers with questions are instructed to contact their Customer Support Representative at http://www.beckmancoulter.com or by phone at (800) 677-7226 in US and Canada. |  | 284 | 284 units total (96 units in US) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Austria, Belgium, Brazil, Canada, Germany, Italy, Japan, Kazakhstan, Kuwait, Mexico, Myanmar, Netherlands, Norway, Panama, Poland, Portugal, Romania, Russian Federation, Spain, Taiwan, Thailand and United Kingdom. | Open, Classified |  |  | 284 | Not_Computer
Z-0567-2016 | 72868 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer PLUS K2EDTA Tubes (Plastic). 13 x 75 mm x |  |  | December 03, 2015 | 2015 | Becton Dickinson &amp; Company | Some of the Blood Collection Tubes in these lots were manufactured with insufficient EDTA additive. An insufficient EDTA additive may lead to erroneous platelet results, specimen recollection, or a delay of treatment. | Process change control | The firm, BD, issued an "URGENT PRODUCT RECALL" notifications/Recall Response Forms to their customers dated 12/3/2015 via email and/or UPS 2nd Day Mail. The notification identified the reason for the recall; how to identify affected product; and actions to be taken. The customers were instructed customers to immediately review their inventory; complete the enclosed Recall Response Form and fax (1-866-873-0312 it to BD or email it to bd4964@stericycle.com; and return all affected products with the completed Recall Response Form following instructions on the enclosed packing instruction, even if you do not have any affected lot in your inventory. If customers have any questions or require assistance with the return of the recalled product and/or availability of replacement product, they were instructed to contact 1-888-628-0732 between 8AM and 5 PM ET Monday through Friday. |  | 29815600 | 29,815,600 units in total | Worldwide Distribution: US (Nationwide) and the countries of: European Union, Peoples Republic of China, Japan and Republic of Korea. | Open, Classified |  |  | 29815600 | Not_Computer
Z-0566-2016 | 72712 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Valve XS Atrium Retractor |  |  | November 17, 2015 | 2015 | Aesculap, Inc. | AAG received a customer complaint that corrosion had been found by a hospitals central sterile service department on the ball joint of an FC424R blade. Investigation revealed that the corrosion was due to incorrect material used in the manufacturing of the disk spring of the ball joint by the supplier | Mixed-up of materials/components | The firm, Aesculap, sent an "URGENT MEDICAL DEVICE CORRECTION OR REMOVAL" letter dated 11/17/2015 to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to review the product recall notification in its entirety; do not destroy any affected product; if you have no inventory remaining, notate that on the Product Removal Acknowledgement Form; contact our customer service department for an RGR at 866-229-3002 or email: SetWorld_Orders.BBMUS_Service@aesculap.com and immediately return affected product; complete and return the Product Removal Acknowledgement Form to Aesculap Quality Assurance department by fax to 610-791-6882 or email to val.strawn@Aesculap.com within two(2) weeks of receipt, even if the total inventory in your possession is zero (0)Sales Representative will make an appointment to come and exchange the effected product with an interium solution (FC360-Atrium Lift Retractor Set) that will allow any scheduled or unexpected cases, which would require the use of this product. The newly manufactured product will ship automatically when it's available.Then a visit will be scheduled to retrieve the loaner FC360R sets.If you have any questions, contact Quality Specialist at 610-984-9414 or email to val.strawn@aesculap.com or Product Manager at 610-984-9150 or email to mark.eppedio@aesculap.com. |  | 41 | 41 | US Distribution to state of: NY. | Open, Classified |  |  | 41 | Not_Computer
Z-0563-2016 | 72542 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ceiling Mounted Surgical Light System TruLight 3000/5000 and Ceiling Mounted Surgical Light System iLED3/iLED5 |  |  | November 02, 2015 | 2015 | Trumpf Medical Systems, Inc. | The firm has received 15 complaints over ten years related to the connection of the spring arm to the top horizontal arm of the lighting system that while positioning the light head enabled the spring arm and light head to fall. | Other | TRUMPF sent an " Urgent Field Safety Information" letter dated December 8, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by the customer:1. Perform an inspection of the lighting systems (see Appendix 1 for instructions). If any of the symptoms listed below are observed, immediately remove the light from operation and contact your Trumpf Medical Service Representative or designee.Signs that indicate a possible fall of the spring arm:The movement of the light and spring arm is difficult for the user,The lights have stopped being operational, and/orA gap is noticed between the spring arm and the central axis (see the attached Appendix 1 for an example).2. Follow the User Manual: No unauthorized servicing of the spring arm is permitted. Snap rings are one use only and replacements must be purchased from your authorized Trumpf Medical partner. If you have any questions concerning this request or the procedure as outlined, please contact TRUMPF Medical Systems, Inc. at 888-474-9359 |  | 20713 | 20,713 | USA (nationwide) Distribution to the states of : AL, AK, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KY, ME, MA, MN, MO, NE, NV, NJ, NM, NY, NC, ND, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI and WY., and Internationally to Canada. | Open, Classified |  |  | 20713 | Not_Computer
Z-0564-2016 | 72658 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Radiographic Transparency for Certain and External Connection Tapered Implants. Images that assist with measurements for dental implants. |  |  | November 16, 2015 | 2015 | Biomet 3i, LLC | Images contained within radiographic transparency ART 1036 are undersized. | Software design (manufacturing process) | BIOMET3i sent an" Urgent Medical Device Recall letter " dated November, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Responsibilities:1. Please review this notice and check your inventory for the affected item by comparing any on-hand ART 1036 Transparencies to a standard 8.5" x 11" sheet of paper. The non-conforming ART 1036 is approximately 7.8" x 10.2".2. Immediately quarantine and remove all affected items from service.3. Complete the attached Business Reply Form and either: a. fax it to 561-514-6316 orb. email it to postmarket@biomet.com orc. return it along with the affected product using the included shipping label4. Return affected product to BIOMET 3i.For assistance or other questions that you may have relative to this notice, please contact BIOMET 3i at 1- 800-342-5454 or 1-561-776-6700. Team members are available to assist you 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 686 | 686 | Worldwide Distribution - US Distribution including Puerto Rico., and to the states of : CA, CO, FL, IL, MA, MI, MO, NY, OH, SC, TX, VA, WI, WV and Hawaii., and to the countries of : Australia, Germany, France, Greece, Northern Mariana Islands, Taiwan and Uruguay. | Open, Classified |  |  | 686 | Not_Computer
Z-0565-2016 | 72810 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | RenovoCath RC 120; |  |  | November 25, 2015 | 2015 | RenovoRx, Inc. | Incorrect version of Instructions for Use were shipped to two US customers. | Labeling Change Control | RenovoRx notified customers by phone and by mail of this issue. This notification requests that the customers remove the affected product from their inventory and remove and replace the incorrect Instructions for Use with the correct Instructions for Use provided by RenovoRx. The customers are requested to return the incorrect Instruction for Use to RenovoRx for disposition and are requested to complete a Recall Confirmation Form enclosed to document the completion of the requested Correction. For further questions, please call (650) 284-4433. |  | 11 | 11 units | US Distribution to the states of : CA and FL. | Terminated | February 29, 2016 | 96 | 11 | Computer
Z-0577-2016 | 72727 | joint, knee, external brace | ITQ | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | ORTHO SPEC GRP CONTENDER POST-OP-KNEE BRACE LITE |  |  | November 12, 2015 | 2015 | Corflex | Knee brace could fail to maintain the locked position and a loss of balance or undesired flexion/extension may occur at the affected knee joint, | Nonconforming Material/Component | Corflex notified customers on November 12, 2015, initially via phone. On December 9, 2015, a follow-up letter dated December 8, 2015, was issued via email with read receipt. Customers have been instructed to return all product to Corflexs Manchester, NH manufacturing facility. This removal includes product that is new, in use and no longer in service. Customers have been provided instructions on how to contact our Customer Service department to initiate the return process and generate a Returns Material Authorization. UPS Call Tags are also being issued to facilitate the return of braces at no cost to the customer.Contact Corflex at service@corflex.com or via phone at (800) 426-7353, Monday through Friday, 8:00 am to 4:30 pm, EST |  | 42 | 42 units | Worldwide Distribution - US (nationwide) and Internationally to Canada , Taiwan , Singapore , and England. | Open, Classified |  |  | 1557 | Not_Computer
Z-0572-2016 | 72710 | No | DSB | Impedance plethysmograph. | Cardiovascular | Cardiovascular | 510(k) | CNSystems TASK FORCE MONITOR 3040i, plethysmograph, impedance. Noninvasively measure and display patient's hemodynamic parameters time synchronized using Impedance Cardiography (ICG), Electrocardiography (ECG), oscillometric Blood Pressure (oscBP) and continuous Blood Pressure (contBP). Furthermore Sympathic Tone and Parasympathic Tone are calculated by power spectral analysis of ECG (heart rate variability) and continuous blood pressure (blood pressure variability) as well as the Baroreceptor Reflex Sensitivity (BRRS). For use by medically trained personnel. It has been designed for diagnoses aiding and must not be used for vital sign monitoring of critically ill patients under absence of additional, suitable self-monitoring devices or medical trained personnel. The device measures continuously the subject's hemodynamic parameters without reporting any diagnosis. It is the physician's responsibility to make proper judgment based on these parameters. |  |  | December 01, 2015 | 2015 | Partners In Medicine Llc | Device non-conformity which may lead to an incorrect oscillometric measurement of upper arm blood pressure values in some subjects - mainly in tall and slim persons (overestimation of systolic blood pressure, underestimation of diastolic blood pressure). | Mixed-up of materials/components | CNSystems sent an Urgent Medical Device Recall Safety Notice dated December 1, 2015, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. CNSystems decided to perform a voluntary recall of the affected devices and to return the devices after repair (free of charge). For questions regarding this recall call 314-324-2470. |  | 30 | 28 worldwide 2 monitors distributed in the U.S. | Nationwide Distribution including MI and CA. | Open, Classified |  |  | 30 | Not_Computer
Z-0581-2016 | 72763 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Natus Quantum System with NeuroWorks Software. Catalog /Part Numbers: 013926. |  |  | November 11, 2015 | 2015 | Natus Neurology DBA Excel Tech., Ltd. (XLTEK) | During an internal testing, Natus Medical Incorporated, identified that in Neuro Works 8.1 with Quantum hardware, the incoming signal is displayed with reversed polarity. | Software Manufacturing/Software Deployment | The firm, Natus Neurology, sent an "Urgent Medical Device Correction Notification" letter on 11/11/2015. The letter described the the problem, the product involved in the recall and actions to be taken. Natus will provide further information once an updated software is released. The consignees were instructed to complete and return the confirmation form by mail to Senior Manager Quality and Regulatory Affairs, Excel-Tech, Ltd (XLTEK) Division, Natus Medical Incorporated, 2568 Bristol Circle, Oakville, ON L6H 5S1 or by email to :ots@natus.com.If you have any questions, contact Senior Manager Quality and Regulatory Affairs at 905-287-5055 or email: sanjay.mehta@natus.com. |  | 56 | 56 (43 US, 13 OUS) | Worldwide Distribution: US Distribution including states of: AL, AZ, AR, CO, FL, GA, MD, MA, MN, MS, NJ, NY, OH, SC, TN, TX and UT; and countries of: INDIA, KOREA, SPAIN, SWITZERLAND, THAILAND, UNITED KINGDOM, and CANADA.. | Open, Classified |  |  | 56 | Computer
Z-0573-2016 | 72866 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew GENESIS(R) II/LEGION(TM) PATELLAR COMPONENT, RESURFACING ROUND, 32MM STANDARD, FOR CEMENTED USE ONLY WITH JOURNEY PEG, STERILE EO, Rx only, REF Number: 71934490 |  |  | November 09, 2015 | 2015 | Smith &amp; Nephew, Inc. | One lot of 35 mm patella implants is labeled as 32 mm. These implants have pegs which are spaced proportionally; therefore the mismatch prevents the pegs from properly aligning with pre-drilled holes during implantation. | Process design | The firm, Smith &amp; Nephew Inc, notified the distributors by Fed- Ex on 10/20/2015 and sent an "Urgent - Product Recall 1st Notification - Urgent R-2015-22" letter dated 11/9/2015 to their customers. The letter described the product, problem, and actions to be taken. The customers were instructed to inspect your inventory and locate any unused devices, quarantine them immediately, complete and return the Inventory Return Certification Form by fax to: +1-901-566-7975 or email to: FieldActions@smith-nephew.com, even if you do not have the product, contact Smith &amp; Nephew's Field Actions Department via email or fax (above) to obtain a return authorization (RA) number, and return any affected product to Smith &amp; Nephew, Attn: Global Field Actions, Global Distribution Center, 3303 E Holmes Road, Memphis, TN 38118.If you have any questions, call Director, Regulatory Affairs 901-399-5520 or email: jason.sells@smith-nephew.com. |  | 20 | 20 units | US Distribution to states of:FL, MO, IN and TX. | Terminated | March 17, 2016 | 129 | 20 | Not_Computer
Z-0478-2016 | 72715 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | The Fuhrman Pleural/Pneumopericardial Drainage Set |  |  | November 17, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling the Fuhrman Pleural &amp; Pneumopericardial Drainage Set due to receiving complaints that it broke off in the pleural cavity while inserting the device into the patient. | Device Design | The firm, Stryker Sustainability Solutions, sent a "URGENT MEDICAL DEVICE RECALL" letter dated 11/17/2015 to its customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to discontinue the use of the product, and complete the Recall Effectiveness Check Form even if no product has been found in inventory. Customers should return the completed Recall Effectiveness Check Form to their local Stryker Sustainability Sales Representative, email to ssspfa@stryker.com, or mail to: Stryker Sustainability Solutions, 1810 West Drake Drive,Tempe, AZ 85283.If the firm indicates that affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Customers will receive credit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433x5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax.If you have any questions, contact Senior Regulatory Affairs Specialist at (888) 888-3433x5506 or email: chelsea.cullen@stryker.com. |  | 34 | 34 devices | US Distribution in the states of: AK, CA, WI, IN, FL, OR, NY, and NC. | Completed |  |  | 34 | Not_Computer
Z-0582-2016 | 72794 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Remote Patient Monitoring System, MDP 2000 Series, Part Numbers: MDP2040-0100, PGM340R3, or PGM340R4. |  |  | November 08, 2015 | 2015 | Vidco, Inc. | Testing at customer site showed unit Remote Patient Monitoring System MDP2040-0100 in a continuous trap condition, not allowing system to reset and reboot. Two customers recently complained of the system freezing and it could only be restarted if the user re-applied power. | Device Design | The firm began notifying consignees on Nov. 8, 2015 and followed up with the Advisory Notice letter, dated November 10, 2015.VIDCO identified the root cause and has corrected software that is currently going through validation. This software will be made available for updating the MDP2040-0100 as soon as it is released. The units may be updated on-site with the software program that will become available on the File Transfer Protocol (FTP) site. If you are unable to perform the procedure yourself, VIDCO will assist you in this effort.The current software version can still be used for monitoring as long as the user confirms the displays are still dynamically writing waves to the screen (i.e. waves will be changing in real-time).Consignees with questions can call (503) 641-1804 or 800-638-4326, or send email to sales@vidcoinc.com |  | 59 | 59 units | US Nationwide distribution in the states of AZ, CA, MD, NM, NJ, and OH. | Open, Classified |  |  | 59 | Computer
Z-0583-2016 | 72814 | No | DZN | Dental hand instrument. | Dental | Dental | 510(K) Exempt | InterGuard 5.5mm Econo Refill, sizes, 4.0mm for short anatomical crowns and 5.5mm, part # 4012 |  |  | December 22, 2015 | 2015 | Ultradent Products, Inc. | Ultradent is recalling InterGuard 5.5mm Econo Refill, lot#s BC2GN and BC4L7 because product was received from the supplier with the floss tether hole not completely punched through and were distributed. | Nonconforming Material/Component | Ultradent Products sent an Urgent Medical Device Correction letter dated December 21, 2015 to affected customers via FedEx. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard remaining inventory. Replacement product will be sent to each of the identified customers upon completion. For questions regarding the recall call 801-553-4366 or 801-553-4220. |  | 157 | 157 | Worldwide Distribution - US Nationwide and the countries of Austria, Canada, Colombia, Denmark, France, Germany, Great Britain, Iceland, and Portugal. | Open, Classified |  |  | 157 | Not_Computer
Z-0584-2016 | 72719 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BIRMINGHAM HIP RESURFACING MODULAR HEAD, REF numbers: |  |  | November 16, 2015 | 2015 | Smith &amp; Nephew, Inc. | The 16 devices distributed in the United States were incorrectly labeled as the monoblock Birmingham Hip Modular Head, with labeling not cleared for marketing in the US. The actual devices were monoblock Modular Femoral Head for hemiarthroplasty use. | Labeling design | The firm sent their field safety notices on 11/16/2015 by letter delivered by Fed Ex (overnight) and email. Non- responding consignees will be contacted within 4 weeks of the initial notice |  | 16 | 16 units in US | Worldwide distribution. US nationwide, France, UAE, Austria, Australia, Belgium, CANADA, Switzerland, Germany, Denmark, Spain, Finland, UK, Hong Kong, ISRAEL, India, Italy, Japan, South Korea, Netherlands, Norway, Poland, PORTUGAL, Sweden, SINGAPORE, Thailand, Turkey, and SOUTH AFRICA. | Open, Classified |  |  | 16 | Not_Computer
Z-0590-2016 | 72750 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGEN System REF 5980-37-01 Tibial Component Size Stemmed 3 Precoat, and REF 5986-37-01 Nonaugmentable Tibial Component Option CR/PS/LPS Size Stemmed 3. Tibial component of implanted knee prosthesis. |  |  | December 03, 2015 | 2015 | Zimmer Manufacturing B.V. | A product complaint was received stating that a carton labeled as a Precoat Stemmed Tibia Size 3 was opened for use and the contents were a non-augmented Stemmed Tibia Plate size 3. | Labeling mix-ups | Zimmer Biomet sent an "Urgent Medical Device Recall-Lot Specific" letter to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to do the following:Review the notification and ensure affected personnel are aware of the contents.2. Locate all affected product identified above and quarantine them immediately.3. Carry out a physical count of all affected product in your territory and complete the Inventory Return Certification Form (Attachment 1). Email a completed copy of Attachment 1 to corporatequality.postmarket@zimmer.com.4. Return the recalled product along with the completed Inventory Return Certification Form (Attachment 1). Clearly mark the outside carton of each product return shipment made as Recall. 5. Please notify Zimmer Biomet of any hospitals to which you have further distributed the affected product. Supply the information for any hospitals that you have identified, as well as the affected surgeons, using the provided spreadsheet template. The template will be emailed to you for completion. Please return to corporate quality.postmarket@zimmerbiomet.com.6. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 768 | 384 distribution events; 1 unit was returned leaving 383 units . | Worldwide Distribution to Germany, Belgium, Denmark, France, UK, Greece, Ireland, Netherlands, Norway, Russian Fed., Saudi Arabia, Sweden, and Thailand. | Open, Classified |  |  | 768 | Not_Computer
Z-0589-2016 | 72652 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Maquet HCU 30 Coolant Type R134A Filling weight 250 g 120 V 50-60 Hz 1920 W |  |  | November 13, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Failure of the HEATER-COOLER UNIT (HCU 30) Type 2 power supply board. A drop of more than 10% voltage can cause overheating and subsequent failing of the R56 thermistor of the power supply board. Board failure typically occurs when the compressor is switched on (prior to patient use); however, it can occur during patient use. This can shut down the compressor or the entire HCU 30 Type 2 unit. | Device Design | MAQUET sent "Product Recall Urgent Medical Device Field Correction" notices and "Medical Device Field Correction Response Forms" dated 11/13/2015 via UPS to affected customers. The notification identified the issue and how to identify the affected unit. The notification instructed customers to always keep a replacement heater-cooler unit on standby in the event of a unit failure; complete the attached Field Correction Response form (fax to 1.973.629.1518 or email to HCU30@maquet.com; and that they would be contacted by a representative of the Maquet Service Team to schedule an onsite service to determine if the power supply board in their HCU 30 Type 2 unit is affected adn replace the power supply board as necessary. If the customer had any questions they are instructed to contact their Maquet Service Representative or Maquet Technical Support at 1.888.627.8383 (Press option 3, followed by Option 1 and then option 1 again), Monday through Friday between the hours of 8:00 am and 5:00 pm. |  | 100 | 100 Units (distributed in United States) | United States Nationwide Distribution | Open, Classified |  |  | 100 | Not_Computer
Z-0585-2016 | 72774 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SYNTHES; 13.5MM Medullary Reamer Head; 352.135 NON STERILE. Part of the Flexible Reamer Set for orthopedic surgical procedures. |  |  | December 01, 2015 | 2015 | Synthes (USA) Products LLC | The 13.5mm Medullary Reamer head, lot #F-17180, is etched as dimension 13.5mm while the actual dimension is 14mm. | Process design | On Dec 1, 2015, Synthes issued a recall notification letter and requested that all consignees review inventory and contact Synthes for a return authorization number. The verification form is to be completed and sent to Stericycle. |  | 65 | 65 | US nationwide distibution. | Open, Classified |  |  | 65 | Not_Computer
Z-0588-2016 | 72739 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific ACUITY Pro Coronary Sinus Guide Catheter, 9 Fr. Model 8104, 8105, 8106, 8107, 8108, 8109, 8110, 8111, 8112, 8113, 8114, 8115, 8116, 8117, 8118, and 8119. |  |  | November 25, 2015 | 2015 | Boston Scientific Corporation | Potential for separation of the proximal handle from the shaft. | Process control | Boston Scientific personnel began retrieving product from hospital shelves on November 25, 2015. The hospital was given a Medical Device Retrieval letter. The letter provided the reason for product retrieval, identified affected product, asked for product to be segregated to ensure that it will not be used, and to contact Boston Scientific for a return authorization number to ship product back to Boston Scientific. A Certificate of Removal was asked to be completed and emailed to BSCFieldActionCenter@bsci.com Questions can be addressed to Technical Services at 1-800-227-3422. |  | 900 | 900 units | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Australia, Austria, Belgium, Colombia, Cyprus, Czech Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Ireland, Italy, Japan, Malaysia, Netherlands, Poland, Portugal, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, and Thailand. | Terminated | April 08, 2016 | 135 | 900 | Not_Computer
Z-0602-2016 | 72840 | No | PFJ | Hygroscopic Laminaria cervical dilator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gynecare Thermachoice III Uterine Balloon Therapy System Ethicon, Inc. |  |  | December 03, 2015 | 2015 | Ethicon, Inc. | Stability data does not substantiate the labeled two-year shelf life of affected product. | Device Design | Ethicon, Inc. sent "Urgent: Medical Device Recall (REMOVAL)" notifications and Business Reply Forms dated 12/3/2015 on same date via UPS Next Day Mail to its customers. The notification informed the customers of the issue with the product; how to identify affected product and action required. The customers were instructed to examine their inventory immediately to determine if they have affected product on hand and quarantine the affected product(s); remove the affected product and communicate the issue to relevant operating room or materials management personnel, or anyone else in their facility who needs to be informed; if any affected product has been forwarded to another facility, contact that facility to arrange return; complete and return the enclosed Business Reply Form (BRF) confirming receipt of this notice within three (3) business days to Stericycle by fax at 1-866-792-5453 or by email at Ethicon7427@stericycle.com, return BRF even if you do not have affected product. Stericycle is handling returns. If the customer requires any assistance with returning product or have any questions, they were instructed to contact the Customer Support Center at 1-877-ETHICON (1-877-384-4266).Ethicon, Inc. issued an update to their 12/3/2015 letter in a notification dated 12/23/2015. The 12/23/2015 notification clarified the lots of the product subject to their recall previously communicated in their 12/3/2015 letter. There are three (3) lots of product subject to the recall that were inadvertently omitted from the original lot listing. The notification stated that "All THERMACHOICEÂż Catheter lots with an expiration date prior to December 10, 2017 are subject to this recall and are required to be returned. The lot number listing of all product subject to the recall is contained in Attachment 1. The expiration date and lot number can be determined by using the Product Identification Tool attached hereto at Attachment 2." The attachments included the updated pr |  | 96124 | 96,124 units | Worldwide Distribution: US (nationwide) including Puerto Rico and countries of: India, Belgium, Argentina, Mexico, Brazil, Canada, Colombia, Ecuador, Singapore, Aruba and Venezuela. | Terminated | May 30, 2017 | 544 | 96124 | Not_Computer
Z-0601-2016 | 72780 | protector, skin pressure | FMP | Skin pressure protectors. | General Hospital | General Hospital | 510(K) Exempt | 3M Coban Self-Adherent Wrap Shipper carton label identified the product as 3M Coban LF Latex Free Catalog #2081, lot number 2018-10AM. |  |  | December 01, 2015 | 2015 | 3M Company - Health Care Business | Product mislabeled as latex free. Actual Catalog number 1581, lot 2020-10AN. contains latex. This could pose a health risk to users with sensitivity or allergic reaction to latex, which could range from mild to severe. | Process control | Consignees were sent via certified mail on 12/1/2015 a 3M letter dated November 25, 2015 and an emal on 12/5/2015 for the updated 3M "Urgent Medical Device" letter dated November 25, 2015. The letter was addressed to 3M Health Care Distributor. The letter described the problem, device involved in the recall and the action required. Advised consignees to quarantine the product, notify their customers and to complete and return the "Product Recall Form". Distributors were also provided with a "Customer" letter for distribution. For questions contact 3M Customer Helpline 1-800-228-3957. |  | 167 | 167 cases (30 rolls/case) | US nationwide distribution. | Open, Classified |  |  | 167 | Not_Computer
Z-0599-2016 | 72707 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | NIPRO SafeTouch II Safety AVF Needle, 17G x 1 Âż inch. Item code FS+173230BC. |  |  | April 01, 2014 | 2014 | Nipro Medical Corporation | Device has the potential to have torn wings. | Process control | The only one Distributor was notified (2014) by e-mail by the firm of the quality deficiency before the product arrived at their facility. It was immediately returned to Nipro Medical when it arrived. |  | 160000 | 160,000 pieces. | US distribution including TN. | Completed |  |  | 160000 | Not_Computer
Z-0603-2016 | 72825 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | AMSCOÂż Warming Cabinet. Manufactured by: STERIS Corporation |  |  | November 20, 2015 | 2015 | Steris Corporation | The firm discovered that information printed on the quick reference magnet is incorrect; specifically, Fahrenheit to Celsius temperature conversions and the default Warming Cabinet temperature. | Process change control | The firm, Steris, sent an "URGENT VOLUNTARY FIELD CORRECTION NOTICE" letter dated 12/14/2015 to its customers. The letter described the product, problem and actions to be taken. The firm included an updated quick reference magnet for each affected unit with the letter. The customers were instructed to replace your existing magnet with the enclosed magnet. If you have questions regarding this matter, please contact Product Manager, at 440-392-7705 or STERIS Customer Service at 1-800-548-4873. |  | 2319 | 2,319 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Bahrain, Canada, China, Colombia, Costa Rica, France, Hong Kong, India, Indonesia, Israel, Kuwait, Lebanon, Mexico, Panama, Peru, Republic of Korea, Russian Federation, Saudi Arabia, Singapore, Spain, Taiwan, Thailand, Turkey, Uganda, United Arab Emirates, Viet Nam &amp; Virgin Islands. | Open, Classified |  |  | 2319 | Computer
Z-0596-2016 | 72879 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Lumify Diagnostic Ultrasound, Catalogue Number: 795216 |  |  | December 09, 2015 | 2015 | Philips Ultrasound, Inc. | Color Flow direction is displayed incorrectly in Lumify 1.0. The system displayed Color Flow direction does not correctly represent the annotated Color Bar or Velocity Markers. (e.g.) When the Color Bar conveys that Red is to be displayed for color flow toward the Transducer, the system displays Blue for color flow toward the Transducer. This could result in misdiagnosis in some studies. | Software design | Philips began mailing the URGENT - Medical Device CorrectionPhilips Model Lumify Ultrasound System, FSN 79500372, dated 2015 DEC 11, to consignees.Consignees are instructed to do the following:1. Check the version of software on your system using the instructions provided above on How to Identify Affected Products. Philips has corrected this issue in version 1.0.1 and higher.2. Connect the tablet to the internet.3. The system will alert you to download the upgrade from Google Play.4. Follow all on screen to upgrade the software version.If consignees cannot promptly connect to the internet to download the upgraded software, Philips recommends that you do not use Color Flow. In accordance with the subscription terms and conditions, you will be required to reconnect the system to the internet within no more than 30 days. If consignees do not connect their device to the internet after 30 days, they will be unable to use the system.Consignees with questions should contact your local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 21 | 21 units in the US only | distributed in CA, CT, ND, NV, OR, RI, TN, and WA. | Open, Classified |  |  | 21 | Computer
Z-0600-2016 | 72753 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ActiV.A.C. 300ml Canister (with Gel) for ActiV.A.C. Therapy System for the application of V.A.C. Therapy in the acute extended and home care settings. The software-controlled therapy unit applies negative pressure wound therapy. The dressing, to which the therapy unit is connected, enables distribution of the negative pressure across the surface of the wound, while the tubing transfers accumulated fluids to the canister. |  |  | November 25, 2015 | 2015 | KCI USA, INC. | There is potential for small holes, the size of a pin, to develop in the pouches of the ActiV.A.C. Sterile Canisters. | Package design/selection | Customer notification letters were sent out the week of November, 23, 2015. The international distributors will be responsible for notification to their sub-consignees. |  | 12728 | 12,728 units | Japan and Australia | Open, Classified |  |  | 12728 | Not_Computer
Z-0594-2016 | 72446 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | BIOMET 3i, PROVIDE IMPRESSION COPING KITS, Catalog # PRK484, PRK485, PRK654, and PRK655 |  |  | September 21, 2015 | 2015 | Biomet 3i, LLC | Inadequate biocompatibility testing. | Process control | The firm, Biomet 3i, sent an "URGENT DEVICE RECALL NOTICE" letter dated September 16, 2015 to both Domestic and Foreign Customers via traceable courier. The letter described the product, problem and actions to be taken. The customers were instructed to review the notice and check your inventory for the affected units; immediately quarantine and remove all affected products from service; if you have product to return, call Biomet 3i Complaints Department at 1-800-443-8166 or 1-561-776-6700 to obtain a CMP# and return affected product to Biomet 3i Returns, 4555 Riverside Dr., Palm Beach Gardens, FL 34410; complete and return the attached Business Reply Form via fax to: +1-561-514-6316 or email to: postmarket@biomet.com, and maintain a copy of the notice for your records.For assistance or any other questions that you may have, please contact Biomet 3i Complaints Department at 1-800-443-8166 or 1-561-776-6700. The Complaints Department is available 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 30111 | 30,111 devices. | Worldwide Distribution: US (nationwide) and countries of: Canada, Argentina, Austria, Australia, Brazil, Chile, China, Colombia, Costa Rica, Germany, Dominican Republic, Spain, France, Great Britain, Hong Kong, Ireland, Israel, India, Korea, Kuwait, Lithuania, Mexico, Netherlands, Portugal, Saudi Arabia, Sweden, Singapore, Taiwan, Venezuela, and South Africa. | Open, Classified |  |  | 30111 | Not_Computer
Z-0597-2016 | 72751 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software version VD10E for Syngo X-Workplace; Picture archiving and communication system. |  |  | November 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Potential post-processing software issue when using tabcard "4D" on X-Workplace with software version VD10E. A too small measurement in the MPRs of a volume when images are acquired with a CT scanner with a tilted gantry. In a 2x2 layout, the 4D tabcard will show an incorrect, too small length measurement. This could result in selecting a device of the wrong size, which then needs to be exchanged. | Software design | Siemens sent an Important Customer Safety Notice dated November 17, 2015. The Notice identified the product, the problem, and the action to be taken by the customer. The potential issue will be remedied with a software update via Update Instructions AX067/15/S. Siemens service organization will contact customers to arrange a date for the installation of the software update. Customers with questions were advised to call 800-888-7436. |  | 73 | 73 units | Nationwide Distribution including:AL, AZ, CA, CO, FL, IA, ID, IL, IN, MA, MD, MI, MN, MO, MT, NC, NY, OH, OK, PA, SC, TN, TX, VA, and WI. | Open, Classified |  |  | 73 | Computer
Z-0595-2016 | 72851 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MED-RX Single Use REF 54-3680R, 8 FR. x 36" (91CM) Polyurethane Feeding Tube For Pediatric Use with Orange Radiopaque Stripe and Enteral Connector, Made in Canada, Benton, Oakville, Ontario, Canada |  |  | December 11, 2015 | 2015 | Benlan, Inc. | The product may be packaged with a 5 or 6.5FR tube instead of an 8FR size. | Under Investigation by firm | Benlan, Inc. sent a Product Recall Notification letter dated December 11, 2015, to all affected customers. Customers were asked to examine their stock and to discontinue/use of the product and return product. The letter included a response form which was to be returned via fax to 905-825-9600. Customers with questions were instructed to call 905-825-9300. |  | 120 | 120 units | Worldwide Distribution - US including GA &amp; Internationally to Canada. | Open, Classified |  |  | 120 | Not_Computer
Z-0604-2016 | 72869 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam - VMAT - Radiation therapy planning system |  |  | December 18, 2015 | 2015 | Elekta, Inc. | When using the option "Tumor Overlap Fraction" in VMAT planning it has been observed that in rare cases the system does include an organ at risk as target volume. This could result in open MLC, and open jaws in areas away from the target volume. | Software design | Elekta sent an Important Field Safety Notice 806-01-ETP-001 to all affected customers on December 18, 2015. The notice informs users of the issue and that should the issue occur it could result in overdosage of organs at risk outside the intended treatment volume. Users are strongly advised to performproper Quality Assurance for all treatment plans before delivery of the first fraction to the patient. To prevent the issue from occurring, the workaround is to not use the Tumor Overlap Fraction for VMAT planning.and not to use the Tumor Overlap Fraction for VMAT planning. The issue will be solved in Oncentra External Beam version 4.5.2. The notice included an acknowledgement form which is to be returned to Elekta. For questions contact your local Elekta representative. |  | 154 | 154 | Worldwide Distribution - US Nationwide in the states of CA, NY, PA, and the countries of Austria, Belarus, Canada, China, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Netherlands, Norway, Poland, South Korea, Spain, Sweden and United Kingdom. | Open, Classified |  |  | 154 | Not_Computer
Z-0611-2016 | 72570 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Siemens Acute Care DDMR TestPak |  |  | November 11, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | Stratus CS Cartridge Barcode labels may not scan on the Stratus CS systems | Labeling design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on November 11, 2015 both in the United States and internationally for communication with affected customers. This notice informs customers of the issue and provides mitigation instructions. Customers with questions were advised to contact their Siemens Customer Care Center or their local Siemens technical support representative. For questions regarding call 781-269-3000. |  | 2647 | 2647 | Worldwide Distribution - US (nationwide), Algeria, Armenia, Australia, Austria, Bangladesh, Bosnia Herzegovina, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Libya, Malaysia, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., and United Kingdom. | Open, Classified |  |  | 12288 | Not_Computer
Z-0606-2016 | 72787 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter INTERLINK System, Y-Type Solution Set, 78 (2.0 m), 2 Injection Sites, Male Luer Lock Adapter, 1C8581 s, |  |  | November 30, 2015 | 2015 | Baxter Healthcare Corp. | Potential for dark loose particulate matter (400 microns or smaller) on the spike of the spike adapter or solution set. | Under Investigation by firm | Baxter sent an Urgent Product Recall letter by letter on 11/30/2015 to affected customers. The recall was extended to the consumer/user level. The letter instructed customers to locate and remove all affected product, contact Baxter Healthcare Center for Service to arrange for return of affected products and credit and complete the enclosed Baxter Customer Reply Form and return it to Baxter by either fax or scanned email. |  | 9600 | 9600 units | US Nationwide Distribution and the country of Canada | Open, Classified |  |  | 96000 | Not_Computer
Z-0624-2016 | 72701 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NIPRO SET / SLIMLINE BLOOD TUBING SET for Hemodialysis with Transducer Protectors and Priming Set |  |  | March 30, 2015 | 2015 | Nipro Medical Corporation | Blood foaming was found during evaluation of the product. | Device Design | NIPRO contacted their consignees by email 9/25/2013 of the affected product. Customers were instructed to return product back to Nipro Medical Warehouse or discard at point of use. Please contact Director of QA/RA at Nipro Medical Corporation for any additional information or return correspondence to JessicaO@nipromed.com / 305.599.7174x249. |  | 1484 | 1,484 pieces. | US Distribution to states of: TN and NJ. | Terminated | February 04, 2016 | 311 | 1484 | Not_Computer
Z-0622-2016 | 72962 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | UniAxial Pedicle Screw &amp; UniAxial Reduction Pedicle Screw |  |  | November 10, 2015 | 2015 | OrthoPediatrics Corp | Response Spine UniAxial Pedicle &amp; UniAxial Reduction Pedicle Screws may have been manufactured incorrectly, thus preventing the driver from fully engaging into the hexagonal drive feature of the screw. | Process change control | The firm, OrthoPediatrics Corp., sent an "URGENT MEDICAL DEVICE RECALL" letter dated 11/13/2015 to affected customers via email. The letter included a description of the reason for the recall, affected product, consignee responsibilities, and actions to be taken. The customers were instructed to take immediate possession of all the identified devices; quarantine them and immediately return them to OrthoPediatrics, and complete and return the Mandatory Reply Form via fax to (574) 269-3692 or email: MF0x@OrthoPediatrics.com.If you have any questions, please contact VP-Regulatory Affairs (574) 268-6379 or email: mfox@orthopediatrics.com. |  | 102 | 102 | US Distribution to states of: AL, AZ, CA, DE, FL, GA, HI, KY, LA, MI, MO, MN, MS, NC, NJ, NY, OH, OR, TN, and TX. | Completed |  |  | 102 | Not_Computer
Z-0623-2016 | 72789 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Cavity 8 Gauge Spine Wand indicated for resection, ablation, and coagulation of soft tissues and hemostasis of blood vessels in spinal procedures. |  |  | October 26, 2015 | 2015 | ArthroCare Corporation | There is cracking in the tray when the device is snapped into place. The crack is through the full thickness of the tray, compromising the sterile barrier. | Under Investigation by firm | Affected consignees will be notified via letter beginning 11/2/15. |  | 262 | 262 units | Worldwide distribution. US nationwide including OH, MO, AZ, AL, CA, FL, SC, GA, and PA; Spain, France, United Kingdom, Germany, Switzerland, Singapore, Russia, Turkey, and Ecuador. | Open, Classified |  |  | 262 | Not_Computer
Z-0625-2016 | 72832 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zimmer, Virage OCT Spinal Fixation System, Item Number: 07.01738.001, a posterior system intended to achieve an implant and construct of the Occipital-Cervical-Thoracic spine (Occiput-T3). |  |  | November 17, 2015 | 2015 | Zimmer Biomet, Inc. | Zimmer Biomet is initiating a recall of the Virage Occipital Eyelet, as it was found through product testing that the staking of the set screw was out of specification and can interfere with the insertion of the occipital rod when the set screw is staked too low. | Process control | Consignees were sent a Zimmer Biomet "Urgent Medical Device Recall" letter dated November 17, 2015. One letter was addressed to Risk Manager and Surgeons and a separate letter was sent to Distributors, Sales Representatives, and Distributor Operations Mangers. The letters described the problem and the product involved in the recall. The letter described the "Risk" and the responsibilities of consignees and requested consignees to quarantine the product. Risk Managers and Surgeons were requested to complete and return the Certificate of Acknowledgment form. Distributors, Sales Representatives, and Distributor Operations Mangers were requested to complete and return the Inventory Return certification Form.For questions consignees can call the customer call center at 1-866-774-6368 between 8:00 am and 5:00 pm EST. |  | 372 | 372 | US: AL, AZ, CA, CO, FL, GA, ID, MI, MN, MO, NV, NJ, OK, OR, PA, SD, TN, TX, WA, WI.OUS: AUSTRALIA, GERMANY, SWITZERLAND, SWEDEN. | Open, Classified |  |  | 372 | Not_Computer
Z-0626-2016 | 72809 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Reverse Humeral Broach, Size 10 Catalog number 804-06-040, Size 14 Catalog number 804-06-044, and Size 16 Catalog number 804-06-046. Orthopedic surgical instrument for use with Encore Shoulder prosthesis. |  |  | December 08, 2015 | 2015 | Encore Medical, Lp | The trial shell mating feature on the broach is undersized. | Nonconforming Material/Component | The recalling firm notified customers of the recall via letter on 12/8/15. |  | 23 | 23 units | Throughout the US to the states: FL, PA, ID, CA, OH, AZ, TN, NV, WA, MS, GA, MA, KS, TX | Terminated | March 01, 2016 | 84 | 23 | Not_Computer
Z-0586-2016 | 72867 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | MOVES System, an emergency ventilator with suction, oxygen concentrator and multi-parameter patient monitoring capabilities. |  |  | November 19, 2015 | 2015 | Thornhill Research Inc | Batteries found unresponsive after storage, not recognized by main device, and are perpetually enabled. | Component design/selection | Thornhill Research sent a Recall Notification letter dated December 15, 2015 to affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to update their device software.For questions customers were instructed to call 416-597-1325. |  | 83 | 83 units | Worldwide Distribution: US distribution in state of: Virginia and to the country of: Singapore. | Open, Classified |  |  | 261 | Computer
Z-0628-2016 | 72885 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | FMC7 FITC, Catalog No. 340918 Analyte Specific Reagent. |  |  | December 14, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Becton, Dickinson and Company (BD) has determined that the FMC7 FITC (ASR) is exhibiting functional degradation. The functional degradation has the effect of reducing the percent positive from 100% down to approximately 87%. | Under Investigation by firm | BD Biosciences sent an Urgent Product Recall letter dated December 2015, to all affected customers to inform them that Becton, Dickinson and Company (BD) has determined that the FMC7 FITC (ASR) is exhibiting functional degradation. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to return the Tracking/Verification form to BD Biosciences. Customers are instructed to contact BD Customer Support at (877) 232-8995 (prompt 3, 3) in the US. Customers outside of the US are instructed to contact their local BD Biosciences or distributor. |  | 1388 | 1,388 units | Nationwide Distribution | Open, Classified |  |  | 1388 | Not_Computer
Z-0627-2016 | 72817 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medrad Source Administration Sets, Catalog number INT CSS. For use with MEDRAD Intego PET Infusion pump Systems, in nuclear medicine diagnostic procedures. |  |  | November 30, 2015 | 2015 | Bayer Healthcare | The firm's investigation of the Medrad Intego indicates that the inner diameter of the saline tubing is out ofspecification and may result in the system failing to position the dose correctlyduring extraction, potentially leading to a low measurement of the extracted dose orto a higher dose being extracted. The occluded SAS tubing may impair the primingfunction and prompt an RP prime fail | Nonconforming Material/Component | Bayer Healthcare sent an "Urgent Medical Device Field Corrective Action" letter dated November 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please immediately take the following steps.Please review your current inventory. Do not use any Source Administration Sets (SAS) with the impacted batch numbers that are included in the index below.If you have unused affected product, immediately quarantine the product and please call Bayer Customer Service at 1-800-633-7231 to receive a Returned Goods Authorization (RGA) number and return the product immediately to Bayer accordingly.Please complete the response form and fax it to Bayer at 1-412-406-0942 or e-mail to randiproductrecalls@bayer.com, whether or not you haveimpacted product.To provide you with new product, please submit a new product order.Please note that due to limited new inventory, there may be a delay in fulfillment.New batches of SAS are currently being shipped and lead times will be dependent on available supply until inventory returns to normal levels. We appreciate your cooperation and sincerely regret any inconvenience. We are committed to providing effective products and service to support your patient care. If you have questions, please contact our customer support team at 1-800-633-7231. |  | 198598 | 198,598 | Nationwide Distribution to the states of : AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MI, MN, MO, NC, NE, NH, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT and WI. | Open, Classified |  |  | 198598 | Not_Computer
Z-0639-2016 | 72540 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Mercury Medical Neo-Tee T-Piece Resuscitator and T-Piece Circuit. |  |  | October 22, 2015 | 2015 | Mercury Enterprises, Inc. dba Mercury Medical | Incorrect use of T-Piece resuscitators will result in complete occlusion of the patient circuit and consequent failure to ventilate the patient. | Device Design | Mercury Medical sent notification letters dated October 22, 2015, to all affected Domestic Consignees by courier and for the foreign consignees the notifications were sent via e-mail.The letter identified the product, the problem, and the action to be taken by the consignee. Customer Immediate Actions - Immediately notify all personnel subject to using these devices to remove protective cap before connecting a mask. Further this filed safety advisory notice should be read in conjuction with the manufacturer supplied instructions for use.In addition, if you further distributed this product, please identify your customers and notify them at once of this product correction. Notificatin to your customers should include a copy of this letter. For questions contact Mercury Medical at 787-573-0088, 727-573-4966 fax or via e-mail at regulatoryaffairs@mercurymed.com |  | 753245 | 753,245 devices | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Bahrain, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Guam, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Netherlands, Norway, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirate, and United Kingdom.. | Open, Classified |  |  | 753245 | Not_Computer
Z-0633-2016 | 72808 | No | JJJ | Beta or gamma counter for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer WIZARD2 2-detector, 550 samples Product Code: 2470-0020 |  |  | December 08, 2015 | 2015 | Perkinelmer | The content of the second #023 barcode ID label in the Barcode ID Label Binder is incorrect. The second barcode ID label #023 symbol incorrectly identifies as #024 when scanned. If measurement protocols have been programmed for both ID #023 and #024, the protocol associated with barcode ID #024 is executed. If the error is undetected, the Gamma Counter may produce erroneous results. | Nonconforming Material/Component | Perkin Elmer sent an Urgent Medical Device Correction letter dated December 7, 2015 to affected customers. The letter identified the affected products, problems and actions to be taken. Customers were instructed to destroy all affected products in their possession. Customers were asked to complete theenclosed Correction Response Form and return it by scan and e-mail to QMresponse.SG@perkinelmer.com . Corrected barcode labels ID #023 will be provided upon request and return of the completed Correction Response Form to PerkinElmer. For further information contact your local PerkinElmer representatives.. |  | 60 | 60 units | Worldwide Distribution - US Nationwide and the countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, MYANMAR, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, and THAILAND. | Terminated | August 04, 2017 | 605 | 430 | Not_Computer
Z-0629-2016 | 72880 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reprocessed Stryker Color Cuff and Color Cuff II Adult Tourniquet Cuffs |  |  | December 17, 2015 | 2015 | Stryker Sustainability Solutions | Reprocessed Stryker Pressure Tourniquet Cuffs were mislabeled, including some labeled as single port when they contained dual port, and vice versa. | Under Investigation by firm | The firm sent out customer notification letters on 12/17/15. The firm requests that customers discontinue the use of the affected lots of the recall product, and other Pressure Tourniquet Cuffs reprocessed by Stryker Sustainability Solutions should be considered as alternatives. A Recall Effectiveness Check Form including the quantity of affected devices shipped to the customer is included with the notification.The firm states to complete the Recall Effectiveness Check Form and indicate if any devices from the affected lots remain in inventory. The letter states the form needs to be completed even if no product is found.If the form indicates that affected devices remain in inventory, a prepaid shippinglabel will be issued for the return of the product. Customers will receive a credit forall affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433X5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online, by regular mail, or by fax. |  | 450 | 450 devices | Nationwide Distribution including the states of AK, NV, VT, FL, CA, CO, WI, IA, MN, OK, PA, MI, OR, KY, and MA. | Open, Classified |  |  | 450 | Not_Computer
Z-0649-2016 | 72925 | No | JIL | Urinary glucose (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | iChem 10 SG Urine Chemistry Strips, PN 800-7005. |  |  | December 08, 2014 | 2014 | Beckman Coulter Inc. | Iris International is recalling the iChem 10 SG Urine Chemistry Strip because of false positive leukocyte result for samples that are truly negative. | Nonconforming Material/Component | An "Important Product Notice," dated 12/8/14 was sent to customers to inform them of the recall by Iris International. The recall letter informs customers that The identified lots of iChem 10 SG Urine Chemistry Strips may have leukocyte pads that are more sensitive at the first reporting range of 1+, corresponding to 25 white blood cells (WBCs) /microliter. The letter provides the impact of the recall and the actions to taken by the customers. Customers are instructed to discard any remaining strips from the affected lots and use another lot number. Customers can either contact their local Customer Service Support or complete and return the Replacement Order form, if replacement product is needed. The enclosed response form should be completed and returned within 10 days. Customers with questions are instructed to refer to http://www.beckmancoulter.com; (800) 854-3633 in US and Canada. Customers outside the US and Canada are instructed to contact their local support representative. |  | 16206 | 16,206 vials total (7,827 vials in US) | Worldwide Distribution -- US, Brazil, Canada, Chile, Ecuador, France, Germany, Greece, Italy, Kuwait, Luxembourg, Malaysia, Philippines, Portugal, Puerto Rico, Singapore, Spain, Switzerland, Turkey, and United Arab Emirates. | Terminated | January 13, 2016 | 401 | 16206 | Not_Computer
Z-0645-2016 | 72698 | No | LYJ |  |  | Neurology | PMA | Cyberonics Vagus Nerve Stimulator VNS Therapy AspireSR Generator Model 106 |  |  | November 18, 2015 | 2015 | Cyberonics, Inc | Certain Model 106 Pulse Generators demonstrate delays in sensing during use of the 'Verify Heartbeat Detection' feature and exhibit the potential for decreased battery longevity. | Manufacturing material removal | Cyberonics sent out written recall notification beginning October 14, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. For further questions, please call 866-882-8804 |  | 29 | 29 units | US Distributions to the states of : IL, NC, KS, and TX., and Internationally to the country of: United Kingdom. | Open, Classified |  |  | 29 | Not_Computer
Z-0646-2016 | 72938 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer Periarticular Locking Plate System 3.5mm Locking Screw 58 mm Length; Zimmer Periarticular Locking Plate System 2.4mm Locking Screw 14 mm Length. |  |  | December 09, 2015 | 2015 | Zimmer Biomet, Inc. | Product on the package label did not match what was inside the package. The label was for part number 00-2359-058-35 (Zimmer Periarticular Locking Plate System 3.5mm Locking Screw 58mm Length) lot number 62918019 while the product inside the package was part number 00-2359-014-24 (Zimmer Periarticular Locking Plate System 2.4mm Locking Screw 14mm Length) lot number 62896038. | Labeling Change Control | On 12/9/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital risk managers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 400 | 400 | Worldwide Distribution-US including the states of CA, NY, PA, AL, FL, GA, ID, WI, MO, SC, NC, and ME, and the countries of Brazil, Dominican Republic, Taiwan, China, Uruguay, Australia, Japan, Malaysia, India, Belgium, France, Italy, Netherlands, Poland, United Arab Emirates, and Canada. | Open, Classified |  |  | 400 | Not_Computer
Z-0650-2016 | 72521 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Integral Centralizer Hip System, Lateralized Femoral Stem (Co-Cr-Mo), |  |  | November 17, 2015 | 2015 | Biomet, Inc. | PMMA is listed as a material on the label but the product does not contain PMMA. | Labeling Change Control | On 11/17/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 21 | 21 | Worldwide distribution. US nationwide including IA, UT, MN, PA, MO, OK, IL, OH, and WA; CANADA, and NETHERLANDS | Open, Classified |  |  | 21 | Not_Computer
Z-0642-2016 | 72597 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System, Catalog Number 6802445, Unique Device Identifier Number 10758750012343, Software Version 3.2 and below; IVD. |  |  | November 03, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly may allow VITROS Systems to sample and process assays while the fluids and reagent temperatures are not in the required temperature range for optimal processing. When this intermittent issue occurs, the VITROS Systems continue to operate and process results without notification to the User, possibly producing erroneous results. | Software design | Ortho-Clinical Diagnostics sent an Urgent Product Correction Letters (Ref. CL2015-211, dated November 3, 2015) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected devices (VITROS 3600 and 5600 Systems). The notification identified the affected products, problem and actions to be taken. Foreign affiliates were informed via e-mail on 11/03/2015 of the issue and instructed to notify their consignees of the issue and required actions.Ortho-Clinical Diagnostics sent an Urgent Product Correction Letters (Ref. CL2015-221, dated November 11, 2015) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected devices (VITROS 5,1 FS and 4600 users). The notification identified the affected products, problem and actions to be taken. Foreign affiliates were informed via e-mail on 11/11/2015 of the issue and instructed to notify their consignees of the issue and required actions. |  | 427 | Domestic: 157, Foreign: 270 | Worldwide Distribution - US Nationwide and the countries of Puerto Rico, and foreign distribution to: Argentina, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Israel, Italy, Japan, Kenya, Mexico, Panama, Philippines, Poland, Romania, Russia, Singapore, Spain, Taiwan, Turkey, and Venezuela. | Open, Classified |  |  | 4399 | Computer
Z-0644-2016 | 72915 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Model NuCARTGN. The NuCART is a mobile display and image management platform used in Minimally Invasive Surgery (MIS) environments to facilitate the routing and ergonomic display of video signals from source devices to flat-panel monitors. It is tailored for use with an OEC C-arm and Endo-Cam in MIS surgery suites, Ambulatory Surgery Centers, Urology, GI and other suites. |  |  | December 04, 2015 | 2015 | CompView Medical, LLC | The attachment of the NuCART Spring Arm to the NuCART horizontal boom may come apart if the securing ring (snap ring) was not installed correctly causing the spring arm to release and detach. | Nonconforming Material/Component | The firm began inspecting devices between 12/4/15 and 12/16/15, and sent the Nucart Product Investigation letter to consignees between 12/11/15 and 12/15/15. Consignees with questions can call the firm at 503-601-5603. |  | 10 | 10 units | US nationwide distribution including AZ, CA, CO, GA, OH, and TN. | Terminated | March 31, 2016 | 118 | 10 | Not_Computer
Z-0647-2016 | 72804 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RT Elements are applications for radiation treatment planning for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. The simulated plan is intended for treatment evaluation for example in tumor board meetings or operating rooms |  |  | November 05, 2015 | 2015 | Brainlab AG | Large objects with fine resolution are potentially displayed cropped when imported into Brainlab Brain Metastases 1.0.0. | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 5, 2015, to all direct accounts informing the of the recall and actions to be taken.User Corrective Action:Users of Brain Metastases Version 1.0.0 or Adaptive Hybrid Surgery Analysis 1.0.0 shall adhere to the following:After import always carefully review and verify all objects for correctness and validity.For reviewing an object, select the image set which was used to outline or to modify the objects in the Data menu. Also compare the contours of an object in Brain Metastases or Adaptive Hybrid Surgery Analysis software to the contours in SmartBrush. Please note that the error will not be visible in SmartBrush. Always verify the correct shape of the objects in Brain Metastases or Adaptive Hybrid Surgery Analysis software. If possible do not use image data sets with large slice distance and/or with a low amount of slices. To retrospectively identify if cropped objects have been used for treatment planning, the object contours of a specific treatment plan can be reviewed in Brainlab Dose Review, where the error will also be visible.Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact affected customers tentatively starting February 2016 to schedule the update. The request that users advise the appropriate personnel working in their department of the content of the Notification letter. If the user facilities require further clarification, thay can contact their local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers)E-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax: Brainlab AG: + 49 89 99 15 68 33 |  | 20 | 20 systems | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, Canada, China, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Mexico, Portugal, South Korea, Spain, Switzerland, Turkey and the United Kingdom. | Open, Classified |  |  | 68 | Computer
Z-0654-2016 | 72944 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Zimmer Universal Locking System 3.5mm Locking Reconstruction Plate |  |  | December 09, 2015 | 2015 | Zimmer Biomet, Inc. | The product on the package label did not match what was inside the package. The label was for part number 00-4936-011-13 lot number 62968956 while the product inside the package was part number 00-4936-008-07 lot number 62968974. | Labeling Change Control | On 12/9/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital risk managers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Specifically, customers were asked to review the notification and ensure all affected personnel are aware of the information; assist the Zimmer Biomet sales representative with the quarantine and removal of any of the identified affected product; and complete and return the Certificate of Acknowledgement form. If after reviewing this notification customers have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 52 | 52 | Worldwide Distribution -- USA, including the states of MN, NY, MS, AL, OH, KY, ME, and NM; and, the countries of Taiwan, Korea, Singapore, El Salvador, Uruguay, Czech Republic, France, Italy, Poland, United Kingdom, South Africa, and Slovakia. | Open, Classified |  |  | 52 | Not_Computer
Z-0655-2016 | 72897 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ergon 2 Articulating Equipment Boom Monitor Brackets |  |  | December 02, 2015 | 2015 | Skytron, Div. The KMW Group, Inc | The hardware that holds the VESA plate and the display interface / monitor to the bracket may fail causing the display interface to fall from the bracket. Over time and with use the threaded rod hardware material can weaken and detach from the bracket. A display interface falling from its mounting location may result in injury to the patient or provider. | Device Design | Skytron sent an Urgent Medical Device Recall letter dated December 02. 2015, to all affected customers via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with questions were instructed to contact at 616-656-2900 with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 286 | 286 | Worldwide Distribution - US including WI, TX, FL, NC, MD, OH, AZ, CA, UT, MT, OK, HI, CO, WY, MO, NJ, IL, NY, MA, IA, DE, PA, MI, IN, and WV and Internationally to Canada, and Japan. | Open, Classified |  |  | 286 | Not_Computer
Z-0653-2016 | 72776 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 30 deg, 8mm Endoscope, IS4000: |  |  | November 23, 2015 | 2015 | Intuitive Surgical, Inc. | This correction to the da Vinci Xi I&A; User Manual communicates proper use of the da Vinci Xi 8 mm Endoscopes in combination with the da Vinci 8 mm and 12 mm cannulae, cannula reducers, and monopolar energy instruments. It informs customers of scenarios that may result in unintended patient harm due to capacitive coupling from monopolar energy sources. | Device Design | Intuitive Surgical sent an Urgent Medical Device Correction letter to affected customers on December 4, 2015,. Letters explained the issue and provided action to follow to prevent the issue. Customers were instructed to do the following:Distribute a copy of this letter to all da Vinci Xi users at your facility.3. Place a copy of this letter with your user manual.4. Complete the attached Acknowledgement Form and return it to Intuitive Surgicalas instructed.5. Retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this Medical Device Notification ,please contact your Clinical Sales Representative or contact Intuitive Surgical CustomerService at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) ormail: customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8AM to 6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST) |  | 2624 | 2624 total all models | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, France, Germany, Hong Kong, India, Israel, Italy, Japan, Monaco, Netherlands, Norway, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Korea, Spain,Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela.. | Terminated | April 05, 2016 | 134 | 2624 | Not_Computer
Z-0659-2016 | 72896 | No | LYJ |  |  | Neurology | PMA | VNS Therapy AspireSR Generator Model 106. |  |  | December 18, 2015 | 2015 | Cyberonics, Inc | Recall being initiated in response to three reports of "Burst Watchdog Timeout" events occurring with the Model 106 AspireSR Generator, resulting in a device reset condition where stimulation output is disabled. | Device Design | The firm notified consignees of the issue via letter on 12/18/15. The letter identified the affected device, the issue involved, and actions to be taken. Physicians are to contact Clinical Technical Support at 866-882-8804 to report if a patient's generator has been disabled due to the issue identified. Users are to complete and return the effectiveness card as soon as possible. If further information is needed, customers can contact Clinical Technical Support at 866-882-8804 or via e-mail at cservices@livanova.com. |  | 4935 | 4,935 units | Worldwide Distribution -- United States, Austria, Belgium, Croatia, Czech Republic, Finland, France, Germany, Iceland, Italy, Netherlands, Norway, Poland, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Cyprus, Israel, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 4935 | Not_Computer
Z-0657-2016 | 72697 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | FMS neXtraÂż 3.5mm Aggressive Cutter Shaver Blades; Product Code: 283315 |  |  | November 20, 2015 | 2015 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product is incorrectly labeled. The blade configuration in the affected lots does not match the description on the label. | Process control | On November 20, 2015, customers were notified of the recall via letter. They were asked to take the following actions: Customer immediate actions:1. Review inventory located within your facility and quarantine any products listed in this product recall notification. Please follow the instructions on the Business Reply Form to return any affected inventory located within your facility; 2. Review, complete, sign and return the attached business reply form to Elizabeth Messana (emessan1@its.jnj.com) in accordance with the directions on the business reply form within 5 business days of receipt of this notification; 3. Forward this notice to anyone in your facility that needs to be informed; 4. If any potentially affected product has been forwarded to another facility, contact Elizabeth Messana at 508-828-3150; 5. Keep a copy of this notice. |  | 45 | 45 units | Worldwide Distribution. US states of TX, OH, MI, and ME; and the countries of FR, DE, SK, IT, NL, CH, and CZ. | Terminated | April 12, 2017 | 509 | 210 | Not_Computer
Z-0630-2016 | 72860 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | arkray SPOTCHEM II Glucose Reagent test strips, Item No. 77101, MEDTEST DX Catalog # 77424. |  |  | December 18, 2015 | 2015 | Arkray Factory USA, Inc. | One lot each of SPOTCHEM II Basic PANEL-1 Reagent Test Strip and SPOTCHEM II Glucose Reagent Test Strip are being recalled due to a lowered reaction rate in the reagent pad for glucose. The affected lots will display a result of approximately 300 to 400 mg/dL for any glucose value that exceeds the upper limit of 416 mg/dL. | Component design/selection | The firm sent an Urgent Medical Device Recall letter to customers beginning 12/18/2015. The letter identified affected product, stated the reason for recall and asked for unused product to be returned. For instructions on returning unused product contact Lesty Jaranilla 1-877-538-8872 x3243. |  | 88 | 88 boxes (25 foil packaged test strips per box) | Distributed in the states of FL, IL, KY, MI, NC, NY, OH, and TN. | Terminated | January 04, 2017 | 383 | 99 | Not_Computer
Z-0661-2016 | 72953 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Fukuda Denshi patient monitor model DS-8100M and DS-8100N. |  |  | December 21, 2015 | 2015 | Fukuda Denshi USA, Inc. | A possible malfunction where a Short-term Battery Error message has appeared and has been traced back to a bad (damaged) Lithium-Ion capacitor (Short-term Backup battery). | Process control | Fukuda Denshi USA, Inc. of Redmond, WA plans to perform 100% field correction of units distributed in the US.The Recall/Field Correction Notice letter (FCN-A01-201512), dated December 21, 2015, and the Customer Confirmation Sheet will be sent to consignees using a clearly red-marked URGENT via Express Mail.The consignees are requested to fill out the Customer Confirmation Sheet and return via fax to 888-224-7090, email to DBlakely@fukuda.com, or send via postal service back to Fukuda Denshi USA, Inc. in a pre-stamped envelope.Consignee can call Doug Blakey directly at 210-240-4521 with any questions or concerns. |  | 95 | 95 units distributed in the US | Distributed in the states of CO, ID, KS, MI, MO, OR, PA, TX, UT, and WI. | Open, Classified |  |  | 95 | Not_Computer
Z-0598-2016 | 71924 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab Cranial Navigation System: An Image Guided Surgery System / Stereotactic. Radiology Departments. |  |  | April 22, 2013 | 2013 | Brainlab AG | Software Error: The effect of setup on overall navigation accuracy could potentially intensify small inaccuracies arising from individual steps of a complex navigation procedure that may cause an inaccurate display of instruments by the navigation system compared to the actual patient anatomy. | Software in the Use Environment | The firm, Brainlab AG, sent out a "Field Safety Notice / Product Notification" CAPA-20130417-000315 Brainlab Cranial Navigation System: Effect of setup on overall navigation accuracy for all versions of software on 4/22/2013 and an update on 5/29/2015.- From April 22, 2013 has been distributed via email to the consignees.- UPDATE from May 29, 2015 was distributed via email to the medical users.An active reply by the consignees will be followed, if necessary via further different communication media.FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter describes the product, problem and actions to be taken. The customers were instructed to adhere to the measures described in document Measures to Improve Cranial Navigation Accuracy in the Product Notification letter when using the Brainlab Cranial Navigation System.Brainlab is requesting that 1. Existing potentially affected customers receive this Product Notification Letter.2. These customers receive the attached supplement to the Instructions for Use regarding the Brainlab Cranial Navigation System in hardcopy version as an amendment to their current User Manuals. Tentative planned timeline for availability: June 2013Additionally to advise the appropriate personnel working in your department of the content of Product Notification letter.If the medical users have the need for further clarification, they are asked to contact their local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33 Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  | 1430 | 1,430 units | Worldwide Distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Latvia, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open, Classified |  |  | 1430 | Computer
Z-0662-2016 | 72878 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Excella II Rod Cat. No.: SR 55100 5.5 x 100 mm LOT 1502 and Excella II Rod Cat. No.: SR 55110 5.5 x 110 mm LOT 1502 |  |  | December 07, 2015 | 2015 | Innovasis, Inc | Innovasis recalled two lots of Excella II titanium rods which were packaged using each other's labels. | Labeling mix-ups | Innovasis sent a Voluntary Product Recall letter, dated December 7, 2015 to customers regarding two lots of Excella II titanium rods which were packaged using each other's labels. Customers were asked to check their inventory and either remove the rod from the mislabeled package and request additional rods, or return the unopened packages immediately. The firm will replace all returned inventory upon receipt. |  | 1519 | 17 units Labeled SR55100 Lot 1502 with SR55110 rods in package | Distributed in US -- CA; KS; OH; &amp; TX. | Terminated | April 07, 2016 | 122 | 1519 | Not_Computer
Z-0660-2016 | 72569 | Yes | LWS |  |  | Cardiovascular | PMA | Optisure Dual Coil Defibrillation Leads, Model No. LDA220, LDA220Q, LDA230Q, LDP220Q. St. Jude Medical. Designed for long term attachment to an Implantable Cardioverter Defibrillator (ICD). |  |  | November 03, 2015 | 2015 | St Jude Medical Cardiac Rhythm Management Division | Leads may have been compromised during the manufacturing process, with cuts in the insulation layer of the SVC shock coil. | Process control | An Important Medical Device Advisory letter was sent to doctors on 11/3/15 to inform them of the recall by St. Jude Medical. The letter informs the doctors that St. Jude Medical identified that during the manufacturing process of a limited number of Optisure leads, a trim technique to remove excess medical adhesive around the SVC shock coil may have introduced damage to the leads insulation. The letter provides the doctors with the recommendations for patients with a potentially impacted Optisure lead connected to a device with Dynamic TX technology and for patients implanted with a potentially impacted Optisure lead connected to a device without DynamicTx technology. Doctors with questions about the advisory are instructed to contact their local Sales Representative or St. Jude Medical Technical Services at (800) 722-3774. |  | 447 | 447 units total (281 units in US) | Worldwide Distribution US (nationwide), Belgium, Switzerland, Germany, France, United Kingdom, India, Italy, Japan, Luxembourg, The Netherlands, Saudi Arabia, and Sweden. | Open, Classified |  |  | 447 | Not_Computer
Z-0686-2016 | 72836 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Wilson Penile Implantation System Retractor Frame, Retractor Stays &amp; Penile Elevation Strap, SINGLE USE ONLY Rx ONLY. |  |  | November 25, 2015 | 2015 | Abeon Medical Corporation | The firm was notified by some of their customers that some of the sealed pouches contain tears compromising the sterility of the product. | Packaging | On 12/17/2015 the firm sent Recall Notification Letters to their customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 575 | 575 units | Worldwide Distribution-US (nationwide) including the states of AL, AR, AZ, CA, CT,FL, GA, IA, IN, KS, KY, MA, MD, MI, MN, NC, NH, NJ, OH, SC, TN, TX, UT, VA &amp; WI., and the countries of Austria, Germany, Italy, Japan, Spain, Sweden &amp; Switzerland. | Open, Classified |  |  | 4180 | Not_Computer
Z-0674-2016 | 72630 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Valeant Pharmaceuticals North America, Small Particle Size Aerosol Generator, SPAG-2 6000 series. Model No. SPAG-2. A nebulizer pneumatic flow system for the administration of Virazole (ribavirin) aerosol. |  |  | September 21, 2015 | 2015 | Valeant Pharmacueticals International | The Serial I.D. Label and some information in the Operation Manual for the SPAG 2 6000 series is incorrect or out of date, including references to accessories that are no longer available. | Finished device change control | Valeant Pharmaceuticals International, Inc. sent Field Correction Notifications /Recall Stock Response Form dated September 22, 2015 through 2nd day air service to their customers. The notification informed the customer about the field correction, how to identify affected product and what actions to take. TThe customers were requested to fill out and return the attached Recall Stock Response Form and return the form via email (bauschandlomb6443@stericycle.com or fax (844.782.5565). A customer service number was provided in case the customer had any related questions or concerns (1.800.321.4576)On October 7, 2015 a second field correction notification was sent to customers. The notification informed the customer about the field correction, how to identify affected product and what actions to take. A customer service number was provided in case the customer had any related questions or concerns (1.800.321.4576). The updated notification clarified that the notification applies to all SPAG-2 6000 model units currently in the market. |  | 1178 | 1178 | USA (nationwide) Distribution and Internationally to Canada and France. | Open, Classified |  |  | 1178 | Not_Computer
Z-0675-2016 | 72997 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | InVision-Plus and InVision-Plus Junior Needleless IV Connector, Disposable, Single Use |  |  | January 05, 2016 | 2016 | Rymed Technologies, LLC | Potential for the female luer component of the InVision-Plus IV connector to crack during use. | Process design | The firm Ry Med Technology LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 4, 2016 and response form to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use and return the product to: RyMed Technologies LLC, 6000 W. William Cannon Drive, Building B, Suite 300 Austin, TX 78749; the user facilities to change the IV connectors in 24 hour when lipids are being administered; if further distributed, contact other accounts, advise of recall and return outstanding stock; complete and return the RECALL NOTIFICATION ASAP and no later than 10 business day via email to: amccutchen@rymedtech.com, fax to: 512-301-7338, Attention Recall Return Notification or mail to RyMed Technologies LLC, Attention Recall Return Notification, 6000 W. William Cannon, Building B, Suite 300 Austin, TX 78749. If you have any questions concerning the recall, contact Director of Quality Assurance/Regulatory Affairs by email to: amccutchen@rymedtech.com or call 512-301-7338x303. If you have questions about returning/replacement of product contact Logistics Specialist by email to: svara@rymedtech.com or call 512-301-7338x309. |  | 496400 | 496,400 units | US Nationwide Distribution. | Open, Classified |  |  | 543000 | Not_Computer
Z-0687-2016 | 73038 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The e.cam Dual Signature Gamma Camera system. Used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588keV. |  |  | December 31, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | To provide operator instruction manuals for the e.cam Dual Signature; Gamma Camera that were not sent in its entirety when the e.cam systems were refurbished and shipped to the customer. | Packaging process control | An updated Customer Advisory Notice Letter will be provided to each impacted consignee. The receipt of the letter will be documented and additional attempts will be made until such receipt is acknowledged. The Customer Advisory Notice asks that the letter be placed with systems instructions for use along with all provided user documentation. It also advises the consignee to forward the notice to the new customer if the equipment was sold.Customers with questions regarding the Customer Advisory Notice can contact their local Service representative or use the following contact number 1-800-888-7436. |  | 10 | 10 | Distributed in the states of Florida, New York, Indiana, Maryland, South Dakota, Missouri, Tennessee, and Georgia, and the countries of Australia and India. | Open, Classified |  |  | 10 | Not_Computer
Z-0672-2016 | 72202 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Accusorb MRI |  |  | October 06, 2014 | 2014 | MWT Materials, Inc. | When not used correctly, Accusorb MRI products may overheat and lead to patient injury. | Device Design | MWT issued a letter dated 10/4/2014 via email to all distributors for further distribution to consignees. Tghe letter identified the affected product, problem and actions to be taken. Questions can be direct to info@MWT-Materials.com. |  | 1 | 1 units | US Nationwide Distribution in the states of NY and MA. | Open, Classified |  |  | 187 | Not_Computer
Z-0683-2016 | 73036 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | IBA Dosimetry GmbH Lift Table Type 4320, Catalog HA03-010#2 and HA03-000#2 |  |  | December 10, 2015 | 2015 | Iba Dosimetry Gmbh | When the electric lift table is connected to the mains power, in a very rare occasion, the table may start to move up in an uncontrolled manner. | Component design/selection | The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email on 12/10/2015 then followed by 12/17/2015 sent by Fed Ex. The firm requested the mains power be disconnected after setting up the table and BP. Replacement part and instructions will follow. |  |  |  | Nationwide and International: Algeria, Angola, Arab Republic of Egypt, Austria, Belarus, Belgium, Bulgaria, Bosnia and Herzegovina, Croatia, Czech Republic, Denmark, Ethiopia, France, Germany, Ghana, Greece, Hungary, Iraq, Ireland, Islamic Republic Iran, Italy, Kazakhstan, Kenya, Kingdom of Saudi Arabia, Kuwait, Lebanon, Libya, Morocco | Open, Classified |  |  |  | Not_Computer
Z-0680-2016 | 72943 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys, 1011-9000-000. |  |  | December 15, 2015 | 2015 | GE Medical Systems, LLC | The Drive Gas Check Valve of certain Anesthesia products could become stuck in a fixed open position, which could cause pressure to build in the mechanical ventilation cycle. If this issue is left unresolved, it could result in excessive or prolonged pressure in the patient breathing circuit during ventilation potentially resulting in barotrauma. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter dated December 15, 2015, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 13 | 13 | Worldwide Distribution - US (nationwide including PR except NH, VT, &amp; WY.Internationally to ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA, BRAZIL, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, CHANA, GIBRALTAR, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LITHUANIA,MALAYSIA, MALDIVES, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI lanka, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM. | Open, Classified |  |  | 4047 | Not_Computer
Z-0689-2016 | 72965 | Yes | NIO |  |  | Cardiovascular | PMA | ZilverÂż 518RX Vascular Stent with Rapid Exchange Delivery System |  |  | December 22, 2015 | 2015 | Cook Medical Incorporated | Manufacturing of the joint between the inner catheter and stylet of the delivery system does not consistently meet specifications, and the joint may separate during deployment. Failure of this joint during a procedure may result in a partial stent deployment. The clinical effect of a partial deployment may be a need for surgical intervention, or may result in vessel trauma or an embolic event | Device Design | Cook Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 22, 2015, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235.On 2/5/2016 a second recall notification was sent to affected consignees, expanding the recall to All Lots of the affected product. |  | 753 | 753 | Nationwide Worldwide | Terminated | July 25, 2017 | 581 | 753 | Not_Computer
Z-0651-2016 | 70765 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Vantage Panoramic X-Ray System used to perform dental and orthodontic radiological imaging |  |  | September 04, 2014 | 2014 | Midmark Corp dba Progeny Inc | Failure of the primary collimator may result in unintentional movement of the device. The unintentional movement will not result in the generation of an error message. | Device Design | MidMark sent an Urgent Medical Device Field Correction letter dated September 4, 2014, to all affected customers. The letter included instructions for distributors to: 1) encourage end users to contact Midmark to arrange the field correction after the final solution becomes available; 2) if a partial image is reported to have occurred with a machine, the machine should be removed from use and the end user should initiate a service request per the normal procedure; 3) quarantine any affected machines in the distributor's inventory and contact Midmark Technical Support within 10 business days to provide the Serial numbers to on-site servicing of the affected machines can be scheduled. Distributors with questions about this market action can contact Midmark Technical Support by calling 888-294-3800 x 3 or by e-mail at imagingtechnsupport@midmark.com. |  | 832 | 832 systems | Worldwide Distribution - US including AL, AR, AZ, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI; and Internationally to Australia, Belgium, Canada, Czech Republic, France, Hungary, Lithuania, New Zealand, Poland, Russia, Saudi Arabia, South Africa, Switzerland, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Open, Classified |  |  | 832 | Not_Computer
Z-0690-2016 | 72904 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lipogems SPA Aspiration Cannulae, intended for aspiration of adipose tissue. LGC 13G (13G x 185mm), Sterile. |  |  | December 15, 2015 | 2015 | Globalmed Logistix | The cannulae may bend during procedure. | Under Investigation by firm | Consignees were notified by telephone and letter via FedEx on/about 12/15/2015. |  | 118 | 118 units | CA, CO, CT, FL, GA, NJ and NY. | Terminated | March 09, 2016 | 85 | 118 | Not_Computer
Z-0694-2016 | 73067 | No | FDF | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fuse 1C Colonoscope |  |  | December 18, 2015 | 2015 | EndoChoice, Inc. | The bending section of the device may partially separate from the insertion tube. Potential for tissue trauma. | Process design | Consignees were notified on12/09/2015 that a sales representative would be contacting them to facilitate the exchange of the colonsocope and the defective device will be returned to EndoChoice. |  | 253 | 253 units | Worldwide Distribution. US nationwide, Colombia, Denmark, France, Germany, Indonesia, Israel, Italy, Japan, Malaysia, Republic of Korea, Spain and Switzerland. | Terminated | March 24, 2016 | 97 | 253 | Not_Computer
Z-0695-2016 | 72927 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Rack Loader; |  |  | December 15, 2015 | 2015 | Bio-Rad Laboratories, Inc. | On a rare occasion, there is a potential to assign a patient result to an incorrect sample ID when running in D10 Rack Loader configuration. | Software design | Urgent Medical Device Correction letters were sent on December 2015 to all affected consignees and subsidiaries. Each subsidiary will be given a Subsidiary Medical Device Correction Notice, Customer Medical Device Correction Notice, and Return Response Form and the number of customers distributed by e-mail. Each Subsidiary much contact customers and translate the Customer Medical Device Correction Notice and Return ResponseForm as needed. A revised letter to subsidiaries and customers was sent on January 20, 2016. |  | 365 | 165 in US, 200 Internationally | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia, Belgium, Brazil, Brunei Darussalam, China, Czech Republic, France, Germany, Greece, India, Israel, Italy, Malaysia, Maldives, New Zealand, Norway, Austria, Poland, Portugal, Russia, Spain, Sweden, Switzerland, Thailand, and UK. | Terminated | May 02, 2016 | 139 | 365 | Not_Computer
Z-0693-2016 | 72926 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler-Peritoneal Dialysis Cycler |  |  | December 15, 2015 | 2015 | Fresenius Medical Care Renal Therapies Group, LLC | The door latch may not fully close and the door may open unexpectedly. The door may make physical contact with anything near the door of the machine, including the patient. May cause injury, delay in treatment, or Increased Intraperitoneal Volume (IIPV) (overfill). | Device Design | Fresenius Medical Care issued on December 15, 2015, a n Important Product Notification to all affected customers informing them of the problem and how to recognize the proper latch closing. Questions regarding the contents of the Customer Notification or questions regarding the issue, contact Technical Services at 1-800-227-2572 and reference the Field Action Number FA-2015-10-W.A reply form is included as part of the customer communication withinstructions to fill out and return via fax or scanned/emailed copy. |  | 14094 | 14,094 units | Nationwide Distribution | Open, Classified |  |  | 14094 | Not_Computer
Z-0697-2016 | 72969 | No | LKP | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | EliA SmDP Well, Article Number 14-5624-01 |  |  | December 28, 2015 | 2015 | Phadia Ab | All equivocal and positive results (greater or equal 7 EliA U/ml) on EliA SmDP well lot 0018 may be incorrect and must be considered invalid. Patient samples in complaint investigations caused unspecific signals up to 22 U/ml on EliA SmDP well lot 0018. The nonspecific signals are not caused by anti-Sm antibodies, nor by streptavidin antibodies. | Nonconforming Material/Component | On 12/28/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions call 1-800-3464364, option #2. |  | 76 | 76 | US Nationwide in the states of NC, GA, NJ, NY, PA, TX, IL, MI, MT, NM, and OR. | Open, Classified |  |  | 76 | Not_Computer
Z-0701-2016 | 72864 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network. |  |  | December 15, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | An issue with the Panorama Central Station may cause the system to spontaneously restart. This may occur after about 49 days of continuous operation. On restart, monitoring will be lost for about 2.5 minutes. Settings and configurations are maintained. | Software design | The firm, Mindray DS USA, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 12/15/2015 to its customers via certified mail. The letter described the product, problem, and actions to be taken. The customers were instructed to route the recall letter to all potential users of your facility. Mindray will resolve the issue by upgrading the system software on all affected Panorama systems in your facility. Please contact your Mindray Service Representative at 1-800-288-2121 (Monday-Friday 8:30am-5:30pm) to arrange upgrade.If you have any questions, please contact Director, Quality Operations and Regulatory Affair at 201-995-8407 or email to: d.arpino@mindray.com. |  | 63 | 63 products | Nationwide Distribution. | Open, Classified |  |  | 63 | Computer
Z-0700-2016 | 72908 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Biosure Ratchet Driver, Product No: 72201888 For delivery and placement of orthopedic screws. Arthroscopic surgical instrument |  |  | April 04, 2013 | 2013 | Smith &amp; Nephew, Inc., Endoscopy Div. | Driver does not meet specification, oversized. Driver may not be able to be fully inserted into the screw, or screw may become stuck on the driver. | Employee error | Smith &amp; Nephew sent an "Urgent-Product Recall 1st. Notification" dated April 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please inspect your inventory and locate all devices from the above listed product and lot numbers and quarantine them immediately. For further questions, please call (978) 749-1000. |  | 9 | 9 | Worldwide Distribution - US Distribution to the states of : AK, AL, CA, CO, FL, IL, IN, MI, NE, NJ, NY, OH, OK, PA, SC, TN and TX., and to the countries of : Austria, Australia, Canada, Dubai, Great Britain, India, Italy, Korea, Malaysia, Netherlands, Portugal, Sweden, Switzerland, Singapore and South Africa. | Open, Classified |  |  | 166 | Not_Computer
Z-0735-2016 | 72910 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | MediChoice Cohesive Bandages, One (1) roll of cohesive bandage per bag, thirty-six (36) bags per case [all model numbers have equivalent packaging] |  |  | November 17, 2015 | 2015 | Owens &amp; Minor Distribution, Inc. | MediChoice sterile cohesive bandages imported and sold under a private brand by Owens &amp; Minor were tested and found to have microbial contamination. | Under Investigation by firm | The firms recall strategy is as follows : recall was initiated by the Owens &amp; Minor Home Office Recall Officer. The Home Office Recall Officer e-mailed the following documents to the Distribution Center (DC) Recall Officer at each affected DC:1. DC Recall Letter (specific instructions for the DCs),2. Recall Inventory Report form (for the DCs to record the amount of affected product quarantined),3. Customer Letter (specific instructions for the customers),4. Recall Return Response Form (for the customers to report back the details of the quantity of affected product in their inventory),5. Attachment 1 (a document that provides examples of the product labeling),6. Recall Checklist (a checklist for the DCs listing the actions that need to be completed), and7. Recall Combine Queries spreadsheet (listing customers and current on-hand inventory at the DCs).The DC Recall Officer at each DC forwarded the information to the DC Customer Service team. The DC Customer Service team notified each impacted customer via e-mail or letter based upon the customers recorded preference.Level to which the recall was initiated: The recall was initiated to the level of the customer (i.e., hospitals, clinics, surgery centers).Follow-up plans: A minimum of three (3) attempts are made to obtain an acknowledgment from the customer that they received the recall notification. An attempt to determine the status of non-responsive businesses will be made. If the non-responsive business is still viable, an attempt will be made to determine the employee responsible for management of recalls and the recall notice will be sent to that employee. If the business is no longer viable, the steps described below for out-of-business accounts will be followed.Sub-recall: No sub-recall is required. These products are supplied to User Facilities and not distributed to commercial sub-accounts.Plans for effective checks:1. Check records to ensure that the initial recall not |  | 64011 | 64011 | Nationwide Distribution. | Open, Classified |  |  | 202327 | Not_Computer
Z-0712-2016 | 72599 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD20 Biplane OR Table; Model Numbers: 722025 722020 |  |  | November 03, 2015 | 2015 | Philips Electronics North America Corporation | The use of Allura Xper, Allura Clarity, Allura CV, Allura Centron and UNIQ systems regarding live images and still images can lead to still images being interpreted as live Images. | Device Design | The firm, Philips, issued "URGENT-Field Safety Notices" dated 10/26/2015 and 11/19/2015 to its customers. The notices described the product, problem and actions to be taken. The customers were instructed to ensure the audible signal (X-ray buzzer) is always enabled, to activate only ONE fluoroscopy pedal at time, and to configure the footswitch to the default Philips configuration. Customers were also instructed to review the information with all members of your staff who need to be aware; retain a copy with the equipment IFU and complete and return the Confirmation Respond form to Mail Stop 4302, 3000 Minuteman Road, Andover, MA 01810 or email: Debra.Gibbs@Philips.com. Philips is tracking the delivery of these notices via certified return receipt.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377. |  | 13297 | 13297 in total | Worldwide Distribution: US (Nationwide) including Puerto Rico and countries of: Albania, Algeria, US Virgin Islands, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Democratic Congo, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Fiji, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius, Mexico, Moldova, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, RÂżunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 13297 | Not_Computer
Z-0731-2016 | 72946 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Baxter, RAPIDFILL Syringe Strip, 10 ml. |  |  | December 21, 2015 | 2015 | Baxter Corporation Englewood | Baxter Healthcare has issued an Urgent Product Recall communication for all lots of RAPIDFILL Syringe Strips due to lack of evidence supporting the packaging claim for 12-month shelf life on released product. | Under Investigation by firm | An Urgent Product Recall notification was sent to consignees via USPS, first class mail on 12/21/15. Baxter Healthcare has asked customers to do the following:l. Locate and remove all affected product from your facility. The product code and lot number can be foundon the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return, credit and alternative products.3. Complete the attached customer reply form and return it to Baxter by either fax or scanned e-mail.4. Communicate the contents of the letter to your staff if affected products are distributed to other facilitiesor departments.5. Conduct a consumer-level recall for dealers, wholesalers, distributors, resellers, or Original EquipmentManufacturers that distributed affected product to other facilities. |  | 4695 | 4,695 | Worldwide Distribution-US (nationwide) and the countries of Canada, Sweden, Germany, and the Netherlands. | Open, Classified |  |  | 4695 | Not_Computer
Z-0732-2016 | 73007 | No | NWV | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | PS-500 Scope Warmer; 10 units per box. |  |  | December 01, 2015 | 2015 | JosNoe Medical, Inc. | Damage to package seals potentially compromising sterility. | Packaging process control | On 12/1/2015, notifications were sent to the affected customers via e-mail. The recall notification included a description of the reason for the recall, the affected product, and instructions for identifying and returning the affected product. |  | 1008 | 1008 boxes | Distributed US (nationwide) and in Canada. | Terminated | March 22, 2016 | 112 | 1008 | Not_Computer
Z-0737-2016 | 72945 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Denali Filter-Jugular/Subclavian DL950J |  |  | March 13, 2015 | 2015 | Bard Peripheral Vascular Inc | IFU missing contraindications: Patients with uncontrolled sepsis and Patients with known hypersensitivity to nickel-titanium alloys. | Labeling Change Control | The firm sent out Medical Device Recall customer notification letters on 03/13/15.The notification letter stated that there is no required action on the customer's part. Customers were told that if the device did not meet their needs then the firm will provide a replacement unit. Effective immediately the firm is implementing a change in the ordering code for the device to accurately reflect the approved information with the IFU and corresponding retrieval kit.Any questions the firm states to call Bard Canada:1-800-268-2862 Western Canada1-800-387-9473 Ontario1-800-387-7851 Eastern Canada |  | 1183 | 1,183 units | Distributed to CA only. | Terminated | February 01, 2016 | 325 | 1183 | Not_Computer
Z-0608-2016 | 72960 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ottobock Modular Polycentric EBS Knee Joint 3R60, 3R60=ST, 3R60=KD, 3R60=HD. Used for the prosthetic fitting of amputations of the lower. limb. |  |  | January 04, 2016 | 2016 | Otto Bock Healthcare GmbH | Otto Bock HealthCare GmbH is recalling Modular Polycentric EBS Knee Joint that were serviced between 27 April 2015 and 17 November 2015 due to a service error where the wrong Loctite wasused during repair. | Process control | Starting Jan 4 2016, In coordination with the Manufacturer, the US Agent will notify customers via telephone call of the correction on behalf of the Manufacturer. The customers will be requested to return the affected devices back to the service center for correction. The devices will be returned to the customer following correction. |  | 40 | 40 units | Nationwide Distribution. | Open, Classified |  |  | 40 | Not_Computer
Z-0736-2016 | 72972 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | HemosIL Plasminogen, Part No. 0020009000, IVD. Instrumentation Laboratory Company. Kit contents: 2 x 2.5 mL Streptokinase reagent and 2 x 2 mL Chromogenic substrate |  |  | December 18, 2015 | 2015 | Instrumentation Laboratory Co. | This Lot is not meeting labeled on-board instrument stability claims of 5 days. | Device Design | Instrumentation Laboratory Co. sent an " Urgent Medical Device Removal" letter dated December 2015, was sent on December 18, 2015. via USPS to US and Canadian customers. The notification requires US and Canadian customers to respond using a Mandatory Response Tracking Form. A Field Safety Notice was e-mailed on December 18, 2015, to the affected IL affiliates, notifying them of the product lot removal due to a stability issue and requesting them to translate the Field Safety Notice, to notify their Competent Authorities and to provide the notification to their affected customer base and secondary distributors. For further questions, please call 1(800) 955-9525. |  | 305 | Domestic: 108 kits, Foreign: 197 kits | Worldwide Distribution - US Distribution and to the countries of : Canada, Argentina, Austria, Belgium, Brazil, China, Czech Republic, France, Germany, Hong Kong, Italy, Lithuania, Panama, Poland, Russia, Saudi Arabia, United Kingdom, and Venezuela. | Open, Classified |  |  | 305 | Not_Computer
Z-0811-2016 | 71008 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope Reusable Blades (GVL 3, 4, 5 and AVL 2, 3, 4, 5). |  |  | April 13, 2015 | 2015 | Verathon, Inc. | Verathon has revised the Operations and Maintenance Manual (OMM) to specifically state compatibility limitations of the reusable video laryngoscope Blade validated through simulateduse cycles, such that exceeding this number of cycles in the specific chemical agent may affect the potential life of the product. | Nonconforming Material/Component | Verathon sent an Urgent Medical Device Safety Alert letter dated April 13, 2015 and Safety Alert Reply Form to their consignees on April 17, 2015Users are informed that Verathon implemented a Safety Alert regarding updates to the following Operations and Maintenance Manuals (OMM):" GVL Operations and Maintenance Manual: 0900-1204-08-60" AVL Operations and Maintenance Manual: 0900-4200-02-60For a copy of the most recent GVL or AVL Operations and Maintenance Manual, please visit http://verathon.com/support.If you have any questions regarding the Operations and Maintenance Manuals revisions noted above, please contact your Verathon representative or Verathon Customer Care at 800.331.2313 or 425.867.1348. You may also email customerservice@verathon.com.PLEASE NOTE: NO ADDITIONAL ACTIONS NEED TO BE TAKEN. |  | 9088 | total 9088 units ( 5786 units in the US and 3302 units International) | Worldwide Distribution. | Completed |  |  | 9088 | Not_Computer
Z-0743-2016 | 73033 | No | KNK | Uric acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Uric Acid CP ref. A11A01670 is an in vitro diagnostic assay for the quantitative determination of uric acid in human serum, plasma and urine based on the enzymatic determination of uric acid using a chromogenic system in the presence of peroxidase and uricase (Trinder method). |  |  | January 11, 2016 | 2016 | Horiba Instruments Inc | N-Acetylcysteine (NAC) present in the blood of patients treated for paracetamol overdose can interfere with the Trinder reaction that uses hydrogen peroxide catalytic on aminoantipyrine and phenol, and can produce falsely low results with the reagents using the Trinder reaction method. | Under Investigation by firm | Customer notification letters were sent on 01/11/16.Customers were provided the following instructions/actions:When performing tests that use the Tinder reaction, the blood draw should be done before treatment is given to the patient.Reagent inserts for the concerned HORIBA Medical Products have been updated and now include the following additional warning: Patients treated with N-Acetylcysteine (NAC) for Paracetamol overdose may generate a false low result. The letter request customers share the information with laboratory staff and retain the notification as part of their laboratory documentation. The letter states it is mandatory for customers to complete and return the enclosed response form within 10 days in order for the firm to verify the understanding of the notice.If any questions contact the local HORIBA Medical representative. |  | 216 | 216 | U.S. distribution to the following; CA, NY, AL, OK, MN, MS, MD, TX, KY, WV, MT, IL, FL, VA, WA, MI, KS, NE, PA, OH, ID, OR, AR, AZ, NM, NC, SC, GA, TNNo foreign distribution. | Open, Classified |  |  | 216 | Not_Computer
Z-0744-2016 | 72913 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Medical Systems, BT2000 Automated Manifold Kit (SKU 014613) and CPT2000 (SKU 014623) Automated Manifold Kits, Sterile R, Rx Only, Made in Mexico. |  |  | December 21, 2015 | 2015 | Acist Medical Systems | Reports related to ingress of air into the manifold. | Nonconforming Material/Component | Acist sent an Urgent Medical Device Recall letter dated December 21, 2015, via e-mail to all affected consignees. The letter was addressed to Distribution Partner or Cardiac Cath Lab / Material Manager in the USA. In Europe, the letter was addresses to Purchasing Department. The letter described the problem and the product involved in the recall. Advised consignees to return affected units and to complete and return the Medical Device Recall Form. Distributors were further advised to send the recall letter to their customers. For questions or concerns, contact in US, ACIST Customer Service 1-888-670-7701 or your ACIST Sales Representative. Europe Direct contact ACIST Europe Customer Service +44 (0) 845 548 12 80 or contact your ACIST Sales Representative, Europe Indirect contact ACIST Europe Customer Service +31 45 750 70 25 or contact your Distributor. Distributors contact: ACIST Europe BV, Office: +31 (0)45 750 7015, Mobile: +31 (0)6 233 648 46. |  | 534754 | 48,614 cartons or 486,140 units | Worldwide Distribution - US including all states except DE, MT, RI, UT, &amp; VT and Internationally to Belgium, France, Germany, Iceland, Ireland, Luxembourg, Netherlands, Reunion, United Kingdom, Australia, Austria, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Estonia, Greece, Hong Kong, Hungary, Islamic Republic of Iran, Israel, Italy, Japan, Republic of Korea, Kuwait, Latvia, Lithuania, Mexico, New Caledonia, Norway, Panama, Poland, Portugal, Saudi Arabia, Serbia, Slovenia, Bosnia and Herzegovina, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, and Venezuela. | Open, Classified |  |  | 534754 | Not_Computer
Z-0745-2016 | 73023 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | DS/US Proton Feature with Pinnacle3 Software Version 10.0.0 and 14.0.0 Radiation Therapy Planning System Model #459800200621 |  |  | January 08, 2016 | 2016 | Philips Medical Systems | In DS/US proton planning, there is a correction factor used internally to the dose engine that is calculating improperly in some cases. This can result in an incorrect dose calculation. | Device Design | Philips sent an "Urgent Medical Device Correction" letter dated 16-DEC-2015 to affected customers. The letter described the Affected Products, Problem Description, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer / User, Actions Planned by Philips and Further Information and Support (contact Philips Customer Care Solutions Center at 1-800-722-9377 ). |  | 3 | 3 | US: Nationwide Distribution in the states of FL, OH, and MO. | Open, Classified |  |  | 3 | Not_Computer
Z-0748-2016 | 72959 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE System; Model Number: 933-599. The AQURE System manages blood gas and immunoassay analyzers. |  |  | December 29, 2015 | 2015 | Radiometer America Inc | The AQURE System has a design error regarding sample type in which sample type may be specified for some results and left blank for other results. The fact that the sample type field is left blank in some cases could lead to misinterpretation of results. | Device Design | Customers were sent a notification letter on 12/29/15.The notification states that the AQURE System can be used, but to eliminate the potential confusion, blank fields for sample type are to be interpreted as not specified. Customers are told that a Radiometer representative will be in contact to schedule a visit or a remote session. During the visit or remote session the representative will run a database script on the AQURE system which will serve for two purposes:1. Update existing patient results where the sample type field is blank to include Notspecified. 2. Install a so-called database-trigger, which will include the text Not specified in ablank sample type field for future patient results.An upgraded version of the AQURE software will include modified device driversto ensure that Not specified is inserted if the device transmits an empty fieldfor sample type. The new software version will be installed by local engineerwhen available. |  | 375 | 375 | Worldwide Distribution-US (nationwide) including the states of PA, GA, MI, OH, VA, FL, AL, TX, OR, OK, TN, NY, and MA, and the countries of Austria, Czech Republic, Denmark, Estonia, United Emirates, Kuwait, Qatar, Saudi Arabia, Oman, Jordan, Finland, France, Germany, Hungry, Ireland, United Kingdom, Italy, Netherlands, Norway, Poland, Spain, Sweden, Switzerland, Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Russia, Singapore, South Africa, Republic of Korea, and Malaysia. | Terminated | March 29, 2016 | 91 | 375 | Computer
Z-0692-2016 | 72914 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Emergency Transport Ventilator |  |  | December 22, 2015 | 2015 | Draeger Medical, Inc. | Drager became aware of situations where the error message "Poti unplugged" was generated causing acoustic and visual alarms, the breathing system releases pressure and the ventilation function stops operating. | Device Design | The firm, Draeger Medical, Inc., sent an "Urgent Medical Device Recall" letter dated 12/21/2015 and a copy of the Instructions for Use (IFU) Supplement to consignees in a method that confirms delivery. The letter described the product, problem and actions to be taken. The consignees were instructed to include a copy of the supplement with their Oxylog IFU; inform all affected users in their facility; and contact Draeger Customer Support at 1-800-543-5047 (press 4 at the prompt) to obtain additional copies of the IFU Supplement Free of Charge.If there are any questions regarding this letter, please contact Michael Kelhart at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger Oxylog ventilator please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt).On 12/09/2016, Draeger mailed an Urgent Medical Device Recall notification letter to affected customers to inform them of a revised new software for the Oxlog 3000/3000plus (version 1.06 for Oxylog 3000 plus and version 1.23 for Oxylog 3000) that reduces the impact of the error condition. Customers will be contacted by a Drager Service Representative to schedule a time to perform the software update free of charge. The new software changes the device behavior in the event of the error (a controller failure). With the new software, in the event of a controller failure, ventilation will not stop. The device will continue to ventilate with the last setting. An alarm will annunciate and be displayed, alerting the user of the controller failure. |  | 402 | 402 units (314 - Oxylog 3000; 88 - Oxylog 3000 Plus) | Nationwide Distribution -- including states of AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | April 05, 2017 | 470 | 402 | Computer
Z-0747-2016 | 73000 | No | HTC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Small Electric Drive (SED). For use in general traumatology. |  |  | August 14, 2014 | 2014 | The Anspach Effort, Inc. | The affected Small Electric Drive (SED) may operate solely in reverse mode, not operate in reverse mode when intended, not function, or unintentionally start. | Nonconforming Material/Component | On 11/13/2014 DePuy Synthes send a letter to all their customers informing them of the recall of all the lots of the Small Electric Drive (SED) with instructions to returned and possible exchange. Contact is Complaint Handling Unit Manager at 561-494-3673 |  | 599 | 599 each | Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CO, FL, ID, IL, KS, LA, MI, NC, NY, TX, WA, VA and HI, and the countries of Germany, Hong Kong, Chile, Israel, Poland, Italy, Turkey, Italy, Singapore, Ireland, Korea, Spain, Norway, Denmark, Great Britain, Finland, Belgium, Netherlands, France, Japan, Sweden, Switzerland, Austria, China, India, Thailand, Russia, New Zealand, Zaire, Australia, Taiwan, and Malaysia. | Terminated | February 05, 2016 | 540 | 599 | Not_Computer
Z-0696-2016 | 73046 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ET-20 Anesthesia Extension Tubes, 2.5 ml approximately, Sterile and Disposable, Wall Medical Inc., Made in China |  |  | January 06, 2016 | 2016 | Global Healthcare Inc | Firm was notified of a potential product defect related to kinking (flattening) of anesthesia extension tubes which may represent a potential health hazard. | Nonconforming Material/Component | Global Healthcare sent an "Urgent: DEVICE RECALL" letter dated January 6, 2015 to its Customer by email. The letter described the product, problem and actions to be taken. The customer was instructed to immediately examine its inventory to determine if you have any of the Model ET-20 on hand; if so, immediately discontinue distributing product; and complete and return the Distributor Device Recall Response Form by email. Additionally, if you have further distributed any of the product, immediately contact your accounts; advise them of recall situation and have them return their outstanding recalled stocks to you. Any returned product will be destroyed by the distributor upon receipt of all affected product.If you require further assistance, please call 770-522-7520. |  | 2000 | 2,000 cases (100 individual units per case) | US Distribution to state of: GA only. | Terminated | March 25, 2016 | 79 | 2000 | Not_Computer
Z-0746-2016 | 73071 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Astral 100, 150 |  |  | January 13, 2016 | 2016 | Resmed Corporation | Resmed Corporation is recalling Astral 100, 150 External Battery because of a start-up issue. | Software design | ResMed sent an Urgent Safety Alert dated January 13, 2016, to all affected customers. The Alert identified the product, the problem, and the action to be taken by the customer. The notification letter provided the following instructions to distributorsField devices can be upgraded using ResMed Service Software or at an authorized ResMed Service Center. Contact Astral Support for further details.The SR2 software update will be available free of charge and should be implemented for all devices within the two-year service cycle or before at the customers preference.To acknowledge receipt of this Astral Safety Alert, submit this form.For any questions call Astral Support:855-245-4640 |  | 14736 | 14,736 | Worldwide Distribution - US including VA, FL, MO, CA, LA, TX, NJ, NH, NC, WV, MI, ID, OK, MI, MN, UT, CT, IL, AZ, UT, KY, MT, PA, RI, TN, DE, NE, LA, NY, SC, WY, IN, VA, GA, CO and Internationally to CA, CO, and AR. | Terminated | April 22, 2016 | 100 | 14736 | Computer
Z-0755-2016 | 72391 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry XPT Systems, Software Version 1.0.3, SMN 11127538 |  |  | October 07, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Multiple Software issues. Auto Start-UP Fail, Calibration Interval Resets when a Reagent Blank is run, Control Definition screen assumes range defined is 2 SD, Laboratory Information System (LIS) communication/Laboratory Automation (LAS) issue, Printer Driver Resets, ISE Calibration Ranges are too conservative for Urine Sodium, Archiving and deletion may fail and Workstation services may restart. | Software design | On October 7, 2015 Siemens Healthcare Diagnostics distributed Urgent Field Safety Notices to their custoemrs via Fed Ex in the U.S. and via email to regional countries for implementation outside of the U.S. Customers were provided a work-around if they have experienced any of the following problems: 1) Intermittent auto-startup function fails, 2) calibration interval resets as though a full calibration was performed, 3) Users may be unaware that the QC definition screen with +/- suggests a 2 Standard Deviation range value, 4) a rescanned and re-queried sample may become stuck in a communication loop and all tests may read as "inprocess" if the the initial reaction process is incomplete, 5) The printer stops printing if multiple printouts are requested in quick succession, 6) the ISE calibration for Urine Sodium (Na) may fail due to the ranges being set conservatively during development, 7) The automatic or manually started, archiving and deletion maintenance activity may fail and 8) A workstation services restart screen may appear under multiple scenarios causing the interface to be temporarily inaccessible. If customers have experienced any of the previously mentioned workflow problems, please perform the following actions: 1. Auto Start Failing: Manually perform the Startup Wash, run controls and any calibrations that are due. 2. Calibration interval resets when an RBL is run: Calibration should be run according to the recommendations in the Instructions for Use (IFU) for each assay. Keep a record of required calibrations and ensure that all assays are calibrated before use. 3. Control definition screen assumes that the range is defined as 2 SD values: When creating QC definitions, enter the values associated with a +/- 2 Standard Deviation (SD) range in the +/- ranges field.4. Laboratory Information System (LIS) communication / Laboratory Automation (LAS) issue: Be aware of errors in the event log and take action as needed. Manual intervention and rerunning |  | 87 | 87 units | Worldwide Distribution: US (nationwide) in states of: AZ, CA, MA, NY, TX &amp; WA, and to countries of: Bahrain, Egypt, Slovakia, France, Spain, Italy, Norway, Singapore, New Zealand, Germany, Australia, Republic Korea, Netherlands, Denmark, Finland &amp; United Kingdom. | Open, Classified |  |  | 87 | Computer
Z-0750-2016 | 73027 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERTÂż FA, bioMÂżrieux, Inc., |  |  | December 24, 2015 | 2015 | bioMerieux, Inc. | Bottles may have been exposed to non-normal shipping conditions which may lead to media degradation and impact growth performance. | Packaging process control | Biomerieux sent a Product Recall Notice letter to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the referenced part and lot number and discard remaining inventory according to their facility's procedure for culture media. Customers were also instructed to contact their local bioMÂżrieux representative for product replacement. |  | 2741 | 2741 units in total | Distributed in the countries of Angola, Austria, Bosnia, Bulgaria, Cameroon, Czech Republic, Democratic Republic of Congo, France, Germany, Greece, Hungary, Italy, Kazakhstan, Kingdom of Saudi Arabia, Kosovo, Lebanon, Netherlands, Niger, Pakistan, Poland, Republic of Belarus, Romania, Russia, Serbia, Sri Lanka, Spain, Sweden, Switzerland, United Arab Emirates, Turkey and Zimbabwe. | Open, Classified |  |  | 2741 | Not_Computer
Z-0756-2016 | 73060 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | API ZYM Bx2 (Ref 70493) |  |  | January 08, 2016 | 2016 | BioMerieux SA | Discrepant identification results related to API identification strips. These discrepant identification results are due to false negative results on the tests of the strips linked to the ZYM B reagent. | Other | BioMÂżrieux sent an Urgent Field Safety Notice letter DATED jANUARY 8, 2015 to affected customers. The letter identified the affected product, problem and action to be taken. dated . The letter indicated that they could continue to use product in inventory with mandatory requirements that before using a ZYM B ampule and each day of use of API strips that they perform a quality control with ATCC strains as described in each package insert of involved API strips. The letter also requested that they distribute this letter to all appropriate personnel in the laboratory, retain a copy for their files, and forward the information to all parties that may use the product. For additional assistance or have any questions, contact your local bioMÂżrieux Customer Service representative. |  | 13271 | 13,271 | Nationwide Distribution in the states of: AL, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, ME, MD, MA, MI, MN, MS, MO, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TX, VT, VA, WA, WV, WI, WY | Terminated | July 14, 2017 | 553 | 13271 | Not_Computer
Z-0754-2016 | 72968 | No | MTF |  |  | Immunology | PMA | ADVIA Centaur PSA Assay; Catalog Number: 06574155/SMN 10310292 (100 test kit) and Catalog Number: 02676506/SMN 10310293 (500 test kit) |  |  | December 29, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | PSA assay is not meeting the High Dose Hook Effect expectation in the Instructions for Use (IFU). | Under Investigation by firm | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on December 30, 2015 to be delivered to customers on December 31, 2015. An Urgent Field Corrective Action (UFCA) was sent out for distribution to all customers outside the United States on December 29, 2015. These notices identified the affected product, problem and actions to be taken. For questions customers were instructed to contact Siemens Customer Care Center or local Siemens technical support representative. |  | 55754 | 55,754 kits | Worldwide - US Nationwide including Puerto Rico, OUS: Albania, Algeria, Angola, Armenia, Austria, Bahrain, Belarus, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iceland, Iraq, Italy, Jordan, Kazakhstan, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Oman, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russian Fed., Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Turkey, Turkmenistan, Unit.Arab Emir., Vatican, AR, AU, BD, BR, CA, CL, CN, CO, DO, EC, GT, HK, ID, IL, IN, JP, KR, LK, MM, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, TH, TW, UY, VN and ZA. | Open, Classified |  |  | 55754 | Not_Computer
Z-0757-2016 | 73059 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus TJF-Q180V flexible gastrointestinal Duodenoscope |  |  | January 08, 2016 | 2016 | Olympus Corporation of the Americas | Olympus America Inc is conducting a voluntary removal/corrective action of all TJF-Q180V duodenoscopes in order to replace the forceps elevator mechanism Olympus is replacing the forceps elevator mechanism with a new forceps elevator design consistent with the design specification in the recently cleared TJF-Q180V 510k. | No Marketing Application | Olympus America Inc. sent an Urgent Notification letter dated January 15, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory for the affected product, Olympus will contact facilities to make arrangements for return of the TJF-Q180V duodenoscope(s) for the elevator mechanism replacement. For questions call (484) 896-5688. |  | 4436 | 4436 | US Nationwide Distribution except PR | Terminated | January 09, 2017 | 367 | 4436 | Not_Computer
Z-0763-2016 | 72966 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Mavig Monitor Suspension System, Model GD60, contained in the following Fluoroscopic imaging systems: Advantx LCV+, Advantx LCN+, Advantx LCLP+, Advantx LCA, INNOVA 2000, INNOVA 4100, INNOVA 3100, INNOVA 3131IQ, INNOVA 2121IQ, INNOVA 2100IQ, INNOVA 3100IQ, and INNOVA 4100IQ |  |  | December 18, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a reported incident in which a Mavig suspension arm with monitor assembly fell to the floor. | Nonconforming Material/Component | Consignees were sent on 12/18/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#12239 dated December 18, 2015. The letter was address to Director of Biomedical Engineering, Director of Radiologyand Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2111 | 2,111 (1078 US; 1033 OUS) | Worldwide Distribution-US (nationwide) including Guam and Puerto Rico and the countries of Algeria, Angola, Argentina, Australia, Bahrain, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, United Kingdom, Uruguay, Venezuela, and Yemen. | Open, Classified |  |  | 2111 | Not_Computer
Z-0766-2016 | 71764 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare TraumaDiagnost; Codes: 72001, 720011, 720013, 720020; A radiological examination unit for emergency work and routine skeletal diagnosis. |  |  | July 22, 2015 | 2015 | Philips Electronics North America Corporation | If arm is not fastened securely it could fall down. | Nonconforming Material/Component | Consignees were notified of the recall via letter sent certified return receipt on 7/22/2015. The letter informs customers that the Instruction for Use (IFU) requires the operator not to use the system and call for service, if a mechanical defect or malfunction is suspected. In this case, it is recommended not to position the Ceiling suspension with the Trauma arm directly above a patient as long as this Field Change Order is not yet installed. Philips plans to repair the affected systems. A Philips Service Engineer will contact customers as soon as the Field Action Kit is ready to be implemented. For questions, contact your local representative or use the Technical Support Line at 1-800-722-9377. When communicating with Philips regarding this issue, please reference Field Change Order 72000014. |  | 661 | 399 distributed, 262 are subject to this event | Distributed in the states of NY, WI, MA, AL, FL, GA, RI, CA, KS, and TX, and the countries of Canada, Australia, Austria, Belgium, France, Germany, Ireland, Italy, Netherlands, New Zealand, Portugal, South Africa, Spain, and Switzerland. | Open, Classified |  |  | 661 | Not_Computer
Z-0759-2016 | 73103 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Patient transport Shuttle with Trendelenburg quick adjustment function (Jupiter, Mars, Titan, TruSystem and Saturn operating table systems). |  |  | January 19, 2016 | 2016 | Trumpf Medical Systems, Inc. | Complaints were received of the table top on the transporter (shuttle) tilting down quickly in a maximum Trendelenburg position. | Component design/selection | TRUMPF Medical sent an Urgent Field Safety Information letter dated January 15, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were instructed to take the following actions:Actions to be taken1. Remove the patient transport Shuttle from operation.2. Trumpf Medical, or a designated service representative, will schedule an appointment with thecustomer to perform service on the impacted device to prevent the issue from occurring.Passing along this information:Please make sure that, in your organization, all users of the patient transport Shuttle as well as any otherpersonnel who must be informed, have been made aware of this safety information.If you have provided the device to third parties, please forward them a copy of this letter, or inform thecontact person listed above.Please store this information at least until this measure has been completed.If you have any questions concerning this request or the procedure as outlined, please contact Trumpf Medical Systems, Inc. at 888-474-9359For questions regarding this recall call 843-534-0606. |  | 1284 | 1284 affected units | Worldwide Distribution - US including MA and TX; Foreign distribution to Austria, Azerbaijan, Belgium, British Virgin Islands, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Japan, Luxemburg, Morocco, Netherlands, New Caledonia, Norway, Poland, Portugal, Qatar, Romania, San Marino, Spain, Sweden and Switzerland. | Open, Classified |  |  | 1284 | Not_Computer
Z-0764-2016 | 72571 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura Xper FD20 Ceiling version only; 722028; imaging applications. |  |  | November 03, 2015 | 2015 | Philips Electronics North America Corporation | Normally, movement of the C-arc is initially stopped by the motor (holding torque) then by the electromagnetic brake, leading to a full stop of the C-arc. Because of improper adjustment and tolerances, when the C-arc is initially stopped in an unbalanced position, the air gap of the brake can become too large, resulting in continued C-arc movement (C-arc is looking for balanced position). | Device Design | The letters for the customers in the United States were sent on 11/3/2015, via certified return receipt and were addressed to the Risk Manager.Philips will perform a test on the 15 affected Allura Xper R8.2 FD 20C systems to see if the motion controller amplifier of the brake motor works correctly. If the test will not pass, the motor will be replaced. The Recall Package is sent to all Key Markets worldwide. The Key Market of an affected country will contact the facility where these systems were installed. |  | 15 | 15 subject to correction and removal | Worldwide Distribution. US states of TN and IN and the countries of Australia, Belgium, Egypt, France, Germany, Japan, Republic of Korea, Netherlands, and Saudi Arabia. | Open, Classified |  |  | 17 | Not_Computer
Z-0762-2016 | 72798 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | VS300 radiographic Cassette Holder or Wall Stand. For use in radiology department. |  |  | December 03, 2015 | 2015 | Del Medical, Inc. | The wall stand cassette holder can suddenly fall to the bottom of its travel, near the ground and contact patient. | Employee error | The firm sent an email followed by a certified undated letter entitled Safety Related Notification- Field Corrective Action VS-200 and VS-300 Wallstands notifying customers of the recall and actions to be taken.To prevent the possible interruption of service, Del Medical has prepared a service kit which must be installed at the next available opportunity.Del Medical is sending service kits to affected customers. Included within the package will be needed components and, a fixture which will facilitate a one person job. The firm expects the procedure to occupy one hour or, less. Also included with each kit will be an acknowledgement form which is intended to document the fix. The form, along with the original parts are to be returned to Del Medical upon completion of the activity to the following address:Del Medical 241 Covington Drive Bloomingdale, Illinois 60108 Attention: Bill Kostecki 1-847-288-7000 |  | 4000 | 4000 | Worldwide Distribution-US (nationwide) and the countries of Canada, Costa Rica, Libya, Mexico, and Philippines. | Open, Classified |  |  | 4000 | Not_Computer
Z-0758-2016 | 73098 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | 2.3mm Turbo-Power OTW, Manual, Model No 423-050 |  |  | January 15, 2016 | 2016 | Spectranetics Corp. | The product is labeled with an expiration date that is past its shelf life. | Error in labeling | Spectranectics sent a Medical Device Field Action letter dated January 15, 2016 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and set aside the affected product and contact Customer Service at 1-800-231-0978 to arrange a product exchange. Customers were asked to document the return on the enclosed Field Action Acknowledgement and Receipt Form. Customers with additional questions were instructed to contact their local Spectranetics sales representative.For questions regarding this recall call 719-447-2318. |  | 8 | 8 units | Nationwide distribution to Florida, North Carolina and Pennsylvania. | Open, Classified |  |  | 8 | Not_Computer
Z-0760-2016 | 72756 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun 5000 and Arctic Sun 2000; Intended for monitoring and controlling patient temperatures. |  |  | November 10, 2015 | 2015 | C.R. Bard, Inc. | Improper targeted temperature therapy was delivered to patients because the Temperature Simulator Key was connected to the patient temperature cable instead of the patient temperature probe. | Use error | Bard Medical Division sent a Safety Alert to each of the customers via FedEx with proof of delivery notice. All consignees have been instructed to carry the notification out to the user-level. |  | 1539 | 1,539 | Nationwide Distribution-DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI. IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | April 21, 2017 | 528 | 1539 | Not_Computer
Z-0768-2016 | 72974 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Mobile Digital Ready Radiographic System |  |  | December 23, 2015 | 2015 | GE Medical Systems, LLC | Potential safety latch mechanism failure. The safety latch, within the column that supports the horizontal arm, is a backup in the unlikely event the primary safety cable breaks. A fall of the arm assembly could result in bodily harm to a person. | Device Design | Consignees were sent on 12/23/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 10906 dated December 18, 2015. The letter was addressed to Director of Biomedical Engineering, Director of Radiology and Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers may continue to use the system. The firm will correct the affected devices and no charge to the customer; and a GE Healthcare representative will contact customers to arrange for the correction. For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2389 | 2389 | Worldwide Distribution -- US, including Washington D.C. Puerto Rico, Virgin Islands, and Guam; and the countries of Afghanistan, Albania, Anguilla, Antigua and Barbuda, Armenia, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea (Republic of), Kuwait, Lebanon, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Papua New Guinea, Paraguay, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Taiwan, Thailand, Trinidad , Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 3410 | Not_Computer
Z-0767-2016 | 72210 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 Radiation therapy system |  |  | November 18, 2015 | 2015 | Mevion Medical Systems, Inc. | The treatment table positioning belt can slip gear positions, resulting in deviation of patient setup and treatment positions from the time of slippage until the discovery of slippage and the halting of treatment. | Labeling design | Customer notification letter was sent via certified mail on November 18, 2015. The notification explains the issue and actions to be taken by the customer. |  | 4 | 4 | US nationwide distribution including FL, NJ, OK, and MO. | Terminated | April 03, 2017 | 502 | 4 | Not_Computer
Z-0786-2016 | 73022 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Bard Distaflo Vascular Bypass Graft; Product Code: DF8006SC; 6 mm x 80 cm. |  |  | January 05, 2016 | 2016 | Bard Peripheral Vascular Inc | Size mislabeling; the graft is a 6 mm diameter graft when the actual packaged product is an 8 mm graft. | Error in labeling | Customer notification letters were sent on 01/05/2016. Customers were asked to not use the affected product, or further distribute it. Customers were asked to check all inventory locations for the affected product code, lot number combination listed. If customers further distributed any of the product they are requested to contact that location, advise them of the recall, and forward the listed instructions. The letter states to remove all identified product from the shelves. If no affected product is going to be returned the firm states to complete and return the attached form stating so. If customers have product it is requested that they also complete the attached form, and state the quantities and lot numbers in stock. The letter requests that customers call the firm's Recall Coordinator, Raye Seisinger, at 1-800-321-4254, option #2 ext. 2501 (M-F 6am to 3pm MST) or email raye.seisinger@crbard.com. Once all information has been verified the customer will be issued a Return Authorization or Consignment Recall Number to facilitate the expedient return of the product. The firm will provide replace product for returned product.Fax the attached form to the firm at 1-800-994-6772. If the customer cannot fax the firm they request that the Recall Coordinator be contact and provided the information verbally.The letter states a mailing label is enclosed to return the affected product. Customers are requested to amek the package as "recalled product" and include the XC or XH number. All products should be returned to the following shipping address:Bard Peripheral Vascular, Inc.1415 W. 3rd St.Tempe, AZ 85281 |  | 4 | 4 units | Distributed in the states of AL, TX, and CA. | Terminated | March 25, 2016 | 80 | 4 | Not_Computer
Z-0771-2016 | 73193 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic; Catalog/part number 032000 and 033000 Cosmetic cover package. |  |  | January 22, 2016 | 2016 | Accuray Incorporated | The gun box mounting bracket may fail to support the weight of the gun box when in vertical (inverted) position. If this failure occurs the gun box may become loose and could come into contact with a patient. | Device Design | Accuray sent an Urgent Safety Advisory Notification letters dated January 22, 2016 to all affected customers by trackable mail method. The letters identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Acknowledgement form. For questions contact Customer Support at 1-877-668-8667 or visit the Accuray Technical Solution Center at http://www.accuray.com/services-support |  | 226 | 226 | Worldwide Distribution. US Nationwide, Venezuela, United Kingdom, Ukraine, Turkey, Thailand, Taiwan, Switzerland, Spain, Saudi Arabia, Russia, Poland, Pakistan, Myanmar, Mexico, Korea, Japan, Italy, Ireland, India, Greece, Germany, France, Finland, Czech Republic, Colombia, China, Canada, and Belgium. | Terminated | October 20, 2016 | 272 | 226 | Not_Computer
Z-0792-2016 | 72937 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho OSI Trios, Table Base; |  |  | December 14, 2015 | 2015 | Mizuho OSI | Table Bases may have been assembled incorrectly, missing a retaining ring from the head end assembly. Over time, the parts could go out of alignment, making clutch engagement or disengagement difficult. If the clutch cannot disengage, the user would be unable to rotate the patient from supine to prone. | Employee error | All consignees were notified by telephone and a recall letter was mailed to each facility as a follow up. Customers were asked to remove tables from operation until they were inspected by OSI service personnel. |  | 33 | 33 | US, South Africa, Japan and United Arab Emirates. | Terminated | February 10, 2016 | 58 | 33 | Not_Computer
Z-0788-2016 | 73141 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Endobon-Xenograft Granules, Zimmer Biomet, Synthetic Bone Grafting Material, used in dental and/or surgical procedures. |  |  | January 21, 2016 | 2016 | Biomet France SARL | Cytotoxicity tests were not passed at 36 months. Use of Endobon-Xenograft Granules past 18 months of its specified shelf life may result in patient illness/injury. | Under Investigation by firm | Zimmer BIOMET France SARL has initiated a voluntary recall via certified mail on January 2, /2016 for the EndobonÂż Xenograft Granules, based on limited (in vitro/bench) results which indicated cytotoxicity tests were not passed at 36 months (real-time) aging. Use of the EndobonÂż Xenograft Granules may result in patient may experiencing soft tissue irritation, up to and including infection. Customers area asked to do the following: 1. Please review this notice.2. Identify, remove from service and segregate all affected Endobon ROX* product in your inventory.3. Complete the attached Business Reply Form (Attachment 1) and:a. Fax it to 877-650-0368 or email it to Biomet5351@stericycle.com.4. If you have product to return, contact BIOMET 3i at the number on the Business Reply Form for an RMA number and return shipping label.a. Return a copy of the Business Reply Form along with any affected product to the address on the Business Reply Form.Replacement product will immediately ship based upon availability.For assistance or other questions that customers may have relative to this notice, please contact BIOMET 3i at 1-800-443-8166 or 1-561-776-6700. Team members are available to assist you 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 145116 | 145,116 units | Worldwide Distribution - US including AL, AK, AR, AK, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MN, MI, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV and Internationally to Canada, Japan, Europe. | Terminated | May 31, 2017 | 496 | 145116 | Not_Computer
Z-0789-2016 | 73091 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Reuseable rotary blade/abrader, product number 899751305. |  |  | November 25, 2015 | 2015 | Richard Wolf Medical Instruments Corp. | Weak weld in the reusable rotary blade/abrader (burr) may result in inner part breaks at the distal or proximal end making the product dysfunctional. | Process control | The firm sent an URGENT Medical Device Recall letter dated 11/25/2015 to consignees informing them of the recall and the actions needed to be taken. These included to not use the affected burr and to destroy the device. The user is also instructed to fill out the Recall Action Response Form to acknowledge receipt of the notification. The form should be scanned to Dsantostefano@richardwolfusa.com or mailed to:Dawn Santostefano, Recall CoordinatorRichard Wolf Medical instruments353 Corporate Woods ParkwayVernon Hill, IL 60061 |  | 10 | Ten units | Distributed in Washington. | Terminated | August 21, 2017 | 635 | 16 | Not_Computer
Z-0787-2016 | 73076 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SIGMA HP MBT Non-Keel Punch Knee Instrument. |  |  | January 08, 2016 | 2016 | DePuy Orthopaedics, Inc. | Increased punch height, causing a 0.063 in (1.6 mm) gap between the trial and the instrument. This may cause the surgeon to select an insert that is too thin. Surgical delay may result if the punch dislodges from the inserter and needs to be removed from the joint space. Insert Dislocation or Spin Out, or Poor Joint Mechanics may result, if not noted during surgery. May require revision surgery. | Component design/selection | On 1/8/2016, URGENT INFORMATION  RECALL NOTICE for Specific Lots of SIGMAÂż High Performance (HP) MBT Non-Keel Punch Knee Instrument notifications were sent to the affected distributors and Medical Professionals with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For clinical questions from surgeons, please contact DePuy Orthopaedics, Inc.s Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m.  5 p.m. EDT). For questions about the device recall information provided, please contact Kim Earle, Senior Recall Coordinator, at 574-371-4917 (M-F; 8 a.m.  5 p.m. EDT). |  | 351 | 351 | Worldwide Distribution-US Nationwide, Armenia, Australia, Austria, Belgium, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hungary, India, Ireland, Israel, Italy, Japan, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, P R China, Portugal, Poland, Russia, Singapore, Slovenia, South Africa, Spain, Sweden, South Korea, Switzerland, Thailand, Turkey, UAE, UK, and Brazil. | Terminated | March 24, 2017 | 441 | 351 | Not_Computer
Z-0772-2016 | 73013 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | IMPLANT,TM, MTX |  |  | January 06, 2016 | 2016 | Zimmer Dental Inc | Information specific to the 3.7 mmD and 4.1 mmD Trabecular Metal Zimmer Dental Implants was omitted from the latest revision of the applicable IFU in the Indications, Contraindications, and Precautions sections. | Labeling Change Control | Zimmer sent an Urgent Device Correction letter dated January 6, 2016 to all affected customers.. The letter identified the affected product, problem and actions to be taken. Customers were asked to do the following; review the notice and check inventory for the affected units, ff they wish to keep the product familiarize themselves with the Technical Bulletin information, if they wish to return the affected product call Zimmer Dental Customer Service at 1-800-854-7019 to obtain a Return Authorization Number and to setup a FedEx pick-up to return the products, and complete Attachment 1, the Business Reply Form.It was requested that customers maintain a copy of the noticeand a signed copy of Attachment 1 for your records. For assistance or any other questions contact Zimmer Dental Customer Service at 1-800-854-7019. Customer Service is available 7:00am to 5:00pm (Pacific), Monday through Friday. |  | 1059 | 1059 | Worldwide Distribution. US Nationwide in the states of NY, CA, KS, NJ, CO, MI, MD, WA, CT, KY, OR, OH, IA, SC, AZ, AK, IN, TN, WI, FL, AL, UT, WY, MA, PA, MN, IL, MI, GA, ND, RI, including Puerto Rico; and the countries of Thailand, France, Germany, Colombia, Hungary, Netherlands, Saudi Arabia, Poland, Spain, and Italy | Open, Classified |  |  | 11631 | Not_Computer
Z-0752-2016 | 73078 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | INFX-8000V Bi-Plane X-Ray Interventional System |  |  | August 14, 2015 | 2015 | Toshiba American Medical Systems Inc | When a fontal DA (Digital Angiography) acquisition was done, scattered x-ray came into the dose meter on the lateral side. As a result, the dose meter sent a minus value to the software. Consequently, the software defined the data as an "abnormal value" and it stopped displaying dose data and the dose data was lost. The following message was displayed, "Dose meter abnormal, Dose info disabled". | Radiation Control for Health and Safety Act | Toshiba's planned action to bring into compliance:1.Affected customers will be notified using the Customer NotificationLetter you provided,2. revised system software will be installed on the systems to prevent this occurrence from happening,3. the corrections will be made free of charge, and4. the corrections will be completed by April 4, 2016.For further questions, please call 1- 800 521-1968. |  | 4 | 4 | US Nationwide Distribution to OH and NY | Open, Classified |  |  | 4 | Computer
Z-0793-2016 | 73084 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew BEAVER Blade 4.0mm |  |  | January 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | Blade exhibiting fracture during hip arthroscopy procedures | Device Design | Smith &amp; Nephew issued letters via Federal Express on January 11, 2016. requested to inspect inventory and locate any unused devices and quarantine them immediately.Complete Inventory Return Certification Form on the following page, indicating the quantities that need to be returned.Contact Smith &amp; Nephews Field Action Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number. |  | 69043 | 69043 units | Worldwide distribution. US Nationwide and countries of AT, AU, BE, CA, CH, CL, CN, CZ, DB, DE, DK, EE, ES, FI, FR, GB, HK, IE, IL, IN, IT, JP, LT, LV, NL, NO, PL, PR, PT, RO, SE, SG, SI, TH, TR, and ZA. | Open, Classified |  |  | 69043 | Not_Computer
Z-0796-2016 | 73130 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | T-piece Nebulizer with in-line connectors, Part number 8915-7-50. Single patient use, non-sterile prescription device designed for use in a hospital or home care environment. |  |  | January 05, 2016 | 2016 | Salter Labs | Reports that the T-piece does not fit on the nebulizer. | Nonconforming Material/Component | A customer letter dated 1/13/16 was sent to all customers who purchased the T-piece Nebulizer with in-line connectors because of reports that the T-piece does not fit on the nebulizer. The letter informs the customers that there has been no reports of interruption or delay in therapy. Customers are instructed to contact Salter Labs Customer Service at 1-800-421-0024 to arrange for thereturn of the product and replacement inventory. Customers are instructed to acknowledge receipt of the communication by completing the attached reply form and faxing it to the number indicated on the attached form. Customers are instructed to contact Salter Labs Customer Support at 800-421-0024 or your local Salter Labs representative to document any issues that you have encountered or to request replacement inventory. |  | 70 | 70 units | Distributed in the states of TX, NY, PA, TN, SC, and OH, and in China. | Terminated | April 21, 2016 | 107 | 70 | Not_Computer
Z-0794-2016 | 72978 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite, a picture archive and communications system. Model Numbers Versions 5.30.0, 5.30.1, 5.30.2, 5.30.2 SB0344, 5.30.3, 5.30.4, 5.30.4 HF0418E, 5.30.5, 5.30.5 HF0404E, 5.30.5 HF0413E, 5.30.5 HF390E, 5.30.5 SB0368, 5.30.6, 5.30.6 HF0405E, 5.30.6 HF0406E, 5.30.7, 5.30.7 HF0429E, 5.30.7 HF0441E, 5.30.7 HF0446E, 5.30.7 HF0470E, 5.30.7 HF0473E, 5.30.8, 5.30.8 HF0477E, 5.30.8 HF0490E, 5.30.8 HF0495 SB0508E, 5.30.8 HF0495E, 5.30.8 HF0495E SB504E, 5.30.8 HF0550, 5.30.8 SB0485E, 5.30.8 SB0486E, 5.30.X SB0381E, 5.35.1, 5.35.1 HF0487E, |  |  | December 17, 2015 | 2015 | Merge Healthcare, Inc. | Potential incorrect Standardized Uptake Values (SUV) measurements in RadSuite. | Software design | Merge Healthcare sent an " Urgent - Medical Device Recall letter dated January 21, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer .If you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter. Merge Healthcare is committed to improve efficiencies and enhancing the quality of healthcare worldwide. If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 164 | 164 | Nationwide Distribution to the states of : MN, WI, IL, PA, AL, MO, MT, AZ, CA, OH, FL, SC, MI, TN, CO, NV, IN, NY, MD, AR, KS, NY, IN, IA, TX, OK, NC, WA and DC. | Open, Classified |  |  | 164 | Not_Computer
Z-0795-2016 | 73053 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | PharmaCaribe PulmoSal 7% Sodium chloride, Inhalation Solution, 7% Sterile, Preservative -free, Non-pyrogenic, pH 7.4 |  |  | December 29, 2015 | 2015 | Pharmacaribe llc | Vials were labeled as USP 7% Hypertonic saline instead of PulmoSal 7% (pH+) | Error in labeling | PharmaCaribe, together with its Contract Manufacturing firm Asept Pak, Inc., have initiated or completed the following actions to issued an Urgent Notice: Field Notification to Medical Device letter dated January t7, 2016. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Determine what is on hand and what has been distributed.Using the Return Notice provided, report this information as per the instructions on the notice. Please record units as either cases or as cartons (if the cases have been opened and are partial). This document is to be returned via mail or email as per the guidance provided. This will inform the Manufacturer of your inventory status and of your intentions regarding product replacement. Cases of this product may be returned for new cases of the same product (different lot) that have the correct embossing on the vials. Instructions for return are provided in the Field Notification Package Cover Letter. Alternatively, customers may choose to be issued a full refund for all returned product. For cases which have already been distributed to the retail/clinician level and/or are in use by patients, it is acceptable to continue to use the product, with the understanding (and proper communication to all parties) that the tabs on the vials are incorrectly labeled.Contact information:Lisa Early, QA Manager Asept Pak, Inc.Phone: 518-651-2026 Fax: 518-651-2046 E-mail: learly@aseptpak.com |  | 7020 | 7,020 vials | US Nationwide Distribution in the states of FL, LA, TX and VA | Terminated | March 28, 2017 | 455 | 7020 | Not_Computer
Z-0797-2016 | 73069 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Firebird Spinal Fixation System Torque Limiting Handle (PN 52-1512) |  |  | January 08, 2016 | 2016 | Orthofix, Inc | There is a possibility that the device may deliver a torque value less than the required setting after initial sterilization. In worst case, the unexpected drop in torque value may lead to the spinal construct loosening and requiring corrective surgical intervention. | Component design/selection | Affected consignees were notified via an Urgent Medical Device Recall Notification letter, dated 1/8/16. The letter identified the affected device and stated the reason for the recall. The recalling firm requests customers to remove the affected devices from their inventory and cease further distribution or use. Customers whom the devices were further distributed to should be notified that the affected devices cannot be used and must be returned to Orthofix per the instructions provided. The attached Acknowledgement Form should be completed and returned. Customers can contact their local Orthofix representative or customer service for further information, replacements, or disposal instructions. |  | 36 | 36 units | Worldwide Distribution -- US, Spain, Italy, and Australia. | Terminated | November 16, 2016 | 313 | 36 | Not_Computer
Z-0688-2016 | 72923 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Moog Curlin Infusion, Ambulatory Infusion Pump |  |  | December 23, 2015 | 2015 | MOOG Medical Devices Group | Moog is recalling Curlin 4000, 6000, and PainSmart Infusion Pumps manufactured or serviced by MMDG between March 18, 2015 and November 6, 2015 because the master pumps used to calibrate them were out of tolerance and may cause an over delivery of medications and solutions at a rate of 1.2 - 6.8% faster than programmed. | Process control | Moog Medical Device Group sent an "Urgent Recall Notice" dated December 22, 2015, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove all of their current Calibration Sets from service before placing these new sets into service. Customers were also instructed to complete and return the Recall Reply Form by fax: 801-264-1051 or email to: MoogMedicalSupport@moog.com and for any questions, please contact: Moog Regulatory Affairs at 801-264-1001. For additional information or clarification call 801-264-1001, ext. 112 or via email at: cdodge@moog.com Monday through Friday, 8:00 AM to 5:00 PM, Mountain Time. |  | 11472 | 11,472 | Worldwide Distribution - US (nationwide), Guam, and countries of: Australia, Canada, Croatia, Turkey, New Zealand, Saudi Arabia, Japan, Korea, and Italy. | Open, Classified |  |  | 11472 | Not_Computer
Z-0805-2016 | 73092 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System versions 11, 13.0, 13.5 and 13.6; |  |  | January 11, 2016 | 2016 | Varian Medical Systems, Inc. | When using PBC 11.0.31 to calculate the dose for a conventional arc field with more than 100 segments for Eclipse versions 11.0, 13.0, 13.5 or 13.6, the displayed dose does not correspond to the calculated Monitor Units (MU). Potential for unintended radiation exposure. | Software design | Varian sent an Urgent Medical Device Correction letters dated January 11, 2016 to all affected customers. Letters were also sent to Varian Sales and Marketing and Service organizations to notify them of the recall. The letter identified the affected product, problem and actions to be taken. For questions customers were instructed to contact their local Varian Medical Systems Customer Support District or Regional Manager. |  | 9499 | 9499 | Worldwide Distribution - US Nationwide and the countries of Albania, Guadeloupe, Norway,Algeria, Guam, Pakistan, Argentina, Guatemala, Panama,Australia, Hong Kong, Paraguay, Austria, Hungary, Peru,Azerbaijan, Iceland, Philippines, Bangladesh, India, Poland,Belgium, Indonesia ,Portugal, Belorussia, Iran, Puerto Rico,Brazil, Iraq, Qatar, Brunei Darussalam, Ireland, Reunion,Bulgaria, Israel, Romania, Burma, Italy, Russia,'Canada, Jamaica, Saudi Arabia, Chile, Japan, Singapore, China, Kazakhstan, Slovakia, Colombia, Kuwait, South Africa, Costa Rica, Latvia, South Korea,Czech Republic, Lebanon, Spain, Denmark, Libya, Sweden, Dominican Republic, Lithuania, Switzerland, Ecuador, Luxembourg, Taiwan, Egypt, Macau, Thailand, El Salvador, Macedonia, Trinidad and Tobago, Estonia, Madagascar, Tunisia, Finland, Malaysia, Turkey, France, Malta, Ukraine, French Polynesia, Mexico, United Arab Emirates, Georgia, Monaco, United Kingdom, Germany, Morocco, Venezuela, Ghana, Netherlands, Vietnam,Greece, New Zealand, and Zimbabwe. | Open, Classified |  |  | 9499 | Computer
Z-0806-2016 | 73086 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdian ViewRay Radiation Therapy System, VIEWRAY Inc. |  |  | December 24, 2015 | 2015 | Viewray Incorporated | The system loaded a completion fraction in the incorrect order after a treatment interruption. | Software design | ViewRay sent an Customer Advisory Notice dated December 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Workaround: If you want to move a completion plan to a later date in the Delivery Calendar then do not schedule it until right before you are ready to treat that completion plan.While we will address this issue in the next release of software we recommend that in the event of a completion plan you always review the plan name loaded on the TPDS and verify the remaining time and beams are as expected with the completion plan.If you have any questions or concerns regarding this issue, please do not hesitate to contact ViewRay Customer Support at support@viewray.com or call the ViewRay" Customer Support center at 855-286-8875. |  | 5 | 5 units | US Distribution to the states of : CA, FL, MO &amp; WI., and Internationally to Republic of Korea. | Open, Classified |  |  | 5 | Computer
Z-0808-2016 | 72802 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Bond Polymer Refine Red Detection DS9390 kits. The common name of the device is detection system for in vitro diagnostic use. The Intended Use states that Bond Polymer Refine Red Detection is a biotin-free, polymeric alkaline phosphatase (AP)-linker antibody conjugate system for the detection of tissue-bound mouse and rabbit IgG and some mouse IgM primary antibodies. It is intended for staining sections of formalin-fixed, paraffin-embedded tissue on the BOND automated system. |  |  | November 27, 2015 | 2015 | Leica Microsystems, Inc. | The Firm informed the importer/distributor, Leica Microsystems, a manufacturing error occurred during the filling of some containers of the Bond Polymer Refine Red Detection DS9390 kits, which may result in inadequate staining. | Mixed-up of materials/components | Leica Biosystems sent an Urgent Medical Device Recall Notice dated November 30, 2015, to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Do not use or continue to use the reagent lots listed above, as they have the potential to result in inadequate tissue staining and lead to weak or false negative pathology results. Please also properly destroy any unused or partially used affected lots of the reagent, and indicate by signing and returning the attached Medical Device Recall Acknowledgement Form that this action has been undertaken. As indicated in the Instructions for Use, the clinical interpretation of any staining or its absence should be complemented by morphological studies using proper controls and should be evaluated within the context of the patients clinical history and other diagnostic tests by a qualified pathologist. The consignees are requested to consult with their laboratory director and physician to determine whether a review of past results obtained from the affected lots is clinically warranted.Direct accounts are requested to forward the notice primarily to the end users where the product has been sold and to all those within the organization who need to be aware of this issue.Periodic effectiveness checks will be conducted to address non responders.For further questions please call ( 847) 405-5413 |  | 3745 | 3745 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Argentina, Australia, Austria, Belgium, Chile, Denmark, France, Germany, Greece, Hong Kong, Italy, Japan , Netherlands, RomaniaRussia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey and UK. | Terminated | January 05, 2017 | 405 | 3745 | Not_Computer
Z-0809-2016 | 73057 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Exactech Acapella One Interbody Fusion Device Parallel. Orthopedic Spinal Surgery Device. |  |  | September 28, 2015 | 2015 | Exactech, Inc. | Incorrect external and internal packaging labels. Exactech has identified nineteen (19) Acapella One Interbody Fusion Device Lordotic (05-090-01-1610) were packaged as Acapella One Interbody Fusion Device Parallel (05-090-02-1610) | Process control | Exactech sent a recall notification letter dated September 28, 2016 to their only 2 consignees. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately cease distribution or use of these products, extend this information to accounts that have this products in their possession, identify and quarentine any of the product and complete and return the attached recall inventory response form to Exactech. Contact: Kaya Davis at Kaya@exac.com or 1-800-392-2832 |  | 19 | 19 pieces | US distribution in the state of FL and PA. | Open, Classified |  |  | 19 | Not_Computer
Z-0807-2016 | 73030 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EVOLIS Microplate System, Catalog # 89601. Part number 89788 for the EVOLIS Operator's Manual. in vitro diagnostic |  |  | December 22, 2015 | 2015 | Bio-Rad Laboratories, Inc. | The error "Washer Reagent Clean Fluid" level low alarmed, which prompted the customer to open the system drawer. When the system drawer was opened the customer noted the plate transport was still running and as a consequence the plate was pushed off the plate transporter. | Use error | Bio-Rad sent an Urgent Product Notification letter, dated December 22, 2015, to consignees via Fed-Ex. The firm recognized the typographical error of the PN number on the first letter and sent a revised notification letter, dated 20 January 2016, with the correct PN 89788 to consignees. The letter identified the affected product, problem and actions to be taken. Consignees are advised that in addition to the WARNING in Chapter 4.9.2 in the EVOLIS Operators Manual (PN 89788) mention is made of, "LED lamps on the front panel of the EVOLIS instrument illuminate when a compartment may be accessed. If an error occurs, (for example a bulk reagent becomes depleted) the system displays an error message, pauses, and allows access to the instrumentcompartments."The plate transport may still be moving the microplate to its final destination when access is allowed. To enhance safety features, your Field Safety Engineer (FSE) will apply a (yellow) warning label to the lower front panel. Labels are currently on order and once received will be provided to each FSE. The label will state: WARNING! Before opening instrument, ensure all moving parts have stopped".Consignees can contact Bio-Rad Technical Support directly at 1-800-224-6723, option #2, and then option #3 or by fax at 425-498-1628. |  | 278 | 278 units | US Nationwide Distribution including Puerto Rico. | Terminated | March 13, 2017 | 447 | 278 | Not_Computer
Z-0810-2016 | 71647 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Actuator for angiographic x-ray monitor ceiling suspension system. Actuator short Assy: THC8AWDS-121 and Actuator long Assy: THC8AWDS-122 used with the following products: 722001 Allura Xper FD 10 C; 722003 Allura Xper FD10; 722005 Allura Xper FD10/10; 722006 Allura Xper FD20; 722008 Allura Xper FD20 Biplane; 722010 Allura Xper FD10; 722011 Allura Xper FD10/10; 722012 Allura Xper FD20; 722013 Allura XPER FD20 BIPLANE; 722015 Allura Xper FD20 OR Table; 722020 Allura Xper FD20 Biplane OR Table; 722022 Allura Xper FD10 OR Table; 722023 Allura Xper FD20 OR Table; 722025 Allura Xper FD20 Biplane OR Table; 722026 Allura Xper FD10; 722027 Allura Xper FD10/10; 722028 Allura Xper FD20; 722029 Allura Xper FD20/10; 722035 Allura Xper FD20 OR Table; 722038 Allura Xper FD20/20; 722039 Allura Xper FD20/20 OR Table; 722058 Allura Xper FD20/15; 722123 Field extensions Xper cardio R7.6; 722124 Field extensions Xper vascular R7.6; 722133 Field ext. Xper cardio systems R7.6; 722134 Field ext. Xper vascular systems R7.6; 72249 INTEGRIS BV5000; 722043 INTEGRIS Allura 15-12 (mono); 722121 Field extensions INTEGRIS cardio systems |  |  | July 06, 2015 | 2015 | Philips Electronics North America Corporation | The actuator assembly became detached and the monitor carriage with a flexvision monitor dropped to its lowest position. | Process control | Philips sent a Field Safety Notice to all affected customers on July 6, 2015. The notice informs consignees of the reason for recall and actions to be taken by the customer/user. On September 15, 2015, Philips plans to issued a mandatory Field Safety Corrective action (FCO72200280) via Certified Return Receipt. This corrective action is the replacement of the affected MCS actuator. For further information, contact your local representative: Technical Support Line: 1-800-722-9377. |  | 3158 | 3158 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of : Canada, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Cayman Islands, Chile, China, Colombia, Costa Rica, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Lebanon, Lithuania, Macedonia, Malaysia, Mexico, Mongolia, Morocco, Myanmar, Nepal, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Pakistan, occupied Palestinian territory, Panama, PERU, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom and Viet Nam. | Open, Classified |  |  | 3158 | Not_Computer
Z-0812-2016 | 72977 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | REF 0038460 Toomey Syringe 70cc Syringe with Catheter Tip and Luer Tip Adapters, Single Use, Rx Only, Latex Free, 50 units. |  |  | January 04, 2016 | 2016 | C.R. Bard, Inc. | Report by a customer of an open package seal. | Packaging process control | BMD sent an Urgent Medical Product Recall letter to each of the customers via FedEx with proof of delivery notice. The leter identified the affected product, problem and actions to be taken. All consignees were instructed to carry out the recall to the user level. The letter included a Recall and Effectiveness Check Form which was to be returned to Bard. |  | 15400 | 15,400 units | US Nationwide Distribution - in the states of: AL, AZ, CA, CO, FL, GA, IL, IN, IA, KS, KY, LA, MD, MI, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, PA, SC, TN, TX, VA, WA, WV, and Puerto Rico | Terminated | June 06, 2017 | 519 | 15400 | Not_Computer
Z-0814-2016 | 73093 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Jadak Barcode Scanner, Part No. 350587-32, 350587-41, 351377-32. |  |  | January 15, 2016 | 2016 | CareFusion 303, Inc. | Customers reported issues when scanning medications with the scanners. When a user scans a medication using the affected barcode scanner, information on a different medication could be communicated to the Pyxis product. | Device Design | BD sent an Urgent Medical Device Recall Notification letter dated January 15, 2015 to all affected customers The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete the enclosed Customer Response Card by selecting the desired deployment option and return the card to CareFusion. Customers with questions are instructed to contact the CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact CareFusion Customer Advocacy at (800) 836-7842. Customers with technical questions are instructed to contact CareFusion Technical Support at (800) 727-6102. |  | 9345 | 9,345 units | Worldwide:Distribution - US Nationwide and the countries of Saudi Arabia, Bahrain, Australia, United Arab Emirates, Qatar, Mexico, Guam and the Bahamas. | Terminated | May 08, 2017 | 479 | 9345 | Not_Computer
Z-0813-2016 | 73224 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Dual Stage Venous Return Catheter, Low-Profile, wire reinforced (36F/46F, Sterile), Model Number: LRD-61046, Expiration 10/31/2018, Product is packed in a sealed pouch, ten (10) pouches per case. Cardiopulmonary bypass catheter. |  |  | January 25, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group initiated field correction for the Dual Stage Venous Return Catheter because the incorrect Instructions For Use (IFU) was packaged in the case. | Employee error | LivaNova sent an Urgent Safety Alert dated January 25, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advice on action to be taken:"Remove and discard the IFU included within the case packaging for the product."A copy of the correct IFU (Part Number 435085202, Revision B) is attached with this letter"Print copies of the IFU and place with the product.For questions regarding this notice, please contact Customer Service at 1-800-650-2623 (Monday to Friday BAM to 5PM MT) or e-mail USFieldSafetyActions@livanova.com.We sincerely apologize for this inconvenience and thank you in advance for your support |  | 60 | 60 | US Distribution to CA., and Internationally to Canada. | Open, Classified |  |  | 60 | Not_Computer
Z-0816-2016 | 73164 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION(TM) HEMI STEPPED, TIBIAL SCREW-ON WEDGE, REF 71421147, QTY 1, 5 MM, SIZE 5-6 LT-LAT / RT-MDL; Tibial knee prosthesis, augment. |  |  | January 18, 2016 | 2016 | Smith &amp; Nephew, Inc. | The package for catalog item, a LEGION Hemi Stepped Tibial Wedge, 5 MM Size 5-6 LT-LAT/ RT-MDL, should include both a screw-on hemi stepped tibial wedge and two tibial wedge screws to affix the wedge to the tibial baseplate. A complaint was received indicating that the screws packaged with the wedge were too long. | Packaging change control | The firm notified their direct consignee of the issue on 01/18/2016 by email. The notice requested return of the product. |  | 8 | 8 units | Distributed in the states of TN, WI, OH, and KS, and in the countries of Belgium, Germany, and United Kingdom. | Open, Classified |  |  | 8 | Not_Computer
Z-0815-2016 | 73120 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | CEA / CD66e Ab-2, Product code RB-368-R7, RB-368-R1; 7 ml each. |  |  | January 18, 2016 | 2016 | Lab Vision Corporation | Certain lots of CEA/CD66e Ab2 were produced with a dilution factor of 1:100 instead of the recommended 1:250. | Nonconforming Material/Component | ThermoFisher sent an Urgent Medical Device recall letter dated January 18, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have any of the affected product to contact Lab Vision with the informationrequested on the attached Acknowledgment Form.Âż Upon identification of the affected lot, please notify Lab Vision of the quantity previously used and the quantity still in inventory.Âż If any of the above listed lots were previously used, please review test results to ensure a valid result was obtained. If a valid result was obtained, no further action is required. If high background staining was seen, standard IHC troubleshooting guidance for high background staining can be used.(See RISK TO HEALTH section of the letter.)Âż Affected product may be returned to Lab Vision. Please indicate on the Acknowledgement Form whether the product will be returned. Upon receipt of the form, RGA information will be sent for product requiring return.Âż If it is preferred to destroy the product at the customer site, please indicate this on the Acknowledgement form.Âż Please use the attached Recall Return Response Acknowledgement &amp; Receipt Form.Customers with questions were instructed to call 269-544-5628. |  | 776 | RB-368-R7 - 30; RB-368-R1 -2 | Worldwide Distribution - US including CA, NY, GA, MS, PA, MA, TX, KY, AZ. and Internationally to Greece, France, Germany, Singapore and Serbia. | Open, Classified |  |  | 776 | Not_Computer
Z-0817-2016 | 73114 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | MEE-1000A Neuromaster, Model No. Accessories: MS-110B (Electric Stimulation), MS-120BK (Current Stimulator Box) JB-116B, JB-132B (Amplifiers) Electromyography. (EMG) |  |  | January 19, 2012 | 2012 | Nihon Kohden America Inc | The MEE-1000A Neuromaster had defective printed circuit boards (PCBs). Gas arrestors on these boards which were intended to provide protection from electrostatic discharge failed to prohibit unintentional current flow. | Device Design | A customer notification letter dated 9/25/12 was sent to customers informing them of a potential malfunction of NIHON KOHDEN amplifier units, models JB-116B andJB-132B and Extension Unit MS-110B, used with the MEE-1000A Neuromaster Neural Function Measuring System. The letter informs the customers of the potential malfunction and that the probability that the malfunction occurs is extremely rare. The letter informs the customers of the preventive measure and the preventive corrective action. Customers are instructed to inform all potential users and operators of the Neural Function Measuring System MEE-1000A Neuromaster about the Corrective Action notice. Customers are instructed that NIHON KOHDEN technical support representative will contact them to arrange an update of their amplifier products. |  | 477 | 477 units total (137 units in US) | Worldwide distribution. US nationwide, Australia, Austria, France, Germany, Greece, Japan, Romania, South Africa, South Korea, Spain, and United Kingdom. | Terminated | February 22, 2016 | 1495 | 477 | Not_Computer
Z-0818-2016 | 72911 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | High Flow Irrigation Extender; Part Number 72201021. |  |  | July 30, 2013 | 2013 | Smith &amp; Nephew, Inc., Endoscopy Div. | During arthroscopic surgery, when the scope is moved inside the cannula, the body of the irrigation extender separates, potentially obstructing visualization and causing leakage of the irrigation fluid. The device is removed and a back-up device is used to complete the surgery. | Under Investigation by firm | Urgent - Product Recall 1st Notification - Urgent R-2013-13 letters, dated July 30, 2013 were sent overnight via Federal Express to customers. The letters identified the affected product, as well as the potential risk involved. Actions to be taken were provided for both hospital and sales representatives. Hospital representatives were asked to inspect their inventory, locate all devices from the list provided, and quarantine them immediately. The Inventory Return Certification Form should be completed and returned, regardless of whether any affected product is on hand or not. Contact the Smith &amp; Nephew Returns Group at 800-343-5717 option 3) or send an e-mail to ProductRecovery@smith-nephew.com. They will provide instructions for returning affected product and receiving a replacement. |  | 203 | 203 units | Worldwide Distribution -- US, including the states of AL, AR, CA, FL, GA, ID, IL, KS, MA, MI, MO, NC, TN, TX, and AK; and, country of Japan. | Open, Classified |  |  | 203 | Not_Computer
Z-0819-2016 | 72749 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Endo GIA Ultra Universal 12 mm Short Stapler Handle |  |  | December 02, 2015 | 2015 | Medtronic | Staplers fail to fire or partially fire and reports of the instrument articulating lever disengaging during use | Process control | Medtronic sent an Urgent Medical Device Recall letters dated December 2 2015 and an Updated Urgent Medical Device Recall letter dated January 13, 2016, to affected customers, via Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473.For questions or concerns, please do not hesitate to contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 364768 | 364,768 units | Worldwide Distribution - US Nationwide and the countries of Canada, Albania, ALGERIA, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Gabon, Georgia, Germany, Greece, Greece, Guatemala, Honduras, Hungary, Iceland, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Mauritius, Mexico, Monaco, Montenegro, Morocco, Namibia, Netherlands, Nicaragua, Nigeria, Norway, Panama, Paraguay, Peru, Poland, Portugal, Romania, Russia, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, UAE, Ukraine, United Kingdom, Uruguay, Venezuela, and Zimbabwe. | Open, Classified |  |  | 1797442 | Not_Computer
Z-0828-2016 | 73180 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Aero-AL Impaction handle Catalog Number 48921007. Stryker Spine Manufactured by : Stryker Spine SAS France |  |  | December 18, 2015 | 2015 | Stryker Spine | Reported jamming of the impaction handle when assembled to the Inserter guide at the threads. | Nonconforming Material/Component | Stryker Spine sent Urgent Product Removal notification letters to branches/agencies on 12/18/2015 via FedEx. The letters provided a description of the affected product, the issue, potential hazards and risk mitigations, and actions to be taken. Customers were asked to examine their inventory and hospital locations to identify the product and reconcile with the affected catalog and lot numbers provided. The Customer Response Form and Product Quantity Tracking Spreadsheet for Health Care Facilities should be completed and returned within five days of receipt of the notice to the e-mail or fax number provided. Questions can be directed to Nabil Riaz - Regulatory Compliance Specialist, Regulatory Compliance &amp; Quality Systems at Nabil.Riaz@Stryker.com. |  | 31 | 31 units | Nationwide Distribution | Open, Classified |  |  | 31 | Not_Computer
Z-0803-2016 | 73251 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Delta XRF Analyzer . This is a Analytical X-ray system |  |  | January 13, 2016 | 2016 | Olympus Scientific Solutions Americas | The Firm has discovered a Software bug. | Radiation Control for Health and Safety Act | Olympus Scientific Solutions Americas ( OSSSA) planned action to bring product into compliance: Notification will be given to purchasers per 21 CFR 1003.21. A software modification will be made available to all users contacted. The letter to affected purchasers makes clear that the product will be brought into compliance without charge and offers the purchaser several options for applying the software update.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Tracking of customer/owner self-installation(s) of the software update should be done to the maximum possible extent via serial-number- controlled downloads, follow-up email reporting, return-mail card or some other similarly effective means. For further questions you may call : (866) 446-6689. |  | 5000 | US - 5000 units | Worldwide Distribution - USA (nationwide) Distribution | Open, Classified |  |  | 5000 | Computer
Z-0826-2016 | 73274 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | software for Syngo Dynamics |  |  | January 29, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a software update that addresses an issue of mixing data from multiple patients. In rare situations, echo trend graphs may mix data from multiple patients. | Software design | The firm, Siemens Healthcare, sent a "Customer Advisory Notice" dated 1/19/2019 to its customers.The notice described the product, problem and actions to be taken. The customers were instructed to create a new department if immediate resolution is necessary, use workplace for reporting and hand over the customer letter before or during installation of field modification(PSL) distributed with Update Instruction SY013/16P. A software update to resolve reported issues is being released via an Update Instruction SY013/16/P. If you have any questions, contact Regulatory Technical Specialist at 610-448-6478 or email: anastasia.sokolova@siemens.com. |  | 151 | 151 | US Nationwide Distribution including states of: AL AR AZ CA CO CT DE FL GA IL IN KS LA MA MD MI MN MO NC ND NE NJ NM NY OH OR PA SC TN TX UT VA WA and WI. | Open, Classified |  |  | 151 | Computer
Z-0824-2016 | 73099 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | BCS Automated Blood Coagulation Analyzer System, Multipurpose system for in vitro coagulation studies, Catalog Numbers: 10454728 (BCS RECONDITIONED), 10454729 (BCS RECONDITIONED), 10454742 (BEHRING COAGULATION SYSTEM), 10459303 (BCS RECONDITIONED), 10460659 (BCS INSTRUMENT), 10461881 (BCS INSTRUMENT), IVD --- Device Listing D018966 |  |  | January 14, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | False short clotting times for PT on BCS and BCS XP with Dade Innovin can occur. This has also an effect for false low PT INR values. These false short clotting times were observed only with turbid, hemolytic and/or icteric samples from intensive care patients. The corresponding reaction curves of these samples showed a pre-peak, which can lead to falsely short clotting time. The result was not flagged by the existing check algorithms. | Software design | Urgent Medical Device Correction Letter (dated 1/12/2016) was sent to US consignees via FedEx on 1/14/2016 and to foreign consignees (distribution is determined at country level). The letter re-emphasizes the importance of checking the interference levels given in the respective Application Sheets of the BCS and BCS XP systems. For questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 1680 | 1680 units total | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bosnia Herzegovina, Brazil, Canada, Chile, China, Croatia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Guadeloupe, Hungary, Italy, Kuwait, Latvia, Lebanon, Libya, Macedonia, Mexico, Montenegro, Netherlands, Peru, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, and Uruguay. | Terminated | March 16, 2017 | 427 | 1680 | Not_Computer
Z-0823-2016 | 73095 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | This reagent is for in vitro diagnostic use. Pax-5 (1EW) monoclonal antibody is intended to be used for the qualitative identification by light microscopy of human Pax-5 in formalin-fixed, paraffin-embedded tissue by immunohistochemical staining using the automated Bond system. |  |  | December 30, 2015 | 2015 | Leica Microsystems, Inc. | The Bond" Polymer Refine Detection and Novolink" Polymer Detection System may not provide adequate staining when using a detection protocol where one of its components - Peroxide Block / Peroxidase Block is applied after the primary marker incubation as specified in the Instructions for Use Following the Instructions for Use. | Component change control | The firm sent out a Medical Device Recall letter dated December 29th 2015 to two-hundred (200) consignees in the United States decribing the recall and the neccesary actions to be taken. These included: - Do not use or continue to use the listed antibody with the specified detection kits.- Order replacement detection kits for use with the listed antibody.- There is no need for retrospective review of tissue staining if appropriate positive controls were properly used and interpreted during the immuno / in-situ staining. -pass this notice primarily to the end users where the product has been sold and to all those within your organization who need to be aware of this issue.-Confirm reciept of the Recall Notice letter by signing and dating the attached Acknowledgement Form and faxing it back to us at 1-847-236-3747, or scan and email it to LMGRA@Leica-Microsystems.com Customer are requested to contact their local Leica representative if you have any questions or concerns, or alternatively please contact us at the following email address: Combination.Products@LeicaBiosystems.comPeter LloydLeica Biosystems Newcastle LtdBalliol Business Park WestBenton LaneNewcastle upon Tyne NE 12 8EWUnited KingdomTelephone +44 191 215 0567Facsimile: +44 191 215 1152 |  | 1700 | 1700 units | US Consignees incluude: AR ,AZ ,CA ,CO ,CT ,DC ,DE ,FL ,GA ,HI ,IA ,IL ,KY ,LA ,MA ,MD ,MI ,MN ,MO ,MT ,NC ,NH ,NJ ,NY ,OH ,OK ,PA ,PR ,SC ,TN ,TX ,UT ,VA ,VT ,WA ,WI and WV. Foreign Consignees include:Argentina , Australia , Brazil , Bulgaria , Chile , China , Ecuador ,Egypt , El Salvador , Finland , France , Georgia , Germany , Greece , India , Italy , Japan , Jordan , Kuwait , Lebanon , Malaysia , Philippines , Poland , Portugal , Romania , Russian Fed. , Singapore , South Korea , Spain , Sweden , Switzerland , Thailand , Tunisia , Turkey , United Kingdom and United Arab Emirates. | Terminated | January 05, 2017 | 372 | 1700 | Not_Computer
Z-0804-2016 | 73137 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Flow-i Anesthesia System |  |  | December 10, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Electrical issues in the control, monitoring or panel sub-systems, caused by the printed circuit boards (PCBs) that control the sub-systems. The problems manifested as multiple re-starts of the sub-systems and, in a few instances, disabling of the system has occurred. Replacement PCBs are now available to correct the issues.Only the Model C20 is affected in the US. | Nonconforming Material/Component | Maquet sent "Urgent Medical Device Field Action Field Correction For Flow-i Anesthesia Systems" letters/Field Correction Response forms dated 12/10/2015 via UPS to affected customers. The notification informed customers of the issue, asked them to distribute the notification to appropriate staff; verify the serial number(s) and informed them they could continue to use the system if they were not experiencing electrical issues. Customers were informed that a Maquet service representative would be contacting them to arrange a time to replace their affected units. Customers were asked to complete and return the enclosed Field Correction Response Form to Maquet at FieldActions@maquet.com or by fax to 973.860.1675.If the customer had any related questions they were instructed to contact Maquet Technical Support at 1.888.627.8383 (option 3; option 1; option 1) Monday through Friday between 8:00 am and 5:00 pm EST. |  | 7 | 7 units within US | US Nationwide in the states of CA, NY, and OR | Terminated | July 06, 2016 | 209 | 7 | Not_Computer
Z-0830-2016 | 73142 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS-FLEX GSF OPT |  |  | January 11, 2016 | 2016 | Zimmer Biomet, Inc. | LDPE bag containing the implant adheres to the highly polished implant surface. | Process control | Zimmer Biomet sent an Urgent Medical Device Recall dated January 11, 2016, to all affected customers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with questions or concerns were instructed to contact customer service at 1-800-348-2759 or contact their Zimmer Biomet Sales Representative.For questions regarding this recall call 800-447-5633. |  | 84396 | 84396 | Worldwide Distribution - US (nationwide) and Internationally to ARGENTINA AUSTRALIAAUSTRIABARBADOSBELGIUMBERMUDABOLIVIABRAZILCARIBBEANCHILECHINACOLOMBIACOST RICADENMARKDOMINICAN REPUBLICDORAL FLECUADOREL SALVADORENGLANDFRANCEGERMANYGUATEMALAHONDURASHONG KONGINDIAISRAELITALYJAMAICAJAPANKOREAMALAYSIAMEXICONEW ZEALANDNICARAGUAPANAMAPERUSINGAPORESPAINSWEDENSWITZERLANDTAIWANTHAILANDVENEZUELAWEST INDIESCanary IslandsCroatiaCyprusCzech RepublicCanada | Terminated | May 23, 2017 | 498 | 738203 | Not_Computer
Z-0863-2016 | 73104 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Symbiq infusion pump is a device used in a health care facility to pump fluids into a patient in a controlled manner. It is intended primarily for use in the hospital setting and can be used in other acute and non-acute areas, such as, but not limited to, Nursing Homes, Mobile Intensive Care, Ambulatory Infusion, Centers, Hospice, Sub acute facilities, Outpatient/Surgical Centers, Long-term Care, Urgent Care, Transport, and Physician Offices. |  |  | December 22, 2015 | 2015 | Hospira Inc. | Potential for the Symbiq pole clamp assemblies to be missing therubber stop pad. | Nonconforming Material/Component | On December 22, 2015 Stericycle, Inc. sent an URGENT MEDICAL DEVICE CORRECTION notification to direct consignees of record via traceable mailing. The notification included a description of the recall, risk to health information and necessary actions to be performed by the direct account. These include: an inspection of the affected devices for missing rubber stop pad on the Symbiq pole clamp assemblies at the direct accounts.If missing rubber stop pads were discovered they were advised to contact Hospira's Advanced Knowledge Center at 1-800-241-4002, option 4, to receive a replacement pole clamp assembly or be directed to a field Service Engineer who will replace the pole clamp.Hospira ALso advised informing healthcare professionals at their institution about the issue and complete the attached reply form and return it by e-mail or fax at: 1-800-441-4100/1-800-241-4002 or at ProductComplaintsPP@hospira.comFinally customers where advised to notify FDA of any adverse events through FDA's MedWatch Adverse Event Reporting program. |  | 6634 | 6634 units. | U.S. consignees nationwide.*** Foreign consignees in: Canada. | Open, Classified |  |  | 6634 | Not_Computer
Z-0858-2016 | 72720 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare IntelliVue Info Center iX, B.0, 866389 |  |  | November 25, 2015 | 2015 | Philips Electronics North America Corporation | Reconstructed ECG leads viewed or printed at the Information Center iX may misrepresent the ECG waveform in specific leads. | Software Design Change | Philips Healthcare issued a Field Safety Notice on 11/25/15 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will provide a software correction free of charge. On 1/28/16 the Philips issued an updated letter to users.Philips has initiated a correction to address this issue. A software correction will be provided to customers with impacted devices at no charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice: Arrhythmia alarms are not affected. ST alarms, ST values and measurements are also not affected and could be used to assist with the decision to obtain a diagnostic ECG.The ECG waves at the bedside are not affected. Printing ECG waves from the bedside is an alternative to printing from the PIIC iX.If you need any further information or support concerning this issue, contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. |  | 1880 | 1880 units | Worldwide Distribution-US (nationwide) and the countries of Algeria Australia Austria Belgium Bulgaria Canada Denmark Estonia Finland France Germany Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Kenya Korea Kuwait Luxembourg Malaysia Mongolia Morocco Netherlands New Zealand Nigeria Norway Poland Qatar Romania Russian Federation Saudi Arabia Singapore Slovenia South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates and United Kingdom. | Open, Classified |  |  | 13120 | Computer
Z-0852-2016 | 73179 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | AccuLIF PL Tubing Set and AccuLIF TL Tubing Set |  |  | December 16, 2015 | 2015 | Stryker Spine | There were two potential interference conditions identified with the way the tubing set attaches to the inserter. | Device Design | Stryker Spine sent an Urgent Product Removal letter dated December 17, 2015 to affected customers via Fed Ex priority overnight. The letter identified the affected product, problem and actions to be taken. Customers are asked to follow instructions included in the letter. The customer response form should be completed and sent back to email via Spine-RegulatoryActions@Stryker.com or Fax 855-632-9049. For questions call (201)749-8389. |  | 413 | 413 units | US Nationwide Distribution | Open, Classified |  |  | 413 | Not_Computer
Z-0853-2016 | 73209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Measurement Module X1 |  |  | January 04, 2016 | 2016 | Philips Electronics North America Corporation | The ST elevation alarm on the Patient Monitor or standalone X2 Measurement Module will not sound when indicated for all chest leads derived using Hexad 12-Lead ECG Monitoring in the Host Monitor | Software Design Change | Philips Healthcare issued the Field Safety Notice on January 4, 2016, and informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice: Until your software is upgraded, please make sure that the ST Analysis is switched ON when using STE measurement in the Hexad 12-lead ECG Monitoring. This can be done by entering the ST Analysis Menu and select ST Analysis to On . For more detailed information, please refer to the Instructions for Use (IFU) of your host monitor or to the X2 Measurement Module IFU.Philips will provide a software correction free of charge. For further questions please call (978) 659-3000. |  | 13300 | 13,300 | Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran ( Islamic Republic of ) Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Morocco, Myanmar(Burma) Netherlands Antilles, Norway, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 13300 | Computer
Z-0855-2016 | 73042 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker Vitallium Wire (6 Pack) Non Sterile. For use in bone procedures. |  |  | December 23, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | The wire packages are correctly marked with a "NON-STERILE" label, however, the enclosed Instructions For Use (IFU) states the device are sterilized via gamma irradiation and should not be resterilized. Since the IFU states the devices are sterile, no instructions for moist heat sterilization are provided. | Under Investigation by firm | Stryker Branches/Agents were notified of the action via email on December 23, 2015. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated December 23, 2015 were sent to Branches/Agents via UPS on December 28, 2015. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated December 23, 2015 were sent to Hospital Risk Managers and doctors via UPS on 12/23/2015. The notification instructed customers of the related issue with the affected product; how to identify affected product; the potential hazards associated with affected product; risk mitigations; and actions needed to be taken.Customers were asked to complete and return the attached Product Recall Acknowledgement Form within 5 days and either email (strykerOrtho7808@stericycle.com or Fax (1-866-672-0627 ) theresponse form back. Customers were instructed to return all affected products to Stryker C/O Stericycle, 2670 Executive Drive, Suite A, Indianapolis, IN 46241 (Attn RA2015-120-Event 7808). A point of contact was provided in case the customer had any questions 201.831.5272 . |  | 3741 | 3,741 units | Nationwide Distribution. | Open, Classified |  |  | 3741 | Not_Computer
Z-0862-2016 | 73041 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Plaza Picture archiving and communication system. |  |  | December 22, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Siemens is releasing an updated software version to address several software issues including RGB images will show "?" since calculation of HU is not possible; save as option enabled; changes in access for loading studies; breast region is now properly fitted to segment boundary when clicking fit breast to screen. | Software design | Siemens sent a Voluntary Correction and Removal Action letter dated December 30, 2015. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that a software update VB10A_HF01/06 is being released simulataneously to eliminate the issue. Customers with questions were instructed to call 1-888-826-9702. |  | 68 | 68 | Nationwide Distribution including ALCACTFLINKSLAMAMDMIMNMOMSNENYOHORPATXWAWI | Open, Classified |  |  | 68 | Computer
Z-0861-2016 | 73168 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Convenience Kits from Centurion that contain an In-Vision- |  |  | January 19, 2016 | 2016 | Centurion Medical Products Corporation | Convenience Kits contain an In-Vision-Plus Needleless IV Connector, code RYM5001 that is being recalled by the manufacturer, RyMed Technologies, LLC. A component molding abnormality may cause female luer threads to crack, leading to potential for the mating tubing set to become detached and for the administration to be interrupted. Other components within these kits are not affected. | Nonconforming Material/Component | The firm, RyMed Technologies, LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 1/4/2016 to the affected customers via certified mail on 1/19/2016 with instructions for identifying and returning the affected product. The recall letter included a description of the reason for the recall, affected product, actions to be taken, and instructions for responding to the formal recall notification. The customers were instructed to discontinue distribution and/or use and return product to: RyMed Technologies, LLC, 6000 W. William Cannon Drive, Building B, Suite 300, Austin, TX 78749; that user facilities change the IV connector within 24 hours when lipids are being administered; if you further distributed any of the lots, contact your accounts, advise them of recall and return their outstanding stock, complete and return enclosed Acknowledgement and Receipt Form ASAP and no later than 10 business days via E-mail: amccutchen@rymedtech.com, Fax: 512-301-7338 or by mail to address above, Attention Recall Return Notification. Please contact Director of Quality Assurance/Regulatory Affairs at (512) 301-7334 Ext.303 or email: ammccutchen@rymedtech.com should you have questions pertaining to recall notification. Questions pertaining to Returning Product/Replacement Product contact Logistics Specialist at (512) 301-7334 Ext.309 or email svara@rymedtech.com. |  | 1130 | 1130 | US Distribution in states of: TX, FL, VA, TN, PA, IL, GA, and CA. | Terminated | July 19, 2016 | 182 | 1130 | Not_Computer
Z-0851-2016 | 73155 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD FMC7 FITC: Model 430918. |  |  | January 25, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | Becton,Dickinson and Company (BD) has determined that the FMC7 FITC (ASR) has an incorrect expiration date listedon the vial label. | Under Investigation by firm | BD Biosciences sent an Urgent Medical Device Recall letters dated January 2016 via certified mail to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed that the remaining stock of the product can be used until 6-30-16. however, if prefer they can discard all remaining stock of the affected lot and BD will replace.The letter request customers complete the Tracking/verification form and returned to BD Biosciences, per the instructions (mail, fax, or email).For questions contact your local BD representative or BD Customer Support at 877-232-8995 (prompt 3,3). |  | 78 | 78 | Worldwide Distribution - US Nationwide in the staates of TX, NY, GA, NJ, AL, MA, IL, PA, DC, FL, CA, MD, OK. and coutry of Taiwan. | Open, Classified |  |  | 78 | Not_Computer
Z-0859-2016 | 73139 | No | LGW |  |  | Neurology | PMA | Specify 5-6-5 Lead, Specify 2x8 Lead, Model Numbers: 39565, 39286. Spinal cord stimulation leads. Part of a neurostimulation system for treating chronic, intractable pain of the trunk and/or limbs-including unilateral or bilateral pain. |  |  | January 11, 2016 | 2016 | Medtronic Neuromodulation | Some Specify 5-6-5 and 2x8 surgical lead product labeling does not adequately distinguish between permanent surgical implant and trialing/screening use. | Labeling design | Consignees were sent a Medtronic "Medical Device Correction" letter dated January 2016. The letter was addressed to Healthcare Professional. The letter described the problem and the product involved in the recall. Advised consignees to ensure their staff of the action that is being recommended. Requested consignees to complete and return the "Physician Confirmation Form". For additional information they can contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekends 7 am - 6 Pm Central time. |  | 51138 | 51,138 (44,006 US; 9,132 OUS) | Worldwide Distribution, including US Nationwide. | Open, Classified |  |  | 51138 | Not_Computer
Z-0860-2016 | 73138 | chair, geriatric | FRJ | Medical chair and table. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Bariatric Power Recliner, Product Model # P9096 |  |  | January 20, 2016 | 2016 | Hill-Rom, Inc. | The recliners were shipped without the foot rest ottoman linkage side shields, which could result in consumer injury. | Under Investigation by firm | Hill-Rom sent an Urgent Field Safety Notice by POD (Proof of Delivery) mailing on January 20, 2016, to all affected customers.. The notice identified the product, the problem, and the action to be taken by the customer. Customers willl be contacted for appointments to update the power chairs with missing linkage side shields. Hill-Rom will be installing the missing linkage side shields to the Hill-Rom P9096 Power Bariatric Recliner to address this issue. Customers were instructed to ensure patients and others do not insert an appendage into the foot rest linkage mechanism during opening or closing. Customers may contact Hill-Rom Technical Support at 800-445-3720, if they have any questions. |  | 216 | 216 | Worldwide Distribution - US including AL, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, VT, WA, WI, WV, WY, VA Medical Center and Internationally to Canada and Mexico.**Consumer Recommended Depth - Consumers/User** | Open, Classified |  |  | 216 | Not_Computer
Z-0865-2016 | 73163 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Flexor Check-Flo Introducer, Catheter introducer, |  |  | January 20, 2016 | 2016 | Cook Inc. | Incorrect packaging. Reports of packages labeled as 6.0Fr actually containing 5.0Fr devices. | Under Investigation by firm | On 1/20/2016, URGENT: MEDICAL DEVICE RECALL notifications were sent to the hospital Risk Management/Recall Administration via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 25 | 25 | Distributed in the states of CA, FL, OK, OH, MI, TX, NJ, and NY. | Open, Classified |  |  | 25 | Not_Computer
Z-0801-2016 | 73056 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Femoral Artery Pressure Monitoring Catheter Sets and Trays |  |  | January 06, 2016 | 2016 | Cook Inc. | A manufacturing process may lead to catheter tip fracture and or separation. Potential adverse events may include loss of device function, medical intervention to retrieve a separated segment, or complications resulting from a separated tip occluding blood flow to end organs. These catheters are used in venous or arterial pressure monitoring, blood sampling, and administration of drugs and fluids. | Process control | On 1/6/2016, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Management/Recall Administration via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 1515 | 1515 | Worldwide Distribution - US Nationwide including DC and PR including the countries of Australia, Germany, China, Chile, Dominican Republic, Panama, Israel, Brazil, Canada | Terminated | August 23, 2017 | 595 | 17827 | Not_Computer
Z-0866-2016 | 72936 | No | DSF | Paper chart recorder. | Cardiovascular | Cardiovascular | 510(K) Exempt | PRN 50-M+, GE part number 2062759-001, manufacturer part number 600-23310-01 |  |  | December 02, 2015 | 2015 | Gsi Group Inc | The item was recalled due to an increased user/patient risk of exposure to fire and/or electrical shock due to the failure of the PRN 50-M+ Digital Writer | Nonconforming Material/Component | On December 2, 2015 GSI Group Corp. distributed Medical Device Correction Notice letter and confirmation forms dated November 18, 2015 to their customers via FedEx. GSI Group Corp. warned customers about a potential safety issue associated with the flammability and fluid ingress rating of the outer enclosure material. Customers are advised to follow all intended usage, cleaning and maintenance instructions that are found in the PRN 50-M+ Digital Writer Technical Manual. It is recommended that customers complete and return the attached Medical Device Correction Confirmation form via email to customerservice@insidesales-solutions.com. GSI Group Corp. will provide a correction at no charge. Once the confirmation form is received, GSI Group Corp. will make arragements to send replacements. Customers with questions can call 781-266-5700. |  | 2257 | 2257 units | Worldwide Distribution - US Nationwide including Puerto Rico; Internationally: Australia, Bahamas, Canada, Germany, Finland, France, Great Britain, Hong Kong, India, Ireland, Italy, Japan, South Korea, Mauritus, Mexico, Peru, Saudi Arabia, Singapore, Taiwan, United Arab Emirates &amp; United Kingdom | Open, Classified |  |  | 2257 | Not_Computer
Z-0753-2016 | 72515 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Precision MPi is an all-digital multipurpose tilt-C x-ray system, intended for a multitude of diagnostic procedures, including radiology, fluoroscopy, interventional procedures, vascular and non-vascular procedures, and specialized applications including angiographic studies. |  |  | July 01, 2008 | 2008 | Regulatory Insight, Inc | It was discovered that the Remote Touch Panel (RTP) of the GE Precision MPi X-ray system may not always boot up as intended and needs to be updated to properly accomplish its intended purpose. | Radiation Control for Health and Safety Act | Nordisk Rontgen Teknik ( NRT ) Planned action to bring defect into compliance:1.Nordisk Rontgen Teknik A/S (NRT) will contact customers and initiate a software and hardware update to correct the defect.2.NRT engineers will install the software and perform testing to ensure the software and hardware update was effective.3.The customer notification letter which includes a statement that NRT will without charge, remedy the defect or bring the product into compliance.For further questions please call (720) 962- 5412. |  | 43 | 43 | US Distribution to the states of :TN, CA, WA, OR, PA, WI, FL, MN, VT, NY, NV, MA, CO, IN, MI, MS and OH. | Terminated | July 11, 2017 | 3297 | 43 | Computer
Z-0864-2016 | 73195 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic Endovascular System, Catalog Number 500-56112. |  |  | January 21, 2016 | 2016 | EKOS Corporation | The EkoSonic kit, including the individually labeled MicroSonic Device (MSD) and Intelligent Drug Delivery Catheter (IDDC) are labeled with a 135 cm working length, and programmed with a 135 cm family code, however the MSD actual working length is 106 cm. | Error in labeling | All affected consignees were contacted by telephone January 20- 22, 2016.The URGENT  VOLUNTARY PRODUCT RECALL letter, dated January 21, 2016, was sent to consignees. Consignees were advised that replacement devices were shipped to their facility on January 21, 2016 and should arrive there on January 22, 2016. Consignees with questions can call: Jocelyn KerstenVice President, Regulatory and Clinical AffairsEKOS Corporation, a BTG International Group Company425-415-3132Jocelyn.kersten@ekoscorp.com |  | 10 | 10 units | Distributed in the states of AZ, CA, MA, MI, OH, and VA. | Terminated | March 15, 2016 | 54 | 10 | Not_Computer
Z-0798-2016 | 72387 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Shuttle Select Slip-Cath Catheter. Angiographic catheter. |  |  | October 07, 2015 | 2015 | Cook Inc. | Reports of catheter tip splits and or separation. Loss of device function, medical intervention to retrieve separated segment, or complications resulting from separated tip occluding blood flow to end organs may result. Recall expanded 10/07/2015 to include additional catheters. | Component design/selection | On 7/2/2015 Cook Incorporated sent, URGENT: MEDICAL DEVICE RECALL notifications dated July 2, 2015, to affected customers with instructions for returning the affected product. All customers were notified viacourier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Should you have anyquestions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or (812) 339-2235. Firm issued press release 8/3/2015.On 10/09/2015 the firm expanded the recall and reissued press release via www.businesswire.com to include additional catheter products and lots as specified in the release. |  | 9519 | 9519 | Worldwide Distribution - US (nationwide including Puerto Rico and US Virgin Islands), Algeria, Antigua, Armenia, Austria, Australia, Bahamas, Belgium, Bangladesh, Bulgaria, Bahrain, Brazil, Brunei Darussalam, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Ghana, Great Britain, Greece Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kuwait, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxemburg, Latvia, Libya, Macau, Malaysia, Malta, Martinique, Mexico, Mongolia, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Switzerland, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 9519 | Not_Computer
Z-0891-2016 | 73088 | No | GCQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus, Model No. CF-H180AL, flexible scope, large diameter, video colonoscope |  |  | January 20, 2016 | 2016 | Integrated Medical Systems Inc | If excessive pressure is applied to the distal end of the endoscope, the internal coil pipe assembly at the distal end of the endoscope can become detached and angulation may not operate as intended. | Nonconforming Material/Component | The firm notified their consignees of the issue by Fed Ex on 01/20/2016. The firm plans to coordinate with the medical facilities for inspection of the endoscopes. |  | 232 | 232 instruments | Distributed US (nationwide) and in Australia. | Open, Classified |  |  | 1583 | Not_Computer
Z-0698-2016 | 72837 | No | KFZ | Therapeutic humidifier for home use. | Anesthesiology | Anesthesiology | 510(K) Exempt | WARM STEAM Personal Steam Inhaler-Adjustable control to regulate flow of steam Soft touch facial &amp; nasal masks" |  |  | November 20, 2015 | 2015 | Bigwall Enterprises | "The public reason for the recall is that the product is being recalled out of an abundance of caution due to one or more consumer complaints, which have been independently investigated and remain unconfirmed. No causal connection between the complaints and the productaÂż"s proper use as instructed have been determined." | Unknown/Undetermined by firm | The firm has only sold the product to Walgreens. Walgreens has notified all of its distribution centers and stores. According to the firm: "Walgreens has issued a stop sale at the store level. Units removed from store shelves and in stock will be sent back to the Walgreens distribution center Valparaiso, Indiana. As units are received from all store locations they will be sent to Bigwall Enterprises." |  | 87503 | 87503 | Nationwide Distribution including Washington, DC, U.S. Virgin Islands, and Puerto Rico. | Open, Classified |  |  | 87503 | Not_Computer
Z-0886-2016 | 73181 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Auto Endo5 Hem-o-lok Automatic Clip Applier, Medium Large Polymer Clips, Rx Only, Sterile, REF 543965 |  |  | February 02, 2016 | 2016 | Teleflex Medical | The clip in the applier may be missing a boss. A boss is the part of the clip that is used to hold the clip within the applier jaws prior to application. | Nonconforming Material/Component | Teleflex sent an Urgent Medical Device Recall Notification letter dated February 2, 2016 to customers. The letter identified the affected product, problem and actions to be taken. A separate letter was sent to distributors advising them to conduct a sub-recall. Customers were instructed to: 1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed in the letter.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1- 855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.This will document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. This will document receipt of the letter.For questions contact your local sales representative or Customer Service at 1-866-246-6990 |  | 543 | 543 ea. | Worldwide Distribution in the state of CA, MI, NC, PA, TX, VA, and the countries of Canada, Korea &amp; Belgium | Terminated | August 11, 2017 | 556 | 543 | Not_Computer
Z-0887-2016 | 73077 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Naturalyte Liquid Acid Concentrate |  |  | January 18, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | On the label calcium (Ca) listed in the right upper corner of the label (red background) incorrectly indicates the product contains 2.5 mEq/L Ca. The concentration indicated in the list of constituents represents the actual calcium concentration of the final dialysate, 3.0 mEq/L. | Error in labeling | The firm, Fresenius Medical Care, issued an "URGENT-MEDICAL DEVICE RECALL" letter dated 1/14/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your stock for the product; if found, discontinue use immediately; place all units in a secure, segregated area; contact your FMCNA Customer Service Team at 1-800-323-5188 for instructions how to return the affected lot (lot# 15STAC072), and promptly complete and return the attached reply form by Fax to: 781-699-9796 Attn: Quality Department or scan and email to: NOTIFYRA@fmc-na.com. Fresenius will pick up and replace your product.If you have any additional medical concerns or questions, please contact Medical Information and Communication at 855-816-2309 or Website: www.fresenius-medinfo.com.Further information or support concerning this issue, contact the Customer Service Care Team at 800-323-5188 and reference the and Support Field Action Number. |  | 848 | 848 (212 cases) | US Distribution to states of: KS, LA, MS, OK, and TX. | Terminated | August 14, 2017 | 574 | 848 | Not_Computer
Z-1018-2016 | 73126 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | CDI H/S Cuvette, non-sterile, Manufactured by Terumo Cardiovascular Group, Perfusion Pack or Smart Perfusion Pack. |  |  | January 19, 2016 | 2016 | Sorin Group USA, Inc. | The CDI System 500 monitor displays the "H/S DISCONNECT AT CUVETTE" error message when the CDI HIS Cuvette does not make a proper connection with the CDI H/S Probe. When this occurs, blood parameter values for HCT, Hgb and S02 do not display. | Component design/selection | A Safety Alert letter dated January 13, 2015 via registered mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to check the communication of the Cuvette prior to initiating Cardiopulmonary Bypass (CPS).The Terumo Safety Alert did not request removal or return of the product. The instructions that Sarin has communicated to the customers is identical to the information from Terumo. Acknowledgement of the letter and understanding ofthe contents is required for all customer responses. |  | 6859 | 6859 | US Nationwide Distribution and Canada. | Open, Classified |  |  | 6859 | Not_Computer
Z-0888-2016 | 73307 | No | OTE |  |  | Radiology | PMA | Mammomat Inspiration with Tomosynthesis functionality |  |  | February 12, 2016 | 2016 | Siemens Medical Solutions USA, Inc | A potential system fatal error may occur during tomosynthesis reconstruction with large breasts that cover nearly the entire detector surface and with a breast thickness larger than 90 mm. The data to be processed for the tomosynthesis application can exceed the capacity of the graphical processing unit (GPU). If the fatal error message occurs, reconstruction of the data will be aborted. The syste | Software design | Siemens sent a Field Safety Notice dated February 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.How will the issue finally be resolved?Siemens is releasing a field modification to resolve this issue as soon as possible. The upgrade software (VB30P) will be available to the affected customers free of charge. A Siemens service engineer will contact you to schedule implementation of the modification for your system. According to regulations set forth in 21 CFR 7.49(d), we advise that you follow the instructions in this notice and where necessary request that you promptly notify and instruct all the staff at your organization who have to be aware of this problem.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed.We thank you for your cooperation. For further questions, please call (610) 448-6478. |  | 76 | 76 units | Nationwide Distribution including Puerto Rico and to the states of : AR, CA, CO, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, NC, NE, NH, NJ, NY, OH, PA, SC, TN, TX, VA and WI. | Open, Classified |  |  | 76 | Computer
Z-0885-2016 | 73258 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | HipLOC CHS Plate/Lag Screw Introducer |  |  | February 16, 2016 | 2016 | Biomet, Inc. | A complaint was received where the Hiploc Plate/Lag Screw Introducer was undersized and would not accept the introducor coupling screw. It appears now that not all parts manufactured to Rev A drawing were reworked. | Process change control | Biomet sent an " URGENT MEDICAL DEVICE RECALL NOTICE" dated February 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The following actions are REQUIRED:Immediately locate and remove the identified device(s) listed below from circulation.Remove the affected product from circulation. Review this notification and ensure that all affected personnel are aware of its contents. Carefully follow the instructions on the enclosed Response Form .Email a copy of the response form to CPWARFieldAction@zimmerbiomet.com prior to return of product.You do not need to complete the Biomet UK response form. Use priority carrier for your shipment.Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA:"Online: www.fda.gov/Safety/MedWatch/HowToReport/default.htm (form available to fax or mail), or"Call (800)FDA-1088Thank you in advance for your assistance and prompt attention. On behalf of Zimmer Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 23 | 23 units | US Distribution to the states of : MO, UT, AZ, IL, PA, KY, MI, NC, NJ, MN and GA., and Internationally to The Netherlands. | Open, Classified |  |  | 23 | Not_Computer
Z-0827-2016 | 73241 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Evita Infinity V500 Ventilator; Babylog Infinity VN500 Ventilator. |  |  | February 04, 2016 | 2016 | Draeger Medical, Inc. | The battery capacity of the optional PS500 power supply unit for the Infinity (ACS) Workstation Critical and Neonatal Care was rapidly and unexpectedly reduced, despite using the most recent power supply firmware version 1.50 released in November 2015 (recall Z-0436-2016). The batteries were prematurely discharged, even though an adequate battery charge status was displayed on the device. | Software design | Draeger Medical, Inc. will send an Urgent Medical Device Recall Letter to consignees in a method that confirms delivery. For consignees that have received new devices with firmware version 1.50 (127 devices) or through device updates (644 devices; reference Recall Z-0436-2016), those devices will be downgraded to firmware version 1.49 and the batteries replaced. Batteries will continue to be replaced for all consignees every six months until a final solution is available. |  | 2553 | 2553 | US Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 2553 | Computer
Z-0884-2016 | 73156 | No | MAY | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker, T2 K-Wire, 3.2x400mm, CoCr, Recon |  |  | January 21, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | During a review of the packaging, it was determined that the seal integrity of the pouch may be decreased. | Under Investigation by firm | Stryker Orthopaedics sent an Urgent Medical Device Removal letter and Customer Response forms dated January 21, 2016 through UPS. The letter identified the affected product, problem, actions to be taken including forwarding notice if applicable, returning all affected products and completing and returning the Customer Response form to Stericycle.by the instructions provided. |  | 147838 | 147,838 units in total | US Nationwide Distribution | Terminated | June 08, 2017 | 504 | 147838 | Not_Computer
Z-0966-2016 | 73024 | No | PIF | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoVive 3s Low Profile Balloon Kits |  |  | December 23, 2015 | 2015 | Xeridiem Mediem Medical Devices Inc | A feeding (reflux) valve leakage issue. The feeding valve appears to close inconsistently, or close in a delayed manner, when an extension set is disconnected from the Low Profile Balloon device, thus allowing feeding solution or gastric contents to leak out of the device in some instances. | Device Design | Xeridiem sent a customer notification letter dated December 23, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customer was provided the following instructions:1. Segregate all devices from the above table that in Boston Scientific Corporation's inventory and restrict them from being distributed to end users.2. Determine the list of end users who have received devices from the above list, notify them of the quality issue, and have these devices returned to Boston Scientific Corporation. Provide Xeridiem Medical Devices with this list of end users.For any questions the customers is instructed to contact the firm.For questions regarding this recall call 520-882-7794. |  | 38 | 38 | Nationwide distribution to MA only. | Open, Classified |  |  | 7171 | Not_Computer
Z-1024-2016 | 73333 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SENSE Body Coil 1.5T with identification; used in the ACS NT, Intera 1.5T and Achieva 1.5T MR system |  |  | July 06, 2015 | 2015 | Philips Electronics North America Corporation | Wrong positioning of the coil cables can create unintended resonances via the patient or elevated coil cable temperature due to coupling to the QBC. The combined use of the Sense coil, high SAR level scanning and placement of the cables contrary to instructions for use can lead to patient burns. | Under Investigation by firm | On 6/18/2015, the Field Safety Notice / Medical Device Correction Notice was sent to customers via certified mailing. The letter explains the issue and provides user instructions. A sleeve is being developed that will be applied to the cable on the Sense Body Coil 1.5T and provide consistent spacing between the cable and the patient, thereby reducing the chance of RF interaction and heating. This sleeve will be distributed as part of a Field Change Order with reference FCO781 00437 for Sense Body Coils in the installed base and will be installed by the Philips Field Service Engineers. |  | 5672 | 5672 | US: NationwideOUS: Canada, AfghanistanAlbaniaAlgeriaAndorraArgentinaAustraliaAustriaAzerbaijanBahamasBangladeshBelarusBelgiumBoliviaBosnia and HerzegovinaBotswanaBrazilBurkina FasoChileChinaColombiaColumbiaCongo, DemocraticCosta RicaCroatiaCyprusCzech RepublicDenmarkDominican RepublicEcuadorEgyptEstoniaEthiopiaFinlandFranceGeorgiaGermanyGhanaGreeceGuadeloupeHong KongHungaryIcelandIndiaIndonesiaIranIraqIrelandIsraelItalyJapanJerseyJordanKazakhstanKenyaKorea, Republic ofKuwaitKyrgyzstanLatviaLebanonLibyaLithuaniaLuxembourgMalaysiaMaltaMartiniqueMauritiusMexicoMonacoMongoliaMoroccoMozambiqueNepalNetherlandsNew CaledoniaNew ZealandNorwayOmanPakistanPalestinian territory, occupiedPanamaPapua New GuineaPeruPhilippinesPolandPortugalPuerto RicoRomaniaRussian FederationRwandaSaudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSri LankaSudanSwedenSwitzerlandSyrian Arab RepublicTaiwanTanzaniaThailandTunisiaTurkeyUgandaUkraineUnited Arab EmiratesUnited KingdomVenezuelaViet NamYemenZimbabwe | Open, Classified |  |  | 5672 | Not_Computer
Z-1022-2016 | 73010 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT &amp; BrightView XCT Upgrade, gamma cameras; For Single Proton Emission Computed Tomography. |  |  | December 17, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm discovered four issues with the BrightView Family systems.Issue 1: During a cardiac non-Auto Body Contouring (ABC) scan, using Cardiac High Resolution (CHR) collimators, the collimator may come in contact with the bottom edge of the patient pallet. If this occurs, collision sensors would pause the scan.Issue 2:During a scan, the gantry motion may halt prematurely. This cause | Software design | On 12/22/2015, the firm sent Customer Information Letters to their consignees.The letter identified the product, problem and actions to be taken by the customers. |  | 427 | 427 units | Worldwide Distribution-US (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, VT, WA, WI, WV &amp; WY, and in DC and PR and in the countries of Algeria, Argentina, Armenia, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Cayman Islands, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestinian, Panama, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom &amp; Viet Nam. | Open, Classified |  |  | 1221 | Not_Computer
Z-1019-2016 | 73265 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Virage OCT Spinal Fixation System. Intended to promote fusion of the Occipital-Cervical-Thoracic spine (Occiput-T3). |  |  | February 03, 2016 | 2016 | Zimmer Spine, Inc. | Zimmer Biomet Spine is initiating a Medical Device Correction regarding the Virage OCT Spinal Fixation System by providing additional guidance in the Virage OCT Surgical Technique Manual. | Under Investigation by firm | Distributors, Hospital Risk Managers and Surgeons were notified via an Urgent Medical Device Correction letter beginning 2/3/2016. The letter identified affected product, described the issue and listed their responsibilities. A Certificate of Acknowledgement form was asked to be returned to Zimmer Biomet via email to corporatequality.postmarket@zimmerbiomet.com. Questions or concerns can be directed to Customer Service at 1-866-774-6368 |  | 15753 | 15,753 | Distributed US (nationwide) and in the countries of Australia and Germany. | Open, Classified |  |  | 15753 | Not_Computer
Z-1023-2016 | 73261 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Custom Perfusion Tubing Pack, Rx, Sterile EO. |  |  | February 04, 2016 | 2016 | Medtronic Perfusion Systems | Medtronic is notifying customers who have or may receive specific lots of Medtronic Perfusion Tubing Packs that include the potentially affected cuvettes manufactured by Terumo Cardiovascular Systems (Terumo CVS). Terumo Cardiovascular Systems (Terumo CVS) has distributed an Urgent Safety Alert regarding the CD IÂż H/S Cuvette. The cuvette is packaged into Medtronic Perfusion Tubing Packs and | Component design/selection | Medtronic sent an Urgent Medical Device Safety Alert dated February 2016 to all consignees on February 4, 2016. The letter described the problem and the product involved in the recall. Advised consignees to review the Terumo CVS Urgent Safety Alert and follow the instructions in the Recommended Action Section of the Urgent Safety Alert letter. Requested consignees to completed the Customer Confirmation Certificate and Email completed form(s) to RS.CFQFCA@Medtronic.com. If email is not available, fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. For questions contact your Medtronic representative or Lifeline Technical Support at 877-526-7890.For questions regarding this recall call 763-391-9000. |  | 9890 | 9890 (7773 US, 2117 OUS) | Worldwide Distribution - US including US: AK, AZ, CA, DE, FL, HI, IL, LA, MD. MA. MN, MS, MO, NE, NJ,. NHY, NC, ND, OH, OR, PAM TN, TX, VA, WA, WV, and WI.Internationally to Australia, Belgium, Canada, Czech Republic, Denmark, Dominican Republic, France, Germany, Israel, Japan, Netherlands, Serbia, South Africa, and Switzerland. | Open, Classified |  |  | 9890 | Not_Computer
Z-1025-2016 | 73159 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Karma Flexx Wheelchair |  |  | January 27, 2016 | 2016 | Karman Healthcare Inc | Karman Healthcare, Inc. is recalling Karma Flexx Wheelchair because incorrect size screws and corresponding hardware were installed on the front caster fork brackets. | Employee error | On 01/27/16 the firm sent out customer notification letters.In the letter customers were informed that the firm was conducting a product update for original purchasers of KM8522. Customers were provided the following procedures for the product update:1. Completed and signed notification must be returned to Karman Healthcare via fax at (626)-581-2335, or e-mail to edward@karmanhealthcare.com.2. Following the received frame, customers may install their current wheels, and footrest onto the newly updated frame. Kindly package the used frame into the carton in the same manner the new carton was shipped.3. A prepaid return shipping label is included with the letter, which customers may use to return the carton; or contact the firm's customer service line at (800)-805-2762 to have a trained representative schedule a pickup at no charge. |  | 35 | 35 | U.S. distribution to the following states -- WA, CA, NY, and TX. | Open, Classified |  |  | 35 | Not_Computer
Z-1027-2016 | 72976 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | NeoCoil 3.0T GEM Flex Coil. |  |  | December 18, 2015 | 2015 | NeoCoil, LLC | Potential for higher than specified surface temperatures. Use of 3.0T GEM Flex system may result in patient warming or burns. | Component change control | Distributor was sent a NeoColi "Urgent Medical Device Recall" letter dated 18 December 2015. The letter described the problem and the product involved in the recall. Consignee was advised to quarantine the product and to identify their customers and provide a list to NeoCoil. NeoCoil will contact each customer for correction and removal of the devices. Requested consignee to complete and return the response form. For questions call Michael Leigh at 261-522-6127. |  | 2121 | 2121 | US Nationwide Distribution | Open, Classified |  |  | 2121 | Not_Computer
Z-1026-2016 | 73308 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | Helioseal Clear Refill 1 x 1.25 g, REF/Product Code/Art. No. 558521AN and Helioseal Clear Assortment 5 x 1.25 g, REF/Product Code/Art. No. 558520AN ---- COMMON/USUAL NAME: Pit &amp; Fissure Sealant |  |  | February 12, 2016 | 2016 | Ivoclar Vivadent, Inc. | Complaints were received claiming the material failed to polymerize under the curing light. After investigation, it was determined that a wrong photo-initiator was used in the production batch U30755 of Helioseal Clear. This failure can lead to incomplete curing of the sealant, if mono-wave LED curing lights are used. | Mixed-up of materials/components | Ivoclar Vivadent sent an Urgent Medical Device Recall letter dated February 11, 2016, to all consignees via Certified Mail, Return Receipt Requested. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to complete, sign and return the attached acknowledgement form and return it to Ivoclar Vivadent via fax (716) 691-2294 or e-mail at recall.us@ivoclarvivadent.com. Contact Ivoclar Vivadent Customer Service at 1-800-533-6825 as soon as possible between the hours of 8:00 am and 6:00 pm EST Monday through Friday and a representative will arrange to pick up your affected materials for return and credit. Please notify your customers of this issue and return the product to us. For questions, please contact Ivoclar Vivadent Customer Service at 800-533-6825. |  | 406 | Domestic: 406 units | Nationwide distribution including NY, PA, FL, IN, IA, TX, and CA. | Terminated | May 08, 2017 | 451 | 406 | Not_Computer
Z-1028-2016 | 73083 | No | MIJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Kopans Breast Lesion Localization Needle |  |  | January 08, 2016 | 2016 | Cook Inc. | Product potentially exposed to body fluids during processing. Potential adverse events that may occur as a result of the foreign matter include blood borne pathogen transmission. However, the risk of transmission of blood borne pathogens is thought to be minimal to non-existent since these devices were sterilized after the potential exposure to body fluids. | Process control | Cook Medical sent an URGENT: MEDICAL DEVICE RECALL dated January 8, 2016, to all affected customers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to do the following:Please review the attached list of affected products and lot numbers shipped to your account, and quarantine any affectedproduct that remains unused.2. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit.3. Please complete the attached Recall Response Form and return to Cook Medical either with the product or separately.4. Please report any Adverse Event to Cook Medical Customer Relations 1-800- 457-4500 or 1-812-339-2235. Mondaythrough Friday between 7:30 a.m. and 5:00 p.m. Eastern Daylight Time or email at FieldActionsNA@Cookmedical.comShould you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 134 | 134 | Nationwide distribution including AZ, CA, GA, IL, MO, MS, MT, NC, NH, NY, OH, OK, and PA.. | Terminated | March 23, 2017 | 440 | 134 | Not_Computer
Z-1030-2016 | 73080 | No | JZG | Beta-2-microglobulin immunological test system. | Immunology | Immunology | 510(K) Exempt | Dimension Vista System |  |  | November 03, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed an increase in the rate of Abnormal Assay errors and calibration failures with the Dimension Vista B2MIC Flex reagent cartridge: Dimension Vista System Flex reagent cartridge and urine stabilizer B2MIC. In the unlikely event that a patient result is generated, a suppressed result would be in conflict with the patients clinical status and other markers of renal injury. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall dated November 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use and discard your remaining inventory of the affected product. Indicate your replacement product needs on the attached Field Correction Effectiveness Check Form and return. For questions call Siemens Customer Care Center- Technical Solutions at 800-441-9250 or your local Siemens technical support representative. |  | 1637 | 1637 units. | Worldwide Distribution and US in the state of CA. | Terminated | January 13, 2017 | 437 | 1637 | Not_Computer
Z-1031-2016 | 73201 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoris Multicompartmental Knee (MCK) System |  |  | January 15, 2016 | 2016 | Mako Surgical Corporation | Reports where the Mylar/Tyvek pouch for RESTORIS MCK Baseplate in the clamshell packaging configuration exhibited damage, compromising the sterile barrier. | Unknown/Undetermined by firm | The firm, Stryker Orthopaedics, sent an "Urgent Product Recall" letter dated 1/15/2016 to all their consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to inspect the device packaging as the damage to the sterile barrier may be notice prior to the procedure; do not use the device if the seal or package is breached and complete and return the attached Product Recall Acknowledgment Form within 5 days by e-mail StrykerOrtho5317@Stericycle.com or fax at: 888-480-2853. Return all affected products available at your location to: Stryker c/o Stericycle, 2670 Executive Dr., Suite A, Indianapolis, IN 46241, Attn. RA2015-047-Event 5317.If you have any questions, feel free to contact Sr. Director QA at 954-628-0502. |  | 146 | 146 | Worldwide Distribution: US (nationwide) including states of: AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, LA, MI, MS, MO, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI, and WY ; and countries of: Italy, Scotland, Turkey, Greece, Germany, and Singapore. | Open, Classified |  |  | 146 | Not_Computer
Z-1029-2016 | 73376 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | ClosureFast Micro Introducer Sheath Set 7Fr/Ch(2.3mm) x 7 cm REF MIS-F707 and ClosureFast Introducer Sheath Set 7Fr/Ch(2.3mm) x 11 cm REF IS-7F11. |  |  | February 17, 2016 | 2016 | Argon Medical Devices, Inc | The incorrect sheath was assembled to the device so that the size 7 products received 11 cm sheaths and the size 11 product received 7 cm sheaths. | Process control | On February 17, a letter was sent to Medtronic notifying them of a voluntary recall. |  | 2000 | 2000 units | Distributed in IL. | Open, Classified |  |  | 2000 | Not_Computer
Z-0829-2016 | 72824 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 9131-001 Defibrillation Electrodes, Lot No. 141125-02 |  |  | January 14, 2016 | 2016 | Cardiac Science Corporation | Defibrillation electrodes may increase electrical impedance over time. If impedance becomes too high, the AED will fail the electrode self-test and it will not be Rescue Ready, the AED STATUS INDICATOR will be RED and the AED will beep. This is normal behavior when the impedance limit is exceeded. The high impedance electrodes require replacement. | Nonconforming Material/Component | Cardiac Science sent an Urgent: Voluntary Medical Device Recall letter dated January 2016 starting on January 14, 2014. The letters were either addressed to Distributor or Customer. The letter described the problem and the product involved in the recall. Requested consignees to replace any electrodes from the lot 141125-02 and return them to Cardiac Science. Furthermore, it advised consignees to contact their customers or members of their facilities who need to be made aware by providing a copy of the notification letter and replacement electrodes if they were a distributor. For questions contact Cardiac Science Technical Support at 1-800-426-0337, option 1 in USA or +1-262-953-3500 OUS or via e-mail a9131cardiacscience.com. For questions regarding this recall call 800-426-0337. |  | 10943 | 10,943 | Worldwide Distribution - US all states including PR except ND &amp; RI.International Distribution to Argentina, Australia, Belgium, Bermuda, Canada, Chile, Colombia, Ecuador, France, Germany, Hong Kong, Iceland, Ireland , Israel, Lebanon, Malaysia, Malta, Norway, Oman, Philippines, Poland, Qatar, Romania, Russian Federation, Singapore, Slovenia, Sweden, Turkey, United Kingdom. | Terminated | December 21, 2016 | 342 | 10943 | Not_Computer
Z-1033-2016 | 73115 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Covidien Kangaroo Polyurethane Feeding Tube Radiopaque Line, Safe Enteral Connections 6.5 Fr/Ch (2.2 mm) x 36" (91 cm), Sterile; Catalog Number: 461438. |  |  | January 21, 2016 | 2016 | Covidien LLC | 8FR 16 PVC tubes were packaged in a 6.5 FR 36 Polyurethane packs and 8FR 20 PVC tubes were packaged in an 8FR 20 Polyurethane packs | Labeling Change Control | Medtronic/Covidien Customers notified via an Urgent Medical Device Recall letter, dated January 22, 2016. The letter identifies the affected product, instructs customers to discontinue use and quarantine any of the product, and return all affected inventory. Customers are also to complete the "Recalled Product Return Form." Questions or concerns can be directed to a Medtronic representative or Customer Service at 800-882-5878. |  | 1010 | 1,010 units | Worldwide Distribution -- US, to the state of CA; and, to Australia and Europe. | Terminated | June 01, 2017 | 497 | 2340 | Not_Computer
Z-1036-2016 | 73237 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | PICC/CVC Secure Dressing Change kit, Ref: 73775, Lot 213403, Sterile Single Use, Medical Action Industries, Inc., Arden, NC. |  |  | February 03, 2016 | 2016 | Medical Action Industries Inc | Misbranding; Due to the presence of CHG in the Tegaderm PICC/CVC dressing change convenience kit, which is not identified on the exterior labeling | Under Investigation by firm | Medical Action Industries, Inc. sent an Urgent Medical Device Recall letter dated February 3, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately examine their inventory and quarantine all product subject to this recall. Consignees were also instructed if product was further distributed to identify their consignees and notify them at once of this product recall. Recalled product should be returned to Medical Action Industries for credit or replacement by contacting quality@owens-minor.com . Consignees were asked to complete and return the enclosed response form. Consignees with questions were instructed to call 828-338-7630. |  | 9 | 9 cases (180 kits) | Nationwide distribution to MO, NE and TX . | Open, Classified |  |  | 9 | Not_Computer
Z-1032-2016 | 73225 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | PINPOINT Endoscopic Fluorescence Imaging System |  |  | January 29, 2016 | 2016 | Novadaq Technologies Inc. | It has been determined that the PINPOINT Operator's Manual does not contain validated parameters for gravity-cycle autoclave sterilization of the PINPOINT laparoscopes and light guide cable in the reprocessing tray. | Process design | Novadaq sent an Urgent Device Correction letter dated February 2, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Recipients of this notice have received shipments that include the inaccurate Operator's Manual. Please follow these instructions:1. Discard any old PINPOINT Operator's Manuals Rev J through M.2. Accept and review the new PINPOINT Operator's Manual (Rev N).3. Complete and sign the enclosed "Acknowledgement of Receipt" form, scan it and email to PI NPOINT@novadaq.com. Any questions or concerns may also be sent to this address. For further questions, please call 1-(844) 668-2327, ( Select Option # 1) |  | 110 | 110 units | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Bangladesh, Canada, Hong Kong, India, and Israel. | Open, Classified |  |  | 110 | Not_Computer
Z-1041-2016 | 73254 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Volcano R100 Sterile Equipment Cover; Sterile; |  |  | February 05, 2016 | 2016 | Volcano Corporation | Volcano Corporation has identified that certain lot numbers of Volcano R-100 Sterile Equipment Covers may have the incorrect expiry date on the outer shipping box; individually wrapped inner product bears the correct expiry date. | Labeling False and Misleading | Urgent Medical Device letters were sent to affected consignees the week of February 8, 2016 by Fed Ex or UPS so delivery can be tracked. For OUS consignees, the same letter will be sent and they will execute delivery to customers and track responses. |  | 1355 | 1355 total | Distributed US (nationwide) and the countries of Australia, South Africa, Canada, New Zealand, Belgium, Japan, and Paraguay. | Terminated | March 21, 2017 | 410 | 1355 | Not_Computer
Z-1045-2016 | 73212 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Getinge Group 1180.19XO Extension Device. Manufacturer MAQUET GmbH, |  |  | January 07, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | MAQUET Medical Systems USA is initiating a field action for correction of the MAQUET Extension Device (Part No 1180.19F0). The Extension Device is manufactured by Maquet GmbH, in Germany, and is a table top module for use with the Maquet MAGNUS operating table system. Maquet GmbH was made aware of a potential issue that can result in the Extension Device becoming unlocked from the table top, causing it to fall down. | Under Investigation by firm | Maquet sent "URGENT PRODUCT RECALL MEDICAL DEVICE CORRECTION" letters/MEDICAL DEVICE CORRECTION RESPONSE FORM dated January 7, 2016, via UPS to the affected customers. The notification informed customers of the issue with the affected product; how to identify affected product; measures to be taken by the users; and that a MAQUET service representative will be contacting them to schedule servicing of the device to install a Support Frame Lock. Customers with questions were asked to contact their local MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383 (Option 2, Option 3) Monday through Friday between 8:00 am and 5:00 pm EDT.Customers were also requested to fill out and return the attached response forms. |  | 249 | 2 within US; 247 OUS | Worldwide Distribution - US including AZ and MO.Foreign distribution to Australia, Austria, Bahrain, Belgium, Brazil, Canada, Denmark, Ecuador, Estonia, Finland, France, Germany, Hong Kong, India, Italy, Japan, Latvia, Malaysia, Netherlands, Norway, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Korea, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | May 25, 2017 | 504 | 249 | Not_Computer
Z-1043-2016 | 72853 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Revision Knee System, EnduRo Knee System |  |  | November 30, 2015 | 2015 | Aesculap, Inc. | Aesculap Inc. US has initiated a recall on Tibial and Femur extension sterile pressfit stem implant packaging that are used for Columbus Revision/Enduro Knee Implants because it is labeled as Cementless. US product is indicated in the US for use with Bone Cement only. | Packaging change control | Aesculap sent an Urgent Medical Device Correction letter to Distributors and Sales Rep. The letter identified the affected product, problem and actions to be taken. Customers were notified to add the over label to product packaging. Telephone conferences with Distributors and Sales Rep reiterated need for representative to inform Surgical Handler of the label when providing product for review. Customers were instructed to complete the attached Product Correction Acknowledgement Form and return to Aesculap Quality Assurance department by faxing the form to 610-791-6882 or e-mail to val.strawn@aesculap.com, two (2) weeks of receipt, even if the total inventory in your possession is zero (0). |  | 1232 | 1232 | US Nationwide Distribution | Open, Classified |  |  | 1232 | Not_Computer
Z-1044-2016 | 72823 | No | LCQ | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Nebulizer Masks Opti-Mist and Opti-Mist Clear |  |  | November 11, 2015 | 2015 | ConvaTec, Inc | The Nebulizer mask malfunctions in that there is reduced or no aerosol output. The malfunction occurs from the reduction in the diameter of a component part inside the nebulizer. | Nonconforming Material/Component | The firm, ConvaTec, sent each consignee affected by the recall an overnight express mail "Urgent Voluntary Recall Notice" and Response Form dated 12/7/2015. The notice described the product, problem and actions to be taken. The recall notice instructed the consignee to immediately stop distributing and quarantine all affected lots; perform a count of all recalled product currently in inventory; complete and return the enclosed Recall Response Form by email at convatec3657@stericyle.com or by fax at 1-888-603-8986, and return all affected product, (ONLY return the lots affected with this recall). If you want replacement product contact customer service by phone at 1-800-634-6003, extension 1103 or by email at respiratoryorders@convatec.com.If you have any questions call 1-877-650-3492. |  | 3605700 | 3,605,700 pieces | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, ME, MA, MN, MS, MO, NV, NH, NJ, NY, NC, OH, OR, PA, SC, TX VA, WA, WI, and countries of: Australia, Belgium, Burma, Canada, Denmark, Egypt, Estonia, Faroe Islands, Finland, France, Germany, Israel, Italy, Jordan, Kenya, Kuwait, Netherlands, Norway, Russia, Saudi Arabia, Sweden, Switzerland, United Kingdom &amp; Vietnam. | Terminated | July 18, 2016 | 250 | 3605700 | Not_Computer
Z-1040-2016 | 73172 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 50001 (Inrad) AccuPlace Drace Stereotaxic Needle Guide; Assists the user with inserting a needle at a predetermined angle for computer tomography. |  |  | February 04, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | During an evaluation of the product packaging, it was determined that there is a potential for the product to cause damage to the packaging, creating a possible breach in sterility. | Package design/selection | Stryker sent URGENT MEDICAL DEVICE RECALL NOTIFICATIONS to Customers via Overnight service. Stryker Sales Representatives were notified via email.1. Immediately review this Recall Notification.2. Check all stock areas and/or operating room storage to determine how manyAccuPlace" Needle Guides are at your facility. Discontinue use of theAccuPlace" Needle Guides.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of thisnotification and identify how many, if any, affected items are currently in yourinventory. Please complete and return the BRF even if you dont have anyaffected product on hand.Note: Your signature on the BRF indicates that you received andunderstand this Notification and have followed the instructions in theNotification.4. If you have further distributed this product, please forward this letter and theattached Business Reply Form (BRF) to all affected locations. Pleaseindicate each location on the BRF.5. Fax the completed BRF to Stryker Instruments Regulatory Department at866-521-2762, or scan and email a copy to kara.spath@stryker.com.6. If the BRF for your facility indicates that recalled product is currently onhand, a FedEx label will be emailed to you. This shipping label should beused to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.In addition, Stryker Instruments sent recall notifications directly to facilities that received recalled product(s) originally shipped by Inrad. Facilities that received products from both Stryker and Inrad will receive one recall notification. The Business Reply Form will be pre-filled to indicate specifically what recalled products were sent to each facility, including the quantity and Product Number(s). |  | 5800 | 5,800 units total | Nationwide Distribution-including the states of AZ, CA, CO, CT, FL, GA, IA, KS, MA, MI, MN, MO, NY, OH, OK, PA, TX, VA, WA, and WI. | Terminated | May 09, 2017 | 460 | 5800 | Not_Computer
Z-1046-2016 | 73298 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | M/L Taper Hip Prosthesis (Item: 00-7711-012-20, Lot: 62937093 |  |  | February 11, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Biomet initiated a lot-specific voluntary recall on 02/11/2016, of the M/L Taper Hip Prosthesis. single manufacturing lot of Size 12.5, Extended Offset, M/L Taper hip stems was incorrectly etched. The device was etched as STD, rather than EXT. The product packaging is labeled correctly as Extended.Offset, and the device is an extended offset stem. | Under Investigation by firm | Zimmer, Inc. initiated a lot-specific voluntary recall on 02/11/2016, of the M/L Taper Hip Prosthesis. Electronic notices were e-mailed to distributors, and letters were mailed via certified mail on 02/11/2016 to distributors and hospitals. A field complaint investigation confirmed that this single manufacturing lot of Size 12.5, Extended Offset, M/L Taper hip stems was incorrectly etched. The device was etched as STD, rather than EXT. The product packaging is labeled correctly as Extended Offset, and the device is an extended offset stem.Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons will be provided with a letter identifying the issue and their responsibilities. These responsibilities include: Assisting the Zimmer Biomet sales representative with the quarantine of the product and completion and return of the Acknowledgement Form. Customers are asked to complete the Certificate of Acknowledgement Form and return to corporatequality.postmarket@zimmerbiomet.com.Customers with have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 9 | 9 units | Distributed in the states of SD, MO, NY, MI, PA, KY and TN and in the country of Argentina. | Open, Classified |  |  | 9 | Not_Computer
Z-1049-2016 | 73248 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery LS, 2286327 |  |  | February 10, 2016 | 2016 | GE Medical Systems, LLC | GE Healthcare has updated the Service Manuals for a number of CT and PET CT Scanners to update an incorrect torque value. | Labeling design | GE Healthcare sent an Urgent Medical Device Correction letter " GEHC Ref# 25467-ABT or 25467" dated February 12, 2016, to all affected consignees . The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology, &amp; Radiologist/Cardiologist. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to ensure that the organization that service their equipment uses the latest version of the Service Manual to maintain the safe and proper performance of their medical device. For questions call in the United States: 1-800-437-1171 or your local GE Healthcare Service Representative. For questions regarding this recall call 262-513-4122. |  | 817 | 817 | Worldwide Distribution - US including all states in including PR except ND, VT &amp; WY.Foreign Distribution: ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, BURKINA FASO, CAMERON, CANADA, CHILE, CHINA, COLOMBIA, DOMINICAN REPUBLIC, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE,HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRA N, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA. | Open, Classified |  |  | 1793 | Not_Computer
Z-1047-2016 | 73151 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Large Volume Paracentesis Kit is packaged in a thermaformed tray which contains slots and snap-in holders for most components. Packed sterile in Tyvek bag. |  |  | January 28, 2016 | 2016 | Gi Supply | Customers have reported that the pinch clamp on the RP Bag was missing. Ascetic fluid from the abdominal cavity may splash User if bag is not clamped prior to removal. | Nonconforming Material/Component | The customers will be made aware of the defect, including pictures provided and are instructed to inspect the RP Bags in the kit prior to use. If the pinch clamp is missing, the customer is instructed to discard the RP Bag and use the sterile RP Bag replacements provided. |  | 390 | 390 Kits | Distributed in DC and the following states CA, FL, IA, KY, MA, MO, NH, NJ, NY, OH, OR, PA, SC, and TN. | Open, Classified |  |  | 390 | Not_Computer
Z-1054-2016 | 73329 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Milex Arcing Diaphragm Size 95; Model Number: MXWS95 |  |  | February 16, 2016 | 2016 | CooperSurgical, Inc. | The products have been identified to contain an incorrect size Diaphragm. | Package design/selection | Cooper Surgical sent an "Urgent - Medical Device Recall" letter dated February 16, 2016, to all affected customers. The letter identified the problem, the product and the action needed to be taken by the customer.Customers were instructed to discontinue use of the affected products and complete the attached Acknowledgement and Receipt form for replacement.For further questions, please call (203) 601-9818. |  | 2 | 2 units | US Distribution to the states of : OR and NY | Terminated | March 29, 2016 | 42 | 2 | Not_Computer
Z-1066-2016 | 72325 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips X-Ray Systems, Allura Xper with R8.2.16 |  |  | October 01, 2015 | 2015 | Philips Electronics North America Corporation | Upon initiating Fluoroscopy the user may encounter a user message Fluoro failed. | Software design | Philips sent a Field Safety Notice to all affected customers on October 1, 2015. The notice informs customers of the issue and the actions planned by Philips. Philips Field Service Engineers are to install Software R8.2.16.1 which addresses the issue. For further information or support concerning the issue, consignees should contact their local Philips Representative or call the Technical Support Line at 800-722-9377. |  | 196 | 196 devices | US Nationwide in the states of : AK, AL, AZ, CA, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, NJ, NY, OH, OR, PA, TN, TX, UT, WI, WV and the country of Canada. | Open, Classified |  |  | 196 | Computer
Z-1065-2016 | 73177 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100, Trilogy 200, Trilogy O2, Trilogy 202, |  |  | January 18, 2016 | 2016 | Philips Respironics | Software Issue | Software design | Philips Respironics has determined that affected devices may continue to be used in accordance with the device Instructions for Use and the Field Safety Notification that is being issued to customers. Philips Respironics will be providing a software update for affected devices free of charge. Customers will be notified by mail when the software update is available for download.For further questions please call ( 724 ) 387-7500. |  | 104508 | 104,508 units | Nationwide Distribution including Puerto Rico and Internationally to the US Virgin Islands. | Terminated | May 26, 2017 | 494 | 104508 | Computer
Z-1064-2016 | 73045 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Super Arrow-Flex Percutaneous Sheath Introducer Sets |  |  | January 12, 2016 | 2016 | Arrow International Inc | Arrow has issued a recall for these products due to potential packaging damage. The sterility of the products cannot be guaranteed. If non-sterile product is used there is a potential for infection to occur. | Packaging process control | Arrow International sent an Urgent Medical Device Recall Notification letter dated January 13, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recall@telefex.com. A customer service representative will contact them with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to the firm. If customers have no affected stock, they should also complete and the enclosed Recall Acknowledgement Form. Customers with questions should call Customer Service at 1-866-246-6990. For questions regarding this recall call 610-378-0131. |  | 56521 | 56,521 units | Nationwide Distribution including ALARAZCACOCTDEFLGAIAIDILINKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPARISCTNTXUTVAVTWAWIWV | Open, Classified |  |  | 56521 | Not_Computer
Z-0911-2016 | 73243 | No | NKM |  |  | Cardiovascular | PMA | MitraClip Clip Delivery System, product number MSK0101. The MitraClip System |  |  | February 04, 2016 | 2016 | Abbott Vascular | Abbott Vascular has recently received reports of cases on Clip Delivery System devices that contain the One-Way Actuator Knob where a user attempted implanting a MitraClip, but the Clip could not be detached from the delivery system due to a mandrel fracture. | Use error | The firm, Abbott Vascular, sent an "URGENT FIELD SAFETY NOTICE/MITRACLIP/CORRECTION" letter dated 2/4/16 to all the customers who received the MitraClip System. The letter described the product, problem and actions to be taken. The Customers are informed that their inventory is acceptable for safe use following the revised IFU steps described in the customer notification letter and that there is no need for the customers to return any product to Abbott Vascular. Customers with questions are instructed to contact their local Abbott Vascular representative.Customers were also instructed complete and return a Training Form/Effectiveness Check via their Abbott Vascular Representative, fax to 1-951-914-5951 or scan and email to AVRegulatoryCompliance@av.abbott.com. The firm is revising the IFU to incorporate the revised procedural steps.A press release dated 2/26/16 will be posted on the firm's website to inform the public that Abbott has initiated a voluntary safety notice regarding the MitraClip Delivery System to reinforce the proper procedures used to operate and deploy the device. Customers with questions or concerns are instructed to contact Abbott Vascular Representative or call Abbott Vascular at (800) 227-9902. |  | 7600 | 7,600 units total (2,300 units in US) | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Argentina Luxembourg Australia Malaysia Austria Mexico Belgium Netherlands Brazil New Zealand Canada Norway Colombia Philippines Corrientes Poland Czech Republic Portugal Denmark Egypt Romania Finland Saudi Arabia France Slovenia Germany Spain Greece Sweden Hong Kong Switzerland Hungary Thailand Indonesia Turkey Israel United Arab Emirates Italy United Kingdom Kuwait and Vietnam. | Terminated | May 12, 2016 | 98 | 7600 | Not_Computer
Z-1081-2016 | 72935 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Stryker Position Pro Mattress. |  |  | December 03, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Potential for power cord to melt; hazards include electric shock and fire. | Component design/selection | Stryker sent an URGENT MEDICAL DEVICE CORRECTION NOTIFICATION letter dated November 10, 2015 to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Normal business hours are Monday-Friday 8 a.m-5 p.m. (ET). |  | 2694 | 2694 | Worldwide Distribution - US Nationwide and the countries of Brazil, Canada, Chile, France, Hong Kong, Italy, Mexico, South Korea, Switzerland, United Kingdom. | Terminated | May 24, 2017 | 538 | 2694 | Not_Computer
Z-1076-2016 | 73008 | No | PHC | Infusion pump. | General Hospital | General Hospital | 510(k) | BD Intelliport Medication Management System Sensor |  |  | December 09, 2015 | 2015 | Becton Dickinson &amp; Company | The sterility of the product cannot be assured. This may result in increased risk of infection. | Under Investigation by firm | Becton Dickinson representatives notified their customers in person and a copy of the "Urgent Product Recall" letter and "Recall Response Form" dated 12/9/2015 was provided. The letter identified the reason for the recall; how to identify affected product; and the actions to be taken. The letter instructed customers to immediately review their inventory; complete the enclosed Recall Response Form and fax (1-201-847-4267 it to BD or email it to Becky_Saggau@bd.com even if you do not have any of the affected lot; and return all affected products with the completed Recall Response Form following instructions on the enclosed packing instruction. If customers have any questions or require assistance with the return of the recalled product and/or availability of replacement product, they were instructed to contact 1-201-847-4267 between 8AM and 5 PM ET Monday through Friday. |  | 250 | 250 units | US Distribution to:California and Utah. | Terminated | January 09, 2017 | 397 | 250 | Not_Computer
Z-1082-2016 | 73140 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Crosslock DVR ePAK -Sterile, Ref 81180050, are complete surgical sets including plates, screws, and instruments for orthopedic fixation |  |  | January 29, 2016 | 2016 | Biomet, Inc. | Components of the DVR ePAKs can oxidize during shipment and storage prior to use. A delay in surgery, less than 30 minutes, may occur if the surgical staff proceeds to re-sterilize the product. If oxidative material would flake off a component contained within the ePAK and enter the surgical field/wound an inflammatory reaction may occur. The inflammatory reaction may result in a foreign body r | Package design/selection | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated 1/29/2016 to the affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The customers were instructed to immediately locate and remove the identified device(s); carefully follow instructions on enclosed Response Form; complete and return Response Form within three (3) business days via email to CPWARFieldAction@zimmerbiomet.com prior to return of product, use priority carrier for your shipment; if you have further distributed the product, you MUST notify hospital personnel of this action. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 21268 | 21,268 | Worldwide distribution: US (Nationwide) and countries of: Panama; Korea; Japan; Netherlands; Ecuador; Chile; Canada; and Australia. | Open, Classified |  |  | 21268 | Not_Computer
Z-1075-2016 | 72979 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number 6801375, Unique Device Identifier Number 10758750001132; and VITROS 5,1 FS Chemistry System (Refurbished), Catalog Number 6801890, Unique Device Identifier Number 10758750001644; IVD. |  |  | December 21, 2015 | 2015 | Ortho-Clinical Diagnostics | Increased U90-382 or 6LU condition codes generated by VITROS 250, 350, 5,1 FS, 4600 and 5600 Chemistry Systems when using Calibrator Kit 9, Lot 954. A trend of complaints regarding customers actions following U90-382 or 6LU condition codes was noted (See RES 72289 - VITROS Calibrator Kit 9, lot 954 recall). | Error in labeling | Ortho Clinical Diagnostics sent an URGENT PRODUCT CORRECTION NOTIFICATION Letter (Ref. CL2015-201, dated December 21, 2015) via FedEx overnight courier and/or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address). The notification letter was sent to all customers who received the affected product. Foreign affiliates were informed by e-mail on 12/20/2015 of the issue and instructed to notify their consignees of the issue and required actions. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Technical Services at 1-800-421-3311. |  | 1405 | US = 994; Foreign = 1405 | Worldwide Distribution - US Nationwide including Puerto Rico, and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela | Open, Classified |  |  | 8387 | Not_Computer
Z-1083-2016 | 73293 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | TF Adaptive Gutta Percha Point Size SM3 50pk, Part No. 815-1541 |  |  | February 08, 2016 | 2016 | Ormco/Sybronendo | Ormco Corporation is recalling the TF Adaptive Gutta Percha Point Size SM3-50pk because the inner generic label is identified as ML3. The products outer package is identified as SM3 which matches the product inside. | Error in labeling | An Urgent Medical Device Recall letter dated February 8, 2016 was sent to all customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and fax back the enclosed acknowledgement form to 909-962-5605 to confirm receipt of the recall notification . |  | 5568 | 5,568 units total (2,070 units in US) | Worldwide Distribution - US Nationwide and the countries of Canada, Australia, Mexico, Europe, Hong Kong, MY, IN, JP, ID, SG, TH | Open, Classified |  |  | 5568 | Not_Computer
Z-1069-2016 | 73365 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Denali Filter-Femoral Delivery System |  |  | February 09, 2016 | 2016 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular, Inc. is recalling Denali Filter-Jugular/Subclavian Delivery System and Denali Filter-Femoral Delivery System because of the stop cocks potentially cracking. | Component design/selection | The firm sent customer notification letters on 02/09/16.The following instructions were provided to customers:1. Do not use or further distribute any affected product.2. Check all inventory locations within the institution for affected product code/lot number combination listed in the recall notice. If further distributed any of the product code/lot numbers, immediately contact that location, advise them of the recall, forward these instructions and have them return the affected product to BPV.3. Remove any identified product from your shelves.4. If have used the affected product, complete and return the attached Recall andEffectiveness Check Form indicating no product will be returned.Once the product affected by this recall has been removed from inventory;5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product that you intend to return. It is extremely important the information is received.6. Call Recall Coordinator Raye Seisinger at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST) or email at raye.seisinger@crbard.com. Once all information has been verified, the Recall Coordinator will issue either a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will provide replacement product for returned product.7. E-mail the completed Recall and Effectiveness Check Form to DenaliRecall@crbard.com or fax it to BPV at 1-800-994-6772. If cannot e-mail or FAX the form, please call the BPV recall coordinator (Raye Seisinger, at 1-800-321-4254 Option #2 Ext 2501) and report the required information verbally.8. A mailing label is enclosed for convenience to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and include the XC or XH number. All products should be returned to the following shipping address:Bard Peripheral Vascular, Inc.1415 W. 3'd |  | 3572 | 3572 | Worldwide Distribution-US (nationwide) including DC and PR, and the states of MN, TX, PA, OH, IL, OR, CA, SC, NC, AK, NY, MD, GA, NJ, AZ, KY, OH, WI, AL, AR, MO, FL, MA, MT, TN, LA, CT, NE, WV, VA, NV, ME, UT, CO, WY, AR, DE, NM, IN, MN, MS, NH, IA, MI, WA, VA, KS, HI, RI, SD, VT, and the countries of Hong Kong, Malaysia, Canada, Chile. | Open, Classified |  |  | 6476 | Not_Computer
Z-1077-2016 | 73263 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics ADVIA Centaur Calibrator E - 2-Pack |  |  | February 04, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Systems Calibrator E Lot 42 and 43 Bias with the Testosterone Assay | Under Investigation by firm | Siemens Healthcare issued an Urgent Medical Device Correction (UMDC) to all affected Siemens Healthcare Diagnostics customers in the United States on February 4, 2016 to be delivered to customers on February 5, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to customers outside the United States on February 4, 2016. These notices inform the customers of the lot to lot bias of Calibrator E with the ADVIA Centaur Testosterone assay and advise customers to move to Calibrator E kit lots ending in 44.Fax-back forms will be used to confirm that the customers received theUrgent Medical Device Correction. Questions, contact your Siemens Customer Care Center or your local Siemens technical supportrepresentative |  | 15224 | 15224 (2-pack) | Worldwide Distribution US (nationwide) and Internationally to:CanadaAR AU BD BR CA CH CL CN CO DO EC GT HK ID IE IL IN JP KR LK MM MX MY NI NP NZ PE PH PY SA SG TH TW UY VN ZA | Terminated | August 01, 2017 | 544 | 22198 | Not_Computer
Z-1067-2016 | 73100 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChem VELOCITY Automated Urine Chemistry, Catalog No. 800-3049, 800-3050, 800-3079, 800-3080, 800-3530, 800-7167, 800-7190, 800-7713, 800-7714, 800-7719, 800-7720 |  |  | January 13, 2016 | 2016 | Beckman Coulter Inc. | Iris International noted loose or missing analyte pads found in several locations (strip vial, strip provider module, strip conveyor system or waste container) that are undetectable by the iChemVELOCITY System. | Equipment maintenance | Iris sent an Urgent Medical Device Recall letter dated January 13, 2016 and an updated notification letter dated February 19, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions or concerns are instructed to contact their local support representative:Âż Via our website, at http://www.beckmancoulter.comÂż By phone, call 800 854-3633 in the United States and Canada. |  | 2661 | 2,661 units total (1,424 units in US) | Affected geographies: Argentina, Australia, Austria, Bangladesh, Bermuda, Bosnia and Herzegovina, Brazil, Brunei Darussalam, Canada, Chile, China, Columbia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mexico, Montenegro, Netherlands, New Zealand, Nigeria, Oman, Panama Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico , Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Republic, United States (including Puerto Rico), and Venezuela | Open, Classified |  |  | 2661 | Not_Computer
Z-1088-2016 | 73303 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm 02 Imaging System |  |  | January 29, 2016 | 2016 | Medtronic Navigation, Inc. | Screws holding the detector panel in as a result the detector panel could make contact with other parts with in the 0-armÂż 02 Imaging System, causing noise and vibration when being positioned or during 3D Imaging could lead to poor image quality | Process control | Medtronic Navigation issued a Field Safety Notice dated January 29, 2016, to notify them of a hardware correction in the unit. A Medtronic service representative will service the system and arrange for an inspection of the system. This representative will properly torque the eight (8) screws if they find them to be loose. In the meantime, if noise or vibration are observed during use of the system customers should contact their Medtronic representative to report the issue. Customers with questions were instructed to call 720-890-3160. For questions regarding this recall call 978-698-6000. |  | 1 | 1 unit | Worldwide Distribution to Austria only. | Terminated | March 10, 2016 | 41 | 1 | Not_Computer
Z-1084-2016 | 73428 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile wDR motorized portable diagnostic X-ray systems. |  |  | January 18, 2016 | 2016 | Sedecal S.A. | This is related to previous 806 Notifications about uncommanded movements of a motorized mobile x-ray system. Reason for Recall: Along with the previously noted encoder and gauges problems and corresponding fixes, we have observed that static electricity can also elicit the unwanted behavior. | Device Design | Philips Medical Systems sent an URGENT - Field Safety Notice letter dated August 6, 2014 was sent to all consignees of record informing them of a potential safety issue on the SEDECAL mobile units manufactured from May 2013 on. The firm requests the direct acounts with affected units should take the following steps:1) The unit is equipped with a safety switch in the drive handle. If the drive handle isreleased the unit will stop; 2) Do not try to correct unintended movement using the controls on the drive handle; 3) Instead, stop the system immediately by releasing the drive handle entirely; 4) Alternatively, the emergency button can be pressed; 5) Take special care when the system crosses metal strips or plates on the floor or enters an elevator, as these are situations where electrostatic discharge is likely to occur; and, 6) Should the customer feel uncertain regarding this action, he/she can contact Philips Healthcare.Philips is preparing a Field Action Kit to replace the existing arrestor strap by a new arrestor with higher resistivity. This will reduce the electromagnetic pulse to a level, which is safe to not disturb the electronics. A Philips Service Engineer will contact the direct accounts to distribute Field Action Kit.The firm states that should direct accounts need to communicate with Philips in regards to this recall that they reference FC0-71200138.If they need any further information or support concerning this issue, please contact their local Philips representative. |  | 349 | 349 systems | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI. | Open, Classified |  |  | 349 | Not_Computer
Z-1089-2016 | 71148 | No | EBI | Denture relining, repairing, or rebasing resin. | Dental | Dental | 510(k) | IMPAK ACRYLIC RESIN LIQUID FOR BOTH IMPAK &amp; IMPAK-PF POWDERS 1 qt. size (946 cc), Part Number 3306, and 11 oz. size (325 cc), Part Number 3748. |  |  | November 18, 2013 | 2013 | CMP Industries, Llc | During the current inspection of CMP Industries, the FDA Investigator discovered that the firm had conducted a recall of Impak Acrylic Liquid Batch 103113 in 2013 due to potential rust without notifying the district. | Nonconforming Material/Component | On November 18, 2013, CMP Customer Service contacted customers by telephone requesting return of the affected product and customers were issued a UPS Call Tag. CMP issued a letter to non-responding customers dated December 9, 2013, requesting return of the product because of the possibility of foreign particles. The letter states, the firm found that one of their processes could deliver a material that would pass their QC tests, but may introduce foreign material into the liquid. A customer service representative will be in contact with customers to arrange the return of the product. |  | 246 | 246 units | Worldwide Distribution -- USA, to the states of MI, NJ, PA, and TX; and, the country of Canada. | Completed |  |  | 246 | Not_Computer
Z-1090-2016 | 73357 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | ROTROL P Control for ROTEM delta Thromboelastometry System, For In Vitro Diagnostic Use Only. Quality control material for monitoring accuracy and precision of test carried out on the ROTEM delta Thromboelastometry System. |  |  | February 18, 2016 | 2016 | TEM Systems Inc | Customer complaints about failed target ranges for alpha angle. Firm has not completed the CAPA identifying the root cause. Projected completion of the CAPA by the firm is July 2016. | Nonconforming Material/Component | An Urgent Important Customer Information notice was sent to accounts on 2/17/2016. The notice requested that they discontinue use of the product and to contact them for a replacement. The letter also included a response form to be returned. |  | 481 | 476 boxes x 5 vials | Distributed in the states of AZ, AR, CA, FL, GA, HI, ID, IL, IN, LA, MD, MA, MN, MO, NE, NJ, NM, NY, PA, OH, SC, SD, TN, TX, UT, WA, WV and the country of Germany. | Open, Classified |  |  | 481 | Not_Computer
Z-1080-2016 | 72961 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Custom Assembly, Custom FCFFM x/8.2 cm Tubing Extension, Model Number: FCA026-001, Manufactured by: Value Plastics dba Nordson MEDICAL, 805 West 71st Street, Loveland, CO 80538, packaged in Bags consisting of 100, 500 or 1,000 stopcocks. |  |  | December 18, 2015 | 2015 | Value Plastics, Inc. | Potential for the parts to crack. | Process control | Customers were initially notified by telephone call; followed by recall notice via email on 12/18/15. The recall letter instructs customers to identify their customers and notify them of the product recall. Returned product will be quarantined and separated from acceptable materials and scrapped. |  | 33000 | 33,000 | 3 Consignees in NY, PA, MN. No foreign, VA/govt. | Open, Classified |  |  | 33000 | Not_Computer
Z-1092-2016 | 73393 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Central Station, Model 96102 is used to provide clinicians with central monitoring of patient data for those patients connected to networked Spacelabs Healthcare patient monitors and telemetry transmitters. |  |  | February 26, 2016 | 2016 | Spacelabs Healthcare Inc | The firm has received one report of values for patient height and weight being switched when input at the Xhibit Central Station, Model 96102, causing a bedside monitor Body Surface Area (BSA) calculation to be in error. | Software in the Use Environment | Spacelabs sent the letter Urgent Medical Device Correction -Spacelabs Healthcare Inc. - Xhibit Central Station, Model 96102, dated 26 February 2016 to US consignees.The firm will send a notification letter to international consignees on March 4, 2016.Consignees who have any questions about this corrective action, please contact Spacelabs at 1-800-522-7025 and select 2 for Technical Support. |  | 1075 | total 1075 units (614 units in the US and 461 outside the US) | distributed in AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MI, MN, MO, MS, MT, NC, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WY and Guam and in the following countries AFGHANISTAN, ARGENTINA, AUSTRALIA, BAHRAIN, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, EL SALVADOR, FRANCE, GERMANY, INDIA, IRAN (ISLAMIC REPUBLIC OF), ITALY, KUWAIT, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, and UNITED KINGDOM. | Open, Classified |  |  | 1075 | Not_Computer
Z-1091-2016 | 72486 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | BladderScan BVI 9600, Model Numbers: 0270-0451, 0270-0452, 0270-0754, and U270-0451. |  |  | November 30, 2015 | 2015 | Verathon, Inc. | The firm is providing customers with an updated Operations and Maintenance Manual for each of the BladderScan BVI 9600 and AortaScan AMI 9700 devices. The updated Operations and Maintenance Manuals clarify that these devices should not be used for the screening, detection, or diagnosis of abdominal aortic aneurysms (AAAs). | Device Design | Verathon sent an URGENT: MEDICAL DEVICE CORRECTION BLADDERSCAN BVI 9600 WITH AORTASCAN MODE &amp; AORTASCAN AMI 9700 letter dated November 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Verathon will manage other international consignees outside Europe internally and follow a similar approach as the firm requested Stericycle to perform. Verathon will instruct the companys foreign distributors to communicate the Customer Correction Letter and provide the revised OMMs to their international customers.This current Medical Device Correction is intended to reinforce the message in the May 2014 communication regarding the potential variance in measurement and to provide you with updated Operation and Maintenance Manuals that further reinforce the devices intended use and indications for use and that revise warning information, including warnings specified in the May 2014 communication.Customers are instructed to destroy all existing Operations and Maintenance Manuals and replace them with updated versions and complete the information on the Medical Device Correction Reply Form and fax or email the completed form within 10 business days to: Fax: (877) 907-9956 or Email: verathon3108@stericycle.com.Customers with any questions can contact your Verathon representative or Verathon Customer Care at (866) 875-0382. Verathon has live agent support on Monday through Friday from 8am to 5 pm Pacific Time.Customers should report any suspected malfunction or adverse event related to any AortaScan AMI 9700 devices and any BladderScan BVI 9600 devices to Verathon Customer Care at (800) 331-2313 |  | 24230400 | 2524 units total (1441 units of BVI 9600 and 820 units of AMI 9700 in US; 209 units of BVI 9600 and 54 units of AMI 9700 OUTSIDE the US | Worldwide Distribution - US (nationwide) and to the countries of : Guam, United Arab Emirates, Australia, Belgium, Canada , China, Czech Republic, France, Germany, Luxembourg, Netherlands, Norway, Saudi Arabia, Sweden , Taiwan, and United Kingdom | Open, Classified |  |  | 24230400 | Not_Computer
Z-1093-2016 | 73173 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Verso Shoulder Forked Retractor (Comprehensive Instrumentation 2 Prong Retractor) |  |  | February 16, 2016 | 2016 | Biomet, Inc. | Zimmer Biomet Inc. Verso Shoulder Forked Retractor (Comprehensive Instrumentation 2 Prong Retractor) supplied by Kirkstall Precision was made according to the W1.4034 (420 stainless) which does not conform to 420 S29 indicated on the print.All units manufactured to date were manufactured using 420 stainless. Therefore, all lots of PN: 402852 are considered in scope. | Mixed-up of materials/components | Zimmer Biomet sent an Urgent Medical Device Recall letter dated February 16, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The following actions are REQUIRED: Immediately locate and remove the identified device(s) listed below from circulation.Carefully follow the instructions on the enclosed Response Form . Email a copy of the response form to CPWARFieldAction@zimmerbiomet.com prior to return of product.Use priority carrier for your shipment.If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed Dear Risk/Recall Manager notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products.Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA:Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1261 | 1261 units | Worldwide Distribution - US including CA, NY, VA, IL, IN, SC, NC, AR, MO, ND, AL, PA, MI, WA, OR, LA, FL, GA, OH, KS, KY, TX, MA, AZ, NJ, UT, IO, NM, WV, MS, NE, and MD OUS: Australia, Chile, Columbia, Costa Rica, Argentina, Korea, Japan, Panama, Malaysia, Puerto Rico, Canada, and The Netherlands. | Terminated | May 30, 2017 | 469 | 1261 | Not_Computer
Z-1035-2016 | 73014 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Titanium Single-Use Video Laryngoscope. It is a single-use video laryngoscope system (GlideScope Titanium System) consists of a sterile-packaged, single-use video laryngoscope, multi-use Smart Cable and portable GlideScope video monitor. |  |  | January 29, 2016 | 2016 | Verathon, Inc. | The firm has become aware of the potential for disruption (flickering) in the video laryngoscopy image when GlideScope Titanium SU laryngoscope blades are in use. Video "flickering" appears as the intermittent break-up of the on-screen video image, appearing as either distorted horizontal or vertical bars of displaced video signal. | Process control | The firm, Verathon, began mailing the "URGENT: MEDICAL DEVICE RECALL"GLIDESCOPE VIDEO LARYNGOSCOPE TITANIUM SINGLE-USE BLADES letter to their consignees on January 29, 2016. The letter describes the product, problem and actions to be taken.Verathon will instruct the company's foreign distributors to communicate the Customer Recall Letter to their international customers.Verathon has contracted with a business services company, Stericycle, to coordinate the initial consignee communications in the United States, Australia and Canada. Stericycle will send the Customer Recall Letters, and conduct consignee follow-up communications for return of the Recall Response Forms.Verathon will replace all affected GlideScope Titanium Single Use (SU) blades that remain in your facility's inventory with new product that has been passed the enhanced screening test for video flickering.Customers were advised to take the following actions:-Fill out the attached Recall Response Form and return it to Verathon by fax (888) 943-5170, or email, verathon5165@stericycle.com. Please return the form even if you do not have any blades subject to the recall.If consignees need to expedite a product replacement contact Verathon representative or Verathon Customer Care at (800) 331-2313 (U.S. and Canada). The firm has live agent support Monday through Friday from 8:00am to 5:00pm Eastern Time. Consignees can also send email to CCTechSupport@verathon.com.Consignees who have any questions about this Product Recall, please contact Verathon representative or call (888) 943-4207 (U.S. and Canada) or send email to verathon5165@stericycle.com.Please report any suspected malfunction or adverse event related to any GlideScope Video Laryngoscope devices to Verathon Customer Care at the telephone numbers above and to FDA's MedWatch Program (1-800-FDA-1088 or www.fda.gov/medwatch). |  | 7733 | 7733 units (6377 units in the US and 1356 units International) | Worldwide Distribution - US (nationwide) and to countries of: Australia, Belgium, Canada, Denmark, France, Germany, Hong Kong, Italy, Luxembourg, Saudi Arabia, Singapore, Spain, Sweden, The Netherlands, U.A.E., and United Kingdom. | Terminated | March 01, 2017 | 397 | 7733 | Not_Computer
Z-1094-2016 | 73034 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Battery Pack |  |  | December 23, 2015 | 2015 | Pro-Dex Inc | The sterile battery may contain particulates within the sterile package. | Process change control | A customer notification letter was sent 12/23/15.The letter instructs the customer to cease distribution and use of the recalled lots. It also recommends that the inventory and distribution records be evaluated immediately to ensure that affected product is removed from any potential point of use.Return product to:Returns Department (Recall complaint ClS-001)Pro-Dex, Inc2361 McGaw AvenueIrvine, California 92614 USAReference the Recall Complaint # on the outside of the shipping box.If any questions or concerns, please contact theQARA group at Pro-Dex at 949 769 3200. |  | 1995 | 1995 total | Distributed in FL. | Terminated | July 21, 2017 | 576 | 1995 | Not_Computer
Z-1057-2016 | 73266 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | FiberOptix Ultra 8 IAB: 8Fr 40cc; Product Code: IAB-05840-LWS |  |  | February 10, 2016 | 2016 | Arrow International, Inc., Division of Teleflex Medical Inc. | The sheath body may become separated from the sheath hub. If the separation occurs, there is a potential for bleeding from the device. | Under Investigation by firm | The firm, Teleflex Arrow International, sent an "Urgent Medical Device Recall Notification" letter dated 2/11/2016 via FedEx to its Customers. The letter described the product, problem and actions to be taken. The customers were informed of the reason for recall and were instructed to take the following actions: 1. If you have any products referenced in the attached enclosure, immediately discontinue use and quarantine those products appropriately. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recalls@teleflex.com. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Arrow International. 3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 13405 | 13,405 US and 33,735 OUS in total | Worldwide Distribution: US (Nationwide) and countries of: Argentina, Australia, Belgium, Brunei Darussalam, Brazil, Bahamas, Canada, Chile, China, Colombia, Ecuador, Guatemala, Hong Kong, Indonesia, India, Japan, Korea (Republic of [South] Korea), Myanmar, Mexico, Malaysia, Nepal, New Zealand, Panama, Peru, Philippines, Pakistan, Singapore, Thailand, Trinidad and Tobago, Taiwan, Uruguay, Viet Nam, and South Africa (Zuid Afrika). | Open, Classified |  |  | 13405 | Not_Computer
Z-1098-2016 | 73421 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Intra operative MRI Systems: Neuro II-SE , Neuro III-SV, iMRI 1.5T A, and iMRI 3T S. |  |  | February 10, 2016 | 2016 | Deerfield Imaging | Emergency helium venting lines (i.e. quench lines) may not have been properly installed. | Device Design | Consignees were sent a IMRIS "Urgent Field Safety Notice" letter dated February 10, 2016. The letter described the problem and the product involved in the recall. The firm also stated that it is important that customers review the portion of the quench line vent described and ensure that it is installed to the specifications supplied as part of the Planning Guide. Customers were reminded that it is important to have an emergency plan in place in the event Helium gas escapes into the magnet room or any other area. For questions customers can contact their local customer support hotline or +1-866-475-0525. |  | 49 | 35 USA, 14 OUS | Worldwide Distribution -- US, including the states of CA, CT, FL, GA, IL, IN, MD, MA, MN, MO, NH, NC, OH, OR, PA, TX, UT, VA, and WI; and ,countries of CHINA, CANADA, AUSTRALIA, FRANCE, JAPAN, and GERMANY. | Terminated | April 05, 2017 | 420 | 49 | Not_Computer
Z-1096-2016 | 73222 | No | NPM | Bone grafting material. | Dental | Dental | 510(k) | NuOSS Cancellous |  |  | August 25, 2015 | 2015 | Collagen Matrix Inc | On 8/25/2015, Collagen Matrix, Inc discovered that one of the 171 units of NuOss Cancellous product shipped to a distributor may have included an incorrect size. | Unknown/Undetermined by firm | Collagen Matrix Inc. notified their sole customer on November 17, 2015.On August 25, 2015, during a review of inventory, tt was discovered that of the 171 units product shipped, 19 units may have been the incorrect size.For further questions please call (201) 405-1477 |  | 171 | 171 units | US Distribution to the state of : MA | Open, Classified |  |  | 171 | Not_Computer
Z-1097-2016 | 73206 | No | NPM | Bone grafting material. | Dental | Dental | 510(k) | NuOss XC Sinus |  |  | May 13, 2013 | 2013 | Collagen Matrix Inc | During the 36 month real time shelf life testing of NuOss XC Sinus (BMCU-Umbrella) to verify the 3 year shelf life designated for the device, 3 of 13 products did not pass the recovery test, which requires that the compressed product fully expand within 1 minute of hydration time. | Process control | The firm, Collagen Matrix Inc., sent an email dated May 14, 2013 to its sole customer. The email described the product, problem and actions to be taken. The customer was instructed to return all product with expiration dates 12/2014 or earlier under Returned Goods, the firm will replace these products with new inventory; and return to Collagen Matrix, Inc., 15 Thornton Road, Oakland, NJ 07436 or quarantine products with expiration date 1/2016 until further notice.If you have any questions,call 201-405-1477 x 304 or email: www.collagenmatrix.com. |  | 347 | 347 units | US Distribution to state of: MA. | Terminated | June 21, 2017 | 1500 | 347 | Not_Computer
Z-1163-2016 | 72900 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Product 65 consists of all product under product code: KTT and same usage: |  |  | January 11, 2016 | 2016 | Zimmer Manufacturing B.V. | As a result of the insufficient sealer calibration data for product packaged in firm's Building II between August 2010 and April 2013. The affected products are sterile. | Packaging | Zimmer Biomet sent an Urgent Medical Device Recall Letter, dated January 11, 2016, to Distributors, Sales Representatives, Operation Managers, and Risk Managers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to 1.Review the notification and ensure that relevant personnel are aware of the contents.2. Assist your Zimmer Biomet sales representative with the quarantine of any affected product.3. Your Zimmer Biomet sales representative will remove the recalled product from your facility.4. Complete and return the attached Certificate of Acknowledgment form to corporatequality.postmarket@zimmerbiomet.com.5. If after reviewing this notification you have further questions or concerns please call Customer Service at 1-800-348-2759, or contact your Zimmer Biomet Sales Representative. |  | 5538 | 5538 | US Nationwide Distribution | Terminated | April 13, 2017 | 458 | 665035 | Not_Computer
Z-1172-2016 | 73253 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Accu-Chek Inform II Base Unit as a part of the Accu-Chek Inform II Blood Glucose Monitoring System |  |  | February 11, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Accu-Chek Inform II Base Unit might produce physical transmission errors in the form of data loss in the communication between the meter and the Data Management Systems (DMS). The issue can lead to the data loss or in the worst case to an erroneous assignment of the patient data (patient mismatch). The issue will only occur at sites using POTCT1-A communication via USB. | Software Design Change | Roche sent an URGENT MEDICAL DEVICE RECALL CORRECTION letter dated February 11, 2016, via UPS Ground (receipt required) to the consignee. The Consignee instructions were as follows:Âż If currently using USB connectivity, follow the instructions outlined in this Urgent Medical Device Correction (UMDC) to determine if your base unit is affected by the issue.Âż Until the new software is available, do not use USB connectivity if the base unit is affected by this issue; switch the connectivity method to Ethernet (TCP/IP) or Wi-Fi.Roche is not requesting return of affected product. Any returned product will be managed according to current procedures.Consignees with questions were instructed to contact the ACCU-CHEK Customer Care Service Center at 1-800-440-3638.For questions regarding this recall call 317-521-3911. |  | 97529 | US 5,604 devices, OUS 91,925 | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV. | Terminated | May 30, 2017 | 474 | 97529 | Computer
Z-1171-2016 | 73129 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Sekure Chemistry Sekisui Diagnostics Acetaminophen L3K Assay, In-vitro Diagnostic |  |  | January 22, 2016 | 2016 | Sekisui Diagnostics P.E.I. Inc. | Crystal formation in the Color Reagent (R2) when left on board automated analyzers causing samples to under recover | Device Design | Sekisui sent an Important Field Correction - Customer Notification letter dated January 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Discard all reagent in house with a current expiration date of 2016-09.2. Update the expiration on all remaining kits on hand with the new expiration dates provided 3. When adding additional reagent to the analyzer, a new reagent wedge should be used.4. If crystals are observed within reagent stored on-board, reagent should be discarded and replaced with fresh reagent.5. Complete, sign and return the attached CONFIRMATION OF NOTIFICATION form by FAX to the number provided. This will indicate receipt of this Field Correction, act as confirm action of compliance with the actions,and allow for issuance of replacement material for expired product on hand (lots 47677, 47694, 47695).Questions contact Technical Services at 1-800-565-0265 or viaemail at PEIDiagnosticTechnical@sekisui-dx.com |  | 14407 | 14,407 kits. | Worldwide Distribution - US Distribution to the states of : AZ, CA, FL, IA, ID, IL, MA, , MN, MT and ND., and to the countries of Canada, France and Hong Kong. | Open, Classified |  |  | 14407 | Not_Computer
Z-1174-2016 | 73183 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | CORDIS AVANTI + CATHETER SHEATH INTRODUCER SYSTEM |  |  | December 31, 2015 | 2015 | Cordis Corporation | Incorrect cannula of the sheath introducer (smaller than intended). | Mixed-up of materials/components | The firm, Cordis, sent an 'Urgent MEDICAL DEVICE CORRECTION (Field Safety Notice) dated 12/31/2015 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:1) Read the Description of the problem section carefully to fully understand theissue involved.2) Check your inventory to determine if you have any remaining affected productin your possession. Check all storage and usage locations. The purpose is toidentify the product, not to remove the product.3) Keep a copy of this notice with any affected product.4) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis by fax to (908) 429-8287 Alternate fax (904) 928-5077. number or scan and email to MiamiFRA@its.jnj.com.5) Share this letter with anyone who needs to be informed in your facility and in any facility where potentially affected devices may have been transferred.6) Maintain awareness of this notice until all affected product has beenconsumed. If vessel dilator to sheath incompatibility is experienced, you canreport the incident through the standard complaint process (1-800-327-7714Option 3).For any health care professional with medical concerns, please contactCordis to speak to a clinician at 1-800-327-7714 Option 3, Mondaythrough Friday from 9:00 AM to 5:00 PM EST.For questions related to the Correction that are not adequately addressedin this letter, please contact the Cordis Field Action Coordinator at (786) 313-2365.For questions related to the Acknowledgement Form and its return, pleasecontact Cordis QA at (786) 313-8730." |  | 1985 | 1,985 units | US Nationwide Distribution including states of:AL, AZ, CA, CO, CT, FL, GA, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NH, NJ,NY, OH, OR, PA, RI, TN, TX, VA, WA, WV, and Hawaii. | Open, Classified |  |  | 1985 | Not_Computer
Z-1173-2016 | 73208 | No | NPM | Bone grafting material. | Dental | Dental | 510(k) | ZCORE Porcine Xenograft Particulate in Syringe |  |  | November 17, 2015 | 2015 | Collagen Matrix Inc | The distributor, Osteogenics Biomedical, notified the firm that the syringes were not working properly and the syringes were being held in quarantine. | Component design/selection | Collagen Matrix Inc. notified their sole customer on 11/17/2015. |  | 200 | 200 units | US in the state of TX | Open, Classified |  |  | 200 | Not_Computer
Z-1176-2016 | 73422 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The Spirit TM Select bed is a Med-Surg bed intended to support and transport |  |  | February 18, 2015 | 2015 | CHG Hospital Beds Inc | It was identified that in some situations the bed exit alarm may not function as intended. The software code for the bed exit system has the potential to auto-reset erroneously. In some situations, the software code does not allow enough time for the weight value to fall to zero once the patient egresses from the bed, in which case, the bed has the potential to reset the bed exit alarm even without a patient on the bed.The bed exit on the Spirit Select bed may incorrectly reset if the Auto Reset option is selected. A patient fall may result in soft tissue or bone injury which may require medical or surgical intervention to prevent permanent disability. If the bed exit alarm incorrectly resets then medical assistance after a fall or during entanglement may be delayed. Patients with dementia or disorientation are at greatest risk. There have been two reported injuries associated with these issues to date. A Medical Risk Assessment determined that there is a remote probability of occurrence of harm requiring medical intervention. | Software design | On 2/18/2016, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 1892 | 1892 | Distributed US (nationwide) and in Canada. | Terminated | May 23, 2017 | 825 | 1892 | Computer
Z-1177-2016 | 73213 | No | HJK | Polymethylmethacrylate (PMMA) diagnostic contact lens. | Ophthalmic | Ophthalmic | 510(k) | Volk G-4 Small Ring, No Flange, Gonio Lens with case, UPC: 811608013602 |  |  | February 02, 2016 | 2016 | Volk Optical Inc | The firm discovered that the incorrect lot number was engraved on the product. | Employee error | Volk Optical sent an Urgent Medical Device Recall letter dated February 4, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Also, if product was further distributed customers should identify their customers and notify them at once of the product recall and include a copy of the letter. The affected devices should be returned to Volk with an assigned Return Merchandise Authorization (RMA) number. Customers should contact Volk at (440) 942-6161 or toll free at (800) 345-8655 as soon as they receive the notification to setup a RMA for their device(s). Customers were instructed to complete and return the enclosed response form as soon as possible. Customers with questions should call (440) 510-0745. |  | 9 | 9 lens | Worldwide Distribution - US (Virginia only) Internationally to Canada, France &amp; Indonesia. | Terminated | April 26, 2017 | 449 | 9 | Not_Computer
Z-1178-2016 | 73267 | No | GDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pilling Modular Verres Body/Obturator, Product Code 140880 |  |  | January 08, 2015 | 2015 | Teleflex Medical | Packaging error in which customers who ordered product code 140880 (Verres cannula for use with a separately available needle tip) were incorrectly provided with product code 140878 (Verres cannula with integral needle tip). | Packaging process control | Teleflex sent an Medical Device Advisory Notification letter dated January 8, 2016, to all affected customers notifying them of the packaging error.Our records indicate that you have received products that are subject to this notification.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at recalls@teleflex.com or 1-866-246-6990. |  | 8 | 8 | US Distribution to the state of : UT | Terminated | April 14, 2017 | 827 | 8 | Not_Computer
Z-1180-2016 | 73009 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Quick Connects. Steris Corporation. Model numbers: QKC1736E, QKC1736INT, QKC1737E, and QKC1737INT. Each Quick Connect contains a flow unit consisting of tubing, tethers, and adapters. For sterilant and rinse water used in endoscope washers. |  |  | December 04, 2015 | 2015 | Steris Corporation | The firm discovered that one of the hysteroscope models listed for processing in four existing Quick Connects has one inlet port rather than two inlet ports as indicated in the Quick Connect processing instructions. | Nonconforming Material/Component | On 1/18/2016 the firm sent Voluntary Field Correction NOtices to their customers. |  | 809 | 809 units | Worldwide distribution. US nationwide, Australia, Bahrain, Bulgaria, Canada, Greece, Italy, Malaysia, Mexico, Qatar, Russia, Spain, and United Kingdom. | Open, Classified |  |  | 809 | Not_Computer
Z-1181-2016 | 73096 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, PB980 Ventilator (980xxxxxxxx). |  |  | January 15, 2016 | 2016 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Graphical user interface (GUI) unresponsive to touch and Loss of primary ventilation under certain circumstances. Covidien Respiratory and Monitoring Solutions, now a part of Medtronic, issued a field corrective action notice for two issues on all models of Puritan Bennett 980 (PB980) ventilator. | Under Investigation by firm | Customers were notified via an Urgent Medical Device Safety Notification, dated February 8, 2016, on 1/14/16. The letter identified the affected device and the issues involved. Consignees are requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Medtronic is investigating the root cause of these conditions and will provide a service update to resolve the issues as soon as corrections can be implemented. Customers are to complete the attached form and return it per the instructions to confirm receipt and understanding of the information. Questions or issues should be directed to the Technical Support Department at 800-255-6774. |  | 1864 | 1,864 units | Worldwide Distribution -- US, including but not limited to the territory of Puerto Rico; and Ireland, Jordan, Saudi Arabia, Bahrain, Australia, Japan, Israel, Turkey, Chile, South Africa, Costa Rica, United Arab Emirates, Qatar, Peru, Panama, Singapore, Kuwait, Lebanon, Brazil, Mexico, Colombia, Russia, Uruguay, Slovenia, Morocco, Argentina, Egypt, and Romania. | Terminated | July 11, 2017 | 543 | 1864 | Not_Computer
Z-1179-2016 | 73238 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Alaris Pump VersaSafe Infusion Set, Model No. 2122-0007 |  |  | February 15, 2016 | 2016 | CareFusion 303, Inc. | CareFusion is recalling the Alaris VersaSafe Infusion Set because of separation and/or leakages at the VersaSafe split septum port and tubing may lead to leakage of the administration set. | Other | A customer notification letter dated 2/15/16 will be sent to customers to inform them that CareFusion is recalling the Alaris VersaSafe Infusion set number 2122-0007, Lot number 13085603. The letter informs the customers that Separation or leakage at the VersaSafe split septum port and tubing may lead to leakage of the administration set. The separation and leakage may occur during infusion. The letter informs customers of the potential risk and actions to be taken. Customers with recall related equations are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PST. Customers with adverse event reports are instructed to contact Customer Advocacy at (888)812-3266, customerfeedback@carefusion.com, 24 hrs/day, 7 days/wk. Customers are instructed to complete and return the enclosed mandatory customer response card to acknowledge the receipt of the customer communication. |  | 5760 | 5,760 units | Michigan | Open, Classified |  |  | 5760 | Not_Computer
Z-1187-2016 | 73438 | No | NEY | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Short Percutaneous Antenna with Thermosphere" Technology |  |  | February 23, 2016 | 2016 | Covidien Medical Products (Shanghai) Manufacturing L.L.C. | Medtronic is voluntarily recalling specific item codes and production lots of Covidien Emprint" Percutaneous Antenna with Thermosphere" Technology due to the potential for the generator to shut down in the precence of high temperature measurements in the Emprint" Percutaneous Antenna. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated February 23, 2016 to all affected customers via Federal Express or Certified Mail. The letter identified the affected product, problem and actions. Customers were advised to quarantine and discontinue use of the affected product and to Return product. For questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 4920 | 4,920 units | Worldwide Distribution - US Nationwide, VA/govt/military consignees. | Terminated | June 21, 2017 | 484 | 4920 | Not_Computer
Z-1182-2016 | 73452 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) products versions 4.6.1 to including 11.9 and McKesson Radiology (MR) products 12.0 and 12.1.1. |  |  | February 18, 2016 | 2016 | Mckesson Medical Immaging | McKesson has identified a design deficiency where under rare circumstances, imported images/studies may re-use a non-unique image directory. This issue may cause incorrect images to be displayed for a patient. | Software design | McKesson sent an Urgent Field Safety Notice dated February 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate actions to take to reduce the risk to patientAll customers receiving this safety notice must contact McKesson Support as soon as possible at 1-800-663-2533 US/Canada and International Toll Free Radiology - 00 800 626 20009:1) To acknowledge that they have read and understood this Field Safety NoticeAnd2) Arrange installation of software update/upgrade.This notice should be distributed to all personnel within your organization who need to be aware of this issue and the required actions, and alert other affiliated parties that are affected by this Field Safety Notice.For further questions, please call (404) 338-3556 |  | 631 | 631 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the states of : AL, AK, AZ, AR, CA, CO, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY., and the countries of : Guam, Great Britain, France, Canada, Ireland and Australia | Open, Classified |  |  | 631 | Computer
Z-1184-2016 | 73403 | No | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | Prime Series Stretchers (Fifth Wheel, Big Wheel and Zoom) |  |  | February 05, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | The hydraulic jack assemblies (PN: 0753-002-101, 0753-002-070, 0753-002-001, and 0753-002-001) were manufactured using the incorrect oil. The oil used does not contain the same wear resistance additives as the specified oils and therefore its use increases the potential for wear which may cause a leak event or jack drift. Testing was done to evaluate the jacks with incorrect oil and leak events we | Mixed-up of materials/components | On 2/5/2016, URGENT MEDICAL DEVICE NOTIFICATION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m.-5 p.m. (ET). |  | 36 | 36 | Distributed in DC and the states of AL, CA, FL, MN, MT, NC, NE, NY, PA, RI, TN, TX, UT, and WV. | Open, Classified |  |  | 36 | Not_Computer
Z-1183-2016 | 73479 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens Syngo Dynamics-a Picture Archiving and Communication System (PACS) Model Numbers: 10091805, 10091807, 10091637, 10091673 |  |  | January 25, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens' conducting a recall due to a potential issue when using the measurement package of the VA10 version of syngo Dynamics. | Software design | The firm, Siemens Healthcare, sent a "Customer Advisory Notice" dated 1/19/2016 to customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions provided on what steps the user can take to avoid the issue and to forward the notice to the new owner, if you sold this device/equipment and its no longer in your possession. Siemens is preparing a modification of syngo Dynamics that will resolve this potential malfunction.If you have any question, contact Regulatory Technician Specialist at 610-448-6478 or email: Anastasia.sokolova@siemens.com. |  | 228 | 228 units | US Nationwide Distribution including states of: MN, NC, NE, FL, TX, CA, AR, MI, MA, IL, DE, VA, SC, OH, CO, WI, MO, NY, CT, UT, PA, WA, NJ, GA, IN, MD, TN, LA, OR, AZ, ND, KS, AL, NM. | Open, Classified |  |  | 228 | Computer
Z-1186-2016 | 73352 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Stockert Heater-Cooler 3T; Item Number 16-02-8, H3T/240V/60Hz ; 16-02-82, H3T208V/60Hz ; 16-02-85, H3T/120V/60Hz. Device Listing E572259 |  |  | August 15, 2012 | 2012 | Sorin Group USA, Inc. | On August 15, 2012, Sorin informed consignees via letter of changes made to create Version 12 of the Heater-Cooler 3T IFU. The revised IFU included enhancements to achieve water quality that meets the standard for German drinking water. | Device Design | Sorin Germany, via Sorin Colorado, distributed an important information letter to its US customers on August 15, 2012 informing them about an update of the cleaning and infection procedure for the U.S. market that were included in Revision 12 of the IFU. |  | 1125 | 1125 units | Worldwide Distribution - US Nationwide, VA/govt/military | Terminated | May 26, 2017 | 1745 | 1125 | Computer
Z-1192-2016 | 73150 | No | LGW |  |  | Neurology | PMA | Algostim LLC, Algovita Implanted Pulse Generator stimulator, Model 2408 ( 3 x 8 channel) |  |  | January 26, 2016 | 2016 | QiG Group LLC | The implantable pulse generator (IPG) model 2408 had suboptimal performance establishing a connection to the external devices unless they were brought into close proximity (as close as 5 cm). The following functions may not be available: Reprogramming the stimulator, Adjusting stimulation parameters (on/off, intensity modifications, stimulation program selection), and Recharging the stimulator. | Nonconforming Material/Component | Consignee was sent on 1/26/2016 a Algostim, LLC "Urgent Field safety Notice" dated January 21, 2016. The letter described the problem and the product involved in the recall. The letter provided Background information as well as described the Clinical Implications, Rate of Occurrence, Recommendations and Mitigations. Advised consignee to deliver the respective letter to each physician who have implanted affected product and collect signed Response forms. |  | 40 | 40 | Distributed in Germany and Luxembourg. | Terminated | April 05, 2017 | 435 | 105 | Not_Computer
Z-1190-2016 | 73363 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | InstinctÂż JavaÂż System (IFU reference: 046WAN0000T) . |  |  | February 12, 2016 | 2016 | Zimmer Spine, Inc. | Zimmer Biomet is initiating a Medical Device Correction Notice to provide updated cleaning and disinfecting instructions for use of the InstinctÂż JavaÂż System (IFU reference: 046WAN0000T) and Universal ClampÂż System (IFU reference: SNA027-N-90001). | Process control | Consignees were sent on 2-12-2016 a Zimmer Biomet "Urgent Medical Device Correction Notice" letter dated February 12, 2016. The letter was either addressed to Hospital Representative or to Distributors, Sales Representatives, and Distribution Operations Managers. The letters described the problem and the product involved in the recall. Advised consignees to immediately utilize the updated and attached Instructions For Use. Requested consignees to complete the "Acknowledgement of Field Notification" form. For questions contact Customer Service at 1-866-774-6368 or e-mail corporatequality.postmarket@zimmerbiomet.com. |  | 2506 | 2506 | US (nationwide)-including the states of AL, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, LA, ,MD, MA, MI, MN,MO, NV, NH, NY, NC, OH, OK,. PA, TN,TX, WA, and WI. | Open, Classified |  |  | 5990 | Not_Computer
Z-1191-2016 | 73314 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si (IS3000) and da Vinci Xi (IS4000) Surgeon Side Console (SSC)Touchpad; Model number 187505-04; Intended to assist in the accurate control of Intuitive Surgical Endoscopic Instruments. The system can also be employed with adjunctive mediastinotomy to perform coronary |  |  | February 16, 2016 | 2016 | Intuitive Surgical, Inc. |  | Other | Urgent Medical Device Recall letters were sent on March 2, 2016. Letters advised users of the issue and that Intuitive Representative will visit each site to replace the affected compact flash card. |  | 542 | 542 (351 SCC Touchpads and 191 PSC touchpads) | US and Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Finland, France,Germany, Hong Kong, India, Israel, Italy, Japan, Malaysia, Monaco, Netherlands, Philippines, Portugal,Qatar, Romania, Russian Federation, Saudi Arabia, South Korea, Sweden, Switzerland, Taiwan, Turkey, United Kingdom and Venezuela. | Open, Classified |  |  | 542 | Not_Computer
Z-1195-2016 | 73204 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Arthrex IBalance TKA Tibial Tray Size 2 - Catalog # AR-503-TTTC, Size 3 - Catalog # AR-503-TTTD, Size 4 -Catalog # AR-503-TTTE, Size 5 - Catalog # AR-503-TTTF, Size 6 - Catalog # AR-503-TTTG , Size 7 -Catalog # AR-503-TTTH, and Size 8 - Catalog # AR-503-TTTJ. |  |  | December 17, 2015 | 2015 | Arthrex, Inc. | Devices reported to have a smooth texture to the outer surface of the metal, which differed from a previous that had a textured outer surface. | Nonconforming Material/Component | On February 5, 2016 Arthrex sent Urgent Medical Device Voluntary Recall letters to their distributors and Surgeons/Medical Facilities reporting the removal. The letter identified the reason for the recall, the potential patient impact, and actions to be taken. Users were to discontinue use of the identified affected devices and follow the instructions provided to return the affected devices that have not been implanted. To return the devices, contact Arthrex at 1-866-267-9138 or by fax at 1-239-591-6943 or via e-mail at Complaints@Arthrex.com. Customers must complete and return the Acknowledgement and Receipt. |  | 2378 | 2,378 units. | Nationwide Distribution -- AR, AZ, CA, CO, DC, DE, FL, GA, IL, IN, KS, MA, MD, MI, MN, MO, NC, NJ, NM, NV, NY, OH, OK, PA, RI, TN, TX, WA, WI and District of Columbia. | Open, Classified |  |  | 2378 | Not_Computer
Z-1175-2016 | 73064 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba DRAD-3000E FPD Wireless System |  |  | January 14, 2016 | 2016 | Toshiba American Medical Systems Inc | When a user performed radiography using the wireless FPD, a message window appeared on themonitor stating that imaging transmission was not completed with no radiographic image. It also indicated to select the "OK" button to re-acquire the image data or to select theÂż'Cancel" button to cancel the re-acquisition. As instructed the user selected "OK" and the same messagewindow appeared. This sam | Radiation Control for Health and Safety Act | Toshiba's planned action: 1."Toshiba America Medical Systems, Inc. (TAMS) will advise contact customers to stop using the system if any abnormalities are found using the system.2. TAMS will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that TAMSwill without charge, remedy the defect or bring the product intocompliance.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and DrugAdministration (FDA) district office coordinator noted below are to beincluded in the notification.2. TAMS will implement this CAP by June 30, 2016.Customers with questions were advised to call 301-796-5910.For questions regarding this recall call 714-730-5000. |  | 28 | 28 | Nationwide distribution to PA, UT, FL, NY, LA, WV, NJ, OH, WI and TX. | Open, Classified |  |  | 28 | Computer
Z-1229-2016 | 73121 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Covidien Dover PREC PRM 18FR SILVER FLY TY |  |  | January 22, 2016 | 2016 | Covidien LLC | Aplicare Castile Soap towelettes included in the Foley Trays and Mid stream specimen kits are contaminated with bacteria | Nonconforming Material/Component | Medtronic/Covidien notfiied customersl via Federal Express on January 22, 2016,, The letter instructs customers to not use the Castile Soap Towelettes contained in the product kits listed on Attachment 1 and discard the affected towelettes. The Castile Soap packet included in these kits does not impact the integrity of the remaining products in the kit. Questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878.For questions regarding this recall call 203-492-5000. |  | 3280 | 3280 | US (nationwide) and to Canada. | Terminated | June 20, 2017 | 515 | 909552 | Not_Computer
Z-1230-2016 | 73384 | No | NCD |  |  | Microbiology | PMA | QFT TB Antigen Tube. |  |  | February 02, 2016 | 2016 | Qiagen Sciences LLC | Potential for obtaining a false positive result due to the possibility of endotoxin presence. | Nonconforming Material/Component | QIAGEN sent an "Urgent - Notification of Market Withdrawal" dated February 2, 2016, to all affected customers. The letter identify the product the problem and the action needed to be taken by the customer.Immediately discontinue use of this lot and discard any remaining inventory. Please acknowledge receipt of this notification and actions taken by completing the information on the next page, and return to: TechService-NA@qiagen.com or fax +1-661-775-7479.QIAGEN apologizes for any inconvenience this market withdrawal has caused, and continues to take every effort possible to ensure the highest quality of product is always available to our valued customers around the globe. Please do not hesitate to contact your local QuantiFERON sales representative or customer service representative should you have any questions or concerns.For further questions, please call ( 240) 686-7500. |  | 152900 | The total amount of product distributed 152,900 Tubes | Worldwide Distribution - US Distribution to the state of : Georgia., and to the countries of : Austria, Belgium, Switzerland, Czech Republic, Germany, Denmark, Estonia, Spain, France, Great Britain, Ireland, Korea (South) Luxembourg, Netherlands, Poland, Portugal, Qatar, Russian (Federation), Singapore, Turkey and South Africa. | Open, Classified |  |  | 152900 | Not_Computer
Z-1233-2016 | 73292 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas p 512 pre-analytical system |  |  | February 23, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Due to a false triggering or detection of the lifting gripper READY signal, sample tubes are not correctly placed back in the Rack Tube Transport (RTT) after the decapping process. Therefore, open sample tubes can be dropped in the cobas p 512, spilling the sample material. | Software design | Roche Diagnostics sent an "Urgent Medical Device Correction" letter via UPS Ground (receipt required) on February 23, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To prevent this malfunction, a software patch has been developed and validated. The installation of this software patch is mandatory and will be performed by a Roche Field Engineering Specialist (FES). Your Roche FES will contact you to schedule the installation of the software patch. Until the software patch is implemented, Roche recommends that instrument operators monitor their systems for the occurrence of Error 140 Gripper has lost tube. If this error occurs, all components in contact with a spill must have extra cleaning. Refer to the cobas p 512 pre-analytical instrument Operators Manual Version 1.5.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 for further questions. |  | 35 | 35 Units | US Distribution including Puerto Rico and to the states of :TX, OH, TN, AZ, WA, MI and GA | Open, Classified |  |  | 35 | Computer
Z-1232-2016 | 73481 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo X Workplace is a medical workstation for real-time viewing, image manipulation, 3D-visualization, communication, and storage of medical images and data on exchange media. It is used for diagnostic image viewing and post processing during interventional procedures. |  |  | February 24, 2016 | 2016 | Siemens Medical Solutions USA, Inc | After importing, the segmentation results appear mirrored at the CARTO system and can't be used for the ablation procedure. | Software design | Siemens sent an Important Customer Safety Notice dated February 24, 2016, to all affected customers notifying them that the issue will be remedied with a software update via AX004/16/S. Following the installation of this update, the segmentation results can be used at the CARTO system as described in the operator manual. Customers with questions were instructed to call 800-888-7436. For questions regarding this recall call 610-448-6461. |  | 6 | 6 | Nationwide Distribution to IL, NY, MT, and MN. | Open, Classified |  |  | 6 | Computer
Z-1244-2016 | 72991 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System (Software Version 3.0.0 and after) |  |  | December 15, 2015 | 2015 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Due to Operator Console software anomaly, a change in treatment completion status of the last patient of the day may be altered from "Completed (or Discontinued)" to "Untreated" under specific conditions. | Software design | Due to Operator Console software anomaly, a change in treatment completion status of the last patient of the day may be altered from "Completed (or Discontinued)" to "Untreated" if all the following conditions are met: (1) At least one patient has been treated on the day; (2) A major error occurs when VeroTM/MHI-TM2000 is operated in a mode, other than Clinical Mode, which means Daily Check Mode or Maintenance Mode. (3) The major error is NOT associated with a function that uses therapeutic/ MV radiation parameters, such as an error in Isocenter Check function, etc. Note 1: If the major error is associated with a function that uses therapeutic/MV radiation parameters, the problem that treatment completion status is changed will never occur. The Operator Console changes the treatment completion status in the course of processing the major error. Note 2: Although treatment completion status is changed to "Untreated", its delivered dose result has been correctly saved into R&V; System. It has not been lost. Advice on Actions by Customer/Users: 1. Please make sure to perform the operations as described below in case a major error occurred in a mode other than Clinical Mode. (1) Cancel the error.(2) Click [Clinical] button on the left of Operator Console screen. (3) Verify, on R&V; System screen, that treatment completion status of the patients who were treated on the day are all correct. The treatment completion status is indicated in different colors as follows depending on the status: Untreated: BLUE, Completed: GRAY, Discontinued: RED. Customers requiring support concerning this issue, please contact the Customer Service Representative: 1-800-597-5911 (Mon - Fri) or via email us.support@brainlab.com. If any other problem is detected, please stop operation and contact the Customer Service representative. |  | 25 | 25 devices (4 domestically in U.S.A. and 21 internationally) | Distributed in the states of Florida, New York, Ohio &amp; Texas, and the countries of France, Germany, Japan, Italy, Korea, &amp; Belgium. | Terminated | April 20, 2017 | 492 | 25 | Computer
Z-1246-2016 | 73197 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Irrigation Clip 30; Intended to position and hold irrigation tubing during surgical procedures. |  |  | November 02, 2015 | 2015 | The Anspach Effort, Inc. | Assembled with the incorrect (smaller) front clips (component 44-0495 instead of 44.0498). | Nonconforming Material/Component | On Nov 2, 2015 Anspach sent a letter to their customers informing them of the recall. Customer Immediate Actions: Please complete the attached reply form confirming your receipt of this letter. Return the completed form by fax to the number or email provided on the form. Returning the form promptly will prevent you from receiving repeat notices. If you distribute any of the products to otherservices or facilities, please forward this information as appropriate. We understand this product recall removal) may have an impact on your facility. The FDA has recommended that healthcare facilities follow their risk management procedures as it relates to this notification. We apologize for any inconvenience that this Product Removal may create and appreciate your cooperation with our request. If you have any questions regarding this correction, please contactthe Complaint Handling Unit Manager at 561-494-3673 or contact your Anspach Sales Representative. |  | 50 | 50 clips | Distributed in the states of PA and MN, and the country of Switzerland. | Open, Classified |  |  | 50 | Not_Computer
Z-1245-2016 | 73316 | refractor, manual, non-powered, including phoropter | HKN | Manual refractor. | Ophthalmic | Ophthalmic | 510(K) Exempt | RT-5100 (Epic 5100) Refractor and RT 3100 Refractor: |  |  | February 29, 2016 | 2016 | Nidek Inc | There has been a reoccurrence of an issue on the RT 5100 and RT 3100 Refractor of the Near Point Chart Arm lowering spontaneously with potential for slight injury to face or nose of patient. | Component design/selection | Nidek sent a Recall letters dated February 29, 2016 to all affected site. The letter identified the affected product, problem and actions to be taken. Nidek Customer Service will be replacing the affected product with new parts to address this recall. For questions call Customer Service at 1-800-722-0219. |  | 2703 | 2703 separate serial numbers | US Nationwide Distribution | Open, Classified |  |  | 2703 | Not_Computer
Z-1247-2016 | 73407 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Neurotherm Curved Radiofrequency (RF) cannnula 10 cm |  |  | February 26, 2013 | 2013 | Neurotherm, Inc. | Straight needle labeled as a curved needle | Process control | NeuroTherm notified consignees by telephone between February 2- March 15, 2013 about the affected product, problem and actions to be taken. Customers were instructed to examine inventory and return devices. |  | 135 | 135 cannulas (27 boxes/5) | US Nationwide Distribution in the states of CA ,CO, OH, OR | Terminated | March 25, 2016 | 1123 | 135 | Not_Computer
Z-1239-2016 | 73315 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Alaris PC Unit, Infusion Pump Model 8000, Part No. TC10005092 |  |  | February 12, 2016 | 2016 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit because a component on the PC unit power supply may cause a "System Error" or "Missing Battery" error code (120.4630). | Under Investigation by firm | A customer notification letter dated 2/25/16 will be sent to all their customers who purchased the Alaris PC unit (Model 8000) to inform them that CareFusion is recalling it because a component on the PC unit power supply board may cause a "System Error" or "Missing Battery" error code (120.4630). The letter informs the customers of the potential risk involved and the actions to be taken. Customers with recall related questions are instructed to call CareFusion Support Center at (888) 562-6018, 7am-4pm PST, Mon-Fri or email at supportcenter@carefusion.com. Customers with clinical inquires, product complaints or clinical troubleshooting are instructed to contact customer advocacy at (888) 812-3266, 24 hrs/day, 7 days/wk or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system are instructed to contact technical support at (888) 812-3229, 6am-5pm PST, Mon-Fri or email at DL-US-INF-TechSupport@carefusion.com. |  | 170 | 170 units | Nationwide in US: AR, AZ, CA, FL, GA, IN, KY, MD, MI, MN, MO, NJ, NM, NY, OH, OR, PA, SC, TN, TX, VA, WI. | Terminated | June 23, 2017 | 497 | 170 | Not_Computer
Z-1237-2016 | 73398 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | UltraxxTM Nephrostomy Balloon Set |  |  | January 27, 2016 | 2016 | Cook Inc. | Cook Medical is voluntarily recalling specific lots of the Ultraxx" Nephrostomy Balloon Set and the AscendÂż AQÂż Ureteral Dilation Balloon Catheter Set because of a labeling mix-up. Cook Medical has received reports of packages labeled as Ultraxx" Nephrostomy Balloon Set actually containing theAscendÂż AQÂż Ureteral Dilation Balloon Catheter Set and vice versa. | Labeling mix-ups | Cook Medical sent an Urgent Medical Device Recall Notification letter dated January 27, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to quarantine any unused affected product. Immediately collect and return all unused affected products to Cook Medical If you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235.For questions regarding this recall call 800-457-4500. |  | 20 | 20 | Worldwide Distribution - US including GA, ID, LA, OR, SC, TX and Internationally to CHINA. | Terminated | May 17, 2017 | 476 | 20 | Not_Computer
Z-1236-2016 | 73366 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Ambu, Aura Gain - Single Use Laryngeal Mask- Sterile |  |  | February 10, 2016 | 2016 | Ambu Inc. | Ambu, Aura Gain laryngeal Mask, was found to have lack of Sterility Assurance. The firm discovered the problem during final quality control at manufacturing site. Damage to the sterile barrier (the pouch) may compromise a medical device's sterility,leading to contamination of the product and,ultimately,patient complications. | Process control | Filed Safety Notice was sent via USPS Certified Mail to all customers on Dec 10,2015. |  | 45990 | Total Numbers of unit distributed: 45990 pes (USA) | Please see attached consignee list | Open, Classified |  |  | 45990 | Not_Computer
Z-1234-2016 | 73337 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BioTex Reusable Adapter Kit, Part Number 401-021-1010, packaged non-sterile in a dedicated case for steam sterilization |  |  | February 22, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation is recalling the Biotex Adapter Kit because it was commercially distributed by BioTex without a cleared premarket [510(k)] from FDA. | No Marketing Application | The firm, Medtronic, sent an "URGENT Medical Device Recall" letter dated 2/2016 to its consignees. The letter described the product, problem and actions to be taken. The identified consignees will be visited by the assigned Medtronic field representative to explain the need for the field action. The consignees are instructed to locate and remove/secure the kit from use. The Medtronic field representative will assist the consignee with identifying and removing the kit, and complete a Customer Visit record for each consignee they are assigned/visit to document the results of their visit. The consignee is instructed to sign the customer visit form presented by the representative. A copy of this record and a letter explaining the FCA will be left with the consignee.Please contact Medtronic Technical Services at 1-800-595-9709 if you have any questions regarding the subject action or the content of the letter. |  | 4 | 4 | US Distribution to states of: FL, NJ, AZ, and NC. | Open, Classified |  |  | 4 | Not_Computer
Z-1235-2016 | 73343 | No | LQE |  |  | Ophthalmic | PMA | KAMRA Inlay, ACI 7000 Corneal Inlay, Part No. 76195-01 and 76195-10 |  |  | February 11, 2016 | 2016 | AcuFocus, Inc. | AcuFocus is recalling KAMRA ACI 7000 Corneal Inlays due to concerns with the labeled shelf life. | Under Investigation by firm | The firm, AcuFocus, sent an "URGENT FIELD SAFETY NOTICE" letter dated 2/11/16 to all their customers to inform them that AcuFocus is recalling 2 lots of KAMRA ACI 7000 Corneal Inlays due to concerns with the labeled shelf life. The described the product, problem and the actions to be taken. Customers are instructed to report any abnormal appearance of the implanted KAMRA inlay to AcuFocus, inc; stop using and remove from your inventory all affected lots; complete and return the attached Customer Recall Response Form via email: customerservice@acufocus.com, Fax: +1 (949) 585-9545 Attn: Recall Coordinator, or mail to: AcuFocus Inc., Attn: Recall Coordinator, 32 Discovery, Suite 200, Irvine CA 92618, within 3 business days even if you have no inventory affected by this recall; and provide the attached Customer Field Safety Notice and Customer Recall Response Form to your customers who have or may have received product from the affected two lots. Customers with product complaint events are instructed to report regarding KAMRA inlays involved in this recall event, and to inform AcuFocus, Inc. by calling +1 (949) 585-9511 ext 608. Customers with any questions related to the return process, are instructed to contact an AcuFocus Customer Service Representative at +1 (949) 585-9511 ext 608. |  | 182 | 182 units | International distribution in countries of: Canada, Germany, Saudi Arabia, Turkey, and United Kingdom. | Open, Classified |  |  | 182 | Not_Computer
Z-1243-2016 | 73052 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000 EstradiolTest Code E2 |  |  | January 13, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | The drug fulvestrant (FaslodexÂż) may cause falsely elevated estradiol results. | Device Design | Siemens issued an Urgent Medical Device Correction (UMDC) to affected Siemens Healthcare Diagnostics customers in the United States on January 13, 2016 to be delivered to customers on January 14, 2016 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on January 13, 2016. The letters identified affected product, problem and the actions to be taken. Customers may continue to use the estradiol assays to report patient results however, the assays should not be used for patients being treated with fulvestrant. Customers where instructed to complete and return the Correction Effectiveness Check attached. For questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 9334 | 9334 units | Worldwide Distribution  US Nationwide , AR, AU, AU, BD, BO, BR, CA, CH, CL, CN, CO, DO, EC, EG, HK, ID, IE, IL, IN, JP, KR, LK, LY, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, SV, TH, TW, UY, VE, VN, and ZA. | Open, Classified |  |  | 54134 | Not_Computer
Z-1266-2016 | 73157 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | The S.C.O.R.E.S. Unit (self contained operating room equipment |  |  | January 13, 2016 | 2016 | PMBS, LLC | Water retention issues. Units with filters that remain wet after the sterilization may compromised sterility. | Device Design | On January 13, 2016 PMBS, LLC sent letters to their customers informing them of the correction. Per product's instructions for use, the SCORES cabinet is to be run for a minimun of a 4 minute sterilization cycle, followed by a 30 munites dry time followed by a 30 minutes door crack time. not all facilities generate completely dry SCORES filters following these instructions. It is imperative that the filters be checked in the Operation Room prior to useing the surgical instruments sterilised within the cabinet. If the filters are wet, do not use the contents. In the letter is a IFU addendum including instructions for running the equiptment with extended cycles. For more information please contact Ms. Michele Mauzerall, CEO at 908-300-7093. |  | 58 | 58 units | FL, GA, MN, NY, OH, OR SC, and TN | Open, Classified |  |  | 58 | Not_Computer
Z-1248-2016 | 72890 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Cardiosave Hybrid and rescue IABP |  |  | December 16, 2015 | 2015 | Maquet Datascope Corp - Cardiac Assist Division | Maquet has recieved information that in some Cardiosave IABPS, the scroll compressor did not meet the specifications for output pressure or vacuum at specific flow rates. | Other | Maquet sent an "Urgent - Medical Device Field Correction Notice and Response form dated December 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers are informed of the issues and provided corrective actions. Customers are asked to completed the attached response form in order to return to 1-973-629-1518, or via email to cardiosavescroll@maquet.com. Questions can be directed 1-800-777-4222, option 3, Monday through Friday 8 am-6 pm. EST. |  | 2350 | 1243 units US; 1107 units OUS | Worldwide Distribution- Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 2350 | Not_Computer
Z-1265-2016 | 73400 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard XP Knee System Porous Femoral Component |  |  | February 05, 2016 | 2016 | Biomet, Inc. | The anterior/posterior (AP) dimension may be oversized by 0.016 . If the femoral component does not fit as tightly as intended, bony ingrowth may not occur causing loosening of the femoral component. A revision surgery may be necessary. | Device Design | On 2/5/2016, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors, hospital risk managers, and implanting surgeons via courier. Distributors were instructed to locate and return the affected product, hospitals to locate and discontinue use, and implanting surgeons to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 890 | 890 | UTNetherlands | Terminated | March 28, 2017 | 417 | 890 | Not_Computer
Z-1249-2016 | 73494 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Reveal LINQ insertable Cardiac Monitor, Model LNQ11. |  |  | February 23, 2016 | 2016 | Medtronic Inc. | Medtronic has identified an issue with the sensitivity of an algorithm used in the Reveal LINQ ICM that may prematurely trigger the Recommended Replacement Time (RRT) alert in some devices. Battery capacity is not affected and the device will continue to support data collection and manual data transmissions. As part of the normal behavior of the device, 30 days after RRT status is reached, Reveal | Software design | Consignees in US were hand delivered a Medtronic "Medical Device Correction" letter dated February 2016. The letter was addressed to Physician or Healthcare Professional. The letter described the Issue &amp; Patient Management Guidance. For questions contact your local Medtronic Representative or Medtronic Diagnostic Technical Services at 800-929-4043.Outside the U.S., the letter will be mailed or hand delivered to impacted consignees. |  | 189009 | 189,009 | Worldwide Distribution-US (nationwide) including DC and the countries of Andorra, Australia, Austria, Belgium, Botswana, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic,Denmark, Faroe Islands , Finland, France, Germany, Greece, Hong Kong,Hungary, Iceland, Indonesia, Iran, Ireland, Israel, Italy, Kazakhstan, Kenya,Kuwait, Latvia, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Namibia,Netherlands, New Caledonia, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore,Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tanzania,Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 189009 | Not_Computer
Z-1253-2016 | 73402 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | 10mm X 12o Solitaire Ti Medium Spacer  Solitaire Anterior Spinal System, Model Number 1400-1230. |  |  | February 12, 2016 | 2016 | Biomet Spine, LLC | Zimmer Biomet Spine initiated a recall of certain Solitaire Ti Spacers because they may have a mislabeled lordotic angle of 12degrees when the angle is actually 6 degrees. | Mixed-up of materials/components | Zimmer Biomet sent an Urgent Medical Device Recall Notice letter dated February 26, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the enclosed Response Form. For questions call 303-465-8929. |  | 10 | 10 | Worldwide Distribution - US Nationwide - one foreign consignee (not Canada). | Open, Classified |  |  | 10 | Not_Computer
Z-1264-2016 | 73543 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clear Vue Prone Breast Treatment Table and Access ClearVision Prone Breast Device |  |  | January 07, 2016 | 2016 | Wfr/aquaplast Corp | Some of the Access ClearVue Prone Breast devices may contain scales that are misaligned from the zero reference with respect to other Access ClearVue Prone Breast devices. This can result in differences in scale orientation. | Process control | Qfix sent a Field Safety Notification letter dated January 7, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached Inventory Survey form and mail it back. For questions call (610) 268-0585. |  | 216 | 18 units of RT-4545CV-17 and 12 units of RT-4545CV-20 | Worldwide Distribution - US Nationwide in the states of NY, PA, CA, IL, and TX. Foreign Countries: Singapore, Australia, Saudi Arabia, and Spain. | Terminated | January 06, 2017 | 365 | 216 | Not_Computer
Z-1250-2016 | 73533 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 |  |  | March 14, 2016 | 2016 | Terumo Cardiovascular Systems Corporation | Specific CDIÂż Blood Parameter Monitoring System 500 devices are being voluntarily recalled because the BPM Sensor Head Assemblys Thermistor, which provides the blood temperature value that results in accurate display values on the monitor, does not meet specification. This may cause inaccurate temperature measurement and inaccurate analyte display values on the CDI System 500 monitor.Terumo Cardiovascular Systems (Terumo CVS) received complaints of inaccurate temperature measurements for specific devices distributed since November 5, 2015. These include both new production devices and recently serviced devices.POTENTIAL HAZARDThere have been no reported illnesses or injuries as a result of this issue.A user who is not aware that the CDI System 500 is displaying inaccurate temperature values may not manage patient temperature appropriately. Using inaccurate information to manage warming and cooling strategies for a procedure could result in prolonged time on bypass. It could also lead to unnecessarily aggressive temperature management, resulting in excessive hyper or hypothermia, with potential neurologic and organ dysfunction, or increased blood component damage.Inaccurate temperature measurement could also cause inaccurate measurements of other BPM values including potassium (K+), pO2, pCO2, and pH. The greater the temperature inaccuracy, the greater the degree of inaccuracy of these other BPM values due to the dependence of their algorithms on temperature for the calculations. Inaccurate measurement of these values could result in inappropriate patient management strategies being employed to address them with the potential to result in moderate patient injury. | Nonconforming Material/Component | Terumo sent an " URGENT MEDICAL DEVICE RECALL" notification dated March 14, 2016, to all affected customers via mail. The letter identified the product the problem and the action needed to be taken by the customer. The Recall Notification has instructions for responding to the formal recall notification. If you have questions, contact Terumo CVS Customer Service: 800.521.2818 Customer Service Hours: Monday  Friday 8 a.m.  6 p.m. ET |  | 222 | 222 | Worldwide Distribution - Nationwide Distribution and to the countries of : Australia, Belgium, Canada, Japan , Singapore and United Arab Emirates. | Terminated | March 28, 2017 | 379 | 222 | Not_Computer
Z-1251-2016 | 73324 | No | MHW | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | IMCO brand Conforming Stretch Gauze Bandages, 4 x 4.1 Yds/10 cm x 3,7 cm, Reorder No. 704-IMC, Bar code 6 86864 04351 5, 96 per case (8 boxes x 12 sterile bandages), Single Use |  |  | February 24, 2016 | 2016 | Amd Ritmed | Product does not meet sterility requirements based on FDA analysis and tested positive for Bacterium Bacillus. | Process control | AMD Ritmed sent an Urgent - Medical Device Recall letter dated February 25, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to examine all stocks immediately to determine if any of the affected product is on hand. If so, please discontinue distributing the product immediately and promptly return.If customers further distributed any of the product, please immediately contact your accounts advise them of the recall situation, and have them return their outstanding recalled stocks to you. Return these stocks as indicated above. Please return the enclosed response form immediately providing the requested information by fax 1-866-871-3455 or by email: AMDCustomerservice@amdritmed.com. Customers with questions can call 1-800-445-0340 or email AMD Customer Service: AMDCustomerservice@amdritmed.com. |  | 126 | 126 cases total (87 cases domestically and 39 cases internationally) | US Distribution including Puerto Rico and to the states of : DE, FL, KY, MI, NE and NJ., and Internationally to the country of Canada. | Open, Classified |  |  | 126 | Not_Computer
Z-1262-2016 | 73504 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The E.CAM gamma camera system |  |  | February 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | We received a customer complaint that Siemens reported via the MedWatch process (1423253-2015-00002). The result of a Hazard Analysis and CAPA investigation, we have determined that for e.cam or Symbia E systems that although the existing user manual is correct, additional supplemental instructions, with the inclusion of visual aids, may further clarify and emphasize the proper collimator change | Labeling design | Siemens sent a Customer Advisory Notice CAN 001 -2016 to all affected customers describing the recall and additional actions to take.Siemens updated user instructions, drafting customer advisory letter, and obtaining internal and external approvals, and translation of this content. Customer Advisory Notice Letter provided to each impacted consignee will document receipt of the letter and additional attempts will be made until such receipt is acknowledged. Instructions provided in letter explain that Hospital staff should inform end users of the enhancements that have been made to the user manual and continue to use your system. It also instruct end user that the advisory notice and addendum instructions should be placed with the system's instructions for use. The system use information should be available to all operators of the system.Siemens advises that if the end user has sold their equipment and it is no longer in their possession, They kindly ask that the end user forward the notice to the new owner of this equipment. They should inform Siemens about the new owner of the equipment.Adverse events or quality problems experienced with the use of this product should be reported to Siemens through the contact information provided below and may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. If the end user has any questions regarding the important notice, they should contact their local Service representative at the contact numbers provided below.-America: 1-800-888-7436-Europe, Middle East, and Africa: +49 9131 940 4000-Asia and Australia: +86 (21) 3811 2121 |  | 3090 | 3090 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico.***** Foreign Consignees: Albania, Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Botswana, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Ghana, Greece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kosovo, Kuwait, Libya, Malaysia, Mauritius, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Oman, P.R. China, Pakistan, Palestinian Ter, Peru, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad,Tobago, Tunisia, Turkey, U.A.E., Ukraine, United Kingdom, Uruguay and Vietnam. | Open, Classified |  |  | 3816 | Not_Computer
Z-1252-2016 | 73386 | No | KYF | Aqueous shunt. | Ophthalmic | Ophthalmic | 510(k) | Molteno3 Glaucoma Implants G-Series: |  |  | February 22, 2016 | 2016 | Innovative Ophthalmic Products, Inc | Innovative Ophthalmic Products (IOP) is recalling the Molteno 3 G-Series Glaucoma Implant the package insert/Instructions for Use (IFU) for the device has been updated to add a caution statement to minimize off-label use. | Labeling False and Misleading | Innovative Ophthalmic Products, Inc ( IOP ) sent an Customer Notification letter dated February 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken. The letter informs the customers that IOP is recalling because the package insert/Instructions for Use (IFU) for the device has been updated to add a caution statement. Specifically, the new IFU caution statement is "Caution should be taken with supra-Tenon implantation of this device as erosion of the conjunctiva is known to result from the raised edge of the oval pressure ridge."For further questions please call (202) 857-6464 |  | 278 | 278 units | Nationwide Distribution | Open, Classified |  |  | 278 | Not_Computer
Z-1268-2016 | 73460 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | Helios HL3T and HLT Track Mounted Dental Lights |  |  | March 08, 2016 | 2016 | Pelton &amp; Crane / Marus / DCI Equipment / KaVo | The firm received two customer complaints local Pelton &amp; Crane distributors that the Helios 3000 track mounted dental lights were missing the two 8/32x3/16 set screws during the initial installation of the dental light. | Nonconforming Material/Component | Pelton &amp; Crane sent an Urgent Medical Device Recall letters to affected customers on March 8, 2016 and will be completed by 3/11/2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Acknowledgement Form in the enclosed self-address stamped envelope to Pelton &amp; Crane. For questions or concerns regarding this recall please call 1-800-659-6560, ext. 7227. |  | 156 | 156 units | US Nationwide Distribution in the states of AZ, AR, CA, CO, CT, FL, IA, ID, IL, IN, MA, MI, MN, NJ, NY, NC, OH, OR, PA, TN, TX, UT, VA, WI and Canada | Open, Classified |  |  | 156 | Not_Computer
Z-1269-2016 | 73417 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Information Manager System, Software Versions 1.0 to 4.41 |  |  | February 11, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the MicroScan LabPro Information Manager System because the software incorrectly allows the operator to manually edit the carbohydrate substrates when manually reading dried overnight gram negative panels with an ID Hold status. | Unknown/Undetermined by firm | An Urgent Medical Device Recall Letter dated 2/11/2016 was sent to all customers who purchased the MicroScan LabPro Information Manager System, Version 1.0 to Version 4.41. Beckman Coulter is expanding the recall that was initiated in July 2015 regarding FA-25776 for the MicroScan LabPro Information Manager System. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Representative: http:www.beckmancoulter.com, (800) 677-7226 in US and Canada. Customers outside of US and Canada, are instructed to contact their local Beckman Coulter representative. |  | 2702 | 2,702 units total (1,032 units in US) | Affected geographies: Bahamas, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, French Guiana, Germany, Ghana, Greece, Guadeloupe, Guatemala, India, Israel, Italy, Japan, Kazakhstan, Lithuania, Mexico, Morocco, Netherlands, Northern Mariana Islands, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Russia, Singapore, South Africa, Spain, Taiwan, Tunisia, Turkey, United States, Uruguay, Virgin Islands, Zimbabwe. | Open, Classified |  |  | 2702 | Computer
Z-1267-2016 | 73448 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | The Spectra-System Dental Implant 2008 system is comprised of dental implant fixtures and prosthetic devices that compose a two-piece implant system. |  |  | March 04, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | Implant Direct Sybron Manufacturing, LLC is recalling GoDirect Implants because some consignees did not receive the correct tool to drive the implant to bone level. | Nonconforming Material/Component | Implant Direct sent an Urgent: Medical Device Field Corrective Action letter dated March 4, 2016, to all affected customers. The letter states that Implant Direct will correct the issue and send the proper tool. The proper tool is included with the notification. The letter requests that customers complete and return the Acknowledgement and Recall Return Form within 48 hours.It customers are an authorized Implant Direct Sybron Manufacturing distributor it is requested that customers are identified and contact them to inform them of the issue within 48 hours of receiving the notification. Distributors should also provide customers with the correct tool.If you have any of the affected product listed above and have any questions contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425.For further questions regarding this please call (818) 444-3300 Ext. 3323 |  | 2033 | 2033 | Worldwide Distribution - Nationwide Distribution to the states of : FL, AZ, IL, MS, TX, CA, ID, AL, MO, MN, NY, CO, GA, ME, NJ, NC, NV, LA, KS, MN, KY, PA, UT, OH, MI, SC, MD, WA, WI., and to the countries of : Canada, Honduras and Dominican Republic. | Open, Classified |  |  | 2033 | Not_Computer
Z-1271-2016 | 73426 | No | LKK |  |  | General Hospital | PMA | Medtronic, Model 8781 Ascenda" Intrathecal Catheter kit, Rx Only, Sterile Ethylene Oxide. |  |  | February 09, 2016 | 2016 | Medtronic Neuromodulation | Medtronic Neuromodulation is initiating a voluntary removal of a single Model 8781 Ascenda" Intrathecal Catheter kit. We have identified that this kit may have been produced with incorrect outer package labels. As a result, it may not contain all components necessary to complete the implant procedure. | Under Investigation by firm | Medtronic sent an "Medical Device Removal" letter dated February 2016. The letter described the problem and the product involved in the recall. The Action Required was to segregate the unused affected product. Their Medtronic Representative will be contacting consignee. For questions contact your Medtronic representative, or call for further questions. (800) 707-0933 |  | 9 | 9 (1 still unused) | Internationally to Japan. | Terminated | August 11, 2016 | 184 | 9 | Not_Computer
Z-1279-2016 | 69339 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Foot Switches used with Philips MultiDiagnost Eleva with Flat Detector; Product Codes: 708037, 708038 |  |  | September 24, 2014 | 2014 | Philips Medical Systems, Inc. | Loss of key image functionality due to a bent pedal of the Footswitch. | Device Design | Field Safety Noticed was mailed on September 24, 2014 via Certified Return Receipt. This FCC consists of a Mechanical Hardware update of the Footswitch in order to prevent bent pedals and therefore prevents an intermittent or continuous inability of making live fluoro images or exposures.Customers with questions were instructed to contact their local Philips representative. |  | 168 | 168 | Worldwide Distribution - US (nationwide) and Internationally to:CZECH REPUBLICALBANIAALGERIAARGENTINAAUSTRALIAAUSTRIAAZARBAIJANBAHAMASBANGLADESHBELGIUMBRAZILBULGARIACAYMAN ISLANDSCHILECHINACOLUMBIACOSTA RICACROATIACUBACZECHOSLOVAKIADENMARKDOMINICAN REPUBECUADOREGYPTESTONIAFINLANDFRANCEGABONGERMANYGREECEHONG KONGHUNGARYICELANDINDIAINDONESIAIRANIRAQIRELANDISRAELITALYJAPANJORDANKAZAKHSTANKENYAKOREA SOUTHKUWAITLATVIALEBANONLIBYALITHUANIAMACEDONIAMALAYSIAMAURITIUSMEXICOMOLDAVIAMONGOLIAMOROCCONEPALNETHERLANDSNETHERLANDS ANTNEW ZEALANDNORWAYOMANPAKISTANPALESTINEPANAMAPapua New GuineaPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUORTO RICOQATARREUNIONROMANIARUSSIAN FEDERATSAUDI ARABIASERBIASINGAPORESLOVENIASOUTH AFRICASPAINSUDANSWEDENSWITZERLANDSYRIATAIWANTANZANIATHAILANDTUNISIATURKEYUKRAINEUNITED ARAB EMIUnited Arab EmiratesUNITED KINGDOMURUGUAYUSAUZBEKISTANVENEZUELAVIETNAMYEMENGUADELOUPESYRIAN ARAB REPUBLIC | Open, Classified |  |  | 4348 | Not_Computer
Z-1276-2016 | 73411 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Estradiol III; Elecsys and cobas e analyzers; REF 06656021 190; M 6.5 mL, R1 9 mL, R2 9 mL, |  |  | March 07, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Due to the risk of a recently identified cross reactivity, the Elecsys Estradiol assay should not be used when monitoring patient being treated with fulvestrant. Customers should be informed that fulvestrant will increase the apparent concentration of estradiol in women being treated with this drug. If treatment with fulvestrant has been altered or discontinued as a result of falsely elevated e | Device Design | Roche sent an Urgent Medical Device Correction letter dated March 7, 2016 via UPS Ground (receipt required) to the consignee list. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to complete the attached fax form and fax it to 1-866-891-1984. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336. |  | 23630 | 23,630 | US Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 125942 | Not_Computer
Z-1308-2016 | 73436 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VERO MHI-TM2000 Linear Accelerator System --- Common Name: Medical Linear Accelerator |  |  | March 07, 2016 | 2016 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | In rare cases the VERO/MHI-TM2000 Operator Console could set an incorrect Gantry angle for the first beam of an IMRT treatment if certain specific conditions are met. If Gantry angle for the first beam of IMRT is incorrectly set and if nobody is aware of it, an erroneous treatment would be given to the patient. | Software design | A Field Safety Notice dated March 2, 2016, was sent to all affected customers via e-mail on March 7, 2016. The Field Safety Notice informed users of the problem and provided a safe workaround until corrected software is available. The correction will be free of charge and a Customer Service Representative will contact all affected customers to schedule its installation. Customers with questions were advised to contact their Customer Service Representative. |  | 29 | 25 Worldwide, US: 4 units | Nationwide Distribution including New York, Texas, Florida, and Ohio | Terminated | April 20, 2017 | 409 | 29 | Computer
Z-1303-2016 | 73477 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens branded SOMATOM Emotion and SOMATOM Sensation families of CT systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | February 15, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is initiating this recall due to a potential malfunction and hence potential hazard to patients and personnel when using the SOMATOM Emotion or SOMATOM Sensation systems in conjunction with the use of the cardio licenses "Cardio Pro" and "Cardio Retro". | Device Design | A Safety Advisory Notice was sent to customers informing them of a potential malfunction and hence potential hazard to patients and personnel when using the SOMATOM Emotion or SOMATOM Sensation systems. Siemens informed them that they will deliver an Update Instruction for their Service Engineers to modify the installed Firmware. This will correct the rare synchronization errors of Scanner Firmware and Scanner Software to start the application of the x-ray in compliance with the optimum triggering point for best image quality. |  | 124 | 124 units | US (nationwide) including Guam and the states of AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, KY, MA, MD, ME, MI, MN, MO, NC, NJ, NM, NY, OH, PA, SC, TN, TX, VA, and WA. | Open, Classified |  |  | 124 | Computer
Z-1278-2016 | 73605 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Cios Alpha, mobile X-ray system |  |  | February 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software issues on Cios Alpha mobile C-Arm system | Software in the Use Environment | Siemens sent an Customer Safety Advisory Notice on February 26, 2016, to inform them of some issues that may arise on the Cios Alpha unit. The letter identified the product the problem and the action needed to be taken by the customer. Users were instructed to change the date format to DDMMYYYY to prevent issues, not to release radiation as long as the DVD icon is blinking and to disconnect and re-connect the trolley if the system is needed immediately. In the meantime they are preparing a modification of Cios Alpha that will resolve the above-mentioned issues. As soon as the modification is available, a service engineer will get in touch with each customer and implement the modification to the Cios Alpha system free of charge. If the device was sold, customers were instructed to forward the recall notification to the new owner.For further questions please call (610) 448-6478. |  | 89 | 89 systems | Nationwide Distribution to the states of : AL, AR, CA, CO, CT, DE, FL, GA, IA, IL, KY, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, WA and WI | Terminated | March 20, 2017 | 395 | 89 | Computer
Z-1299-2016 | 73548 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | system, x-ray, angiographic, model # 10094135, 10094137, 10094139, 10094141, 10280959 |  |  | February 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc | In Artis zeego systems, angulations in the vicinity of the C-arm collision area can cause the cable inlet to become clamped and can result in mechanical damage. In Artis systems with A100Plus generators, an attempt to resume operation following detection of a fault (such as a short circuit in the X-ray tube) can result in the failure of a module in the high-voltage generator. | Device Design | The firm, Siemens, sent an "Important Customer Safety Notice" dated 2/16/2016 to its customers. The notice informed them of a corrective action that will be performed to prevent a possible hazard to patients, operators, other persons or equipment. The notice described the product, problem and actions to be taken. Siemens plans to contact each customer to arrange a date to add additional protection mechanisms by means of a system software update.If you have any questions, contact Engineer, Regulatory Technical Specialist at 610-448-6461 or email: Meredith.adams@siemens.com. |  | 552 | 552 units | US Nationwide Distribution including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, GU, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and PUERTO RICO. | Terminated | June 21, 2017 | 481 | 552 | Computer
Z-1309-2016 | 73550 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdian ViewRay Radiation Therapy System, ViewRay Treatment Planning and Delivery System (also known as the MRIdianÂż System) is indicated for stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body when radiation treatment is indicated |  |  | March 07, 2016 | 2016 | Viewray Incorporated | When editing the isocenter or the couch position of the plan while in the treatment workflow (in the Points screen) and re-optimizing, the software will not prompt the user to shift the couch to the new isocenter. Asa result there is the potential to deliver dose to the initial isocenter rather than the new location. | Software design | ViewRay sent an Urgent Device Correction letter dated March 7, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact ViewRay Customer Support at support@viewray.com or call the ViewRay Customer Support center at 855-286-8875. |  | 5 | 5 units | Worldwide Distribution - US including CA, FL, MO &amp; WI., and International distribution to the Republic of Korea and The Netherlands. | Open, Classified |  |  | 5 | Computer
Z-1311-2016 | 73589 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/zeego (system, x-ray, angiographic) for single and biplane diagnostic imaging and interventional procedures such as cardiac and neuro angiography, general and rotational angiography, operating room angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | February 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Possible leak in the cooling system, fluids could potentially seep into the equipment cabinet. This leak can occur sporadically in affected systems. | Other | Siemens mailed out an Important Customer Safety Notice to customers informing them of a corrective action that will be performed to prevent a possible leak in the detector cooling system of Axiom Artis and Artis zee systems. The notice stated the affected plastic couplings will be replaced with metal couplings to prevent fluid leaks. Siemen's service organization will contact each customer to arrange a date to perform the corrective action. Customers asked to call 1-800-888-7436 to schedule an appointment for the corrective action to be implemented. |  | 1757 | 1757 systems | Nationwide distribution: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and Puerto Rico | Open, Classified |  |  | 1757 | Not_Computer
Z-1282-2016 | 73606 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis One, Interventional, Fluoroscopic, x-ray system |  |  | March 11, 2016 | 2016 | Siemens Medical Solutions USA, Inc | After importing segmentation results of the left atrium created on the Artis One system, the possibility exists for the results to be mirrored to an electro-anatomical 3D mapping system, CARTO from Biosense Webster Inc. | Software in the Use Environment | Siemens mailed an Important Customer Safety Notice to customers to inform them of the issue and the corrective actions they plan to take to address the problem. The letter stated that Siemens service organization will contact customers to arrange a date for the installation of the software update. Customers were instructed to call Siemens service organization for an appointment at 1-800-888-7436. |  | 1 | 1 angiography system | Nationwide Distribution to Michigan only | Open, Classified |  |  | 1 | Computer
Z-1281-2016 | 73546 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | CIOS ALPHA (Image-intensified fluoroscope x-ray system) |  |  | February 16, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens found a defective part in their production line which could increase risk of an electrical shock for service personnel when working inside the device during system maintenance. The risk of an electrical shock is only present when a service technician ignores several steps listed in the service manual as precautionary actions during device maintenance. | Nonconforming Material/Component | Siemens sent a Customer Safety Advisory Notice dated February 18, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that Siemens is planning to replace the affected parts as a precautionary measure. A Siemens service engineer will contact each customer to schedule implementation of this precautionary measure. |  | 4 | 4 units | US Nationwide Distributed in the states of AR, CA, NE and UT | Open, Classified |  |  | 4 | Not_Computer
Z-1283-2016 | 73672 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis zee/zeego, system, x-ray, angiographic |  |  | February 18, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to a leak in the cooling system, fluids could potentially seep into the equipment cabinet. Loss of coolant can result in a cooling system failure and subsequent system failure. This leak can occur sporadically in affected systems. | Unknown/Undetermined by firm | Siemens sent an Important Customer Safety Notice to customers on February 18, 2016 informing them that their service organization will contact them to arrange a date to perform the corrective action (replacement of affected plastic couplings with metal couplings to prevent fluid leaks). Customers were advised to contact Siemens service organization for an earlier appointment at 1-800-888-7436. |  | 986 | 986 systems in total | US Nationwide Distribution in the states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY including Puerto Rico. | Terminated | August 08, 2017 | 537 | 986 | Not_Computer
Z-1289-2016 | 73380 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Rebound Cartilage |  |  | February 24, 2016 | 2016 | Ossur H / F | Ossur Americas is recalling , Knee Shield / Patella Cup, Knee Undersleeve Gray, NecLoc Extrication Collar, Rebound Cartilage, Rebound Cartilage Custom, Rebound Air Walker, Rebound Diabetic Walker, Back Support Suspenders Clip on Black, Back Support Suspenders Sew On Black, Extreme Custom, Unloader Custom XT, Unloader Custom Lite, Unloader Custom, and Papoose Infant Spine Immobilizers because the p | Labeling False and Misleading | The firm sent out customer notification letters on 02/26/16.The notification letter instructed customers to include the latex caution notice in any existing inventory of affected products. The latex caution notice can be retrieved at the conclusion of the response survey or by going directly tosafetyalerts.ossur.com. Should customers require Âżssur to provide pre-printed notices for existing inventory, contact Âżssur customer service. For the Âżssur Miami JÂż Cervical Collars with lot numbers before MX150727 customers are instructed to remove and discard the assist strap, which contains natural rubber latex, located in the patient pack.Recipients of the notice with should take the following actions:1. Pass the notice to those who need to be aware within the organization or to any organization where the affected devices have been transferred.2. Maintain awareness on this notification for an appropriate period.3. If have further distributed the products, identify customers and notify them at once of the product alert. It is recommended to include a copy of the alert notification letter.4. If any of your customers are currently wearing a product, check for any skin irritations or other reactions consistent with natural rubber latex allergies.5. To verify have received the notification, visit safetyalerts.ossur.com to complete the Response Survey.If any questions, contact Âżssur Customer Service at (800) 233-6263.The notification letter for the Papoose Infant Spine Immobilizer contained the following additional actions/instructions:Examine inventory, quarantine products subject to the recall and contact customer service for a return authorization to return the product to Âżssur.If any of customers are currently wearing a product, it is recommended replacing the unit with one shipped after the transition lot number MX150727. Check for any skin irritations or other reactions consistent with natural rubber latex allergies. |  | 1030 | 1030 | U.S. distribution to the following; AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, RI, TN, TX, UT, VA, VT, WA, WI, WV | Open, Classified |  |  | 27899 | Not_Computer
Z-1312-2016 | 73600 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Interlocking Bolt, 151 mm (General and Plastics Surgery Devices) Guide, Surgical Instrument |  |  | February 16, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall of certain lots of the Interlocking Bolt which are a part of the Synthes 4.5 mm LCP Proximal Tibia Plate System. The Synthes 4.5 mm LCP Proximal Tibia Plate are indicated for the treatment of nonunions, malunions and fractures of the proximal tibia including: simple comminuted, lateral wedge, depression, medial wedge, bicondylar, combi | Device Design | DePuy Synthes sent an Urgent Notice to inform customers that they were initiating a voluntary recall of lot numbers: 1776197, 5706588, 1782340 and 1868678 of the Interlocking Bolt which are part of the Synthes 4.5mm LCP Proximal Tibia Plate System. Customers were asked to review their inventory for the affected lots, call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verifcation Section (page 3 of the letter) identifying that the affected product has been located and return it to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. A copy of the Verification Section of the letter was to be faxed to 866-737-4701 or emailed to: Synthes2856@stericycle.com. |  | 42 | 42 plate systems | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Open, Classified |  |  | 42 | Not_Computer
Z-1313-2016 | 73325 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | NEL GBS Medium 20 ml in 100x15mm plate (Standard Petri Dish) |  |  | December 21, 2015 | 2015 | Northeast Laboratory Services, Inc. | Product marketed without a 510 (k) | No Marketing Application | Northeast Laboratory services issued recall letter dated 2/29/16and advised users to: Immediately examine your inventory for all lots of these products you may have in stock. Please destroy all remaining product in stock. Users requested to fill out the questionnaire immediately and return the questionnaire via email or Fax.Previously Northeast Laboratory Services notified users by letter dated 12/21/15 that the firm has discontinued the products effective January 2016. Alternative products were referenced in the letter for users to transition to replacement product. |  | 9300 | 9300 | Nationwide Distribution. | Terminated | April 07, 2016 | 108 | 42040 | Not_Computer
Z-1315-2016 | 73381 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TeleSentry Battery, 3.7 Li-Ion Polymer Battery, CAUTION Do not expose to high temperature. Do not disassemble. Dispose of properly. Use specified charger only. 60pC/140pF The ScottCare Corporation. |  |  | February 25, 2016 | 2016 | Scottcare Corporation | The firm received a complaint of the battery exploding while recharging. | Device Design | On 2/26/2016 the firm sent "Urgent: TeleSentry Battery Recall" notification letters to their customers. The letter identified the affected device and the reason for the recall. Customers were asked to immediately examine their inventory and quarantine the affected device. If the product had been further distributed, then a copy of the letter should be provided to those customers. Customers are to follow the steps provided for returning the affected device, including contacting ScottCare Service at 1-800-243-9412. In addition, the attached response form should be completed and returned. Questions should be directed to ScottCare Customer Service at 1-800-243-9412. |  | 9067 | 9,067 batteries | Worldwide Distribution -- USA, including the states of CA, CT, DE, FL, IL, KS, KY, MA, MD, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, PA, SC, TN, TX, &amp; WV; and, the country of Canada. | Open, Classified |  |  | 9067 | Not_Computer
Z-1231-2016 | 73364 | No | RCE |  |  |  |  | GE Inspection Technologies, LP, Cabinet X-ray system used for industrial quality control. |  |  | December 16, 2015 | 2015 | GE Inspection Technologies, LP | GE Inspection Technologies, LP ("GE") has discovered that certain of its cabinet x-ray systems may not fully comply with the Food and Drug Administration's (FDA) Cabinet XÂż ray System performance standards, 21 C.F.R. Âż 1020.40(c)(1)(i) | Vendor change control | GE Inspection Technologies Planned Action to bring into Compliance:1. GE Inspection Technologies will use their updated adhesive procedure to install new brackets to the bottom left and right lead strips in the control door. The new procedure and brackets were developed in November 2013.2. A second notification of customers with affected systems will be sent within5 days of approval of this CAP. This notification will describe the repair above and propose that a GE technician visit each site and install the described repair.3. Repairs will be completed within 120 days after approval of this CAP and will be made free of charge. For further questions, please call (717) 447-1278.; |  | 55 | 55 US Distribution | US Distribution to the states of : TX, MD, MI, CA, TN, NH, NY, AL, NC, ID, MA, AZ, IN, FL, PA, NJ, WI, VA, MN, MD CO, GA, SC, OH and DC. | Open, Classified |  |  | 55 | Not_Computer
Z-1274-2016 | 73283 | No | PGX |  | Microbiology | Microbiology | 510(k) | Simplexa Group A Strep Direct (MOL2850) with Direct Amplification Disc (MOL1451, MOL1452, MOL1455) |  |  | February 10, 2016 | 2016 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Direct Amplification Discs (DAD) supplied with the Simplexa HSV 1 &amp; 2 Direct (MOL2150), Simplexa Flu A/B &amp; RSV Direct (MOL2650) and Simplexa Group A Strep Direct kits (MOL2850) because it may result in a false negative result, a false positive result or Error code (s) 500 or 505, and it may result in higher invalid rates. | Under Investigation by firm | A customer notification letter dated 2/10/16 was sent to all their customers who received the Direct Amplification Discs (DAD) supplied with the Simplexa HSV 1 &amp; 2 Direct (MOL2150), Simplexa Flu A/B &amp; RSV Direct (MOL2650) and Simplexa Group A Strep Direct kits (MOL2850). The letter informs the customers that Focus Diagnostics is recalling MOL2150 and MOL2650 kits because it may result in a false negative result, a false positive result or Error Code(s) 500 or 505. Additionally, some customers are reporting insufficient specimen volume errors which may result in higher invalid rates. In those cases, the error results in a Report Note Statement of Insufficient Specimen Volume . The letter informs the customers of the recommendations to be taken for the identified kits and the risks involved for positive results, negative results, and delayed results. The letter informs the customers of the actions to be taken. Customers are instructed to sign the achnowledgement form and email the form to Technicalinfo@focusdx.com or fax back to Focus Diagnostics Technical Services at (562) 240-6526 within 10 business days.On 2/11/16, Focus Diagnostics sent to their customers an amendment to the customer notification letter to clarify what is considered an "early Ct", Focus would consider Internal Control (IC) Ct's less than 25 as "early Ct". Customers are informed that if an internal control Ct is less than 25, additional testing may be necessary.On 2/15/16, Focus Diagnostics sent to their customers an amendment to the customer notification letter to inform them to determine if they had instances of early Cts less than 25, then to please refer to Section III  Analyzing Runs, pages 6-10 and 6-15 in the Focus Diagnostics, Operator Manual for Integrated Cycler Studio 6.0  For Use with IVD Assays in the USA. Also, for a more detailed analysis, customers are instructed to view excerpts from the ICS Operator Manual covering Section III, pages 6-10 and 6-16.On 3/17/16, |  | 41 | 41 units | Worldwide Distribution - US Nationwide in the states of: TX, CA, MA, NY, WA, MI, ME, IN, NJ, WV, OH, PA, HI, NH, CO, MN, FL, SC, IN, IA, TN, OR, AZ, WI, TN, IL. and the countries of: Australia, Brazil, Canada, Dominican Republic, Denmark, Ecuador, France, Germany, Israel, Kuwait, Spain, Sweden, Slovenia, Thailand, and United Kingdom. | Terminated | June 16, 2016 | 127 | 18614 | Not_Computer
Z-1316-2016 | 73281 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | evaluemed Infant Heel Warmer 301-1223; Distributed by Tri-anim |  |  | February 10, 2016 | 2016 | Philips Electronics North America Corporation | eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses bear no expiration date, nor any warning that product used beyond its expected shelf life may warm to temperature above its specification. | Labeling design | Philips sent an Urgent Medical Device Recall leter dated January 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: Actions for Tri-Anim Customers to Take1.) Discontinue use of eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses.2.) Check for any inventory of eValueMed Infant Heel Warmers or eValueMed Infant Transport Mattresses.3.) Destroy all products on hand (e.g., by activating or puncturing each unit). Make a note of the total number of cases of products that were destroyed as this will be needed for step 4 below.4.) Complete and sign the Customer Reply Form provided in this letter or Tri-anims version thereof, to notify Tri-anim of how many eValueMed Infant Heel Warmers and eValueMed Infant Transport Mattresses were destroyed. Submit your signed reply form only after you have completed all steps. Please do not return any product to Philips.Customers with questions should call 336-449-5230 or 262-798-9802.For questions regarding this recall call 978-659-3000. |  | 539.156 | 539.156 | Nationwide distribution to Ohio only. | Terminated | August 01, 2017 | 538 | 539.156 | Not_Computer
Z-1318-2016 | 73489 | No | KRR | Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | All instrumentation associated with the Prelude PF Resurfacing Knee System. |  |  | March 09, 2016 | 2016 | Biomet, Inc. | Inadequate design control. | Device Design | Zimmer Biomet sent an "Urgent Medical Device Recall Notice" dated March 9, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The following actions are REQUIRED:1. Immediately locate and remove the identified device(s) listed below from circulation.2. Carefully follow the instructions on the enclosed Response Form .3.Email a copy of the response form to CPW ARFieldAction@zimmerbiomet.com prior to return of product.4. Use priority carrier for your shipment.Please confirm receipt of this notice by sending back the response form within three (3) business days. Thank you in advance for your assistance and prompt attention. On behalf of Zimmer Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 576 | 576 | US Distribution to the states of : CA, CO, GA, IN, KS, MI, MO, NC, NJ, OH, TN and TX. | Terminated | May 23, 2017 | 440 | 576 | Not_Computer
Z-1319-2016 | 73334 | apparatus, melting point, paraffin | IDT | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Richard-Allan Scientific Paraffin Type 3 |  |  | February 16, 2016 | 2016 | Richard-Allan Scientific Company | Multiple complaints were received from customers indicating that they ordered product 8335(Type 3 Paraffin) but when they opened the box, they found product 8336 (Type 6 Paraffin) inside. | Labeling mix-ups | ThermoFisher sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 15, 2016, to all affected customers. The letter contained the following instructions: If you have the above lot in your possession, please determine whether the bags are labeled as Type 3 (PN 8335) or Type 6 (PN 8336).Âż If the product is labeled as Type 3 (PN 8335) please complete the attached Response form with the quantity in inventory and return the form to Richard-Allan Scientific. The product may be used and no further action is required.Âż If the product is labeled as Type 6 (PN 8336) please complete the attached Response form with the quantity in inventory. The product may be used as is or returned to Richard-Allan Scientific. Please indicate on the Response form if the product will be used or a return is desired. Return the completed form to Richard-Allan Scientific.Âż If affected product has been used previously, please complete the attached Response form with the quantity that was used. No further action is required.By returning the attached Acknowledgment Form will facilitate reporting of this matter to the FDA. Customers with questions were instructed to call 269-544-5628. |  | 29 | 29 cases (10 bags per case) | Worldwide distribution to AZ , CA , FL , GA , IL , MI , MN , NM , NY , OH , TX , UT and Internationally to Germany and the United Kingdom.6 distributed to OUS accounts: Germany, and the United Kingdom | Open, Classified |  |  | 29 | Not_Computer
Z-1328-2016 | 73373 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MicroAire PALÂż Cannula |  |  | February 08, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire has initiated a recall on product PAL-R4011XL (40cm x 4mm Del Vecchio French Rasp cannula), due to our receipt of three (3) complaints that it bent or broke during surgery.Accordingly, the PAL-R4011XL has also been put on production hold until further notice. | Nonconforming Material/Component | The firm is sending letters to all customers who have received PAL-R4011XL fromAugust 2013- October 2015. The firm sent these letters by UPS to expedite the delivery and for traceability. |  | 145 | 145 manufactured and distributed | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 145 | Not_Computer
Z-1325-2016 | 73487 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008 Series Hemodialysis Delivery Systems: 2008T, 2008K2, 2008K and 2008K@home: |  |  | March 10, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Added Warning Statement: Fresenius 2008 Series Hemodialysis Machines False Blood Leak Alarm When Dialyzing Patients Treated with Hydroxocobalamin (or any form of Vitamin B-12) | Device Design | Fresenius Medical Care issued an Urgent Medical Device Alert dated March 10, 2016, to all affected consignees. Consignees were instructed to place notification and Warning Addendum with their 2008Âż Series Hemodialysis Machine Operators Manual. A reply form was included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.Further information or support concerning this issue, customers were instructed to contact Fresenius Technical Services at 800-227-2572. |  | 8384 | 8384 units (8369 US) (15 OUS) | Nationwide Distribution | Open, Classified |  |  | 8384 | Not_Computer
Z-1323-2016 | 73259 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroAire PAL-R4011XL - 40cm x 4 mm Del Vecchio French Rasp aspiration cannula |  |  | February 08, 2016 | 2016 | MicroAire Surgical Instruments, LLC | Complaints that the cannula bent/broke during surgery. | Component design/selection | The firm is sending letters to all customers who have received PAL-R4011XL from August 2013- October 2015. The firm sent these letters by UPS to expedite the delivery and for traceability. |  | 37 | 37 | Worldwide Distribution. US nationwide including Puerto Rico, Australia, Bangladesh, Czech Republic, India, Japan, Lebanon, New Zealand, Norway, Singapore, Switzerland, Taiwan, and Turkey. | Open, Classified |  |  | 37 | Not_Computer
Z-1324-2016 | 73387 | Yes | FJI | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Hemoflow F3 Dialyzer |  |  | February 26, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Dialyzer header leak due to possible improper torque | Process control | The firm, Fresenius Medical Care Renal Therapies Group, LLC issued an "URGENT MEDICAL DEVICE RECALL" letter dated 2/24/16 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of your Hemoflow F3 Dialyzer; place all affected units in a secure, segregated area; contact your FMCRTG, LLC Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product, and fill out and return the attached Reply Form via fax to: 781-699-9796 ATTN: Quality Department or scan and email to: NOTIFYRA@fmc-na.com at your earliest convenience.If you have any additional medical concerns or questions, please contact Medical Information and Communication at 855-616-2309 or Website: www.fresenius-medinfo.com. |  | 1650 | 1,650 units | US Nationwide Distribution including states of: AZ, CA, CT, DC, GA, IL, IN, MI, MN, MO, MS, NY, OH, PA, SC, SD, TX, and WA. | Terminated | June 26, 2017 | 486 | 1650 | Not_Computer
Z-1320-2016 | 73497 | No | OHG | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | Gomco Style Circumcision Clamp Trays is an instrument used to compress the foreskin of the penis during circumcision of a male infant. |  |  | February 29, 2016 | 2016 | Cardinal Health | Two lots of Gomco Style Circumcision Clamp Trays, SSI-0034 and 10-1100 contained a 1.3cm circumcision clamp. They both were labelled as containing a 1.1cm clamp. | Labeling Change Control | Cardinal Health sent an URGENT: PRODUCT RECALL letter dated February 29, 2016, to affected consignees explaining the reason for the product recall and actions that should be taken by the end user. These included the following actions:1. STOP using affected product.2. RETURN the enclosed acknowledgment form via fax to 847-689-9101 or 614-652- 9648, whether or not you have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.3. NOTIFY any customers to whom you may have distributed product affected by this recall.4. CONTACT the appropriate Customer Service group to arrange for return of any affected product you may have:" Hospital800.964.5227" Distributor800.635.6021 The firm asks that consignees who have question or requests of special assistance relating to the recall should contact Cardinal Health Customer Advocacy at 800.292.9332.For questions regarding this recall call 847-473-1500. |  | 440 | SSI-0034: 380 units and 10-1100: 60 units | Nationwide Distribution including AR, KS, WV, , NC, TX, AL, MS and VA. | Open, Classified |  |  | 440 | Not_Computer
Z-1332-2016 | 73456 | No | MCH | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | STD KIT GK260, Standard Set 360 Z |  |  | December 22, 2015 | 2015 | Applied Medical Resources Corp | Applied Medical Resources Corporation is recalling CA500 Epix Universal Clip Applier, kit models STD GK258 AME Standard Kit 44, STD KIT GK213, AME Standard Kit 15, and STD KIT GK260, Standard Set 360 Z, which contain the affected CA500 lot, because of a potential loading mechanism nonconformance. | Nonconforming Material/Component | Customer notification letters were sent on December 22, 2015 for the U.S. and on January 6, 2016 for those in Europe.The customer notification letter requested that the following actions be taken:- Check inventory for recalled product.- Complete the attached Recall Notification Confirmation Form to acknowledge therecall, indicate if facility is returning or has already used this lot, and provide a no-charge PO number if replacements requested.- Return the Recall Notification Confirmation Form to Applied Medical by emailing torecall60684384@appliedmedical.com or faxing to (949) 713-8936.- Return affected product and a copy of the Recall Notification Confirmation Form to Applied Medical For product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by email at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or by email at lcontursi@appliedmedical.com. |  |  |  | U.S .Nationwide Distribution in the states of GA, AL, LA, TX, MN, AK, CA, ID, OR, UT, WY, NM, and AZ.Foreign distirbution to the following; Austria, Belgium, Denmarl, Canada, Japan, France, Germany, Ireland, Italy, Greece, Netherlands, Russia, Spain, Switzerland, Turkey, Great Britain | Terminated | October 03, 2016 | 286 |  | Not_Computer
Z-1322-2016 | 73012 | No | LNR |  |  | Gastroenterology/Urology | PMA | The THERAKOS CELLEX Photopheresis System is a photoimmune therapy. Catalog CLXUSA. |  |  | December 21, 2015 | 2015 | Therakos Inc | The Mallinckrodt (Therakos) post-market monitoring process detected an increased trend in tubing leaks within the pump tubing organizer, associated with a portion of Kits from Lot D 342. | Nonconforming Material/Component | Therakos sent an Urgent Medical Device Recall letter dated December 23, 2015, to all affected customers by UPS Overnight Mail. The letter identified the product, the problem, and the action to be taken by the customer. The consignees were instructed to immediately identify, discontinue use and return affected product. Customers with questions were instructed to call Therakos Customer Service at 1-855-422-9115, option 4.For questions regarding this recall call 908-238-6813. |  | 729 | 729 units | Nationwide Distribution to AL, AZ, CA, CO, DC, FL, ID, IL, IN, MA, ME, MN, MO, NC, NJ, NY, OH, PA, SC, TN, TX, UT, VT, AND WI. | Open, Classified |  |  | 729 | Not_Computer
Z-1326-2016 | 73561 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Titanium Low Profile Neuro Screw, Self Drilling, 3mm, part numbers: 400.833, 400.833E |  |  | February 23, 2016 | 2016 | Synthes (USA) Products LLC | Out of specification at the thread. The thread height of these screws is under-sized. The cross section of the affected area is less than that of conforming screws such that the thread height of the screws is undersized. | Nonconforming Material/Component | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated February 23, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were notified by mail of a potential patient impact. They were instructed to review existing inventories and call DePuy Synthes at 1800-479-6329 to obtain a Return Authorization (RA) Number and replacement order. After completing the verification section on page 4 of the letter, they were advised to return the verification section (page 3) with the product to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. A copy of the completed verification section is to sent to DePuy Synthes by: Fax 877-931-0643 or scan/email Synthes8097@stericycle.com. |  | 535 | 500 units distributed in USA, and 35 units distributed OUS | Worldwide Distribution - US Nationwide in the states of MA, VA, TN, GA, WA, IA, NJ, TN, PA, NY, CO, CT, VT, and the countries of Canada and Switzerland. | Open, Classified |  |  | 685 | Not_Computer
Z-1321-2016 | 73525 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TOPAZ MicroDebrider 0.8 mm REF AC4040-01; TOPAZ MicroDebrider IFS 0.8 mm REF ACH4040-01; TOPAZ MicroDebrider 0.8 mm REF Q6000-01; TOPAZ EPF MicroDebrider 0.8 mm 45 REF Q6002-01 |  |  | March 03, 2016 | 2016 | ArthroCare Corporation | During functional testing of devices after real time aging, it was determined that product failed to meet shelf-life at the expiration date indicated on the product labels. | Package design/selection | Consignees were notified via letter on 3/3/2016. Consignees were instructed to quarantine the recalled products and complete an Inventory Return Certification Form. |  | 26903 | 26,903 devices | Worldwide Distribution-US (nationwide) and the following countries-Austria, Belgium, Bulgaria, Canada, Denmark, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Ireland, Israel, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Spain, Sweden, Switzerland, China, Hong Kong, India, Lebanon, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arabic Emirates, Vietnam, Australia, Japan, New Zealand, Brazil, and Chile. | Terminated | May 17, 2017 | 440 | 26903 | Not_Computer
Z-1341-2016 | 73498 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona The Personalized Knee System Articular Surface Posterior Stabilized (PS) Left Height 18mm Sterile For use in total knee arthroplasty. REF 42-5114-008-18 |  |  | March 17, 2016 | 2016 | Zimmer Manufacturing B.V. | The firm received a customer complaint establishing that the product (Persona PS Articular Surfaces)information shown on the box differed by sides. One side showed P/N: 42-5114-008-18/ Lot: 62632101 and the other side showed P/N: 42-5114-005-14/ Lot: 62646580. | Employee error | Distributors, Sales Representatives, Operation Managers, and Hospital Risk Managers will be provided with an "Urgent Medical Device Recall-Lot Specific" letter. |  |  | TBD | Distributed in US (AL, AZ, CA, GA, HI, ID, KS, LA, MA, MN, MO, NY, OH, OK, PA, TX, UT, VA, and WI) and foreign countries (Korea, Japan, Germany, Singapore, Austria, Italy, South Africa, and Spain). | Terminated | March 14, 2017 | 362 |  | Not_Computer
Z-1343-2016 | 72957 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | NAMIC CONVENIENCE KIT SELF REGIONAL HEALTHCARE PV LEG PACK, REF Catalog No. 651919311, UPN Product No. H7496519193111 |  |  | November 03, 2015 | 2015 | AngioDynamics Inc. | Procedure Products initiated the recall due to the inner lumen of the needles not meeting the intended specification. Specifically, at least a portion of the affected lot is unable to accept a 0.038" guidewire as intended. In some cases the inner lumen was found to be 0.036". | Process control | On November 03, 2015, Navilyst Medical distributed Urgent Voluntary Medical Device Recall notices to their customers via courier service. Customers were advised to immediately remove the recall product from inventory and to segregate the products in a secure location for return to AngioDynamics, Inc. Customers should immediately forward a copy of the recall notification letter to all sites if these products were further distributed. If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday - Friday EST) to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return). The Reply Verification Tracking form can be submitted via fax to: Attn: Convenience Kits - Procedure Products Needles 1-800-782-1357 or via email to rdenino@angiodynamics.com or sbaxter@angiodynamics.com. |  | 136 | 136 kits | US Nationwide in the states of CA, IL, MA, SC, TX | Terminated | March 16, 2017 | 499 | 289 | Not_Computer
Z-1348-2016 | 73429 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Unique Device Identifier No. 10758750002740 &amp; 10758750007110. |  |  | February 16, 2016 | 2016 | Ortho-Clinical Diagnostics | The two sets of wires (4 total wires) exiting from the EMI filters were switched, creating a situation where a module that is intended to be de-energized will remain energized only if the associated AC power cord is unplugged. | Process control | On February 16, 2016, Ortho Clinical Diagnostics distributed Urgent Product Correction Notification notices to their customers via FedEx overnight courier, e-communications and via USPS. Customers are advised that there is no impact to the results generated with an affected system. As a precaution, your Ortho-trained service representative will contact you to schedule an inspection of the wiring within the system at your facility. Following the inspection, reconfiguration of the wiring on your system will be performed if appropriate. Please store a copy of the notification with your user documentation until the wires within your system(s) are inspected. In addition, complete the receipt form to indicate that you have been informed of this anomaly. Please return the form to Joe Falvo via email at confirmationadmin@its.jnj.com or via fax to 1-888-557-3759 or 1-585-453-4110. Customers with questions can contact the Technical Solutions Center at 1-800-421-3311 anytime. |  | 2101 | 2,101 Total units (1,043 domestically &amp; 1,058 internationally) | Worldwide Distribution-US (nationwide) including Puerto Ricoand the countries of Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy &amp; Spain. | Open, Classified |  |  | 2101 | Not_Computer
Z-1275-2016 | 73553 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Guardian II Hemostasis Valve, models 8210 and 8211. The following models are only available Outside the US: FH101, FH101-T, FH101-25, and FH101-50. |  |  | March 04, 2016 | 2016 | Vascular Solutions, Inc. | Vascular Solutions became aware of a potential problem with the click version of the Guardian II hemostasis valves. The low pressure seal may not close properly, which may allow air to be introduced into the device and may lead to risk of an air embolism. | Employee error | Consignees were sent on 3/4/2016 via first class mail a Vascular Solutions "Urgent Medical Device Recall" letter dated March 3, 2016. The letter described the problem and the product involved in the recall. It included the "Immediate Action Required of Customer / User" which instructed the customer to identify the product, remove it from inventory, place it in a secure area, and complete and return the VSI Account Inventory Form to the Customer Service Department. |  | 5283 | 5283 | Nationwide Distribution in the states of AL, AR, CA, CO, DE, FL, GA, ID, IL IN, IA,KY, LA, MD, MA, MI, MO, NE,NV, NH, NY, NC, OH, PA, SD, TN, TX, UT, VA, WA, and WV.OUS: Ireland | Terminated | October 19, 2016 | 229 | 5283 | Computer
Z-1349-2016 | 73475 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | A Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. |  |  | January 18, 2016 | 2016 | Siemens Medical Solutions USA, Inc | To inform users about the possible incorrect values for Distance Measurements when using certain modalities in combination with syngo Imaging. | Software design | Siemens mailed a Customer Advisory Notice to their customers instructing them to check their modalities for the usage of the mentioned DICOM attributes with different values. They recommended upgrading the workplace(s) to syngo Imaging Studio version VB36D if they have such modalities and it cannot be ensured by the modality provider that only one attribute is filled or both attributes are filled with the same value. Customers were asked to contact their local service organization to clarify the compatibility of their syngo Imaging server version with workplace version VB36D. |  | 49 | 49 imaging systems | US (nationwide) including the states of AL, CA, CO, CT, FL, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OK, OR, PA, TN, TX, and WI. | Terminated | August 03, 2017 | 563 | 49 | Not_Computer
Z-1270-2016 | 73468 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Medical Fixed eye Safety filters Model Integre S LP5532. |  |  | February 01, 2016 | 2016 | Laserex Systems Inc. | The Firm discovered that there is a design error which is traceable back to 2012 when a design change on the safety filter optic adaptor was implemented. | Device Design | Ellex Medical Pty Ltd. Planned Action to bring into Compliance:1.Field Corrective Action replacing all affected safety filters of all units manufactured between July 2012 and February 2016, 2.All repairs will be made at no cost to the purchaser; and3.The corrections will be completed by March 30, 2017.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. For further questions please call (800) 824-7444. |  | 74 | 72 units plus 2 spare units | Worldwide Distribution - US Distribution | Open, Classified |  |  | 74 | Not_Computer
Z-1347-2016 | 73555 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA Pioneer MR Systems |  |  | March 01, 2016 | 2016 | GE Medical Systems, LLC | Ferrous material in the SIGNA Pioneer Table Control Box (TCB) could cause unexpected attraction to the MR system's high strengthmagnetic field during servicing activities. This could create a potential risk of serious injury to service personnel who attempt to inappropriately remove the TCB fromthe magnetic field. There have been no events reported as a result of this issue. Clinicalscanni | Labeling Change Control | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 68002 dated March 1, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions or concerns regarding this notification,contact GE Healthcare Service at 18004371171 or your local Service Representative. |  | 16 | 16 (2 US; 14 OUS) | Worldwide Distribution - US Nationwide in the states of SC, NC and the countries of Australia, Canada, France. Germany, Japan and Turkey. | Open, Classified |  |  | 16 | Not_Computer
Z-1357-2016 | 72922 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BARDEX I.C. Anti-Infective 2-Way 3cc Foley Catheter, Rx only. Intended for urinary bladder drainage in female patients requiring catheterization for management of incontinence, voiding dysfunction and surgical procedure. |  |  | December 18, 2015 | 2015 | C.R. Bard, Inc. | Labeling error: Catalog contains Instructions for Use (IFU) error; incorrect inflation volume. | Error in labeling | C. R. Bard, Inc. sent a "NOTIFICATION OF MISLABELING" letter, dated 12/18/2015 to all affected consignees. The letter identifies the product, problem, and actions to be taken by the customers. |  | 1428000 | 1,428,000 units (total) | Distributed US (nationwide) and the countries of Canada and Australia. | Open, Classified |  |  | 1428000 | Not_Computer
Z-1360-2016 | 73579 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Arthrex Deltoid Ligament Reconstruction Implant System. |  |  | January 11, 2016 | 2016 | Arthrex, Inc. | Potential for a component contained in the Implant System to be non-sterile. | Other | On 01/22/2016, all distributors and end user facilities were sent an Urgent Medical Device Voluntary Recall letter by Arthrex, Inc. The firm is voluntarily recalling and replacing the referenced device. Immediately discontinue distributing and use of the indicated product and return to Arthrex. Additionally the Recall Acknowledgement of Receipt card should be completed and returned. Contact Customer Service at 800-934-4404 for replacement and return information. |  | 117 | 117 devices | Distributed to the states of AR, AZ, CA, CO, FL, IL, IN, KS, KY, LA, MD, MO, NC, NJ, NV, NY, OH, PA, SD, TN, TX, and WI. | Open, Classified |  |  | 117 | Not_Computer
Z-1351-2016 | 73338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64:Computed Tomography X-ray Systems, Intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes |  |  | February 05, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm was notified of a software error in which the system may not automatically send all image/data series to remote devices. | Software design | Philips Medical Systems sent a Customer Information dated February 5, 2016, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. To correct the issue Philips will release field change orders (FCO) 72800652 and 72800644 to install a software update on the affected systems.Customers with questions were instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer CareSolutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts).For questions regarding this recall call 440-483-2997. |  | 285 | 285 units | Worldwide Distribution - US to the District of Columbia: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA &amp; WV.Product was also shipped to the following countries: Afghanistan, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Cuba, Denmark, Egypt, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kyrgyzstan, Lebanon, Lithuania, Macedonia, Malawi, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestinian, Philippines, Poland, Republic of Korea, RÂżunion, Romania, Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 1002 | Computer
Z-1350-2016 | 73639 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS, a family of CT systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission dta from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | March 22, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens became aware that the front cover on some Siemens CT system was not properly welded. It is possible for the welded-in threaded bolt to break out of the ground plate of the front cover if the cover is opened during service. This could not happen during clinical operation. Therefore, there is no potential risk for the user or patient and the CT System has no risk for a malfunction during clinical operation | Other | Siemens sent an "Informational Only" notice to customers on 03/22/2016 informing them to be aware that the front cover on some Siemens CT systems was not properly welded. The letter was only to notify customers that the issue is being addressed with the completion of a service visit. In addition, Siemen plans to replace each cover and hand delivered each notification to customers at the time the service is performed. |  | 1 | 1 CT system | Distributed to: Cincinnati, OH. | Terminated | August 07, 2017 | 503 | 1 | Not_Computer
Z-1361-2016 | 73410 | No | GXD | Radiofrequency lesion generator. | Neurology | Neurology | 510(k) | Neurotherm NT2000iX RF Generator |  |  | October 09, 2013 | 2013 | Neurotherm, Inc. | NT2000iX software shipped with the international setting turned on. The international settings include access to Corodotomy, Bi-Polar, and No Temperature modes were not cleared in the United States for use. | Software Manufacturing/Software Deployment | Neurotherm notified consignees via telephone and guided through the process to update the software manually. Thecalls were initiated by Neurotherm representatives on October 09,2013. |  | 25 | 25 | CA, CO, IA, ME, NJ, NY, SC, TX, WI | Terminated | April 14, 2016 | 918 | 25 | Computer
Z-1362-2016 | 73171 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | USE (Ebony) PTA .014" rapid exchange (RX) Peripheral Dilatation Catheter. Distributed by US Endovascular, St. Louis, MO. |  |  | February 11, 2016 | 2016 | Us Endovascular | Product may kink during use rendering the product unusable. | Device Design | NATEC Medical ltd.sent an Urgent Medical Device Advisory Notice dated March 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Warnings:This device should be used only by physicians who are experienced and have a thorough understanding of the clinical and technical aspects of angioplasty.What action do you need to take:1. Always use a guiding catheter when contralateral access is needed.2. Always insert cautiously the PTA catheter inside the guiding catheter, and if any resistance is felt, the cause should be determined via fluoroscopy.3. For contralateral procedure, the USE014OTW has to be preferred to Ebony 014 Rx.NATEC Medical Ltd and our distributor US Endovascular are committed to keeping you and your healthcare professional informed of issues and solutions concerning our products and services. Should you have any questions regarding this or any other matter regarding your USE 014Rx PTA catheter, please do not hesitate to contact our distributor: US Endovascular at (314) 341-4930. |  | 843 | 843 catheters | US Distribution to the states of : AR, CA, FL, LA, MO, MI, MS, NY, TN, and TX. | Open, Classified |  |  | 843 | Not_Computer
Z-1363-2016 | 72894 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Medix NatalCare ST-LX Intensive Care Incubator. Neonatal incubator. |  |  | January 08, 2016 | 2016 | Natus Medical Incorporated | NatalCare LX Incubators were manufactured using a defective oval port hinge assembly with the result that the screws holding the port in place can become loose and possibly result in the port detaching from the incubator. | Process control | Natus of Seattle, WA began sending the Urgent Field Safety Notice NatalCare LX Incubator letter (DOC-017939A) to consignees on 1/8/16.A repair kit to correct this issue is now available for repair of affected units. Consignees should complete the NatalCare LX Incubator Oval Port Hinge Repair Kit Order Form and FAX it 847-949-6031 or scan and email this form to Natus_Quality_Programs@natus.com.Alternately you can telephone Natus Technical Service at 888-496-2887 and order the kit from a specialist. Natus will send the kit(s) free-of-charge to the address you specify. Each upgrade will require up to Âż - 1 hour of time and require simple hand tools.If you have any questions, please contact Technical Service at 888-496-2887 or send email to: customer.service@natus.com |  | 75 | 75 units (65 units in the US and 10 units in Pakistan) | Distribution US nationwide and Pakistan. | Open, Classified |  |  | 75 | Not_Computer
Z-1336-2016 | 73412 | No | MDS |  |  | Clinical Chemistry | PMA | Dexcom G4 PLATINUM (Professional) Receiver |  |  | February 23, 2016 | 2016 | Dexcom Inc | Customers may not receive an intended audible alarm or alert if relying on hearing the alarm or alert. | Device Design | Dexcom sent out a customer notification letter by certified mail to all direct customers on 02/23/16 notifying customers that they may not receive an intended audible alarm or alert if relying on hearing the alarm or alert. As a result customers may not detect a severe hypoglycemic (low glucose) or hyperglycemic (high glucose) event.The letter instructs the customer to peridically test the alarms and alerts on their receiver to make sure that the alarms and alerts are functioning properly (especially if the receiver gets wet of is dropped). Receiver part numbers and instructions for testing the alarms and alerts can be found on the back of the notification as well as in the User's Guide.Should customers find that alarms and alerts are not functioning properly or if any questions regarding the notification contact the deciated hotline tool free at 844-607-8398 or dial 858-291-1700.The healthcare professional notification that was sent out provided similar instructions as the direct customer notificaiton in which it is recommended that healthcare professionals periodically check the speaker function by following the testing instructions attached in the notification.The distributor notification letter instructed them to take the following actions:Hold shipments of Dexcom receivers and sticker all existing Dexcom receiver boxes on the outside front of the box with the enclosed notification stickers (see the attached visual guide for sticker placement). Once completed can begin shipping the products.In order to ensure that the notification is timely received by patients select one of the following two (2) options by no later than February 25, 2016:Send, via certified mail, the notification directly to those patients that have purchased a Dexcom receiver from you. If elect this option, Dexcom will provide the required quantity of notifications enclosed in certified mail envelopes. In order to facilitate this effort email the number of notifi |  | 5722 | 5,722 | Worldwide Distribution-US (nationwide), Canada, Italy, Netherlands, Belgium, Norway, New Zealand, Poland, Sweden, Denmark, Slovenia, Turkey, South Africa, India, Israel, Hungary, Hong Kong, UK, Ireland, France, Finland, Spain, Portugal, Colombia, Chile, Germany, Austria, Switzerland, Kuwait, United Arab Emirates, Australia, Czech Republic, Bahrain, Lebanon, and Qatar. | Terminated | August 11, 2017 | 535 | 270111 | Not_Computer
Z-1364-2016 | 73528 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Diana Automated Compounding System, Diana Onco Plus, Item No. ASN220 |  |  | March 08, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has received a report of compounded volumes not meeting the accuracy specification by the Diana Compounding System when mixing volumes under 5ml. | Use error | ICU Medical, Inc. sent customer notification letter dated March 8, 2016, to customers to inform them that ICU Medical Inc. has received a report of compounded volumes not meeting the accuracy specification by the Diana Compounding System when mixing volumes under 5ml. The letter informs the customers that ICU Medical will provide an updated Operator's Manual to provide clarity on the difference between Operating Range and Accuracy Range as soon as they are available. Customers are informed that the updated specifications are provided with the customer communication and to be used until the Operators Manual is provided. Customers are instructed to return the completed Medical Device Safety Communication Response form to ICU Medical via Fax (801) 264-1755 or by email at recall@icumed.com. Customers with any questions or require assistance relating to the customer notification letter are instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time:(866)8299025 and select option 8, or email the following address:productreturns@icumed.com. |  | 36 | 36 units | Nationwide Distribution to Nationwide in AL, AZ, CA, CO, CT, DC, GA, IL, MD, MI, MN, MO, NC, NH, NJ, NY, OH, OK, PA, TN, TX, VA..Worldwide: AU, BE, BR, CA, CH, DE, DZ, EE, EG, ES, FI, FR, HK, HR, IL, IT, JO, KR, MX, NL, PL, SA, SE, SI, TR. | Open, Classified |  |  | 196 | Not_Computer
Z-1371-2016 | 73390 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX-i oxygenators |  |  | February 23, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | Maquet Cardiopulmonary (MCP) is conducting a voluntary recall based on the potential risk of exposure to endotoxin for various products. | Unknown/Undetermined by firm | The firm, Maquet Inc., sent an "URGENT MEDICAL DEVICE RECALL" Letter/Response Form dated 2/23/2016 via Fed Ex to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review the list to determine if you have the affected products; monitor patients for signs; monitor and treat signs according to your protocols and clinical judgments; support shock symptoms and maintain circulatory and hemodynamics per your facility's protocol and care guidelines, and complete and and return the enclosed Response Form by email to: MCPrecall2016@maquet.com or fax to: 1-973-629-1518 within three (3) business days.For any questions, please contact your Maquet sales and representative or Maquet Customer Service at 1-888-627-8383 (press option 2, followed by option 2) Monday through Friday, between the hours of 8:00am and 6:00pm EST. |  | 72000 | 72,000 units in total | US Nationwide Distribution. | Open, Classified |  |  | 72000 | Not_Computer
Z-1377-2016 | 73649 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System (IS4000) user manual English and OUS. |  |  | March 16, 2016 | 2016 | Intuitive Surgical, Inc. | Intuitive Surgical has found that the use of instruments other than the Cardiac Probe Grasper (e.g. LargeNeedle Driver Instrument) to manipulate the ablation probes during cardiac surgery can result in an increase of microscopic metallic particulates generated from the contact with the probes. | Device Design | Intuitive sent a Field Safety Notice/Urgent Medical Device Correction letters dated March 22, 2016 by trackable mail method to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. In addition to existing instructions for use, observe the following warning when manipulating the cardiac ablation probe, Use only the Cardiac Probe Grasper to manipulate a cardiac ablation probe.2. Distribute a copy of this letter to all da Vinci S, Si, and Xi users at your facility.3. Place a copy of this letter with your user manual.4. Complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed.5. Retain a copy of this letter and the Acknowledgement Form for your files.Intuitive Surgical will provide a user manual addendum to incorporate the warning related to this risk.For questions contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service. |  | 248 | 248 - total all models | Worldwide Distribution - US Nationwide in the countries Australia, Belgium, Brazil, Canada, Chile, Finland, France, Germany, Greece, Iceland, India, Israel, Italy, Netherlands, Qatar, Russia, Singapore, South Korea, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 248 | Not_Computer
Z-1375-2016 | 73630 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | Boston Scientific TSX Transseptal Needle, Model Number M004TSX60o; |  |  | March 15, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily conducting a recall of one lot of the TSX" Transseptal Needle due to a discrepancy between the labeling and the actual configuration of the device in the packaging. The curve dimensional values on the labels of this lot incorrectly indicate a 50Âż curve, rather than the correct 86Âż curve | Software design | The consignee was provided verbal and written notification of the field action on March 15, 2016. There was a total of one(1) consignee, an OEM customer in the USA, who received 25 units. Written notification included product identification information, instruction to immediately quarantine any devices and discontinue distribution, and instructions to initiate a product recall in accordance with their company's established policies for customers to whom they shipped the affected Merit product. An email notification was sent to the Merit sales representative of the affected account. The sales representative was provided a Customer Response Form applicable to their account. The sales representative was instructed to assist the account with the completion of the Customer Response Form and the return of the affected units back to the Merit Field Assurance Department. |  | 25 | 25 | Distributed to one customer in the US. | Open, Classified |  |  | 25 | Not_Computer
Z-1370-2016 | 73719 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, PT Graphix Magnet J-TIP Guidewire |  |  | March 28, 2016 | 2016 | Boston Scientific Corporation | Boston Scientific is voluntarily implementing a Medical Device Recall of one batch of PT Graphix Magnet J-TIP Guidewires. Boston Scientific has found that, while the batch is labeled as a J-tip model, the guidewires are straight and do not have the pre-formed J-tip. The most severe injury expected to occur is a limited procedure delay caused by the need to replace the guidewire with one that has | Process control | Boston Scientific sent an "Urgent Medical Device Recall" letter dated March 28, 2016. The letter was addressed to Cath Lab Materials Manager / Field Action Contact. The letter described the problem and the product affected by the recall. Advised consignees to cease immediately further distribution or use of the product and to notify their customers. They also requested consignees to complete and return "The Reply Verification Tracking Form" For questions contact your local Boston scientific sales Representative of the Boston Scientific Filed Action Center.. |  | 47 | 47 units | US nationwide in the states of ME, LA, and TN | Open, Classified |  |  | 47 | Not_Computer
Z-1381-2016 | 73383 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys CS2 anesthesia.. |  |  | February 17, 2016 | 2016 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue involving a momentary, self-correcting anesthetic agent bolus when using 21% oxygen on all Aisys CS2 and upgraded Aisys anesthesia devices. | Nonconforming Material/Component | Consignees were sent on 2/17/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 34075 dated February 17, 2016.The letter was addressed to Chief of Anesthesia, Director of Biomedical / Clinical Engineering, and Health Care Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GEHealthcare Service at 1-800-437-1171 or your local Service Representative. |  | 5496 | 5496 | US: AL, AZ, AR, CA, CO, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI.OUS: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, CANDA, CHILE, CHINA, COLOMBIA, CROATIA, CRÂżCHE REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY,GREECE, HELSINKI, HONG KONG, HUNGARIA, ICELAND, INDIA, IRAQ, IRELAND, ITA LY, JAPAN, JORDAN, KAZAKHSTAN, KOREA, KUWAIT, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, PERU, POLAND, PORTUGAL QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM,URUGUAY, VIETNAM. | Open, Classified |  |  | 6790 | Not_Computer
Z-1380-2016 | 73665 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ DIAGNOSTIC ULTRASOUND SYSTEM, Model EPIQ 5C, EPIC 5G, EPIQ 5W, EPIQ 7C, EPIC 7GC, and EPIQ 7W. |  |  | March 28, 2016 | 2016 | Philips Ultrasound, Inc. | The fasteners securing the control panel assembly to the base of the Philips EPIQ Ultrasound System may loosen over time, which could subsequently lead to the detachment of the entire assembly from the ultrasound system. | Device Design | The firm, Philips, sent an "Urgent-Medical Device Correction" Philips EPIQ Ultrasound System (MDC 79500381/2), letter dated 2016 MAR 23 to consignees on 3/28/16. The letter describes the product, problem and actions to be taken. The customers were instructed to review the information with all members of your staff who need to be aware of the contents of this communication.If, at any time, the control panel assembly on your ultrasound system wobbles or feels loose, stop using your system immediately and contact your local Philips representative or Philips Customer Service at 1-800-722-9377. Otherwise, you may continue to use your system.Philips will contact all EPIQ customers to arrange for service to replace the fasteners connecting the control panel assembly to the system with fasteners less likely to loosen with repeated handling. This service will be performed free of charge. If you need any further information or support concerning this issue, please contact your local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 11085 | 11,085 units total (4909 units in the US and 6176 units outside the US) | Worldwide Distribution: US (nationwide) including Washington, D.C., and countries of: Algeria, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bermuda, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Guiana, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kazakhstan, Kenya, Korea, Republic of, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mayotte, Mexico, Monaco, Mongolia, Morocco, Myanmar, Netherlands, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, State of, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, RÂżÂżunion, Romania, Russia , Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, and Viet Nam. | Open, Classified |  |  | 11085 | Not_Computer
Z-1378-2016 | 73610 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Spetzler Claw Tip, Universal - 5450-800-311 |  |  | March 22, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The Spetzler Claw Tip was packaged with the incorrect Tip Sleeve Assembly. The Tip Sleeve Assembly included in the packaging is the incorrect size.Risk to Health:If the Tip Sleeve Assembly does not fit the Spetzler Claw" Tip, a surgical delay may occur while an alternate Tip Sleeve Assembly is retrieved. | Mixed-up of materials/components | On 3/22/2016, URGENT MEDICAL DEVICE RECALL NOTIFICATION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Kara Spath269-389-4518kara.spath@stryker.com |  | 140 | 140 units | Distributed in the states of AZ, CA, CO, CT, IL, IN, MD, MI, MN, NY, PA, SC, TN, WA, WI, and VA, and in the NETHERLANDS. | Terminated | July 19, 2016 | 119 | 140 | Not_Computer
Z-1379-2016 | 73595 | No | LWH | Surgical sponge scale. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | SurgiCounter scanner, part of the SurgiCount Safety-Sponge System. |  |  | March 21, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The affected scanner does not correctly interact with the SC360 software. When docked pre/postoperatively, the affected SurgiCounter scanner may display an error message (code=2). Also, the SurgiCounter scanner software version number is not appropriately displayed in the SC360 application during installation. | Software design | Stryker sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letters dated March 21, 2016 to affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Kelly Jo Davis269-389-2921kellyjo.davis@stryker.com |  | 286 | 286 | US Nationwide Distribution in the states of AR, AZ, CA, CT, FL, GA, IN, LA, MD, MN, MO, MS, NJ, NM, NY, OH, SC, SD, TN, TX and UT. | Terminated | September 06, 2016 | 169 | 286 | Computer
Z-1397-2016 | 72992 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Catalog Number: 9001771 MAYO STAND COVER STERILE |  |  | January 07, 2016 | 2016 | Customed, Inc | Recall initiated due to inadequate validation/qualification of sterilization processes, and package integrity for products released for distribution as sterile. | Packaging | Recall Notification Letter sent to direct accounts via certified mail with return receipt on January 7, 2016 and January 8, 2016. |  | 540 | 540 | Distributed only in Puerto Rico. | Terminated | May 03, 2017 | 482 | 62224 | Not_Computer
Z-1388-2016 | 73054 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | The Concerto Basic Shower Trolley is intended for assisted hygiene care, especially showering and bathing of residents in care environments such as senior/assisted living, group home, special care, nursing homes, hospitals and home care. The Concerto Basic Shower Trolley is intended for indoor use. |  |  | December 21, 2015 | 2015 | Arjo, Inc. dba ArjoHuntleigh | Reports of the bolts connecting the stretcher to the hydraulic piston and the stretcher tilting mechanism becoming loose. | Process change control | An Urgent Field Safety Notice and Customer Response form was mailed to affected customers describing the reason fro the recall and the actions to be taken. The instruction to customers is to review the Field Safety Notice and complete the Customer Response Form to acknowledge receipt and understanding of the Field Safety Notice and the required actions.These actions taken are as follows:1. Ensure that all caregivers and users of the ArjoHuntleigh Concerto or Basic Shower Trolley are made aware of this Field Safety Notice (FSN) and all affected Concerto Shower Trolleys at your facility will be available for inspection by an ArjoHuntleigh Service Technician. 2. Ensure that a copy of this field safety notice (FSN) is placed in the Concerto/Basic Shower Trolley Instructions for Use manual for future reference and training. 3. Complete and sign the enclosed Customer Response Form and return this form to the ArjoHuntleigh Regulatory Department. Note: ArjoHuntleigh will contact the designated facility representative listed on the Customer Response Form to schedule service to inspect and upgrade your device, free of charge. ArjoHuntleigh advises customers to distribute the notice to those individuals who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. Additionally the they should maintain awareness of the notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. The Notice includes a Note for the customer to inform FDA of adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online, by regular mail or by fax. The letter states that if the cusstomer has additional questions or they require assistance completing the Customer Response Form, please contact ArjoHuntleigh at 1-800-323-1245, ext. 57985 or via email at AHNA-RA |  | 1596 | 1,596 units | US Consignees; 222 units distributed in the following states: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV. ***Foreign Consignees; 1374 units distributed in the following countries: Australia, Austria, Belgium, Bosnia &amp; Hercegovina, Brazil, CANADA, China, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, Iceland, Ireland, Italy, Japan, Lithuania, Luxemburg, Netherlands, New Zealand, Nigeria, Norway, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan and UK. | Open, Classified |  |  | 1596 | Not_Computer
Z-1394-2016 | 73430 | No | NFG | Medical image communications device. | Radiology | Ophthalmic | 510(K) Exempt | VERION Reference Unit (Vision Planner) |  |  | March 01, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting a voluntary medical device correction of all VERION Reference Units (Vision Planner) that are shared with the Alcon LenSx Laser System after receiving reports concerning the inclusion of unplanned arcuates on printed, saved (.pdf) or exported surgical plans. | Software change control | Alcon sent an Urgent Medical Device Correction letter dated February 29, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that an Alcon Technical Services Engineers will contact customers, schedule Service Calls, and perform onsite corrections. For questions call 1-800-862-5266. |  | 97 | 97 units | Worldwide Distribution - US Nationwide throughout the US and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Czech Republic, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, France, Georgia, Greece, Guatlemala, Hungary, India, Indoesia, Isreal, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Nicaragua, Norway, Panama, Philippines, Portugal, Romania, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, UK, Ukrain, United Arab Emirates | Terminated | April 13, 2017 | 408 | 97 | Not_Computer
Z-1395-2016 | 73454 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | BARD ArcticGel Pads |  |  | March 04, 2016 | 2016 | C.R. Bard, Inc. | Complaints were received of restricted/inaccurate flow rate. | Process control | Bard Medical sent an Urgent Medical Product Recall letter dated March 4, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to not use any of the affected products and to return any unused product Bard Medical Division via the provided Customer Recall and Effectiveness Form. BMD will be providing a replacement for the returned product. If the product was further distributed, the consignee was instructed to identify their customers and notify them of the product recall. |  | 21412 | 21,412 units | Worldwide Distribution - Domestic: US Nationwide; Foreign: Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Brazil Canada, Chile, China, Colombia, Croatia, Cyprus, Denmark, Egypt, Finland, France, Germany, Greece, Oman Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, , South Korea, Kuwait, Macedonia, Malaysia, Mexico, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, The Netherlands, Turkey, United Arab Emirates, Ukraine, United Kingdom, Uruguay and Vietnam. | Open, Classified |  |  | 21412 | Not_Computer
Z-1392-2016 | 73552 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Oblique Posterior Atraumatic Lumbar (OPAL) System, Polymeric spinal fusion cage, non-sterile, OPAL Spacer 10 mm x 24 mm, 11 mm Height - Revolve |  |  | February 18, 2016 | 2016 | Synthes (USA) Products LLC | The label on the outer packaging does not match the product inside and intended use of the device. | Labeling mix-ups | DePuy Synthes sent an Urgent Notice letter dated February 18, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review inventory, immediately remove the affected lots from their stock and complete the attached Verification Section. For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant |  | 6 | 6 units | Distributed in the US: OH and ND | Open, Classified |  |  | 6 | Not_Computer
Z-1389-2016 | 73568 | No | MNL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Cygnus Medical First StepBedside ADDWATER Kit,non-sterile. |  |  | March 17, 2016 | 2016 | Cygnus Medical | Foreign material in container | Packaging | Cygnus Medical LLC sent an Recall Notice letter certified mail to direct customers and receivers on March 17, 2016. The letter identified the product the problem and the action needed to be taken by the customer. The letter was preceded with an email or fax to expedite notification. Customers to respond via fax or email regarding the quantities in their possession, including a response where the quantity is zero. Returns to made to firm. Contact Quality Manager Cygnus Medical LLC, phone: (203) 488-4554 x 105 fax: (203) 764-1858.Email: qualitycontrol@madpoly.com.For further questions, please call (203) 488-4554. |  | 4680 | 39 cases of 120 units | US Distribution to the states of : CA , CO, FL, IL. MN, MT, NY, OH, PA, TX, VA, VT, WA and WI. | Open, Classified |  |  | 4680 | Not_Computer
Z-1390-2016 | 73597 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | BD-Retic-Count Reticulocyte Reagent System; |  |  | March 18, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | The carton label for one lot of BD Retic-Count Reticulocyte Reagent System had the incorrect storage conditions. The labels on the reagent as well as the Instructions for Use are correct. | Under Investigation by firm | BD Biosciences sent an Urgent Product Recall letter on March 18, 2016, to all affected customers by certified mail. The letter identified the product the problem and the action needed to be taken by the customer. Customer Actions:1. Examine your data files obtained using reagents from the affected lots for low performance.2. Discard all remaining stock of BD Retic-Counter reticulocyte reagent system, Catalog Number 349204, lot number 5175559, and request replacement product. BD will replace product at no cost.3. As a requirement of our Regulatory Tracking Process, we request that you complete the enclosed Tracking/Verification form , even if you no longer have the affected reagents listed in your inventory.Please return the Tracking/Verification form to BD Biosciences, per the instructions (mail, fax, or email).If you require further assistance, please contact BD Customer Support at 877.232.8995 (prompt 3, 3) in the United States. For customers outside the US, contact your local BD Biosciences representative or distributor. |  | 205 | 205 | Worldwide Distribution - US Distribution to the states of : CT, MN, FL, NY, CA, MA, DC, UT, VA and IN., and to the countries of : Belgium, Brazil, Canada, China, India, Japan, Korea, Singapore and Taiwan. | Open, Classified |  |  | 205 | Not_Computer
Z-1497-2016 | 73339 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMT G-JetÂż Low-Profile Transgastric-Jejunal Feeding Device 14F 1.5cm 45cm SINGLE USE ONLY Rx ONLY STERILE. |  |  | January 26, 2016 | 2016 | Applied Medical Technology Inc | The firm received complaints that the product is shorter in length than indicated on the label. | Employee error | On 1/26/2016 the firm telephoned their customers notifying them of the recall. On 2/19/2016 the firm sent recall notification letters to their customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 20 | 20 units | Distributed in the states of AZ, CA, DE, IN, MI, NM, PA &amp; VA. | Open, Classified |  |  | 20 | Not_Computer
Z-1391-2016 | 73544 | No | PLV | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | X-Guide Surgical Navigation System, X-Clip Bulk Kit |  |  | February 23, 2016 | 2016 | X-NAV Technologies, LLC | The X-Guide system utilizes an X-Clip (PN P007665) which contains fiducials to assist with registration of the patient anatomy. The X-Clip is a device worn intraorally by a patient during the CT scan and guided surgery procedures. The X-Clip contains three fiducial markers made from a Stainless Steel material. The X-Nav specification requires these fiducials be fabricated from Stainless Steel 3 | Nonconforming Material/Component | The firm, X-Nav Technologies, LLC, sent a "Medical Device Recall" letter date 2/23/2016 via mail and email to customers indicating the issue and risks involved. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory of X-Clips and remove any product containing the affected lots; complete the enclosed Acknowledgement and Receipt form to record the lot numbers located on the labeling for parts and discard affected product and return form to firm by Email to: Vigilance@X-Navtech.com or by Mail to: X-NAV TECHNOLOGIES, LLC, 1555 BUSTARD ROAD, SUITE 75,LANSDALE, PA. 19446, ATTN: Director, Quality Assurance and Regulatory Affairs, even if you have already used or no longer have possession of the X-Clips identified. Should you have any questions regarding this notification, please contact our Customer Service representatives at 267.436.0420 from 8:30a.m. to 5:30p.m. (EST) Monday-Friday. |  | 845 | 845 units (P007665) | US Distribution to: FL, OR, NE, PA, NM, FL, VA, OH, WA, CA, MD, LA, ID, CN, TN, and DC. | Open, Classified |  |  | 845 | Not_Computer
Z-1503-2016 | 73685 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ad-Tech Electrode Connection System/Cables. Lightweight TECH-ATTACH Cable Models: |  |  | March 29, 2016 | 2016 | Ad-Tech Medical Instrument Corporation | Ad-Tech is recalling unopened/unused sterile cables intended to be used only with all compatible Ad-Tech electrodes for the purpose of connecting these electrodes to third party monitors/stimulators. This recall has been initiated due to package failures (i.e. compromised sterile barriers) that were identified during the execution of a shipping simulation investigation. | Package design/selection | Consignees were sent a "Medical Device Recall" letter dated March 29, 2016. The letter described the problem and the product being recalled. Advised consignees to immediately examine their inventory, quarantine and return the product to AD-Tech. If the product has been further distributed identify customers and notify them of the recall. Requested consignees to complete and return the "Medical Device Recall Return Response" form to their Ad-Tech Clinical Specialist at customersupport@adtechmedical.com; FAX: 262-634-5668 or Telephone 262-634-1555. For questions contact Ad-Tech Clinical Specialist.. |  | 197 | 197 | Distributed in the states of GA and MI, and the countries of FRANCE, UNITED KINGDOM, COLOMBIA, and BRAZIL. | Terminated | August 24, 2016 | 148 | 197 | Not_Computer
Z-1499-2016 | 73557 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica Microsystems (Schweiz) Ag. The common name of the device is surgical microscope. It has been assigned the product code FSO by the FDA and is classified as a Class I medical device. |  |  | January 11, 2016 | 2016 | Leica Microsystems, Inc. | An electronic component of the Leica M525 OH4 and Leica M720 OH5 surgical microscopes with 400W illumination could potentially overheat, resulting in smoke and smell emanating from the ventilation openings of the illumination units located opposite and away from the surgeon / patient side of the surgical microscope. | Nonconforming Material/Component | The URGENT MEDICAL DEVICE RECALL letter was sent via USPS mail and email on January 25th 2016 to consignees describing the reason for the recall and the necessary steps to take. These included the following:Continue to use the system according to the User Manual. Both illumination units should be checked prior to beginning any operation case. The Leica M525 OH4 and Leica M720 OH5 provide two redundant illumination units. Independent from the issue explained, the user should be reminded that if they notice a failure in the illumination unit that they are currently using, they should switch to the back-up illumination at the next opportunity during surgery by using the "Change Lamp" page 37 for M525 OH4 and page 40 for M720 OH5 button on the control unit "Main" screen (see Picture 2 below) as described in user manual. Changing lamps will interrupt the illumination for less than 5 seconds.They should contact their Leica Microsystems representative for repair of the defective illumination unit before using the instrument for the next surgery.Note: The user manuals of the Leica M525 OH4 and Leica M720 OH5 state:Âż Never begin an operation with only one functioning xenon lamp! Âż Like any other instrument in the operating theater this system may fail. Leica Microsystems therefore recommends that a backup system be kept ready during the operation. The recall further describes transmission of the field notice:The notice must be passed on to all those within their organization who need to be aware of it and to any organization where the affected devices has been sold, rented, leased or otherwise transferred.The end users are requested to confirm receipt of the notice by signing, dating and returning the attached Acknowledgement Form' with the serial number of their instrument back to Leica Microsystems by email to:LMGRA@leica-microsystems.comwithin 10 working days after receipt of field safety notice. |  | 381 | 381 units | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of MA, CA, SC, NJ, TX, CO, FL, PA, NY, MO, WI, NC, UT, ID, AZ, MI, KY, CT, OR, , IL, AL, MN, MT, OH, NE, GA, IN, SD, TN, OK, AK, KS, NV, NM, , ND, VA, WA, WY and MD, and the countries of China, Belgium, France, Germany, Great Britain, Italy, Norway, Portugal, Spain, Sweden, Switzerland, India, Japan, Argentina, Chile, Colombia, Mexico, Peru, Belarus, Egypt, Greece, Morocco, Poland, Qatar, Romania, Russian Fed., Saudi Arabia, South Africa, Turkey, Canada, , Australia, Malaysia and Taiwan. | Terminated | April 24, 2017 | 469 | 381 | Not_Computer
Z-1504-2016 | 73776 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Extended Brilliance Workspace NM Nuclear medicine image display and processing application suite. AutoSPECT Pro Software Application. Philips Medical Systems, Cleveland, OH. |  |  | March 04, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The AutoSPECT Pro application was only designed to reconstruct cardiac SPECT data obtained with detectors positioned at 90Âż or 180Âż relative to one another. However, certain gamma cameras allow for other relative detector angles. Data acquired at these other angles will not be correctly reconstructed by AutoSPECT Pro and the results will likely be erroneous. | Software design | On March 4, 2016, the firm sent an URGENT - Field Safety Notice Medical Device Correction letter to their customers. The letter identified the affected device, the reason for the recall, and under what circumstances the problem occurs. Customers are to contact Philips as described if they are unsure whether their gamma cameras supports the modes. The firm will implement the AutoSPECT Pro software upgrade and contact customers to schedule the appointment. If further assistance or support is needed, customers are to contact their local Philips representative. |  | 126 | 126 Units | Worldwide Distribution -- US, including the states of CO, DE, FL, GA, IL, IN, KS, MO, NC, NV, NY, OH, OR, PA, TN &amp; TX; and, the countries of Australia, Austria, Belgium, Canada, China, Denmark, Egypt, France, Germany, Greece, Ireland, Israel, Italy, Lebanon, Malaysia, Saudi Arabia, Slovenia, South Africa, Switzerland, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 196 | Computer
Z-1508-2016 | 73202 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Robotic Arm Interactive Orthopedic System (RIO) |  |  | April 30, 2014 | 2014 | Mako Surgical Corporation | Revision of the field manual | Process control | As per our conversation with firm representative on March 23, 2016, they stated "no notification was sent to the customers", and referred to Ref: Part 806 File number: RA 2015-017 where it states "Stryker Orthopaedics has completed a retrospective review of other revisions to the field service manual, and we have confirmed there are no other similar events for which Part 806 reports should have been submitted". For further questions, please call (866) 647-6256.Internal actions taken only. |  | 249 | 249 | Worldwide Distribution - Nationwide Distribution and to the countries of : Australia, Italy, Japan, Hong Kong, Scotland, Germany, Turkey , Greece, Singapore, Taiwan and South Korea. | Open, Classified |  |  | 249 | Not_Computer
Z-1509-2016 | 73643 | No | MZO |  |  | Microbiology | PMA | OraQuick HCV Visual Reference Panel |  |  | January 15, 2016 | 2016 | OraSure Technologies, Inc. | OraSure Technologies, Inc. discovered the package insert included with the OraQuick HCV Rapid Antibody test Visual Reference Panel may be incorrect due to a complaint received from one of their customers. The customer reported they received the Ora Quick Ebola Visual Reference Panel instead of an OraQuick HCV Visual Reference Panel. | Unknown/Undetermined by firm | Letters, dated January 15, 2016, and correct HCV VRP insert were sent to all customers who received HCV VRP lot 6648737. The letter stated the reason for the recall; and, asked users to discard the Package Insert that was included with the product and replace it with the corrected one provided with the letter. Questions regarding the recall should be directed to OraSure Sales Representative or our Customer Care Department at 1-800-ORASURE. |  | 244 | 244 Visual Reference Panel Kits | Nationwide Distribution -- AK, AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KY, LA, MA, MI, MO, MT, NC, ND, NE, NH, NM, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, and WI. | Open, Classified |  |  | 244 | Not_Computer
Z-1507-2016 | 73505 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | VariAx Compression Plating System |  |  | October 10, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received two reports indicating that during screw insertion under a high lateral angle a bone screw went through a oblong hole. | Device Design | Stryker Orthopaedics sent notification letters/product accountability forms on October 10, 2013, via email. Customers are provided the product issue, hazards and risk mitigation along with actions needed. Customers should receive the affected products and complete the BRF and return to Stryker. Questions can be directed to Stryker. For further questions, please call (201) 831-5000. |  | 118 | 118 units | Nationwide Distribution | Terminated | April 21, 2016 | 924 | 118 | Not_Computer
Z-1342-2016 | 73633 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Fetch 2 Aspiration Catheter, Catalog No. 109400-001. |  |  | March 22, 2016 | 2016 | Boston Scientific Corporation | Boston Scientific is recalling all models of Fetch 2 Aspiration Catheter as a result of receiving reports from the field of catheter shaft breaks; including cases requiring the use of a snare or other intervention to retrieve the http://webapps.cdrh.fda.gov/division-tracking/images/trash.pngbroken shaft from the patient. | Process control | The firm,Boston Scientific, "Urgent Medical Device Recall Notice - Immediate Action Required" letter dated 3/22/2016. The letter was addressed to Material Manger / Field Action Contact. The letter described the problem and the product involved in the recall and actions to be taken. The customers were instructed to immediately cease further distribution or use of the product and complete and return the Reply Verification Tracking Form via email: BSCFieldActionCenter@bsci.com or fax to: BSC Field Action Center at 1-866-213-1806. Distributors should notify their customers of the recall and inform them that this recall is to the hospital level. For questions regarding this communication, please contact your local Boston Scientific Sales Representative or the Boston Scientific Field Action Center at 763-494-1133 or email: BSCFieldActionCenter@bsci.com. |  | 21155 | 21,155 units | Worldwide Distribution - US (nationwide) including DC &amp; VI except DE, ID, SD, and WY and the countries of Australia, Austria, Belgium, Bulgaria, Canada, Chile, Croatia, Cyprus, Ecuador, France, Germany, Great Britain, Hungary, Italy, Japan, Jordan, Korea, Malaysia, Mauritius, Mexico, Netherlands, Philippines, Portugal, South Africa, Spain, Sweden, Thailand, Trinidad, Tobago, and Uruguay. | Terminated | August 11, 2017 | 507 | 21155 | Not_Computer
Z-1510-2016 | 73382 | component, traction, non-invasive | KQZ | Noninvasive traction component. | Orthopedic | Orthopedic | 510(K) Exempt | Ossur Lil Angel Vest P2 Kodel. Pediatric sized vest is a support structure used with the halo system for traumatic or chronic disorders requiring traction along the cervical spine. |  |  | March 14, 2016 | 2016 | Ossur H / F | There is a potential for partial or complete fracture of the plastic stabilizer strap on the vest. | Component design/selection | Customer notification letter were sent on 03/14/16 via traceable method.The notification asked customers to examine their inventory, quarantine product subject to the recall and contact customer service for a return authorization to return the product to Ossur.Customers were asked to take the following actions:1. Pass the notice to those who need to be aware within the organization or to any organization where the potentially affected devices have been transferred.2. Maintain awareness on the notice for an appropriate period.3. If have further distributed this product identify the customers and notify them at once of this product alert. It is recommended that a copy of the notice is included.4. If any customers are currently wearing a product it is recommended a detailed inspection of the product vest structurebe conducted to ensure there is no visible cracks on the vest. |  | 435 | 435 total all models | Distribution US Nationwide and Hong Kong | Open, Classified |  |  | 435 | Not_Computer
Z-1516-2016 | 73705 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | GDC-10 360 SOFT 10MM X 20CM SR Detachable Coil |  |  | March 21, 2016 | 2016 | Stryker Neurovascular | The incorrect DFU was packaged with two lots of GDC 360 degree Detachable Coils. | Under Investigation by firm | The firm, Stryker Neurovascular, sent an "URGENT: Field Safety Notification" dated 3/21/2016 via courier to its consignees/customers. The notification describes the product, problem and actions to be taken. The customers were instructed to immediately check your inventory; follow the instructions listed in the Directions For Use (DFU); circulate and maintain the notice internally to all interested/affected parties; inform Stryker if any of the subject devices have been distributed; complete and return the FIELD SAFETY CORRECTIVE ACTION AKNOWLEDGMENT FORM to Stryker via fax at 1 (866) 876-4355 within 7 calendar days, as well as notify any other organizations that the device may have been distributed to.Should you have any queries concerning this matter contact Quality Manager at 510-413-2593 or email: Geraldine.ahern@stryker.com |  | 316 | 316 total, both models | Worldwide Distribution: US (nationwide) including states of: OH, CO, WA, IL, MI, CA, TX and countries of: Hong Kong, Italy, Germany, France, Spain, Australia, Canada, Sweden, India, Turkey, Cyprus, Belgium, Switzerland, Czech Republic, Uruguay, Mexico, Saudi Arabia, Russia, Brazil, Netherlands, and Ukraine. | Open, Classified |  |  | 316 | Not_Computer
Z-1522-2016 | 73655 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Emerald Cleaner, product number 09H46-02. |  |  | March 24, 2016 | 2016 | Abbott Laboratories | Three lots of CELL-DYN Emerald Cleaner may have occurrences where the analyzer generates Quality Control Out of Range Low for parameters RBC and PLT. | Under Investigation by firm | All customers who were shipped the affected lots were sent a Product Recall letter on March 24, 2016 via Federal Express.Letters were titled Product Recall - Immediate Action Required.An additional four lots were discovered to be affected by the recall and the firm send letters to customers affected by these lots on June 27, 2016 |  | 19774 | 10,364 plus 9410 additional units for recall expansion. | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Austria, Bosnia &amp; Herze, Cameroon, Egypt, Germany, Indonesia, Iraq, Lebanon, Mali, Montenegro, Netherlands, Philippines, Tanzania, Thailand, UAE, New Zealand, Malaysia, Mexico, South Korea, Jamaica, Turks &amp; Caicos, Italy, Canada, Vietnam, Singapore, Argentina, Colombia, France, Paraguay, Australia, Uruguay, Bahrain, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Ethiopia, Finland, Greece, Switzerland, United Kingdom, Brazil, Hong, Kong, Hungary, India, Ireland, Israel, Kazakhstan, Kenya, Kuwait, Latvia, Lithuania, Luxembourg, Turkey, Jamaica, St. Martin, Nigeria, Oman, Pakistan, Philippines, Poland, Qatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, Guatemala, Peru, and Uruguay. | Terminated | January 04, 2017 | 286 | 19774 | Not_Computer
Z-1519-2016 | 73593 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Definition AS; Intended to produce cross-sectional images of the body. Computed tomography (CT) x-ray system. |  |  | March 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software bugs in these SOMATOM systems could possibly cause scan aborts, rescans of patients with additional dose, delayed diagnosis and as worst case scenarios, could possibly cause the need of additional contrast medium, accidental x-radiation to the user or impede customer workflow. | Software design | Siemens initiated a Customer Advisory Notice on 03/07/2016 to inform customers about actions for bug-fixing the following systems: SOMATOM Definition AS, SOMATOM Definition Flash and SOMATOM Definition Edge, which could possibly cause scan aborts, res cans of patients with additional dose, delayed diagnosis and as worst case scenarios, could possibly cause the need of additional contrast medium. Siemens mailed certified as well as via email distribution a Customer Advisory Notice to their customers to inform them about corrective actions for bug-fixing the following systems running with software version VA44A_SP3a, VA44A_SP3b, VA44A_SP3c, correction of potential safety issues. The issues described will be corrected with software VA44_SP4, handled as update CT051/15/P. Following the installation of this update, the mentioned malfunctions will be corrected. |  | 1096 | 1096 | Nationwide Distribution-including DC, and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Open, Classified |  |  | 1096 | Computer
Z-1523-2016 | 73752 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Mach 4 Mouse Probe vial within the Mach 4 Universal HRP-Polymer Kit; |  |  | March 31, 2016 | 2016 | Biocare Medical, LLC | Biocare Medical has identified a single lot (082715) of its Mach 4 HRP Polymer kit in which some few Mach 4 Mouse Probe vials may have been mislabeled as HRP-Polymer. | Labeling mix-ups | Biocare Medical sent a Medical Device Recall Notification letter dated March 31, 2016 to affected US customers by USPS. Foreign consignees were also notified per their various government regulations. The letter identified the affected product, problem and actions to be taken. Customers were requested to complete and return by April 15, 2016 the enclosed Return Response Form. |  | 93 | 93 | US Nationwide Distribution in the states of: KS, CA, AZ, OR, FL, KS, CO, MA, KY, NY, IL, OR, MO, NC, TX, PA, OK. SC, TN, PA, TN, GA, and OR.Hungary, Canada, Boliva | Open, Classified |  |  | 93 | Not_Computer
Z-1530-2016 | 73509 | No | PCD | Powered patient transport. | Physical Medicine | Physical Medicine | 510(K) Exempt | Acorn 180 Curved Stairlift |  |  | February 17, 2016 | 2016 | Acorn Stairlifts Inc | Aluminum rivets holding the base squab to the framer under prolonged customer use were seen to elongate and subsequently break. | Component change control | Customers were notified by phone calls, documents mailed with a Rationalization Action plan to follow. Contact information 1-888-563-0410 x1500 or 1400 |  | 10897 | 1,290 USA and 9,607 Foreign | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, CO, DE, FL, GA, IL, IN, KS, KY, LA, MA, ME, MD, MI, MO, NH, NJ, NY, NC, OH, PA, RI, SC, TN, TX, UT, VT, VI, WA, WV, WI, and District of Columbia and to the countries of UK and Germany. | Open, Classified |  |  | 10897 | Not_Computer
Z-1500-2016 | 73588 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 560 Hematology Systems, Siemens Material Number (SMN) 11170842, IVD |  |  | March 10, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Software anomaly; Siemens identified that software version 1.4.2133 on the ADVIA 560 Hematology System does not trigger the following flags: The G or L morphology flags for immature granulocytes (IG) and atypical lymphocytes (ATYP), respectively. The results from patient samples which have immature granulocytes or atypical lymphocytes will not generate the flags when they should. As defined in t | Process design | Siemens sent an Urgent Medical Device Correction Letter dated March 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. .Actions to be Taken by the Customer:To ensure that the patient sample amount for processing is met, follow the requirements for the amount of sample required in the ADVIA 560 Hematology System Operators Guide. Review all patient results for abnormal white blood cell results and process as defined by your lab criteria for abnormal samples and confirm using manual differentials as needed.Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.Please review this letter with your Medical Director.Your Siemens Customer Service Engineer will contact you to schedule time for installation of the 1.4.2333 software when it is available.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (312) 275-7795. |  | 72 | US: 8 systems; Foreign: 64 systems | Worldwide Distribution - US Distribution including Puerto Rico and to the states of : California, Illinois, Washington and Wisconsin., and to the countries of : Angola, Austria, Canada, Chile, Colombia, Czech Republic, France, Germany, India, Italy, Kenya, Netherlands, Norway, Portugal, Singapore, Slovenia, Spain, Turkey, Uganda and United Kingdom. | Open, Classified |  |  | 72 | Computer
Z-1525-2016 | 73604 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Self-CathÂż Pediatric Catheter Size CH 06, Product # 5044101400, Catakig # 306. |  |  | March 22, 2016 | 2016 | Coloplast Manufacturing US, LLC | Coloplast Corp. is voluntarily recalling a single lot of Self-Cath Pediatric French Size CH 06 cathetersbecause a small portion of the lot was identified to have missing eyelets. This recall is based upon acomplaint Coloplast received regarding a catheter malfunction. | Nonconforming Material/Component | Coloplast sent an "Urgent Medical Device Recall" letter dated March 22, 2016 to customers. The letter described the problem and the product involved in the recall. Advised consignees to immediately examine, quarantine and return the affected product. If the product has been further distributed, consignees should identify their customers and notify them of the recall. It also requested consignees to complete and return the Acknowledgement Form. For questions contact Coloplast Customer Service directly at 800-533-0464. |  | 31320 | 31,320 | US Nationwide Distribution in the states of: AL, AR, C, FL, IL KY, LA, MD, MA, MI, MN, MO, NJ,NY, NC, OK,OR, PA, TN, TX, VA and WA. | Open, Classified |  |  | 31320 | Not_Computer
Z-1529-2016 | 73602 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 for Proton Radiation Therapy |  |  | March 15, 2016 | 2016 | Mevion Medical Systems, Inc. | Possibility of Gantry Motion with Aperture or Compensator only partially inserted. An Aperture may fall out during Gantry motion. It is possible that it can hit a patient and cause traumatic injury or death. | Device Design | Customer notification letters were sent to customers on March 15, 2016, by certified mail, return receipt requested, as well as by email. The letter explains the problem, description of potential harm, and the cause containment, and mitigation. The cause of the malfunctioning sensors was discovered to be a non-conforming limit switch assembly. A Mevion Field Service Engineer will immediately inspect all such parts in the field for similar non-conformance, and replace those parts if necessary. In addition, a new periodic preventive maintenance check will be performed to individually test each of the redundant sensors at all sites. In the interim period, on a daily basis, until the inspection, repair, and the preventive maintenance check is completed at any site, Mevion Field Service will check that the Applicator Door interlock sensors are working properly. Mevion will also update the inspection procedure for incoming parts to ensure that any non-conformance is identified early in the manufacturing process. |  | 5 | 5 | US nationwide distribution to MO, OK, FL, and NJ. | Open, Classified |  |  | 5 | Not_Computer
Z-1524-2016 | 73549 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Cervical Spine Locking Plate (CSLP) System, 4mm Cancellous Self Tapping Screw, Length 16mm Lot # 9666776 and 4mm Cortex Standard Screw, Length 16mm Lot # 9671879 |  |  | January 06, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a Voluntary Medical Device Recall of lot # 9666776 and 9671879 of Cancellous and Cortex 4.0 mm Screws, which are a part of the Cervical Spine Locking Plate (CSLP) System due to discovery of a single lot of 450.136 screws (4mm Cancellous Self Tapping Screw, Length 16mm, lot 9666776) being swapped with a single lot of 450.016 (4mm Cortex Standard Screw, length 16 mm, lot | Packaging process control | DePuy Synthesis sent an Urgent Notice Medical Device Recall letter dated January 6, 2016 to customers. The letter identified the affected product, problem and actions to be taken. The letter requested customers to complete and return the attached Verification Section (page 3 of the letter). For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 290003280 | 30 units of Lot # 9666776 and 11 units of Lot # 9671879 | Worldwide Distribution - US Nationwide in the states of NC, CA, NY, TX, NJ, CO and the countries of: Ontario and Canada. | Open, Classified |  |  | 290003280 | Not_Computer
Z-1540-2016 | 73609 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TC-PLUS(TM), Primary Fixed Tibial Component VKS Right, SIZE 2 RIGHT, STERILE R, REF 12 000 009, S&N; 7500264, Smith &amp; Nephew Orthopaedics AG |  |  | March 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | The Affected Devices may contain a dimensional deviation that could potentially result in thesurgeon being unable to seat the polyethylene (PE) Insert on the affected VKS tibial baseplate during surgery. | Process control | The firm, Smith &amp; Nephew, issued an "URGENT FIELD SAFETY NOTICE: Medical Device Field Safety Corrective Action/Recall" letter dated 3/10/2016 to its local QA/RA personnel/customers in the affected countries by email on 03/11/2016. The letter described the product, problem and actions to be taken. The customers were instructed to locate and quarantine affected unused devices immediately; return quarantined product to your national Smith &amp; Nephew agency/distributor; complete and return the Return Slip via fax to your national Smith &amp; Nephew agency/distributor; and pass on this safety information to all those who need to be aware in your organization. The local QA/RA staff will then notify end customers according to local procedures. In Germany as well as in Switzerland, sales representatives will inform the affected customers in-person and retrieve affected devices from their stock.If you have any questions, call 978-749-1330 or email: joe.metzger@smith-nephew.com |  | 1 | 1 unit | International Distribution only in countries of: Germany and Switzerland. | Terminated | April 10, 2017 | 395 | 191 | Not_Computer
Z-1553-2016 | 73349 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Pentra 400 (version 5.0.8 or lower) |  |  | February 19, 2016 | 2016 | Horiba Instruments Inc | Horiba Instruments, Inc. is recalling ABX Pentra 400( version 5.0.8 or lower) and Pentra C400 (version 1.1.2 or lower) because clinical chemistry analyzer malfunctions when the following certain alarms appear on the system. | Software design | HORIBA Medical sent an "Urgent Safety Notice" dated February 19, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new software version to solve the issue is currently under development; however the firm requests customers use the provided procedure until the new software version is installed on the system, or be extremely vigilant with results of one of the alarms mentioned appears on the system. The letter instructs customers to share the information with laboratory staff and retain the notification as part of laboratory documentation. The letter states it is mandatory for customers to complete and return the enclosed response form within 10 days, in order to verify understanding of it.Any questions or concerns contact the local Horiba Medical representative.For further questions, please call (949) 250-4811. |  | 201 | 201 | USA (nationwide Distribution) to the states of : CA, NY, AL,OK, MS, MD, NV, MN, TX, KY, WV, MI, NY, MT, IL, FL,WA, ID, MO, WY, NJ, VA, KY, GA, PA, TN, MS, KS, AZ, NE, SC, OR, OH, IA, RI, MI and NC. | Open, Classified |  |  | 201 | Computer
Z-1555-2016 | 71667 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit One A-C Powered Hospital Bed |  |  | July 02, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Customer complaints associated with faulty brake system | Device Design | Stryker Corp.sent an Urgent Medical Device Recall letter on July 21, 2015, via FedEx to all affected customers. The letter will identified the product the problem and the action needed to be taken by the customer. Customers are instructed the following actions: 1.Locate the units listed on the attached business reply form.2.Remove these units from service. (Note: The original mattress purchased with the bed is not within scope of this action. The mattress will be compatible with your replacement option.)3.Your Stryker Sales Representative will contact your facility to coordinate removal of the recalled beds and present replacement options.4.Return the enclosed business reply form to your Stryker Sales Representative to confirm receipt of this notification or fax (269)488-8691 or email productfieldaction@stryker.com 5.If you have loaned or sold any of the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form.6.If you have disposed of any of the beds and they are no longer in use, please advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form.*For additional information, please refer to the products operations and/or maintenance manuals. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax.If there are any questions or concerns, please contact Stryker Customer Service (1-800-327-0770) during normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 143 | 143 total | US Distribution to the states of : AL, AZ, FL, GA, HI, IA, IL, KS, LA, MI, MN, MO, NE, NJ, OH, PA, SC, TN and TX., and Internationally to the countries of : Canada and UK. | Terminated | March 24, 2017 | 631 | 143 | Not_Computer
Z-1556-2016 | 73131 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit One A-C Powered Hospital Bed |  |  | July 02, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | Inaccurate scale systems | Device Design | Stryker Corp. will initiate the recall by distributing letters on 07/21/2015 via FEDEX. Customers are asked to do the following: 1. Locate the units listed on the attached business reply form. 2. Remove these units from service. (Note: The original mattress purchased with the bed is not within scope of this action. The mattress will be compatible with your replacement option.) 3. Your Stryker Sales Representative will contact your facility to coordinate removal of the recalled beds and present replacement options. 4. Return the enclosedbusiness reply form to your Stryker Sales Representative to confirm receipt of this notification or fax (269)488 -8691 or email productfieldaction@stryker.com 5. If you have loaned or sold any of the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form. 6. If you have disposed of any of the beds and they are no longer in use, please advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form. *For additional information, please refer to the products operations and/or maintenance manuals. Adverse reactions or quality problems experienced with the use of this product may be reported tothe FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. If there are any questions or concerns, please contact Stryker Customer Service (1-800-327-0770) during normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 143 | 143 | Distributed in the states of AL, AZ, FL, GA, HI, IA, IL, KS, LA, MI, MN, MO, NE, NJ, OH, PA, SC, TN, and TX, and the countries of Canada and UK. | Terminated | March 24, 2017 | 631 | 143 | Not_Computer
Z-1528-2016 | 73205 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trinity Acetabular Shell |  |  | January 29, 2016 | 2016 | Corin USA Limited | Potential for abrasion of the sterile barrier pouches of the new packaging system. | Packaging process control | Corin sent an Urgent Medical Device Recall letter dated February 3, 2016. The letter identified the affected product, problem and actions to be taken. On January 29, 2016, Eight (8) of 11 distributors were spoken to during a Corin National Sales Meeting informing them of the recall. The 3 remaining distributors were notified by phone calls same date. Also, on February 2, 2016 Corin sent a Recall Letter to their Foreign consignee informating them of the recall of the Trinity Acetabular Shell and the actions to be take by the costumers including the return of the Customer Response Form and Contact information as Corin LTD Compliance Director at +44 (0) 1285 649190 Monday thru Friday 8:30am to 5:00pm UK Time |  | 629 | 629 devices in total | Worldwide Distribution - US Nationwide in the states of CA, FL, ID, TX, OK, OH, MA, and NC and the countries of Australia Belgium, Egypt, Germany, Ghana, Italy, Japan, Poland, South Africa and UK. | Open, Classified |  |  | 629 | Not_Computer
Z-1562-2016 | 73617 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys and Aisys CS2. |  |  | March 15, 2016 | 2016 | GE Medical Systems, LLC | The vacuum suction tubing could be kinked within the machine resulting in suction less than the 20 lpm required by ISO 10079-3. This could result in inadequate suction performance to clear the airway and inadequate ventilation of the patient. | Employee error | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" GEHC Ref# 34076 letter, dated March 15, 2016. The letter was addressed to the Chief of Anesthesia / Health Care Administrator / Risk Manager and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter states that users can continue to use all functions of the device. However, the vacuum suction option can only be used if it passes the evaluation. A GE representative will contact customers to arrange for the correction at no cost to the customer. For questions or concerns regarding this notification, contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 952 | 952 (488 US; 464 OUS) | Worldwide Distribution -- US, to the states of AZ, CA, DC, FL, GA, ID, IL IN, KY, MD, MI, MN, MO, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI; and, to the countries of AUSTRALIA, BELGIUM. BOLIVIA, CANADA, CHILE, DENMARK, FINLAND, FRANCE, GERMANY, HUNGARIA, INDIA, IRELAND, JAPAN, KAZAKHSTAN, KOREA, KUWAIT, NETHERLANDS, NEW ZEALAND, SAUDI ARABIA, SPAIN, SWITZERLAND, THAILAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 952 | Not_Computer
Z-1561-2016 | 73761 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Radionuclide Radiation Therapy System |  |  | April 19, 2016 | 2016 | Elekta, Inc. | The latches of the frame adapter can be locked even if the locating pins of the frame adapter is not inserted into the corresponding holes in the coordinate frame. | Use error | Elekta sent an Important User Notice (100-03-202-27) letter to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Acknowledgement Form attached. For questions contact your local Elekta office. |  | 217 | 217 | Worldwide Distribution - US Nationwide in the states of AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, KY, LA, MA, MI, MN, MS, MO, MT, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WV, WI, and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Iraq, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom | Open, Classified |  |  | 217 | Not_Computer
Z-1559-2016 | 73783 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Trypticase TM Soy Agar with 5% Sheep Blood (TSA IITM) // MacConkey II Agar |  |  | March 28, 2016 | 2016 | Becton Dickinson &amp; Co. | Trypticase TM Soy Agar with 5% Sheep Blood (TSA IITM) // MacConkey II Agar is contaminated with Listeria Monocytogenes. The device is a non-sterile prepared plated media. | Material/Component Contamination | BD sent an Urgent Product Recall letter dated March 2016 to all US customers via UPS or email. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected lots and discard any remaining product. BD will replace the discard inventory. Customers were asked to complete and fax the attached form to : Attention Regulatory Compliance, Fax: 410-316-4258 or email to BDRC2@bd.com. |  | 88300 | 88,300 plates (883 cartons) | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, KS, LA, MA, MD, ME, MI, MN, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TX, UT, VA, WA. | Terminated | November 14, 2016 | 231 | 88300 | Not_Computer
Z-1558-2016 | 73838 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force, Computed Tomography x-ray system |  |  | April 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The Neonate Head protocol with the pediatric kernel Hp38 could result in artefacts and possibly lead to a misdiagnosis (i.e. either non-existing blood or liquid is mimicked in the images or actual existing blood or liquid is not depicted as expected). There is also a risk of a potential misdiagnosis in using this protocol for surgery planning. Investigations revealed an incorrect parameterization of the reconstruction algorithm applied for the pediatric head kernel as cause of the problem. | Software design | Siemens distributed a Safety Advisory Notice dated April 1, 2016, to affected customers The letter identified the affected product, problem, how to avoid potential risks, how the issue will be resolved. and actions to be taken. Customers were asked to observe the safety notice and comply with the corresponding measures until further notice. |  | 37 | 37 CTsystems | US Nationwide Distribution : NC, NH, IL, MA, WA, PA, MD, MN, TX, NY, SC, OH, CA, KY, MT, IA | Terminated | April 17, 2017 | 381 | 37 | Not_Computer
Z-1552-2016 | 73706 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3, PowerCross" 0.018" OTW PTA Dilatation Catheter, model: AB18W030200150, Sterile EO, Rx Only. |  |  | March 29, 2016 | 2016 | ev3, Inc. | Medtronic is conducting a voluntary Urgent Medical Device Recall of two lots of the PowerCross 0.018" Âż Over the Wire (OTW) Percutaneous Transluminal Angioplasty (PTA) Dilatation Catheter because the product is mislabeled. | Labeling False and Misleading | Consignees were sent on 3/29/2016 a Medtronic "Urgent Medical Device Recall" letter dated March 2016. The letter was addressed to Risk Manger or Health Care Professional.The letter described the problem and the product involved in the recall. |  | 27 | 27 (25 US, 2 OUS) | US: CA, GA, LA, MI, MS, MO,NV,NH. NY, OH, PA, SD, TX.OUS: SINGAPORE | Terminated | May 19, 2017 | 416 | 27 | Not_Computer
Z-1557-2016 | 73673 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | eCareCoordinator allows the customer to schedule patient tasks (i.e., take weight measurement, take blood pressure). |  |  | March 10, 2016 | 2016 | Philips Visicu | The adherence check generates a software error.Two issues occur as a result of the software error: 1) No adherence flag is generated for that day; 2) No task is generated for the date 14 days later. An error message is generated and forwarded to operations. Subsequently, for days in which no task was generated, patient will not get a reminder on the tablet to take a measurement. If the pa | Under Investigation by firm | 1. Field Safety Notice(FSN) will be sent to all customers using eCareCoordinator v.1.3 platform. The FSN describes the problem and the circumstances in which the design defects occur and the action planned by Philips to correct the problem. The FSN will also detail the functionality change in detail.2. A software correction will be made and released. Customers will be upgraded to eCC v.1.4. The v.1.4 is scheduled to be released on or about 8 March 2016. The defect will have been corrected and tested. |  | 9 | 9 | US Distribution -- to the stated of FL, MA, MI, MS, PA, and KS. | Open, Classified |  |  | 9 | Computer
Z-1560-2016 | 73857 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion Lightning CT System iStation cover, SID: 70000246, Item Code: TSX-0354A/2D.03 |  |  | April 12, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems, Inc. is recalling the TSX-035A because of a potential problem with the Aquilion Lightning CT System iStation cover falling off of the gantry. | Device Design | An Urgent Medical Device Correction letter dated 4/12/16 was sent to a customer to bring to the customer's attention with a potential problem with the Aquilion Lightning CT System iStation cover. The letter informs the customer that the iStation cover may come off due to deficient dimensional geomtry. The letter informs the customer of the problem identified and corrective action to be taken. The customer is instructed to complete and return the attached form and fax it to (877) 349-3054. If the customer has any questions regarding the recall letter, then the customer is instructed to contact Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 1 | 1 unit | US in the state Nevada | Open, Classified |  |  | 1 | Not_Computer
Z-1563-2016 | 73117 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M Steri-Vac Gas Sterilizer/Aerator, models 5XL and 8XL. |  |  | January 22, 2016 | 2016 | 3M Company - Health Care Business | 3M discovered a valve with incorrectly labeled "input" and "output" orientation. This defect can result in the valve being installed backwards, creating a remote possibility of valve failure. A failure may result in EO being vented outside the sealed chamber during a limited period in the EO injection phase of the cycle. | Nonconforming Material/Component | 3M sent an Urgent Medical Device Field Correction letter to customers beginning 01/22/2016. The letter identified affected product, and stated that a 3M Service representative will contact the consignee to inspect the sterilizer. Questions can be directed to 3M Health Care Service Line at 1-800-228-3957. |  | 74 | 74 valves | AL, CO, CT, FL, GA, IN, NC, NJ, NY, OK, OR, PA, RI, VA, WI, and OUS to include: Argentina, Canada, Chile, China, Columbia, El Salvador, Germany, India, Indonesia, Japan, Korea, Mexico, Peru, Philippines, South Africa, Switzerland, Thailand, Trinidad, Turkey, United Arab Emirates, Vietnam. | Open, Classified |  |  | 74 | Not_Computer
Z-1567-2016 | 73660 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific Ultraflex Tracheobronchial Distal Release Stent System 12mm X 30mm |  |  | March 18, 2016 | 2016 | Boston Scientific Corporation | Mislabeled: 10mm diameter x 30mm length Ultraflex Tracheobronchial stents were labeled as 12mm diameter x 30mm | Process control | Boston Scientific issued notification letters dated March 17, 2016. The notification instructs the customer to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. Customers with questions should call 866-868-4004. For questions regarding this recall call 508-382-9555. |  | 4 | 4 units | Nationwide Distribution to AL, CA, MA, and TX. | Open, Classified |  |  | 4 | Not_Computer
Z-1566-2016 | 73674 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | PneumoSure XL High Flow Insufflator and PneumoSure High Flow Insufflator; |  |  | March 28, 2016 | 2016 | Stryker Endoscopy | When operating the device in high flow mode (40L/min, 15 mmHg) and connecting to a house gas inlet source, within seconds of insufflating the flow rate would drop to 0L/min and the red Check Gas Supply symbol would appear on the display, immediately suspending insufflation. | Nonconforming Material/Component | Urgent Medical Device Recall letters were sent to domestic customers on March 28, 2016 by certified mail. Stryker International distribution sites were notified and are responsible for providing information to affected customers. |  | 489 | 489 | Worldwide distribution. US nationwide including Puerto Rico, China, Hong Kong, India, Japan, Korea, Romania, Vietnam, Singapore, Taiwan, Netherlands, Peru, Australia, Canada, Mexico, Argentina, Chile, Poland, South Africa, and Guatemala. | Open, Classified |  |  | 489 | Not_Computer
Z-1564-2016 | 73677 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Magellan Robotic Catheter Accessory Kits, |  |  | March 25, 2016 | 2016 | Hansen Medical Inc | Firm's quarterly dose audit testing showed that one batch was radiated below the augmented dose and therefore did not meet the augmented dose of 28.4 kGy. | Under Investigation by firm | Hansen Medical sent an Urgent Notice Product Recall notification dated March 28, 2016, to all affected customers. The notification identified the product, the problem, and the action to be taken by the customer. Customers were instructed to provide the product to the Hansen Medical field representative to assure credit appropriately applied to their account. Customers with questions were instructed to contact their local Hansen Medical representative or call 650-404-5581. |  | 8 | 8 kits | Nationwide Distribution to OH, TX. FL, CA and GA. | Open, Classified |  |  | 8 | Not_Computer
Z-1565-2016 | 73435 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automated peritoneal dialysis (APD) cycler |  |  | February 16, 2016 | 2016 | Baxter Healthcare Corp. | Baxter Healthcare Corporation is sending this communication to inform you of incomplete instructions in the AMIA Automated Peritoneal Dialysis (PO) System Clinician Guide for calculating the total recommended solution therapy volume. Specifically, the instructions do not specify the need for an extra 200 mL of PD solution in order to prime the patient line and for air purge operations. The tot | Under Investigation by firm | An Urgent Device Correction communication was sent to the one affected customer via U.S.P.S., first class mail on 02/11/2016.A Clinician Guide Addendum was created to instruct clinicians of the extra 200 mL of PD solution required to prime the patient line and for air purge operations. This Addendum was provided as an attachment to the Urgent Device Correction communication letter. |  | 10 | 10 units | US Consignee: NY ***Foreign Consignee: Ecuador | Open, Classified |  |  | 10 | Not_Computer
Z-1570-2016 | 73748 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Trucount Absolute Counting Tubes: |  |  | March 31, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | One lot of lot of BD Trucount tubes may contain tubes from another lot. There is a 3.3% difference in bead counts between the two lots found within some of the pouches and presents a low risk to patient safety or impact to data if this product is used. | Packaging | BD Biosciences sent an Urgent Product Recall letter dated March 2016 to all affected customers on March 31, 2016 by Certified Mail, e-mail and direct delivery by Sales Associates.Letters instructed that any remaining inventory of the affected lot be discarded and that date files obtained using reagents from the affected lot be reviewed for the possibility of incorrect data.Customers with questions were instructed to contact BD Customer Support at 877-232-8995 (prompt3, 3) in the united State. For customers outside the US, customers should contact their local BD Biosciences representative or distributor.For questions regarding this recall call 408-954-6080. |  | 1743 | 1743; 467 US and 1276 International | Worldwide Distribution - US including MD, CA GA. WA. PA. CO. TN. FL. NC. DE. VA, CT, LA. NC, NY, OH, MN, NH and Internationally to China, Japan, and India. | Open, Classified |  |  | 1743 | Not_Computer
Z-1568-2016 | 73736 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler 45 and Stapler 30 instruments used on the da Vinci Xi systems (IS4000) with p5 software. |  |  | March 25, 2016 | 2016 | Intuitive Surgical, Inc. | Potential for unexpected motion of the Xi Stapler jaws on the da Vinci Xi System with p5 software relating to a combination of the p5 software and the surgeon quickly transitioning from the clamp to the fire pedal during use. | Software design | Intuitive Surgical sent an Urgent Product Safety notification to customers via e-mail on March 25, 2016. Customers were advised to STOP USE on the EndoWrist Stapler 45 and 30 instruments on da Vinci XI systems with p5 software. Product letters followed and were mailed on March 31, 2016 by Fed Ex.Urgent Medical Device letters were sent via Fed Ex on March 31, 2016 (dated April 1, 2016). These letters reiterated what was explained in the previous e-mail and provide more information of the unexpected motion issue. In addition to the stop use request, letters requested that all personnel be fully informed of the issue.Further information or support concerning this issue, please contact your Clinical Sales Representative or contact ISI Customer Service at the number:North America: 800-876-1310 Option 3 (6 am to 5pm PST) |  | 43 | 43 | US Nationwide Distribution in the states of: WI, FL, OH, TX, CA, AZ, IL, MD, MN, NY, OK, HI, MO, IA, KS, PA, and AL. | Open, Classified |  |  | 43 | Computer
Z-1571-2016 | 73688 | No | JSJ | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Hardy Diagnostics GBS Detect |  |  | March 21, 2016 | 2016 | Hardy Diagnostics | Performance failure; poor hemolytic reaction of non or weakly hemolytic group B streptococcus (GBS) with the target QC organism, Streptococcus agalactiae (ATCC 13813) due to product deterioration. | Under Investigation by firm | On 03/21/2016 customers were notified by phone, and were sent a notification letter.Customers were instructed via phone call to discard remaining plates of the implicated lot, and to notify the firm of how many plates are left. The firm stated that they will be issuing replacements. The firm stated that they will be sending written notification as well. The firm requested full name and e-mail or fax number in order to follow-up with the written notification.The customer notification letter instructs customers to complete, sign, scan/e-mail or fax back enclosed form stating compliance with the action regarding the specific lot. The firm requests the form be returned to the attention of the Quality Assurance Department at techservice@hardydiagnostics.com, fax number 805-614-9274. If the lab is unable to be reached via telephone and would like replacements or credit contact Customer Service Department at 800-266-2222, option 1. Replacements will be sent out at no charge.Any questions contact Technical Services Department at 800-266-2222, option 2 or via e-mail at techservice@hardydiagnostics.com. |  | 3530 | 3,530 (353 pk/10) | U.S. Distribution to the following states: AL, SC, FL, CT, VA, WI, MA, MI, WV, OH, IN, IL, NC, and TX. | Open, Classified |  |  | 3530 | Not_Computer
Z-1569-2016 | 73616 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System configured with Thermo Scientific Recapper Module (230V or 110V -- COMMON/USUAL NAME: enGen (TM) Track System -- NOTE: Outside the US the system is also known as VITROS Automation Solution (VAS) TC Automation System (configured with Thermo Scientific Recapper Module from VITROS Automation Solutions), Product Code 6844300 --- The enGen (TM) Lab Automation System is a modification to VITROS Analyzers that includes the addition of an automated track sub-system (TC Automation) for pre-analytical and post-analytical sample and data management. In addition to an Automated Track Sub-system, the enGen (TM) System includes one or more VITROS Chemistry, Immunodiagnostic and/or Integrated Systems (clinical analyzers) as well as a Data Management Subsystem (Middleware), and a Software Configuration (gsb) File. The enGen System is typically interfaced with a Laboratory Information System (LIS) via connectivity facilitated at the Data Management Subsystem level. The enGen (TM) Laboratory Automation System contains a module (i.e. enGen (TM)Laboratory Automation Recapper Module) that applies caps to the sample collection container after the sample has been aspirated. In this module, an electrical cable supplies power to the cap vibratory feeder and is supposed to be located above the vibratory feeder support plate. The cap vibratory feeder sits upon a feeder support plate. Un-capped sample containers pass underneath this support plate prior to being re-capped for storage. The Recapper Module is supplied by Thermo Fisher Scientific. |  |  | March 17, 2016 | 2016 | Ortho-Clinical Diagnostics | The electrical cable that supplies power to the cap vibratory feeder of the Recapper Module may loosen from its correct position above the vibration feeder support plate and migrate to under the support plate. If this occurs, the uncapped sample tubes may hit the cable, and sample fluid (e.g., serum) may splash from the tube and into another tube or onto the under carriage of the support plate. | Component design/selection | On 3/17/2016 URGENT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2016-074, dated 3/17/2016)) was sent notifying the users of the issue and advise them of the required actions. The notification was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. Foreign affiliates were informed by email on 3/17/2016 of the issue and instructed to notify their consignees of the issue and required actions. |  | 49 | US: 19, Foreign: 30 | Distributed in the states of MD, NY, ID, NC, CO, OH, PA, TN, MS, MO, CA, DE, FL, GA, AZ, NJ, and WV and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, France, India, Italy, Norway, Portugal, Saudi Arabia, Spain, Sweden, Thailand, and United Kingdom. | Open, Classified |  |  | 49 | Not_Computer
Z-1578-2016 | 73680 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hand-Held Scanner USB IT3800 |  |  | March 03, 2016 | 2016 | Roche Molecular Systems, Inc. | The hand-held barcode scanner model IT3800 used with the COBAS AmpliPrep instrument mis-identified a sample barcode ID. | Labeling design | Roche Molecular Systems sent an "Urgent Medical Device Correction (UMDC) notification letter dated March 3, 2016, to their affected customers. Customers were instructed to provide the UMDC to other affected parties who may need to be aware of this issue. Customers were asked to complete the attached fax form and fax it to 1-866-737-5396. Customers with questions should contact Roche Support Network Customer Support Center at 1-800-526-1247. For questions regarding this recall call 908-253-7200. |  | 6939 | 6,939 pieces | Nationwide Distribution | Terminated | May 25, 2017 | 448 | 6939 | Not_Computer
Z-1575-2016 | 73693 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network |  |  | March 09, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Panorama Central Station including the work station View Station, View Station and the eGateway will revert to the year 2000 when Daylight Savings Time (DST) occurs on March 13, 2016, or under various conditions. | Unknown/Undetermined by firm | Mindray DS USA Inc. issued a Medical Device Service Notice dated March 9, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers requiring technical assistance were instructed to contact Mindray's Technical Support team at 1-800-288-2121. Mindray will address the systems response to DST in a future software update which will be provided to the customer at no cost. |  | 1492 | 1447 units US, 45 units OUS | Nationwide Distribution | Open, Classified |  |  | 1492 | Computer
Z-1573-2016 | 73741 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System PN: 989605858501 |  |  | March 24, 2016 | 2016 | Philips and Neusoft Medical Systems Co., Ltd. | The following issues are found in NeuViz 16 systems with software version 1.1.4.21425 and version 1.1.4.21426: 1) During the filming operation on MX 16-slice console software, the clipboard used for copying and pasting images is not cleared between patients. If the operator fails to copy the current patient's images before pasting, a previous patient's image may be present in the clipboard and be | Under Investigation by firm | Affected consignees will be notified of the recall via letter on March 24, 2016. Mandatory Field Change Order and FCO kits will be released to all affected systems in June, 2016. |  | 37 | 37 units | NC, OH, NE, SC, TX, LA, PR, MO, FL, CT | Open, Classified |  |  | 37 | Computer
Z-1579-2016 | 73747 | No | NUA | Cystic fibrosis transmembrane conductance regulator (CFTR) gene mutation detection system. | Immunology | Immunology | 510(k) | CF InPlex ASR Card; Catalog Number: 95-0501. |  |  | March 31, 2016 | 2016 | Hologic, Inc | Complaints of false positive HET mutations and leaking cards; could lead to unnecessary invasive testing during pregnancy and the associated risks of such invasive testing. | Nonconforming Material/Component | Customer notification letters were sent 03/31/16. Customers were instructed to immediately discontinue the use of the affected lot numbers.If a HET or MUT result was reported with the affected lots, Hologic recommends that customers notify the recipient of the result and request the patient be retested with another test method.Any HET or MUT results that were previously retested and reported out as negative do not require further testing.The firm requests that customers segregate any remaining inventory of all affected lots listed on the Customer Response Form. After segregation is complete reconcile the number of units (ASR Card or IVD Molecular Test Kit) which remain in inventory upon receipt of notification against the number originally shipped to site, and destroy all remaining inventory of the affected lots.Document the following in the appropriate fields on the Customer Response Form:1. If have zero of the affected products on hand check the first box.2. If have some inventory of the affected products on hand check the second box and then document the quantity on hand below, in the Qty On-Hand column.Note: When counting the number of units remaining in inventory count each unit (ASR Card box or IVD Molecular Test Kit), whether partial or full, as a unit.3. Print name, phone and email address, in case the firm needs to follow up.4. Sign and date the form. Signing acknowledges compliance with this recall notification, including destroying all remaining inventory of the affected products.When this documentation is complete fax the CRF to Hologic Technical Support using fax number: 858-410-8250. Alternatively, you can scan the document(s) and email to Hologic Technical Support at molecularsupport@hologic.com.If laboratory has shared any of the affected cards noted in the CRF with another laboratory contact Hologic Technical Support immediately using the phone number below, so that Hologic may contact the other laboratory t |  | 2580 | 2580 total | Distributed in the states of OH, AL, IL, AZ, AR, NJ, TX, GA, CO, DE, NC, MI, FL, ME, CA, IN, OR, KS, KY, MO, WA, NY, OK, WV, CT, IA, LA, RI, MA, and PA, and in the country of Canada. | Open, Classified |  |  | 2580 | Not_Computer
Z-1576-2016 | 73650 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Drew Scientific Inc. Drew3-PAC. This D3 Reagent Pack includes: diluent reagent 5L, enzymatic cleaner 750 mL, and a lysing reagent 125 mL. |  |  | March 10, 2016 | 2016 | Jas Diagnostics/Drew Scientific | The reagent pack may result in elevated platelet background during Drew3 Hematology Analyzer routine check. | Process control | JAS Diagnostics/ Drew Scientifics sent an Important Product Advisory letter dated March 10, 2016, to all affected customers. The domestic customers were notified via letter and international customers were notified via email. Customers were instructed to contact Drew Scientific Inc. Customer Service or Technical Service at 1-800-433-0945 for product replacement. The firm suggested a temporary workaround of performing five background runs three times per day in order to ensure background counts are running within specification. |  | 7691 | 7,691 | Worldwide Distribution - US Distribution including the states of : AL, AR, CA, FL, GA, IA, IN, KS, KY, LA, MA, MN, MO, MS, NC, NE, NY, SC, TN, and TX. and to the countries of : Bolivia, Cairo, Columbia, Ecuador, Italy, Moscow, Netherlands, Pakistan, Peru, Slovenia, Spain, and Venezuela. | Open, Classified |  |  | 7691 | Not_Computer
Z-1574-2016 | 73770 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 5.0 mm Periprosthetic Locking Screw Self-tapping 12 mm Sterile; |  |  | April 04, 2016 | 2016 | Synthes (USA) Products LLC | This lot of the 5.0 mm Periprosthetic Locking Screw, Self-tapping, 12mm Sterile, was shipped after the expiration date of May 31, 2012. The label on the product has the correct expiration date. | Process control | DePuy Synthes mailed to customers an Urgent Notice informing them they were voluntarily recalling 5.0 mm Periprosthetic Locking Screw Self-tapping 12 mm-Sterile, part # 02.221.462S, Lot # 5539094, expiration date May 31, 2012, due to the units being shipped after the expiration date of May 31, 2012. Customers were instructed to call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verification Section on page 3 by checking the appropriate box and return the form and product to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. |  | 15 | 15 units | Distributed in IL and IN. | Open, Classified |  |  | 15 | Not_Computer
Z-1572-2016 | 73886 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON X700 Ultrasound System with software version 1.0.04. |  |  | September 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Due to a communication error between the software and V5Ms transducers rotation function, ACUSON X700 ultrasound systems at software versions 1.1.04, display an IMG_15 error message causes the system to lock up, which requires a reboot to recover system operation. | Software design | September 5, 2015 we initiated a correction via update program US020/15/S where we visited all four units and updated their software to version 1.0.05 |  | 4 | 4 | Distributed in Germany, U.A.E., Hungary and Brazil. | Terminated | April 27, 2016 | 235 | 4 | Computer
Z-1581-2016 | 73779 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | reVive Light Therapy Pain Relief Device (Model Number: RVPNSYS) |  |  | March 30, 2016 | 2016 | LED Technologies, LLC | Incorrect Power Supply in the package, rendering the device inoperable. | Nonconforming Material/Component | Customers were notified on March 30, 2016 via telephone call. LED Technologies, Inc. provided the consignee with a letter to distribute to customers who purchased the reVive Light Therapy Pain Relief device, and included instructions to receive a replacement of the correct power supply. |  | 32 | 32 units | Distributed to 1 consignee located in the United States. | Open, Classified |  |  | 32 | Not_Computer
Z-1583-2016 | 73512 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | RESTORIS PST RIO Offset Shell Impactor |  |  | February 23, 2016 | 2016 | Mako Surgical Corporation | Potential disassociation of the orientation pin. | Nonconforming Material/Component | On March 4, 2016 Stryker Orthopaedics initiated a voluntary product recall of the RESTORIS PST RIO Offset Shell Impactor informing the customer in an Urgent Product Recall to do the following: 1. Inspect all metal instruments for any damage or difficulty disassembling/assembling instruments2. Return all affected products at their location to Stryker C/O Stericycle, 2670 Executive Dr. Suite A, Indianapolis, IN, 46241, Attn of Event 8763, Clayton Odor.3. Complete and sign the Business Reply Form and send it back via e-mail or fax within five business days.Contact 866-241-0741 or via e-mail strykerortho8763@stericycle.com |  | 96 | 96 | Nationwide Distribution -- AL, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MN, MS, NC, NH, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, WA, and WI. | Open, Classified |  |  | 96 | Not_Computer
Z-1582-2016 | 73676 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Pleur-evac Chest Drainage System, Dry Suction/Dry Seal Control, REF A-6000-08LF, Rx Only, Teleflex Medical. |  |  | March 30, 2016 | 2016 | Teleflex Medical | The label on the Tyvek bag is missing. This is used by customers to identify material code, lot number, and expiration date once the product is removed from the shipper box. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification dated March 30, 2016 to customers and distributors. The letter identified the affected product, problem and actions to be taken. The letter requested that they discontinue use and quarantine any products. The letter instructed customers to complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com.. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 12124 | 12,124 ea. | Worldwide Distribution - US Nationwide in the states of: AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, LA, MD, MA, MI, MS, MO. MT, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WV, Puerto Rico, Belgium, Canada, Chile, Dominican Republic &amp; Germany | Open, Classified |  |  | 12124 | Not_Computer
Z-1585-2016 | 73745 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | X-Ray Diagnostic Table, Model No. YSF-300 |  |  | March 31, 2016 | 2016 | Shimadzu Medical Systems | Report of unintentional movement of table to reverse tilting. | Device Design | Shimadzu sent an Urgent Voluntary Medical Device Recall Notice to all affected customers on April 8, 2016. Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray Diagnostic Table. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contacted Shimadzu National Technical Support at (800) 228-1429.Shimadzu will send out an updated customer notification letter to inform customers that the liquid penetration may cause an electrical short circuit, causing the table to unintentionally tilt in reverse. If this unintentional reverse tilting occurs, there is a potential risk of injury to the patient and/or the user.For questions regarding this recall call 310-217-8855, ext 139. |  | 185 | 185 units | Nationwide Distribution | Terminated | January 02, 2017 | 277 | 185 | Not_Computer
Z-1584-2016 | 73581 | No | MQJ | Assisted reproduction micromanipulators and microinjectors. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Integra 3; Model Numbers: 6-54-100, 6-54-110, 6-54-150, 6-54-160 |  |  | March 11, 2016 | 2016 | RESEARCH INSTRUMENTS LTD - SITE B | The units have been fitted with a potentially faulty power supply cable. | Device Design | Research Instruments sent a recall notification dated 03/11/2016, via UPS with confirmed delivery receipt, to all affected customers. The letter identified the affected product, problem and the action needed to be taken by the customer. The letter asked customers to contact Research Instruments First Line Support if they had any questions. |  | 380 | 380 units | Worldwide Distribution - US Nationwide in the states of: MA, NY and countries of : Algeria, Australia, Austria, Bahrain, Belgium, British Virgin Islands, Bulgaria, Canada, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hong Kong, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Nigeria, Norway, Pakistan, Peru, Poland, Romania, Serbia, Singapore, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Taiwan, Thailand, Tunisia, Turkey, UAE, Ukraine, United Kingdom, and Vietnam, | Open, Classified |  |  | 380 | Not_Computer
Z-1587-2016 | 73486 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Livongo Health In TouchÂż Blood Glucose Monitoring System (BGMS), is an Over-The-Counter (OTC) system designed for the self- monitoring of blood glucose by persons with diabetes in home settings to aid in their diabetes management. |  |  | February 18, 2016 | 2016 | Livongo Health | The instructional insert for the In Touch Control Solution for use with the In Touch meter and In Touch strips as a quality control check did not contain the limitation of sensor accuracy: (1) with the use of acetaminophen (Tylenol)(2) with the use of Ascorbic Acid (vitamin C) supplementation(3) in the presence of uric acid | Error in labeling | The recall was conducted in three distinct events. Members were first sent an e-mail of the Recall Notification - Livongo Test Strip Insert (Event 1) containing information regarding the test strip insert discrepancy. Following the e-mail distribution, a hard copy of the same information was distributed (Event 2) to members via United States Postal\ Service (USPS) mail. Lastly, a second e-mail (Event 3) containing the same information was sent to members. The distribution list for the second (hardcopy distribution) and third (e-mail distribution) events was a subset of the initial member distribution. |  | 63000 | 63,000 units | US consignees: Arizona, California, Delaware, Florida, Georgia, Illinois, Kentucky, Maryland, Michigan, Minnesota, Mississippi, Missouri, Montana, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Virginia and Washington. | Open, Classified |  |  | 63000 | Not_Computer
Z-1586-2016 | 73433 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AnalystÂż MD Version 1.6.1 and 1.6.2 Software used with the following instruments: API 3200MD" LC/MS/MS System, Instrument Part Number: 5024501; 3200MD QTRAPÂż LC/MS/MS System, Instrument Part Number: 5024500; Triple Quad" 4500MD LC/MS/MS System, Instrument Part Number: 5031257; QTRAPÂż 4500MD LC/MS/MS System, Instrument Part Number: 5031231 |  |  | February 24, 2016 | 2016 | Ab Sciex | Wrong quantitative results may be displayed in a report from the device, which may potentially lead to an incorrect patient diagnosis. | Software design | Sciex sent an Urgent Medical Device Correction letter dated February 29, 2016, to all affected US customers via Fedex on March 2, 2016. International customers were notified beginning on March 3, 2016. The letter informed customers of the issue and requests that they take temporary steps to eliminate the potential for erroneous events. In addition, a software update is in development and is expected to be available in approximately 4 weeks. Customers with questions were instructed to contact SCIEX at 1-289-982-2531. |  | 279 | 279 | Worldwide Distribution - US (Nationwide) and Internationally to the following countries:AUSTRIACANADACHINAFRANCEGERMANYINDIAIRELANDITALYKOREAMALAYSIANETHERLANDSNORWAYRUSSIASINGAPORESPAINSWITZERLANDUNITED KINGDOM | Open, Classified |  |  | 279 | Computer
Z-1589-2016 | 73974 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE, Revolution CT. |  |  | April 20, 2016 | 2016 | GE Medical Systems, LLC | An out of tolerance torque tool was used on three joints during the manufacturing process. If the X-ray Tube becomes loose on the CT gantry during operation, this could result in beam tracking errors, potential scan aborts and/or a possible image artifact. There is a secondary safety mechanism to prevent components from being expelled from the gantry if the bolts fail. | Process control | GE Healthcare Systems sent an Urgent Medical Device Correction letter dated April 20, 2016, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 26 | 26 (US = 4;OUS = 22) | Worldwide Distribution - US including FL, GA, NC, and TX and Internationally to AUSTRALIA, CANADA, CHINA, FRANCE, GERMANY, ITALY, JAPAN, REPUBLIC OF KOREA, NORWAY, PERU, SWEDEN UNITED ARAB EMIRATES, UNITED KINGDOM. | Open, Classified |  |  | 26 | Not_Computer
Z-1588-2016 | 73199 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | MAKOplasty Hip Array Kit Laminate 150230 |  |  | November 25, 2015 | 2015 | Mako Surgical Corporation | Stryker Orthpaedics became aware that there is the potential for the 210540 MAKOplasty Hip Array Kit Laminate 150230 to list two part numbers for the RIO Base Array, which is item #5 on the product table. | Component change control | On 12/09/2015, Product Field Action Letters and Product Accountability forms were sent to Branch/Agencies via overnight courier with a return receipt; Hospital Risk Management were sent Notification letters and a Product Recall Acknowledgement form with return receipt on 12/9;/2015. |  | 10 | 10 units. | CO, KY, MN, WA and WV | Open, Classified |  |  | 10 | Not_Computer
Z-1591-2016 | 73862 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Pentaray NAV High-Density Mapping ECO Catheter, Catalog No. D128207, D128208, D128209, D128210, D128211, D128212 |  |  | April 14, 2016 | 2016 | Biosense Webster, Inc. | Biosense Webster is recalling the Pentaray Catheters to clarify the contraindication language in the Instructions For Use (IFU) and product labeling for this catheter relative to patients with prosthetic valves. | Use error | The firm, Biosense Webster, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 4/14/2016 to its customers. The letter described the product, problem and actions to be taken. The letter informs customers that the recall action is NOT a product removal, and that they may continue to use PENTARAY Catheters in accordance with the updated contraindication as stated in the recall letter. Customers are instructed to read and follow the Labeling Correction Notice carefully; pass on this letter to anyone in your facility that needs to be informed; maintain a copy of this letter with the PENTARAY Catheters, and sign and return the attached ACKNOWLEDGEMENT FORM immediately by email: BWI_FA_Coordinator@its.jnj.com, fax: 909-839-8514 or mail: Biosense Webster, Inc., Attn: FCA Coordinator,15715 Arrow Highway, Irwindale, CA 91706. The updated the IFU to clarify the contraindication statement as follows: "Do not use PENTARAY Catheters in patients with prosthetic valves". For questions related to this medical device correction, please contact your BWI sales representative or call 909-839-8500. |  | 59672 | 59,672 units | Worldwide Distribution: US (Nationwide) and countries of: Canada, China, Hong Kong, Indonesia, India, Japan, Korea, Malaysia, New Zealand, Singapore, Thailand, Taiwan, United Arab Emirates, Austria, Belgium, Bulgaria, Switzerland, Czech Republic, Germany, Denmark, Algeria, Spain, Finland, France, United Kingdom, Greece, Hungary, Ireland, Israel, Iraq, Iceland, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Sweden, Turkey, South Africa. | Open, Classified |  |  | 73584 | Not_Computer
Z-1502-2016 | 73634 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The Spacelabs Healthcare Xhibit Telemetry Receiver (XTR), Model 96280 is used to transmit patient(s) physiological measurements and data to a monitor. |  |  | March 28, 2016 | 2016 | Spacelabs Healthcare Inc | Reports of an abnormally high frequency of squelch events on Xhibit Telemetry Receivers (XTR), which is indicative of the loss of the telemetry data signal. In some cases, the Xhibit Central Monitor was also reported to display an Xhibit Offline error message for all patients being monitored. | Other | On March 28, 2016, a customer letter will be sent via priority service, return receipt requested, to all U.S. customers. On April 4, 2016, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record. Spacelabs recommends that you brief your staff regarding the failure modes listed above. If you continue to use affected Telemetry Receiver(s), please note:**When there is a loss of signal, a signal loss alarm is generated and a triangular-shaped squelch waveform is displayed for any channel exhibiting signal loss. If unacknowledged, this alarm will escalate to medium priority after five minutes. After an additional five minutes, this alarm will escalate to high priority. This failure is usually transient.**When the XTR goes offline, all telemetry waveform(s) associated with the XTR will disappear from the Central Station and an offline message will appear in each telemetry waveform zone. This failure is not recoverable. Please circulate this notification to all persons affected by the information it contains and add a copy to the appropriate monitor user manual until such time as corrections have been implemented by Spacelabs. Spacelabs will contact you to schedule a convenient time for Spacelabs to update your Xhibit Telemetry Receivers at no cost. If you have any questions about this corrective action, please contact Spacelabs at 1-800-522-7025 and select 2 for Technical Support. |  | 115227 | 217 units of XTRs and 531 units of PCBAs (164 XTRs and 425 PCBAs in US and 53 XTRs and 106 PCBAs outside the US) | distributed in AL, AZ, CA, CO, CT, FL, ID, IN, KY, MI, MN, MO, MS, MT, SC, TX, WY and in the following countries: AUSTRALIA, CANADA, CHINA, FRANCE, NETHERLANDS, PHILIPPINES, ROMANIA, and UNITED KINGDOM. | Open, Classified |  |  | 115227 | Not_Computer
Z-1498-2016 | 73612 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | iVue and iVue 500 (now called iScan) OCT systems with software version 2016.0.0.127 (USA only) and 2016.1.0.127 (out of US) |  |  | March 16, 2016 | 2016 | Optovue, Inc. | A software anomaly was detected that affects Optovue iVue and iVue 500 with Normative Database that may display images of the same eye rather than right and left if user scrolls on the map. | Under Investigation by firm | Recall Letters were sent to customers on March 22, 2016 to inform them of the recall. |  | 40 | 40 | PA, CA, IL, HI. MO, CO, GA, ME, IN, AZ, WV, VA, FL, AR, UT, NE, IN, NV, LA, TN, KS, OK. Foreign distribution forthcoming | Terminated | April 04, 2017 | 384 | 40 | Computer
Z-1595-2016 | 73601 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TCM CombiM |  |  | November 18, 2015 | 2015 | Radiometer America Inc | Radiometer America, Inc. is recalling TCM CombiM on rare occasions they may have an error that can cause the alarm to sound higher than the pre-set maximum volume. | Software design | RadioMeter sent an Product Notification letter out on November 18, 2015, January 10, 2016 and March 29, 2016.The notification sent on November 18, 2015 and January 10, 2016, instructed customers a temporary solution of switching off the analyzer and then on again in order to re-establish the pre-set alarm volume. If alarm sound is higher than the pre-set volume customers are instructed to contact their local Radiometer service representative. If any questions contact Radiometer America Technical Support at 1-800-736--600 opt. 4.The notification sent on March 29, 2016, provided the same instructions and temporary solution. The notification also informed customers that the firm will be in contact with them to discuss the upgrade process. |  | 1488 | 1488 | Worldwide Distribution - USA (nationwide) Distribution to the states of : MO, NE, AL, PA, IL, MN, TX, ID, CO, UT, NV, CA, NY, RI, OR, NJ, AZ, GA, IN, VA, LA, NM, MD and AR., and to the countries of : Canada, Austria, Czech Republic, Denmark, Finland, France, Germany, Hungary, Ireland, United Kingdom, Italy, Netherlands, Norway, Spain, Sweden, United Arab Emirates, Kuwait, Qatar, Saudi Arabia, Oman, Jordan, Croatia, Malta, Iceland, Brazil, Thailand, Philippines, Peru, Honduras, Mexico, Kenya, Israel, Romania, Russia, Poland, Turkey, Australia, New Zealand, China, Hong Kong, South Africa, India, Korea, Singapore and Taiwan. | Open, Classified |  |  | 1488 | Not_Computer
Z-1592-2016 | 73854 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound imaging system with software version VB10C and using transesophageal (TEE) transducer Z6Ms, V5Ms or V7M. |  |  | April 08, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | While imaging with a transesophageal (TEE) transducer (Z6Ms, V5Ms, or V7M), user may lose the ability to control the color region of interest, the pulsed wave or continuous wave Doppler gate or cursor, the M-mode cursor, the 2D field of view, or the RES region of interest with the trackball. | Software design | Under Siemens update program US013/16/S a Customer Advisory Notification will be mailed to all affected customers by Certified mail with return receipt beginning March 30, 2016. Letters provide customers information on how to avoid the issue until the software patch is installed. |  | 87 | 87 | Worldwide distribution. US nationwide, Austria, Canada, China, Germany, India, Italy, Spain, Sweden, and Switzerland. | Open, Classified |  |  | 87 | Computer
Z-1594-2016 | 73742 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CARESTREAM Image Suite V4; Image Suite V4: DICOM STORE SCP: REF/Catalog # 1056191, DICOM STORE SCP/FOR IMAGE SUITE V4: REF/Catalog # 6566988; Image Suite V4 Bundles: IMAGESUITE STANDALONE PACS: REF/Catalog # 1741289, IMAGESUITE STANDALONE PACS FOR INDIA: REF/Catalog # 1741297; Carestream PRO (Image Suite V4): PRO Medical Wireless GOS System-Desktop: REF/Catalog # 1741891, PRO Wireless System Laptop: REF/Catalog # 1741925, PRO Wireless System - w/o Computer: REF/Catalog # 1741933, PRO Tethered System Desktop: REF/Catalog # 1741941, PRO Tethered System Laptop: REF/Catalog # 1741958, PRO Tether System - w/o Computer: REF/Catalog # 1741966, PRO Fixed System - w/o Computer: REF/Catalog # 1741974, PRO Medical Wireless CsI System-Desktop: REF/Catalog # 1741982, PRO Wireless System Laptop: REF/Catalog # 1742006, PRO Wireless System - w/o Computer: REF/Catalog # 1742014, PRO Tethered System Desktop: REF/Catalog # 1742022, PRO Tethered System Laptop: REF/Catalog # 1742055, PRO Tether System - w/o Computer: Catalog # 1742063; PRO Fixed System - w/o Computer: REF/Catalog # 1742089 -- Made in U.S.A. by: Carestream Health, Inc., 150 Verona Street, Rochester, NY 14608 |  |  | April 15, 2016 | 2016 | Carestream Health Inc | Carestream Health received a complaint related to CARESTREAM Image Suite 4 from a foreign hospital stating that the annotation on the overlay is displayed as "L (Left)", when it should be "R(Right)". | Software design | Carestream Health sent an URGENT MEDICAL DEVICE CORRECTION letter dated April 6, 2016, and Consignee Notification Acknowledgement form to the US consignees on April 15, 2016, via Certified Mail, Return Receipt Requested. Foreign consignees are being notified based on the individual country regulations. Customers who have this Image Suite configuration but do not use it for CT or MR Imaging do not need to take any action. Customers who use the product for CT or MR Images are asked to make users aware of the problem and take additional steps to check correct orientation of any markers, until the software is updated. Customers with questions or concerns were instructed to contact the Carestream Customer Care Center in the US at 1-800-328-2910. Customers outside the US were instructed to contact their local Service Support number.For questions regarding this recall call 1-800-328-2917. |  | 306 | US: 37 units, Foreign: 269 units | NY and OH for further distribution nationwide; Foreign distribution to the following countries: Afghanistan, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia &amp; Herzgovena, Brazil, Bulgaria, Cambodia, Canada, Chad, Chile, China, Colombia, Croatia, Czech Rep., Djibouti, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Italy, Jamaica, Japan, Jordan, Kuwait, Kyrgyzstan, Laos, Latvia, Lebanon, Liechtenstein, Lithuania, Macao, Malaysia, Mexico, Montenegro, Myanmar (Burma), Nepal, Netherlands, New Zealand, Niger, Norway, Oman, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Scotland, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Taiwan, thailand, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Vietnam. | Terminated | March 16, 2017 | 335 | 2855 | Computer
Z-1600-2016 | 73884 | No | LGW |  |  | Neurology | PMA | Nuvectra, Algovita Spinal Cord Stimulation System Clinical Programmer, Model 4500. For use with Algovita Spinal Cord Stimulation Systems. The version of the software in the Clinician Programmer is v1.1.5. |  |  | April 06, 2016 | 2016 | Nuvectra | If used, the optional Swap feature copies programming parameters from the external pulse generator (EPG) as well as unintended calibration data to the implantable pulse generator (lPG). This may result in the user being unable to recharge the lPG.This safety notice only applies to CPs Model 4500. | Software design | Nuvectra sent an "Urgent Field Safety Notice" letter dated April 6, 2016 to customers. The letter identified the affected product, problem, recommendations and Mitigations, and actions to be taken. |  | 17 | 17 OUS | Germany | Open, Classified |  |  | 17 | Not_Computer
Z-1597-2016 | 73767 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | INFINITI Ultrasound Pack. Catalog Numbers: 8065751716, 8065751717, 8065751718, 8065751720, 8065751721, 8065751723. |  |  | October 06, 2014 | 2014 | Alcon Research, Ltd. | The recalling firm received complaints related to leakage from the tubing near the white luer of the irrigation line for the INFINITI Ultrasound Packs. | Process change control | The recalling firm notified affected customers on 10/6/2014 via a customer letter with an attached Medical Device Removal Response Form. On 10/23/2014, a second notification was sent to 76 customers who did not respond to the initial notification. |  | 6597 | 6597 packs | Nationwide Distribution. | Terminated | March 31, 2017 | 907 | 6597 | Not_Computer
Z-1596-2016 | 73585 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Positive Susceptibility Test Cards (AST-P605), REF 22 325. 20 Cards/Kit. bioMerieux, Inc, Durham, North Carolina 27712. |  |  | February 24, 2016 | 2016 | Biomerieux Inc | Product may lead to false susceptible results of Staphylococcus spp., Enterococcus spp., and S. agalactiae to Minocycline. | Other | Biomerieux, Inc. sent an Urgent Notice Product Correction on February 24, 2016, to all affected customers. The recalling firm notified their foreign subsidiary of the Field Safety Corrective Action (FSCA) by email. The local subsidiary in the country impacted has informed users impacted and provided instructions to resolve the issue.Customers were instructed to confirm the Field Safety Corrective Action letter has been distributed and reviewed by all appropriate personnel within theirorganization.Âż For the referenced test kit (VITEKÂż 2 ASTP605, Lot 485368220), refrain from reporting Minocycline (MNO) results for any organism until a new lot is implemented in your laboratory.Âżo Please note that a bioART rule can be created to suppress the analysis/reporting of Minocycline (reference VITEKÂż 2 Online Software User Manual). The bioART rule can be disabled upon implementation of a new ASTP605 lot number.Customers were instructed tostore the letter with their bioMÂżrieux instrument documentation.Âż Complete the attached Acknowledgement Form and return it to their local bioMÂżrieux representative.For questions regarding this recall call 314-731-8526. |  | 6740 | 337 Kits of 20 Cards/Kit | Worldwide Distribution to China  Taiwan region; No US/military/government distribution. | Terminated | May 26, 2017 | 457 | 6740 | Computer
Z-1599-2016 | 73668 | No | IXW | Automatic radiographic film processor. | Radiology | Radiology | 510(K) Exempt | Smart CR Digital Computed Radiography System; Model Number: CR-IR346RU |  |  | April 01, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Reports of failure of the device's power supply. | Device Design | Fujifilm sent an Medical Device Correction letter dated March 22, 2016, via FedEx to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Instructions for turning the unit off when not in use as an interim action. FMSU will conduct a field replacement of the power supply in all installed units.For further questions, please call (888) 385-4633. |  | 2062 | 2,062 devices | Nationwide Distribution. | Open, Classified |  |  | 2062 | Not_Computer
Z-1598-2016 | 73947 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera/Skyra Systems, whole body magnetic resonance imaging systems, Model numbers 10432914, 10432915. Docking station top cover (part no# 10591430) of the removable patient table. |  |  | April 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The docking station top cover of the removable patient table may contain a magnetic metal plate. When docking/undocking the table, the magnetic top cover may become attracted to the magnet. | Vendor change control | Siemens mailed to customers on April 13, 2016, a Customer Safety Advisory Notice, MR013/16/S, to inform them of the actions to take to prevent a possible hazard to persons and equipment. Customers were informed that they will be contacted by one of Siemens' service organization who will schedule an appointment to perform a correction, which is to exchange the top cover. |  | 8 | 8 systems | US nationwide distribution to OH, FL, TN, MN, OH, NY, VA, and MN. | Terminated | April 17, 2017 | 369 | 8 | Not_Computer
Z-1536-2016 | 73744 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Dialog+ Hemodialysis machines: Conductivity Sensor (BIC) Item 3456103A |  |  | April 01, 2016 | 2016 | B Braun Medical Inc | Potential leakage of conductivity sensors of the Dialog Dialysis machines. | Process control | B. Braun mailed an Urgent Medical Device Correction notice to affected customers informing them of the Dialog+Hemodialysis System due to the potential for Bicarbonate and End conductivity cells (3456102A and 3456103A respectively) to exhibit cracks. Notice provided customers with actions for Qualified Service Technicians, Clinical Personnel Administering Dialysis treatment and adverse event reporting. All Dialog+Hemodialysis machines will be inspected at the customer's site by the customer Biomed, who has been trained and qualified by B. Braun Avitum. The inspection of each machine will be performed in accordance to a Field Service Information (FSI). Customers with questions were instructed to contact the Technical Support Team at 800-621-0445 or the Customers Service Team at 800-848-2066.For questions regarding this recall call 610-997-4717. |  | 147 | 147 | Nationwide Distribution to ND, OH, CA, VA, FL, LA, SC, IL, UT, AZ, KY, IA, WI, OK, TN, SD, AL, NJ, TX, NM, NY, KS, WA, MO, CO, PA, CT, DE, GA, MD, VA, Puerto Rico, and St. Thomas USVI | Terminated | July 11, 2017 | 466 | 1033 | Not_Computer
Z-1605-2016 | 73780 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic CareLink" Monitor (2490C) and Medtronic CareLink Express" Monitor (2020B) |  |  | March 31, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | A recent firmware update developed by Medtronic for the 2490C CareLink Monitors and 2020B CareLink Express Monitors included incorrect data on the country analog modem dial-up configuration table used to dial into the CareLink Network. This incorrect data resulted in the removal of the 0 prefix necessary for select countries as well as an incorrect alignment of phone numbers for other select cou | Software design (manufacturing process) | Medtronic sent an "Urgent Field Safety Notice" dated April 2016to their customers. The letter was addressed to Healthcare Professional. The letter identified the affected product, problem and customer actions to be taken. For questions contact your local Medtronic Representative. |  | 5038 | 5038 | AUSTRIA, AUSTRALIA, BELGIUM, SWITZERLAND, GERMANY, DENMARK, FRANCE, GREAT BRITAIN, HUNGARY, JAPAN, NETHERLANDS, NEW ZEALAND, POLAND, PORTUGAL, SWEDEN. | Open, Classified |  |  | 5038 | Not_Computer
Z-1601-2016 | 73907 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Dynamics; Kinetdx Picture Archiving and Communications System |  |  | April 14, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens has identified changes in default carry forward workflow using VA10B_HF03 version of Syngo Dynamics. If the template design has measurements inside a carry forward-configured report section, this will result in carrying forward the entire section of the report including data from previous studies. This may result in the physician making analysis based on the old measurements. | Software design | Siemens mailed out a "Customer Advisory Notice" #SY009/16/P to customers informing them of the issue and what to do if it was to occur. Siemens informed customers that they are working on the "carry forward" feature enhancement and is planning to release a software update VA10B_HF06 and they should follow the instructions in the advisory notice until the update is performed. |  | 84 | 84 units | Nationwide: AZ, CA, IL, IN, MA, MD, MI, MN, NJ, NY, OH, OR, PA, TN, TX, UT, WA, WI | Terminated | March 20, 2017 | 340 | 84 | Computer
Z-1607-2016 | 74005 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Toshiba American Medical Systems Celesteion, Item Code: PCA-9000A/2, 9000A/3B, 9000A/2B |  |  | April 24, 2016 | 2016 | Toshiba American Medical Systems Inc | Potential degradation of detector performance. | Device Design | Toshiba sent an Urgent Medical Device Correction letter dated April 24, 2016 to all affected customers. The letter informed customers that Toshiba America Medical Systems is bringing attention to a potential problem with the Celestion PET/CT System. The letter informed the customers of a potential degradation of detector performance and informed the customers of the problems and the corrective actions to be taken. Customers were instructed to complete and return the attached form and fax it to (877) 349-3054 or be sent via email to regulatoryaffairs@toshiba.com. Customers with questions were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 3 | 3 units | US Distribution to WI and NV. | Open, Classified |  |  | 3 | Not_Computer
Z-1602-2016 | 73869 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Dissecting Tool Legend 10 cm 4mm BA DIAM. |  |  | April 08, 2016 | 2016 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Incorrect color coding and attachment reference were included on the labeling of 10BA40D dissecting tools with lot number H5237248. | Labeling mix-ups | A notification letter was provided to each consignee, translated as necessary, on April 8, 2016. |  | 81 | 81 tools | Distributed in France, Ireland, Netherlands, Portugal, Switzerland, and United Kingdom. | Open, Classified |  |  | 81 | Not_Computer
Z-1604-2016 | 73641 | stretcher, patient restraint | NZD | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Poly-bags containing the strap labeled with "Adult" sticker. |  |  | March 10, 2016 | 2016 | Ferno-Washington Inc | On March 9, 2016, the firm was contacted regarding the possibility of incorrect labeling on the outside poly- bag of the traction splint straps. The firm investigated and determined its supplier had labeled the outside packaging for the pediatric and adult traction splints straps incorrectly. The packaging labeled Adult contained a pediatric splint strap and the packaging labeled Pediatric conta | Process control | Ferno-Washington, Inc , has contacted all customers via telephone on March 10, 2016. All products will be replaced with new products.For further questions, please call: (937) 283-2885. |  | 270 | 270 | Worldwide Distribution - US Distribution to the states of : AK ,AZ ,CA, CO, FL ,GA ,HI, IL ,IN ,KY ,MA ,MI ,MS ,NC ,NE ,NJ ,NY ,OH,ON,TX ,VA ,WI , and WV. There are no government accounts for this recall. There is one Canadian account for this recall. There are no Mexican accounts for this recall. and to the countries of : UAE, MALAYSIA, UNITED KINGDOM, SLOVAKIA, SOUTH AFRICA, QATAR and AUSTRALIA.. | Open, Classified |  |  | 423 | Not_Computer
Z-1606-2016 | 73631 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit, Model 8015 |  |  | March 24, 2016 | 2016 | CareFusion 303, Inc. | The Alaris PC units model 8015 may display a system error code 133.6080 due to failure with the super capacitor (C245) at power up on the Alaris PC unit logic boards. | Nonconforming Material/Component | The firm, Carefusion, sent a "Medical Device Recall Notification" letter dated March 30, 2016 to its customers. The letter described the product, problem, the potential risk involved and the required action for users. The customers were instructed to do the following: If you observe System Error code 133.6080, remove the PC unit from use and contact the CareFusion Support Center, and promptly complete and return the enclosed Customer Response Card. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018 from 7am-4pm PST, M-F, or by email at supportcenter@carefusion.com. Customers with clinical inquiries, product complaints, or clinical troubleshooting are instructed to contact Customer Advocacy at (888) 812-3266, 24 hrs/day, 7 days/week or by email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229, 6am-5pm PST, M-F or by email at dl-us-inf-techsupport@carefusion.com. |  | 23397 | 23,397 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Great Britain and Canada. | Open, Classified |  |  | 23397 | Not_Computer
Z-1608-2016 | 73318 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TLCÂż Self Retaining Retractor System Contents 1-RETRACTOR FRAME 6-ELASTIC STAYS STERILE EO Rx ONLY |  |  | February 17, 2016 | 2016 | Applied Medical Technology Inc | The firm received a complaint of an open pouch which breaches the sterilization validation. | Packaging | AMT sent an Urgent Medical Device Recall letter dated February 17, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Service at 440-717-4000. |  | 5334 | 5,334 units | Product was shipped to the following states: AL, AR, AZ, CA, CT, FL, GA, IA, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NH, NJ, OH, PA, SC, TN,TX, UT, VA &amp; WI.Product was also shipped to the following countries: Austria, Denmark, Germany, Italy, Japan, Norway, Spain, Sweden, Switzerland &amp; United Kingdom. | Open, Classified |  |  | 12451 | Not_Computer
Z-1620-2016 | 73882 | No | FFT | Stomach pH electrode. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Given Imaging BravoÂżÂż pH capsule delivery device, 1-pack. |  |  | April 15, 2016 | 2016 | Covidien LLC | Potential of an allergic reaction in a patient with nickel sensitivity. | Device Design | The firm, Medtronic, sent an "URGENT: Medical Device Recall" letter dated 4/15/2016 to its Customers via Federal Express. The letter describes the product, problem and actions to be taken. The customer was instructed to promptly forward the information from the letter to recipients, if you have distributed any of the affected devices; and complete and return the Acknowledgement and Receipt Form by email to: Quality Compliance MITGFCA@Covidien.com or Fax to: (203) 492-7719. Medtronic is adding the following statement to the BravoÂż pH monitoring system User Manual and product package insert so that physicians are aware of the nickel content and the potential of an allergic reaction in a patient with nickel sensitivity. "The Bravo capsule contains a trocar needle that is made of stainless steel. Use caution in patients with known sensitivities or allergies to the metals that are contained including chromium, nickel, copper, cobalt and iron. The BravoÂż pH test lasts from 48-96 hours."If you have any questions, please contact Technical Service at 1-800-448-3644, option 3. |  | 1305 | 1,305 units | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Belgium, Brazil, Canada, Cyprus, El Salvador, Greece, Hong Kong, Israel, Mexico, Panama, Singapore, South Korea, and Thailand. | Terminated | July 13, 2017 | 454 | 14200 | Not_Computer
Z-1618-2016 | 73594 | No | LAO |  |  | Toxicology | 510(k) | Syva Emit Methotrexate Assay, is a homogeneous enzyme immunoassay intended for use in the quantitative analysis of methotrexate in human serum or plasma. |  |  | March 08, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics confirmed that Syva Emit Methotrexate Application (DVMTX.1) for Dimension Vista System does not meet On-Board Stability claim of 20 hours. Testing using reagent at the end of the 20 hours On-Board Stability has shown that values of methotrexate for patient and quality control samples increased. | Unknown/Undetermined by firm | Siemens mailed an Urgent Medical Device Correction letter to customers informing them that they were withdrawing the Emit Methotrexate Application on the Dimension Vista System. Customers were asked to contact their local Siemens sales representative for alternative testing solutions for Methotrexate. Siemens advised customers to change the On-Board Stability for the Methotrexate application on the Dimension Vista System to four hours: alternatively, customers can run the Methotrexate application in batch mode, along with QC. It was also requested in the notification that customers complete and return the Field Correction Effectiveness Check Form that was attached with the notification within 30 days. |  | 18364 | 18364 units | Units were shipped to MI, FL, NJ, NY, DC, CA, GA, MI, MT, VA, NC, TX, PA, Australia, Belgium, Czech Republic, Finland, Japan, Republic of Korea, Italy, Saudi Arabia, Slovakia, Slovenia, Spain and Switzerland. | Terminated | August 07, 2017 | 517 | 18364 | Not_Computer
Z-1616-2016 | 73657 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Atlas Spine Torque Limiting Handle 6N-m, P/N 70089 Model # ITN-001-000 |  |  | July 26, 2013 | 2013 | Atlas Spine, Inc. | The device was not appropriately set to mitigate the risk of implant fracture upon rotation inside the disc space. | Device Design | Atlas Spine, Inc. informed all consignees of the recall and sub-recall on July 26, 2013, in person, electronically, and/or via email. Firm takes the responsibility of tracking of set return and inventory control of the product.For questions regarding this recall call 561-354-1108, ext 102. |  | 16 | 16 | Nationwide Distribution to FL, MO, and WV. | Completed |  |  | 16 | Not_Computer
Z-1621-2016 | 73737 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Direct Drive Clip Applier CA090, indicated for ligation of tubular structures or vessels in laparoscopic and general surgical procedures. |  |  | April 01, 2016 | 2016 | Applied Medical Resources Corp | Increased customer complaints indicating inconsistent clip application; may lead to unoccluded vessels. | Nonconforming Material/Component | Customer notification letters were sent on 04/01/16 via UPS.The notification requests that customers check their inventory for recalled products, complete the attached Recall Notification Confirmation Form to acknowledge the recall and indicate if the facility is returing or has already used any of the products. If customers are a distributor it is requested that they notify any facilities to which have been distributed the affected product. Customers are asked to return the Recall Notification Confirmation Form to Applied Medical by emailing to recall60698158@appliedmedical.com or faxing to (949) 713-8921. Customers are to return affected product and a copy of the Recall Notification Confirmation Form to Applied Medical.For product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by email at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or by email at lcontursi@appliedmedical.com. |  | 559274 | 559,274 | U.S. distribution to the following: IL, IN, WY, PA, CA, LA, MN, AZ, OH, OR, CT, WI, NC, AK, NE, MI, KS, FL, MN, TN, SC, MA, MD, KY, TX, GA, TX, NJ, MO, CO, MS, MT, ID, ME, VA, NY, UT, MI, AL, PR, NM, DE, OK, WA, DC, NH, ARForeign distribution to the following; Australia, Spain, New Zealand, Italy, Israel, Estonia, United Arab Emirates, Germany, Netherlands, France, Kuwait, Korea, Qatar, Mexico, Switzerland, Austria, Bahrain, Lebanon, Brazil, Belgium, United Kingdom, Finland, Japan, Argentina, Ireland, Iran, Denmark, Norway, Colombia, Serbia, Uruguay, Sweden, Canada, Costa Rica, South Africa, Czech Republic, Chile, Finland, India, Thailand, Malaysia | Open, Classified |  |  | 559274 | Not_Computer
Z-1613-2016 | 73977 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS  Diagnostic Viewer versions 8.3.7, 8.4.2, 8.4.3, and 8.4.4. Novarad Corporation |  |  | November 16, 2015 | 2015 | Novarad Corporation | The SUV values that are being calculated in the PET/CT fusion tool are incorrect. | Software design | Consignees were notified on 11/16/15 via letter to contact Novarad Technical Support to arrange for an upgrade to a corrected version of NovaPACS software. |  | 2386 | 2,386 | Nationwide. 3 Canadian and 33 foreign consignees. No VA/gov/military. | Open, Classified |  |  | 2386 | Computer
Z-1622-2016 | 73592 | No | CDQ | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista Blood Urea Nitrogen (BUN) Flex reagent cartridge |  |  | February 16, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has determined that Dimension Vista Blood Urea Nitrogen (BUN) lots 15215AE, 15243BB, 15264BA, 15299BB, 15300BA and 15341AC may exhibit inaccurate patient and/or Quality Control results. The issue has been isolated to Flex reagen cartridges manufactured from specific mold cavities. | Nonconforming Material/Component | Siemens Healthcare sent an Urgent Medical Device Correction letter dated February 2016, to all affected customers.It was determined that Dimension Vista Blood Urea Nitrogen (BUN) lots 15215AE, 15243BB, 15299BB, 15300BA, 15320BB and 15341AC may exhibit inaccurate patient and/or Quality Control results. The issue has been isolated to Flex reagent cartridges manufacture from specific mold/cavity identifies as described in Table 2 of the Urgent Medical Device Correction notification sent to customers.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (302) 631-0528. |  | 18668 | 18668 units | Worldwide distribution - US (nationwide) distribution including Puerto Rico and to the states of : AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. , and to the countries of : Canada, United Arab Emir, Australia, Austria, Belgium, Switzerland, Bahamas, Czech Republic, Germany, Denmark, Spain, Finland, France, Italy, Netherland, Norway, Portugal, Slovakia, Slovenia, Japan, Republic of Korea, New Zealand. | Open, Classified |  |  | 18668 | Not_Computer
Z-1617-2016 | 73827 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun Temperature Management System, 5000, 100-120V and 220-240V |  |  | November 09, 2015 | 2015 | C.R. Bard, Inc. | The product may contain electronic components that lead to a premature drainage of the internal control panel coin cell battery responsible for maintaining the system close and Static Random Access Memory (SRAM). This premature coin cell battery drainage could render the device unresponsive upon System startup. | Unknown/Undetermined by firm | BARD sent an Urgent Medical Product Recall to all affected consignees on November 9, 2015. The letter indicates that a technician will contract the firm for the evaluation and potential replacement of the affected electronic components. The technician will fill out the Field Action Effectiveness Check Form after evaluation.Customers with questions were instructed to call 1-770-784-6471 or email bmd@crbard.com.For questions regarding this recall call 770-784-6220. |  | 2613 | 2613 | Worldwide Distribution to AL, AK, AZ, AR, CA, CO, CT, DEL, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Finland, France, Germany, Gibraltar, Greece, Guam, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Oman, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom | Terminated | June 06, 2017 | 575 | 2613 | Not_Computer
Z-1615-2016 | 73801 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur CP Immunoassay System, Siemens Material Number (SMN) 10309710, IVD. --- CLASSIFICATION NAME: Immunoassay Analyzers. |  |  | March 30, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | The firm identified an issue in sample processing when using EZee-Nest Insert Cups due to two inaccuracies that exist in the ADVIA Centaur CP Operators Guide. | Use error | Siemens sent an Urgent Medical Device Correction letter dated March 2016, to all affected customers. The letter was sent to the US customers via Fed Ex beginning on March 30, 2016., and were e-mailed dated March 20, 2016, to regional countries for implementation outside the US. Field service personnel were sent a support bulletin describing the issue and instructing them how to manage customer questions. " Complete and return the Field Correction Effectiveness Check form attached to this letter within 30 days." Please review this letter with your Medical Director.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (914) 524-2868. |  | 3679 | US: 442 units; Foreign: 3237 units | Worldwide Distribution - USA (nationwide) and to the countries of : Afghanistan, Albania, Algeria, Angola, Argentina, Australia, Austria, Bangladesh, Bosnia Herzeg., Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Kenya, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mexico, Netherlands, New Caledonia, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uganda, United Kingdom, Uruguay, Vietnam, Yemen. | Terminated | March 29, 2017 | 364 | 3679 | Not_Computer
Z-1623-2016 | 73856 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Femoral Replacement (patient specific, custom made orthopedic implant device). PIN 19960 |  |  | April 12, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | Incorrect component used to manufacture distal femoral replacement. | Mixed-up of materials/components | A notification was sent to the prescribing physician on 04/12/2016. |  | 1 | 1 | US distribution to OH. | Open, Classified |  |  | 1 | Not_Computer
Z-1626-2016 | 73835 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Launcher Guiding Catheter |  |  | April 04, 2016 | 2016 | Stryker Sustainability Solutions | Device packaging may not be sealed. | Process control | Customer notification letter was sent 04/04/2016.The customer was provided the following instructions:Use of the affected lots of the recalled product should be discontinued. A RecallEffectiveness Check Form has been enclosed. Note that the recall is limited to the specific lot number shipped to facility.Once inventory has been checked complete the Recall EffectivenessCheck Form. The form needs to be completed even if no affected product isfound.Return the completed and signed Recall Effectiveness Check Form utilizingone of the following methods:Return to your local Stryker Sustainability Sales RepresentativeEmail to ssspfa@stryker.com, orMail to:Stryker Sustainability Solutions181 O West Drake DriveTempe, AZ. 85283Attn: Jodie RueckertIf the form indicates that any affected devices remain in inventory, a prepaidshipping label will be issued for the return of the product. Customers will receive acredit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433x5555http://www.stryker.com/productexperience/The FDA's MedWatch Adverse Event Reporting program either online, byregular mail, or by fax. |  | 2 | 2 | US distribution to KY. | Open, Classified |  |  | 6 | Not_Computer
Z-1624-2016 | 73816 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton G5 with software version between v2.41, v2.42, v2.50. |  |  | March 09, 2016 | 2016 | Hamilton Medical, Inc. | After performing the suctioning maneuver, including disconnecting the patient, suctioning , and reconnecting the patient, the preset pattern of ventilation many not continue as expected. | Software design | Hamilton Medical Inc. sent an" Medical Device Field Safety Corrective Action" letter dated March 7, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Hamilton Medical then had to compile a list of to whom they had distributed the affected ventilators and sent letters on April 6, 2016. Letters included the Medical Device Field Safety Corrective Action from Hamilton AG. Letters included immediate actions to be take by use and that reps will be visiting each site to upgrade ventilators to v 2.60 to alleviate the issue.Distributors will update software and will provide an updated Operator's manual as soon as it is available.Questions may be directed to 800-426-6331, Ext, 215. |  | 1115 | 1115 | US distribution only. | Terminated | June 29, 2017 | 477 | 1115 | Computer
Z-1633-2016 | 73442 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Dimension clinical chemistry system, Sirolimus (SIRO) Flex reagent cartridge (DF306 SMN 10464331) Lot EB6064, Open Well Instability |  |  | February 09, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received customer complaints regarding QC shifts and imprecision after one day when using DimensionÂż SIRO lot EB6064. They have confirmed SIRO lot # EB6064 does not meet the open well stability claim of 2 days. | Under Investigation by firm | Siemens sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be Taken by the Customer:Discontinue use of and discard the kit lot listed in Table 1.Please review this letter with your Medical Director.Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days.If you have received any complaints of illness or adverse events associated with the product listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions, please call (302) 631-0528. |  | 394 | 394 | Worldwide distribution and US including Puerto Rico and to the states of : TX, FL, CA, LA, DC, GA, IL, AR, WI, MI and MD. | Open, Classified |  |  | 394 | Not_Computer
Z-1648-2016 | 73669 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Covidien Kangaroo Connect Enteral Feeding Pump |  |  | April 01, 2016 | 2016 | Medtronic | Kangaroo Connect Feeding Pump Occlusion alarms fail to alarm | Device Design | Medtronic(formerly) Covidien notified accounts via Urgent Medical Device Recall letter, dated 4/1/16, that specific serial numbers with software versions 1.12 and 1.13 have a software anomaly. The firm shipped replacement units with software V 1.14. Affected devices must be returned to the firm once the replacement units have been received. A Recalled Product Return Form required to be completed and returned. Questions or concerns, contact Medtronic representative or Customer Service at (800) 882-5878. |  | 223 | 223 units | Worldwide Distribution -- USA, Australia, Canada, France, and Singapore. | Open, Classified |  |  | 368 | Computer
Z-1632-2016 | 73837 | Yes | LWS |  |  | Cardiovascular | PMA | Ellipse ICD (Implantable Cardioverter Defibrillator) Model CD2411-36Q, CD2411-36C, CD1411-36Q, CD1411-36C |  |  | March 31, 2016 | 2016 | St Jude Medical Inc. | St. Jude Medical is recalling the Ellipse ICD (Implantable Cardioverter Defibrillator) due to the potential inability to deliver high voltage therapy. | Other | A customer letter dated 4/7/16 was sent to all customers to inform them that St. Jude Medical is recalling the Ellipse ICD (Implantable Cardioverter Defibrillator) due to the potential inability to deliver high voltage therapy. The letter informs the customers that the recall involves information related to 127 units of our Ellipse ICD device distributed worldwide but not implanted in patients. Customers with any questions about the advisory are instructed to contact their local Sales Representative or St. Jude Medical Technical Services at 800-722-3774, 24 hours a day, seven days a week. |  | 127 | 127 units | Worldwide Distribution - US Nationwide in the states of CT, NJ, SC, FL, GA and countries of: Finland, France, Germany, Great Britain, Greece, Israel, Italy, Japan, Kuwait, Lebanon, Netherlands, Northern Ireland, Poland, Romania, Slovakia, Spain, Tunis, United Arab Emirates, and Vietnam. | Open, Classified |  |  | 127 | Not_Computer
Z-1634-2016 | 73846 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | CinchLock SS Knotless Anchor with Inserter; |  |  | April 12, 2016 | 2016 | Stryker Corporation | Stryker Endoscopy is recalling the CinchLock SS Knotless Anchor with Inserter and the CinchLock Flex Anchor with Inserter after 13 reports of s of an internal wire breakingupon deployment of the anchor. | Device Design | Urgent Medical Device Recalls letters were sent on April 12, 2016 to US customers by certified mail. Stryker International distribution sites will be notified by e-mail. Letters describe the reason for recall with the affected part and lot numbers. Any recalled product should be segregated and returned to Stryker. |  | 3100 | 3100 total device, both model numbers. | Distributed US (nationwide) and in Australia, Switzerland, Netherlands, Spain and Sweden. | Open, Classified |  |  | 3100 | Not_Computer
Z-1638-2016 | 73578 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Ascom Mobile Monitoring Gateway (versions 4.1.1 and 4.2.0) and Ascom Unite Connect for Nurse Call (version 2.2.0). |  |  | December 10, 2015 | 2015 | Ascom US, INC. | Due to a malfunction of software, the secondary module will not automatically take over messaging if the primary module fails after a software upgrade to one of the other affected versions. Messages from GE Carescape/Nurse Call will not forward to handset until primary module is restored. | Software design | On January 4, 2016 , the firm is initiating a field corrective action to update the software in the installed base and resynchronize the primary and secondary servers. All customers were sent a letter January 4, 2016, informing them of the voluntary recall of Ascom Mobile Monitoring Gateway and Unite Connect for Nurse Call. The reason stated for why the recall was initiated was that the secondary module would not take over the messaging function when the primary module failed, following a software upgrade. Ascom instructed customers to check if their software was one of the affected versions they should contact their local Ascom representative to review the MMG configuration and determine if the system is subject to malfunction. The same instructions were directed towards those who were not able to determine their software version. If it is determined that the program is subject to malfunction, the primary and secondary modules should be re-synchronized. A test of module redundancy should be conducted on re-synchronized modules. The second module should take over for the primary module within thirty seconds. Model redundancy should be tested frequently and following system upgrades. For more information call (919) 535-6143 , 8am till 4:30pm. |  | 319 | 319 | CA, CT, FL, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MO, NJ, NM, OH, OK, PA, TN, TX, VA, WA, WI and Hawaii.Australia, Austria, Belgium, Canada, France, Germany, Netherlands, Poland, Qatar, Sweden, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 319 | Computer
Z-1637-2016 | 73873 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Axial Torque Limiting Driver (1st generation) |  |  | February 02, 2015 | 2015 | Bradshaw Medical Inc. | A component failure could cause a situation where the applied torque was significantly higher than the calibrated torque setting. | Device Design | Consignees were sent a Bradshaw Medical "Product Withdrawal Notice" letter, date marked starting February 2, 2015. The letter specified the Part Number, Lot Number (s) and Shipped date for each consignee. The letter described the product and advised consignees to discontinue use of the product. Requested consignees to complete and return the Customer /Distribution verification Form to BMI Regulatory Affairs FAX (262) 925-1380 and to further send the "Withdrawal Notice and Customer Verification Form" to all customers to whom they have distributed the product. |  | 3760 | 3,760 units | US Nationwide Distribution in the states of: CA, FL, IN, MA, MI, OH, PA, TN UT, VA and the country of UNITED KINGDOM | Open, Classified |  |  | 3760 | Not_Computer
Z-1627-2016 | 73908 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | GEM(TM), FLOWCOUPLER(R), 2.0 mm, STERILE EO, Rx only, 1 FLOWCOUPLER, REF GEM2752-FC |  |  | April 21, 2016 | 2016 | Baxter Healthcare Corp. | Instructions for use booklet may puncture the outer Tyvek lid. | Package design/selection | Baxter sent an Urgent Product Recall letter dated April 21, 2016 to customer. The letter identified the affected product, problem and actions to be taken. The letters requested that direct consignees contact Synovis MCA to arrange for return and credit. Companies that further distributed the product were asked to notify other facilities or departments. For questions contact Synovis MCA at 800-510-3318 or 205-941-0111. |  | 315 | 315 units | Nationwide and Germany, Sweden, Netherlands, Israel, Denmark, United Kingdom, Canada | Open, Classified |  |  | 2058 | Not_Computer
Z-1654-2016 | 73667 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CHEMISTRY ENZYMATIC CREATININE_2 |  |  | March 17, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has become aware of N-Acetylcysteine (NAC) and Metamizole (Dipyrone) interference with Trinder and Trinder-like reaction assays. Siemens has confirmed that falsely depressed results may occur on samples drawn from patients receiving N-Acetylcysteine (NAC) or Metamizole. NAC is the accepted antidote for acetaminophen toxicity and is justified in patients at signific | Device Design | Siemens mailed an Urgent Medical Device Correction letter to customers making them aware of the problem. Special instructions were also given for the customer to review the letter with their Medical Director and ensure Venipuncture occurs before drug administration of NAC or Metamizole as indicated in the letter. Baseline assay values before administration of NAC or Metamizole therapy would not be affected. Customers were asked to complete and return the Field Correction Effectiveness Check Form attached with the letter within 30 days. Furthermore, if customers received any complaints of illness or adverse events associated with the specified lots, they were to immediately contact their local Siemens Customer Care Center or local Siemens technical support representative. |  | 137422 | 137422 units total | Nationwide: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Puerto Rico, St. Thomas, VI, Guam, Pacific Mariana Islands | Open, Classified |  |  | 137422 | Not_Computer
Z-1650-2016 | 73832 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 64 Multi-Slice CT Scanner System (consist if two variants: NeuViz 64e, NeuViz 64i) |  |  | May 02, 2015 | 2015 | Neusoft Medical Systems Co., Ltd. | It was found by R&D; in April, 2014 that there was a defect in NeuViz 64 system software version 1.0.5+P09. The icons of side decubitus didn't meet the patient position description, when the scanner position was set to "Right" in the system setting - scanner options. | Device Design | An upgraded software version (1.0.5.2619+P11) was released on May 7, 2015. The FCO was released to the service engineers of Neusoft Medical Systems USA, which is subsidiary of the reporting firm, and the engineers performed FCO for each affected device together with the distributors' engineers. The corrections (FCOs to upgrade software) on all the affected systems in the USA have been completed. At present (4/27/16) there is no longer this issue on the systems, so no further actions are needed for customers/distributors. However, the recalling firm shall inform the direct accounts of this fix before May 5, 2016. The recalling firm will submit the communication evidence ASAP once they have evidence of the communication. For further questions, please call (281) 453-1205. |  | 7 | 7 units | US Distribution - Including Puerto Rico and the states of IL, CT, SC, NE | Open, Classified |  |  | 7 | Computer
Z-1647-2016 | 73413 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 Gram Positive Susceptibility Test Cards (AST-P646 REF 420144), 20 cards/kit. |  |  | February 24, 2016 | 2016 | Biomerieux Inc | Product insert fails to identify performance limitation related to EUCAST breakpoints for Teicolplanin. | Labeling Change Control | Biomerieux sent an Urgent Product Correction Notice letter via email on February 24, 2016 to affected customers . The letter identified the affected product, problem and actions to be taken. Customers were asked to complete the attached Acknowledgement Form (Attachment A) and return it to Biomerieux, Inc. For questions contact your local bioMÂżrieux representative. |  | 1035 | 1,035 kits | Foreign distribution in the countries of Czech Republic, Hong Kong (China), Hungary, Japan, Poland, South Africa, Switzerland, and Turkey. | Open, Classified |  |  | 37476 | Not_Computer
Z-1577-2016 | 73701 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Battery Pack Used with the Capnostream"20 and Capnostream"20p Patient Monitor, Battery Packs Model Number - 016400, 010520. |  |  | April 13, 2016 | 2016 | Medtronic | Medtronic Respiratory &amp; Monitoring Solutions Announces a Voluntary Recall of a Battery Pack Used with the Capnostream"20 and Capnostream"20p Patient Monitor Due to the Battery Pack Might Experience Thermal Runaway with the Potential of Fire Risk to the Monitor. | Process change control | Medtronic sent an Urgent Medical Device Recall letter dated April 2016, by postal mail or courier to all affected customers. The letter instructed users to remove and properly dispose of current batteries. If product was further distributed customers were instructed to forward the recall letter. Customers were asked to completed and return the Acknowledgement and Receipt Form by fax or email. Customers with questions should contact Technical Services at 800-635-5267, option 1, ten 1 and 1 or by email at HQTSWEB@COVIDIEN.com.For questions regarding this recall call 303-876-8859. |  | 9817 | 9,817 batteries (8,170 with monitors and 1,647 as spare parts) | Nationwide Distribution to AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Product was also distributed worldwide. There is 1 VA customer. | Open, Classified |  |  | 9817 | Not_Computer
Z-1669-2016 | 73765 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lumenis M22 System Model Number: GA-0005400 (M22 IPL + YAG Module) with Acne Filter (KT-1014971). |  |  | November 17, 2015 | 2015 | Lumenis Ltd | Lumenis Ltd Announces a Field Action of the M22 IPL Acne Filters for the Lumenis M22 IPL Hand Piece due to the Risk of Superficial Burns When Using the Device. | Device Design | Customers were notified on November 17, 2015 by Customer Notification Letter and customers were asked to return the device. |  | 33 | 33 filters | Distributed in the states of GA, CA, CO, CT, KY, LA, MA, MN, NJ, NY, RI, and SC, and the countries of Italy, Germany, France, and China. | Terminated | November 30, 2016 | 379 | 33 | Not_Computer
Z-1670-2016 | 73635 | No | CDT | Triglyceride test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Triglyceride_2, concentrated; TRIG_c; Catalog # SMN # 10697575. |  |  | March 07, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the ADVIA Chemistry Triglyceride concentrated reagent (TRIG_c) does not meet Instructions For Use (IFU) linearity claims at the upper limit of the assay range as it approaches end of shelf life for lots # 348297 and 359932. | Under Investigation by firm | Siemens mailed an Urgent Field Safety Notice, dated March 2016, to customers confirming the ADVIA Triglyceride concentrated reagent (TRIG_c) does not meet Instructions for Use (IFU) linearity claim at the upper limit of the assay range as it approaches end of shelf life for lots # 348297 and 359932. Siemens asked customers to discontinue use of and discard kits from the two referenced lots, complete and return the Field Correction Effectiveness Check attached to the letter within 30 days. Customers were to review their inventory to determine their laboratory 's replacement needs. Customers are asked to retain the letter for their records and forward a copy to anyone that received the product. Customers may contact Siemens Customers Service or their local technical support. |  | 7433 | 7,433 units (4,228 for lot 348297 and 3,199 for lot # 359932) | Worldwide Distribution -- USA, including the states of CA, TN, UT, OK, FL, MI, MN, NY, CT, NC, KS, TX, CO, WA, NM, and OR; and, the countries of Austria, Australia, Belgium, Brazil, Bulgaria, Denmark, Estonia, France, Germany. Great Britain, Italy, Latvia, Netherland, Norway, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Czech Republic, India, New Zealand, Thailand, Turkey, and Russian Fed. | Open, Classified |  |  | 7433 | Not_Computer
Z-1671-2016 | 73786 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB Sciex Triple Quad 4500MD LC/MS/MS System, Part No. 5031257; and QTRAP 4500MD LC/MS/MS System, Part No. 5031231. |  |  | March 25, 2016 | 2016 | Ab Sciex | Under certain conditions, mass spectrometers may report incorrect quantitative results. | Under Investigation by firm | AB SCIEX issued Urgent: Medical Device Correction letter, dated 3/25/16, advising user of the problem and advising not to use the system until corrected. The letter also discussed the risk to health, as well as how to recognize if the issue has occurred. Sciex Field Service Employee (FSE) will be sent to customer site to perform the field corrective action on the affected system. Questions regarding the letter should be directed to SCIEX at 1-289-982-2531. |  | 2 | 2 units | Distributed to the state of NC. | Completed |  |  | 2 | Not_Computer
Z-1686-2016 | 73694 | No | LSE | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | Diamedix Is-EBV-VCA IgG Test Kit |  |  | October 20, 2014 | 2014 | Diamedix Corporation | A gel-like substance was found in kit component, sample diluent , that could potentially compromise product performance. | Unknown/Undetermined by firm | The firm, Diamedix, sent letters dated October 14, October 20 and December 5, 2014 to its customers directly informing them of the correction action. Notification of all domestic customers via Certified Mail. Notification fo the international customers via Federal Express and/or e-mail. With the letter, a replacement bottle of Sample B Diluent was sent and the customer was advised to discard any unused bottles of affected product. New lot number labels were provided to identify kits with the replacement Sample B Diluent. Customers with questions were told they could contact Diamedix Technical Services (1-800-327-4565) if they had any questions or required any additional guidance regarding the correction. |  | 588 | 588 kits | Worldwide distribution: US nationwide including states of: AZ, CA, CO, FL, GA, IL, IN, KY, MA, MI, MO, NC, NE, NH, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, WI and Puerto Rico. and Country of: ITALY. | Open, Classified |  |  | 1379 | Not_Computer
Z-1672-2016 | 74096 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visualase Cooled Laser Applicator System (VCLAS). The VCLAS system consists of the Visualase Laser Diffusing Fiber (LDF) and the Visualase Cooled Catheter System (CCS). |  |  | May 09, 2016 | 2016 | Medtronic Navigation, Inc. | Potential for the Visualase Cooled Laser Applicator System (VCLAS) catheters to melt during procedures. | Labeling design | Medtronic will be sending a Medical Device Safety Notification letter FedEx on May 18, 2016. Each customer will be asked to complete and return a response form to Medtronic. This is a safety alert notification and does not require return of devices.We are committed to patient safety and appreciate your prompt attention to this matter. If you have any questions regarding this action, please contact Medtronic Technical Services at 1-800-595-9709. |  | 120 | 120 | Distributed within the US, Netherlands, Canada and Australia and to govt/VA/military consignees. | Open, Classified |  |  | 120 | Not_Computer
Z-1676-2016 | 73981 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Periarticular plating system, cancellous bone screw, partially threaded, hex head, 4.0 mm diameter, 16 mm length |  |  | May 04, 2016 | 2016 | Zimmer Manufacturing B.V. | A small hole may be present in the inner tray of specific lots of sterile Hip and Trauma Implant Screws. | Packaging change control | All distributors will be notified via electronic mail. Distributors with affected lots will also be notified via courier. Hospitals risk managers and surgeons will be notified via recall notification letter, dated May 4, 2016. |  | 99 | 99 | Worldwide distribution. US nationwide, Korea, Japan, Argentina, Brazil, Singapore, Taiwan, China, El Salvador, Australia, India, Hong Kong, Chile, Germany, Italy, France, Spain, UK, Sweden, Canada. | Terminated | July 05, 2017 | 427 | 11400 | Not_Computer
Z-1614-2016 | 73849 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | NeuroBlate System SideFire Select 2.2 mm Directional Laser Probe, Sterile EO. Model #s SFS000-01,SFS002-01, SFS122-01, SFS222-01, SFS322-01, SFS422-01, and SFS522-01. |  |  | March 25, 2016 | 2016 | Monteris Medical Corp | Monteris received a report that the sapphire lens on a 2.2SF probe fractured during a LITT procedure, with resulting introduction of C02 into the patient's brain. | Device Design | Monteris Medical Corporation field personnel visited each consignee starting March 25, 2016 and April 1, 2016. The devices were physically removed from each site. Customers from whom the devices were removed will either receive a credit or receive replacements of other models of laser delivery probes. |  | 121 | 121 | Nationwide Distribution in the states of KS, MO, NC, OH, and VA. | Terminated | March 10, 2017 | 350 | 121 | Not_Computer
Z-1689-2016 | 73194 | No | LFZ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | Diamedix Is-CMV IgG Test Kit |  |  | March 26, 2015 | 2015 | Diamedix Corporation | Product contained an incorrect substrate. | Other | On March 26, 2015 Erba Diagnostics sent a communication to their clients informing them of the recall and sending t hem a replacement sustrate HRP to replaced the wrong one sent on the kit. |  | 150 | 150 kits | AZ, CA, FL, GA, KY, MA, MI, MN, MO, NC, NE, NY, OH, OK, OR, TN, TX, VA, WI and Puerto Rico. | Open, Classified |  |  | 150 | Not_Computer
Z-1688-2016 | 73883 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | LenSx Laser System. Indicated for use in patients undergoing cataract surgery and refractive surgery. |  |  | April 14, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting a voluntary recall of all LenSx Lasers with Corneal Flap Capability after receiving reports concerning incomplete corneal flap creation, which manifests as an un-dissected area(s) within the side cut. | Software design (manufacturing process) | Recall notification letters were sent to consignees on April 14, 2016. The letter details precautions to be taken by the surgeon and requests the consignee sign and return an acknowledgement form. |  | 81 | 81 units | United States, Albania, Australia, Brazil, Chile, Costa Rica, Czech Republic, France, Germany, Greece, Hungary, India, Italy, Romania, Slovakia, Spain, Taiwan, Thailand, Turkey, UAE, Ukraine | Open, Classified |  |  | 81 | Not_Computer
Z-1690-2016 | 73691 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 1) 24 G x 0.75 in. BD Angiocath" Autoguard" shielded IV |  |  | March 18, 2016 | 2016 | Becton Dickinson &amp; Company | The device may have a defect in the catheter. In some instances this defect could result in catheterseparation or breakage. | Device Design | BD isssued an urgent voluntary product recall lettter/recall response card on 3/18/2016. |  | 1305150 | 1,305,150 units | Nationwide | Terminated | July 19, 2017 | 488 | 1305150 | Not_Computer
Z-1693-2016 | 73777 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | 23G and 25G Constellation Totalplus Paks provides access into the posterior segment of the eye for surgical instruments within the frame of 23G and 25G Pars-Plana-Vitrectomy (PPV) surgeries. |  |  | August 11, 2015 | 2015 | Alcon Research, Ltd. | The product has potential to leak beyond their design specification. | Process control | The affected consignees were contacted via letter on 8/11/15 to communicate the potential for leakage. |  | 225000 | 225,000 units | United States, Argentina, Armenia, Australia, Azerbaijan, Bangladesh, Brbados, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Curacao, Cypress, Denmark, Ecuador, Finland, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Lsreal, Italy, Japan, Korea, Lithuania, Malta, Mexico, Monaco, Netherlands, New Zealand, Pakistam, Philippines, Poland, Prtougal, Puerto Rico, Romania, Russia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay | Terminated | April 20, 2017 | 618 | 339149 | Not_Computer
Z-1668-2016 | 73738 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | ANSPACH, Rx only, Otologic Curved Micro (OCCM) Diamond Ball Burr, 2.0 MM, Part # OCM7-2SD |  |  | March 28, 2016 | 2016 | The Anspach Effort, Inc. | Certain lots of the Otologic Curved Micro (OCM) Diamond Ball Burrs may be missing the diamond coating and the ball may be the incorrect size. | Process control | Consignees were sent an Urgent Notice of this Medical Device Recall on 3/28/2016 for their product Otologic Curved Micro (OCM) Diamond Ball Burr, with the following actions required: We have record that your facility has the product(s) subject to this recall (removal). DePuy Synthes asks that you review your inventory and immediately remove the affected lots from stock.Please take the following actions: If you DO HAVE any of the identified devices, please take the following steps: Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Please contact Product Support at (800) 327-6887, option 1, to arrange for return of any unused product. A credit will be issued for returned product. Send a copy of the completed Verification Section (page 3) to Product Support by fax to (561) 627-2682. Send a copy of the completed Verification Section (page 3) to the Complaint Handling Unit byscan/email to RA-DPYUS-chu.pb@ITS.JNJ.com. If you DO NOT HAVE the identified product, please take the following steps: Complete the Verification Section by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledgesyour receipt of medical device recall (removal) information.Âż Return the documents to The Anspach Effort by:Âż Fax: (561) 627-2682Âż Scan/email: RA-DPYUS-chu.pb@ITS.JNJ.comÂż Giving them to your DePuy Synthes Sales ConsultantÂż Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual of the notification. |  | 27 | 27 devices | US Distribution in the states of KY and TX. | Open, Classified |  |  | 27 | Not_Computer
Z-1636-2016 | 72350 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | ABG-HM-1 Hummi Micro Draw Blood Transfer Device |  |  | October 01, 2015 | 2015 | Hummingbird Med | Hummingbird Med Devices, Inc. is recalling ABG-HM-1 Hummi Micro Draw Blood Transfer because the connection between the "Hummi" Y connector and the yellow cannula hub may not be sufficiently secure enough and could separate while in use. | Component design/selection | On 10/01/15 the firm sent out customer notification letters.The letter states for the following actions to be taken:1. Customers using the Hummi Micro Draw device are recommended not to use the identified lot.2. A local distributor/sales representative will be contacting customers to arrange a return of the identified lot for credit and arrangement of replacement.3. Replacement product should be available within 10 days of 10/30/15.4. During the period of product unavailability customers might want to choose to temporarily to return their previous method for blood drawing.The firm requests a Reply Form be completed and returned via fax at 949-583-2775 or email to JorgeHaider@hummingbirdmed.com.On 11/19/2015 the firm sent out a second notification letter involving additional recall lots. The firm provided the same instructions to customers as their previous notification; however stated replacement product should be available by mid December 2015. |  | 11500 | 11500 units | Distributed in the states of CA, KY, MD, GA, and IL. | Terminated | January 26, 2017 | 483 | 11500 | Not_Computer
Z-1700-2016 | 73441 | No | DHA | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 1) VITROSÂż Immunodiagnostic Products Total Beta-hCG II Reagent Pack |  |  | February 18, 2016 | 2016 | Ortho-Clinical Diagnostics | When testing was conducted using certain lots, VITROSÂż Systems generated results within the measuring range for samples known to not contain measurable hCG. The system reported results up to approximately 7.40 mIU/mL (IU/L) for patient samples that should have been less than the measuring range of the assay (&lt;2.39 mIU/mL (IU/L)). | Under Investigation by firm | On February 18, 2016, Ortho Clinical Diagnostics distributed Urgent Product Correction Notification letters and Confirmation of Receipt Forms to their customers (Ref. CL2016-043) via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail. Customers are advised to immediately discontinue using and discard all remaining inventory of the listed lots. Ortho Clinical Diagnostics will replace all remaining inventory or provide account credits as indicated on the confirmation of receipt form. Customers are advised to review previously reported results using the affected lots and to discuss these findings with your Laboratory Medical Director to determine the appropriate course of action. Post the notification by your VITROS System or with your user documentation. Complete and return the Confirmation of Receipt form to Joe Falvo via email at confirmationadmin@its.jnj.com or by fax to 1-888-557-3759 or to 1-585-453-4110. Clients with questions can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 14780 | 14,780 pieces | Nationwide Distribution. | Open, Classified |  |  | 14780 | Not_Computer
Z-1702-2016 | 73815 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | AngioVac Circuit Pack (Tandem Packs), REF/Catalog No. 25-187, UPN H965251870, STERILE, Rx Only, Single Use Only -- The AngioVac Circuit and AngioVac Circuit Pack are designed to be used with the AngioVac Cannula. |  |  | March 21, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | Maquet Cardiopulmonary, the manufacturer of the RotaFlow Centrifugal Pump, determined that the products may possess levels of endotoxin that exceed regulatory limits. AngioDynamics confirmed that the affected RotaFlow Centrifugal Pumps were included in packaged AngioVac Circuit Packs. | Environmental control | On 3/21/2016, AngioDynamics sent recall notification packages (dated 3/21/2016) to the consignees via Federal Express. Consignees are instructed to follow the instructions provided on the Maquet Cardiopulmonary recall letter (dated 2/23/2016) and return the Reply Verification Tracking Form, provided in the recall notification. |  | 377 | Domestic: 377 kits | Nationwide Distribution. | Open, Classified |  |  | 716 | Not_Computer
Z-1703-2016 | 74171 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | WHIN Infusion Set - 90 Huber Needle and Wing (22 ga/1in). Intravenous delivery through intravascular ports, 20 units per case, Catalog Number 471734. |  |  | April 26, 2016 | 2016 | B. Braun Medical, Inc. | Potential for the tip of the Huber needle on the WHIN Infusion Set to protrude from the needle guard. | Component design/selection | The firm, B. Braun Medical Inc., sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 4/26/2016 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to evaluate their current inventory for the affected lots; fill out the Product Removal Acknowledgement form and record the total number of units found, and return the form to B. Braun Medical Inc. Quality Assurance department by faxing the form to 610-849-1197 or email to PA_Quality Assurance.BBMUS_Service@bbraun.com within two weeks of receipt. A BBMI Customer Support Representative will contact the customer to provide instructions for handling the affected product and arrange for return to BBMI.Should you have any questions or concerns regarding the attached information, please contact our Customer Support Department at (800) 227-2862. |  | 5980 | 5,980 units [Lot # 0061482060 (3,340 units); Lot # 0061486821 (2,640 units)] | Worldwide distribution: US Distribution to states of: IN, NV, MA, IL, AR, TX, OH, CA, GA, FL, MS, NJ, AL, PA, TN, KY, VA, SC, and country of: Canada. | Open, Classified |  |  | 5980 | Not_Computer
Z-1699-2016 | 73678 | [6 empty] | Orapharma (Onpharma) Onset Mixing Pen Rx only, Product/Part Number 1054-02. |  |  | January 29, 2016 | 2016 | Valeant Pharmacueticals International | It has been determined that the Onset Mixing Pen has an optimal use period of 18 months from start of use. However, the current labeling and IFU does not stipulate this time period. | Labeling False and Misleading | Valeant sent a letter entitled "Field Correction to Orapharma Onset Mixing Pen," dated 1/29/2016, to customers via UPS. The letter identified the affected device and the issue involved. The firm recommends that users stop using the device if it had been in use for more than 18 months. Customers were to refer to the list of the affected lots to determine if their device needed to be replaced. If the device was one of the ones identified, users were to contact Valeant Pharmaceuticals to arrange for a replacement at no charge. Questions should be directed to customer service 800-321-4576. |  | 7716 | 7,716 | Nationwide Distribution | Open, Classified |  |  | 7716 | Not_Computer
Z-1704-2016 | 73823 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Mouse Monoclonal Antibody Calretinin, NCL-L-CALRET-566. |  |  | March 22, 2016 | 2016 | Leica Microsystems, Inc. | Novocastra Liquid Mouse Monoclonal Antibody Calretinin when used in combination with specific Lots of antibody diluent will give reduced staining. | Nonconforming Material/Component | A Medical Device Recall Notification letter was sent via e-mail on April 5, 2015 to the consignees describing the nature of the recall and action that needed to be taken. The firm requested that the consignees take the following actions: Only use Novocastra Liquid Mouse Monoclonal Antibody Calretinin with the lot numbers of diluent specified.There is no need for retrospective review of tissue staining if appropriate positive controls were properly used and interpreted during the immuno / in situ staining.Complete and return the attached acknowledgment.The firm requests that consignees contact their local Leica representative if they have any questions or concerns, or alternatively please contact the firm at the following e-mail address; LMGRA@leica-microsystems.com |  | 290 | 290 units | Worldwide Distribution -- U.S., to the following states: AK, CA, FL, GA, MN, MT, NJ, OH, TX and WA; and, the country of Canada. | Open, Classified |  |  | 290 | Not_Computer
Z-1709-2016 | 73945 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2 mm Three-fluted drill bit, sterile, Catalog ID # 03.010.060S |  |  | April 18, 2016 | 2016 | Synthes (USA) Products LLC | The sterile packaging may be compromised. A protective cap is to be secured on the tip of the drill bit. If the protective cap is not secured, the drill bit may puncture the sterile barrier. | Device Design | DePuy Synthes sent an Urgent Notice Medical Device Letter dated April 19, 2016 to customers. The identified the affected product, problem and actions to be taken. The letter instructed customers to examine their inventory to determine if they currently have the affected product. Customers were asked to complete and return the Verification Section of the recall notification letter. For questions call 610-719-5450. |  | 7363 | 7363 units | US Nationwide distribution in the states of: AE, AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX,UT, VA, VT, WA, WI | Open, Classified |  |  | 8637 | Not_Computer
Z-1708-2016 | 73983 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Daig Supreme Fixed Curve Diagnostic Electrophysiology Catheters; Model #: 401450, 401877. Indicated for temporary intracardiac sensing, recording, stimulation, and electrophysiology mapping of cardiac structures. |  |  | April 28, 2016 | 2016 | Stryker Sustainability Solutions | Mislabeled for size | Under Investigation by firm | The firm sent out customer notification letters on 04/28/16.Customers are instructed to discontinue the use of affected lots of recalled product. Other EP Catheters reprocessed by Stryker Sustainability Solutions should be considered as alternatives. A Recall Effectiveness Check Form including the quantity of affected devices shipped to facility is enclosed.Please complete the Recall Effectiveness Check Form and indicate if any devices from the affected lots remain in your inventory. This form needs to be completed even if no affected product is found.Please return the completed and signed Recall Effectiveness Check Form to your local Stryker Sustainability Sales Representative or email to SSSPFA@stryker.comor mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ, 85283Attn: Jodie RueckertIf the form indicates that affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Customers will receive a credit for all affected devices returned.Adverse reactions or quality problems experienced with the use of this productmay be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433X5555- http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online,by regular mail, or by fax.On 05/12/16 the firm sent out a clarification letter to emphasize that use of affected lots of product should be discontinued. In addition, SSS would like to note that, use of devices larger than expected may result in excess force applied to the blood vessel. The risk to the patient can vary from a prolongation in procedure time, vessel rupture, organ damage or other unforeseen consequences. Some of these injuries may require surgery or other procedures to remedy. Any questions contact the Stryker Sustainability Solutions Sales Representative or Regulatory Affairs. |  | 18 | 18 devices | Distributed in the state of NY and the country of Canada. | Open, Classified |  |  | 18 | Not_Computer
Z-1707-2016 | 73968 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery IGS 740. Indicated for use in generating fluoroscopic and rotational images of human anatomy. |  |  | April 15, 2016 | 2016 | GE Medical Systems, LLC | Potential non-recoverable loss of image acquisition. The affected Discovery systems may experience multiple X-Ray abort errors before or during a real-time fluoroscopic Interventional procedure. | Device Design | Consignees were sent on 4/15/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#12246 dated April 14, 2016. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 25 | 25 | Worldwide Distribution-US (nationwide) including the states of AL, CA, CO, CT, IL, IN, MD, MN, MS, MO, NE,NV,NH, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI, and the countries of ALGERIA, BELGIUM CANADA, CHINA, FRANCE, GERMANY, HUNGARY, IRELAND, ITALY, JAPAN, MEXICO, NETHERLANDS, PERU, SAUDI ARABIA, SINGAPORE, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 79 | Not_Computer
Z-1705-2016 | 74087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE System, for patient positioning, with assistance of a 30 Tracking System in a radiotherapy environment. |  |  | May 12, 2016 | 2016 | Elekta, Inc. | If by mistake the initial pre-treatment imaging was performed BEFORE the HexaPOD was moved to the DRIVE (*START) position, iGUIDE offers the possibility to branch into a specific catch-up workflow, which was implemented to use the initial scan data. | Software design | Elekta sent an Important Field Safety Notice that was distributed to all affected customers on May 12, 2016. The notice informs end users about the problem and the potential clinical impact. It also gives a recommendation how the risk can be mitigated. The notice also advises the customer that a patch will be released that will introduce a patch that avoids an inaccurate position and thus eliminates an unsafe situation.Corrective Action #2: Permanent Solution  Software UpgradeA software patch will be released that will correct the behavior of the iGUIDE software. The target release date for this patch is January 2017.Service teams will have 6 months from the date of release to correct all affected devices in the field. For further questions, call (770) 300-9725. |  | 20 | 20 | IL, LA, Austria, Australia, Botswana, Germany, Denmark, France, India, Japan | Open, Classified |  |  | 20 | Computer
Z-1720-2016 | 73949 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Accriva- APTT Cuvette for use on the Hemochron Jr. test system, catalog #J103, lot B6JCA012. |  |  | April 22, 2016 | 2016 | Accriva Diagnostics Inc., dba ITC, dba Accumetrics | Clot activating reagent mis-positioned or visibly absent in the cuvette's test channel causing APTT test to either have an error code or an erroneously high result. | Process design | Accriva sent an " Urgent Field Safety Notice" letter dated April 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to take the following actions:1. Forward communication to all those within organization who need to be aware of this matter. In addition, forward notification to all other organizations where affected devices may have been transferred to.2. Check inventory to determine if you have any APTT Test Cuvettes, LotNumber: B6JCA0123. If have inventory of APTT Test Cuvettes, with the subject Lot Number destroy the product and submit the attached Destruction Form along with your request for credit.4. If have inventory DO NOT SHIP TO CUSTOMERS any kits labeled with APTT Test Cuvettes and the subject Lot Number.5. If have shipped this product to customers please inform them immediately to STOP using the product. Provide the customers with a copy of this notification, request for the destruction of any unused product and the return of the attached forms.6. Follow the instructions on the attached Field Corrective Action Form, complete the form and return it as instructed on the form.7. If have questions, please contact Accriva Diagnostics directly contact information is provided below.Accriva Diagnostic, Attn: Regulatory AffairsUSA Tel: 858.263.2347Email: Fieldaction16-007@accriva.com |  | 135 | 135 boxes | U.S. distribution to the following; state: VA. Foreign distribution to the following; China, Italy, Germany. | Open, Classified |  |  | 135 | Not_Computer
Z-1813-2016 | 73984 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano s5, s5i, CORE and CORE Mobile systems with software version 3.4 and v3.4 software kits. |  |  | April 25, 2016 | 2016 | Volcano Corporation | Volcano has become aware of an incompatibility issue between Impacted Systems and hospital network scans. In specific circumstances, an impacted System will encounter unexpected data from the hospital network and be forced to reboot | Software in the Use Environment | The firm, Philips, sent an "URGENT MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter dated 4/25/2016 to its customers beginning April 29, 2016 which include a technical service bulletin to explain how to assure the incompatibility issue does not affect system performance. The customers were instructed to do the following:-If you have an Impacted System but it is not connected to a network, this issue does not affect your system. Additionally, if network scans are not performed, this issue does not affect your system.-If you have an Impacted System that is connected to a network and staff performs network scans while the Impacted System is in operation, this issue does affect you and the circumstances exist for a manual reboot to occur during an active procedure. To avoid this, the following corrections must promptly be taken: 1. Disconnect your Impacted Systems from the facilities network. If you need to reconnect it while not being used in a procedure, make sure it is disconnected again before starting any procedure; OR 2. Do not permit your staff to perform network scans while the Impacted Systems are in operation.-Please complete, sign, and return the attached form indicating that you received this Field Corrective Action notification via Fax to Volcano Customer Service at (916) 638-8812 or email to CS@volcanocorp.com.If you have any questions, please call (800) 228-4728, option 5. |  | 5875 | 5875 (2685 in US; 3190 ROW) | Worldwide Distribution: US (nationwide) and countries of: Japan, South Africa, Germany, Sweden, Belgium, France, Italy, Spain, India, Russian Federation, Hungary, Turkey, Poland, Portugal, Costa Rica, Finland, Great Britain, Croatia, India, Chile, Thailand, Mauritius, Iraq, Denmark, Norway, Ireland, Netherlands, Bulgaria, Czech Republic, Luxembourg, UAE, Austria, Poland, Pakistan, Sri Lanka, Kazakhstan, Slovakia, Saudi Arabia, Unite Arab Emirates, Egypt., Jordan, Kuwait, Lebanon, Bulgaria, Romania, Latvia, Lithuania, Bosnia and Herzegovina, Greece, Malta, Cyprus, Romania, Tunisia, Kenya, Algeria, Israel, Jordan, Hong Kong, Republic of Korea, Malaysia, New Zealand, Venezuela, Brazil, Philippines, Australia, Taiwan, Peru, Singapore, Panama, Mexico, Argentina, Costa Rica, Dominican Republic, China, and Paraguay, | Open, Classified |  |  | 5875 | Not_Computer
Z-1711-2016 | 73870 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SMR glenosphere impactor/extractor; Used to implant the glenosphere in the SMR Reverse Shoulder replacement. |  |  | April 08, 2016 | 2016 | Limacorporate S.p.A | Complaints of intra-operative breakage of the glenosphere impactors/extractors. | Component design/selection | The affected firms received written notification to all distributors. The letter instructed the consignee to identify and remove all affected product in stock and return parts to Lima USA. |  | 314 | 314 units | Distributed US (nationwide) and in Australia. | Open, Classified |  |  | 314 | Not_Computer
Z-1814-2016 | 74025 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | PerFuse Decompression Instrument PN: 800-0541, Lot:100650. Designed to access the femoral head for decompression. |  |  | April 11, 2016 | 2016 | Biomet, Inc. | The Trocar and Plunger Assemblies are missing from the instrument, which could result in a delay in surgery greater than 30 minutes. | Process control | Zimmer Biomet has initiated a voluntary recall of the PerFuse Decompression Instrument, following an investigation which identified that the Trocar and Plunger Assemblies are missing from the instrument. The firm contacted customers via telephone and e-mail on March 28, 2016; and official recall notices were mailed via FedEx on 4/11/2016. Customers were instructed to quarantine recalled items, Complete the Certificate of Acknowledgment and email a copy toCPWARFieldAction@zimmerbiomet.com; Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA: www.fda.gov/Safety/MedWatch/HowToReport/default.htm (form available to fax or mail), or Call (800)FDA-1088 .Questions related to this recall should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 22 | 22 | Distributed in the US states of CA, GA, MN, and OH, and in the Netherlands. | Terminated | May 30, 2017 | 414 | 22 | Not_Computer
Z-1812-2016 | 73959 | No | CAD | Anesthetic vaporizer. | Anesthesiology | Anesthesiology | 510(k) | Draeger D-Vapor/ D-Vapor 3000 Vaporizers with Baxter Desflurane Agent and Bottles. Catalog # M35500 (D-Vapor) and M36700 (D-Vapor 3000). Anesthesia Vaporizer |  |  | April 21, 2016 | 2016 | Draeger Medical, Inc. | Operators noticed unwanted release of evaporating anesthetic agent at the end of the filling procedure. Eye contact with volatile anesthetics can result in serious irritation. Immediate eye flushing and consultation of a physician is recommended. Updated Material Safety Data Sheets (MSDS) for Desflurane recommend wearing safety glasses when handling these substances. | Process change control | Drager mailed a "Medical Device Safety Alert" dated April 21, 2016, to customers informing them of the problem and provided instructions for the "correct filling procedure" and recommendations on how to prevent evaporation and spraying of Desflurane after bottle release to resolve the "High Friction of gas bottle valve" problem. |  | 16078 | 16078 units distributed within the United States | US nationwide distribution, including Puerto Rico and Guam. | Open, Classified |  |  | 16078 | Not_Computer
Z-1712-2016 | 73575 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Therapy Treatment Planning System Stand-alone Software 3.0, 3.5, 4.0, 4.5 and 4.7., designed for treatment planning and analysis of radiation therapy. |  |  | February 10, 2016 | 2016 | RAYSEARCH LABORATORIES AB | A software issue with editing tools that use the left mouse button held down, for drawing in and interacting with the patient views and beams eye views. The views can become unsynchronized with the stored data if simultaneously right clicking, pressing Ctrl-S or Ctrl-Z while the left mouse button is held down. This bug does not affect dose computations, which are based on the stored system state. | Other | On February 10, 2016, RaySearch Laboratories distributed Field Safety, Medical Device Correction notices and Reply Forms to their customers via email. To determine if the software version you are using is affected, open the About RayStation dialog in the RayStation application and check if the build number reported. This bug does not affect dose computations. Dose is always computed based on the stored system state. Actions to be taken by the user include the following: Always release the left mouse button before using other keys or mouse buttons, If the warning message "Temporary items should not be visible on approval" is displayed or if there is some suspicion that the bug may have triggered in structure definition, either review the structures in the Plan approval dialog, or close and reload the patient case and review the structures again. All customers should educate planning staff and all users about this workaround. All customers are advised to inspect product and identify all installed units with the referenced software version number, then confirm that you have read and understood the notice by completing and returning the reply form. The completed reply form should be returned to Freddie Cardel, freddie.cardel@raysearchlabs.com, fax 888-501-7195. Customers with questions can call 1-877-778-3849, available east coast business hours all US working days. |  | 1264 | 1,264 Total units (552 units domestically &amp; 711 units internationally) | **Domestic** AR, AZ, CA, CT, CO, DE, FL, GA, IL, LA, MI, MD, MS, MA, NY, NV, NJ, OH, OR, PA, NC, RI, TN, WA &amp; Washington D.C; **Internationally** Algeria, Australia, Austria, Belgium, Canada, China, France, French Polynesia, Germany, Iran, Italy, Japan, Netherlands, New Zealand, Norway, Republic of Korea, Spain, Sweden, Switzerland, Thailand &amp; United Kingdom | Terminated | April 24, 2017 | 439 | 1264 | Computer
Z-1829-2016 | 74031 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Streamline OCT Occipito-Cervico-Thoracic System, Crosslink Counter Torque Sleeve. Spinal fixation device instrument. |  |  | April 22, 2016 | 2016 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | Potential for a product defect where the cannula may be obstructed. This obstruction could prevent a driver from going through, which increases the potential for the driver to slip, resulting in possible patient injury. | Process design | RTI sent an Urgent Voluntary Product Notification dated April 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.RTI Surgical is conducting a voluntary recall of the Streamline OCT Occipito-Cervico-Thoracic System, Crosslink Counter Torque Sleeve due to the potential for a product defect where the cannula may be obstructed. This obstruction could prevent a driver from going through, which increases the potential for the driver to slip resulting in possible patient injury.For further questions, please call (906) 226-4489. |  | 17 | 17 devices | US nationwide distribution in CA, FL, IN, MI, MO, NC, NY, and TX. | Open, Classified |  |  | 17 | Not_Computer
Z-1830-2016 | 73703 | No | OWD |  |  | Molecular Genetics | PMA | cobasÂż EGFR Mutation Test, v2 and cobasÂż cfDNA Sample |  |  | March 15, 2016 | 2016 | Roche Molecular Systems, Inc. | An error was found within the Hungarian translations of the cobasÂż EGFR Mutation Test v2 Instructions for Use (M/N07340761001-01HU, Doc Rev. 1.0, Dated 08/2015) and the cobasÂż cfDNA Sample Preparation Kit Instructions for Use(M/N 07573758001-01HU, Doc. Rev. 1.0, Dated 05/2015). | Packaging | Roche Molecular Diagnostics Inc. issued an Urgent Field Safety Notice dated March 15, 2016 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Actions to be taken by the customer:" If you have utilized 2 ÂżL of DNA PBB during testing of plasma specimens, the samples must be retested using the correct volume (2mL DNA PBB).Communication of this Field Safety Corrective ActionThis notice must be passed on to all those who need to be aware within your organization or to any organization/individual where the potentially affected devices have been distributed/supplied.Please maintain awareness of this notice and resulting action for an appropriate period to ensure the effectiveness of the corrective action.The undersigned confirms that this notice has been notified to the appropriate Regulatory Agency.For questions regarding this recall call 908-253-7200. |  | 8 | 8 kits | Hungary | Open, Classified |  |  | 8 | Not_Computer
Z-1816-2016 | 74038 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Various trauma and sports medicine instruments and implants. |  |  | April 01, 2016 | 2016 | Biomet, Inc. | Insufficient seal strength on sterile barrier packaging.The seal could be jeopardized causing the sterile barrier to be compromised and the product could become nonsterile.If the product becomes non-sterile and is used or implanted, it may lead to an infection. | Equipment maintenance | Zimmer Biomet sent an Urgent Medical Device Recall Notice dated April 1, 2016, to all affected customers. Customers were instructed to locate and return the affected product, and implanting surgeons were instructed to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  |  |  | Worldwide Distribution - US (Nationwide) and globally to Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Ecaudor, The Netherlands, Japan, Korea, Malaysia, Mexico, Peru. | Terminated | May 23, 2017 | 417 |  | Not_Computer
Z-1784-2016 | 73734 | No | HBA | Neurosurgical paddie. | Neurology | Neurology | 510(k) | American Surgical Delicot 4 mm x 50 mm |  |  | March 31, 2016 | 2016 | American Surgical Company | Product packaging defective compromising sterility | Packaging process control | American Surgical issued an Urgent Recall and Field Safety Notice letterdated March 31, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were provided instruction to address affected devices in their inventory. They should address this by either selecting Action A or B: Action A: Users should perform a check of the device before use and, as stated in the IFU, patties should not be used if in an open or damaged pouch. For the affected lots, special attention should be given to lower portion of the pouch where the Tyvek and clear film are sealed.Action B: Place the affected lots in quarantine and contact American Surgical Company for an RMA and replacement. All affected devices are eligible for replacement by ASC.Complete and return confirmation of actions described in Field Safety Notice.Contact:Tel: 1-800-343-0060 Option 2, Fax: 781-595-5460Email: CustomerService@AmericanSurgical.com |  | 99 | 99 units | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Canada, France, Germany, Greece, Hong Kong, Hong Kong, Israel, Italy, Italy, Japan, Netherlands, New Zealand, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, and United Kingdom. | Open, Classified |  |  | 7701 | Not_Computer
Z-1831-2016 | 73923 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADD Administration Sets with Flow Stop, Product Reorder No: 21-7321-01, 21-7321-24, 21-7322-01, 21-7322-24, 21-7323-24, 21-7324-01, 21-7324-24, 21-7333-24, 21-7336-01, 21-7336-24, 21-7339-01, 21-7339-24, 21-7359-01, 21-7359-24, 21-7383-01, 21-7383-24, 21-7390-01, 21-7390-24, 21-7391-01, 21-7391-24, 21-7394-01, 21-7394-24, 21-7395-24. |  |  | April 22, 2016 | 2016 | Smiths Medical ASD, Inc. | Smiths Medical became aware that under delivery of medication may occur on CADD Administration Sets with Flow Stop. | Nonconforming Material/Component | Smiths Medical sent an "Urgent Medical Device Field Correction Action" letter dated April 20, 2016. The letter was addressed to "Clinicians who oversee the use of the CADDÂż Administration Sets with Flow Stop and Distributors thereof". The letter described the Affected devices, Reason for Field Correction Action, Risk to Health, and Instructions to Customers (Distributors and Clinicians). Requested consignees to complete and return the "Field Corrective Action Confirmation Form". For questions regarding this notification, contact Smiths Medical Customer Service Department at 1-866-831-8399. |  | 13696830 | 13,696,830 units | US: ALL STATES IN CONTINENTAL USA INCLUDING DC AND PR EXCEPT DE.OUS:UNITED ARAB EMIRATES, AUSTRIA BELGIUM, CANADA, CHINA, GERMANY, DENMARK, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, HONG KONG, IRELAND, INDIA, ITALY, JAPAN, LUXEMBOURG, MAURITIUS, NEW CALEDONIA, NEW ZEALAND, PANAMA, PORTUGAL, SAUDI ARABIA, SWEDEN, SINGAPORE, TAIWAN, VIET NAAM, SOUTH AFRICA. | Open, Classified |  |  | 13696830 | Not_Computer
Z-1835-2016 | 73377 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Video Duodenoscope |  |  | February 22, 2016 | 2016 | Pentax Medical Company | Pentax Medical learned of reports of carbapenem-resistant Enterobacteriaceae (CRE) infection at a medical facility. | Other | Pentax Inc. issued an "Urgent Field Correction" Letter dated February 19, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter provides information regarding the updated Reprocessing instruction for use along with some main changes in the Operation IFU. Instructions are also provided regarding prior versions of the Reprocessing and Operation Instructions for Use. A web posting went live on 2/19/2016.For further questions, please call:Tel:800-431-5880 ext. 2064 (8:30AM- 5:00 PM, Monday-Friday, EST) |  | 423 | 423 units | Nationwide | Open, Classified |  |  | 423 | Not_Computer
Z-1834-2016 | 73661 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pneumatic Impactor, General Surgery, With Microaire Hose Connector |  |  | March 11, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire discovered that 20 suspect Impactor instruments ( MicroAire Pneumatic Impactor) were shipped to end users and were missing the lock and unlock symbol markings on them. | Other | MicroAire sent an Urgent Medical Device Recall letter dated March 28, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.We request that you please respond to this letter promptly to determine how many of these units you have and will be returning to MicroAire. Upon return of the instrument to MicroAire, MicroAire will provide you with replacements.Please complete and send back to MicroAire page two of this letter, indicating that you do have this instrument at your facility and whether you will be returning the instrument for replacement.You can email the form to MicroAire Customer Service at notification@microaire.com or fax the form to 1-800-648-4309.You can contact MicroAire Customer Service at 1-800-722-0822 in the United States or 1-434-975-8000 outside the United States, between the hours of 8:30 AM and 5:30 PM (ET) Monday through Friday if you have any questions. |  | 20 | 20 units | United States | Open, Classified |  |  | 20 | Not_Computer
Z-1836-2016 | 74188 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SEAL,IS4000 PORTS,5MM-8MM,BOX of 10; Model numbers: 470361-03, 470361-04, 470361-05, 470361-06; Part number: 470361. The da Vinci Xi Surgical System cannula sterile single use disposable cap for each cannula . When attached to the cannula, the seal maintains insufflation when the cannula is empty and when the instrument is inserted in the patient during a surgical procedure. |  |  | May 09, 2016 | 2016 | Intuitive Surgical, Inc. | Potential for particulates in insufflation stopcocks on the da Vinci Xi Surgical System cannula seals. The particulates could be introduced into the cannulalumen when insufflation is connected during a surgical procedure. | Manufacturing material removal | Upon discovery of the issue, Intuitive sent Urgent Product Safety Notices to all customers by e-mail to initially alert them of the situation. This will be followed by a formal letter that is scheduled to arrive on May 23, 2016. Both notification request that users Do NOT use the da Vinci Xi 5 mm8 mm Universal Cannula Seal or the da Vinci Xi 12 mm &amp; Stapler Universal Cannula Seal for insufflation, or open the stopcock during intracardiac procedures. |  | 46 | 46, 191 boxes - total both sizes | Worldwide distribution. US, Australia, Austria, Belgium, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Monaco, Netherlands, Norway, Puerto Rico, Portugal, Qatar, Romania, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 46 | Not_Computer
Z-1842-2016 | 73123 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hip Preservation System Signature Series PreBent Burs under the following labels: 1) PreBent Spherical Bur, 5.5mm x 19cm Catalog Number HPS-HB11, 2) PreBent Polishing Bur, 5.5mm x 19cm Catalog Number HPS-HB12 and 3) PreBent Oval Bur, 6.0mm x 19cm Catalog Number HPS-HB13 |  |  | December 21, 2015 | 2015 | ConMed Corporation | Metal shavings released from burs during use are due to contact between the bur and the hood of the PreBent Bur assembly. During the manufacturing process an assembly error occurred which caused inadvertent mixing of two sizes of bearings. | Mixed-up of materials/components | ConMed sent an Urgent Medical Device Recall Notification letter dated December 17, 2015, to all affected customers. Customers are urged to stop use and return these devices immediately. Customers who have further distributed these items are urged to contact all of those organizations immediately and provide a copy of the recall notification letter. It is imperative that all end users of these devices receive the recall notice and respond immediately. Customers with questions can call 315-624-3237 or email largohpshb@conmed.com. The devices should be returned to ConMed Corporation NCR-47057 11155 Concept Blvd. Largo, FL 33773. For questions regarding this recall call 315- 624-3533. |  | 991 | 991 pieces | Worldwide Distribution - Nationwide to CA, CO, FL, IN, MD, MN, NC, NY, OH, TX, UT &amp; VA; International: Canada, Wales | Open, Classified |  |  | 991 | Not_Computer
Z-1848-2016 | 73427 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Kerlix" X-Ray Detectable Laparotomy Sponges, Tray Pack, 16" x 17" (40.6 cm x 43.2 cm) |  |  | March 03, 2016 | 2016 | Covidien LLC | Compromised sterility due to breach of barrier | Packaging process control | Covidien initiated recall by letter on March 4, 2016. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Any questions or concerns, contact your Medtronic representative or Customer Service at (800) 882-5878 |  | 114720 | 114720 | Distributed US (nationwide) and the countries of Canada, Austria, Belgium, Columbia, Croatia, France, Germany, Guadeloupe, Ireland, Italy, La Reunion, New Zealand, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain,Sweden and United Kingdom. | Terminated | June 15, 2017 | 469 | 8748710 | Not_Computer
Z-1860-2016 | 73469 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Endoscope; ED-530XT Operation Manual |  |  | December 23, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | This correction is in response to publicized reports of multi-drug resistant bacteria on endoscopes used for Endoscopic Retrograde Cholangiopancreatogram (ERCP) procedures. | Other | Fujifilm issued a urgent field correction/Tracking/Verification form on 12/23/2015. |  | 1473 | 1473 units | Nationwide Distribution. | Open, Classified |  |  | 3031 | Not_Computer
Z-1844-2016 | 74128 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | EZ Slide Battery/line powered hydraulic table with removable leg section and 21 top slide |  |  | April 06, 2016 | 2016 | Skytron, Div. The KMW Group, Inc | The potential exists within the identified tables that incorrectly manufactured parts have been used in assembly. If a table is found to have these incorrectly manufactured parts, it is possible for the tables elevation column slide cap to become detached from the elevation column and could result in the loss of control of the table top and injury to the patient or staff. | Reprocessing Controls | Skytron sent an Urgent Medical Device Recall letter dated A,pril 4, 2016, to all affected customers. On April 6, 2016, URGENT: Medical Device Recall notifications were sent to the affected hospitals and distributors via certified mail. The notification instructed the hospital to identify and remove the affected product from use. The notification instructs the distributor to contact the hospital to schedule replacement. The recall notification included a description of the reason for the recall, affected product, and consignee responsibilities. The notification did not include instructions for responding to the formal recall notification, however the company did request acknowledgement via follow-up email.Please call Technical Services, at 616-656-1187 or 616-656-6403 with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 458 | 458 | Nationwide Distribution to AL, AZ, CO, IA, WI, PA, TN, MA, TX, MI, FL, MS, LA, and CA. | Terminated | July 20, 2017 | 470 | 458 | Not_Computer
Z-1843-2016 | 73664 | No | LDH |  |  | General Hospital | 510(k) | AllergiEnd ST-9 Multiple Skin Test Applicator (Item 1000) and AllergiEnd ST-9 Multiple Well Test Tray (Item (1001). Contents: 8 applicators to perform up to 72 skin tests. |  |  | March 22, 2016 | 2016 | Medscience Inc | During an FDA inspection it was found that the products are marketed without a cleared 510k. | No Marketing Application | MedScience sent an Urgent Recall letter dated March 28, 2016, to their customers to inform them of the product recall. The letter instructed customers to discontinue and refrain from further use of recalled products and to contact MedScience to arrange the return of products still in possession by the consignee. MedScience set up a phone line and email address for customers to ask questions and obtain further direction: (800) 393-8817 and recall@medscienceinc.com. |  | 11400 | 11,400 (Item 1000) &amp; 380 (Item 1001) | Nationwide Distribution to AZ, CA, CO, CT, FL, GA, LA, MD, MI, NJ, NY, PA, TX, VA, and WA. | Open, Classified |  |  | 11400 | Not_Computer
Z-1865-2016 | 73826 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The ARCHITECT Estradiol Kit |  |  | March 24, 2016 | 2016 | Abbott Laboratories | Interaction of Fulvestrant with the ARCHITECT Estradiol assay, leading to falsely elevated estradiol results in patients treated with Fulvestrant. | Device Design | The firm sent a Product Correction letter dated 3/24/2016 to all affected consignees describing the nature of the recall as well as instructions to follow. The necessary actions included the following:Consignees are asked to review the letter with their Medical Director. Patients undergoing Fulvestrant therapy should not be tested with the ARCHITECT Estradiol assay. Complete the Customer Reply Form. If the consignee has forwarded the product listed to other laboratories, a copy of Product Correction letter should be provided to them.The customer is asked to retain the letter for their laboratory records.Customers with questions should contact either Customer Service at 1-877-4ABBOTT for the US; or, the local area customer service for outside of the US. |  | 39346 | 39,346 units | Worldwide Distribution -- US, to the following states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY; and, the countries of ALBANIA, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVI, BRAZIL, BRUNEI, BULGARIA, CAMEROON, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GAZA &amp; JERICHO, GEORGIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COAST, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MALDIVES, MALI, MAURITIUS, MEXICO, MONTENEGRO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REPUBLIC OF YEMEN, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TOGO, TRINIDAD &amp; TOBAGO, TUNISIA, TURKEY, UAE, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM and ZIMBABWE. | Open, Classified |  |  | 39346 | Not_Computer
Z-1867-2016 | 73684 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Integra MicroFrance Monopolar and Bipolar Electrosurgical Instruments |  |  | March 07, 2016 | 2016 | Integra LifeSciences Corp. | Reports relating to unintended patient or operator burns or operator electric shock for some of Integra's MicroFrance electrosurgical instruments. | Other | Integra sent an "Urgent Medical Device Correction" Field Safety Notice and an addendum dated March 7, 2016 to their customers via traceable courier service. |  | 2487 | 2,487 Instruments | AR, CA, CO, FL, IN, KY, MA, MI, NC, NJ, NY, NM, OH, TN, TX, WA and WI | Open, Classified |  |  | 32633 | Not_Computer
Z-1846-2016 | 73871 | No | NUJ | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Reprocessed Microline Scissor Tips intended to be used with a reusable hand piece and designed for minimally invasive and open surgical procedures to facilitate cutting, preparation, mobilization, and coagulation of tissue. Model #'s MIC3112, MIC3122, MIC3132, MIC3142, MIC3152, MIC3162. |  |  | April 26, 2016 | 2016 | STERILMED, INC. | Sterlilmed is recalling all unexpired Scissor Tips because of a potential manufacturing defect that can possibly result in a patient injury. | Process design | Consignees were sent on /26/2016 a Sterilmed "Urgent Medical Device Recall" letter dated April 26, 2016. The letter was addressed to either Materials Manager, Risk Manager and Operating Room Director or Distributor. The letter described the problem, the product involved in the recall and the actions required. Advised consignees to immediately quarantine and return the affected product and to complete and return the "Business Reply Form" to Sterilcycle.Advised Distributors to forward the letter to their customers if product was sent to them.For questions contact Sterilmed's Customer Care Team at 1-888-541-0078 . |  | 11326 | 11,326 | US: AL, AR, CA, CO, FL, GA, IN,IA, KY, MD, MA, MI, MN, MO, NV, NJ, NC, OH, OK, OR, PA,, SC, TN, TX, VA, WA, WV, WI.Canada | Terminated | April 05, 2017 | 344 | 11326 | Not_Computer
Z-1872-2016 | 73961 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-ray Systems |  |  | April 08, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm became aware of a problem where the system may not map Varian drive after CT user logout/login. | Software design | The firm sent a Customer Information letters dated March 8, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. For information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 158 | 158 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY. Product was also shipped to the following countries: Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, Colombia, Cuba, Estonia, Finland, France, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania,Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 1097 | Not_Computer
Z-1845-2016 | 73640 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA Duodenovideoscope Olympus TJF Type 160F/VF |  |  | March 15, 2016 | 2016 | Olympus Corporation of the Americas | New reprocessing instructions for the Olympus TJF-160/VF duodenoscopes, consisting of revised manual cleaning and high level disinfection procedures they plan to implement as well as a new cleaning brush (MAJ-1534), which is enclosed with the notification. | Under Investigation by firm | Olympus notified customers on 03/14/2016 via an Urgent Safety Notification letter that referenced discontinuation of previously distributed copies of the TJF-160F/VF Reprocessing Manual. Customers were asked to inspect their inventory and discard any existing inventory of TJF-160F/VF Reprocessing Manuals. They were asked to update their information with the enclosed TJF-160F/VF Reprocessing Manual, which contains new brushing and flushing steps for the TJF-160F/VF elevator mechanisms and the forceps elevator recesses. Instuctions for cleaning and a new cleaning brush (used for cleaning the device) was included with the notification. Customers were asked to ensure all reprocessing personnel are completely knowledgeable and throughly training on the new reprocessing instructions in the new Reprocessing Manual. |  |  |  | Nationwide: | Open, Classified |  |  |  | Not_Computer
Z-1839-2016 | 74139 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | COSEAL Surgical Sealant Kit, 4 mL, Product Code: 934071; For use in vascular reconstructions to achieve adjunctive hemostasis by mechanically sealing areas of leakage. |  |  | May 12, 2016 | 2016 | Baxter Healthcare Corp | Potential for incomplete dissolution of the polyethylene glycol (PEG) component during the reconstitution of the product, which may affect the consistency of the hydrogel formation during use. | Nonconforming Material/Component | An Urgent Product Recall letter dated 5/13/16 was sent to customers to inform them that Baxter Healthcare Corporation is issuing a voluntary product recall for the product codes and lots listed below due to the potential for incomplete dissolution of the polyethylene glycol (PEG) component during the reconstitution of the product, which may affect the consistency of the hydrogel formation during use. The letter provides the customers with the list of affected products, hazards involved, and actions to be taken. Customers with questions regarding the recall communication, are instructed to contact Baxter Product Surveillance at (800) 437-5176, 8-5pm, Monday-Friday. |  | 6804 | 6,804 units | Distributed US (nationwide) including Puerto Rico and in the Bahamas, Canada, and Singapore. | Terminated | March 23, 2017 | 315 | 13206 | Not_Computer
Z-1841-2016 | 74079 | No | GDZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Miltex Premium Grade Self-Lock Chisel Blade Handle, Catalogue No. 4-405, manual surgical instrument. |  |  | March 11, 2016 | 2016 | Integra York PA, Inc | One lot of self-lock chisel blade handles have a wider slot that holds the chisel blade that may result in the inability to firmly secure the blade in the handle. | Nonconforming Material/Component | Integra mailed out an Urgent: Voluntary Medical Device Recall letter to customers informing them to evaluate their inventory for the affected product. If product was shipped to customers, they were asked to identify their customers and to stop shipping immediately. Integra requested that customers complete the attached form that was sent with the notification indicating whether or not they had the affected product on hand and customer service will contact them to provide a Return Authorization number for the product's return. Customers could reach Customer Service at 866-854-8300. |  | 678 | 678 Chisel Blade Holders | Worldwide distribution. US nationwide, Mexico, and Korea. | Terminated | April 17, 2017 | 402 | 678 | Not_Computer
Z-1871-2016 | 73704 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ergon III Series Skyboom Monitor Bracket. Part # 3FCM1 Flatscreen Bracket, single monitor mount, 27inch max. screen size |  |  | March 28, 2016 | 2016 | Skytron, Div. The KMW Group, Inc | Potential for the weld to fail at the pivot point of the monitor bracket. This weld failure could cause the display interface and bracket to fall from its mounted location, resulting in injury. | Process design | Skytron LLC initiated a voluntary recall of twenty-four monitor brackets used as a mounting point for display interfaces utilized in clinical applications for viewing purposes and to provide an articulating point of movement on the ceiling mounted radial arm suspension system of the Ergon 3 equipment pendant /booms via telephone calls to customers, and written notices sent via certified USPS mail on 04/01/2016.All potentially affected users will receive a field correction notification directly from Skytron, LLC along with the contact information of their local Skytron distributor. Skytron distributors must contact identified users within the universe of affected products to schedule a time replace the affected monitor bracket(s).Skytron will provide a replacement of the monitor bracket. The replacement monitor bracket(s) and installation must be conducted by you, the Skytron distributor. The affected monitor bracket(s) and Service Report indicating that the monitor bracket(s) was replaced must be returned to Skytron.Please contact Larry Perez, Vice President of Technical Services, at 616-656-1187 or lperez@skytron.us, or Thao Selleck, Complaint Resolution Analyst, at 616-656-6403 or tselleck@skytron.us with any questions. For after-hours support, please call 1-800-SKYTRON. |  | 24 | 24 | Worldwide distribution. US states: VA, VT, UT, GA, NJ, and IL; Japan. | Terminated | July 25, 2017 | 484 | 24 | Not_Computer
Z-1883-2016 | 73948 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Torque Limiting Screwdriver: |  |  | April 26, 2016 | 2016 | Greatbatch Medical | Greatbatch Medical has initiated a global Field Recallfor Torque Limiting Devices. Greatbatch identified through laboratory testing that the required sterility assurance level (SAL) of 10-6 is notachieved on torque limiting devices when subjected to steam sterilization cycles identified in the InstructionFor Use (IFU) provided with the device. | Device Design | Consignees were sent on 4/26/2016 a Greatbatch "Urgent Medical Device Recall" letter dated April 26, 2016. The letter described the product involved in the recall, the reason, and the risk to health. The letter advised consignees to identity and quarantine the product and to complete and return the Filed Recall Response Form. For questions contact your local Greatbatch Medical representative, at 1-763-951-8235, or e-mail FieldActionCenter@Greatbatch.com. |  | 1110 | 1110 | US: IN, MI, TX.OUS: GERMANY, FRANCE, ITALY, SWITZERLAND. | Open, Classified |  |  | 14128 | Not_Computer
Z-1881-2016 | 73978 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Firebird NXG Spinal Fixation System Set Screw Driver used to insert a set screw into a spinal fixation construct. Rx Only Provided non Sterile. |  |  | April 27, 2016 | 2016 | Orthofix, Inc | There is a possibility that the subject set screw driver could disengage from the torque limiting handle when the counter torque wrench is allowed to contact the counter torque handle. | Device Design | Affected distributors will be provided a written notification of field correction via USPS Certified Mail or FedEx on 4/27/16. Distributors will be additionally instructed to provide field correction notifcations to hospitals, surgeons, or other distributors that they may have further ditributed the product to. |  | 36 | 36 units | KS, NV, TX, CA, ID, FL, MA, MI, NJ, AZ, NH, SC, UT, WA | Terminated | November 23, 2016 | 210 | 36 | Not_Computer
Z-1880-2016 | 73766 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | STAT-Check Resuscitation Bag |  |  | March 15, 2016 | 2016 | Ventlab LLC | Stat-check CO2 indicators are non-functional. The indicator is yellow when not in use and should be purple when not in use. The yellow indicates a 5% CO2 level, cardiac massage and ventiliation being performed correctly. The color will not change during use and will not be an indicator of CO2 level or improper performance of cardiac message and ventilation. | Process control | Ventlab sent an URGENT: MEDICAL DEVICE PRODUCT REMOVAL notification dated March 14, 2016, on March 15, 2016, electronically to all affected customers. Customers were instructed to return the affected product and notify affected consignees of the recall. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this action, please call Ventlab at 1-844-635-5326, or email at PFA@ventlab.com. |  | 2430 | 2430 | Nationwide distribution to AZ, MS, NY, MO, FL, WI, TX, FL, and CA. | Open, Classified |  |  | 2430 | Not_Computer
Z-1876-2016 | 73802 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastr Liquid Mouse Monoclonal Antibody Renal Cell Carcinoma Marker (RCC) is recommended for use as part of a panel of antibodies in the definition of tumors of proximal nephrogenic differentiation. |  |  | April 05, 2016 | 2016 | Leica Microsystems, Inc. | Testing of retained units indicated that Lot numbers of CDX2 and RCC indicated are not stable up to the expiry date on the product labeling. | Nonconforming Material/Component | An Urgent: Medical Device Recall Notification dated 4/5/2016 was sent to consignees with the affected lots of CDX2 and RCC marker. The customers were notified of the problems as well as the following actions to be taken by the consignees: Do not use or continue to use the reagent lots listed above, as they have the potential to result in inadequate tissue staining and lead to weak or false negative pathology results.Appropriately destroy any unused or partially used affected lots of the reagentÂż Complete the attached Recall Notification Acknowledgement Form that this action has been undertaken.The firm also request that the notification is passed to all those within their organization who need to be aware of this issue.The firm requests consignees to return the completed acknowledgement form to LMGRA@Leica-Microsystems.com to receive credit for affected lots that have been destroyed by the consignees laboratory. |  | 270 | 270 units | Distributed in the following states AR, CA, CT, FL, ID, IL, IN, MT, OH, PA, TN and TX, and in Canada. | Open, Classified |  |  | 720 | Not_Computer
Z-1885-2016 | 73885 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 Intelligent Lab System running on software versions V.3.6.1 SP1 or V.3.6.2, Device listing # D011374, All serial numbers and lots are affected. |  |  | March 24, 2016 | 2016 | Siemens Healthcare Diagnostics Inc. | Discrepant patient results on Dimension Vista Intelligent Lab Systems. Siemens Healthcare confirmed a software defect which, in a very specific set of circumstances, results in the Dimension Vista System omitting an aliquot probe rinse between sample aspirations when processing tubes in Sample Racks that are front loaded on the Dimension Vista System. | Software design | Siemens sent an Urgent Medical Device Correction letter dated March 24, 2016, to all affected customers. Customers were advised to follow the instructions contained in the "Urgent Medical Device Correction" notification letter to help minimize or eliminate the impact that this specific software defect could cause until a software correction could be made (i.e instructions to minimize the impact of the software defect due to "Bad Sample Detect " False Transition" errors when processing front loaded samples) and (how to eliminate the impact of the software defect when processing front loaded samples). Customers were also asked to review the recall notification letter with their Medical Director and complete and return the Field Correction Effectiveness Check Form within 30 days. Customers with questions were advised to contact Siemens Customer Care Center - Technical Solutions at 800-441-9250 or their local Siemens technical support representative.For questions regarding this recall call 302-631-8248. |  | 1301 | 1301 units | Nationwide Distribution to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, , ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico. | Terminated | March 13, 2017 | 354 | 5202 | Computer
Z-1878-2016 | 74004 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | Replacement HEPA filter for Natus Sterile-Drier Model 43 and Model 44. Filter Catalogue Part Number: 54343, 54344, and 100105-01. |  |  | April 22, 2016 | 2016 | Natus Medical Incorporated | The replacement HEPA Filters have been manufactured with sealant that may not be properly cured. Improperly cured sealant can emit a toxic gas when heated, and can drip down during use, contaminating the interior chamber of the Sterile Drier and its contents. | Nonconforming Material/Component | Natus sent an " URGENT RECALL NOTIFICATION/ FIELD SAFETY NOTICE" Olympic Sterile Driers Model 43 &amp; 44, Replacement HEPA Filter Contamination letter (dated April 22, 2016) was sent to all affected customers on May 2, 2016. for the first set of customers and then sent on May 12, 2016, for the second set of customers, who purchased 57 replacement filters. Customers are advised "do not install any replacement HEPA filter " if the serial number is one of the 21 shown as affected serial numbers on the letter. The URGENT RECALL NOTIFICATION/ FIELD SAFETY NOTICE Olympic Sterile Driers Model 43 &amp; 44, Potential HEPA Filter Contamination (dated May, 2016 and released on May 12, 2016,) was sent to customers on May 16, 2016, who purchased 11 sterile driers with the installed filters. Customers are advised "do not use the Sterile drier " if it contains a HEPA filter that is identified with one of 21 affected serial numbers. Customers are advised not to install any defective replacement HEPA filters or do not use sterile drier if it currently contains defective HEPA filter. Please discontinue use of the Sterile Drier immediately if it contains defective HEPA filter as heat may cause the uncured adhesive to drip into the chamber.Customers contact Natus to obtain the RMA number and return unused filters and Natus will provide a replacement filter.Please contact Natus at Seattle_Technical_Service@natus.com or call 866-940-7143 (option 1, Olympic products) for next steps regarding a contaminated Sterile Drier. |  | 21 | 21 filters | distributed in the US including AZ, CA, FL, IL, KS, MA, MD, MN, NC, NH, NY, OH, OK, PA, VA, WA, WA DC,and WI and in Canada and Singapore. | Open, Classified |  |  | 21 | Not_Computer
Z-1870-2016 | 73743 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | General Surgery Pack, part number 006880-3. This is a custom convenience surgical instrument disposable kit. |  |  | March 21, 2016 | 2016 | Windstone Medical Packaging, Inc. | These packs contain item #4037 Kerlix X-ray Detectable Laparotomy Sponges which were recalled by Medtronic due to the potential for sterile barrier to be comprised at the packaging seal. | Packaging | Aligned Medical Solutions (AMS) sent the Urgent Medical Device Recall letter, dated March 21, 2016, and the Recall Reply Form to their consignees via email.AMS informed consignees that their AMS packs (surgical kits) contained Kerlix X-Ray Detectable Laparotomy Sponge which was recalled by Medtronic.Consignees were advised to follow the over-labeling instructions and dispose the Kerlix X-Ray Detectable Laparotomy Sponge at the time the surgical kit is opened for use. Consigneeswere instructed to follow the "Instruction for placing Recall Notice Label to Packs" and label any kits remaining inventory.Consignees should return the Reply Form via fax at 407-865-7456 or call 321-527-7714 for any questions about this recall. |  | 170 | 170 units | Nationwide distribution to Louisiana and Montana. | Open, Classified |  |  | 1192 | Not_Computer
Z-1612-2016 | 73825 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator |  |  | April 14, 2016 | 2016 | Philips Electronics North America Corporation | MRx monitor/defibrillators could fail to charge because the therapy printed circuit board may have been loaded with an incorrect electronic component. | Employee error | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 2016, on April 14, 2016, to all affected customers. The letter informed customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of providing a replacement MRx unit, free of charge, for all units affected by this issue. When the replacement unit arrives, please use the return shipping label and packaging provided with the replacement MRx unit to return the affected unit to Philips. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: You can continue to use your MRx prior to receiving a replacement unit. But, if you observe a red X in the Ready for Use (RFU) indicator, you should follow the troubleshooting instructions in the MRx IFU.Customers with questions were instructed to contact their local Philips representative or call 1-800-722-9377. |  | 17 | 17 | Worldwide Distribution - US including OH, CA, NC and Internationally to China, Iraq, Japan, and Korea. | Open, Classified |  |  | 17 | Not_Computer
Z-1906-2016 | 73986 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Non-Vented Dispensing Pin with One-Way Valve and Luer Lock Ultrasite Valve |  |  | April 14, 2016 | 2016 | B. Braun Medical, Inc. | BBMI recently identified a potential for a tear in the outer blister packaging on a limited number of units. The tears have been identified in the blister package paper, where the unit label writing is located. | Process control | B. Braun notified customers via letter to inform them of the reason for the recall, the risks and affected lots involved were identified and customers were asked to discontinue use of the product and complete the Product Removal Acknowledgement form and return to B. Braun Medical Inc. Quality Assurance department by faxing it to (610)849-1197 or e-mail to: PA_QualityAssurance.BBMUS_Service@bbraun.com and a BBMI Customer Support Representative will contact the consignee to provide instructions for handling the affected product and arrange for its return to BBMI. |  | 11000 | Approximately 11,000 units available in commercial channels | US Nationwide Distribution in the states of: AL, AR, AZ, BC, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN. TX, UT, VA, VT, WA, WI, WV and Puerto Rico | Terminated | March 13, 2017 | 333 | 11000 | Not_Computer
Z-1900-2016 | 73200 | No | MRA |  |  | Orthopedic | PMA | Novation Ceramic AHS Cluster-Hole Shell, Plasma Coated, 54mm, O.D. (Acetabular Cup), Catalog Number 140-01-54; For use in primary total hip arthroplasty. |  |  | April 12, 2016 | 2016 | Exactech, Inc. | Devices were inadvertently sent to packaging and finished goods prior to the PMA site-change supplement approval. | Packaging process control | Customers were notified 4/12/2016 with a Product Recall Advisory Notice of affected products.In order to comply with applicable regulations and Exactech policies: " Immediately cease distribution or use of these products. " Extend this information to your accounts that may have this product in their possession. " Identify and quarantine any of the subject devices in your inventory " Complete and Return the attached Recall Inventory Response Form to Exactech, per the instructions provided on the notice itself. According to policy, please contact your Exactech inventory representative within 5 business days to confirm quantities at your location. If you have any inventory restocking questions related to this issue, please contact Kaya Davis at kaya@exac.com or 1-800-392-2832. Please return the products to the following address:Exactech Distribution Center 2411 NW 66th Court Gainesville, Florida 32653Attention: Kaya Davis kaya@exac.com 352-337-3915 (Fax) |  | 3 | 3 devices | Distributed in FL, NJ, NY, OH, OK, RI, TN, VA and HI. | Open, Classified |  |  | 20 | Not_Computer
Z-1879-2016 | 74164 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | May 18, 2016 | 2016 | Elekta, Inc. | When DICOM exporting a 3D Monaco plan and the "Composite Field Sequencing" (CFS) checkbox is selected, the Y jaws will snap to the port shape when they should remain where they were defined. | Software design | Field Safety Notice 382-01-MON-004 was sent to customers on May 18, 2016. The notice informs users of the specific product and version numbers affected by both issues, and any work around that can be used to avoid the issue. The customers were instructed to complete and return the Field Safety Notice Acknowledgement form to their local Elekta as soon as possible. |  | 1401 | 1401 units | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, FL, GA, IL, IN, KY, MI, MN, MO, MT, NV, NH, NJ, NY, NC, OK, OR, PA, SD, TX, VA, WA, WI, and the countries of: Alger, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Belarus, Botswana, Brazil, Bulgaria, Canada, China, Columbia, Croatia, Cyprus, Czech Republic, Ecuador Egypt, Estonia, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Morocco, Nabibia, Netherlands, New Zealand, Norway, Panama, Peru, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka,Sweden, Switzerland, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, Zimbabwe | Open, Classified |  |  | 1401 | Not_Computer
Z-1691-2016 | 74106 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Sedecal SA Mobile Diagnost w DR x-ray system |  |  | March 23, 2015 | 2015 | Sedecal USA, Inc. | Due to a software defect, the system may sporadically apply the default x ray exposure parameters for an adult ( patient type : Normal ) even though the patient type " Newborn" was selected and is displayed in the generator control area of the Eleva User Interface. | Radiation Control for Health and Safety Act | Sedecal USA has developed a new software release to address this issue. The distributor of this product, Philip, will modify the systems in the field to correct the defect at no cost to the customer.1. You will contact customers and initiate a software update to correct the two defects.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (847) 394-6960. |  | 145 | US - 145 | US Distribution - TX, OH, NY, MD, IN, PA, DE, WA, NH, AZ, GA, NC, UT, VT, CA, LA, NJ, MN, FL, WI, HI, IA, MA, ME, TN, CT, ID, MI, AR and KS. | Open, Classified |  |  | 145 | Computer
Z-1833-2016 | 73799 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ is a dual-line volumetric infusion system designed to meet the growing demand for hospital-wide, as well as alternate site and home healthcare, standardization. With its primary line, secondary line, and piggyback fluid delivery capability, the Plum A+ is suited for a wide range of medical/surgical and critical care applications. Full compatibility with LifeCare Plum Series administration sets and accessories. |  |  | March 25, 2016 | 2016 | Hospira Inc. | Firm's investigation determined that the Proximal Occlusion Alarms observed in the field were attributed to the mechanism calibration process being performed with an out of specification metal cassette, resulting in an offset in proximal pressure. | Process control | Hospira Inc, sent an "URGENT MEDICAL DEVICE CORRECTION" letter on March 31, 2016, to the direct consignees of record via traceable mailing which included a data capture log containing a list of devices located with each customer. Delivery will be confirmed via traceable mailing. Consignees who further distributed the product were asked to notify their customers. Please take the following actions related to Urgent Medical Device Correction letter notification: 1. Inform potential users in your organization of this notification.2. Complete the attached reply form and return it to the fax number or e-mailaddress on the form, even if you do not have the affected product.3. If you have distributed the product further, notify your accounts that receivedthe product identified above of this notification and ask them to contactStericycle at 1-877-479-8088 (M-F, 8am-5pm ET) to receive a reply form.Hospira will contact you to make arrangements for your impacted Plum 360 I Plum A+ devices to be re-calibrated to correct this issue.: |  | 2567 | 2567 Plum A+ infusion pumps. | U.S. distribution : Natrionwide. Foreign :Brazil, Canada, Chile, Colombia, Ecuador, Guatemala, Mexico, Oman, Peru, Taiwan, Turkey, United Kingdom and Uruguay. | Open, Classified |  |  | 4057 | Not_Computer
Z-1909-2016 | 73979 | No | PFX |  | Immunology | Immunology | 510(k) | CytoScanÂż Dx Assay. The CytoScan Dx Assay consists of five reagent modules, wash buffers (WS), a microarray kit, and analysis software. |  |  | April 14, 2016 | 2016 | Affymetrix Inc | During a recent Ohio site inspection by the FDA it was discovered that warning statements needed to be included with the product shipment. | Labeling design | A letter dated April 14, 2016 was sent to customers. The letter stated the reason for the recall. It also informed customers that there is no impact on the performance of the assay. Customers were asked to either visit the IVD web portal (www.affymetrix.com/ivd) and review the warning statements; or, request a hard copy free of charge. |  | 507 | 507 kits | Distributed in the US to the states of AR, CA, DC, MI, MA, MN, NC, NM, NY, PA, SC, TN, TX, WA and WV. | Open, Classified |  |  | 507 | Not_Computer
Z-1911-2016 | 74140 | No | LHJ | Staphylococcus aureus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Sure Vue Color Staph 150T, Sure Vue Color Staph 300T, Sure Vue SELECT Staph, biokit SELECT Staph 75T |  |  | February 19, 2016 | 2016 | Integrated Science Systems Inc | Customer complaints of very weak reactions with the kit, or false negative results with different Staph aureus cultures. | Process control | Integrated Science Systems sent a Recall Notification letter dated March 29, 2016. Integrated Science Systems requested that unopened kits in inventory be returned. Once the root cause has been determined, the kits will be reworked and returned to INOVA for distribution. For questions call 706-922-8686. |  | 2625 | 2625 kits | Nationwide distribution to CA only. | Open, Classified |  |  | 2625 | Not_Computer
Z-1695-2016 | 74083 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare DigitalDiagnost stationary X-ray system |  |  | November 10, 2015 | 2015 | Philips Healthcare | The detector may signal that it is ready for acquisition when it actually is not, resulting in failure to properly acquire the X-ray image. | Radiation Control for Health and Safety Act | Philips Healthcare plans to correct the defect at no cost to customers by implementing a field correction involving the installation of new software.1. You will contact customers and initiate a software update to correct the defect.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call 1-800- 722-9377. |  | 160 | 160 | USA (nationwide) Distribution to the states of : ( DuraDiagnost) to the states of : OH, CA, WA and NE.(DigitalDiagnost) to the states of : AK, AZ, AL, CA, CT, CO, DE, FL, GA, HI, ID, KY, LA, MA, MN, MD, MS, ME, MO, NE, NY, ND, OH and PA. | Terminated | May 03, 2017 | 540 | 165 | Computer
Z-1910-2016 | 74070 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker FlowPort II Adapter: Model number: 00CAT00778 |  |  | May 04, 2016 | 2016 | Stryker Corporation | Complaints were received for the Stryker FlowPort II Adapter, and investigation found that the scope sat in the adapter 180 degrees in the wrong direction | Device Design | A recall letter was not sent as all affected devices have been returned. A ship hold was placed on the device on October 16, 2015. |  | 21 | 21 units | GA, UT, MD, NJ, MT | Terminated | June 03, 2016 | 30 | 21 | Not_Computer
Z-1715-2016 | 74149 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Brilliance iCT SP Computed Tomography X-Ray System |  |  | September 06, 2015 | 2015 | Philips Healthcare | Software Defects resulting in: (1) sagittal result shortened for axial scans; (2) single series displayed on console viewer when Creating multi-planner reconstruction; (3) scan length changes with a change in field of view; (4) DoseRight algorithm miscalculations leading to incorrect CTDI values; (5) scan length changes during subsequent axial results, (6) surview scan lengths near 135mm or 184mm | Radiation Control for Health and Safety Act | Philips plans to initiate a field corrective action that will install a new software release to address the issue. Philips will modify the systems in the field to correct the defect at no cost to the customer.1. You will contact customers and initiate a software update to correct the defect.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (1-800) 722-9377, Option 5 |  | 423 | US-423 | US Distribution - AZ, PA, FL, DC, W VA, NC, SC, AR, OK, OH, NY, NV, TX, PA, IN, CA, MN, MA, DC, MS, MD, OR, IL, CO, MO, VT, IA, AL, LA, GA and WA. | Open, Classified |  |  | 423 | Computer
Z-1912-2016 | 74080 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Synthes 8.0 mm Flexible Shaft Reamer, 360 mm, Orthopedic manual surgical instrument. Part number 351.43 |  |  | May 02, 2016 | 2016 | Synthes (USA) Products LLC | Devices did not pass the biological safety evaluation for cytotoxicity following exposure to test conditions. The high growth inhibition levels observed during testing could be attributed to corrosion of the device at solder points. This could potentially be reproduced during use and reprocessing. | Process design | Depuy Synthes sent an Urgent Notice: Medical Device Recall letter on 05/02/2016 to all affected consignees. Consignees were instructed to do the following: review your inventory immediately to see if any of the affected products were being stored, and call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement order, to complete the Verification Section (page 3) and return with the product to Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. |  | 244 | 244 units | US nationwide distribution. | Open, Classified |  |  | 244 | Not_Computer
Z-1915-2016 | 74196 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Revolution CT |  |  | May 09, 2016 | 2016 | GE Healthcare, LLC | GEHC has become aware of a potential safety issue in the Revolution CT Patient Table that can result in unintended motion in cases involving large patients. The Revolution CT cradle can descend on its own after activation by the user. This can put the patient at risk for harm. The harm can involve the impinging or crushing of a hand if caught between the cradle and CT inner gantry bore. No injurie | Process control | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter ,GEHC Ref# 25468, dated May 9, 2016 to its Consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions; continue to use your Revolution CT product; follow the Patient Positioning warning outlined in the Revolution CT User Manual; to prevent pinching or crushing of the patients extremities, keep the patients hands and feet away from the edge of the moving table top/cradle and its surrounding equipment, or between table base and side panels of the table; and take special care when positioning physically large patients. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 197 | 197 (US = 68; OUS = 129) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CT, FL, GA, HI, IL, IN, KS, NE, NY,NC, OH, OR, PA, TN, TX, UT, A, WA, WI; and countries of: Australia, Belgium, Brazil, Canada, China, Denmark, Estonia, France, Germany, Hong Kong, Indonesia, Italy, Japan, Korea (Republic of), Kuwait, Mexico, New Zealand, Norway, Peru, Poland, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 197 | Not_Computer
Z-1914-2016 | 73887 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System, Catalog Number 6802413, Unique Device Identifier No. 10758750002740; and VITROS 5600 Integrated System (refurbished), Catalog Number 6802915, Unique Device Identifier No. 10758750007110; IVD. |  |  | April 06, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics confirmed a software timing anomaly in which the VITROS 3600 and 5600 Systems may unexpectedly aspirate from or dispense into an unintended sample that may lead to erroneous or contaminated results. | Software design | On 4/6/2016,URGENT PRODUCT CORRECTION NOTIFICATION letters (Ref. CL2016-076, dated 4/06/2016) were sent via FedEx overnight and/or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. Follow enclosed instructions to help decrease the probability of the software anomaly until the software update is installed. Discuss any concerns with previously report results with their Laboratory Medical Director to determine the appropriate course of action. Install Software Version 3.2.3 at their earliest convenience. Foreign affiliates were informed of the issue by e-mail on 4/6/2016, and instructed to notify their consignees of the issue and required actions. |  | 7881 | VITROS 5600: US: 1161 units, Foreign: 1063 units; VITROS 5600 Refurbished: US: 19 units, Foreign: 38 units | Nationwide, Puerto Rico and Foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 8686 | Computer
Z-1921-2016 | 73658 | No | JTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Oxoid Nitrocefin (+ Reconstitution Fluid Batch) REF: SR0112C, Oxoid Ltd. For the rapid chromogenic detection of beta-lactamase activity in bacterial cells. in vitro diagnostic |  |  | March 24, 2016 | 2016 | Remel Inc | Individual bottles of product may produce weak beta-lactamase reactions. Use may result in delayed or false negative reporting. | Under Investigation by firm | Customers were notified by letter beginning on March 24, 2016. |  | 2 | 2 units were imported and distributed domestically | US distribution in MO. | Open, Classified |  |  | 2 | Not_Computer
Z-1920-2016 | 74053 | No | JEY | Bone plate. | Dental | Dental | 510(k) | 1.5 mm Rapid Resorb Orbt1 F1 P1 w/Bending Template / 35 mm - Sterile, Rapid Resorbable Fixation System. 1.5 mm Rapid Resorbable Orbital Floor Plate with Bending Template. |  |  | May 02, 2016 | 2016 | Synthes (USA) Products LLC | Labeling errors. The outer packge label expiration date may exceed the expiration date of the subcomponents. | Under Investigation by firm | DePuy Synthes sent an Urgent Medical Device Recall Notice dated May 2, 2016, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory immediately and remove the affected lots from their stock. Customers were instructed to call 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product., customers were also asked to complete the Verification Section and scan/email to Synthes4096@stericycle.com. Customers with questions were instructed to call 610-719-5450. |  | 14 | 14 units | Nationwide Distribution to DE, NC, CA, OK, PA, SC, WI, CO, NJ, TX, UT, IL, MO, MI, AL, KS. | Open, Classified |  |  | 55 | Not_Computer
Z-1917-2016 | 74148 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Trabecular Metal Tibia; Persona Trabecular Metal Tibial Plate Instruments and Modular Brackets; Used to facilitate the implantation of joint replacement implants. |  |  | May 11, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Inc. is initiating a recall of specific Persona" Trabecular Metal" Tibial Plate Instruments and Modular Brackets, due to incidents of radiolucent lines and loosening in implants. | Component design/selection | Zimmer Inc. is initiating a voluntary recall of Persona" Trabecular Metal" Tibia / Tibial Plate Instruments and Modular Brackets following an increase in complaints of radiolucent lines and loosening, and earlier recall of the Persona" Trabecular Metal" Tibial Plates . All sizes and lots of the affecteddevices are being removed from distribution. Recall notices were issued to affected distributors, hospitals, and surgeons on 05/11/2015 via mail. Customers may contact the following with any questions: 1-877-946-2761 between 8:00 am and 5:00pm EST. Customers are asked to do the following: 1. Review the notification and ensure affected personnel are aware of the contents.2. Assist Zimmer Biomet sales representative with the quarantine of any affected product.3. Zimmer Biomet sales representative will remove the recalled product from customer facilities.4. Complete the Certificate of Acknowledgement Form and return to: corporatequality.postmarket@zimmerbiomet.com. |  | 10256 | 10,256 | Worldwide Distribution -- US, including the states of AK, AL, AZ, CA, FL, GA, HI, ID, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI; Foreign: Australia, New Zealand, Korea,Austria, Belgium, Germany, Italy, France, Luxembourg, Netherlands, South Africa, Spain, Sweden, Switzerland, and United Arab Emirates. | Terminated | June 02, 2017 | 387 | 10256 | Not_Computer
Z-1922-2016 | 74125 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Actifuse ABX/Actifuse MIS System |  |  | August 18, 2015 | 2015 | Apatech Limited | Baxter Healthcare is issuing a voluntary recall for all lots of Actifuse ABX and Actifuse MIS System products with expiry before July 29, 2017 due to the possibility that the products may have endotoxin levels above specification criteria. | Process control | The firm, Baxter, sent an "Urgent Product Recall Notification" letter dated 8/18/2015 and a follow-up letter date 4/29/2016 to consignees via FedEx, overnight mail. Baxter also notified affected customers by phone. The letters described the product, problem and actions to be take. Baxter Healthcare has instructed customers to do the following: 1. Locate and remove all affected product from their facilities.2. Contact Baxter Healthcare Center for Service to arrange for return and credit at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time, Monday through Friday.3. Complete the attached customer reply form and return it to Baxter by either fax it to 224-270-5457 or scanned e-mail to fca@baxter.com.4. If you distribute this product to other facilities or departments with your institution, notify all affected parties of this communication.For general questions regarding this communication, please call Baxter's Therapeutic Medical Affairs at 805-657-0007 or 862-432-7941. |  | 11696 | 11696 units | Worldwide Distribution: United States: Nationwide (including Puerto Rico) and countries of: Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, Great Britain, Ireland, Italy, The Netherlands, Norway, Sweden, Slovenia, Czech Republic, Kuwait, Mexico, Brazil, Australia, Taiwan, and New Zealand. | Open, Classified |  |  | 11696 | Not_Computer
Z-1924-2016 | 74176 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Tearaway Introducer, Model # VS203, VS303, 510 K # 130687, packaged individually in a pouch, 5 pouches per carton, lot # MBZ140, MBZL450, MBZV930 MBZZ490 |  |  | May 05, 2016 | 2016 | Medical Components, Inc dba MedComp | This recall has been initiated due to the product labeled with the incorrect expiration date. | Labeling mix-ups | MedComp sent a Product Recall letter dated May 5, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to evaluate their inventory and quarantine for the return of all un-used affected product. They were asked to contact Customer Service for a Return Goods Authorization (RGA) number at 215-256-9191. |  | 3122 | VS203 Lot# MBZX140 (100 units); VS303 Lot # MBZL450 (60 units); VS303 Lot # MBZV930 (65 units); VS303 Lot # MBZZ490 (33 units). | US Distributed to: FL, TX, NC | Terminated | June 16, 2017 | 407 | 3122 | Not_Computer
Z-1926-2016 | 74059 | No | LGW |  |  | Neurology | PMA | Medtronic RestoreSensor, Model 37714; and, RestoreSensor SureScan, (MRI), Model 97714, Implantable Neurostimulators. |  |  | May 03, 2016 | 2016 | Medtronic Neuromodulation | Loss of therapy during recharging of device. Reportedly, a charging session was terminated prior to obtaining recharge threshold voltage of 3.615, which triggered rapid battery depletion state. As a result, the device depleted to 1.925 volts in one to two days rather than the typical 30 days. Insufficient coupling between recharger and device during recharge session was found to be key factor. | Under Investigation by firm | Consignees were sent on 5/3/2016 a Medtronic "Urgent Medical Device Safety notification" letter dated May 2016. The letter was addressed to "Healthcare Professional". The letter described the problem and the product involved in the recall. The letter listed the recommendation to patients to follow current recharge instructions, paying particular attention to Charging Efficiency and Battery Charge Level indicators on the recharger. Consignees were also requested to complete and return the "Physician Reply Form:.For questions contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekdays 7 am - 6 pm CST. |  | 60941 | 60,941 | Worldwide Distribution | Terminated | August 04, 2017 | 458 | 60941 | Not_Computer
Z-1916-2016 | 74107 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Universal SlingBars |  |  | March 08, 2016 | 2016 | Hill-Rom, Inc. | The center bolt of the sling bar, which connects the bar to the patient lift, has failed during use. If the sling bar is not used as intended the bolt may be weakened. If the bolt is weakened there is a potential risk for breakage with the result of a free fall of the patient. This hazard could cause potentially minor to catastrophic injuries to a patient. | Labeling design | On 3/8/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected US customers via email and courier. On 3/24/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected Canadian customers via email and courier. On 3/29/2016, MEDICAL DEVICE CORRECTION notifications were sent to the affected Global customers via email and Docapost. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The notification provided new instructions for users with a lift manufactured prior to 11/19/2013, instructed the user on how to inspect for damage, and offers replacement at a 50% reduction in cost. If you have any questions concerning this Field Safety Notice, please contact Hill-Rom Technical Support, your distributor, or your Hill-Rom representative. |  | 163561 | 163,561 | Worldwide and US Nationwide Distribution AE, AO, AR , AT , AU , AZ , BD , BE , BG , BH , BM , BN , BR , CA , CH , CL , CN , CO , CS , CZ , DE , DK , EG , ES , FI , FR , GB , GR , HK , HR , HU , ID , IE , IL , IN , IR , IS , IT , JO , JP , KR , KW , LB , LI , LK , LT , LU , LV , LY , MA , MT , MX , MY , NL , NO , NZ , OM , PH , PK , PL , PT , QA , RO , RS , RU, SA , SE , SG , SI , SK , SY , TH , TN , TR , TW , UA , US , UY and ZA. | Terminated | May 15, 2017 | 433 | 163561 | Not_Computer
Z-1923-2016 | 74002 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | 4 ml FX Sodium Fluoride/Potassium Oxalate tube, Ref 454297 |  |  | April 26, 2016 | 2016 | Greiner Bio-One North America, Inc. | The 4 ml FX Sodium Fluoride/Potassium Oxalate tube may have no or low additive (sodium fluoride and potassium oxalate), which can effect analytical results for glucose and lactate. | Process control | Greiner Bio-One North America, Inc. sent an Urgent Product Recall letter dated April 26, 2016, to all affected consignees. The letter requested that they stop distribution of the product and to provide Greiner with a list of their sub-accounts. The letter included a Product Disposition/Site Confirmation form which is to be returned to Greiner Bio-one. Consignees with questions were instructed to call 800-515-8112. For questions regarding this recall call 704-261-7823. |  | 250800 | 250,800 pieces (209 cases) | Nationwide Distribution to CA, FL, IL, KY, NY, PA, VA and Puerto Rico | Open, Classified |  |  | 250800 | Not_Computer
Z-1925-2016 | 73999 | No | LWG | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Table Cushion (HybriGel Pads) |  |  | April 19, 2016 | 2016 | Berchtold Corp. | The Surgical Table Cushion (HybriGel Pads) could potentially contain staples that's secure two underlying fabrics, which would be captured on certain scans on patients, such as x-ray. | Nonconforming Material/Component | Berchtold Corporation sent an Urgent Medical Device Recall letter dated April 19, 2016, to all affected customers. The consignees were instructed to return the enclosed acknowledgment form by email to ComRecallCoordinator@stryker.com and then discard the HybiGel Pads. |  | 227 | 227 units | Nationwide Distribution | Terminated | March 31, 2017 | 346 | 227 | Not_Computer
Z-1935-2016 | 74225 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | MEGAÂż 7.5Fr. 30cc Intra-Aortic Balloon Catheter with Insertion Kit &amp; STATLOCKÂż |  |  | May 13, 2016 | 2016 | Datascope Corporation | It was discovered during post-quarantine packaging on 2/16/2016 that the IFU for the STATLOCK Sheath Stabilization device for Percutaneous Sheath Introducer (Vendor PN 2403097-1405R) was erroneously packaged with the Mega and Sensation Plus IAB Catheter Kits instead of the IFU for the STATLOCK Catheter Stabilization device for IAB catheters (Vendor PN 24030905-1406R). | Unknown/Undetermined by firm | Maquet Inc sent an Urgent Medical Device Field Correction Letter/Response Form dated 5/13/2016 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 31030 | 31,030 units (US); 15,604 units (OUS) | Worldwide Distribution. | Open, Classified |  |  | 31030 | Not_Computer
Z-1928-2016 | 74215 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima 1.5T MR430s MR Scanner, a diagnostic imaging device |  |  | May 09, 2016 | 2016 | GE Healthcare, LLC | The gas venting may not occur properly. A site reported a magnet quenching with subsequent venting of the cryogen gas into the magnet room. Cryogen gas that vents into the magnet room could cause critical deprivation of oxygen to users and patients. | Nonconforming Material/Component | GE Healthcare sent an " Urgent Medical Device Correction" letter dated May 9, 2016, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Detail, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 101 | 101 | US: AL, AZ, CA, CO, FL, ID, IL, IA,ME, NE, NV,NH, NJ, NY, NC, OH, PA, SC, SD, TN, TX, UT,VA, WA, WI.OUS:ArgentinaAustraliaAustriaBrazilCanadaChileColombiaDenmarkEgyptFinlandFranceGermanyGreeceHong KongIcelandItalyJapanKorea (Republic of)KuwaitMalaysiaNetherlandsNorwayPeruPolandPortugalRussiaSingaporeSouth AfricaSpainSwedenSwitzerlandUkraineUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 198 | Not_Computer
Z-1929-2016 | 74265 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac Vero is a Patient Positioning System for Radiation therapy. |  |  | May 04, 2016 | 2016 | Brainlab AG | Potentially incorrect positioning when using Implanted Marker Detection with Brainlab ExacTrac Vero 3.5 | Device Design | A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION was sent to customers on May 4, 2016 and distributed via email. An active reply by the consignees will be followed, if necessary via further different communication media.As of now for new installations- the Product Notification letter is handed over to the customer during installation (before clinical use) by the according Brainlab Service representative,- a signature is required by the customer that the contained User Corrective Action will be implemented before clinical use by the facility (to be adhered to by the users).Brainlab Instructions and actions required to be performed by the customer are described below.1. For Positioning and Verification, carefully review plausibility of implant marker fusions resulting in rotation angles of more than 10Âż. If in doubt consider the following options before proceeding:- Deselect one or several markers defined in the CT data set, if you suspect these markers to have migrated- Apply bony fusion instead of marker based fusion- Apply ConeBeam CT correction and verification instead of marker based fusion- Acquire a new CT scan and repeat treatment planning and definition of implanted marker positions based on the new CT data. For further details regarding the suggested proceedings, please consult the according Clinical User Guide. 2. For Dynamic Tracking continue to follow the instructions of the Clinical User Guide regarding the review of the correlation model based on implanted marker detections. Specifically: - Examine the correctness of localization of markers and targets in the X-ray sequence.- Verify the consistency of the Target Detection Curve and/or Correlation Model Curve. Irregularities of the curves are an indication for incorrect implanted marker detection. In this case please consider the options described in 1. before proceeding.If the customers require further clarification, they should contact their local Brainlab Customer |  | 27 | 27 Systems | Distributed in the states of Florida, New York, Texas and Ohio, and in the countries of Belgium, France, Germany, Italy, Japan and South Korea. | Open, Classified |  |  | 27 | Not_Computer
Z-1933-2016 | 73921 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholesterol, Catalog No. OSR6x16 |  |  | April 13, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has identified that Dipyrone (Metamizole) may cause erroneous, false low results in the following assays: Cholesterol OSR6X16, Triglyceride OSR6X118, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. | Device Design | Beckman Coulter sent an Urgent Medical Device Recall letter dated April 13, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions regarding the customer notification letter are instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support; (800) 223-0130 in the US and Canada. Outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative. |  | 43877 | 43,877 units total (10,496 units in US) | Worldwide Distribution - US Nationwide including Puerto Rico and the countries of: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia-Herzegovina, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guadeloupe, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Lesotho, Liberia, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Zambia, and Zimbabwe. | Open, Classified |  |  | 43877 | Not_Computer
Z-1942-2016 | 74172 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Norian Drillable Inject 3 CC-Sterile; catalog # 07.704.003S |  |  | May 11, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall of unexpired and unopened part and lot numbers for the Norian Drillable Injects-Sterile due to the rotary pouch within the referenced lots and associated part numbers may potentially be labeled with the incorrect powder volume/size; however, the rotary pouch contains the correct powder volume/size according to the outer label. | Under Investigation by firm | DePuy Synthes sent an Urgent Medical Device Recall letter dated May 9, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with the affected products were asked to call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s), Complete the Verification Section (page 3 of this letter) and Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section by:Fax: 855-695-8597 or Scan/email: Synthes6797@stericycle.com. If they had any questions they were asked to call 610-719-5450. For questions regarding this recall call 610-719-5443. |  | 1246 | 1246 units | Worldwide Distribution - Nationwide to AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD., ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, UT, VA, VT, WA, WI, WV, WY and Internationally to Canada... | Open, Classified |  |  | 6156 | Not_Computer
Z-1698-2016 | 74189 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ellipse ND: YAG Laser applicator for Ellipse Nordlys system |  |  | May 05, 2016 | 2016 | Ellipse A/S | Failure to provide the required instructions specifying a procedure and schedule for calibration of the measurement system. | Radiation Control for Health and Safety Act | Ellipse A/S has evaluated our instructions in the Service Manual and brought those parts into the Operator's Manual as well which we believe is relevant and usable for the end-user of the product. The details in efficiently implementing the distribution of the manual has been evaluated and coordinated with our service provider in the US.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Any future customer notification letters to be sent must comply with the requirements of 21 CFR 1003.21. This includes, but is not limited to, the prescribed packaging and the statement that you will remedy the defect or bring the product into compliance at no charge.For further questions, please call (760) 798-9550. |  | 25 | US - 25 | Nationwide Distribution to the states of : NJ, PA, MI, CO, CA, FL, OK, NY, NH, CT, AZ and MO. | Open, Classified |  |  | 39 | Not_Computer
Z-1940-2016 | 73998 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | CENTURION FMS (Fluidics Management System) Pack for the CENTURION Vision System and identified Custom Paks |  |  | April 28, 2016 | 2016 | Alcon Research, Ltd. | Alcon is conducting this Voluntary Medical Device Removal for specific lots of its CENTURION FMS packs due to the possible presence of a molding irregularity found within the aspiration luer which may result in a reduction of vacuum within the aspiration line. | Process control | The affected consignees were sent a recall notification letter on 4/28/16. |  | 156712 | 156,712 units | US, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colunbia, Croatia, Cyprus, Czech Republic, Denmark, Finland, Germany, Greece, Hong Kong, Hungary, India, Ireland, Italy, Japan, Korea, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Russia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay | Terminated | March 31, 2017 | 337 | 156712 | Not_Computer
Z-1939-2016 | 74170 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Maestro Air Motor |  |  | April 28, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Due to the buildup of grease within the hose and motor, the product may require multiple attempts to start. The grease may also migrate into the attachment end of the motor. There is a negligible possibility that the orange grease may leak from the attachment end of the motor and contaminate the surgical wound resulting in soft tissue inflammation. | Component design/selection | Stryker sent an "Urgent Medical Device Recall Notification" dated April 28, 2016, to all affected customers. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments at (269) 389-4518. |  | 189 | 189 | Domestic only: ALARCAFLGAILINKSKYMDMIMONCNMOHOKSCTNTXVAWAWI | Terminated | March 24, 2017 | 330 | 189 | Not_Computer
Z-1938-2016 | 74034 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Sterile Sheridan/HVT Cuffed, Murphy Eye, Tracheal Tube, Ref: V5-10314, ID: 7.0 mm, Rx Only, Manufactured by: Teleflex Incorporated. |  |  | May 05, 2016 | 2016 | Teleflex Medical | The labeling of some units indicates that it contains a size 7mm tracheal tube, but may contain a size 7.5mm tracheal tube. | Error in labeling | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" letter dated May 5, 2016 to its accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue distribution and quarantine product; communicate this recall to any of your customers who have received product; have your customers return to you any affected product with Recall Acknowledgement Form; contact Teleflex Medical Customer Service for more information at 1-866-246-6990; and complete and return Recall Acknowledgment Form via fax to 1-855-419-850, Attn: Customer Service or email to recalls@teleflex.com, even if you and your customers have no affected stock/product. Teleflex will provide you with instructions for the return of product. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 6390 | 6,390 ea. | US Distribution to states of: AL, AZ, CO, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NY, NC, OK, PA, RI, TX and UT. | Open, Classified |  |  | 6390 | Not_Computer
Z-1941-2016 | 74213 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Signature TKA GDE/MDL Set 04-05 |  |  | May 03, 2016 | 2016 | Materialise USA LLC | A patient specific case of Signature Personalized Patient Care System Knee Guides included a femur guide which does not match the patient associated with the case.If the issue is not identified prior to surgery, the personalized patient specific guide would provide an improper fit and would not be able to be used. | Process control | Materialise sent an URGENT: Signature Personalized Patient Care System Knee Guides Notification dated May 3, 2016. It was sent to the affected customer via email with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Contact Materialise customer service for additional information or to find a status of the second patient specific surgical guide Monday through Friday, 8:00AM to 4:30 PM, Eastern Time.734-259-7016 |  | 1 | 1 | One consignee, Indiana. | Completed |  |  | 1 | Not_Computer
Z-1937-2016 | 73828 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Henry Schein Single Use Syringe Luer-Lock, 10 mL, Intended for use by health care professionals for general purpose fluid aspiration/injection |  |  | March 25, 2016 | 2016 | Becton Dickinson &amp; Company | BD 10 mL Syringe Luer-Lok Tip with BD PrecisionGlide Needle 21G x 1 (0.8 mm x 25 mm) and Henry Schein 10 mL Single Use Syringe Luer-Lock are being recalled due to possible adulteration and non-conforming manufacturing practices. | Error in labeling | Becton Dickinson sent via email and UPS Ground an "Urgent Voluntary Product Recall" Letter dated March 25, 2016 along with a Recall Response Form and packaging instructions to their sole customer. Customers were instructed to immediately review their inventory , quarantine any affected product, and immediately discontinue shipment of the affected product. Customers should complete the Recall Response Card from and fax it to BD at 1-201-847-6990. Return all affected product with the completed Recall Response Card. If product was further distributed please identify your customers and notify them immediately of the product recall. Customers were also instructed to complete the Recall Response Form even if they do not have any affected lots in their inventory. Customers with questions were instructed to call 1-201-847-4267. For questions regarding this recall call 201-847-5612. |  | 811200 | 811,200 units | Nationwide Distribution to NY only | Open, Classified |  |  | 987200 | Not_Computer
Z-1949-2016 | 74082 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | LCP Dia-Meta Volar Distal Radius Plate 15 H Shaft/LT-Sterile; Indicated for fractures, osteotomies, and non-unions of the radius and other small bones. |  |  | May 02, 2016 | 2016 | Synthes (USA) Products LLC | Product incorrectly packaged. The labels on the outside of the box and on the LCP Dia-Meta Volar Distal Radius Plate inside the box are correct; however, the preprinted box which contains the LCP-Dia-Meta Volar Distal Radius Plate is incorrect. The box is preprinted for a 4.5 mm LCP Condylar Plate - Distal Femur. | Under Investigation by firm | DePuy Synthes sent out an Urgent Notice to customers informing them of a package mixup and the impact it could have on users. Customers were asked to evaluate their inventory and immediately remove the affected lot (9890811) from stock. Customers were asked to call DePuy Synthes at 1800-479-6329 to obtain a Return Authorization Number for return of the affected product, fill out the verification section of the letter and mail with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Any questions, they were said to call 610-719-5450 or to contact their DePuy Synthes Sales Consultant. |  | 6 | 6 units | Distributed to one customer: Colorado | Open, Classified |  |  | 6 | Not_Computer
Z-1945-2016 | 70054 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | The XN-Series modules (XN-10, XN-20) are quantitative multi-parameter automated hematology analyzers intended for in vitro diagnostic use in screening patient populations found in clinical laboratories. |  |  | December 18, 2014 | 2014 | Sysmex Corporation | There is a potential for elevated hemoglobin (Hgb) and impedance platelet count (PLT-I) values. | Device Design | Sysmex Corporation sent a Product Notification Field Correction letter were to all affected customers of record. The letters included instructions for customers to: 1) Review results with an unexpected increase in Hgb, MCH, MCHC, and/or PLT-I. The increase may result in a low or normal value being artificially elevated to either a normal result or a result above the reference ranges, or a high value being artificially increased. There is negligble impact to the RBC when the issue occurs; 2) The relationship with an elevated Hgb would cause the indices (MCH and MCHC) values to appear to be abnormally elevated and may result in the results being flagged for verification or indicated to be outside of reference ranges; 3) Review any "Turbidity/Hgb Interf?" flag which will automatically trigger at an MCHC of 37.5 g/dL; 4) If Hgb is affected, the PLT-I may be also. Review any PLT-I result when an unexpected increase in MCH/MCHC or "Turbidity/Hgb Interf?" flag occurs; 5) Gross interference is likely to trigger a flag, however there is a potential any interference could increase the Hgb or the PLT-I count from a lower level to an "normal" level. Follow laboratory protocol for confirming unexpected results. Customers were further advised that they will be contacted by a Sysmex Field Representative within 180 days to schedule an appointment to replace the rinse cup on the analyzer(s) and that version 00-16 software will be installed to change the rinsing sequence for the aspiration probe. Customers with questions are requested to e-mail the Sysmex Technical Assistance Center. Urgent questions can be answered by calling the Technical Assistance Center at 888-879-7639 (24 hours a day / 7 days a week). |  | 1137 | 1,137 units (US) 54 units (Canada) 9 inventory | Worldwide Distribution - US to AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY + District of Columbia + Guam; Internationally to Australia, Bahrain, Bangladesh, Cambodia, Canada, China, Germany, Hong Kong, India, Indonesia, Malaysia, Maldives, Mongolia, Mynamar, Nepal, New Zealand, Pakistan, Philippines, Singapore, South Korea, Sri Lanka, Taiwan, Thailand, and Vietnam.THE CORRECTION IS LIMITED TO XN-10/XN-20 UNITS DISTRIBUTED BY SYSMEX AMERICA, INC., INSTALLED AT END-USER SITES IN US, PUERTO RICO, AND CANADA. | Terminated | July 25, 2016 | 585 | 1137 | Computer
Z-1948-2016 | 74178 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | Ariol is an automated scanning microscope and image analysis system. It is intended for in vitro diagnostic use as an aid to the pathologist in the detection, classification, and counting of cells of interest based on particular color, intensity, size, pattern, and shape. |  |  | May 06, 2016 | 2016 | Leica Biosystems Richmond Inc. | An issue during the manufacturing process caused systems to have an improperly activated Windows 7 OS, even though a valid Windows 7 license was associated with each manufactured system. | Nonconforming Material/Component | According to the firm a decision to conduct a Field Corrective Actions of the Ariol Image Analysis and Scanning Systems was done out of an abundance of caution. Leica Biosystems Richmond sent Medical device Field Correction notification letters to the customers of the potentially affected product. The initial written communication was sent to the customers via first class USPS on May 6, 2016. The Medical Device Field Correction letter describes the issues identified with the Ariol Image Analysis and Scanning Systems and the Windows 7 operating system. The letter provided the customer with actions they are required to take which include the following: 1) whether they have a properly activated Windows 7 operating system. 2) If they do not have a properly activated operating system they should to contact Leica Technical Support for additional information and to request a service appointment for correction to their system. For USA customers, they ask that they call 18005374669. International customers are asked to go to http://www.leicabiosystems.com/contactus/contactusonline/ and then choose your country.3) Customers are asked to confirm receipt of the notice by signing and dating the attached Acknowledgement Form and emailing a copy to RichRA@leicamicrosystems.com or faxing it back to us at 18472363800.Service personnel will be sent to the affected customer locations to replace the hard drives of the affected systems. |  | 33 | 33 systems | U.S. consignees: Nationwide; Foreign Consignees: China, Japan, Russia, Italy and England. | Open, Classified |  |  | 33 | Not_Computer
Z-1947-2016 | 74097 | Yes | NIM |  |  | Cardiovascular | PMA | Cordis PRECISE (R) RX Nitinol Stent System (Biliary) |  |  | May 04, 2016 | 2016 | Cordis Corporation | Inability to deploy the stent or partial stent deployment. | Component design/selection | Cordis sent an Urgent Medical Device Recall (Removal) letter dated May 4, 2016, to all affected customers.Action to be taken by the customer: 1) Immediately check your inventory to confirm whether you have any units from affected lots in your possession. Identify and set aside any units from the affected lot in a manner that ensures the affected product will not be used. Check all storage and usage locations.2) Review, complete, sign and return the enclosed Acknowledgement Form directly to Cordis at the fax number on the form. Note that you will receive multiple letters and Acknowledgement Forms if your facility received affected product at both a legacy Ship-To account number and a Cardinal Health ShipÂż To account number at the same address. In that scenario, please sign and return both forms.3) Return all affected products to the address listed on the form, with reference to Return Authorization Number ( i.e. RA #)#which is listed on the form. Replacement devices will be shipped after the affected product is returned. Consigned Inventory will also be replaced after the product is returned and an analysis of SKU mix needed to properly support the customer is conducted.4) Share this letter with others in your facility who need to be made aware of this recall, including the areas in some facilities where these products are frequently used (see list above) Contact any other facilities that have beenprovided with units of affected lots. Maintain awareness of this notice until all affected product has been returned to Cordis.5) Keep a copy of this notice with any affected product until returned.For questions related to assistance returning product or billing concerns or questions please contact your sales representative or Cordis Customer Service at 1-800-327-7714 Option 1, Monday through Friday from 7:00AMto 8:00 PM EDT. |  | 7500 | 4,400 units in US; 3,100 units to Foreign Countries | NATIONWIDE and FOREIGN CONSIGNEES: Japan, Taiwan, Thailand and South Korea. | Open, Classified |  |  | 7500 | Not_Computer
Z-1955-2016 | 73536 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00720504022 ACETAB.CUPW/SH,0DEG,22MMX40MM |  |  | April 15, 2016 | 2016 | Zimmer Trabecular Metal Technology, Inc. | Zimmer Biomet is initiating a recall of specific lots of foil pouched, sterile Trabecular Metal implants that were manufactured between April 2011 and October 2012. The foil pouches are part of a dual barrier system to provide sterile integrity. | Unknown/Undetermined by firm | Zimmer issued an Urgent Medical Device Recall Letter dated 4/15/2016. |  | 19884 | 19,884 total | Nationwide and Foreign. | Open, Classified |  |  | 19884 | Not_Computer
Z-1888-2016 | 74111 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | AVEA Comprehensive ventilator, 17310-xx (xx = 0 - 14) |  |  | May 17, 2016 | 2016 | Carefusion 211 Inc dba Carefusion | CareFusion has identified a potential risk associated with AVEA Ventilator caused by an incorrect F1 fuse on the AVEA Ventilator Transducer Communication Alarm (TCA) Board which may fail/blow, causing loss of power to the User Interface Module (UIM). | Under Investigation by firm | CareFusion sent a Field Safety Notice dated May 17, 2016, to all affected customers. The notice informed customers that CareFusion has identified a potential risk associated with AVEA Ventilator caused by an incorrect F1 fuse on the AVEA Ventilator Transducer Communication Alarm (TCA) Board which may fail/blow, causing loss of power to the User Interface Module (UIM). This malfunction creates a blank AVEA Ventilator UIM and INOP condition during pre-use check or during use with cessation of ventilation to the patient. By design an audible clinical alarm is activated. Customers were informed of the problems identified and the actions to be taken. Customers were not required to return the affected devices or suspend use, customers would receive an copy of the identified affected serial numbers, as well as a response card, customers were requested to promptly return the response card to expedite the correction process and acknowledge receipt of the notification, Customers would be contacted by CareFusion to arrange onsite remediation of the affected devices, in the interim if any AVEA ventilator unit in their facility exhibits a blank AVEA Ventilator UIM and INOP condition during pre-use check or during use with cessation of ventilation, immediately remove the ventilator from service, provide alternate ventilation and contact CareFusion Technical Support per the contact information listed below to report the issue. Customers with questions should call 888-562-6016, for technical support 800-231-2466 or support.vent.us@carefusion.com. |  | 922 | 922 units total (501 units in US) | Worldwide Distribution - US (Nationwide) and the countries of Australia, Hungary, Paraguay, Belgium, Indonesia, Philippines, Bolivia, Ireland, Poland, Bermuda, India, Saudi Arabia, Canada, Japan, Singapore, Chile, Jordon, South Africa, China, Kuwait, Taiwan, Ecuador, Malaysia, Thailand, Egypt, Mexico, Turkey, France, Netherlands, United Kingdom, Guatemala, and Nicaragua. | Open, Classified |  |  | 922 | Not_Computer
Z-1957-2016 | 74156 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Red Head 2 Reusable Fat Harvesting Canister, Product # RH-2, Non-Sterile. |  |  | April 21, 2016 | 2016 | Miami Fat Supply, Inc | During an FDA inspection it was found that the products are marketed without a cleared 510k, nor is approved for multiple use. | No Marketing Application | Customers were sent an Urgent: Medical Device Recall/Correction letter, dated 04/19/2016, on 04/21/2016 that served as a notification of the Red Head 2 voluntary recall. Customers were informed that there were previous complaints of compromised product integrity and performance following sterilization. Lids are subject to fracture and fail after steam sterilization. Customers were advised to destroy used product and to ship unused product back to the manufacturer. Customers were to return an acknowledgement form to confirm their notification of the voluntary recall. |  | 443 | 443 units | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, FL, GA, IL, MD, MI, MO, NM, NY, OH, TX, UT, and VA; and, the country of Dominican Republic. | Open, Classified |  |  | 443 | Not_Computer
Z-1989-2016 | 74045 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Excelsior Disposable Syringe Heparin Lock Flush, USP syringe (5 mL Fill in 10 mL syringe). For flushing of IV catheters and IV tubing only. |  |  | April 28, 2016 | 2016 | Excelsior Medical Corp | Certain lots of Heparin Lock Flush, USP Syringes (5 ml Fill in 10 ml Syringe) failed Anti-Factor IIa assay for potency during testing of retention samples for heparin potency. | Nonconforming Material/Component | Excelsior Medical Inc. sent a recall letter on 4/29/2016 via certified mail. Customers are instructed to discontinue using the product subject to this recall. Customers are requested to complete the customer reply form and fax it (866) 807-3714, so arrangements can be made to return the product to Excelsior. |  | 184500 | 184,500 | US Nationwide distribution. | Terminated | August 08, 2017 | 467 | 184500 | Not_Computer
Z-1990-2016 | 74207 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker PowerPRO (XT, IT, or TL) Cots |  |  | May 05, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | Complaints of smoke coming from the foot end of the cot. This reportedly occurred when either the battery was placed into the cot or shortly after placing the battery into the cot. In addition to the alleged smoking, one of the reports had also noted that when the user attempted to lower the cot (prior to the alleged smoking) the device went into high speed retract and then a few moments later i | Vendor change control | On 5/5/2016, IMPORTANT MEDICAL DEVICE NOTIFICATION notifications dated 5/3/2016 were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 770 | 770 | Distributed US (nationwide) and the countries of Australia, Canada, Hong Kong, Spain, Switzerland, and United Kingdom. | Open, Classified |  |  | 770 | Not_Computer
Z-1958-2016 | 74084 | No | PLF | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | BioSync Anatomic Cotton Wedge, 16 x 4.5 mm.; Used for internal bone fixation for bone fractures, fusions or osteotomies in the ankle and foot. |  |  | April 22, 2016 | 2016 | Arthrex, Inc. | Potential for polystyrene particulate presence. | Material/Component Contamination | On 04/22/2016 Arthrex contacted all distributors that received potentially affected devices and instructed them to place the product in quarantine. Distributors were then instructed to return affected devices to Arthrex. Surgeons who have implanted the device will also be receiving a recall notification. Surgeons, risk managers, and distributors received mailed recall notification identifying the recalled products and batches, the reason for recall, and instructions for returning affected product. Customers are to call Arthrex's Customer Service Line (800-934-4404) to obtain return authorization and further instructions. Customers were also required to complete a "Recall Acknowledgement of Receipt" postal card to return to Arthrex, Inc. in an enclosed envelope. |  | 138 | 138 total products | Distributed in the states of CA, CO, FL, GA, IL, IN, KS, MD, NC, NY, PA, SC, VA, WA, and WI. | Open, Classified |  |  | 138 | Not_Computer
Z-1964-2016 | 74067 | No | KYB | Intraocular lens guide. | Ophthalmic | Ophthalmic | 510(K) Exempt | STAAR AQ Cartridge-FP, guide and delivery system for Intraocular lenses. |  |  | May 02, 2016 | 2016 | Staar Surgical Co. | The STAAR AQ Cartridge-FP may malfunction and may crack during loading or delivery of the IOL (Intraocular Lenses). | Process control | STAAR sent a Medical Device Recall Letter dated May 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. STAAR Surgical is recalling the STAAR AQ Cartridge-FP because it may malfunction and may crack during loading or delivery of the IOL. Customers with questions are instructed to contact their STAAR customer service representative at (800) 352-7842. |  | 2612 | 2,612 units | US | Open, Classified |  |  | 2612 | Not_Computer
Z-1966-2016 | 73598 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Vaxcel PICC with PASV Catheter Kit under the following labels: 1) 3F, UPN H965454350, Rx ONLY, 2) 4F, UPN M001454550, Rx ONLY &amp; 3) 5F, UPN M001454700, Rx ONLY. |  |  | March 14, 2016 | 2016 | Navilyst Medical, Inc., an AngioDyamics Company | PICC catheters that contain valve(s) manufactured prior to July 15, 2014 lack a manufacturing specification change that reduces the incidence of hemolysis during blood sampling through the PICC catheter. | Device Design | Navilyst Medical distributed Urgent Voluntary Medical Device Recall notification letters on March 14, 2016, and Reply Tracking forms to their customers via courier service. Customers were asked to immediately remove any affected (recalled) product from their inventory (whether in labs, central supply, shipping and receiving or any other location). Segregate this product in a secure location for return to Navilyst Medical. Immediately forward a copy of therecall notice to all sites to which you have distributed affected product. If affected product is located in your institution, please call Navilyst Medical Customer Service at 1-800-772-6446 Monday to Friday, 8 am - 7 pm, EST to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) by fax to 1-800-782-1357.For questions regarding this recall call 518-792-4112. |  | 2716 | 2,716 units | Worldwide Distribution US (Nationwide) Internationally to AU, GR, SE, SA, IN, CA, GB, BR, BE, NL, ES, KR, HK, AE, IT and ES. | Open, Classified |  |  | 17283 | Not_Computer
Z-1991-2016 | 74229 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Piccolo Composite Tibia and Femoral Nailing Systems - Ball Tip Guide Wire for fixation of fractures that occur in and between the proximal and distal third of the long bones being treated. |  |  | May 10, 2016 | 2016 | CarboFix Orthopedics, Ltd. | Firm received complaints where the Piccolo Composite Ball Tip Guide Wire penetrated the internal package, and thus compromised the product sterility. | Under Investigation by firm | On May 10, 2016, an advisory notice was issued to all distributors that currently may possess (or supplied to hospital) the product. The notice was sent to a representative of the distributor companies via an e-mail (using confirm receipts for delivery and notification when the email is read), together with a letter to hospital(s) for the representative to use in case the distributor supplied the product to the hospital. The letter requested that the product be returned and also requested a sub-recall. |  | 638 | 638 | AL, AZ, CA, CO, DC, FL, GA, IN, KS, KY, MD, MI, NE, NV, NY, OH, OK, TX, WA | Open, Classified |  |  | 638 | Not_Computer
Z-1965-2016 | 74114 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TS Series Depth Gauge for Integra Ti6 Internal Fixation System 2.0/2.5mm Screws |  |  | May 10, 2016 | 2016 | Integra LifeSciences Corp. | As a result of an internal investigation of a non-conformance it has been determined: 1) Depth Gauges from Lot UP15H2 exceed the overall length specification requirements. 2) Depth Gauges from prior lot numbers manufactured before Integra acquired the product line reflect an incorrect number for the first mm depth marking on the gauge. The number should be 4, it is 6. 3) The space / holder | Process design | The firm, Integra Lifesciences, sent an "Urgent: Voluntary Medical Device Recall" letter dated 5/10/16 to affected consignees/customers via traceable courier service on 5/10/16. The letter describes the product, problem and actions to be taken. The customers were instructed to stop using and distributing product immediately; complete and return the Recall Acknowledgement and Return Form via fax to 1-609-275-9445 or email to FCA2@integralife.com, even if you do not have any of the product.Should you have any questions, please contact Customer Service at 1-866-854-8300. |  | 71 | 71 units | US Distribution to states of: CO, FL, IL, MI, NJ, NY, SC, TX, and VA. | Open, Classified |  |  | 71 | Not_Computer
Z-2001-2016 | 73169 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE Phototherapy system, Catalogue Part Number: 010066. |  |  | April 11, 2016 | 2016 | Natus Medical Incorporated | The firm is adding the warning statement to the Instruction for Use for the neoBLUE LED Phototherapy System and provide a laminated Quick Guide to affix to each device. | Labeling False and Misleading | The URGENT: FIELD SAFETY NOTICE PRODUCT LABELING CHANGEneoBLUE LED Phototherapy System, dated April 2016, was went to US customers between April 11, 2016 and April 27, 2016. Natus distribution partners outside the US will receive the Letter to Distributors and they are requested to contact their downstream customers to distribute the information. The estimated date of notifying international consignees is May 31, 2016.Enclosed with this letter, customers would receive:1) Quick Guides (2 each) which provide operating steps, information about measuring intensity, and warnings (PN 051693)2) User manual addendum3) Customer reply formNatus requests that this information be distributed to all associated parties within your organization that use or service this product. Please note: this communication is in addition to any prior labeling changes that you may have received from Natus.Customers should submit the Reply Form via fax to 847-949-6031 or email to mailto:Natus_Quality_Programs@natus.com.If there are any questions regarding this letter and addendum, please call Natus Technical Service at 888-496-2887. |  | 13541 | 13541 units total (7840 units in the US and 5701 units outside the US) | Worldwide Distributed - US Nationwide in the US including American Samoa, Guam Puerto Rico and in the following countries:Argentina, Australia, Austria, Bangladesh, Belgium, Bermuda, Bolivia, Brunei Darussalam, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, , Kenya, , Kuwait, Latvia, Lithuania, Malaysia, Moldova Republic of, Morocco, Namibia, Netherlands, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, Virgin Islands (British), and Yemen. | Open, Classified |  |  | 13541 | Not_Computer
Z-1996-2016 | 71372 | No | GEL |  |  | General &amp; Plastic Surgery | 510(k) | EVARREST Fibrin Sealant Patch, Manufactured by: Omrix Biopharmaceuticals, Ltd, Israel. |  |  | April 11, 2016 | 2016 | Ethicon, Inc. | Ethicon is recalling EVARREST Fibrin Sealant Patch because the characters "0-bad" were printed in the expiration date field on the foil pouch labels of the Fibrin Patch units instead of the valid expiration date (2016-12-28). | Error in labeling | Ethicon sent an "Urgent: Product Recall (Removal)," dated April 14, 2016 via UPS next day mail to their affected customers. The letter identified the affected product and the reason for the recall. The letter also states that product from the affected lot is still within the expiration date, and health care practitioners, who have treated patients with the product should continue to follow patients in the usual manner. Users are to examine their inventory and quarantine the affected product. The affected product should be returned per the instructions outlined in the letter. Ethicon will provide replacement product upon receipt. Customers are to complete and return the Business Reply Form. If customers need assistance returning the product, they can contact Stericycle at 1-866-761-9472. If there are questions regarding the replacement product, customers can contact Ethicon Customer Support Center at 1-877-ETHICON. |  | 228 | 228 units | Nationwide Distribution | Terminated | August 02, 2017 | 478 | 228 | Not_Computer
Z-1994-2016 | 73203 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Universal Battery Charger II (UBC II) |  |  | January 15, 2016 | 2016 | The Anspach Effort, Inc. | The device is reported to enter safe mode and is no longer able to charge batteries. | Process design | DePuy Synthes sent an "Urgent Notice- Medical Device Recall (Removal) letter dated January 15, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter informing them of the recall (removal) and asking them to return the UBC II when it display the blue flashing lights that indicate service and to contact DePuy Synthes Customer Support at 800-327-6887 option 2 for return device and arrange to repair or replacement. Any other question please contact the Complaint Handling Unit Manager at 561-494-3673. |  | 1992 | 1,992 | USA Distribution to the states of : AR, AZ, CA, CO, FL, IL, IN, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OK, OR, PA, SC, TN, TX, VA, WA, WI and Hawaii | Open, Classified |  |  | 1992 | Not_Computer
Z-2002-2016 | 74341 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Small Patient ECC Pack, REF BEQ-TOP 24100, Product Usage: |  |  | May 09, 2016 | 2016 | Datascope Corporation | Custom Tubing Set BEQ-TOP-24100 (Small Patient ECC Pack) was identified as possessing a level of endotoxin that exceeded the regulatory limits. | Under Investigation by firm | Maquet sent via express mail and hand delivered a recall letter and response form dated May 9, 2016 to their sole customer. |  | 17 | 17 units | US in the state of GA | Open, Classified |  |  | 17 | Not_Computer
Z-1993-2016 | 73473 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK15 monitor/defibrillator with End-Tidal C02 (EtC02) feature installed. |  |  | January 27, 2016 | 2016 | Physio-Control, Inc. | The firm became aware that when using EtC02 in the kPa or % setting and in a situation where the reading is above 9.9 kPa, the display of the LIFEPAK 15 respiratory rate may partially obscure a portion of the leading digit of the EtC02 value. This affects all LIFEPAK 15 with an EtC02 feature installed and configured to the kPa or % setting. | Software design | Physio-Control sent the letter with the title Information Regarding LIFEPAK 15 MONITOR/DEFIBRILLATOR WITH End-tidal CO2 OPTION, dated January 2016, to customers in the United Kingdom (UK) as REQUESTED by the Medicines and Healthcare product Regulatory Agency (MHRA) on January 27, 2016.Physio initiated a voluntary field action on February 23, 2016 and started notifying US consignees on March 22, 2016 and started notifying international consignees on April 18, 2016. The Information Regarding LIFEPAK 15 MONITOR/ DEFIBRILLATOR WITH End-tidal CO2 OPTION letter, dated March 2016, was sent.Customer Letters began mailing in EMEA region, beginning on 06/20/2016 via designated Courier.Customer Letters began mailing in Canada on 09/09/2016 via FedEx.Customer Letters began mailing in Latin America on 09/16/2016 via email.Customer Letters began mailing in APAC region on 09/22/2016 via email.Second Customer letters with Confirmation Sheets began mailing in the US October 2016 via FedEx and USPS Certified Mail.Actions should be taken by users:1. Please forward this information to all of your sites, trainers and users that may have a LIFEPAK 15 monitor/defibrillator.2. If you use your LIFEPAK 15 monitor/defibrillator to display EtC02 values you should also display the EtC02 waveform which includes the range along the right side of the display.3. If you use either kPa or % settings, activate EtC02 alarms to identify EtC02 values outside the normal range.4. If your protocols allow this option, consider changing the EtC02 to measure in the mmHg setting.The firm reported that the revised software for LifePak15 version 4 (v4) and the customer install instructions would be available on LIFENET by April 25, 2016.The revised software for LifePak 15 versions 1 and 2 (v1 and v2) and the customer install instructions are estimated to be available on LIFENET in mid May 2016.Customers who have any questions about this subject, please contac |  | 44798 | 44798 units total (23189 units in US; 21609 units international) | Worldwide Distribution - US (nationwide) including Guam and Puerto Rico and Internationally to Âżland Islands, Algeria, Argentina, Aruba, Australia, Austria, Azerbaijan, Bahrain, Barbados, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Brunei Darussalam, Cambodia, Canada, Cayman Islands, Chile, China, Costa Rica, Croatia, CuraÂżao, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Eritrea, Faroe Islands, Finland, France, French Guiana, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guinea, Guyana, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Islamic Republic of, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, the former Yugoslav Republic of, Malaysia, Maldives, Malta, Martinique, Mauritius, Mayotte, Mexico, Morocco, Mozambique, Namibia, Netherlands, New Caledonia, New Zealand, Norway, Oman, Pakistan, Panama, Papua New Guinea, Poland, Portugal, Qatar, RÂżunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Province of China, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Yemen. | Terminated | March 31, 2017 | 429 | 44798 | Computer
Z-1992-2016 | 73692 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Protura Robotic Couch System Software, versions 1.6.1 (Part Number MT6XSM1.6.1), 1.7.0 (Part Number MT6XSM1.7.0), and 1.7.1 (Part Number MT6XSM1.7.1), CIVCO, Orange City, IA 51041 |  |  | March 31, 2016 | 2016 | Med Tec Inc | When using the recalled product in split shift mode, the non-sub-millimeter positioning has been found to be inconsistent with the setting parameters. | Software design | Civco Medical Solutions sent a "Urgent Field Correction Notice" dated March 31, 2016, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer.The software's other configuration, "All 6 Degrees Protura" does not have the same translational issues. The firm send notification letters on February 23, 2016. In this letter, the Product Manager verified each site's software is configured to this operational mode which operates without the identified problem. In a later and separate notification on March 31, 2016, the firm informed customers that a new software version will be installed to address the problem. No timeframe was given in the letter.For further questions, please call ( 712) 737-8688. |  | 9 | 9 | Worldwide Distribution - US Distribution including the countries of Switzerland, Germany.. | Open, Classified |  |  | 9 | Computer
Z-2000-2016 | 74239 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BIOGRAPH mCT S(20) - 4R, MATERIAL NUMBER 10528958 |  |  | April 25, 2016 | 2016 | Siemens Medical Solutions USA Inc. | The bracket that is used to lift the cover has bolts that were not welded in potentially causing the cover to fall down during service. | Process design | Siemens sent a Customer Safety Advisory Letter to all affected customers. The letter identified the affected product, problem and actions to be taken. Siemens service organization will be contacting customers to schedule the update as per the attached Update Instruction and will be provided the replacement bracket to resolve the issue. For questions contact your Siemens representative. |  | 1 | 1 unit | Worldwide Distribution - US Nationwide and the countries of China and Malaysia. | Open, Classified |  |  | 4 | Not_Computer
Z-1995-2016 | 74197 | No | NVP | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Flexichamber Anti-Static Valved Collapsible Holding Chamber, used by patients to administer aerosolized medication from most pressurized Metered Dose Inhalers(pMDIs). |  |  | May 16, 2016 | 2016 | FSC Laboratories, Inc | Product had torn/loose seal around the mouthpiece. | Process control | Letters were sent to consignees on 5/16/2016 informing them of the recall and requesting that they immediately discontinue use of the product. The letter included a reply form to be returned to FSC Laboratories. |  | 1435 | 1435 units | US Nationwide distribution. | Open, Classified |  |  | 1435 | Not_Computer
Z-2003-2016 | 74206 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Dome Centric component (Reference 01.04227.005) from the Anatomical Shoulder (AS) System 2.0. |  |  | May 17, 2016 | 2016 | Zimmer Gmbh | In some cases it has been difficult or not possible to disassemble the adjusted Dome Centric from the AS Humeral Rasp after initial positioning of the AS Humeral Trial Head which resulted in the whole construct of the Dome Centric with the still assembled AS Humeral Rasp taken out of the humeral canal. This could cause a delay in surgery time or the surgeon could decide to close the wound without finishing the surgery, increasing the infection risk or second use of anesthesia necessary. | Device Design | Zimmer Biomet sent an URGENT MEDICAL DEVICE RECALL  LOT SPECIFIC notification dated May 17, 2016, to all affected consignees via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Consignees were instructed to review the notification and ensure affected personnel were aware of the contents. Assist their Zimmer Bioment sales representative with the quarantine of any affected product from the table above. Complete the Certificate of Acknowledgement Form and return to corporatequality.postmarket@zimmerbiomet.com.If after reviewing this notification consignees have further questions or concerns please call the customer call center at 1- 800-348-2759 between 8:00 am and 8:00 pm EST. |  | 177 | 177 | Nationwide Distribution to AZ, CA, CO, GA, IL, KS, KY, MA, MD, MO, OH, PA, TN, and TX. | Open, Classified |  |  | 177 | Not_Computer
Z-2004-2016 | 74263 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Musculoskeletal Transplant Foundation, AFT Diverted Tube (3/4/Filled), indicated for treatment of surgically-created osseous defects or osseous defects created from traumatic injury. It can be used in the: extremities, pelvis, and spine. Product Code: 227005. |  |  | April 21, 2016 | 2016 | Musculoskeletal Transplant Foundation, Inc. | AFT Diverted Tubes are being recalled due to a customer complaint involving one unit which had an unsealed Tyvek pouch. The non-sterile foil pouch may introduce contamination into the sterile field. | Packaging process control | Musculoskeletal Transplant Foundation (MTF) sent an Voluntary Field Action Notification Letter dated April 21, 2016, and via telephone to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were notified of the situation with the product and provided instructions on how to handle the product. Questions can be directed to Customer Service Customer at 1-800-433-6576. |  | 114 | 114 units | nationwide | Open, Classified |  |  | 114 | Not_Computer
Z-2038-2016 | 74179 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | SIOOA I (EP I84) Rabbit Monoclonal Antibody; |  |  | May 10, 2016 | 2016 | Cell Marque Corporation | Cell Marque became aware of some lots of SIOOA I (EP I84) Rabbit Monoclonal Antibody, catalogue numbers 408R-14, 408R-15, and 408R-16, that exhibited an unacceptable IHC test result. | Unknown/Undetermined by firm | Product recall notifications were sent to customers on May 12, 2016. Letters advised that all unused or partially-used vials be destroyed. Customers should contact customer service for replacement and credit. |  | 23 | 23 units | Distributed in the states of MA, CA, and WA and the countries of Australia, Denmark, Turkey, Germany, Russia, Slovenia, Netherlands and Zambia. | Open, Classified |  |  | 23 | Not_Computer
Z-2039-2016 | 74245 | No | JOY | Automated cell-locating device. | Hematology | Hematology | 510(k) | CytoVision Image Analysis and Capture System, an automated cell-locating device. The Cytovision is a rapid metaphase finder, image acquisition and computer aided chromosome analysis system which assists the operator in viewing chromosomes and looking for cellular anomalies. CytoVision enables a qualified Cytogeneticist to rapidly and accurately analyze the chromosome banding pattern. |  |  | May 02, 2016 | 2016 | Leica Biosystems Richmond Inc. | Systems have an improperly activated Windows 7 OS, even though a valid Windows 7 license was associated with each manufactured system. The system appears to be unlicensed. These systems cannot be re-activated using standard Windows licensing procedures and are requiring checks of each potentially affected system to confirm Windows 7 OS is properly activated. | Software Manufacturing/Software Deployment | Leica Biosystems Richmond sent Medical device Field Correction notification letters to the customers of the potentially affected product. The initial written communication was sent to the customers via first class USPS on May 2, 2016. According to the firm a decision to conduct a Field Corrective Actions of the Ariol Image Analysis and Scanning Systems was done out of an abundance of caution. The Medical Device Field Correction letter describes the issues identified with the CytoVision Image Analysis and Capture System and the Windows 7 operating system. The letter provided the customer with actions they are required to take which include the following: 1) whether they have a properly activated Windows 7 operating system. 2) If they do not have a properly activated operating system they should to contact Leica Technical Support for additional information and to request a service appointment for correction to their system. For USA customers, they ask that they call 18005374669. International customers are asked to go tohttp://www.leicabiosystems.com/contactus/contactusonline/ and then choose your country. 3) Customers are asked to confirm receipt of the notice by signing and dating the attached Acknowledgement Form and emailing a copy to RichRA@leicamicrosystems.com or faxing it back to us at 18472363800. Service personnel will be sent to the affected customer locations to replace the hard drives of the affected systems. |  | 283 | 283 systems | Worldwide distribution. US nationwide; Portugal, Spain, Thailand, Italy, China, Romania, Egypt, India, France, Hungary, Mexico, Germany, Norway, Turkey, Pakistan, Malaysia, UK, Luxembourg, Israel, Canada, Switzerland, Taiwan, South Africa, Saudi Arabia, Korea, Singapore and Japan. | Open, Classified |  |  | 283 | Not_Computer
Z-2040-2016 | 74113 | No | GMJ | Histoplasma capsulatum serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Histoplasma Yeast Complement Fixation Antigen Dilute Ref# H30150 |  |  | May 09, 2016 | 2016 | Immuno-Mycologics, Inc | The device was found to have bacterial contamination. | Process control | Immuno-Mycologics, Inc sent an Urgent Medical Device Recall letter dated May 9, 2016, to all affected customers. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their stock and destroy affected stock on hand. Immediately complete the attached Acknowledgement and Receipt Form even if they do not have affected stock. Customers who require replacement indicate the number of vials needed in the attached Acknowledgement Form. Customers with questions were instructed to call 1-405-360-4669. |  | 467 | 467 units | Nationwide Distribution to UT, MA, OH, IN, NC, NY, VA, CA, IA and KY. | Terminated | January 10, 2017 | 246 | 467 | Not_Computer
Z-2041-2016 | 74036 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Healthcare ADVIA Centaur Tnl-Ultra Assay . In-Vitro Diagnostic for quantitative determination of cardiac troponin For in vitro diagnostic use in the quantitative determination of cardiac troponin I (cTnl) in serum, heparinized plasma, or EDTA plasma using the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT/ADVIA Centaur CP systems |  |  | May 02, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Exhibits a greater than 10% change in results in samples with biotin levels up to 10 ng/mL (41 nmol/L). Instructions for Use (IFU) states that specimens that have up to 10 ng/mL (41 nmol/L) of biotin | Component design/selection | Siemens Healthcare sent an Urgent Medical Device Correction letter dated May 2, 2016, The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to continue use of the affected product, review the letter with their doctor, complete and return the Field Correction Effectiveness Check attached to the letter. If customers received any complaints of illness or adverse events associated wit the affected product they should immediately contact their local Siemens Customer Care Center or their local Siemens technical support representative. Customers with questions should contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 508-668-5000. |  | 16033 | SMN 10317708 = 16434 kits and SMN 10317709 = 16033 kits | Worldwide Distribution - US (Nationwide) and Internationally to Albania AustriaBelgium Bulgaria Burkina Faso Canada Croatia Cyprus Czech Republic Denmark Egypt Estonia Finland France French Polynesia Georgia Germany Great Britain Greece Guadeloupe Hungary Iceland Italy Kazakhstan Latvia Lithuania Luxembourg Macedonia Malta Martinique Netherlands Norway Poland Portugal Romania Russian Federation Serbia Slovakia Slovenia Spain Sweden Switzerland Tunisia Turkey United Arab Emirates and Uzbekistan. | Open, Classified |  |  | 16033 | Not_Computer
Z-2044-2016 | 74241 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Crit Line in a Clip (CLiC). CliC P/N: CL 10041001 |  |  | May 24, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | Use of the Blood Pressure Graph located on the Crit-Line screen of the 2008T Hemodialysis Machine may result in a frozen screen, possibly causing the blood pump to stop. | Software design | Fresenius Medical Care Renal Therapies Group, LLC issued Recall Letter issued 5/24/16 via certified mail. Clinics actively using the CLiC option have been verbally communicated to stop using the CLiC option at . The other clinics that did not activate the CLiC option have been verbally communicated to not activate the CLiC option at this time. A Software update to all 2008T Hemodialysis machines will be required.A reply form is included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.For technical Inquiries, please contact Technical Services at 1-800-227-2572. Additional medical concerns or questions, please contact the Medical Information and Communication line at 855-616-2309 or website www.fresenius-medinfo.com. |  | 875 | 875 units | Distributed US (nationwide) | Open, Classified |  |  | 875 | Computer
Z-2042-2016 | 74230 | No | GAW | Nonabsorbable polypropylene surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sharpoint Polypropylene Blue Monofilament 6"/15cm 9-0 USP, Product No. J2558N |  |  | May 20, 2016 | 2016 | SSC- Surgical Specialties Corporation | Surgical Specialties is recalling Sharpoint Polyproplyene Blue Monofilament 6"/15cm 9-0 USP because it may contain Black Nylon 9-0 suture. | Process control | A customer notification letter dated 5/20/16 was sent to all customers to inform them that Surgical Specialties is recalling Sharpoint Polyproplyene Blue Monofilament 6"/15cm 9-0 USP because it may contain Black Nylon 9-0 suture. Customers are instructed to complete the attached customer acknowledgement form and return via fax to Surgical Specialties' Quality Assurance department at (610) 404-4010 or return via email to kknappenberger@surgicalspecialties.com. Customers with any questions are instructed to contact Kelly Knappenberger at (484) 220-2402 or Kris Liberacki at (262) 994-8083. |  | 36 | 36 boxes (432 units) | TX, ME, MA, MI, OH, KS, MN, KS, TN, GA | Open, Classified |  |  | 36 | Not_Computer
Z-2005-2016 | 74192 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT-X 4R, MATERIAL NUMBER 10248670 |  |  | April 22, 2016 | 2016 | Siemens Medical Solutions USA Inc. | Siemens Medical Solutions, Molecular Imaging has become aware of a potential for unexpected SUV values if a non-Siemens phantom is used for calibration. | Nonconforming Material/Component | Siemens sent a Customer Safety Advisory Letter on April 22, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local Siemens representative at the contact numbers provided in the letter. |  | 83 | 83 units | Worldwide Distribution - US Nationwide | Terminated | May 08, 2017 | 381 | 1075 | Not_Computer
Z-2045-2016 | 74191 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Myriad handpiece |  |  | February 24, 2014 | 2014 | Nico Corp. | This recall has been initiated due to a defective component within the handpiece which willcause the handpiece to fail. The failure will result in the inability to cut tissue with thehandpiece, rendering it non-functional. Should this failure occur, a replacement Myriadhandpiece will be required. If a replacement handpiece is not available, alternate instrumentation for tissue removal will be necessary. | Device Design | NICO sent an Urgent: NICO Myriad Handpiece Recall Notification on February 24, 2014, to all affected customers requesting return of the affected product. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this matter, please call (317) 660-7118, ext. 104. |  | 63 | 63 units | US Distribution to the states of : OR, IL, IA, NC, TX, CA, GA, OH, WI, AR, MI, MD, IN. | Completed |  |  | 63 | Not_Computer
Z-2043-2016 | 74372 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fujifilm Synapse PACS software version 4.4.000, |  |  | May 10, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Synapse cannot display image files, DICOM SR files, and/or Annotation files. The "Image Not Loaded" message displays instead | Software design | The firm, FujiFilm, sent an "MEDICAL DEVICE CORRECTION" letter dated May 10, 2016 via Federal Express to its customers. The letter describes the product, problem and actions to be taken.The customers were instructed to: Extend the Ad hoc deletion retention period or discontinue use of Ad Hoc Deletion, and complete and submit the attached Customer Response Form to Fujifilm via fax to (203)-251-7863. Affected systems will be corrected via on-site visits by FMSU Customer Service Engineers upgrading to Version 4.4.100.Questions or concerns, contact the FUJI FILM Customer Support Center at 1-888-FUJIMED Option #1 . |  | 27 | 27 units | US Distribution to states of: CA, FL, KS, NE, NY, and PA. | Open, Classified |  |  | 27 | Not_Computer
Z-2047-2016 | 74015 | No | MCF | Manual toothbrush. | Dental | Dental | 510(K) Exempt | Germ Terminator, Toothbrush Steam Sterilization for Oral Health. Model GT100. The Germ Terminator uses steam heat to sanitize two manual toothbrushes or two toothbrush heads. Water is poured into a reservoir in the device's housing. The steam cycle is followed by a drying cycle. |  |  | April 05, 2016 | 2016 | Ceg Enterprises Llc | CEG Enterprises, LLC is recalling Germ Terminator Toothbrush Sanitizer GT100 due to the following issues: UL label believed to be not genuine; discoloration of power cords (could indicate some degradation); reports of excessive heat. | Unknown/Undetermined by firm | CEG Enterprises sent an Urgent Germ Terminator Toothbrush Sanitizer GT100 Recall letter dated April 19, 2016, via UPS and email to all affected customers. The letter identified the affected product, the reason for recall, stated to quarantine product, and to return all remaining product. The letter also instructed that Tuesday Morning retail stores post the recall notification letter on the retail shelves where the product was offered for sale, as well as on their website. Consumers are encouraged to return the device to a Tuesday Morning retail store. Questions can be directed to 612-337-1821 |  | 100548 | 100,548 units | Nationwide Distribution through Tuesday Morning retail stores. | Open, Classified |  |  | 100548 | Not_Computer
Z-2046-2016 | 74208 | tray, impression, preformed | EHY | Preformed impression tray. | Dental | Dental | 510(K) Exempt | 6534-08N InterActiveÂż Implant Open-Tray Transfer Narrow |  |  | May 19, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | The incorrect sized product, InterActive Implant Open-Tray Transfer Wide, 6534-08W, was packaged instead of the InterActive Implant Open-tray Transfer Narrow, 6534-08N, causing a possible issue with the impression. This discrepancy may lead to the possibility of an oversized crown fabrication. | Packaging | The firm sent out customer notificaiton letters on 05/19/16 via FedEx.Customers were asked to reivew the table to determine if they have any of the affected product in inventory. Customers were being informed that the proper product was being sent with the notificaiton. Customers were asked to complete and return the the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized Implant Direct Sybron Manufacturing distributor, the firm requests that they identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours of receipt of this notification in order to provide the customers with the correct tool. |  | 48 | 48 | Worldwide Distribution - U.S. Nationwide in the states of OR, NE, VA, FL, MO, WA, TX and the countries of CZ, GB, NL, MF, DK, IT, IE, DE | Open, Classified |  |  | 48 | Not_Computer
Z-2048-2016 | 74276 | Yes | NIP |  |  | Cardiovascular | PMA | EV3 EverFiex" Self-expanding Peripheral Stent System, Sterilized using Ethylene Oxide, Rx Only, PRB35-08-040-080. |  |  | May 23, 2016 | 2016 | Ev3, Inc. | Medtronic is issuing a medical device recall to return product potentially associated with an outer boxmislabel issue for two Protege EverFiex" Self-expanding Peripheral Stent Systems, model PRB35-08-040-080, from lot A164580. The size on the outer box is labeled as 8x40x80mm while potentially containing a6x40x120mm device as indicated by both the inner pouch and device labels. This potential issue is limited totwo EverFiex stent system devices from lot A164580 that were repackaged by Medtronic and distributed.This does not affect any other Medtronic product or implantable devices. | Labeling mix-ups | Medtronic sent an "Urgent Medical Device Recall" letter dated May 2016, to all affected customers. The letter described the problem and the product involved in the recall. Consignees were instructed to identify, quarantine and return unused product to Medtronic. Consignees were requested to complete the Customer Confirmation Certificate and email it to RS.CFQFCAÂżMedtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. Customers with questions were instructed to call 651-367-0612.For question regarding this recall call 763-398-7000. |  | 2 | 2 | Nationwide Distribution to LA and WI. | Open, Classified |  |  | 2 | Not_Computer
Z-2049-2016 | 74226 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PrismaSate BGK2/0 Dialysis Solution for Continuous Renal Replacement Therapy |  |  | May 23, 2016 | 2016 | Baxter Healthcare Corp. | Presence of leaks near top of the PrismaSate bags | Other | Baxter Healthcare Corporation sent an Urgent Product Recall communication to affected customers via U.S.P.S. first class mail on 05/23/2016. |  | 10054 | 10,054 units | Distributed to Columbia and the following US states: AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, WI | Terminated | March 28, 2017 | 309 | 10054 | Not_Computer
Z-2056-2016 | 63990 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids |  |  | February 04, 2013 | 2013 | Hospira Inc. | Frequent Proximal Occlusion Alarms (POAs) occurring on Symbiq devices with list numbers 16026 and 16027. When there are false alarms, the pump will continue to alarm even though there is no occlusion present. False POAs could result in a delay/interruption in therapy, and require the clinician to reset the alarm multiple times or to replace the pump. | Software design | Hospira sent letters on May 2, 2013, to customers informing them of the discontinuation of sales, service, and support for the Symbiq pump were sent via traceable mailing to all affected customers. For questions regarding this recall call 224-212-4892. |  | 15285 | 15,285 | Worldwide Distribution - USA (nationwide) | Completed |  |  | 35596 | Not_Computer
Z-2051-2016 | 74445 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ViperWire Advance Peripheral Guide Wire, Sterilized with Ethylene Oxide, Rx Only. Part number 72023-02, Model Number VPR-GW-17. Five (5) guide wire pouches in one (1) guide wire box. For use with CSI Peripheral Orbital Atherectomy Devices (OAD), including Stealth 360 and Diamondback 360. Cardiovascular. |  |  | June 13, 2016 | 2016 | Cardiovascular Systems Inc | The pouch label was missing the use by date (UBD) of 2018-04. | Process control | Consignees were sent on 6/13/16 a Cardiovascular Systems, Inc. "Urgent Correction" letter date June 10, 2016. The letter described the problem and the affected product. Advised consignees to set affected product aside for their sales representative which will be applying the correct pouch label to the affected product. Requested consignees to fill and sign the Customer Acknowledgement Form that the sales representative will be presenting to them. For questions contact Quality at Cardiovascular Systems, Inc. 1-651-259-2068 1225 Old Highway 8 NW Saint Paul, MN 55112. |  | 190 | 190 units (38 - 5 pack shelf cartons) | US Nationwide distribution. | Open, Classified |  |  | 190 | Not_Computer
Z-2050-2016 | 73608 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Bigger-Better-Bladder With 3/8" ID tubing, ITEM BBB38 |  |  | March 03, 2016 | 2016 | Circulatory Technology Inc | The seal between the balloon and the housing may leak. | Device Design | Circulatory Technology Inc. sent an Urgent Product Recall notices dated March 3, 2016, to their customers via email. Circulatory Technology Inc. has identified this action as a voluntary recall, however they did not withdraw the product from the market. Instead, Circulatory Technology Inc. has advised customers to check the integrity of the seal between both the inlet and outlet tube and the housing as per the Instructions For Use. In addition, for this lot, Circulatory Technology Inc. recommends that you bend the inlet as well as the outlet tube away from the housing before conducting the leak test. There should be no visible separation between the tubing and the housing. Customers with questions can call 516-624-2424. |  | 72 | 72 units | US Distribution to the states of: CA, IN, MI, NY, PA, TX &amp; VA | Open, Classified |  |  | 72 | Not_Computer
Z-2054-2016 | 73395 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Sodium Chloride Inhalation Solution, USP 7%, 60 x 4 mL Sterile Unit-Dose Vials, Rx Only, NDC 0487-9007-60. |  |  | November 16, 2015 | 2015 | Nephron Pharmaceuticals Corp. | Lack of Assurance of Sterility: Internal dye ingress test failure to assure the integrity of the vial. | Process control | Nephron Pharmaceuticals Corporation sent an Urgent Drug Recall letter to all affected customers on November 16, 2015. The letter identified the product, the problem, and the action to be taken by the customer. NPC requested customers remove the affected lots from their inventory, including store shelves. Customers with questions should call 1-844-374-6841 |  | 13692 | 13,692 cartons | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, FL, GA, IL, KY, LA, MA, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA, WI, WV and Internationally to Canada. | Terminated | January 09, 2017 | 420 | 24900 | Not_Computer
Z-2052-2016 | 74378 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | June 17, 2016 | 2016 | Elekta, Inc. | It is possible that that the selected Plan-of-the Day could be accidently changed by scrolling the mouse wheel up and down and inadvertently selecting another plan. | Software design | Elekta sent an Important Field Safety Notice to all customers on June 17, 2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. A software patch will be issued for Mosaiq version 2.62 in order to the correct the defect. For further questions please call (404) 993-5886. |  | 69 | 69 | US Distribution including Puerto Rico and to the states of : AK, AZ, CA, FL, GA, IL, IN, LA, MD, MI, MO, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and to the countries of : Canada, Denmark, Finland, Germany, Israel, Japan, Netherlands, Norway, Sweden, United Kingdom | Open, Classified |  |  | 69 | Computer
Z-2072-2016 | 73894 | No | CAL | Flexible laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Intubation Scope |  |  | April 06, 2016 | 2016 | Pentax Medical Company | Pentax has become aware that operational/cleaning accessories and therapeutic devices (e.g., clips, stents, balloons, etc.) can become lodged in an endoscope's instrument channel. | Device Design | Pentax sent an Urgent Field Correction letter dated April 6, 2016 to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were reminded to carefully check that all accessories are intact, that no parts have fallen off and become lodged within the endoscope's instrument/suction channel and to ensure that any therapeutic devices passed through the instrument channel are accounted for after use. If the instrument/suction channel does become blocked or clogged due to the accumulation of debris, an accessory that cannot be removed, or other cause, do NOT attempt to correct the blockage or continue to use the endoscope. Customers were instructed to contact their local PENTAX Medical Service facility to have the endoscope repaired. Customers with questions were instructed to call 800-431-5880, ext 2064. |  | 1220 | 1,220 units | Worldwide Distribution - US (Nation Wide) including Puerto Rico and Internationally to Denmark, Guam | Open, Classified |  |  | 31293 | Not_Computer
Z-2057-2016 | 73584 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | dynarex Specimen Containers, Screw-On, Leak-Resistant Lid, Tamper-Evident Seal, Sterile, Single Use Only |  |  | March 03, 2016 | 2016 | Dynarex Corporation | Exterior carton is mis-labeled as Sterile. Non-sterile, bulk specimen containers were packed in incorrect exterior carton. | Labeling False and Misleading | On March 2, 2016, dynarex Corporate Headquarters distributed Voluntary Device Recall notices and Recall Verification Forms to their customers via email. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to examine their stock immediately to determine if they have any of the lots listed on hand. If so, discontinue distributing and quarantine the product immediately. Promptly contact the Dynarex Credit &amp; Return representative Avi Celnick (845-365-8200 Ext 6644) to arrange for product return. Customers account will be credited for all returned products. All customers are advised to complete the Voluntary Recall Verification Form and returm immediately via fax or e-mail to Vijay Sachdev (Fax 845-365-8238; v.sachdev@dynarex.com). |  | 9975 | 399 cases/25 sleeves/20 cups &amp; lids per sleeve | US Nationwide in the states of AL, AR, FL, GA, MA, NC, NY, NJ, NH, PA, OH, TX &amp; VA; International: Canada | Open, Classified |  |  | 9975 | Not_Computer
Z-2084-2016 | 74224 | No | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | K-wire, Smooth Double-pointed 0.8 x 90mm, a component of the METASURG Capture Screw System |  |  | May 20, 2016 | 2016 | Integra LifeSciences Corp. | The affected K-wires do not meet specification requirements. Specifically, a) K-wires were made of 316LVM stainless steel instead of 304 stainless steel. b) K-wires have passivated surface finish instead of electro-polish surface finish. c) K-wires have incorrect length and diameter. | Nonconforming Material/Component | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 20, 2016, to all affected consignees regarding the recall via traceable courier service. The letter identified the product, the problem, and the action to be taken by the customer. Customers who have supplied affected product should forward a copy of the Recall Notice to them. Customers should review their inventory for affected product, complete the form as to whether they have any affected product or not, return the form by email or fax as indicated on the form. Customers with questions should contact Customer Service at 1-800-654-2873. For questions regarding this recall call 609-936-6822. |  | 182 | 182 units | Nationwide Distribution to OK, CA, TX, WA, GA, NY, MS, NY, PA | Open, Classified |  |  | 182 | Not_Computer
Z-2064-2016 | 74142 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris System PC Unit Model 8015 with software versions 9.17 and 9.19; |  |  | May 12, 2016 | 2016 | CareFusion 303, Inc. | A patient weight can be populated incorrectly under certain conditions when using the RESTORE feature to restore infusions running on the Alaris LVP module model 8100 and the Alaris Syringe module model 8110. | Software design | An Urgent Medical Device Recall Letter was sent to customers on 5/31/16 to inform them that BD had identified an issue with the Alaris System PC unit model 8015 with software versions 9.17 and 9.19. Firm has received reports that a patient weight can be populated incorrectly under certain conditions when using the RESTORE feature to restore infusions running on the Alaris LVP module model 8100 and the Alaris Syringe module model 8110. The letter informs the customers of the potential risk and the required action for users. Customers are informed that BD will contact them within 60 days to initiate the scheduling process for the software upgrade installation. |  | 69693 | 69,693 units | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 69693 | Computer
Z-2081-2016 | 74282 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | CENTURION Vision System |  |  | May 24, 2016 | 2016 | Alcon Research, Ltd. | The grounding pin within the Power Entry Switch on certain systems was susceptible to detachment when the AC power cable is removed. | Nonconforming Material/Component | Alcon has already notified the affected customer verbally on May 24, 2016, and via written correspondence on June 1, 2016. The device was corrected on May 26, 2016, and the Power Entry Switch that was removed from the customers CENTURION Vision System did not exhibit a grounding pin susceptible to detachment. For further questions, please call (817) 293-0450. |  | 1 | 1 unit | US Distribution to Illinois | Terminated | April 13, 2017 | 324 | 1 | Not_Computer
Z-2063-2016 | 74234 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Liquid stable 2 part Homocysteine Reagent, for in vitro quantitative determination of total homocysteine in human serum and plasma - H7575-57 kit includes one Rl and one R2 bottle and one bottle of calibrator level 1 and one bottle of calibrator level 2. |  |  | April 03, 2015 | 2015 | Medtest Holdings, Inc. | Deterioration in performance of the Liquid Stable (LS) 2-Part HOMOCYSTEINE REAGENT resulting in under recovery.The manufacturer has indicated that the HOMOCYSTEINE REAGENT lot listed above shows deterioration in assay performance manifesting itself by under recovery of Homocysteine in Controls and Patient Samples.The cause of this deterioration has been attributed to a loss of reductant activity in Reagent 1.Testing performed indicated that the affected lot is at the limit of acceptable clinical performance. This observation was consistent across the reportable range of the assay. For this reason, do not continue to use affected reagent lot. Test data generated to date will be acceptable where validity control specifiCations have been met. | Unknown/Undetermined by firm | On 4/3/2015, Urgent Field Safety Notice notifications were sent to the affected customers via email with instructions for identifying and destroying the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 13 | 13 | MI, MN, NJ, IL, NY.US Only. | Completed |  |  | 13 | Not_Computer
Z-2085-2016 | 74275 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Saw Sagittal Catalog Number: ES 300; Saw Sagittal Refurbished Catalog Number: ES300-R; Saw Saggital, Finger Catalog Number: ES310; Saw Saggital, Finger Refurbished Catalog Number: ES310-R; Used to dissect bone in surgical applications. |  |  | May 13, 2016 | 2016 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Medtronic Powered Surgical Solutions testing of the Sagittal Saw Rocker Arm Seal identified a potential for leakage of fluid into the motor during use. | Device Design | Recall notification letters were sent to consignees on 05/13/2016. |  | 368 | 368 devices | Worldwide Distribution-Including the states of AL, AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, MA, MD, NC, NH, NY, SC, TN, TX, and UT. | Terminated | June 13, 2017 | 396 | 368 | Not_Computer
Z-2062-2016 | 66372 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | *** 1) Product labeled in part: 19683-28; LIFESHIELD; MICRODRIP SOLUSET WITH 1.2 MICRON FILTER 150ML BURETTE PLUMSET, CONVERTIBLE PIN, CLAVE SECONDARY PORT, OPTION-LOK, NON-DEHP; NOMINAL LENGTH: 118 IN; PRIMING VOLUME: 14 ML; |  |  | September 09, 2013 | 2013 | Hospira Inc. | Intravenous fluid may leak at the CLAVE secondary port of Hospira PlumSet IV administration sets. | Device Design | Hospira sent an URGENT MEDICAL DEVICE CORECTION letters dated September 9, 2013 to customers. The letter identified the affected product, problem and actions to be taken. the letter included additional instructions on the proper handling/use of the CLAVE devices. The letters request that all incidents of separation or breakage of CLAVE from the secondary port of the cassette be reported to Hospira by calling 1-800-441-410 (Monday - Friday, 8:00 AM - 5:00 PM CST) or by e-mail at ProductComplaintPP@hospira.com. Customers are further requested to complete and return a response form included wih the URGENT MEDICAL DEVICE CORRECTION letter and to notify any customers of this advisory. Hospira will also be supplementing the instructions for CLAVE use as an interim measure until a new method for bonding the CLAVE to the secondary port of the Plum cassettes can be devised and implemented. |  | 3920932 | 3,920,932 units in total | US (Nationwide) Distribution | Completed |  |  | 10212296 | Not_Computer
Z-2099-2016 | 74408 | No | NYQ |  |  | Immunology | PMA | Dako HER2 CISH pharmDx Kit, product code: SK 109 |  |  | June 10, 2016 | 2016 | Dako North America Inc. | An incorrect HER2 probe concentration was used when producing the probe component (Vial 3) in this kit lot. This has resulted in the HER2 probe concentration being too low. | Process control | Dako sent a recall notification dated June 10, 2016 to all customers. The letter identified the affected product, problem, the investigation findings, the affected results, and the actions to be taken. Customers are instructed to complete and return the enclosed device recall form to Dako QA Vigilance by Dako.dkvigilance@agilent.com. Customers with questions regarding the notification, are instructed to contact their sales representative. |  | 31 | 31 kits | Worldwide Distribution - US Nationwide in the states of: PA, CA, AL, NC and countries of Austria, Bosnia and Herzegovina, Belgium, Bulgaria, Brazil, Canada, Switzerland, Germany, Spain, France, Greece, Italy, Macedonia,Qatar, Romania, Serbia, Sweden, and Turkey. | Open, Classified |  |  | 31 | Not_Computer
Z-2100-2016 | 74198 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK CR Plus and/or LIFEPAK EXPRESS are Automatic External Defibrillator(s) and are non-wearable. |  |  | May 25, 2016 | 2016 | Physio-Control, Inc. | LIFEPAK CR Plus Automated External Defibrillators (AED) or LIFEPAK EXPRESS AED may fail to initiate voice prompts when the ON/OFF button is pressed and the lid is opened due to an internal component (reed switch) that can intermittently become fixed in the closed position. A defibrillator in this condition will fail to deliver a shock, with the potential result that therapy is not delivered and a patient is not resuscitated. | Component change control | The firm, Physio Control, sent an "URGENT MEDICAL DEVICE CORRECTION-ACTION REQUIRED"- LifePak CR Plus AE and LifePak Express AED letter and the Confirmation sheet, dated May 2016, to US customers on 5/25/16. Physio will notify international consignees during the week of 6/13/16.The letter described the product, problem and the actions to be taken. The customers were instructed to-URGENTLY bring this letter to the immediate attention of the person(s) responsible for maintaining/monitoring your Automated External Defibrillators (AEDs); follow the instructions on the Confirmation Sheet that outline specific actions to take for the serial numbers listed, and if you have a routine check process, please continue this process. If you have not established a routine check process, please refer to section 5 of the Operating Instructions for recommended actions. In addition, customers should follow the instructions on the Confirmation Sheet and submit this form to Physio by: -Fax to: 1-866-448-9567-Email to: rsrecall@physio-control.com -Mail to: Physio-Control, Inc. P.O. Box 970006, Dept. 15N Redmond, WA, 98073-9706Physio-Control will contact customers with LIFEPAK CR Plus and LIFEPAK EXPRESS AEDs that contain the affected reed switch assembly. A device correction including provision of loaner devices and replacement of the reed switch component will be arranged for all affected devices. Customers who have any questions regarding this notification, please call Physio-Control at 1-800-442-1142, 6:00 a.m. to 4:00 p.m. (Pacific), Monday  Friday. |  | 25178 | 25178 units (10,418 in US and 14,760 outside US) | Worldwide Distribution-US (nationwide) including Guam and Puerto Rico and countries of: Argentina, Austria, Australia, Bahamas, Bangladesh, Belgium, Brazil, Brunei Darussalami, Canada, Chile, China, Croatia, Cyprus. Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Lithuania, Luxembourg, Malta, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan (Province of China), Turkey, United Arab Emirates, United Kingdom, New Caledonia, and New Zealand,. | Open, Classified |  |  | 25178 | Not_Computer
Z-2094-2016 | 74271 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION(TM) HK FEMORAL ASSEMBLY, SIZE 5 LEFT, A/P 62 MM, M/L 70 MM, REF 71421375 |  |  | May 26, 2016 | 2016 | Smith &amp; Nephew, Inc. | The surface inside of the femoral component, where cement is applied for adhesion, is below specification. | Process control | The firm sent an Urgent Product Recall letter dated May 26, 2016 to customers, via email and FedEx Overnight delivery . The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. Inspect inventory and locate any unused devices from the listed product and batch numbers, and quarantine them immediately.2. Complete the Inventory Return Certification Form on the following page, indicating the quantities that need to be returned. If you do not have product to return, place an x in the column No Product to Return. Include your contact information in the spaces provided at the bottom of the form.3. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return.4. Return any affected product to the address listed on the Inventory Return Certification Form. Indicate the RA number on your return shipment. |  | 10 | 10 units | Worldwide Distribution and US Nationwide including Puerto Rico and the countries of Australia, Austria, Belgium, Canada, France, Germany, Italy, Netherlands, South Africa, Sweden, and UK. | Open, Classified |  |  | 171 | Not_Computer
Z-2089-2016 | 73629 | No | EBI | Denture relining, repairing, or rebasing resin. | Dental | Dental | 510(k) | IMPAK Liquid Acrylic Resin sold under the following labels: |  |  | March 02, 2016 | 2016 | CMP Industries, Llc | Possible presence of foreign material (rust). | Material/Component Contamination | CMP Industries/Nobilium sent an Urgent IMPAK Recall letter dated March 2, 2016, to all affected customers via certified mail with return receipt in an attempt to verify if customers were in possession of the affected product. Customers were advised to immediately examine inventory and to quarantine product subject to the recall. In addition, if customers further distributed the product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of the notice. If customers have any of the product in their posession subject to recall, please contact the firm by calling 518-434-3147, ext 144, Monday - Friday 8-4:30 pm, EST. |  | 2596 | 236 pieces (25 pieces/11 oz, 211 pieces/1 Quart) | Worldwide Distribution - US including CA, MI, LA, OH, NY, IN, WA; Internationally: to CA | Terminated | March 16, 2017 | 379 | 2596 | Not_Computer
Z-2088-2016 | 74390 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza, Picture archiving and communication system (PACS), Model Numbers - 10863171, 10863172, 10863173, |  |  | May 09, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software error in previous software versions in which two references for the same image may exist in the database. | Software design | Siemens sent an Customer Information letter dated May 10, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter with software instruction to customers informing them about an upcoming software VB10B under(SY003/16/P) update change and issue that the new software version will address. For further questions, please call (610) 448-6478. |  | 74 | 74 systems | US Distribution to the states of : AL, CA, CT, FL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NE, NY, OH, OR, PA, TX, WI | Terminated | March 20, 2017 | 315 | 74 | Computer
Z-2087-2016 | 74280 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hettich Centrifuge Modules: |  |  | May 11, 2016 | 2016 | Siemens Healthcare Diagnostics Inc. | Premature bucket failures | Under Investigation by firm | Siemens mailed to customers an Urgent Medical Device Correction (UMDC#LAI16-02.A.US) Urgent Field Safety Notification letter informing customers of the problem with recommendations to follow the existing maintenance instructions in the product operator manuals as well as a new requirement to spin at cycle speeds not exceeding 3500 rpm. |  | 487 | 487 | Nationwide: AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA, WI, WV. | Open, Classified |  |  | 487 | Not_Computer
Z-2097-2016 | 74061 | No | EBI | Denture relining, repairing, or rebasing resin. | Dental | Dental | 510(k) | IMPAK Liquid Acrylic Resin sold under the following labels: 1) IMPAK ELASTIC ACRYLIC RESIN LIQUID FOR MOUTHGUARDS &amp; BRUXISM SPLINTS, Contents: 1 Quart (946 ML), REF 3306, Rx ONLY; 2) IMPAK ELASTIC ACRYLIC RESIN LIQUID FOR MOUTHGUARDS &amp; BRUXISM SPLINTS, Contents: 11 oz (325 ML), REF 3748, Rx ONLY, 3) BRUX-EZE AN ELASTIC ACRYLIC RESIN FOR THE CONSTRUCTION OF DENTAL TREATMENT APPLIANCES, 1 Quart Liquid, PN 7450; Indicated for relining a denture surface. |  |  | April 28, 2016 | 2016 | CMP Industries, Llc | Possible presence of foreign material (rust). | Material/Component Contamination | On April 28, 2016, Nobilium, Division of CMP Industries LLC distributed Urgent Recall notices and Response Forms to their customers via courier service. Customers were advised to immediately examine inventory and to quarantine product subject to the recall. In addition, if customers further distributed the product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of the notice. If customers have any of the product in their possession subject to recall, please contact the firm by calling John Nicpon at 518-434-3147, ext. 144. Hours of operation are Monday - Friday 8-4:30 pm, EST. Return completed recall response forms to nicpon@nobilium.com. |  | 77 | 77 units Total | Domestic - MI, MN, NJ, NM, NY &amp; PA; International - Canada | Open, Classified |  |  | 77 | Not_Computer
Z-2086-2016 | 74173 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Cadence Size 5 Left Tibial Tray Implant, designed to treat ankle arthritis through replacement of the ankle joint with a prosthesis, thereby reducing pain, restoring alignment, and allowing for movement at the replaced joint. Catalog Number: 10207205; Lot Number #QJ0248 |  |  | April 29, 2016 | 2016 | Integra LifeSciences Corp. | As a result of an internal review of all labeling for the Cadence System, it was determined that the outer package label for the Size 5 Left Tibial Tray Implant reflects the material composition incorrectly. The labeling incorrectly reflects the material as Cobalt Chromium instead of Titanium. | Labeling mix-ups | On 4/29/16, US consignees were notified by electronic mail delivered by traceable method. |  | 12 | 12 units | NC and PA | Open, Classified |  |  | 12 | Not_Computer
Z-2105-2016 | 73918 | tips and pads, cane, crutch and walker | INP | Cane, crutch, and walker tips and pads. | Physical Medicine | Physical Medicine | 510(K) Exempt | Universal Quick Adjust under the following labels: 1) Aluminum crutch, adult, 1 pair per case, 2) Aluminum crutch, tall adult, 1 pair per case |  |  | February 17, 2016 | 2016 | Medical Depot Inc. | Tip crutch failure involving the bottom of the metal portion which comes in contact with ground | Nonconforming Material/Component | On February 17, 2016, drive DeVilbiss Healthcare issued a preliminary notice their customers informing them of the situation via courier service. Secondly, On February 25, 2016, drive DeVilbiss Healthcare distributed another notice informing customers that the potential issue was confirmed to a limited number of lots. Customers were advised to quarantine all products in inventory as replacement products will be provided. The company notified customers that they will commence shipping orders for the product on Monday, February 29, 2016. Customers were also advised that a substantial inventory of replacement crutch tips for distribution will begin on or about March 7, 2016. On March 11, 2016, drive DeVilbiss distributed a third notice to their customers via courier service advising them of customer reimbursement for time and labor costs associated with this action. Drive Devilbiss Healthcare will reimburse based on the total number of tips requested to be shipped. Lastly, on May 12, 2016, drive DeVilbiss distributed a fourth Urgent Medical Device notice and response form to their customers via email. drive DeVilbiss advised customers to: 1) quarantine any stock that has not been inspected already and compare serial numbers of stocks to list provided, 2) any product received after 3/7/16 has been fully inspected and has an approved crutch tip that is free of defect, 3) if products were further distributed, contact those customers who received product between 7/2015 to 12/2015 and determine if items are still being used, 4) If the customer is still using this item, contact your sales representative so the return and replacement of this product can be processed and 5) complete and return the response form vial email. Customers with questions can contact by telephone 516-998-4600 ext. 4422. |  | 123125 | 123,125 pairs | Nationwide Distribution | Open, Classified |  |  | 123125 | Not_Computer
Z-2114-2016 | 74511 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128, model #728323; Intended to produce cross- sectional images of the body. |  |  | June 21, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Supplier design change contributed to a thinner plastic which was used as part of the cover on the devices. The thinner plastic cover may crack causing injury to patients, operators, by-standers and service personnel. | Nonconforming Material/Component | Philips is informing the consignee of this issue through:Customer Information Letter (CIL) 72800656 .International consignees will be contacted by Philips Healthcare Global Market representatives, who will distribute the CIL. A Field Service Engineer (FSE) will contact the customers to schedule a time to conduct the appropriate correction through Field Change Order (FCO) 72800656 once it is available. This FCO will address the above identified issue.A response form sent to the customer via certified mail. This package will also contain a return receipt. |  | 14 | 14 units | Distributed in China, Denmark, Germany, Hungary, Italy, Philippines, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 29 | Not_Computer
Z-2110-2016 | 74507 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Big Bore Radiology tomography X-ray system scanner, model number 728244. |  |  | June 14, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Socket Head Cap Screws (SHCS) used in the Z-Axis Adjustment Plate were found to be made of stainless steel rather than the specified alloy steel. | Nonconforming Material/Component | Philips Medical Systems (Cleveland), Inc. sent a Customer Information letter dated June 14, 2016, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. No action is required of the customer for this issue. Systems that may have been manufactured with these stainless steel fasteners still meet the safety specifications required for the adjustment cradle and may continue to be used clinically. For further information or support customers were instructed to call 1-800-722-9377.For questions regarding this recall call 404-483-2015. |  | 89 | 89 | Worldwide distribution to Belgium, China, Egypt, Germany, India, Korea, Republic of, Myanmar, Nigeria, Peru, Russian Federation, Rwanda, Singapore, Spain, Sweden, Taiwan, Thailand, and United Kingdom | Open, Classified |  |  | 136 | Not_Computer
Z-2107-2016 | 74422 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | SYNGO Breast Care, visualization and image enhancement tools to aid radiologist in the review of digital Mammography images and tomosynthesis datasets. |  |  | June 06, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software issues. Siemens is voluntarily initiating a recall after they became aware of the following system behavior: 1) At times, the view and laterality marker is overlaid by patient demographic information, thus becoming difficult to read.2) The facility name and address are not shown per default (in the current software versions this information is displayed only when the reader chooses an appropriate image text display mode). | Other | Siemens mailed customers a letter to inform them about improvements that were made to the SYNGO Breast Care reading software to correct certain behaviors that were being seen during use. Customers were asked to upgrade their systems to this new version. |  | 29 | 29 units | Distributed to: CA,NY,TX,OH,CO,NY,TX,CA,NE,NJ,TX,FL,IL,TX,MO,CA,PA,FL,NJ,MO,ND | Open, Classified |  |  | 29 | Computer
Z-2108-2016 | 74393 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Free T3 Assay, Ref No. A13422 |  |  | June 08, 2016 | 2016 | Beckman Coulter Inc. | Demonstrate an upward shift in patient results. | Device Design | Beckman Coulter sent an " Urgent Medical Device Recall" letter dated June 8, 2016, to all affected customers. Beckman Coulter has determined through customer feedback and internal testing that the Access Free T3 Reagent lots listed above demonstrate an upward shift in patient results. The preliminary results indicate this upward shift may be related to a June 2015 formulation design change that was introduced to improve the Access Free T3 open reagent pack stability. The upward shift in patient test results, therefore, is expected to be maintained for all future lots. The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions are instructed to contact Customer Technical Support Center at http:www.beckmancoulter.com, call (800) 854-3633 in US and Canada. Customers outside US and Canada, are instructed to contact local Beckman Coulter representative. |  | 205744 | 205,744 kits total (13,358 kits in US) | Worldwide Distribution - US Distribution and to the countries of : Algeria, Australia, Austria, Bangladesh, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Djibouti, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Ghana, Gibraltar, Greece, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Zambia, Zimbabwe | Terminated | April 18, 2017 | 314 | 205744 | Not_Computer
Z-2120-2016 | 74273 | No | GZO | Galvanic skin response measurement device. | Neurology | Neurology | 510(K) Exempt | The ZYTO Tower is the input device to program the software with various virtual items. |  |  | November 23, 2015 | 2015 | ZYTO Technologies, Inc. | ZYTO Technologies Inc. Announces a Voluntary Recall of the ZYTO Tower and ZYTO Select and Elite Software Due to Claims Exceeding the 510(k) Clearance. | Labeling False and Misleading | Consignees were notified on 11/23/2015 via letter and email of an update to the ZYTO Select and Elite Software that would disable all tower functionality.. |  | 1252 | 1252 total | Nationwide. Canadian and other foreign consignees. No VA/govt/military. | Open, Classified |  |  | 1252 | Computer
Z-2118-2016 | 74427 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128-Computed Tomography X-ray system |  |  | April 01, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Ingenuity Core128 systems running software version 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images. A patient rescan may be required if the images cannot be used for interpretation due to the swirl-like artifact. | Software design (manufacturing process) | Philips Medical Systems (Cleveland) Inc. sent an Urgent Field Safety Notice letter to affected customers , via certified letter with receipts returned for tracking purposes. The letter identified the affected product, problem and actions to be taken. Philips Healthcare is implementing software update version 3.5.5 to correct the problem. For questions contact your local Philips representative or local Philips Healthcare office. |  | 19 | 19 | US Nationwide Distribution in the states of CA, FL, ID, KS, MI, NY,OH,TX | Open, Classified |  |  | 262 | Computer
Z-2119-2016 | 74363 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Calibrator Kit 29, REF/Product Code 680 2344, used in conjunction with VITROS Chemistry Products dTIBC Reagent Generation 30 (GEN 30), REF 680 2001, and Assay Data Disk (ADD), REF/Product Code 6801876 (Unique Device No. 10758750001576), Data Release Versions (DRVs) 5870 through 5883 inclusive; IVD --- Ortho-Clinical Diagnostics |  |  | May 24, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics started receiving customer complaints for biased results when using VITROS dTIBC Reagent GEN 30 product. Ortho's investigation confirmed that incorrect calibration mathematics were assigned to the Calibrator Kit 29 Lots 2915 and 2995 supporting dTIBC Reagent GEN 30 product. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated May 24, 2016, via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (Ortho Direct, Drop-Ship and US Federal Government) to inform them of the issue and advised them to discontinue using the affected product until Assay Data Diskette (ADD) Data Realease Version (DRV) 5884 or above was installed on their VITROS System. Foreign affiliates were informed of the issue via e-mail on 5/24/2016 and instructed to notify their consignees of the issue and required actions.Customers with questions were advised to contact Ortho's Customer Technical Services at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 842 | US: 1517; Foreign: 842 | Worldwide Distribution - US (Nationwide, Puerto Rico) and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, Poland, and Venezuela. | Open, Classified |  |  | 842 | Not_Computer
Z-2124-2016 | 74391 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetic Resonance System G-scan Brio (Part Number /Model No.101002000) |  |  | June 06, 2016 | 2016 | Esaote S.p.A. | The device is being recalled because dynamic loading forces from larger patients can fracture the extremity portion of the patient table. Additionally, a weight limit is added to the patient seat for weight-bearing examinations. | Other | Esaote initiated a voluntary recall of the G-scan Brio Magnetic Resonance System on 06/06/2016 with letters sent via certified, due to a potential weakness on the extremity areas of the patient table that could lead to a stress fracture of the table. The firm has instructed customers to refer to the Medical Device Safety Notification provided with the recall notice, date and sign this letter in the space provided on the last page and send it back via email to quality@esaoteusa.com. If customers have questions regarding the recall, please call 1-317-813-6030. For additional information, please contact us by phone during normal business hours, Monday-Friday from 8:00am  5:00pm Eastern Standard Time, via email at quality@esaoteusa.com, or at any time by visiting our website. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online at www.fda.gov/medwatch/report.htm, by regular mail using the postage-paid, pre-addressed Form FDA3500 available at www.fda.gov/MedWatchgetforms.htm or by fax to 1-800-FDA-0178. |  | 40 | 11 USA; 29 OUS | Domestic: FL, TX, OH, MN, NV; Foreign: No customers in Canada;No VA/DOD. | Terminated | May 30, 2017 | 358 | 40 | Not_Computer
Z-1908-2016 | 73461 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | ABG-HM-1 Hummi Micro Draw Blood Transfer Device; Used for blood transfer and collection from peripheral arterial line catheters. |  |  | March 11, 2015 | 2015 | Hummingbird Med | Hummingbird Med Devices, Inc. is recalling the Hummi Micro Draw because of insertion difficulties. | Process design | The firm, Hummingbird Med Devices, Inc., sent a "Safety Alert" on March 3, 2015 and Addendum on March 11, 2015, to all affected customers. The notification sent on March 3, 2015, instructed customers that: 1) the Hummi Micro device is not recommended to be used with the Baxter Interlink device with the Hummi Micro draw blood transfer device as the Interlink T Connector septum may be damaged.2) the use of Nexus Medical Micro-T Split Septum T Connector NMT8046 with all uses of the Hummi Micro Draw Blood Transfer device. The connector has been tested and validated for use in combination with the Ameritus Medical's Hummi Micro-Draw Blood Transfer Device.3) Distributors of the Hummi Micro Draw Device contact and inform all Hummi Micro-Draw Blood Transfer Device users to recommend change to the Nexus Split Septum Micro-T Connector NMT8046 for all uses with the Hummi Micro Draw device. Inservice materials and training on proper Hummi Micro Draw device blunt cannula insertion into the connector's septum will be provided.4) Distributors complete the attached form indicating hospitals and date contacted, name, and contact information of responsible person and resulting outcome. and5) Scan and email the attached form to: jorgehaider@humigbirdmed.com or fax to: 1-949-583-2775. On March 11, 2015, the firm sent out a second email to their distributor to inform customers again that only the split septum T-Connector recommended for use with the Hummi Micro Draw is the Nexus Micro T split septum T-Connector. The Nexus Micro T-Connector is designed to provide improved insertion force and eliminate potential leaking while providing a microbial barrier when used with the Hummi Micro Draw device.The Hummi Micro device manufactured in 2015 now has improved lubrication on the blunt tube which makes insertion into smaller septum opening of the Mexud Micro T-Connector smooth and easier to accomplish. The firm states in the letter that the distributor will offer |  | 62400 | 62,400 | US Distribution in the state of CA. | Terminated | February 07, 2017 | 699 | 62400 | Not_Computer
Z-2126-2016 | 73198 | No | MOQ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Battery Power Line II Battery Oscillator, DePuy Synthes Power Tools, The Anspach Effort Inc. |  |  | November 25, 2015 | 2015 | The Anspach Effort, Inc. | Has the incorrect operating directions on the mode switch ring, i.e. its is labeled "FWD, LOCK, REV" instead of "ON, LOCK, ON". | Labeling design | DePuy Synthes sent an Urgent Notice Recall Notification dated November 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to review their inventory and immediately remove any affected devices, complete the verification section of the letter and return it to the Anspach Effort Inc., and return any affected product to their DePuy Synthes Sales Consultant. Any questions regarding this recall / removal contact the Complain Handling Unit Manager at (561) 494-3673 or their DePuy Synthes Sales Consultant.For further questions regarding this, please call (561) 627-1080. |  | 20 | 20 hand pieces | US Distribution to the state of : VA and Internationally to:Thailand and UK | Open, Classified |  |  | 20 | Not_Computer
Z-2125-2016 | 74344 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic, MyCareLink" Patient Monitor, Model 24950, Rx Only. |  |  | May 26, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | Recently, a new software version was automatically sent to a subset of Model 24950 MyCareLink Monitors. After release, Medtronic identified an issue with the software that prevents implanted device data from being available to clinicians on the CareLink" Network. While the transmission appears successful to the patient, the transmitted data, including CareAlerts, are not visible to the clinic. | Software design | Medtronic sent an Urgent Field Safety Notice dated May 2016 to all affected customers. The letter described the issue and the product involved in the recall. The letter also included the "Customer Actions" for consignees to follow. Requested consignees to complete and return the Customer Confirmation Certificate to |  | 262 | 262 | Worldwide Distribution to Austria, Belgium, China , Czech Republic , Denmark , Finland , France, Germany, Hong Kong, India , Ireland , Israel, Italy , Malaysia, Netherlands, Norway, Poland, Portugal , Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, UAE, and United Kingdom. | Open, Classified |  |  | 262 | Computer
Z-2123-2016 | 74300 | No | EOX | Esophagoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Covidien Cytosponge Cell Collection Device. |  |  | June 06, 2016 | 2016 | Covidien, LLC | Medtronic is recalling all lots of the Covidien Cytosponge Cell Collection device after two reports of the device detaching from the removal string during withdrawal from the patient's esophagus. | Process control | The firm, Medtronic, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 6, 2016 to affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and discontinue use; return affected product to Medtronic/Covidien, Quality Department, 540 Oakmead Parkway, Sunnyvale CA 94085; complete and return the RECALLED PRODUCT RETURN FORM by email to Customer Service feedback.customerservice@Covidien or fax to (800)-895-6140 if product purchased directly from Medtronic; or by email to Quality Compliance MITGFCA@Covidien.com or fax to (203) 492-7719 if Product purchased through distributor or you have "no inventory"; and to promptly alert any other person or facilities if they have further distributed the product. All affected lots must be returned.If you have any questions or concerns, please do not hesitate to contact yourMedtronic representative or Customer Service at (800) 882-5878. |  | 620 | 620 total units (260 US; 360 OUS) | Worldwide Distribution-US Distribution to TN, NC, IL, CA, MN, OR and NY; and countries of: United Kingdom and Iran. | Open, Classified |  |  | 620 | Not_Computer
Z-2128-2016 | 74317 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Novation Press-Fit Splined RDD, Extended Offset, HA Coated, 12/14 Taper Femoral Stems, Size 18. |  |  | May 26, 2016 | 2016 | Exactech, Inc. | Labeled with the incorrect Global Trade Item Number (GTIN). The GTIN on the label identifies the devices as another. | Under Investigation by firm | Exatech sent an Important Notification letter dated May 16, 2016, to all affected customers. The letter notified customers of the product's incorrect Global Trade Item Number. Customers were instructed to cease distribution/use of the product, extend the information to accounts possessing the recalled product, identify and quarantine any devices in inventory, and to complete and return the Recall Inventory Response Form acknowledging that they received the recall notification. Customers were instructed to call 800-392-2832 with any questions. For questions regarding this recall call 352-377-1140. |  | 14 | 14 | Nationwide Distribuiton to OK, ME, NY, &amp; Hawaii. | Open, Classified |  |  | 22 | Not_Computer
Z-2133-2016 | 74329 | Yes | CBP | Nonrebreathing valve. | Anesthesiology | Anesthesiology | 510(k) | 5195GM, Cadillac First Aid Kit, #23459698, brand General Motors. No UPC. |  |  | May 16, 2016 | 2016 | Lifeline First Aid LLC | Cadillac First Aid Kit is recalled because it contains the CPR Shield with One Way Valve and Barrier Filter (mask) and this mask is cleared only for prescription use and is not cleared for over-the-counter (OTC) use. | Other | The firm sent the PRODUCT RECALL NOTICE- LIFELINE FIRST AID KITS WHICH INCLUDE CPR MASK letter, dated 5/11/2016, to their consignees.CPR shields will be replaced without charge and acceptable Lifeline First Aid kits will be re-worked and returned using our standard RGA process. In addition, the firm sent the PRODUCT RECALL NOTICE- CADILLAC FIRST AID KIT letter, dated May 31, 2016, to General Motors. General Motors will be sending out a dealership bulletin to inform their customers of this issue. Once confirmed, replacement masks will be sent to the requested dealerships and made available to be ordered (at no cost) as needed.Please complete and fax the attached Customer Response Form to Lifeline at (503)783-3721.Consignees with questions should call Lifeline First Aid, LLC at (503)783-3660. |  | 2975 | 2975 kits | Distributed in AZ, CA, CO, FL, GA, IL, KY, MD, MN, MO, MT, NE, NJ, NY, OK, OR, PA, SC, TN, UT, VA, WA, WA, and WI. Also distributed in New Zealand. | Open, Classified |  |  | 103905311 | Not_Computer
Z-2135-2016 | 74454 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MEXL-3010/A5- Vantage Titan 3T Magnetic Resonance Imaging System |  |  | April 14, 2016 | 2016 | Toshiba American Medical Systems Inc | A potential problem associated to the gradientcoil (ASGC) in the Titan 3T MRI System. | Device Design | An Urgent Medical Device Notification letter was sent to customers to alert them to a potential problem associated to the gradient coil (ASGC) in the Titan 3T MRI System. The letter informs the customers of the problem identify and the corrective action. Customers are informed that their Toshiba service representative will contact them for an appointment to install the modifiedsoftware on their system when it becomes available. Customers are instructed to complete the attached customer response form and fax it to the toll-free number shown at the top of the form; sent via email to TAMS-RegulatoryAffairs@toshiba.com. Customers with any questions are instructed to contact Paul Biggins at (800) 421-1968 or to contact their local Toshibarepresentative at (800) 521-1968. |  | 33 | 33 units | US Nationwide Distribution | Terminated | March 30, 2017 | 350 | 33 | Computer
Z-2131-2016 | 74442 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750w |  |  | June 08, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with the patient bore heating on the Discovery MR750w.A small area on the bore surface, which is normally accessible to the patient, can get warm enough to cause a serious burn in the event the patient touches the bore and proper padding is not used.No injury has been reported to GE as a result of this issue. | Nonconforming Material/Component | Consignees were sent on 6/8/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 60901 dated June 8, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 471 | 471 (184 US; 287 OUS) | US: AK, AZ, CA, COM, CT , DC, FL, GA, GU, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NV, NH, NJ, NY, NC, OH,OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WY.OUS:AlgeriaArgentinaAustraliaAustriaBahrainBelarusBelgiumBermudaBrazilBulgariaCanadaChinaEgyptFranceGermanyHungaryIndiaIndonesiaIrelandItalyJapanKazakhstanKorea (Republic of)KuwaitLebanonMexicoNetherlandsNew ZealandPolandQatarRomaniaRussiaRussian FederationSaudi ArabiaSingaporeSpainSwedenSwitzerlandTaiwanThailandTurkeyUnited Arab EmiratesUnited KingdomUruguay | Open, Classified |  |  | 471 | Not_Computer
Z-2134-2016 | 74515 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Image1 SPIES X-LINK Camera Control Unit (CCU), Product Model No. TC301, when used with Image1 SPIES D1 camera head, Product Model No. TH116 &amp; TH115, and DC1 Video Mediastinoscope |  |  | June 20, 2016 | 2016 | Karl Storz Endoscopy | Potential electrical interference issue when the Image1 SPIES X-LINK and Image1 SPIES D1 camera head are used in conjunction with the DCI Video Mediastinocope and an electro-surgery generator (ESG). | Device Design | KARL STORZ sent an Customer Notification letter dated June 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. KARL STORZ has become aware become aware of a potential electrical interference issue when the lmage1 SPIES X-LINK and lmage1 SPIES D1 camera head are used in conjunction with the DCI Video Mediastinoscope and an electro-surgery generator (ESG). Customers are instructed to complete the attached acknowledgement Form and to return it to KARL STORZ. Customers with questions are instructed to contact Technical Support Dept. at (800) 421-0837, press 2, and then press 1. |  | 619 | 619 units total (90 units in US) | Worldwide Distribution - US Distribution to the states of : IL, NM, TX, MD, TN, PA, FL, KY, MO, NY, MN, DC, MA, WI, CT, MI, ID, CO, GA, NJ and AR., and to the countries of : Argentina, Australia, Belgium, Bulgaria, China, Germany, Finland, Great Britain, India, Italy and Japan. | Terminated | January 05, 2017 | 199 | 619 | Not_Computer
Z-2136-2016 | 74362 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW, ARROWG+ARD AND ARROWG+ARD BLUE PLUS PRESSURE INJECTABLE CENTRAL VENOUS CATHETER. |  |  | June 02, 2016 | 2016 | Arrow International Inc | Arrow is recalling these products due to the possibility that although the kit is labeled as containing a 20 cm catheter, a 16 cm catheter may be packaged in the kit. | Process control | Arrow mailed an Urgent Medical Device Recall Notification letter to customers requesting they immediately discontinue distribution and quarantine any products with the affected lot numbers. A Recall Acknowledgement Form was included with the notification letter for customers to complete and email to recalls @teleflex.com indicating the amount of product on hand. |  | 1170 | 1, 170 units | US Nationwide Distribution: in the states of MA, CA, IL, TN, IL | Open, Classified |  |  | 1170 | Not_Computer
Z-2083-2016 | 74158 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-arm 02 surgical x-ray imaging system |  |  | March 17, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | Review of on-system electronic product labeling and Instr uctions for Use in response to an internal management concern uncovered gaps in compliance with the requirements for information to be provided to users in 21CFR part 1 020.30(h). | Radiation Control for Health and Safety Act | Medtronic planned action to bring device into Compliance:1. You will contact customers and provide an Errata sheet to them updating the missing and incorrect information.2. Medtronic field service personnel will perform radiation output measurements via validated procedures to attain compliance with 21CRF part 1020.32(k)(6).3. You will update the systems software-based instructions for use by a software update which includes the information supplied by the Errata sheet.For questions call (800)-595-9709. |  | 63 | US - 63 | Nationwide Distribution | Open, Classified |  |  | 63 | Computer
Z-2130-2016 | 74311 | No | NVE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The System 83 Plus, Endoscope Washer/Disinfector |  |  | May 06, 2016 | 2016 | Custom Ultrasonics, Inc. | To warn customers that Custom Ultrasonics, Inc. System 83 Plus AERs should not be used for cleaning and/or high-level disinfection of duodenoscopes until further notice. | Process control | Custom Ultrasonics, Inc. sent an Urgent Medical Device Recall letter dated May 6, 2016, to all affected customers. The letter advised customers not to use the System 83 Plus AERs for cleaning and/or high-level disinfection of duodenoscopes until further notice. In addition, they sent a warning label for customers to affix in a prominent location on the System 83 Plus AER's on or before June 3, 2016 stating the following: "This device is not indicated for reprocessing of duodenoscopes. Do not reprocess any duodenoscopes in this device until further notice. For alternative reprocessing options, please contact the duodenoscope manufacturer."For questions regarding this recall call 215-364-1477. |  | 2345 | 2,345 units | Nationwide Distribution to IN, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD., ME, MI, MN, MO,MS, MT, NC, NE, NH, NJ, NM, NV, NY OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT,WA, WI, WV, and WY. | Terminated | April 18, 2017 | 347 | 2345 | Not_Computer
Z-2129-2016 | 74464 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | PoleStar system (N20 / N30) Warning Lamps Control Box, Warning Lamps Control Box Cables, Thermoflex Chiller |  |  | June 20, 2016 | 2016 | Medtronic Navigation | Medtronic Navigation is recalling certain components of the PoleStar system (N20 / N30) because of misplaced ground pins which may cause electrical shock to the user. | Process change control | Medtronic sent a Field Safety Notice letter dated August 1, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that Medtronic will perform an inspection on site to verify and correct the grounding connection of the PoleStar System. For questions contact Medtronic Navigation Technical Services at 1-800-595-9709. |  | 52 | 52 | Worldwide Distribution - US Nationwide. | Terminated | June 21, 2017 | 366 | 52 | Not_Computer
Z-2082-2016 | 74153 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm Imaging mobile x-ray system. |  |  | March 17, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | It has been discovered that replacement x-ray generator and system motion batteries could become damaged during shipment, resulting in unacceptable risk of failure. | Radiation Control for Health and Safety Act | Medtronic's planned action to bring device into Compliance. 1. You will provide an Errata sheet to customers updating the missing and incorrect information.2. Field Service Engineers (FSEs) will perform radiation output measurements via validated procedures to attain compliance with21CFR part 1020.32(k)(6).3. You will inspect and test X-ray generator and system motion batteries for damage, in accordance with the plan submitted to CDRH. Damaged batteries will be replaced.4. You will update the systems software-based instructions for use by a software update which includes the information supplied by the Errata sheet.For further questions, please call (800) 595-9709. |  | 581 | US - 581 | Nationwide Distribution | Open, Classified |  |  | 581 | Computer
Z-2138-2016 | 74418 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Ti Button Plate 7 Holes, Fastener, Fixation, nondegradable, Soft Tissue |  |  | May 31, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating a recall due to affected devices being distributed and or labeled with MR Safety information that does not meet the current ASTM standard. | Under Investigation by firm | Synthes mailed to customers (Hospital Personnel and Sales Consultants) an URGENT FIELD SAFETY NOTIFICATION letter informing them the insert they received, which contained incorrect MR Safety Information, does not meet the current ASTM standards. Customers were asked to review the package insert provided with the notification letter, discard the outdated package insert, complete the attached verification Section (page 3 of this notification) and send a copy of the completed Verification Section by Fax: 844-782-5568 or scan/email: Synthes7923@stericycle.com. For further questions call (610) 719-6500. |  | 6255 | 6,255 units | USA Distribution to the states of : AK,AL,AR,AZ,CA,CH,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,ON,OR,PA,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI, &amp; WV | Terminated | April 25, 2017 | 329 | 6255 | Not_Computer
Z-2137-2016 | 74078 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | HeaterBaum Foot &amp; Ankle Heating Pad, Heating Soft Compression Therapeutic Boot with Precise Temperature Control Thermostat, 110 V with AC Adapter. |  |  | January 19, 2016 | 2016 | Ergoactives, Llc | Products are marketed without a cleared 510k. | No Marketing Application | Customers were notified by phone on 01/19/2016. Two follow-up notifications were sent to customers via e-mail. Customers were asked to immediately examine their inventory for the affected product; and, to notify and identify any customers they may have further distributed to. Ergoactives LLC requests that all Heaterbaum products be sent back to their headquarters: 20725 NE 16th Avenue, Suite A-26, Miami, FL 33179, along with the order number and note of where the product was purchased. The enclosed response form should be completed and returned as well. If customers have any questions, they can call Ergoactives LLC at (305)359-5031. Customers will receive a full refund for the return of their Heaterbaum products. |  | 278 | 278 units | Worldwide Distribution -- US, including the states of AL, AR, CA, FL, GA, KS, MA, MD, ME, MI, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, PA, TN, TX, UT, VA, and WA; and, the countries of Ecuador &amp; Venezuela. | Open, Classified |  |  | 278 | Not_Computer
Z-2143-2016 | 74354 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System, Product Code engen, Software Versions enGen Select 3.2 and Above and enGen Custom 2.0 and Above; IVD |  |  | May 16, 2016 | 2016 | Ortho-Clinical Diagnostics | If the user selects the Comment or Interpretation option (in the enGEN .gsb file), and the numerical results fall between Negative and Reactive, the IM incorrectly displays the interpretive text as Borderline and sends the incorrect text to the LIS. The IFU states the interpretation term for these numerical result values is Retest?.Âż The difference in terminology exists for US Markets only. | Software design | Ortho Clinical Diagnostics sent an "IMPORTANT PRODUCT CORRECTION NOTIFICATION" Customer Letter dated May 16, 2016, to all affected customers in the US with enGen Systems via FedEx or US Priority Mail (for Post Office Box addresses) or by e-mail to subscribers of ORTHO PLUS e-Communications to advise them of the issue and the required actions . Required Actions: 1. Use the numerical test results to assess actions to be taken for these assays, per the IFU.2. Post this notification by your enGen System or with your user documentation.3. Complete and return the Confirmation of Receipt form by 25 May 2016. with the Instructions for Use (IFU) directions).If you have any questions regarding this notification, please call the Ortho Technical Solutions Center at any time at 1-800-421-3311. |  | 30 | US: 30 units | Nationwide | Open, Classified |  |  | 30 | Not_Computer
Z-2145-2016 | 74346 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROS Immunodiagnostic Products AFP Reagent Pack, REF/Product Code 680 0784, Unique Device Identifier No. 10758750000876, IVD; --- Ortho-Clinical Diagnostics Percoed, Bridgend, UK --- Please note: the VITROS Immunodiagnostic Products AFP Calibrators, Unique Device Identifier No. 10758750000913, REF 680 0863, were included in the communication because they are lot-linked to the Reagent Packs and cannot be used with any other lot of reagent. There is no known failure with the Calibrators. |  |  | May 09, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho Clinical Diagnostics (OCD) confirmed that specific lots of VITROS Immunodiagnostic Products AFP Reagent Pack product exhibit an increase in the reaction signal over time potentially resulting in the inability to calibrate and use the product. | Nonconforming Material/Component | The firm, Ortho Clinical Diagnostics, sent an "IMPORTANT PRODUCT CORRECTION NOTIFICATION" Customer Letter (Ref. CL2016-098) dated 5/9/2016, via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received the affected product. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue distributing and discard the affected lots; complete and return the Confirmation of Receipt form by May 18, 2016 via Fax to 1.888.557.3759 or 1.585.453.4110 or scan to PDF and email to: ConfirmationAdmin@its.jnj.com, and forward this notification if you have distributed these systems outside of your facility.If you have any questions, contact Customer Technical Services at 1-800-421-3311. |  | 13577 | US: Lot 6471 - 282, Lot 6480 - 344 | US Nationwide Distribution | Open, Classified |  |  | 13577 | Not_Computer
Z-1887-2016 | 73636 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System: HeartWare Battery, Lithium Ion, 14.8 V, Rx only. Catalog 1650 (US) and 1650, 1650-DE and A00035 (OUS) |  |  | January 07, 2016 | 2016 | HeartWare Inc | Heartware Ventricular Assist System Battery Cell experiencing premature depletion. | Under Investigation by firm | The firm, HeartWare, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 7, 2016, to all their consignees informing them that of the HeartWareÂż HVAD Battery recall (device exchange). The letter described the product, problem and actions to be taken. The customers were instructed to:1. Promptly complete the attached Acknowledgment Form and return it to HeartWare no later than 30 days from the date of this letter via email to HeartWare7786@stericycle.com or provide verbal Acknowledgement by calling Customer Service line , 1(844)491-7871; or fax to +1 (844)721-3044..2. Immediately identify and quarantine affected batteries in hospital inventory, and apply a RECALLED PRODUCT  DO NOT USE sticker (provided with this notice) as illustrated below. Replace quarantined inventory AC Adapters and batteries by requesting a Return Goods Authorization (RGA) through a qualified representative or directly through HeartWare Customer Service (via telephone 1-877-367-4823 or email cs@heartware.com). A qualified representative will update the Controller software for affected Controller units in inventory.3. Identify, collect and quarantine affected product in patient possession. Arrange for current patients to bring their Batteries to a clinic appointment as soon as possible, (at least within the next three months), with urgency dependent upon individual patient circumstances. During their regularly scheduled appointment, a qualified representative will assist with quarantining and replacing affected product. Apply a RECALLED PRODUCT  DO NOT USE sticker as illustrated on affected batteries. Replace affected patient product using hospital inventory, and replace consumed hospital inventory by requesting an RGA as described above.4. Return affected product to HeartWare. Your HeartWare representative will assist you with returning affected product to HeartWare.5. Complete and return the attached Completion Form no later than 6 months from the date of this let |  | 18631 | 18, 631 | US Nationwide distribution in the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, and Washington D.C. | Open, Classified |  |  | 18631 | Computer
Z-2146-2016 | 74356 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | QuickSlide MGS-80, Automated Mutli-Slide Gram Stainer/GramPRO-80. |  |  | May 24, 2016 | 2016 | Hardy Diagnostics | Discrepant culture and gram stain results, which could potentially lead to misdiagnosis or improper treatment of a patient. | Under Investigation by firm | The firm began notifying consignees via phone on 05/23/16. The firm also sent out Urgent Notification letters, dated June 15, 2016.The firm's phone script instructs customers to decrease the number of rinses on the slide and increasing the number of rinses in the chamber for better outcomes and performance with the instrument. Reducing the rinses on the slide may prevent bacteria from coming off the slide, and adding rinses to the empty chamber will help ensure that the chamber is clean before the next slide is introduced.The firm's customer notification letter recommended an immediate update to the automated rinse cycle to decrease the number of slide rinses and increase the number of empty chamber rinses to improve instrument performance. A step-by-step procedure for updating the rinse cycle settings was discussed via phone and should now be implemented.In addition to the changes, the firm is updating the GramPro 80 User Manual to reflect new recommendations, and it is strongly recommended that all users perform methanol fixation of slides for automated staining and discontinue the use of heat fixation to prevent cross contamination of the unit. Once updates to the User Manual have been completed, the firm will provide an updated revision for customer records.For further information or technical assistance, contact QuickSlide Technical Team at gibbsd@HardyDiagnostics.com. |  | 25 | 25 | U.S. Distribution to the following states: CA, TX, NE, OK, LA, NC, PA, AZ, NY, PA, VA, NM, MI, and WV. | Open, Classified |  |  | 25 | Not_Computer
Z-2147-2016 | 74490 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | AORTIC ARCH CANNULA, Catheter, Cannula And Tubing, Vascular, Cardiopulmonary Bypass (part numbers NA-55X7 and NA-55X8) |  |  | June 22, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA is recalling Aortic Arch Cannulae (part numbers NA-55X7 and NA-55X8) because the distal end may be less than specified. The defective cannula are showing an angle that is about 15-20 degrees instead of approximately 45 degrees. | Process control | The firm, LivaNova, sent an "Urgent Safety Alert" field communication letter dated June 22, 2016 via certified mail to consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to: check your inventory to determine if you own Cannulae, if no inventory, complete the response form and return it per instructions indicated on form.; check the tip angle to determine if it is conforming to specifications, if specifications not met, file a complaint through your normal process whether you decide to keep it, use it or return it to Sorin Group; If you are not willing to use it, return affected product parts to Sorin Group USA,; contact Customer Service at 1-800-650-2623 or email CustomerService@livanova,com to have a RMA issued, and complete and return the Customer Response Form no later than July 20, 2016 via Fax at 303-467-6502 or by e-mail to USFSN@livanova.com. If you need support for the execution of these actions, please contact your LivaNova sales representative or call Customer Service at 1-800-650-2623 (Monday to Friday 8AM to 5PM MT).For questions regarding this notice, please contact Customer Service at 1-800-650-2623 (Monday to Friday 8AM to 5PM MT) or e-mail USFSN@livanova.com. |  | 5740 | 5740 | Worldwide Distribution: US (nationwide) to states of: LA, Ml, MO, OH, TN, and TX; and countries of: Italy and Canada. | Terminated | March 27, 2017 | 278 | 5740 | Not_Computer
Z-2144-2016 | 74522 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OfficePACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Potential data loss occurs as a result of the software archiving not working properly. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated January 26, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to provide a copy of the notification to all users. For additional questions customers were advised to call Merge Customer Service at (87) 741-5369. For questions regarding this recall call 262-367-0700. |  | 216 | 216 sites have the affected software | The software was distributed nationwide to medical facilities. There was no government/military/foreign distribution. | Open, Classified |  |  | 216 | Computer
Z-2142-2016 | 68318 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The LifeCare Pump Model 4200 PCA Infuser makes possible the practical implementation of patient-controlled analgesia. The health care professional sets the prescribed dose, rates, and limits for the drug to be infused. The infuser can be set to operate in one of three modes: PCA Mode, Continuous Mode, and PCA + Continuous Mode. The LifeCare Pump Model 4200 PCA Infuser is designed with a number of mechanical, electrical, and software safeguards that prevent patient injury or drug abuse. |  |  | March 08, 2013 | 2013 | Hospira Inc. | Hospira has received reports of PCA pumps not detecting distal occlusions, one of which resulted in a serious injury. The issue is caused by normal wear and tear on the Half Nut (the component/nut that travels up and down the lead screw) which prevents it from properly detecting the pressure build-up associated with a distal occlusion. | Component design/selection | Hospira sent an IMPORTANT DEVICE INFORMATION letter dated March 8, 2014 to all affected customers.. The letter advised customers to: 1) Perform the Performance Verification Test (PVT) Occlusion Test as defined in the PCA Technical Service Manual (TSM); 2) If the device does not pass the test, remove it from clinical service and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue; and 3) Perform the appropriate troubleshooting and repair activities defined by your facility, which may include returning the device to Hospira for further diagnosing and servicing. The letter advised the Hospira is in the process of establishing a useful life for the half nut and to determine when it will require replacement. A new Technical Service Manual will be issued in late 2014 that adds a requirement for annual PVT Occlusion testing. The System Operating Manual will include new instructions to reduce the excessive wear on the half nut resulting from incorrectly using the vial information. Customers who further distributed the devices are requested to notify their customers of the recall and ask them to call Stericycle at 888-965-5801 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. Questions about this recall should be direct to Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 6268 | 6,268 infusion pump systems | Nationwide Distribution to AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY) + District of Columbia + Puerto Rico + Virgin Islands; *** FOREIGN: | Completed |  |  | 106513 | Not_Computer
Z-2149-2016 | 74585 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Surface Monitoring System{OSMS, Varian Private Label} This is used on Varian's TrueBeam EDGE; |  |  | June 23, 2016 | 2016 | Varian Medical Systems, Inc. | Varian Medical Systems has received reports that unintended changes were made to planned couch parameters, specifically the couch rotation parameter, during patient set-up. This can occur when a user sets a different couch rotation from the plan couch rotation while performing patient alignment with Optical Surface Monitoring System [OSMS]. | Software design | The firm, Visionrt, sent an "URGENT MEDICAL DEVICE CORRECTION NOTIFICATION-URGENT FIELD SAFETY NOTIFICATION" letter dated June 24, 2016 to all affected customers. Letters describe the product, problem and actions to be taken. The customers were instructed as follows: DO NOT USE the "Move Couch" function if the couch rotation value set in the AlignRT device DOES NOT match the planned couch rotation for the treatment field; retain a copy of this document along with your most current product labeling; pass the letter to all those who need to be aware of it within their organization; always confirm that the patient is correctly aligned prior to commencing treatment; is not required to return their devices; promptly inform Vision RT if they believe that patient harm occurred, and complete Appendix 1- Customer Acknowledgment Response and return it via email to corrections@visionrt.com.Should you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866-778-2379 from the US) or as per http://www.visionrt.com/contact/details. |  | 189 | 166 OSMS in US, 23 OSMO - International, | Worldwide Distribution: US (nationwide) and countries of: Austria, Azerbaijan, Brunei Darussalam, Bulgaria, Denmark, Georgia, Germany, India, Italy, Malaysia, Morocco, Norway, Portugal, Switzerland, and United Kingdom. | Open, Classified |  |  | 211 | Not_Computer
Z-2158-2016 | 74283 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5 mm and 4.0 mm Cannulated Screw, Product code HWC, Device Listing Number D029094, Screw, Fixation Bone |  |  | February 16, 2016 | 2016 | Synthes (USA) Products LLC | Labeling does not match the cleared indications for use in the United States and Canada. | Labeling False and Misleading | DePuy Synthes initiated a Field Safety Notice to remove the affected Technique Guides and marketing literature from the field. Sales Consultants and Synthes Trauma Marketing will receive training on affected technique guides and marketing literature which promoted off label use. All technique guides and literature have been put on hold and removed from DePuy Synthes website. Effectiveness of this field action will be tracked using their training system. They will confirm training when completed as a means to measure effectiveness. |  | 19497844 | 19,497,844 units in total | US Nationwide Distribution in the states of :AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, N/A, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | June 01, 2017 | 471 | 19497844 | Not_Computer
Z-2148-2016 | 74030 | [6 empty] | Breathe Right Clear, 30 ct. carton., UPC 7-57145-00246-7 (US), UPC 9-300673-871749 (AU/NZ) |  |  | April 25, 2016 | 2016 | GSK Consumer Healthcare | Some cartons used in the secondary packaging of Breathe Right Clear, Large, 30 ct. Nasal Strips were incorrectly labeled for cartons intended for the AU/NZ market. | Error in labeling | GSK Consumer Healthcare sent an Urgent - Product Recall Letter dated April 4, 2016, to all affected customers and an updated recall letter dated April 25, 2016, via FedEx.The recall letters identified the product and problem. The April 25, 2016, was sent to further clarify, the lot number on the Breathe Right carton. The customers were asked to immediately examine their inventory, quarantine and cease distribution of the specific lot of the affected product. Customers were asked to carry out a physical count and record the data on the response form. The response form should be returned to GSK within five (5) business days, even if you do not have the recalled lot. The response form should be faxed to (336) 499-8795. If customer needs shipping assistance or questions about the recall process, they can call 1-844-722-2104, and refer to Event Number 535STRIPS16. If any customer has questions, they can direct them to GSK Consumer Healthcare Consumer Relations team at 800-858-6673. |  | 1440 | 1,440 cases ( 34, 560 retail - units ) | US Nationwide | Terminated | July 19, 2017 | 450 | 1440 | Not_Computer
Z-2159-2016 | 74510 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OfficePACS software. The firm name on the label is Merge Healthcare, Heartland, WI. |  |  | January 31, 2016 | 2016 | Merge Healthcare, Inc. | Potential data loss occurs as a result of product archiving not working properly. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/31/2016 explaining the reason for recall and instructing the customer to provide a copy of the notification to all users. |  | 49 | 49 sites have the affected software | Distribution was made to medical facilities located in AL, AZ, CA, CO, CT, FL GA, ID, IL, KY, MA, MO, NE, NH, NJ, NY, NC, OK, PA, SC, TN, VA, and WA. There was no government/military/foreign distribution. | Open, Classified |  |  | 49 | Not_Computer
Z-2169-2016 | 74425 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Target Detachable Coils: |  |  | June 06, 2016 | 2016 | Stryker Neurovascular | Stryker Neurovascular has become aware that some units of Target Nano product do not have the intendedstretch resistance. The sutures inside the coil that provide stretch resistance may have been damaged during manufacturing which can lead to an increased occurrence of coil stretching. | Process change control | Stryker Neurovascular sent an Urgent Medical Device Voluntary Recall letter dated June 3, 2016, to all affected customers via FED Ex. Letter includes recall letter, list of impacted lot numbers, customer acknowledgement form and frequently asked questions. Customers were instructed to immediately check their inventory. segregate the effected inventory in a secure location for return to Stryker, circulate the Field Safety Notice internally to all interested/affected parties. Inform Stryker if any of the subject devices have been distributed to other organizations. Customers were also instructed to complete the attached customer response form and return the completed form to their nominated Stryker Representative. |  | 45302 | 45,302 (8,188 in the U.S.) - total, all model numbers | Worldwide Distribution - US (nationwide) and Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Canary Islands, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Italia, Japan,Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Mexico,Netherlands, Norway, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, United States, Uruguay, and Vietnam. | Open, Classified |  |  | 45302 | Not_Computer
Z-2151-2016 | 74556 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS  versions 7.4, 7.5, 7.6, and 8.0, Diagnostic Viewer. Distributed by Novarad Corporation. Picture archiving and communication system (PACS) |  |  | May 09, 2013 | 2013 | Novarad Corporation | Potential for images to be flipped while streaming, which could incorrectly display image orientation markers. | Software design | Consignees were notified on 05/09/13 via letter and email to upgrade software and to turn off streaming function. |  | 2695 | 2695 | Worldwide distribution including US nationwide. No Canadian distribution. | Open, Classified |  |  | 2695 | Computer
Z-2202-2016 | 74489 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Femoral Arterial Line |  |  | June 02, 2016 | 2016 | Arrow International Inc | Labeling error. Some shipping cartons labeled as ASK-04018-VUH contained ASK-04020-VUH kits; and some shipping cartons labeled as ASK-04020-VUH contained ASK-04018-VUH kits. | Labeling mix-ups | Arrow International mailed an Urgent Medical Device Recall Notification letter and acknowledgement form dated June 20, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1. Immediately discontinue distribution and quarantine any products with the product codes and lot numbers listed above.2. Inspect affected product within your control to identify if the shipping corrugate label correctly identifies the products inside the shipping carton. Products with an incorrect label on the shipping carton can be returned to Arrow. 3. Once you have completed inspection of all of the products from your own inventory, check the box on the enclosed Acknowledgment Form that indicates that you have completed the inspection and fax it to 1-855-419-8507, Attn: Customer Service, or email it to recalls@teleflex.com.For questions contact your local sales representative or Customer Service at1-866-246-6990. |  | 50 | 50 units in total | US distribution in the states of TN and KY. | Terminated | August 21, 2017 | 445 | 50 | Not_Computer
Z-2191-2016 | 74416 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Part # SD313.011, Matrix Neuro Screwdriver Blade MQC/MR Conditional/Short, Lot # UR75220, mfg 7/18/2007 |  |  | May 31, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating this recall due to affected devices that were distributed, etched and or labeled with MR Safety information not meeting the current ASTM standard. | Labeling False and Misleading | DePuy Synthes mailed an URGENT MEDICAL DEVICE FIELD SAFETY NOTIFICATION letter to their customers (Hospital Personnel, including Imaging Department Personnel and DePuy Synthes Colleague) to inform them of the recall and actions to take. The letter identified the affected, problem and actions to be taken. The letter instructed customers not to use the affected Instruments which are labeled or etched MR Safe and MR Conditional during MRI scans. Customers were asked to review the information contained in the Field Safety Notification letter and complete the Verification Section located on page 3 of this notification. For questions call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 18 | 18 units | US Nationwide Distribution in the states of: AR,CA,CO,CT,DC,ID,IL,KY,LA,MA,MD,MN,MO,MS,MT,NY,PA and TX. | Open, Classified |  |  | 749 | Not_Computer
Z-2203-2016 | 74154 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobasÂż Liat Influenza A/B Quality Control Kit; For use with cobasÂż Liat Influenza A/B Assay. |  |  | May 17, 2016 | 2016 | Roche Molecular Systems, Inc. | Kit labeled with the incorrect expiration date. | Incorrect or no expiration date | An Urgent Medical Device Recall (UMDR) notice (Customer Letter) was sent to the 35 individual affected customer sites by UPS on May 17 2016.Customers were asked to take the following actions: " Discontinue use of cobasÂż Liat Influenza A/B Quality Control kit lot numbers 51222E, 60121H, 60201D, and 60229F, and discard them per your local regulations. " Complete the enclosed UMDC faxback form (6941-00-0516) and fax it to 1-844-449-8506. " File this UMDC for future reference. |  | 194 | 194 distributed kits | Distributed in the states of: TX, NY, OR, HI, MD, PA, NC, MA, IA, WA, OR, AK, ME, MN, VA, MT, PA, and LA. | Open, Classified |  |  | 194 | Not_Computer
Z-2207-2016 | 74389 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HX2" Temperature Management Systems; |  |  | June 08, 2016 | 2016 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) is conducting a voluntary recall of its heater-cooler devices due to the distributor no longer selling the Interchlor sanitizer (Product Code 205) referenced in its heater-cooler cleaning instructions. | Nonconforming Material/Component | Terumo Cardiovascular Systems (Terumo CVS) initiated a voluntary recall on 06/08/2016, by mailing certifed letters to customers. Terumo CVS is notifying all users of its heater-cooler devices that the distributor, Intercon Chemical Company, no longer sells the Interchlor sanitizer (Product Code 205) referenced in its heater-cooler cleaning instructions. However, the identical sanitizer is available from other sources.This notice applies to the following Terumo CVS heater-cooler devices:Âż Sarns" TCM and TCM II Cooling and Heating SystemsÂż Sarns" Dual Cooler-HeatersÂż HX2" Temperature Management SystemsTerumo CVS will provide users with updated user instructions regarding the 10% Sodium Hypochlorite sanitization solution. Instructions will be provided under separate cover, when available; updated manual and cleaning guide will be completed in December 2016.Customers are asked to do the following: Review this notice and ensure that all users receive notice of this issue.2. Return the attached Customer Response Form acknowledging receipt of this notice.3. Source a replacement for the Interchlor sanitizer. See information above or call Terumo CVS Technical Support for assistance.If you have questions about this notification, please contact your Terumo representative or Terumo CVS Technical Support:800.441.3220, extension 6932Technical Support Hours:Monday  Friday8 a.m.  6 p.m. ET(24/7 Emergency Hours) |  | 6412 | 6412 units | Domestic: AK, AL, AR, AZ, CA , CO, CT, DC , DE , FL, GA, HI, IA, ID, IL, IN, KS, KY,LA , MA, MD, ME, MI , MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV , NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY Foreign: Argentina, Australia, Belgium,Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, England, Finland, France, GermanyGreece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Lebanon, MalaysiaMalta, Mexico, Netherlands, New Zealand, Norway, Philippines, Portugal, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates , Uruguay , Venezuela, Vietnam | Terminated | August 04, 2017 | 422 | 6412 | Not_Computer
Z-2204-2016 | 74562 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | 5.0mm Round Fluted Bur, Super Long |  |  | June 16, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments initiated a voluntary recall of specific lots of Round Fluted Burs, due to tarnishing or corrosion which may be present on the recalled burs and could result in a foreign body reaction (inflammation) necessitating surgical intervention. | Under Investigation by firm | Stryker initiated a voluntary recall of the Stryker Round Fluted Burs (used with Stryker CORE" System) on June 16, 2016, via certified mail due to the potential for the devices to have tarnishing or corrosion present. Non-responding accounts will be contacted by phone, email, and/or fax as necessary in an effort to obtain signed Business Reply Forms. Affected product will be removed from the field. Further corrective and preventative actions will be determined through the associated CAPA. The existing label will be defaced, packaging will be opened and products will be destroyed. Actions to be taken by the Customer/User:1. Immediately review this Recall Notification. 2. Check all stock areas and/or operating room storage to determine how many affected Round Fluted Burs from the affected lots are at your facility. Quarantine and discontinue use of the recalled Round Fluted Burs.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you dont have any affected product on hand. Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4 If you have further distributed this product, please forward this letter and the attached BRF to all affected locations. Please indicate each location on the BRF. 5. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762. 6. If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. Upon receipt of the recalled product, a credit will be issued to your account.Report any serious adverse events or product quality problems to Stryker Instruments: 1-800-253-3210 Health care professionals and consumers |  | 25 | 25 | US Distribution to the states of : CA, IL, SDForeign: Canada, Netherlands | Terminated | June 06, 2017 | 355 | 25 | Not_Computer
Z-2211-2016 | 74297 | No | OIJ | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(K) Exempt | Mammotome Revolve Dual Vacuum Assist Biopsy System Model number: MST1009 &amp; MST0815. Each disposable probe is packaged within a thermoformed tray and covered with a pre-printed (product code and graphics) Tyvek lid. Each tray is labeled on-demand after sealing with a lot number and expiration date. Trays are kitted into sets of five probes and packaged within a sales carton. Two sales cartons are then packed into a corrugated case and shipped to the global distribution center. Device is labeled as sterile. |  |  | May 19, 2016 | 2016 | Devicor Medical Products Inc | A 10 Gauge needle, 9cm long was shipped in an 8 Gauge needle, 15 cm package. | Packaging | An Urgent: Medical Device Recall of Mammotome Revolve Dual Vacuum Assist Biopsy Probe due to Potential Product Code Mislabel communication was sent to the customers via 1st class USPS on June 6, 2016. The labeling and the Customer response form was also attached. Customers were asked to examine their inventory and quarantine the affected product. If the product was further distributed, those customers are to be identified and notified. The letter stated that the assigned sales rep would be in contact to support the identification, inspection and replacement of the recalled products. Customers are to complete and return the enclosed response form. Questions are to be directed to Mammatome Customer Service 1-877-926-2666. |  | 360 | F11537107D1 = 360; F11536423D1 = 80; F11537109D1 = 360 | Worldwide Distribution -- USA, including the states of AL, AR, CA, CT, DC, FL, GA, IL, KS, KY, LA, MA, MD, ME, MI, MN, MS, NC, NJ, NY, OH, PA, SC, TN, WA and WI; and, the countries of Canada and Japan, Korea, and The Netherlands. | Terminated | August 03, 2017 | 441 | 360 | Not_Computer
Z-2213-2016 | 73696 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | AMD-RITMED brand Conforming Stretch Gauze Bandages, Sterile, 6 in x 4/1 yds./15 cm x 3.7 cm, Reorder Number/Code D606, Bar code 6 86864 04230 3, Made in China -- Manufactured for &amp; distributed by: AMD-RITMED INC. Montreal Quebec H8T 3J8 -- US Distribution: AMD-RITMED INC. Tonawanda, NY 14150 -- Type of packaging: 1 roll/pouch, 12 pouches/inner box, 4 inner box/case, 48 rolls per case. |  |  | March 30, 2016 | 2016 | Amd-Ritmed, Inc. | The product lot did not meet sterility requirements based on FDA sampling and analysis. | Material/Component Contamination | The firm, AMD-Ritmed, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 3/30/2016 and Medical Device Recall Response Form to the US distributors/customers via e-mail. The letter described the product, problem and actions to be taken. The customers were instructed to examine your stock to determine if you have any; immediately discontinue distributing product; promptly return product via parcel post to Tonawanda/Buffalo warehouse, ATTENTION: RETURNED GOODS, AMD-Ritmed Inc., 295 Firetower Drive, Tonawanda, NY 14150; if you have further distributed the product, immediately contact your accounts, advise them of the recall situation and have them return their outstanding stock, and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE-Acknowledgment &amp; Receipt Form via email to : AMDCUSTOMERSERVICE@AMDRITMED.COM or fax to: 1-866-871-3455, ATTN: AMD-Ritmed Customer Service.If you have any questions, please contact us at (1-800-445-0340) or at AMD Customer Service AMDCustomerService@amdritmed.com. |  | 107 | US: 107 cases (48 rolls per case) | US Distribution in states of: NY, CA, and TN. | Terminated | July 15, 2016 | 107 | 107 | Not_Computer
Z-2206-2016 | 74618 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 3.0, RayStation 3.5, RayStation 4.0, RayStation 4.5, RayStation 4.7 and RayStation 5 --- Radiation Therapy Treatment Planning System --- designed for treatment planning and analysis of radiation therapy, provides treatment unit set up parameters and estimates dose distributions. |  |  | July 01, 2016 | 2016 | RAYSEARCH LABORATORIES AB | An issue was found with the evaluation of biological clinical goals in RayStation 3.0, RayStation 3.5, RayStation 4.0, RayStation 4.5, RayStation 4.7 and RayStation 5. Biological clinical goals for an adapted plan based on another planning CT than the original planning CT will show incorrect values. -- To the best of the firm's knowledge, the issue has not caused any patient mistreatment or other incidents. However, the user must be aware of the following information to avoid incorrect plan evaluation during treatment planning. | Under Investigation by firm | RaySearch Laboratories AB sent Field Safety Notice/Medical Device Correction #11905 Letter (RSL-D-61-298, dated 7/1/2016) to end users via e-mail on 7/1/2016. The letter informs end users of the issue and provides actions to be taken by the users (work around). |  | 170 | US: 170 units | California, Connecticut, Delaware, Florida, Hawaii, Maine, Missouri, New York, Ohio, Texas and Washington | Open, Classified |  |  | 170 | Not_Computer
Z-2216-2016 | 74094 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobasÂż LiatÂż Analyzer |  |  | May 09, 2016 | 2016 | Roche Molecular Systems, Inc. | Device was released with default pcal/gain settings forphotometers instead of the pcal/gain settings that are required for release. The default settings can potentially lead to the generation of erroneous or invalid results. | Under Investigation by firm | Customers were notified of the recall via phone and email sent on 5/9/2016. They were instructed to take the following actions: Immediately discontinue use of the affected instruments until they have been reworked by a Roche Field Engineering Specialist. A Roche representative will contact you to schedule this service. Complete the attached fax form and fax it to 1-317-521-4815. File this Urgent Medical Device Correction (UMDC) for future reference. Please contact the Roche Support Network Customer Support Center 1-800-800-5973 if you have questions. |  | 2 | 2 | Distributed in NC and WA. | Completed |  |  | 2 | Not_Computer
Z-2214-2016 | 74415 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Depuy Synthes Trauma- Trochanteric Fixation Nail Advanced (TFNA) Blade/Screw Guide Sleeve, Surgical Instrument |  |  | June 07, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is initiating a voluntary medical device recall due to Blade/Screw Guide Sleeves (Part # 03.037.017) may have been manufactured with a slightly larger than specified outer diameter which could prevent the sleeve from going through the TFNA aiming arm properly, which may cause surgical delay. | Nonconforming Material/Component | DePuy Synthes sent an Urgent Notice Medical Device Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to conduct functional testing to inspect the instrument to determine if the Guide Sleeve is able to fit through the aiming arm properly. Customers were asked to only return the device if the product does not pass.Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. |  | 2064 | 2,064 units | Nationwide Distribution toAK,AL,AR,AZ,CA,CO,CT,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,PR,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI,WV, and WY | Terminated | April 25, 2017 | 322 | 2064 | Not_Computer
Z-2222-2016 | 74437 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | T030-000 -30" 20 Drop IV Admin Secondary Set. Packaged in a Tyvek pouch, with 50 pouches per case. |  |  | June 14, 2016 | 2016 | WalkMed Infusion, LLC | WalkMed Infusion, LLC Announces a Voluntary Field Action of Multiple Triton Administration Sets Due to the Potential Presence of Particulate In or On the Administration Sets. | Nonconforming Material/Component | Terumo sent a notification letter dated June 8, 2016 to customers. The letter identified the affected product, problem and action to be taken. Customers were instructed to 1. Review this notice and ensure that all users receive notice of this issue. 2. Return the attached Customer Response Form acknowledging receipt of this notice. 3. Source a replacement for the Interchlor sanitizer. See information above or call Terumo CVS Technical Support for assistance. 4. Terumo CVS will provide users with updated user instructions regarding the 10% Sodium Hypochlorite sanitization solution. Instructions will be provided under separate cover, when available.For questions contact your Terumo representative or contact Terumo CVS Technical Support at 800.441.3220, extension 6932. |  | 2453 | 2453 cases | US Nationwide Distribution in the states of AL, AZ, CA, FL, IL, MA, MS, NJ, NV, NY, OH, PA, TN, UT, and WA. | Open, Classified |  |  | 24249 | Not_Computer
Z-2217-2016 | 74652 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Carestream Touch Prime, Catalog Number 1738830, and Carestream Touch Prime XE, Catalog Number 1738822 --- Common Name: Touch Ultrasound |  |  | July 08, 2016 | 2016 | Carestream Health Inc | A software issue related to the generic volume measurement functionality could result in the volume of the subject anatomy being overestimated. For example, measurement results of the anatomy may appear enlarged or distended when it actually is not. | Software design | The firm, Carestream Health Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" Letter dated 7/7/2016 and response form the end users/customers on 7/8/2016 via Certified Mail, Return Receipt Requested. The letter described the product, problem and actions to be taken. The Customers were instructed not use the generic volume measurement tool until Carestream has updated the system software. In the interim, calculate volumes by using the specific volume tool from the desired calculation package. If you have distributed the device outside your facility, please alert your customer(s) of this field correction and contact Carestream Customer Care Center as listed above.Carestream has created a software update for Touch Ultrasound to resolve this issue and will update your software. Carestreams representative will contact you and schedule a convenient time to upgrade the software.For questions or concerns, please contact the Carestream Customer Care Center in the US at 1-800-328-2910 (available on a 24/7 basis). Outside of the US, please call your local Service support number. |  | 7 | 7 Units (US: 5 units, OUS: 2 units) | Worldwide Distribution: US Distribution to states of: GA, IA, and TX; and country of: Italy. | Terminated | May 02, 2017 | 298 | 7 | Computer
Z-2219-2016 | 74436 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The pump (with cushioning foam inserts) is packaged in a single pump box. Four pump boxes are placed in an over-shipper for distribution. |  |  | June 14, 2016 | 2016 | WalkMed Infusion, LLC | WalkMed Infusion, LLC Announces a Voluntary Field Action of the Triton Infusion Pump (model 300000) and Triton FP Infusion Pump (model 400000) Due to Unapproved Changes to the Software and Specifications of the Triton Infusion Pump (model 300000) and Unapproved Changes to the Intended Use of the Triton FP Infusion Pump (model 400000). | Device Design | WalkMed sent an Urgent : Medical Device Recall letter dated June 14, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Users: WalkMed recommends the following actions be taken with respect to the recalled product:Our records indicate that you have received Triton pumps. Immediately review your inventory and determine whether you have any of the Triton or Triton fp infusion pumps (model numbers 300000 and 400000).Please complete and promptly return the enclosed recall response form, even if you do not have any affected stock subject to this recall notice.WalkMed recommends that users stop use of and replace the Triton and Triton fp Infusion pumps with alternative devices as soon as possible.To ensure that patients who use WalkMed Triton infusion pumps can continue to receive infusion therapies without interruption, WalkMed will support customers in their transition to alternative devices. To ensure continuity of patient care, WalkMed will continue to provide IV administration sets to existing customers, provide preventive maintenance and service for distributed pumps and make available any parts necessary for the repair or service of units for up to 3 months from the date of this recall notice.Contact WalkMed at (303) 420-9569 to arrange for shipment of the product being returned.If an owner of a Triton or Triton fp infusion pump is interested in information about options for a replacement product, please contact WalkMed. For further questions please call (303) 420-9569. |  | 2482 | 2482 units | Nationwide to AL, AZ, CA, FL, IL, KS, MA, MD, MS,NJ, OH, PA, TN, TX, UT, and WA. No foreign/VA/govt/military. | Terminated | April 28, 2017 | 318 | 5448 | Computer
Z-2188-2016 | 74473 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Angio WorkStation: XIDF-AWS801v6.00, v6.01, and V6.10 System: INFX-8000V Bi-Plane System |  |  | February 24, 2016 | 2016 | Toshiba American Medical Systems Inc | During a procedure the Peak Skin Dose (PSD) value was displayed on the DTS larger than the DTS expected value. It was found that the software incorrectly calculated the PSD value when the x-ray condition of the frontal plane and the lateral plane are the same. It incorrectly used in the calculation a dose value from a previous exposure. | Radiation Control for Health and Safety Act | Toshiba Medical Systems (TAMS) planned action to bring product into Compliance:1.TAMS will install the software and perform testing to ensure the software update was effective.2. The customer notification letter which includes a statement that TAMS will, without charge, remedy the defect or bring the product into compliance.. If you have any questions regarding this, please feel free to contact your local Toshiba representative at (800) 521-1968. |  | 23 | 23 | US Distribution to the states of : NC, NY, TX, GA, CA, AZ, IL, FL,MA, MO and DE. | Terminated | April 25, 2017 | 426 | 23 | Computer
Z-2218-2016 | 74449 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith and Nephew TWINFIX QUICK-T Knot Pusher Suture Cutter; Part Number: 7209417, for arthroscopic use. |  |  | June 15, 2016 | 2016 | Smith &amp; Nephew, Inc. | The device could be subject to breaches of its sterile packaging. | Package design/selection | Urgent Recall Notification Letters were sent via Federal Express on June 15, 2016, to the customers listed in the firm's distribution report. Customers were asked to inventory their devices and contact the firm for return authorizations. |  | 563 | 563 | Worldwide distribution. US nationwide, AT, AU, DE, FR, GB, JP, PT, and RU. | Open, Classified |  |  | 563 | Not_Computer
Z-2238-2016 | 74639 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | 2.3MM Tapered Router, |  |  | June 24, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | A variation in flute depth on the routers was observed. The variation in flute depth may cause the routers to bemore susceptible to breakages.Risk to Health:Injury to critical soft tissue requiring surgical intervention may occur. | Process design | Stryker sent an Urgent Medical Device Recall Notification letter dated June 24, 2016, to the affected customers via courier, and to Sales Representatives and international affiliates via e-mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments at 269-389-2921. |  | 3163 | 3,163 | Worldwide Distribution - US (Nationwide) Netherlands, Canada, Japan, Switzerland, and United Kingdom | Terminated | March 23, 2017 | 272 | 3163 | Not_Computer
Z-2239-2016 | 74295 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | ATEC CANISTER. Suction Canister with Lid 490 cc., RX Only; Suction canister used to collect fluid during a breast biopsy. |  |  | May 31, 2016 | 2016 | Hologic, Inc | Cracks ATEC Canister Lids. | Process control | Customers were notified of the recall via certified letter beginning on May 31, 2016. The letter requests that customers immediately inventory your supplies of the Recalled product in all appropriate areas and cease use of all affected stock. In addition, they are asked to complete the enclosed Recall Response Card indicating whether you have any of the Recalled product on hand, as well as the quantity of product that you have in stock. |  | 16870 | 16,870 | Worldwide distribution. US: AK AL AR AZ CA CO CT DC DE FL GA HI IA ID IL IN KS KY LA MA MD ME MI MN MO MS MT NC ND NE NH NJ NM NV NY OH OK OR PA PR RI SC SD TN TX UT VA VT WA WI WV WY OUS: United Arab Emirates, Australia, Argentina, Belgium, Canada, Switzerland, Chile, Colombia, Germany, Denmark, France, Great Britain, Hong Kong, Ireland, Israel, Italy, Korea, Kuwait, Lebanon, Malaysia, Netherlands, New Zealand, Qatar, Singapore, Thailand, Taiwan, and South Africa. | Terminated | May 24, 2017 | 358 | 16870 | Not_Computer
Z-2243-2016 | 74547 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Parker Trach-Vac Endotracheal Tubes. Tyvek pouches, 10 per box. |  |  | June 02, 2016 | 2016 | Parker Medical | Parker Medical announces a voluntary field action for the Parker Trach-Vac Endotracheal Tubes - (Endotracheal tube,with Parker Flex-Tip,Sub-Glottic Suction, High Volume low Pressure) due to tube kinking during use. | Nonconforming Material/Component | Customers were notified on June 2, 2016 that a recall was initiated via telephone. The information communicated identified the affected product, problem and actions to be taken. Customers were instructed to segregate the affected product for immediate pick-up. A Sales Representatives went to each hospital and retrieved all unused product. |  | 5500 | 5,500 units | US Nationwide Distribution in the states of KY, MO, FL, TN, and TX. | Terminated | January 31, 2017 | 243 | 5500 | Not_Computer
Z-2241-2016 | 74468 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension VistaÂż Homocysteine Flex Reagent Cartridge (HCYS), Device Listing No.: D017878 is an in vitro diagnostic test for the quantitative measurement of total homocysteine in human serum, heparinized plasma and EDTA plasma on the Dimension Vista System. |  |  | May 17, 2016 | 2016 | CSL Behring GmbH | Reagent lots # 15243MA and 15243MB shows a reduced stability once opened that does not meet the Instructions For Use claim of seven (7) days. A maximum bias of -37% was observed. | Under Investigation by firm | Siemens mailed an Urgent Medical Device Recall to all customers who received the affected lots notifying them of the following: 1) Discontinue use and discard any remaining inventory of the affected lots, and 2) Indicate their replacement product needs within 30 days. |  | 898 | 898 units | Distributed to: AL, AR, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WI, WV | Open, Classified |  |  | 898 | Not_Computer
Z-2242-2016 | 74484 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | MODIFIED LINEAR STANDARD/LATERAL NECK TRIALS WITH SPRING SIDES REF S-200635 and S-200636 |  |  | June 21, 2016 | 2016 | Encore Medical, Lp | Reported failure of the instrument. The retaining ring component of a neck trial became disassembled from the body of the neck. | Process design | The recalling firm notified affected consignees via letter on 6/21/16. The consignees were instructed to return the affected devices. |  | 14 | 14 (7 sets) | CO, ID, TX | Terminated | January 10, 2017 | 203 | 14 | Not_Computer
Z-2237-2016 | 74486 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Replace Select Tapered TiU NP 3.5x11.5mm, Article No. 36105 |  |  | June 22, 2016 | 2016 | Nobel Biocare Usa Llc | Incorrect top label on an implant package. | Labeling False and Misleading | Nobel Biocare sent a customer notification letter dated June 23, 2016 to all affected customers to inform them that Nobel Biocare has detected that the lot#13030677 of Replace Select Tapered TiU NP 3.5 x 11.5 mm (product#36105) carries an incorrect top label. The letter informs the customers that the incorrect top label states "Replace NP 3.5 x 11.5 CC" (ie. Conical Connection). The product, Replace Select Tapered TiU NP 3.5 x 11.5 mm, has a Tri-lobe connection. Customers were instructed to complete the attached Quality Message Acknowledgement form. Customer with questions were instructed to contact customer service at 714-282-4808. |  | 53 | 53 units total (33 units in US) | Worldwide Distribution - US including OR, OH, NY and Internationally to Taiwan and BrazilNon US: Taiwan, SC | Completed |  |  | 53 | Not_Computer
Z-2244-2016 | 74513 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiographic Injector and Syringe, MEDRAD MRXperion MR Injection System, MEDRAD MRXperion Sterile Disposable MRI Kit |  |  | June 09, 2016 | 2016 | Bayer Healthcare | Bayer Healthcare is initiating this recall due to complaints that were received from customer sites describing a 4205 error message when the injector is used with a 3T scanner. | Device Design | Bayer sent an Urgent Medical Device Field Corrective Action letter dated June 8, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to take the following steps should a 4205 error message is encountered. For questions contact Bayer Customer Support Team at 1-800-633-7237. |  | 42 | 42 units | Worldwide and the US state of NEW YORK | Open, Classified |  |  | 42 | Not_Computer
Z-2240-2016 | 74576 | No | JDO | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TSP Hip Fracture Plating System |  |  | June 10, 2016 | 2016 | Biomet, Inc. | The screw may fracture at the junction of the shaft and the head during implantation. The patient may retain a foreign body if the fractured screw cannot be removed. Soft tissue irritation will not occur as the retained portion of the screw is below the surface of the plate. A delay in surgery less than 30 minutes may occur. | Device Design | Zimmer Biomet sent an Urgent Medical Device Recall letter dated June 10, 2016, to all affected customers via courier. Distributors and hospitals were instructed to locate and return the affected product, and implanting surgeons were instructed to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-6725, Monday through Friday, 8 a.m. to 5 p.m. |  | 5458 | 5458 units | Worldwide Distribution - US (Nationwide) Internationally to Australia, The Netherlands, and Japan | Terminated | July 20, 2017 | 405 | 5458 | Not_Computer
Z-2247-2016 | 74559 | [6 empty] | The Synchron Systems No Foam DxC (Reference No. B64130) affects the UniCel DxC System (chemistry analyzer). All lots of No Foam DxC reagent are affected by this recall. |  |  | June 27, 2016 | 2016 | Beckman Coulter Inc. |  | Under Investigation by firm | A Customer notification letter dated 6/27/16 informs their customers that Beckman Coulter has identified an incompatibility between No Foam DxC reagent (REF B64130) and the DxC hydropneumatic parts associated with No Foam delivery. The letter informs their customers that the DxC No Foam reagent is intended to prevent formation of foam and/or bubbles in the DxC waste system of the hydropneumatic subsystems. Over time, this incompatibility may lead to: Breaks in No Foam Cap Assembly (PN A84827) and Cracks in No Foam Waste Collector (PN A60007/A64665). The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the recall notice are instructed to contact our Customer Support Center Via http://www.beckmancoulter.com/customersupport/supportÂż Via phone, call 1-800-854-3633 in the United States and Canada. Customers outside the United States and Canada, are instructed to contact their local Beckman Coulter Representative. |  | 4221 | 4,221 units total (2,033 units in US) | Affected geographies: United Arab Emirates, Curacao, Angola, Argentina, Australia, Bangladesh, Belgium, Bulgaria, Brunei, Canada, Switzerland, Colombia, Costa Rica, Czech Republic, Germany, Ecuador, Egypt, Eritrea, Spain, Finland, France, United Kingdom, Greece, Guatemala, Hong Kong, Ireland, India, Iraq, Italy, Jordan, Japan, Korea, Republic of, Kuwait, Lebanon, Lithuania, Morocco, Macao, Mexico, Malaysia, Netherlands, New Zealand, Oman, Panama, Philippines, Pakistan, Poland, Portugal, Paraguay, Qatar, Romania, Serbia, Russia, Rwanda, Saudi Arabia, Singapore, Slovakia, Swaziland, Thailand, Tunisia, Turkey, Taiwan, China, United States of America (including Puerto Rico), Viet Nam, Mayotte, South Africa. | Open, Classified |  |  | 4221 | Not_Computer
Z-2245-2016 | 74620 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via, picture archiving and communications system software controlled. |  |  | June 23, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Incorrect values for the volume calculation. Software update VB30B via Update Instructions SY018/16/P to resolve software errors. | Software design | Siemens released and mailed to customers a notice with software update VB30B instructions via update instruction SY018/16/P for customers to follow until the update can be made. The letter also described the software issues involved; and asked customers to include the letter in the operator's manual until the update is performed. Customers are to ensure that all necessary personnel are aware of the situation. |  | 50 | 50 units | Nationwide Distribution -- AR, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, ND, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TX, and UT. | Open, Classified |  |  | 50 | Computer
Z-2248-2016 | 74264 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OT1000 Series Orthopedic Surgical Tables |  |  | May 23, 2016 | 2016 | Steris Corporation | STERIS has identified that the pivot pin assembly located on the table top may become loose preventing the table from articulating properly. | Other | Recall Notification Letters to the affected Consignees was sent on May 23, 2016 via FedEx with tracking numbers for delivery confirmation. STERIS Field Service Technicians will visit each domestic Customer location to perform the inspection and reinstallation of the pivot pin assemblies. For international direct accounts who have further distributed the surgical table, Steris asked to identify the Customers, and to notify them of this product field correction, and that not to operate the tables until the inspection has been done of the pivot pin assembly on all affected tables. |  | 148 | 148 | AL;AK, AZ; AR; CA; CO; CT; DC; FL; GA; IL; IN; IA; NY; NJ; NC; OH; OK; OR; PA; PR; SC; TX; UT and VA. Foreign accounts: Canada; France; Korea; Panama; Qatar and Thailand. | Completed |  |  | 148 | Not_Computer
Z-2249-2016 | 74249 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Amerigel Hydrogel Wound Dressing Advanced Formula for wound management 1 oz. |  |  | May 20, 2016 | 2016 | Amerx Health Care Corp. | Product did not meet specifications for Zinc Acetate (OOS). | Other | Initial recall notification was sent to distributors, informing them of the voluntary recall and providing instructions on expanding the recall to the wholesale level. Distributors were instructed to identify recalled product in their inventory, quarantine any recalled product, and to complete and fax back a Recall Response Form. Contact information was provided for those with questions: 1-800-448-9599 and recall@amerxhc.com. |  | 6993 | 6,993 Units | AR, AZ, CA, CO, CT, FL, GA, IL, IN, MA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, &amp; Alaska. Bahrain &amp; Colombia. | Open, Classified |  |  | 6993 | Not_Computer
Z-2246-2016 | 74308 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Model 7305P-D Powered Suction Unit |  |  | May 23, 2016 | 2016 | DeVilbiss Healthcare LLC | A small percentage of units of 7305 Series Vacu-Aide Suction Units are experiencing premature failure during operation. | Component design/selection | DeVilbiss Healthcare sent an Urgent - Medical Device Recall letter on May 23, 2016, to all affected customers notifying them of the issue. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to isolate all of the affected units included on the serial number list and return for replacement. For further questions, please call (814) 443-7602. |  | 48 | 48 units | US Distribution to the states of : NC, OH, TX, MO, TN, WI, KY, FL, LA, MN, IN, and SC | Terminated | May 30, 2017 | 372 | 48 | Not_Computer
Z-2250-2016 | 74604 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Convenience Kits from Centurion that |  |  | April 29, 2015 | 2015 | Centurion Medical Products Corporation | Kits from Centurion contain a Medtronic/Covidien Light Glove that is being recalled by the manufacturer, Medtronic.According to the notice received from Medtronic, the recall was initiatedbecause some light covers may have splits or holes, which may result in the transfer ofmicroorganisms from the light handle into the patient wound. | Nonconforming Material/Component | On 4/28/2015, URGENT PRODUCT RECALL NOTICE notifications were sent to the affected customers via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 483 | 483 | KS, IN, IA, KY, MT, SD | Completed |  |  | 483 | Not_Computer
Z-2252-2016 | 74537 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Trucount Tubes: |  |  | June 24, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | One lot of BD Trucount Absolute Counting Tubes (IVD) may not be sealed properly and exposure to ambient environment can rapidly degrade the performance of the product. | Packaging process control | Urgent Product Correction notifications letters were sent on June 24, 2016 to all affected customers. Letters identify the issue and product and provide illustrations of the Kit Box and pouch labels.Correction letters identified the issue and the product involved with photos of product labels. BD letters states that they want to reinforce the need to visually inspect the desiccant contained in each pouch . If the desiccant has turned from blue to lavender, discard the remaining tubes.The acknowledgement form should be completed and returned regardless if any inventory remains. If you require further assistance, please contact BD Customer Support at 855.236.2772 (prompt 3) in the Unites States. For customers outside the US, contact your local BD Biosciences representative or distributor |  | 844 | 328 US, 516 International | USA Distribution to the states of : NC, WA, WI, NJ, PR, MD, NY, MA, CA, NY , MD, KY, IN, VA, SD, PA, OH, AR, AZ and Foreign to : Canada, Belgium and Chile | Open, Classified |  |  | 844 | Not_Computer
Z-2256-2016 | 74309 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIAÂż Centaur XPT system software versions V1.0.1 (Bundle 1.0.912 SMN 10819704), V1.0.2 (Bundle 1.0.1086 SMN 11219806), V1.0.3 (Bundle 1.0.1108 SMN 11220781) and V1.1 (Bundle 1.1.243 SMN 11221979). |  |  | May 26, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Eight (8) issues were identified which may affect the results generated by the system software version. | Under Investigation by firm | On May 26, 2016, Siemens Healthcare Diagnostics Inc. distributed an Urgent Medical Device Correction notices to their customers via courier service. The letter identified the affected product, problem and actions to be taken. Customers are advised to complete and return the Field Correction Effectiveness Check form within 30 days via fax to (312) 275-7795. For questions contact your local Siemens technical support representative. |  | 446 | 446 units Total (3 domestically &amp; 443 internationally) | Worldwide Distribution - US Nationwide in the states of MD, MN, MS and the countries of Australia, Austria, Bangladesh, Belgium, Canada, Chile, Czech Republic, Denmark, Egypt, Finland, France, Georgia, Germany, Guadeloupe, Hong Kong, Hungary, India, Ireland, Italy, Japan, Jordan, Latvia, Malaysia, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Qatar, Republic Korea, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, U.A.E, United Kingdom &amp; Vietnam. | Open, Classified |  |  | 446 | Computer
Z-2253-2016 | 74447 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Anspach G1 High Speed Electric Handpiece Air Cooled (eG1A) |  |  | June 15, 2016 | 2016 | The Anspach Effort, Inc. | The graphics at the attachment interface indicating locked and unlocked position are reversed. | Employee error | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE-RECALL (Removal)" letter dated June 15, 2016 to its customers. The letter described the product, problem and actions to taken. The customers were instructed to review your inventory and immediately remove the product subject to this recall from stock; If you DO HAVE any of the identified devices, please take the following steps:" Ensure anyone in your facility impacted by this notification reads this letter carefully.Customer Letter Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating product subject to the recall has been located. Indicate the number of devices found." The RMA information has been filled in for your convenience." Please include your name, title, address, telephone number, signature and date in the spaces provided." Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Department by:" Fax: (561) 627-2682" Scan/email: DPYUS-PowerToolsFieldActions@its.jnj.com" Giving it to your DePuy Synthes Sales Consultant." This return documentation acknowledges your receipt of medical device removal information." Return the affected product to your DePuy Synthes Sales Consultant" Keep this notice visibly posted for awareness until all products subject to this recall has been returned. While processing your returns, please maintain a copy of this notice with the product subject to this action and keep a copy for your records.If you DO NOT HAVE the identified product, please take the following steps:" Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Include your name, title, address, telephone number, signature and date in the spaces provided." Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Department by Fax: (561) 627-2682" Scan/email: DPYUS-PowerToolsFiel |  | 18 | 18 | Worldwide Distribution: US Distribution to states of: GA and FL; and country of: Switzerland. | Open, Classified |  |  | 18 | Not_Computer
Z-2254-2016 | 74221 | No | OWD |  |  | Molecular Genetics | PMA | cobas EGFR Mutation Test V2 |  |  | April 05, 2016 | 2016 | Roche Molecular Systems, Inc. | During internal studies using contrived plasma specimens, several mutations (L8568R, Exon 19 deletion, T790M) inconsistently generated "No Mutation Detected" (ie False Negative ) result with the cobas EGFR mutation Text V2.0 when utilizing the cobas cfDNA Sample Preparation kit. | Nonconforming Material/Component | Roche Molecular Diagnostics sent an Urgent Field Safety Notice letter dated April 5, 2016 to affected customers. Then letter identified the affected product, problem and actions to be taken. The Instructions for use are provided for the user until new ones are available.. |  | 1244 | 735 devices EGFR; 509 devices cfDNA | Foriegn Distribution Austria, Belgium, Croatia, CzechRepublic, Denmark, France, Germany,Greece, Hungary , Italy, Lithuania,Luxembourg, Netherlands, Norway, Poland,Romania, Slovakia, Slovenia, Spain,Switzerland, Turkey, United KingdomcfDNA: Austria, Belgium, Croatia, CzechRepublic, Denmark, France, Germany,Greece, Hungary , Italy, Lithuania,Luxembourg, Netherlands, Poland,Romania, Slovakia, Slovenia, Spain,Switzerland, Turkey, United Kingdom | Open, Classified |  |  | 1244 | Not_Computer
Z-2265-2016 | 74505 | No | KTN | Infectious mononucleosis immunological test system. | Immunology | Immunology | 510(k) | Acceava Mono II Rapid Test Strip (Whole Blood, Serum, Plasma), Catalog No. 92404. Mononucleosis Rapid Test Device a simple test that utilizes an extract of bovine erythrocytes to qualitatively and selectively detect IM heterophile antibodies in whole blood, serum or plasma. |  |  | June 22, 2016 | 2016 | Alere San Diego, Inc. | The Positive Control information referencing human plasma as the antibody material described in the product labeling and packaging is incorrect. | Unknown/Undetermined by firm | A Medical Device Correction letter dated 6/22/16 will be sent to all customers to inform them that Alere is recalling the MONO Rapid Test products because the Positive Control Information referencing human plasma as the antibody material described in the product labeling and packaging is incorrect. Customers are informed that the labeling states the Positive Control contains the following: Diluted human plasma containing IM heterophile antibodies, 0.09% sodium azide. The Positive Control contains the following: Goat anti-mono antibody, 0.09% sodium azide. Customers are instructed to complete and fax or email the enclosed verification form within 10 days. Customers with questions are instructed to contact Alere San Diego at (844) 804-4945 or email at verfications.ts@alere.com. |  |  |  | US | Open, Classified |  |  |  | Not_Computer
Z-2096-2016 | 74301 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Soft Vu Omni Flush Angiographic Catheter |  |  | June 01, 2016 | 2016 | Stryker Sustainability Solutions | Product defect; tip of two (2) Angiographic Catheters model # 10732203 broke off, prior to being used on the patient. | Incorrect or no expiration date | URGENT MEDICAL DEVICE RECALL RESTERILIZED ANGIODYNAMICS SOFT VU OMNI FLUSH ANGIOGRAPHIC CATHETER customer notification letters were sent 06/01/16.Customers were instructed to discontinue use of the recalled product immediately. Customers were told to complete the Recall Effectiveness Check Form and list the affected lots shipped to the facility. The form should be completed, signed, and returned to the local Stryker Sustainability Sales Representative, emailed to ssspfa@stryker.com, or mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ 85283Attn: Jodie Rueckert Customers will receive a credit for all affected devices returned. |  | 167 | 167 | Worldwide Distribution -- USA, to the states of TX, UT, PA, CA, VT, MA, WA, VA, NY, MD, and CT; and, the country of CANADA. | Terminated | October 19, 2016 | 140 | 167 | Not_Computer
Z-2212-2016 | 74465 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | DIAMOND-FLEX CIRCULAR RETRACTOR, ANGLED, 40MM 5MM, Product Code 89-6114, Lot Code(s): 878971; 879366; 879404, and Date Code: C16. |  |  | July 01, 2016 | 2016 | Carefusion 2200 Inc | BD, formerly CareFusion, has identified a potential risk associated with a weld failure which could result in the wire protruding thru the tip of the instrument when articulated. If this failure were to occur while in use in a procedure it has the potential to damages tissue or organs | Material/Component Contamination | BD, sent an "URGENT: Medical Device/ Safety Alert/Recall Notification" letter dated 6/30/2016 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to return the affected units, along with the enclosed CUSTOMER RESPONSE FORM to: BD formerly CareFusion, 75 North Fairway Drive, Vernon Hills, Illinois, 60061 Attn: Customer Advocacy; to expedite the correction process and acknowledge receipt of the notification. The firm will issue a credit upon receipt of the returned affected product.For questions and support 1-800-323-9088 Prompt 3 or email: GMB-US-Complaint-Intake@carefusion.com. |  | 10 | 10 | Worldwide Distribution - US, including the states of WA and MD; and, the country of Japan. | Terminated | March 29, 2017 | 271 | 10 | Not_Computer
Z-2272-2016 | 74569 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC Synchron Clinical Systems and UniCel DxC Synchron Access Integrated Clinical Systems |  |  | April 22, 2014 | 2014 | Beckman Coulter Inc. | Possible leaking of No Foam solution from the No Foam Bottle Assembly on the UniCel DxC Synchron Systems. | Nonconforming Material/Component | An Important Product Notice dated 4/22/14 was sent to all their customers to inform them that Beckman Coulter has identified a manufacturing issue with a plastic fitting on the No Foam Bottle Assembly which may crack and cause No Foam solution to leak. The letter informs the customers of the impact, action to be taken, and resolution. Customers are asked to inspect the No Foam Bottle Assembly for leaks prior to refilling the bottle. If leaks occur, customers should contact Customer Support Center following the instructions provided. Customers are instructed to complete the enclosed response form within 10 days. Customers with questions are instructed to contact the firm's Customer Support Center- Via our website, http://www.beckmancoulter.com/customersupport/support- Via phone, call 1-800-854-3633 in the United States and Canada- Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 3540 | 3,540 units total (1,067 units in US) | Worldwide Distribution -- United Arab Emirates, Angola, Australia, Belgium, Bulgaria, Canada, Switzerland, China, Colombia, Costa Rica , Czech Republic, Germany, Denmark, Egypt, Eritrea, Spain, Finland, France, United Kingdom, Greece, Guatemala, Hong Kong, Indonesia, Ireland, India, Italy, Jordan, Japan, Korea, Republic of, Kuwait, Lebanon, Lithuania, Morocco, Mexico, Malaysia, Namibia, Netherlands, Norway, New Zealand, Oman, Panama, Philippines, Pakistan, Poland, Portugal, Paraguay, Qatar, Romania, Serbia, Russian Federation, Rwanda, Saudi Arabia, Sweden, Singapore, Slovakia, Swaziland, Thailand, Tunisia, Turkey, Taiwan, US (including Puerto Rico), Vietnam, South Africa, and Zimbabwe. | Open, Classified |  |  | 3540 | Not_Computer
Z-2263-2016 | 74043 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Dual Chamber Container fill with gravity/attached transfer sets, 250 mL / 1250 mL, 300 mL / 1700 mL, 500 mL / 2500 mL, 500 mL / 3500 mL |  |  | June 23, 2016 | 2016 | The Metrix Company | An issue was identified after a redesign of the auto compounding device. The compounder cannot meet the designed accuracy specification when dispensing into bags. After reviewing the initial design of the device, it was determined the dispensing volume validation was not properly done and none of the compounders on the market can meet the stated specifications. | Device Design | On 6/23/2016, the recalling firm sent written recall notices to their distributors and end users notifying them of the recall. Customers and consignees were instructed to not use the devices, locate and quarantine all affected devices, and to complete the response form. |  | 18474 | 18,474 case | Product distributed to medical facilities in the following US states: AL, AZ, CA, CO, CT, FL, IA, IL, KS, LA, MA, MD, MN, NC, OH, PA, PR, TN, TX, UT, and WA. Product was also distributed internationally to the following countries: Australia, Canada, Dominican Republic, Malaysia, Turkey, and United Kingdom | Open, Classified |  |  | 180049 | Not_Computer
Z-2251-2016 | 74637 | No | MOT |  |  | Radiology | 510(k) | Best Theratronics cabinet x-ray systems intended for use in the irradiation of blood and blood products (packaged in transfusion bags) to inactivate T-lymphocytes for the prevention of Graft Versus Host Disease. |  |  | May 06, 2016 | 2016 | BEST THERATRONICS LTD. | It was discovered that both the new device and the predicate device failed to comply with performance standard. | Radiation Control for Health and Safety Act | Best Theratronics ( BTL) planned action to bring into compliance:BTL proposes that the repairs are to be completed during annual routinepreventive maintenance visits at the user's site to minimize downtime, which disrupts the user's ability to irradiate blood. A standard repair kit has been developed. For further information, please contact Best Theratronics Customer Service at 1-866-792-8598 |  | 128 | 128 | Nationwide | Terminated | August 01, 2017 | 452 | 128 | Not_Computer
Z-2273-2016 | 74355 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products Ca Slides, GEN 29, REF/Product Code 145 0261, Unique Device Identifier No. 10758750009114 or 20758750009111, IVD -- Sales Unit: 300 slides/pack |  |  | May 18, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) confirmed a potential for biased calcium (Ca) results to be generated when using VITROS Ca Slides, GEN29, and the VITROS 250/350/950/5,1 FS, 4600 Chemistry System or the VITROS 5600 Integrated System. An increase in negatively biased calcium results complaints prompted the investigation that identified this product issue. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an URGENT PRODUCT CORRECTION NOTIFICATION letter dated May 18, 2016 to affected customers. . The letter identified the affected product, problem and actions to be taken. The letter instructed to discontinue use of the affected lot(s) and discard all remaining inventory. Customers were asked to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. |  | 32382 | US: 27,625 sales units; Foreign: 4757 sales units | Worldwide Distribution - US Nationwide in the countries of Puerto Rico and Foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, Poland and Venezuela. | Open, Classified |  |  | 32382 | Not_Computer
Z-2274-2016 | 74567 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Allegra X-5 rotor yoke, Catalog No. (Instruments) B30589 and B30590 (Rotor) B30593. Product Labeling: IFU B29071AA. |  |  | February 05, 2016 | 2016 | Beckman Coulter Inc. | A defective rotor yoke on Allegra X-5 allows improper rotor seating on the hub resulting in possible rotor failure if the buckets are loaded in an imbalanced configuration. | Nonconforming Material/Component | Beckman Coulter sent an Important Product Notice Recall letter dated February 5, 2016 to customer. The letter identified the affected product problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form. Customers with questions are instructed to contact our Customer Support Center" Via website, http://www.beckmancoulter.com/customersupport/support" Via phone, call 800-369-0333 the United States and Canada" Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 62 | 62 units total (33 units in US) | Worldwide Distribution - US Nationwide (including Puerto Rico) and the countries of: Australia, Canada, Colombia, Germany, Guatemala, Netherlands, Singapore, Tunisia, , Spain | Open, Classified |  |  | 62 | Not_Computer
Z-2236-2016 | 74554 | No | RCE |  |  |  |  | GE Inspection Technologies, LP - Phoenix v¦tome¦x m240 cabinet x-ray systems. |  |  | March 02, 2016 | 2016 | GE Inspection Technologies, LP | Radiation emissions which exceeded the limit in the performance standard were detected during maintenance of one system. The identified cause was insufficient attachment of shielding in the door. Jostling during delivery was identified as the probable cause of this failure to comply with the performance standard. | Radiation Control for Health and Safety Act | GE plans to notify its customers with the systems through a TIL. The TIL will inform customers that GE will add lead strips to the inside window frame under the lead glass window of the main cabinet door. GE recommends that it visit each customer site and repair the system by adding lead strips along the window frame of the cabinet door, beneath the lead glass window, within 90 days following FDA's approval of the repair plan. |  | 4 | 4 | US Distribution to the states of : FL, TN, and TX. | Terminated | August 07, 2017 | 523 | 4 | Not_Computer
Z-2279-2016 | 73985 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultrasound Gastroscope |  |  | April 06, 2016 | 2016 | Pentax Medical Company | Pentax Medical Video Processor EPK-i5010 has been shipped with a reusable Pentax Medical water bottle assembly (OS-H4) which has not been cleared for use with this devices in the United States. | Device Design | Pentax Medical sent an "Urgent Field Correction Letter and Response Form" dated April 6, 2016 via Certified Mail to their affected customers. The recall letter identified the product, problem and action to be taken by customer. The customer is asked to complete the response form even if the customers does not have any OS-H2/OS-H4 water bottles to discard. If customer needs assistance in completing the response form, they can contact their local sales representative. For questions regarding the recall, customers can contact Mr. Paul Silva, Pentax Americas Recall Coordinator at: (800) 431-5880 ext. 2064, from (8:30 am - 5:00 pm, Monday - Friday, EST), via fax: (201) 799-4063 or via email to: recall.coordinator@pentaxmedical.com. |  | 819 | 819 units | US Nationwide Distribution | Terminated | July 17, 2017 | 467 | 23755 | Not_Computer
Z-2288-2016 | 74563 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AnastoCiip Universal Clip Remover (6 Pack} |  |  | June 22, 2016 | 2016 | LeMaitre Vascular, Inc. | US marketing brochure for the AnastoClip AC ClosureSystem states that the Anastoclip catalog numberscan be used for dura closure, but the catalog numbers are only intended for cardiovascular use. | Unknown/Undetermined by firm | The firm, LeMaitre Vascular, Inc., sent an "URGENT: Medical Device Correction" letter dated June 20, 2016 to its distributors and hospitals via Priority mail. The letter describes the product, problem and actions to be taken. The customers were informed of the intended use of the devices and the correct catalog numbers to order for the dura application. The customers were instructed to return unused devices for an exchange or refund; destroy any Anastoclip related marketing materials that you have received from Lemaitre between October 25- 2015 and June 15, 2016, and complete and return the reply form for reconciliation of this recall via fax at 781-221-2223, email: CSUS@LEMAITRE.COM or by mail. LeMaitre Vascular, Inc. will replace any affected device with a new device labeled for dura closure or issue a credit per request.Questions concerning this safety notice, please contact me at 781-221-2266 ext. 183. |  | 55 | 55 units (6 pack) | US Nationwide Distribution | Open, Classified |  |  | 1051 | Not_Computer
Z-2281-2016 | 74500 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Trident Constrained Liner Inserter/Impactor Tip, 22 mm, 28 mm, and 32 mm. Non Sterile |  |  | June 03, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Impactors do not meet the required sterility assurance level (SAL) of 10-6 when using the recommended steam sterilization method. | Process design | Stryker notified its branches/agencies of this recall by e-mail on May 26, 2016. They were asked to quarantine the affected devices. Stryker sent a recall letter and business reply form dated June 3, 2016 to their affected customers via UPS. Stericycle will be handling the returns. |  | 1864 | 1,864 units | Worldwide distribution. US nationwide (AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, NJ, NY, OH, PA, TX, WA, WI, WV), Sweden, South Africa, Chile, Australia, New Zealand, Canada, United Kingdom, Switzerland, Colombia, and Poland. | Open, Classified |  |  | 1864 | Not_Computer
Z-2290-2016 | 74596 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie Q7 wheelchair. Model EIR4 |  |  | July 01, 2016 | 2016 | Sunrise Medical (US) LLC | There is an observed failure of the backrest bracket which can fatigue over the life time of the product. When this occurs there is the possibility of the bracket fracturingcausing the backrest to detach from the chair frame. | Device Design | The firm began calling dealers to alert them to the recall and obtain primary contacts information. Dealers were then sent their list of Quickie Q7 wheelchairs that records show being sold by them and dealers are asked to return a list of chairs they are able to locate for correction. The first dealer letter was sent via e-mail on July 7, 2016. |  | 13852 | 12,142 devices in US, 1,710 devices internationally. | US and Mexico, Japan, Chile, Kuwait, Columbia, Hong Kong, China, Uruguay, South Africa, Costa Rica, South Korea, Singapore, New Zealand, El Salvador, Bahrain, Uruguay, Lebanon, United Arab Emirates, Brazil, Argentina, Netherlands Antilles, British Virgin Islands, | Open, Classified |  |  | 13852 | Not_Computer
Z-2282-2016 | 74452 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instructions for Use (IFU) for Progesterone 3 Calibrator Set, repackaged as one set/package. |  |  | June 07, 2016 | 2016 | Tosoh Bioscience, Inc. | Instructions For Use had incorrect lot number and calibrator values | Error in labeling | Tosoh Bioscience sent an Urgent Correction Notification letter dated June 7, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call Tosoh Technical Support at 1-800-248-6764. |  | 123 | 41 single sets of Progesterone 3 calibrators | US Nationwide Distribution in the states of AL; CA; CO; FL; KY; MA; NY; OH; SC; TX; UT; VA and WI. | Open, Classified |  |  | 123 | Not_Computer
Z-2291-2016 | 74529 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Molift Air 205, Molift Air 205 Propulsion, Molift Air IRC, Molift Air 205 IRC Propulsion, Molift Air 300, Molift Air 300 Propulsion, Molift Air 300 IRC, Molift Air 300 IRC Propulsion |  |  | June 24, 2016 | 2016 | Moller Vital | Potential for the bolt on the suspension coupling on the Molift Air to come loose from the housing. | Device Design | Moller Vital sent a Field Safety Corrective Action on June 22, 2016. Field correction instructions were provided to affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to read the instruction, and return the recipient confirmation, report all serial numbers distributed in their region/country to Etac Supply Gjovik, order the upgrade kit from Etac Supply Gjovik, Customers were also instructed to follow up with the customers and users and confirm that all hoists in their market are upgraded. Customers with questions were instructed to contact Etac Supply Gjovik.at molift.qa@etac.com. |  | 516 | 516 units (374 US) | Worldwide Distribution - US (nationwide) and Internationally to SZ, TR, BE, CH, DE, DK, ES, FI, FR, GB, IS LV, NL, NO, SE, SI, AU, CN, IL, JP, NZ, IN, QA, PE, AE | Terminated | June 23, 2017 | 364 | 516 | Not_Computer
Z-2292-2016 | 74693 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Avenir MÂżller Stem 4 Standard |  |  | June 30, 2016 | 2016 | Zimmer Gmbh | Product mix-up. The Avenir MÂżller Stem 6 lateral uncemented might be placed in the packaging of the Avenir MÂżller Stem 4 standard uncemented and vice versa.Potential for delay in surgery, and/or potential for use of incorrect implant which may in turn lead dismetry in patient and possibly the need for revision. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall letter dated June 30, 2016 to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification immediately and ensure affected personnel are aware of the contents and assist Zimmer Biomet sales representative with the quarantine of any affected device. Your Zimmer Biomet sales representative will remove any affected device. Customers were also instructed to complete the Certification of Acknowledgement form and return to corporatequality.postmarket@zimmerbiomet.com. Customers with further questions or concerns were instructed to contact the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST.For questions regarding this recall call 800-348-2759. |  | 2 | 2 | Nationwide Distribution to OH and NM | Terminated | March 27, 2017 | 270 | 2 | Not_Computer
Z-2294-2016 | 74222 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Endotine Transbleph 3.5 (CFD-080-0167) |  |  | May 06, 2016 | 2016 | MicroAire Surgical Instruments, LLC | MicroAire has initiated a limited recall (29 Units)on the Endotine Transbleph 3.5 (CFD-080-0167), Lot # 285373, due to minor mislabeling ( the front label is correct, but the side label lists an incorrect part number, i.e. CFD-080-0177). | Labeling Change Control | The firm is sending recall letters to the customers who have received the Endotine Transbleph 3.5 (CFD-080-0167), lot #285373. Letters sent by UPS to expedite the delivery. |  | 29 | 29 | The product was distributed in the United States and Sweden | Completed |  |  | 29 | Not_Computer
Z-2293-2016 | 74668 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | July 15, 2016 | 2016 | Elekta, Inc. | It is possible that a change to an Order Set will not be saved in the current open Care Plan. This can result in a wrong chemo dose in the order. | Software Design Change | Elekta sent Important Field Safety Notice 371-03-MSQ-001 to all customers on 7/15/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. The letter also included a response form which is to be returned to Elekta. |  | 383 | 383 | Worldwide distribution. US nationwide, Antigua and Barbuda, Australia, Austria, Bahamas, Bahrain, Canada, Czech Republic, Germany, Hungary, South Africa, Spain, Sweden, Turkey, and United Kingdom | Open, Classified |  |  | 383 | Not_Computer
Z-2296-2016 | 74583 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Integra External Fixation System Rocker Bottom Model # 12225400. |  |  | July 01, 2016 | 2016 | Integra LifeSciences Corp. | The silicone insert embedded in the rocker bottom has the potential to become dislodged or loose from the rocker bottom mold. | Device Design | All affected consignees were notified via an "Urgent: Voluntary Medical Device Recall" letter sent on 7/1/16. The letter identified the affected product as well as the reason for the recall. Customers were asked to check their traceability records for shipments of the affected product. If the affected parts are found, customers are to stop use and distribution. A copy of the letter should be provided to any customers that the product may have been distributed to. The attached "Acknowledgement and Return Form" should be completed and returned regardless if customers have affected product or not. Questions regarding the instructions should be directed to Customer Service at 1-800-654-2873. |  | 27 | 27 units | Distributed in the US to the states of FL, VA, OK, CO, DC, and IN. | Open, Classified |  |  | 27 | Not_Computer
Z-2298-2016 | 74524 | No | KLS | Theophylline test system. | Toxicology | Toxicology | 510(k) | SyvaÂż EMITÂż 2000 Theophylline - Catalog # 4P019UL /SMN#10445324 is a homogeneous enzyme immunoassay intended for use in the quantitative analysis of theophylline in human serum or plasma. |  |  | June 17, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the SyvaÂż EMITÂż 2000 Theophylline reagent lot G3, when run ona Beckman AU Clinical Chemistry System, and the ADVIAÂż Chemistry Theophylline_2 (THEO_2) reagent lot334018, may exhibit increased imprecision for commercially available Bio-Rad Liquichek TDM Quality ControlLevels and patient samples. | Process control | Siemens Healthcare Diagnostics mailed a letter to customers on June 17, 2016, asking that they discontinue use of and discard the affected kit lots identified in the recall notification letter. Customers were also asked to complete and return the Field Correction Effectiveness Check/Product Replacement Form that was mailed with the Recall Notification Letter. |  | 1004 | 1,004 units | Distributed to: CA, CT, DE, FL, GA, IA, IN, KS, KY, LA, MA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, PA, TN, VA, WA | Terminated | March 15, 2017 | 271 | 1662 | Not_Computer
Z-2300-2016 | 64758 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Imagecast PACS with Centricity RIS-IC versions prior to 10.6 Update Package 18 |  |  | February 18, 2013 | 2013 | Ge Healthcare It | A software defect was discovered that causes images to be out of context with clinical information. | Software design | GE sent an "Urgent Medical Device Correction" letter to all affected customers on April 2, 2013.The letter identified the product the problem and the action customers need to take.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI #85204.For further questions, please call ( 847) 277-4468. |  | 13 | 13 | US Nationwide Distribution. | Open, Classified |  |  | 13 | Computer
Z-2295-2016 | 74523 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Babytherm 8004/8010, Infant Warmer, Class II device, 510 K K971198, Catalog # FR00105 (Babytherm 8010), 2M30404 (Babytherm 8004), non sterile, software controlled. |  |  | June 22, 2016 | 2016 | Draeger Medical, Inc. | Risk associated with operating the device in skin temperature mode with temperature sensors attached but no patient connected. | Nonconforming Material/Component | Drager mailed to customers an Urgent Medical Device Recall Notification Letter dated June 22, 2016 to affected customers. The letter identified the affected product, problem information on how to operate the device and replacement of parts information. For questions regarding this letter please call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger Babytherm 8004/8010 please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 1599 | 1,599 units | US Nationwide Distribution including Puerto Rico and the states of AK,AL,AR,CA,CO,CT,DE,FL,GA,HI,IA,IL,IN,KS,KY,LA,MA,MD,MI,MN,MO,MS,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,TN,TS,TX,UT,VA,WA,WI,WV, and WY. | Open, Classified |  |  | 1599 | Not_Computer
Z-2297-2016 | 74617 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SPHERICAL HEAD REFLECTION(R) 6.5 MM CANCELLOUS SCREW, REF 71332525, QTY: (1), STERILE R, smith&nephew; |  |  | July 06, 2016 | 2016 | Smith &amp; Nephew, Inc. | The labels correctly state that the product should be a 25mm screw, but the box contained a longer, 35mm screw. | Packaging process control | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated July 6, 2016, to all affected consignees via email and FedEx Overnight delivery on July 6, 2016. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were asked to inspect their inventory and located any unused affected devices and quarantine them immediately, complete the Inventory Return Certification Form and fax to 1-901-566-7975 indicating the quantities that need to be returned. Consignees should contact Smith &amp; Nephew's Field Actions Department via email at FieldActions@smith-nephew.com to obtain a return authorization (RA) number.For questions regarding this recall call 901-399-5520. |  | 185 | 185 units | Worldwide Distribution - US (nationwide) and Internationally to AU, CA, GB, IN, ZA, and TR | Terminated | July 01, 2017 | 360 | 185 | Not_Computer
Z-2302-2016 | 74366 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Total T4 Reagent Pack, REF/Product Code 874 4468, Unique Identifier No. 10758750005147, IVD --- Ortho-Clinical Diagnostics, |  |  | June 02, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho-Clinical Diagnostics confirmed that sample results generated from VITROS Immunodiagnostic Products Total T4 Reagent Packs may show a negative bias prior to the open pack storage limit of &lt;= 8 weeks that is located in the Instructions for Use (IFU) for VITROS Immunodiagnostic Products Total T4 Reagent Pack. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter (Ref. CL2016-112) dated June 2, 2016, to all affected customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received the affected product to inform them that sample results generated using VITROS Total T4 Reagent Packs that are opened for more than 4 weeks may produce a negative bias and advised them to only use the Total T4 Reagent Pack up to 4 weeks after the pack has been opened. Foreign affiliates were informed of the issue by e-mail on 6June 2, 2016, and instructed to notify their consignees of the issue and required actions. Customers with questions were instructed to contact Global Technical Solutions Center at 1-800-421-3311. |  | 278 | US: 12,015; Foreign: 55 Sales Units to Canada, 223 Sales Units shipped to Bermuda | Worldwide Distribution - US (Nationwide including Puerto Rico) and foreign distribution to Canada and Bermuda | Open, Classified |  |  | 278 | Not_Computer
Z-2303-2016 | 74651 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, Wisconsin. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | A migrated study that has annotations will display the annotations, but when a prior study is loaded that should have annotations, the second study will not display annotations. | Software design | The recalling firm, Merge, issued an "URGENT: MEDICAL DEVICE RECALL" letter dated 1/27/2016 via e-mail on 1/30/2016 or via certified mail if they did not have an e-mail address for the customers. The letter described the product, problem and actions to be taken. The customers were instructed to not discontinue use of Merge PACS; to open prior comparison studies in the secondary viewer to check if there were annotations and read the report or visit note pertaining to that study to see if there are noted findings that were not otherwise apparent; to receive an upgrade call Merge Customer Service at (877) 741-5369 or send an email to support@merge.com to schedule the upgrade (Reference this call #2016-010); complete the online form: http://www.merge.com/Landing-Pages/Customer-Upgrade-Decline-Form.aspx no later than 2/22/2016; If you decide to decline the maintenance release, complete the attached Customer Decline Response form and return to Merge Healthcare via recall@merge.com, and ensure all users of the product are provided with this notification. If you have further distributed this product, identify your customers and notify them at once of this product recall.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 208 | 208 sites have the affected versions | Worldwide Distribution: US (Nationwide) and countries of: Australia, Belgium, Canada, Jordan, New Zealand, and the United Kingdom. | Open, Classified |  |  | 208 | Not_Computer
Z-2255-2016 | 74593 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Model: Analytical X-Ray System X -Ray Fluorescence. |  |  | June 22, 2016 | 2016 | Olympus Scientific Solutions Americas | Trigger lock option does not always prevent x-ray generation. Specifically if the unit is in test mode or repeat test mode then the trigger lock may not prevent x-ray initiation when the trigger is pulled. The correction is via repair by a software update. | Radiation Control for Health and Safety Act | The firm notified customers via an "Important - Electronic Product Radiation Warning" letter. It described the issue; and, told customers to not rely on the Trigger Lock feature until the software upgrade had been completed. The letter also provided instructions for customers to not only determine their software version; but, also instructions to bring their affected device into compliance. The firm will remedy the issue without charge. Customers can contact their local representative or OSSA at ani.technicalsupport@olympus-ossa.com, phone: +1 866 446 6689, fax: +1 781 419 3980, with questions or if additional information is needed. |  | 212 | 2, 210 units distributed in the USA | Nationwide Distribution. | Open, Classified |  |  | 212 | Computer
Z-2301-2016 | 74648 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Studies coming over via telmed were missing patient's DOB, procedure, and referring physician. | Software design | The recalling firm issued Urgent: Medical Device Recall letters dated 1/27/2016 and 5/26/2016. The letter informed the customer of the affected product, the nature of the issue, the potential harm involved, and actions to be taken. The firm released a fix for the issue; and, customers were asked to complete and return the enclosed form by June 17, 2016. Questions can be directed to recall@merge.com. |  | 109 | 109 sites have the affected versions | Worldwide Distribution -- USA, New Zealand, and United Kingdom. | Open, Classified |  |  | 109 | Not_Computer
Z-2304-2016 | 74312 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | ONLINE TDM Vancomycin 100 tests; ONLINE TDM Vancomycin 200 tests; Hitachi Vancomycin (Modular P); Hitachi Vancomycin (917, MOD P) |  |  | May 26, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche Diagnostics Operations, Inc. has issued a voluntary recall for the Online TDM Vancomycin assay on the cobas c 311/501/502 analyzers and the Modular Analytics P module due method sheets that state an incorrect method comparison against the COBAS INTEGRA 800 analyzer. | Under Investigation by firm | Roche sent an Urgent Medical Device Correction letter dated May 26, 2016, to all affected customers via UPS Ground (receipt required) . The letter identified the product the problem and the action needed to be taken by the customer.Customers will be informed to disregard the method comparison data in the package insert and to follow the instructions outlined in UMDC 16-083. Non-responding accounts will be defined as users for which confirmation of receipt of the Urgent Medical Device Correction cannot be obtained. These accounts will be monitored on an ongoing basis and follow-up attempts will bemade accordingly. Affected product under Roche control was placed on hold. Affected product will be reworked to add a sticker referring to customers to UMDC 16-083. Roche is not requesting return of the affected product. Any returned product will be managed according to current procedures. The incorrect method comparison will be removed from the method sheets. Updated method sheets will be released at a later time. Customers may contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336. |  | 53897 | 53,897 Units distributed | Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN,TX, UT, VA, VT, WA, WI, WV, WYForeign:NoneVA/DOD: See list below | Terminated | March 23, 2017 | 301 | 53897 | Not_Computer
Z-2312-2016 | 74477 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | 3T 16ch Flex SPEEDER Coil, Large, NeoCoil Part Number NC046200, Toshiba Medical Systems Model MJAJ-222A. The NeoCoil 3T 16ch Flex SPEEDER Coils are tuned to receive RF frequency corresponding to the proton precession in a 3 tesla magnetic field, which is governed by the Larmor equation. |  |  | June 17, 2016 | 2016 | NeoCoil, LLC | There is a potential for higher than specified surface temperatures involving the 3T 16ch Flex SPEEDER Large Coil. | Nonconforming Material/Component | NeoCoil sent an "Urgent Medical Device Recall" letter dated 17 June 2016. The letter described the problem and the product involved in the recall. Advised consignee to quarantine the product and if the product was further distributed to identify the customers and provide the list to NeoCoil. The recall will be carried to the user level. Requested consignee to complete and return the response form. For questions call 262-522-6127. |  | 39 | 39 (11 US, 28 OUS) | Worldwide Distribution - US and Internationally to Georgia, Germany, Korea, Netherlands, Pakistan, Spain, Japan, and the United Kingdom. | Open, Classified |  |  | 39 | Not_Computer
Z-2311-2016 | 74600 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | AmediCheck Instant Test Cup, Part No. C-C0501A-3A |  |  | July 05, 2016 | 2016 | Branan Medical Corporation | Branan Medical is recalling the AmediCheck Instant Test Cup, a 5 panel drug cup, because it may erroneously contain 12 panel cups due to a packaging error. | Labeling mix-ups | BMC sent an Urgent Medical Device Recall Notice letter on July 6, 2016 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the form even if they do not have affected product. The letter informs the customers that replacement product will be provided free of charge for any unused portion of the affected lot based on completion and return of the attached Customer Verification Form. Customers are instructed to return the completed Customer Verification within 10 business days of receiptby FAX at 949-598-7166, or by email at branan.complaints@brananmedical.com. Customers with anyquestions about the information contained in the notification are instructed to contact Branan Medical Corporation at 949-598-7166 or by e-mail atbranan.complaints@brananmedical.com. |  | 96 | 96 units | US Nationwide distribution in the states of: TN, NC, FL, and PA | Open, Classified |  |  | 96 | Not_Computer
Z-2307-2016 | 74731 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva; x-ray system. Product Codes: 708032, 708036 |  |  | July 19, 2016 | 2016 | Philips Electronics North America Corporation | The two bolts that attach the clamping plate to the lateral shaft may break, which could cause the c-arm to move laterally forward, slide out of its bearings, and sag | Device Design | On July 19, 2016, Philips sent a Field Safety Notice to all affected customers. Philips Field Service Engineers will add an independent safety catch free of charge. |  | 993 | 993 total | Worldwide distribution. US nationwide, Canada, Algeria, ARGENTINA, Australia, Austria, Belgium, China, Colombia, Czech Republic, Denmark, Egypt, Estonia, Faroe Islands, Finland, France, French Guiana, Germany, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Jersey, Jordan, Republic of Korea, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Philippines, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, TURKEY, Ukraine, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 993 | Not_Computer
Z-2308-2016 | 74568 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU480 Chemistry Analyzers, Reference No. B28624 |  |  | July 30, 2015 | 2015 | Beckman Coulter Inc. | Beckman coulter has received complaints regarding leaking and cracked C cups. When the cup cracks, coolant can leak onto the connectors of the Printed Circuit Boards (PCB) in the analyzer, resulting in a short-circuit causing the instrument to go into a stop mode. | Nonconforming Material/Component | In March 2011, the material and molding process of the Coolant cup (C cup) was changed. BEC did not issue a customer notification letter, as we do not believe that a notificationletter was either warranted or of any potential use or value to our customers, because the coolant cup ( C cup ) is an internal component that only a BEC Field Service Engineer (FSE) can access and inspect  i.e., the C cup is not in a customer-accessible part of theanalyzer. Thus, there are no specific actions the customer could take to inspect, identify, or mitigate the potential for a crack or leakage of the C cup. For questions regarding this recall call 714-961-4933. |  | 3355 | 3,355 units total (617 units in US) | Worldwide Distribution - US (nationwide (including Guam and Puerto Rico), and Internationally to Albania, Algeria, Australia, Austria, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Canada, China, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Liberia, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, United States Uruguay, Viet Nam, Zambia, Zimbabwe. | Open, Classified |  |  | 3355 | Not_Computer
Z-2200-2016 | 74424 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Temporary Pacing Lead System, Model 6416, Sterile EO. |  |  | June 16, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | This Pacing Lead System does not comply with Section 8.5.2.3 of IEC 60601-1, and corresponding provisions of FDA 21 CFR 898, which relate to design standards to prevent connecting a patients lead to a possible hazardous voltage. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated June 2016 to affected customers. The letter was addressed to Health Care Professional. The letter described the problem and the product involved in the recall. The letter described the "Customer Actions" which advised consignees to identify, quarantine and return affected product to Medtronic. Customers were instructed to complete the enclosed Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or scan and e-mail to RS.CFQFCA@medtronic.com. For questions, please contact your Medtronic CRHF Sales Representative. |  | 11697 | 11697 (5523 OUS, 6174 US) | Worldwide Distribution. US nationwide, Albania, Algeria, Australia, Botswana, Bulgaria, Canada, Colombia, France, Gabon, Germany, Guam, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, Kazakhstan, Macedonia, Mauritius, Netherlands, Netherlands Antilles, Norway, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Slovenia, South Africa, Spain, St Maarten, Sweden, Switzerland, Tajikistan, Turkey, Uganda, United Arab Emirates, United Kingdom, and US Virgin Islands. | Open, Classified |  |  | 11697 | Not_Computer
Z-2323-2016 | 74548 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | V-TEK TWIST D 3.5X90MM 30MM Can 12MM Rnd; The twist drill is used to drill holes into bones/tissues for foot, ankle and hand implant systems. |  |  | June 27, 2016 | 2016 | Zimmer Gmbh | Zimmer Biomet initiated a voluntary recall of the Twist Drill &amp; V-Tek Twist Drill due to to a potential tolerance stack up identified between the twist drill and the mating Kirschner Wire (K-Wire) component. | Under Investigation by firm | Zimmer GmbH initiated a voluntary recall of the Normed - TWIST DRILL Cannulated &amp; Normed- V-TEK TWIST DRILL Cannulated, with written notices on June 27, 2016 via certified mail due to a potential tolerance stack up identified between the twist drill and the mating Kirschner Wire (K-Wire) component. Customers are asked to do the following: Review the notification and ensure affected personnel are aware of the contents.2. Locate all affected product identified above and quarantine them immediately.3. Return any affected product within your possession. Clearly mark the outside of all return packages, Recall, and include a copy of the Inventory Return Certification form (Attachment 1) with your returnshipment(s).4. If after reviewing this notification you have further questions or concerns please call the customercall center at 1-800-348-2759 between 8:00 am and 8:00 pm EST.Adverse reactions or quality problems experienced with the use of this product may be reported to theFDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax." Online:www.fda.gov/medwatch/report.htm" Regular Mail: use postage-paid, pre-addressed Form FDA 3500 available at: www.fda.gov/MedWatch/getforms.htm.Mail to address on the pre-addressed form." Fax: 1-800-FDA-0178Affected units returned as part of this recall activity will be scrapped |  | 169 | 169 | Domestic: CO, ILForeign: Argentina, Costa Rica, India, Malta, Japan, Philippines, Vietnam, Portugal, Dominican Republic, Argentina, Australia, Russia, Saudi Arabia, Ireland, Ecuador, China, Singapore, Indonesia, Canada, Mexico, South Korea , Spain, Netherlands, Sweden, Malaysia, New Zealand, Switzerland, El Salvador, England, Norway, Finland, Venezuela, Taiwan, Italy, Colombia, Greece, Thailand, France, Israel, Jamaica, Japan, Jordan, Germany, Korea, Lebanon, Belgium, Hong Kong, Philippines, Iran, Brazil, Turkey, United Arab Emirates (UAE), Uruguay | Terminated | May 30, 2017 | 337 | 169 | Not_Computer
Z-2316-2016 | 74715 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Synergy Spine. The software application is sent in CD format with an IFU, wrapped in plastic with a label for shipping purposes. |  |  | July 18, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. announces a voluntary field action for the Medtronic Navigation StealthStation Software applications affected by Neurologica BodyTom/CereTom floor-based scanners. | Software design | Medtronic sent an Urgent Medical Device Safety Alert to all affected customers on August 20, 2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached confirmation form and return to RS.NavFCA@medtronic.com or fax it to Medtronic Technical Service at 651-367-7075. For questions contact Neurologica Technical Support at 1-888-564-8561 (US/Canada) or +1-978-564-8561 (International) or email support@neurologica.com |  | 27 | 27 in total | Worldwide Distribution - US Nationwide in the state of: NY, AZ, AR, NC, CA, OH, MI, IN, MA, TX, OH, OR, CO, MO, UT, IL. and the countries of: India, Australia, and Russian Federation. | Open, Classified |  |  | 27 | Computer
Z-2320-2016 | 74624 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | RePlantÂż Angled Abutment |  |  | July 05, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | Issue with design specification; oversized condition. | Process control | Implant Direct sent an Urgent: Medical Device Field Corrective Action letter dated July 6, 2016, to all affected customers. The letter requests that it be reviewed and determined if customer has any affected product in inventory. Customers are asked to complete and return the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized Implant Direct Sybron Manufacturing distributor, they request that customers identify those customers that may have been shipped the affected product lot and contact these customers to inform them of this issue within fortyeight (48) hours of receipt of the notification in order to provide the customers with the correct tool.If customers have any of the affected product listed the firm will send a replacement part. If any questions contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425. |  | 205 | 205 | U.S. distribution to the following; CA, WA, UT, GA, AZ, NE, OR, MI, TX, WI, FL, AL, AR, MO, VA, MA, TN, LA, NC, KS, NY, SC, OK.Foreign distribution to the following; CA, DE, IT, GB, PL, DE, GB, NL, GR, ES, HR, BG. | Open, Classified |  |  | 205 | Not_Computer
Z-2325-2016 | 74691 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | BK Medical ApS Ultrasound System Scanner bk2300 - |  |  | June 28, 2016 | 2016 | B-K Medical A/S | In the system (bk3500) software generic volume measurement (HxWxLxl) function the user may beunaware of the calculation factor (1) and assume the generic volume calculation factor is similar to that ofthe other (prostate, adenoma, testis, kidney, and bladder) volume calculation factors provided in thesoftware. | Software design | BK Medical sent an Important Corrective Action Notification letter dated June 27, 2016, to all affected customers. The letter explains the issue and asks customers to discontinue the use of the generic volume measurement and contact BK medical ApS Service for assistance. Customers with questions were instructed to call 800-645-4442. |  | 6 | 6 | Worldwide Distribution - US to TX, FL, and MA. Internationally to Australia | Open, Classified |  |  | 6 | Computer
Z-2329-2016 | 74599 | No | MGB |  |  | Cardiovascular | PMA | 6F/7F MynxGrip Vascular Closure Device; |  |  | July 13, 2016 | 2016 | AccessClosure, Inc., A Cardinal Health Company | A total of 70 6F/7F Mynx Grip devices were shipped from a distributor using improper transport/packaging configurations. | Process control | Cardinal Health sent an Urgent Medical Device letters dated July 15, 2016 to affected customers. The letters provided the reason for recall and the exact lot number and quantity shipped to each location. Customers were instructed to quarantined affected products and complete and return the Field Action response form indicating the number of devices that will be returned. Instructions for returning the devices are included with an RMA number and Fed Ex airbill. Questions should be directed to Customer Advocacy at 408-610-6570 between 9 am and 5 pm PST. |  | 55 | 55 units | US in the states of CA, and MS. | Terminated | August 15, 2016 | 33 | 55 | Not_Computer
Z-2333-2016 | 74610 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SureStep Tray, with Add-a-Foley, COMPLETE CARE, Drainage Bag, and STATLOCK Stabilization |  |  | June 29, 2016 | 2016 | C.R. Bard, Inc. | Labeling problem | Labeling mix-ups | Bard Medical sent an Urgent Medical Device Product Recall letter dated June 29, 2016. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers examine their inventory and quarantine the product. The letter also requested a sub-recall. The letter included a reply form which is to be returned to Bard. Customers were also instructed if product was further distributed they identify customers and notify them at once. Customers with questions should contact the BMD Recall Coordinator at 1-800-793-8110. |  | 530 | 530 units | Nationwide Distribution to CA, FL, IL, IA, KY, NC, MN, NC, OR, RI, SD, TX, UT, and WA. | Open, Classified |  |  | 530 | Not_Computer
Z-2327-2016 | 74669 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Horizon 24 Large Titanium Clips with Adhesive Backing, Ref 004204, Rx Only, Sterile. |  |  | August 08, 2013 | 2013 | Teleflex Medical | The manufacturing and expiration dates are not printed on the blister pack label. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification dated August 8, 2013, to all affected customers. The letter requested that they quarantine the product and to return it to Teleflex. An acknowledgment form was included with the letter which also was to be returned to Teleflex. A separate letter was sent to distributors requesting they conduct a sub-recall.If you have any further questions, please call (610) 378-0131. |  | 8220 | 8,220 ea. | USA Distribution to the states of :AL, AZ, CA, CT, FL, IL, KY, MD, MI, MN, NM, NY, NC, OR, PA, SC, TX, VA, WA, WI, and Internationally to Canada | Completed |  |  | 8220 | Not_Computer
Z-2328-2016 | 74751 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Patient Monitor |  |  | July 26, 2016 | 2016 | Philips Electronics North America Corporation | If an affected Patient Monitor has been powered on continuously for several months, any displayed waveforms will contain outdated data and therefore fail to reflect the patients current condition. | Nonconforming Material/Component | Philips sent an Urgent Medical Device Correction letter dated July 20, 2016 to all affected customers. The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Customers were informed that Philips will provide a software correction free of charge. For any further information or support concerning this issue, contact your local Philips representative at 1-800-722-9377. |  | 50495 | 50,495 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 50495 | Computer
Z-2332-2016 | 74560 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Spirit Select Bed, A-C Powered Hospital Bed |  |  | June 28, 2016 | 2016 | CHG Hospital Beds Inc | Stryker Medical is initiating a voluntary recall of the Spirit Select and Spirit Plus A-C powered hospital beds due to reports of hi-lo actuators broken at the mount ends, which could cause the bed to unexpectedly lower resulting in patient injury. | Component design/selection | The firm, Stryker Medical, issued an "URGENT MEDICAL DEVICE RECALL" letter dated 7/8/2016 by FedEx to its customers and included a revised preventive maintenance checklist. The letter described the product, problem and actions to be taken. The customers were instructed to: locate the units listed on the attached business reply form; remove these units from service, if not possible to remove units from service, place units in the height position according to instructions; file the revised preventative maintenance checklist; return the Business Reply Form to confirm receipt of the notification by fax (269)488-8691 or email productfieldaction@stryker.com. If customers have loaned or sold any of the beds listed in this letter, please forward a copy of the recall notice to the new users and advise us of their new location in the space provided on the business reply form. Your Stryker Field service Representative will contact your facility to add support brackets to your beds. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax."Online: www.fda.gov/medwatch/report.htm "Regular Mail: use postage-paid FDA form 3500 available at: www.fda.gov/MedWatch/getforms.htm. Mail to MedWatch , P.O. Box 3002, Rockville, MD 20847-3002 "Fax: 1-800-FDA-0178 If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 4308 | 4,308 units total | Worldwide Distribution: US (nationwide) including states of: AK , AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, and WI; and country of: Canada | Terminated | June 08, 2017 | 345 | 4308 | Not_Computer
Z-2330-2016 | 74608 | [6 empty] | EnVision FLEX+, Mouse, High pH (Link), Product No. K8002, K8004. |  |  | July 06, 2016 | 2016 | Dako Denmark A/S | The pH of Target Retrieval Solution in the kit measures 4-5 instead of the correct pH of 9.1. | Process control | A recall notification letter dated 7/6/16 was sent to customers to inform them that Dako is recalling EnVision FLEX+ kit because the pH of Target Retrieval Solution in the kit measures 4-5 instead of the correct pH of 9.1. The letter informs the customers that with a pH of 4-5 instead of pH 9 there could be weak staining, false negative or false positive non-specific staining depending on the antibody used. There is the possibility that this staining could lead to a false result. Customers are to discard the affected product, determine the impact as instructed, and confirmed that they have received replacement kits. Customers are instructed to complete the Device Recall Form and contact local sales representative for any questions regarding the recall notification. |  | 147 | 147 kits | Worldwide Distribution -- US, Canada, and Colombia. | Open, Classified |  |  | 147 | Not_Computer
Z-2326-2016 | 74565 | No | FTR | Silicone gel-filled breast prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Mentor MemoryGel Breast Implant; Indicated for females for breast augmentation and breast reconstruction. |  |  | July 01, 2016 | 2016 | Mentor Texas, LP. | The box of 300cc MemoryGel Breast Implant was labeled with null manufacturing and expiration dates. | Labeling mix-ups | The affected consignees were notified via customer letter on 7/1/16, and a follow-up with respective sales representative via visit. |  | 35 | 35 units | Distributed in the states of NY, WI, CA, TX, CO, NC, NJ, OH, TN, RI, and IL. | Open, Classified |  |  | 35 | Not_Computer
Z-2341-2016 | 74577 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge HEMO software. |  |  | July 02, 2012 | 2012 | Merge Healthcare, Inc. | The International Normalized Ratio (INR) value displayed in the study report does not match the value that is imported on the pre-procedure labs screen from HL7 | Software design | The firm issued an "Important Product Information" bulletin to customers dated 7/2/2012 via e-mail letting them know there was an upgrade available. On 1/30/2016, the recalling firm initiated phone calls to their customers letting them know there was an upgrade available and if they declined it, they would like a written response. |  | 37 | 37 | The medical facilities who currently have the affected versions are located nationwide. There was government distribution but no military or foreign distribution. | Terminated | June 26, 2017 | 1820 | 37 | Not_Computer
Z-2343-2016 | 74549 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | 91496 Ultraview SL Command Module, Options A, B, C, and I. |  |  | July 01, 2016 | 2016 | Spacelabs Healthcare Inc | The firm has received multiple reports of the Non-Invasive Blood Pressure (NIBP) parameter becoming non-functional with an associated message (No Reading). In addition, there is one of the following three error messages: Inflate Error, HW Error, or No Data. | Vendor change control | Spacelabs expanded the recall and sent an Urgent - Medical Device Correction letter dated August 30, 2016, to additional affected customers via priority mail (return receipt requested). On September 6, 2016, a customer letter (translated as necessary) was sent via email to all international subsidiaries and distributors of record.**********************************************************************************************Spacelabs Healthcare Inc, sent an Urgent - Medical Device Correction letter dated July 1, 2016, to all affected customers (return receipt requested). On July 8, 2016, a customer letter (translated as necessary) was sent via email to all international subsidiaries and distributors of record.Spacelabs Healthcare will contact customers to schedule a convenient time for a Spacelabs service technician to update their affected Command Module(s) at no cost.For additional information or technical assistance, please contact:Technical Support 1-800-522-7025 and select (2) for Technical Support |  | 5580 | 3801 modules and 116 PCBAs &amp; valves distributed in the US; 1297 modules and 366 PCBAs &amp; valves distributed outside the US | Worldwide Distribution - USA Distribution and to the countries of :AFGHANISTAN, ARGENTINA, AUSTRALIA, BAHRAIN, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM; CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DOMINICAN REPUBLIC; ECUADOR, EGYPT, ESTONIA, FRANCE, GERMANY, GUATEMALA, HONG KONG, INDIA, IRAN (ISLAMIC REPUBLIC OF), ISRAEL, ITALY, JORDAN, KUWAIT, MEXICO, MOROCCO, NETHERLANDS, NICARAGUA, PAKISTAN, PALESTINE, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, Slovakia, SPAIN, SRI LANKA, SWITZERLAND, TAIWAN, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VIET NAM. | Terminated | May 22, 2017 | 325 | 5580 | Not_Computer
Z-2215-2016 | 74469 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000, Disposable Pediatric Patient Circuit, Catalog Number: 5704964 |  |  | May 31, 2016 | 2016 | Draeger Medical, Inc. | Isolated episodes of leakage at the check valve were found | Device Design | Drager sent an Urgent Medical Device Recall letter dated May 31, 2016 and Customer Reply and Order Card to all customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to inspect stock and dispose of affected products. For questions regarding this letter call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger device/accessory contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 42922 | 42,922 units worldwide (US distribution 1,530) | US Nationwide Distribution including states of: AK, CA, FL, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NY,OH,OK,OR,PA,TN,TX,UT,VA,WA, and WI. | Open, Classified |  |  | 42922 | Not_Computer
Z-2337-2016 | 74636 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Air Dermatome II Width Plate, 2 inch. Part # 885120200. Width plate is packaged inside of a padded-foam corrugate box. A product label is applied to the lid of the box. The width plates, when sold individually, are placed in a labeled poly zipper bag, which is then placed into a corrugated shipper carton or padded envelope. |  |  | July 04, 2016 | 2016 | Zimmer Surgical Inc | The surface coating applied to the device is blistering, peeling and discoloring after usage over time | Process design | The firm, Zimmer Biomet Surgical, sent an "URGENT MEDICAL DEVICE RECALL CORRECTION NOTICE" dated 7/6/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:1. Carefully review this letter and ensure all users of the Zimmer Air Dermatome II at your facility have been informed of this notice.2. Inspect the Zimmer Air Dermatome II devices within your facility and indicate on the attached Certificate of Acknowledgment the condition of the device. Complete and sign the attached Certificate of Acknowledgment and return it to corporatequality.postmarket@zimmerbiomet.com.3. Before each use, continue to examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are notpresent, you may continue to use. Immediately after each use, re-examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are present, discontinue use. In the event that alternate devices or therapies are not available, surgeons may, at their discretion, determine the use of the device as a medical necessity and continue to use. After use, follow your facilitys wound irrigation protocol or Zimmer Biomet recommends copious irrigation with a normal saline solution or immersion of the graft in a bowl of normal saline. Zimmer Biomet recommends irrigating the donor site with saline as well. Carefully examine the skin graft and donor site for gray coating flakes and continue to irrigate as necessary. Report this as an adverse event to Zimmer Biomet Surgical at SurgicalRegulatoryReporting@zimmerbiomet.com.If after reviewing this notification you have further questions or concerns please call 1-330-364-0989 between 8:00 am and 5:00 pm EST, Monday through Friday. |  | 90 | 90 units | US Nationwide Distribution including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, and WI. | Terminated | June 21, 2017 | 352 | 1270 | Not_Computer
Z-2346-2016 | 74402 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice whole body computed tomography X-ray system. Imaging diagnostic tool. |  |  | April 20, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The following issues have been found in MX 16-slice systems with software version 1.1.4.21426:1. During the filming operation on MX16-slice console software, the clipboard used for copying and pasting images is not cleared between patients. If the operator fails to copy the current patients images before pasting, a previous patients images may be present in the clipboard and be copied onto the film of the current patient. 2. During the Bolus Tracking scan, if the Auto Voice in Tracker scan is enabled, the Tracker scan will be aborted unexpectedly and the diagnostic scan after the tracker scan needs to be manually started.3. If the scan protocol with SAS (Spiral Auto Start) function is selected to plan the scan, the SAS option may not be displayed on the Contrast tab of the scan protocol parameter area or the SAS option is displayed but not enabled as pre-configured in the protocol. This issue only occurs on the first helical scan after system startup that applies SAS function | Software design | On April 20, 2016 the firm sent Field Safety Notices to their customers. Philips will release a Field Change Order to upgrade the software free of charge. A Philips Field Engineer will contact their customers to schedule this correction. |  | 899 | 899 units | Products were sold in the following states: AL, AR, CA, FL, GA, IL, IN, MS, NC, ND, NJ, NV, NY, OH, PA, SC, TN &amp; VA.Products were also shipped to the following countries: Albania, ALGERIA, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Burundi, Cape Verde, Chile, China, Colombia, Congo, Democratic, Costa Rica, Cyprus, Dominican Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Hungary, India, INDONESIA, Ireland, Italy, Japan, Jordan, Latvia, Lebanon, Libyan Arab Jamahiri, Lithuania, Malawi, Mali, Mexico, Morocco, Myanmar, Namibia, Nepal, Netherlands, Nigeria, Oman, Palestine, Paraguay, Philippines, Poland, Romania, Russia, Senegal, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, United Kingdom, Venezuela, Viet Nam &amp; Zambia. | Open, Classified |  |  | 979 | Computer
Z-2344-2016 | 74655 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego, Artis Q/ Q.zen, stand alone system, software controlled |  |  | July 07, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to an error in the 19 Live Display, image reproduction may fail in the examination room and the potential exists for the loss of images immediately after system startup. | Nonconforming Material/Component | Siemens sent an Important Customer Safety Notice dated July 7, 2016, by certified mail to customers to inform them of the corrective action they plan to take to prevent potential loss of the image display for live images in the examination room after system startup. Siemens service organization will contact each customer to arrange a date to resolve the issue with a software update of the monitor firmware via AX012/16/S. For questions regarding this recall call 610-448-6461. |  | 51 | 51 units | Nationwide Distribution to AL,AZ,CA,CO,IA,IL,MA,MD,MI,MN,MO,NE,NJ,NY,OH,OR,PA,RI,TX,VA,WA,WI,WV | Open, Classified |  |  | 51 | Computer
Z-2352-2016 | 74764 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Estradiol Reagent Pack, Unique Device Identifier No. 10758750005017, REF/Catalog Number/Product Code 855 2630; IVD -- Ortho-Clinical Diagnostics, Pencoed, Bridgend CF35 5PZ, UK --- The VITROS Estradiol test is performed using the VITROS Estradiol Reagent Pack and Calibrators on VITROS ECi/ECiQ Immunodiagnostic Systems, VITROS 3600 Immunodiagnostic Systems and VITROS 5600 Integrated Systems. |  |  | July 14, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | Ortho Clinical Diagnostics has issued a recall of their VITROS immunodiagnostics Estradiol Reagent assay. The firm was made aware of a potential cross-reactivity between the drug Fulvestrant and theVITROS Immunodiagnostic Estradiol Reagent assay. The firm conducted an internal investigation and was able to confirm biased positive (ranging from 1303-4046%) on Estradiol results, obtained from postmenopausal females taking Fulvestrant. | Device Design | On 7/14/2016, URGENT PRODUCT CORRECTION NOTIFICATION Customer Letter (Ref. CL2016-141, dated 7/14/2016) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received product since 6/1/2015 to inform them that Fulvestrant is an interferent of the Estradiol Reagent Pack and causes positively biased results. Discuss any concerns you may have regarding previously reported estradiol results with your Laboratory Medical Director to determine the appropriate course of action. The results from this or any other diagnostic test should be used and interpreted only within the context of the overall clinical picture. For questions, contact Customer Technical Services at 1-800-421-3311. Foreign affiliates were informed of the issue via e-mail on 7/14/2016 and instructed to notify their consignees of the issue. |  | 7597 | Domestic: 1620; Foreign: 7597 | Nationwide, Puerto Rico, and foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Poland, Singapore, Spain, Venezuela. | Open, Classified |  |  | 7597 | Not_Computer
Z-2350-2016 | 74705 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | CURVED-TIP STAPLER 30,IS4000, model number 470530; |  |  | July 14, 2016 | 2016 | Intuitive Surgical, Inc. | An increase in the number of da Vinci EndoWrist Stapler complaints where the customer experienced aStapler Firing Failed message as well as review of da Vinci Xi Surgical System logs found an increased trend | Software design | Field Safety Notifications - Urgent Medical Device Corrections and Removal letters were sent on July 15, 2016 (dated July 14, 2016 advised users of the Stapler Firing Failed message and actions to take should the message be received. |  | 2065 | 2065 total | US, including Puerto Rico and Australia, Austria, Belgium, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Japan, Monaco, Netherlands, Norway, Portugal, , Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Venezuela. | Terminated | January 03, 2017 | 173 | 2065 | Not_Computer
Z-2345-2016 | 74609 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Medivance ArcticGel Neonatal Pad |  |  | June 30, 2016 | 2016 | C.R. Bard, Inc. | An internal review of an available product found that the Information For Use only contained English and Japanese languages when a CE-marked 23 language version was expected. | Other | A Customer Communication Correction Letter was sent to customers on June 30, 2016. The letter requested that they examine inventory and remove the original IFU and substitute a copy of the IFU that was included with the Correction letter. The letter also requested a sub-recall, and included a reply form which was to be completed and returned. If you have any further questions please call (401) 825-8690. |  | 64 | 64 | US Distribution to the states of : CA, FL, GA, IL, IN, MI, MN, MO, NY, OR, PA, TX, WA, and Internationally to Germany, Italy and Switzerland | Terminated | June 06, 2017 | 341 | 64 | Computer
Z-2351-2016 | 74557 | formalin, neutral buffered | IFP | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Richard Allan Scientific 10% Neutral Buffered Formalin, |  |  | June 13, 2016 | 2016 | Richard-Allan Scientific Company | potential contamination with potassium hydroxide may have an impact when performing immunohistochemistry testing. It is unknownwhether the contamination will affect tissue antigenicity. If it were to affect tissue antigenicity, theoutcome would be risk of false positive or false negative test results which may lead to seriousinjury. There is the possibility the patient may need to be re-biopsied. | Nonconforming Material/Component | On 6/13/2016, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected customers via certified mail with instructions for identifying and disposing of the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 18 | 18 cases | Distributed in the states of FL, PA, NY, MD, CA, and AZ. | Open, Classified |  |  | 18 | Not_Computer
Z-2353-2016 | 73255 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | S3 MedSurg Bed, Model 3005.Stryker Medical |  |  | February 02, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | CPU board failures cause fowler (backrest) electronic controls to stop working. | Component change control | Stryker Medical initiated a voluntary recall of the S3 MedSurg Bed, Model 3005, with notices sent on February 2, 2016 via certified mail. Stryker medical identified that Fowler does not raise up to desired angle; and the Fowler does not raise electronically, but manual function remains operational through the manual CPR release. Actions Needed for Customers:1. Locate the bed(s) listed by serial number on the attached business reply form.2. A Stryker ProCare Representative will contact your facility to schedule a time for repair.3. Fax (269)488-8691 or email productfieldaction@stryker.com the enclosed business reply form to confirm receipt of this notification.4. If you have further distributed the beds listed in this letter, please forward a copy of this notice to the new users and advise us of their new location. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Normal business hours are Monday-Friday 8 a.m-5 p.m. (ET). |  | 8490 | 8490 | Distribution US nationwide, Canada, and Mexico. | Terminated | June 12, 2017 | 496 | 8490 | Not_Computer
Z-2365-2016 | 74581 | No | NFB | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Invacare Precise Rx Pediatric Flowmeter Model Number: IRCPF16 AW. Packed one per carton. |  |  | June 21, 2016 | 2016 | Invacare Corporation | The rear gradation scale was screen printed upside down on the flowmeter gauge. | Process control | lnvacare Corporation sent an Urgent Medical Device Recall letter dated June 27, 2016, to all affected consignees with return receipt to notify them of the recall. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees with questions were instructed to call 877-413-6008. |  | 395 | 395 units | Worldwide Distribution - US to CA, CO, CT, FL, GA, IA, ID, IL,IN,KS,KY,LA,MA,MD,MI,MS,MT,NC,NE,NJ,NM,NY,OH,PA,SC,TN,TX,UT,VA,WA and Internationally to Canada, Germany and New Zealand. | Open, Classified |  |  | 395 | Not_Computer
Z-2363-2016 | 74776 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Inc., Hartland, WI. |  |  | February 05, 2016 | 2016 | Merge Healthcare, Inc. | When measuring a lesion on an unmagnified mammography image then performing the same measurement on an image magnified by the Hologic imager (different image), the measurements are not the same. | Software design | The recalling firm's customer support notified customers via phone call beginning 2/5/2016 to inform them of the issue. Customers were informed that a software fix that corrects the issue was available. Customers were asked not to use the measurements on mammo images until upgraded. If customers want an upgrade, then someone would contact them to schedule the upgrade. |  | 49 | 49 sites potentially have the affected versions | Worldwide Distribution -- USA and Australia. | Open, Classified |  |  | 49 | Computer
Z-2375-2016 | 74782 | No | DZI | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Instructions for Use (IFU) and Manual NobPrl CC Guided Surgery Kit |  |  | July 20, 2016 | 2016 | Nobel Biocare Usa Llc | Incorrect drill protocol in the Instructions for Use (IFU) | Labeling design | The firm, Nobel Biocare, sent "IMPORTANT-QUALITY MESSAGE" Field Correction Letters and response forms dated July 22, 2016. The letter describes the product, problem and actions to be taken and informs the customers that the titles of the IFU and the Manual are "NobelGuide for NobelParallel Conical Connection" (ID number 2002) and "Guided Surgery with Nobel Parallel Conical Connection" (ID number 79423). The Customers are instructed to review if you have a printed version of the IFU or the Manual related to the article and discard them; and to complete the attached "Quality Message Acknowledgement" form and fax to 714-998-9348 or email to: joann.radovich@nobelbiocare.com. Customers are informed that an English version of the corrected IFU is attached to the letter and translated versions will be available on the Nobel Biocare website for electronic IFU's http://www.nobelbiocare.com/IFU within 30 days. If you have further questions, contact our customer service department at 714-322-5001. |  | 231 | 231 kits | Worldwide Distribution: US (nationwide) and Countries of: Austria, Belgium, Canada, Germany, Denmark, Spain, France, Great Britain, Italy, Netherlands, Norway, Poland, Portugal, Russia, and Turkey. | Open, Classified |  |  | 231 | Not_Computer
Z-2342-2016 | 74761 | [6 empty] | NucliSENS easyMAG Magnetic Silica |  |  | July 08, 2016 | 2016 | BioMerieux SA | Several customer complaints about amplification performance issues when using different batches of MagSil with the NucliSENS easyMAG extraction platform. | Under Investigation by firm | Consignees were notified of the recall by letter on July 18, 2016. The letter informed them of what actions need to be taken if they experience problems. The letter included a reply form to be returned. Subsidiaries and distributors were notified electronically on 7/8/2016. |  | 7845 | 7,845 units | Worldwide Distribution - US Nationwide in the states of: AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MO, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, and the countries of Angola, Argentina, Austria, Australia, Austria, Botswana, Brazil, Burkina Faso, Cambodia, China, Columbia, Djibouti Antilles, Estonia, Georgia, Germany, Greece, India, Israel, Italy, Kuwait, Latvia, Malaysia, Mexico, Netherlands, Poland, Qatar, Russia, Senegal, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Zambia, and Zimbabwe | Open, Classified |  |  | 7845 | Not_Computer
Z-2377-2016 | 74634 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Modular Handle Triathlon Instruments, Non-Sterile, REF 6541-4-808, Lot 123, NDC M546654148080S |  |  | June 23, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedic is voluntarily recalling the Triathlon Modular Handle because analysis revealed that the press fit specifications between the dowel pin and the mating hole in the shaft were not being achieved potentially resulting in disassociation from the instrument. | Device Design | Stryker sent an Urgent Product Recall letter dated June 24, 2016 to Branches/Agencies by e-mail and Business Reply Forms was sent to Branches/Agents/Hospital Risk Management via UPS (with return receipt) on June 24, 2016. The letter identified the affected product, problem and actions to be taken. For questions call (201) 831-6693. |  | 12469 | 12,469 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 12469 | Not_Computer
Z-2373-2016 | 74622 | disk, abrasive | EHJ | Abrasive device and accessories. | Dental | Dental | 510(K) Exempt | OptraPol Polishing Discs, Rx ONLY: OptraPol Refill Cup /10: REF/Product ID 634459, OptraPol Refill Small Flame/10: REF/Product ID 634458, OptraPol Assortment: REF/Product ID 634464 --Device Listing # B106952. One step polishing system for all common dental composites and amalgam. |  |  | July 11, 2016 | 2016 | Ivoclar Vivadent, Inc. | After a few seconds of polishing with moderate force, parts of the OptraPol cup are breaking away. OptraPol Small Flame is also affected. Two complaints were filed referring to OptraPol, claiming that the abrasive polishers wear faster than expected. | Under Investigation by firm | URGENT MEDICAL DEVICE RECALL Letter (dated 7/11/2016) and Response Form were sent to customers on 7/11/2016 via Certified Mail, Return Receipt, Requested. The letter instructs customers to notify their customers of the recall if the product had been distributed and to return the product to Ivoclar Vivadent. For questions, please contact Ivoclar Vivadent Customer Service at 1-800-533-6825. |  | 86 | US: 86 units total | Distributed in the states of CA, FL, GA, IA, IL, IN, KS, KY, MA, MI, NV, NY, OH, PA, TX, and WA. | Open, Classified |  |  | 86 | Not_Computer
Z-2378-2016 | 74788 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Tritest CD3 FITC/CD4 PE/CD45 PerCP (IVD); |  |  | July 18, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | Internal review found that some vials of CD3/CD4/CD45 state that the product contains gelatin. The product does not contain gelatin and should state in buffer. | Labeling Change Control | BD Life Sciences sent an Urgent Product notification letter dated July 18, 2016 to all affected customers via UPS and/or e-mail. The letter informed the user of the issue and states that no action is required since product performance is not impacted. The response form is to be returned to BD as instructed. Customers requiring further assistance were instructed to contact BD Customer Support at 855-236-2772.For questions regarding this recall call 408-954-6080. |  | 14089 | 14,089 | Nationwide distribution to AR, CA, CO, IL, IN, LA, MA, MD, Ml, MN, NC, NJ, NM, NY, OR, PA, TN, TX, WA | Open, Classified |  |  | 14089 | Not_Computer
Z-2374-2016 | 74681 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | Servo Humidifier 163, Model 01-06-8125-8, REF 6419365 |  |  | December 03, 2015 | 2015 | Teleflex Medical | Cracks may occur at the 15mm and 22mm ISO female connector. | Release of Material/Component prior to receiving test results | Inmed, a division of Teleflex sent an Urgent Field Safety Notice letter dated December 3, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check inventory, cease use and distribution of stock of the affected product batch and quarantine immediately.. The letter included an acknowledgment form which was to be completed and returned your MAQUET representative.For further information or support concerning this issue, contact your local MAQUET representative at fieldactions@maquet.com. For questions contact your local sales representative or Customer Service. |  | 2700 | 2,700 ea. | US Nationwide Distribution in the states of FL, KS and WV. | Terminated | April 14, 2017 | 498 | 2700 | Not_Computer
Z-2376-2016 | 74615 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Volker 2080, 2082, S 280 and S 282 series beds with lateral release button. |  |  | June 14, 2016 | 2016 | Hill-Rom, Inc. | The current construction of the side rail latch may not be able to support the full patient weight and may give way allowing the side rail to abruptly lower and the patient to fall, potentially resulting in harm. To date there has been one report of a related injury. | Device Design | Hill-Rom sent an URGENT: MEDICAL DEVICE CORRECTION letter dated June 14, 2016 to the affected US and Canadian customers via email and courier, international customers. The letter included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The letter requests the number of beds in the consignees possession so that repair kits can be provided. If you have any questions regarding this safety notice, contact Hill-Rom Technical Support at 1-800-445-3720, your distributor, your Hill-Rom representative or contact us on www.hill-rom.com. |  | 24262 | 24262 | Worldwide Distribution - US Nationwide and the countries of: Canada, AUSTRALIA, CHINA, HONG KONG, THAILAND, POLAND, RUSSIAN FEDERATION, SLOVENIA, UNITED ARAB EMIRATES, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, ICELAND, IRELAND, ISRAEL, ITALY, LUXEMBOURG, NETHERLANDS, NORWAY,SPAIN, SWITZERLAND, and UNITED KINGDOM. | Terminated | May 15, 2017 | 335 | 24262 | Not_Computer
Z-2371-2016 | 73750 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Computed Tomography X-ray systems is intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. |  |  | March 24, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software defect causing intermittently slow response of Host. | Software change control | On 3/24/2016 the firm mailed Customer Information letters to their customers. |  | 29 | 29 units | Product was sent to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY.Product was also sent to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Cuba, Denmark, Djibouti, Egypt, Estonia, Finland, France, Germany, Ghana, Gibraltar, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 635 | Computer
Z-2380-2016 | 74131 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 CT Model 728231; To produce cross-sectional images of the body. |  |  | April 29, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software issues found in software versions v4.1.3/4.1.4/4.1.5 in the Philips Brilliance 64 products that could affect the performance of the equipment. | Software design | The customers will be sent a letter with information about, what the problems are and under what circumstances they may occur, the actions customers can take to minimize the effect of the problem and the actions planned by Philips to correct the problem. To correct these issues, a Philips field service representative will install a software update on the affected systems. Reference field change order (FCO) 72800652. |  | 150 | 150 | Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WV &amp; WY.Product was also shipped to the following countries: Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, Colombia, Cuba, Estonia, Finland, France, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, RÂżunion, Russian Federation, Senegal, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 1403 | Computer
Z-2379-2016 | 74773 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. Product was distributed under the AMICAS label, Brighton, MA, but labeling was later changed to Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | The patient name in the Halo title bar and the thumbnails do not match the name on displayed images. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016 explaining the reason for recall and instructing the customer to provide a copy of the notification to all users. A second letter was issued to nonresponders dated 5/25/2016 |  | 93 | 93 sites have the affected version | The software was distributed to medical facilities nationwide. Government distribution was made to MN. Foreign distribution was made to Australia and to the United Kingdom. There was no military distribution. | Open, Classified |  |  | 93 | Not_Computer
Z-2391-2016 | 74553 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Coolife Sinergy Radiofrequency Kit, SIK-17-75-4 (109971504); Kit used in conjunction with a Radiofrequency Generator to create radiofrequency lesions in nervous tissue. |  |  | June 16, 2016 | 2016 | Halyard Health, Inc | Mismatch between the length of the RF electrode (probe) and the cannula (introducer), | Process control | All distributors and healthcare facilities that were directly shipped recalled devices were notified of the recall beginning 6/16/2016 by email and by FedEx. They were instructed to examine their stock and to quarantine and discontinue use of the product. The letter contained a response form which was to faxed back to Halyard. |  | 51 | 51 kits | Distributed in the states of CT, IL, MA, MI, NV, NH, NC, and TX, and the country of Canada. | Open, Classified |  |  | 51 | Not_Computer
Z-2389-2016 | 74719 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4mm Titanium Locking Screws Self-Tapping with Stardrive Recess 12mm, Synthes 2.4mm Titanium Locking Screws Self-Tapping with Stardrive Recess 12mm, Plate, Fixation, Bone, Orthopedic Devices |  |  | June 28, 2016 | 2016 | Synthes (USA) Products LLC | It was reported that the specified lots of the 2.4mm Titanium Locking Screw Self-Tapping with Stardrive Recess 12mm can break between the head and the shaft. This resulted from the neck feature being too close to the StarDrive feature leaving a thin wall between neck and drive. | Under Investigation by firm | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated July 14, 2016 to affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove and return all unpackaged Locking Screws Self-Tapping with Stardrive Recess 12mm in any affected kits along with any screws still in the original packaging with the part number and lot number identified in the recall notification. Customers were also instructed to call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verification Section and indicate the number of devices found and return with product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132 by fax to 877-523-9109 or scan/email: Synthes5246@stericycle.com, even if you do not have the product. Note: if the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 41 | 41 units | US Distribution to states of: AR, CA, CO, DE, FL, GA, IA, IL, KY, LA, MN, NJ, OK, PA, SD, TN, and WA. | Open, Classified |  |  | 41 | Not_Computer
Z-2386-2016 | 74757 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Integra External Fixation System Universal Wire Fixation Bolt - 17.5 mm Catalogue #12224218 ; Integra External Fixation System Universal Wire Fixation Bolt - 23 mm Catalogue #12224219 |  |  | July 20, 2016 | 2016 | Integra LifeSciences Corp. | Through the investigation of reported complaints Integra verified that there have been instances where the Universal Wire Fixation Bolts the Slotted Post broke during use at the surgical site. | Device Design | Integra sent an Urgent - Voluntary Medical Device Recall letter dated July 20, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Integra is asking you to take the following actions:1. Please check your IntegraÂż External Fixation System inventory to determine if you have the 17.5mm or 23mm UniversalWire Fixation Bolts and/or the large or small Slotted Posts.2. If you do have 17.5mm or 23mm Universal Wire Fixation Bolts and/or large or small Slotted Posts, stop using them immediately.3. Complete the attached "Acknowledgement and Return Form" and check the box: I do have affected products on the list and record the lot number.4. Or, complete the attached "Acknowledgement and Return Form" and check the box: I do not have affected products.5. Complete the other information as indicated on this form. Keep a copy of the form for your records.Return the completed "Acknowledgement and Return Form" by email or fax indicating your receipt and review of this notification.When your form is received, if you have noted you have affected products, an Integra Representative will contact you and provide you with directions to return the product, as well as input an order to replace the quantity you indicated on the form.Receipt of this form ensures that Integra has achieved a level of effectiveness in communicating this information. We recommend you also maintain a copy of this notification and signed copy of the acknowledgement form for your records. Regulatory agencies such as the FDA perform audits of field actions of this nature to verify that our customers have been notified and understand the nature of the field action.Should you have any questions regarding these instructions, please contact Sales Operations at 888-601-0203, option 2. We apologize for any inconvenience this may cause and thank you for your cooperation in this effort.Fo |  | 495 | 495 units | US Distribution to the states of : FL, VA, OK, CO, DC and IN. | Open, Classified |  |  | 606 | Not_Computer
Z-2400-2016 | 73577 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | InvacareÂż Bariatric Bed with mattress and rails , Junction Boxes, |  |  | March 15, 2016 | 2016 | Invacare Corporation | The firm recently identified a quality issue with the Junction box used in the Invacare Bariatric Bed. They discovered that the Junction Box may omit sparks which could result in a fire in the presence of flammable materials. | Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall letter dated May 12, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call Invacares recall support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. |  | 2702 | 2,702 units | US Nationwide Distribution including DC and PR. Internationally Canada. | Open, Classified |  |  | 8947 | Not_Computer
Z-2404-2016 | 74664 | container, liquid medication, graduated | KYW | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Medicine Cup, 1 OZ. Liquid medication dispenser to issue a measured amount of liquid medication. |  |  | May 11, 2016 | 2016 | Medline Industries Inc | There is 5 mL printed next to the increment on the medicine cups instead of 10 mL. The cup has the correct mL markings for the other gradations, but sequentially, the markings are listed as 2.5 mL, 5 mL, 7.5 mL, and 5 mL. | Process control | Medline Industries sent a an URGENT RECALL letter dated 5/11/2016 to all affected consignees via certified mail. The URGENT RECALL letter described the reason for the recall and the actions to take. |  | 903 | 903 Cases and 4,090 Sleeves (8,605,000 Medicine Cups) | Continental U.S. | Open, Classified |  |  | 903 | Not_Computer
Z-2390-2016 | 74619 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Specialty Triathlon Tibial Alignment Handle with Secondary Lock, Non-Sterile, |  |  | June 23, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | It was reported that the secondary locking mechanism, and its corresponding components, disassociated from the Specialty Triathlon Tibial Alignment Handle with Secondary Lock Assembly during surgery. Upon further investigation, it was discovered that the weld, which was intended to hold the secondary locking mechanism together, did not meet the weld size specified on the engineering drawing. | Process control | Stryker sent an Urgent Product Recall letter dated June 24, 2016, with a Business Reply Form attached to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to inform users of the Urgent Product Recall letter and forward the notice to all those individuals who need to be aware within their organization. Customers were instructed to return all affected product at their location to:Stryker C/O Stericycle2670 Executive Dr., Suite AIndianapolis, IN 46241Customers were also instructed to complete and sign the enclosed Business Reply Form and fax it to 888-912-7352 or email to Stericycle at strykerortho4582@stericycle.com. Customers with questions were advised to call (201) 831-6693.For questions regarding this recall call 201-831-5000. |  | 47 | 47 units | Nationwide Distribution to FL, IN, LA, NC, PA, TN and UT | Open, Classified |  |  | 47 | Not_Computer
Z-2525-2016 | 74534 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | StratusÂż CS STAT Fluorometric Analyzer-microprocessor-controlled instrument |  |  | June 22, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Software defect, where either an Above Assay Range or an inaccurate value could potentially be reported, in the Stratus CS Acute Care Diagnostics System regarding TestPak calibration and quality control (QC) status when a TestPak is due to expire within 48 hours. | Software design | Siemens Healthcare issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on June 22, 2016 both in the United States and Outside the US for communication with affected customers. This notice informs customers of the issue and provides mitigation instructions. Customers are to contact Siemens Customer Care Center or their local Siemens technical support representative at 800-405-6473. Siemens Healthcare is currently developing a software update to address this issue and will be providing new information as it becomes available |  | 977 | 977 units | Worldwide Distribution -- USA, U.A.E., Algeria, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Canada, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Namibia, P.R. China, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., and United Kingdom. | Open, Classified |  |  | 977 | Computer
Z-2396-2016 | 74743 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Orthopedic manual surgical instrument |  |  | July 01, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Malfunctioning of the Exeter Rasp/trial introducer/extractor Handle. | Device Design | Stryker notified Branches/Agencies of the Urgent Medical Device Recall by e-mail on July 1, 2016 and an "Urgent Medical Device Recall Letter and Business Reply Form" was sent to branches, agents and hospital risk management via UPS (with return receipt) on July 5, 2016. The letter identified then affected product, problem and actions to be taken. For questions call (201) 831-6693. |  | 241 | 241 units | Worldwide Distribution - US Nationwide | Terminated | June 08, 2017 | 342 | 241 | Not_Computer
Z-2405-2016 | 74724 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) |  |  |  | July 15, 2016 | 2016 | Topcon Medical Laser Systems, Inc | Topcon discovered an anomaly with titration, only when EpM is turned on during the titration process.This event may occur if the user turns on EpM, turns off the Landmark (LM) function (with any EpM supported pattern) and then enters titration mode. | Device Design | The firm, PASCALS Synthesis, sent an "Urgent Field Safety notification" letter dated July 15, 2016 to all affected sites by traceable method. The letter described the product, problem and actions to be taken. The customers were instructed to Acknowledge the receipt of the field notice via an official letter; locate all systems with Endpoint Management software feature enabled; for all affected units- inform the physician, hospital or institution customers of the addendum and additional warnings issued as a result of this notice; obtain signed confirmation from the end-user regarding receipt of the addendum and additional warnings; send the signed confirmation from ALL end-users to TMLS, no later than October 15, 2016; and complete to the following record "Products lines affected" for all affected units and either fax to (408) 235-8259 or scan and email to tmlsinfo@topcon.com.Contact Customer Care if you have questions at: 1.408.235.8200 or email: tmlsinfo@topcon.com. |  | 205 | 205 units (21 in US, 184 OUS) | Worldwide Distribution: US (nationwide) and countries of: Canada, Mexico, Japan, Netherlands, Algeria, Australia, Brazil, Chile, China, France, Guatemala, Hong Kong, Italy, New Zealand, Pakistan, Philippines, Singapore, Taiwan, and United Kingdom. | Open, Classified |  |  | 205 | Computer
Z-2395-2016 | 74716 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Q AND Q.ZEN- MODULAR ANGIOGRAPHIC SYSTEM, interventional fluoroscopic x-ray system22 |  |  | July 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The possibility exists for contaminants, in the form of biomass, to develop in the cooling system of Artis systems resulting in damage to the pump system. | Under Investigation by firm | Siemens has implemented a Customer Safety advisory Notice to all affected customer via Update Instruction AX007/15/S. Additional, a corrective action will be implemented via Update Instructions AX006/15/S and AX003/16/S. |  | 14298 | 11,111, worldwide and 3,187 in the USA | Nationwide Distribution to: AK,AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,PR,RI,SC,SD,TN,TX,UT,VA,WA,WI,WV,WY, Puerto Rico and Guam | Open, Classified |  |  | 14298 | Not_Computer
Z-2398-2016 | 74687 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | ZystonÂż Transform Implant Kit. Model Number PCR8200-3101. Device labeled sterile. |  |  | June 30, 2016 | 2016 | Zimmer Biomet Spine, Inc | Zimmer Bioment Spine, Inc. announces a voluntary field action for the ZystonÂż Transform Instrument Case and ZystonÂż Transform Implant Kit because the locking mechanism on the implant is not functioning as intended. | Under Investigation by firm | Verbal notification of the decision to recall the product and related instructions was provided on June 30, 2016. Additionally, written notification of the decision to recall the product and related instructions will be provided to each of the individuals responsible for locations that have affected inventory under their control. Clear instructions and materials necessary to support the accountability and retrieval of product from the field will be provided with the recall notification. In order to ensure that the locations were alerted of the recall in a timely fashion, supplemental email notifications will also be sent to each recall notification addressee before distributing the recall documentation via Federal Express. |  | 5 | 5 | Nationwide to PA, TX, TN, NV, VA, WV, and KY. No Canadian, other foreign, VA/gov/military consignees. | Open, Classified |  |  | 5 | Not_Computer
Z-2392-2016 | 74714 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi Scenaria Computed Tomography System Hitachi Medical Systems America, Twinsburg, OH 44087 |  |  | June 30, 2015 | 2015 | Hitachi Medical Systems America Inc | The firm discovered that the centrifugal force applied to the power supply may cause it to fail over time. The exam performed at the time of the failure will likely be lost, resulting in the need to rescan the patient. | Component design/selection | Hitachi sent an Device Correction Notice dated April 1, 2015, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Customers were informed that there is a problem with the Hitachi Medical Systems America, Inc.CT system; the issue is with a power supply in the CT gantry sub-system. During fast rotation (i.e., 35 sec), the centrifugal force applied to the power supply may cause it to fail over time.For further questions, please call (330) 425-1313 ext. 3720. |  | 33 | 33 units | US Distribution to the states of : CA, FL, GA, IA, IL, IN, MD, NC, NY, OH, PA, SD, TN, TX &amp; WY., and Internationally to Brazil and Mexico. | Open, Classified |  |  | 33 | Not_Computer
Z-2535-2016 | 74799 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Implant, Âż 4.1mm RC, SLAÂż 12mm, TiZr, NTP |  |  | July 14, 2016 | 2016 | Straumann Manufacturing, Inc. | Misalignment with the Loxim transfer piece markings, the positioning of the implant will be incorrect for optimal restoration | Under Investigation by firm | Straumann notified customers via telephone on Thursday, July 14, 2016. Customers with an e-mail address on filewere also e-mailed on Friday, July 15, 2016.Straumann notified customers were sent a hard copy letter via UPS on Saturday, July 16, 2016 for delivery on Monday, July 18, 2016.The customers have been notified not to use these implants and to return unused implants to Straumann. The customers that have used the implants have been asked to provide the restorative situation for the implants. Depending on the restorative situation, additional instructions may be necessary to restore the implants. |  | 33 | 33 pieces | Nationwide Distribution. | Open, Classified |  |  | 103 | Not_Computer
Z-2470-2016 | 73184 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Implant EP Pick-Up Coping |  |  | November 25, 2015 | 2015 | Biomet 3i, LLC | Pouches may not have been sealed during packing. | Packaging | Biomet sent an Urgent Device Recall Notice to all affected consignees the week of November 30, 2015, communicating the issue and customer responsibilities. International consignees were notified via traceable courier or email week of November 30, 2015. Customers were instructed to return affected product by calling BIOMET at 1-800-342-5454 to obtain a RMA number to return products. Customers were also asked to complete the Business Reply Form and fax to 561-514-6316.For questions regarding this recall call 561-776-6700. |  | 1 | 1, 648,273 devices (all products subject to this recall) | Worldwide Distribution Nationwide andHawaii, Guam, Puerto Rico and Virgin Islands (U.S.) .FOREIGN CONSIGNEES: Albania, Andorra, Argentina, Australia, Austria, Barbados, Belgium, Bolivia, Brazil, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Denmark, Dominican Republic, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Korea (South), Kuwait, Lebanon, Liechtenstein, Lithuania, Martinique, Mexico, Mongolia, Netherlands, Northern Mariana Islands, Paraguay, Peru, Poland, Portugal, Qatar, Reunion, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tabago, Turkey, United Arab Emirates, Uruguay and Viet Nam. | Open, Classified |  |  | 1648274 | Not_Computer
Z-2526-2016 | 74520 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB Sciex API 3200MD" LC/MS/MS System. In-Vitro Diagnostic |  |  | June 16, 2016 | 2016 | Ab Sciex | MultiQuant MD software where under certain conditions a usercan be presented with incorrect quantitative results when using the Sum Multiple Ions feature. | Software design | AB SCIEX issued recall notification by letter dated June 16, 2016, to all affected customers. The letter identified the problem, provided temporary measures until a permanent software is available. A response form was to be completed and returned. questions regarding this noticecontact SCI EX at +1 289 982 2712. |  | 124 | 124 units | Worldwide Distribution - US (Nationwide) Internationally to Canada, Austria, China, Netherlands, Norway, India, Korea, Germany, France, Switzerland, Ireland, Italy, Singapore, Russia, UK, Malaysia, | Open, Classified |  |  | 275 | Computer
Z-2532-2016 | 74734 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iConnect Access used with Ortho PACS software provides medical specialists with access to diagnostic quality images, reports, and various types of patient data over conventional TCP/IP (e.g., internet) networks. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Studies that are viewed in iConnect Access that originate on Merge PACS / Ortho PACS may not display the correct current patient demographics if there has been a PDE (Patient / Study Demographic Edit) performed after the images were ingested in PACS. | Software design | The recalling firm issued a letter dated 1/28/2016 via e-mail on 1/31/2016 or via certified mail if they did not have an e-mail address for the customer. |  | 146 | 146 sites may have the affected version for both products | Distribution was made to medical facilities nationwide and to one foreign medical facility in New Zealand. There was no military or government distribution. | Open, Classified |  |  | 146 | Not_Computer
Z-2530-2016 | 74533 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Perioscopy Sheath: Part number: 1008; |  |  | June 24, 2016 | 2016 | Danville Materials, Inc. | The Perioscopy Sheaths are currently provided with a label that indicates the Sheath has a 3 year shelf-life/expiry date. Danville Materials has become aware that validation data are not available to support the 3-year shelf-life/expiration dating claim. | Under Investigation by firm | Recall letters were sent on June 24 to Distributor OEM and customers by traceable delivery method. Letters informed customers of the issue and the affect part and lot numbers. Pro duct is to be quarantined and returned. Distributors are instructed to contact all customers and provide a copy of the recall letter. The response form should be completed and returned by each consignee. |  | 2123 | ~ 2123 units | US and New Zealand, Australia, Canada, Japan, Israel, United Kingdom, Italy, Mexico. | Open, Classified |  |  | 2123 | Not_Computer
Z-2533-2016 | 74625 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Synovis VASCU-GUARD Peripheral Vascular Patch in the following sizes and product codes: |  |  | June 24, 2016 | 2016 | Baxter Healthcare Corp. | Baxter is issuing a safety alert in response to postmarketing reports received for the VASC U-GUARD Peripheral Vascular Patch. Baxter has received reports for intraoperative or postoperative bleeding episodes, which required additional clinical intervention. One report involved a case with a fatal outcome. UPDATE 9/1/2016 Baxter issued a notice retrieving 6 lots of VASCU-GUARD due to the increased postmarketing reports related to intraoperative and postoperative bleeding. | Under Investigation by firm | Baxter sent a Safety Alert to customers via USPS overnight delivery beginning 6/24/2016. The letter told of the reported adverse events, asked for product that product immediately be discontinued from use, to segregate/quarantine all affected lots, and to return the enclosed customer reply form. General questions regarding the Safety Alert should be directed to Baxter Product Surveillance at 1-800-437-5176, between 8:00 am - 5:00 pm Central Time Monday thru Friday. Clinical inquiries can be directed to Baxter's Medical information at 1-800-933-0303, Monday thru Friday 8:00 am - 5:00 pm central time.UPDATE 9/1/2016 Baxter issued an Urgent Product Recall letter beginning 9/1/2016. This letter explained the issue and identified affected product that was to be removed from the consignee and returned to Baxter. Customer is to contact Baxter Healthcare Center for Service to arrange for return and credit. This can be done at 888-229-0001 Monday thru Friday 7:00 am - 6:00 pm Central Time. General questions regarding the recall can be addressed to Baxter Product Surveillance at 800-437-5179 Monday thru Friday, 8:00 am - 5:00 pm Central Time. |  | 294867 | 32,763 eaches UPDATE 9/1/2016 2,042 eaches | Nationwide including Puerto Rico, and outside the US to include: Australia, Austria, Finland, Germany, Greece, Ireland, Italy, the Netherlands, Norway, Poland, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, United Kingdom. | Terminated | June 23, 2017 | 364 | 294867 | Not_Computer
Z-2536-2016 | 74582 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Stryker AutoPlex System. Model Numbers 0607-687-000, 0605-887-000, 0605-687-000, and 0605-683-000. For bone cement. |  |  | June 14, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | The Piston Head can become separated from the Delivery Piston, blocking the injection assembly Valve resulting in a cement backflow towards the injector handle. Potential for delay in surgery if additional cement needs to be prepared for the injection procedure. | Process control | On 6/13/2016, URGENT MEDICAL DEVICE RECALL NOTIFICATION notifications were sent to the affected customers via courier and sales representatives via email. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Jennifer Olson269-389-2644jennifer.olson@stryker.com |  | 4841 | 4,841 boxes (9,682 units) | Worldwide distribution. US nationwide, France, Netherlands, Spain, Italy, Canada, Israel, Romania, Switzerland, and United Kingdom. | Terminated | March 27, 2017 | 286 | 4841 | Not_Computer
Z-2539-2016 | 74672 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TRYX Antibacterial Envelope Rx Only |  |  | June 29, 2016 | 2016 | TYRX Inc. | TRYX products are being recalled since the processes of spaying, welding, drying oven and polymer were not adequately validated. | Process control | Medtronic sent an Urgent Medical Device Recall letter dated June 2016 to affected customers. The letter identified the affected product problem and actions to be taken. Customers are asked to immediately remove and quarantine all unused product that remains in inventory, return unused product to Medtronic and contacted the local Medtronic representative or Customer Service at 800-848-9300 to assist with the return and credit of unused product. Medtronic will provide credit for all non-expired, unused product. The customer should complete the attached Customer Confirmation Certificate and email to RS.CFQFCA@Medtronic.com or fax to Medtronic at 651-367-0612 Attention Customer Focused Quality. |  | 3460 | 3460 units | US Nationwide Distribution | Open, Classified |  |  | 6647 | Not_Computer
Z-2538-2016 | 74683 | No | HCF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Weck EFx Classic Fascial Closure System, Rx Only, Sterile, |  |  | March 24, 2016 | 2016 | Teleflex Medical | Incorrect expiration date was printed on the product label. | Unknown/Undetermined by firm | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 24, 2016, to all affected consignees. The letter requested that consignees check their stock, immediately discontinue use and quarantine any products. To return affected product consignees were instructed to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, ATTN: Customer Service or email to recalls@teleflex.com. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 45 | 45 | Nationwide Distribution to CA, CO, IA, KS, MD, MI, NY, NC, PA | Open, Classified |  |  | 45 | Not_Computer
Z-2547-2016 | 74663 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard TruGuide Disposable Coaxial Biopsy Needle  Mixed Product |  |  | July 06, 2016 | 2016 | Bard Peripheral Vascular Inc | May contain two trocar tip stylets with coaxials or two blunt tip stylets instead of one of each. | Packaging | Bard Peripheral Vascular (BVP) sent an Urgent Medical Device Recall Notification letter dated July 6, 2016 to all affected customers via FedEx with proof of delivery notification.. The letter identified the product, the problem, and the action to be taken by the customer. The notification provided customers the following instructions:1. Do not use or further distribute any affected product.2. Check all inventory locations within your institution for affected product code I lot number combination listed in the recall notice. If you have further distributed any of the product code I lot numbers, please immediately contact that location, advise them of the recall, forward these instructions and have them return the affected product to BPV.3. Please remove any identified product from shelves.4. If used the affected product, complete and return the attached Recall and Effectiveness Check Form indicating no product will be returned.Once the product affected by recall has been removed from inventory;5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product in stock even if no longer have possession of the recalled product.6. Call the firm's Recall Coordinator Raye Seisinger at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST) or email at raye.seisinger@crbard.com. Once all information has been verified, the Recall Coordinator will issue a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will provide a credit for your returned product.7. Fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. If cannot FAX the form, please call the BPV recall coordinator at 1-800-321-4254 Option #2 Ext 2501) and report the required information verbally.8. A mailing label is enclosed to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and i |  | 1250 | 1250 units | Nationwide Distribution to the following; CA, LA, VA, IA, AR, SC, PA, MA, IL, MO.Foreign distribution to the following; Canada, Belgium, Australia, Taiwan, China, Argentina. | Terminated | November 17, 2016 | 134 | 1250 | Not_Computer
Z-2548-2016 | 74803 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Xpert MRSA |  |  | July 18, 2016 | 2016 | Cepheid | Internal investigation found the certain lots of MRSA kits have the potential for performance issues at lower positive sample concentrations. | Nonconforming Material/Component | Recall letters dated July 18, 2016 were sent to affected customers on July 19, 2016 by Federal Express. Letters advised users of the affected product and asked that any remaining inventory be destroyed. |  | 2000067041 | 68 kit for lot 1000037539 and 127 kit s for lot 1000029307. | US only: PA, DE, FL, SC, ME, CA, IL, NC, TX, WA, NY, TN, FL, OR, OK, MO, MS, GA, NM, IN, OH, MD, VA, CO, AZ, and NC. | Open, Classified |  |  | 2000067041 | Not_Computer
Z-2542-2016 | 74876 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Xario 100 Diagnostic Ultrasound System, TUS-X100; Xario 200 Diagnostic Ultrasound System, TUS-X200. |  |  | August 04, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems (TAMS) is recalling the Xario Diagnostics Ultrasound System because it may become hot because of a software error. | Software design | An Urgent Medical Device Correction letter was sent to all customers to bring to attention a potential problem of transducer heating with your Xario Ultrasound System. The letter informs the customers that the surface of the acoustic lens of a transducer may become hot because of a software error. The error occurs when a transducer is disconnected, reconnected, or another transducer is selected. The letter informs the customers of error message and the corrective action. Customers are instructed to complete the attached form and fax it to (877) 349-3054 or TAMS-RegulatoryAffairs@toshiba.com. Customer with any questions are instructed to call (714) 730-5000 or their Toshiba service representative at (800) 521-1968. |  | 93 | 93 units | US Distribution | Open, Classified |  |  | 93 | Computer
Z-2546-2016 | 74796 | Yes | NPT |  |  | Cardiovascular | PMA | Medtronic EnVeo R Delivery Catheter System (DCS), Models: ENVEOR-L, ENVEOR-L-C, and ENVEOR-US. |  |  | August 03, 2016 | 2016 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Potential for vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). | Device Design | Firm initiated a recall on 8/3/16 by sending an Urgent Medical Device Correction letter to inform them of the potential for vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). The letter informs the customers of the recommendations to reduce the rate of vascular trauma events and the procedural uses. The letter informs the customers that the Evolut R System Instructions for Use (IFU) will also be updated appropriately consistent with the Dear Doctor communication. Patients with questions can contact Medtronic Cardiovascular Patient Services at 877-526-7890 (Monday-Friday, 7:30am-5pm Central Time). Customers with any questions are instructed to contact their Medtronic Field Representative or Lifeline Technical Services at 1-877-526-7890.On 8/19/16, firm sent a copy of a customer notification letter that will be sent to their customers to inform them that of the communication letter that was released in August 2016 by Medtronic regarding potential for aortic vascular trauma with the use of the Medtronic EnVeo R Delivery Catheter System (DCS). The letter informs the customers that FDA classify the recall as a Class II recall and the recall does not require customers to return the product. The recall letter informs the customers that Medtronic has taken further action by updating all customer training materials and the Instructions for Use (IFU). |  | 7319 | 7,319 units | Worldwide Distribution -- US, Australia, Austria, Belgium, Brazil, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Kazakhstan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malaysia, Morocco, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 31, 2017 | 240 | 7319 | Not_Computer
Z-2551-2016 | 74846 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Unity Z3D software. |  |  | July 23, 2015 | 2015 | Merge Healthcare, Inc. | The software is unable to accurately determine the calcium score of scans with a slice thickness not equal to 3 mm. | Software design | The firm, Merge, sent a "canned" email on 7/23/2015 to its customers which also contained a personal message. The recalling firm issued a second letter entitled "URGENT: MEDICAL DEVICE RECALL" dated 1/27/2016 via e-mail on 1/30/2016. The letters described the product, problem and actions to be taken. The letter informed customers that an upgrade is available. The customers were instructed to call Merge Customer Service at (877) 741-5369 or send an email to support@merge.com to schedule the upgrade-Reference this recall #2016-14. If you decide to decline this maintenance release, complete the attached form and return to Merge Healthcare via recall@merge.com. Alternatively, complete the online form: http://www.merge.com/Landing-Pages/Customer-Upgrade-Decline-Form.aspx no later than February 22, 2016. Ensure that all users of the product are provided with this notification. If you have further distributed this, identify your customers and notify them at once of this product recall and provide a copy of the recall notification letter.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 9 | 9 sites potentially have the affected version | US Distribution to states of: CA, PA, TX, and MT. | Open, Classified |  |  | 9 | Computer
Z-2550-2016 | 74736 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Quattro Air FFM MED  AMER, Product Code: 62702, full face mask for use with active-exhaust-valve ventilators |  |  | July 12, 2016 | 2016 | Resmed Corporation | Masks in this lot, labelled and sold as 62702 vented Quattro Air FFM size Medium actually contain a non-vented (NV) Quattro Air FFM Large mask. | Packaging process control | An urgent product recall notification letter was sent to all customers on 7/12/16 to inform them that ResMed is recalling NV Quattro Air &amp; Quattro Air Mask due to a mislabeling Issue. The letter informs the customers that during the investigation of a product complaint, ResMed identified that a small number of masks labelled and sold as 62702 Quattro Air FFM MED  AMER, Lot 1142623 actually contain an NV Quattro Air FFM Large mask. The letter informs the customers to notify anyone they may have sold the products to. Customers with any questions are instructed to call the ResMed Customer Service team: 858-836-5968. |  | 42 | 42 units | US nationwide distribution | Open, Classified |  |  | 42 | Not_Computer
Z-2549-2016 | 74704 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DS USA Inc., Passport Series Patient Monitor intended for monitoring, displaying, reviewing, alarming, and transferring of multiple physiological parameters including ECG (3-lead, 5-lead or 12-lead selectable, arrythmia detection, ST segment analysis, and heart rate (HR), Respiration rate. |  |  | June 23, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified a potential leakage issue with the CO2 Module's Dryline II Water Trap used with the Passport Series Patient Monitor. | Under Investigation by firm | Mindray DS USA inc sent a recall letter on 6/23/2016 to their customers. |  | 360 | 360 units | Nationwide, Canada | Open, Classified |  |  | 360 | Not_Computer
Z-2545-2016 | 74880 | No | MGZ | External vein stripper. | Cardiovascular | Cardiovascular | 510(k) | 1.5 mm HYDRO LeMaitre Valvulotome, 40 cm. Sterile |  |  | August 03, 2016 | 2016 | LeMaitre Vascular, Inc. | Wire which deploys the loop may be come detached at the handle causing the device to be inoperable | Process control | LeMaitre Vascular, Inc. sent an Urgent Medical Device Recall letter dated June 29, 2016 to customers via priority mail. The letter identified the affected product, problem and actions to be taken. Customers were requested to locate, quarantine and return the devices. Customers were instructed to complete and return the response form to LeMaitre Vascular, Inc. as a record of notification and reconciliation. Any unused product will be requested to be returned and replaced. For questions call 781-221-2266 ext. 183. |  | 50 | 50 units | Worldwide Distribution -  US Nationwide and in the countries of Australia, BELGIUM, BRAZIL, CANADA, Czech Republic, DENMARK, FINLAND, FRANCE, GERMANY, Italy, Japan, NETHERLANDS, Russia, Spain, SWEDEN, Switzerland, and UK. | Open, Classified |  |  | 2422 | Not_Computer
Z-2579-2016 | 64785 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS Workstation, versions 3.1.1.x through 3.2.1.x Picture Archiving and Communication System |  |  | November 15, 2012 | 2012 | GE Healthcare It | While merging exams in a test instance, merges carry forward in the production system, when the middle tier is configured to the wrong IP address of the Centricity Exam Manager. When a current and historical exam are opened at the same time, the system asynchronously tries to access the operation specifying how each image should be grouped, causing random image-series grouping errors. | Software design | GE sent an Urgent Medical Device Correction letter dated November 15, 2012. The software has been modified to correct these defects. The modified software was released and installed in user sites under GE Healthcare FMI #85196. |  | 498 | 498 units | Worldwide distribution including US nationwide, Puerto Rico, Australia, Austria, Belgium, Bermuda, CANADA, Denmark, Egypt, ENGLAND, France, Germany, India, Ireland, Hong Kong, Korea, Kuwait, Italy, Malta, Netherlands, Portugal, Qatar, Saudi Arabia, SCOTLAND, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, and UK. | Open, Classified |  |  | 498 | Computer
Z-2580-2016 | 74828 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products TSH Reagent Pack, REF/Catalog Number/Product Code 191 2997, Unique Device Identifier No. 10758750000227, IVD --- Ortho-Clinical Diagnostics Pencoed Bridgend CF35 5PZ UK |  |  | July 28, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | A field correction was issued by the firm for VITROS TSH reagent lot 5040 due to a higher than expected frequency of customer calibration curves falling outside of the calibration quality parameters and therefore unusable to process samples. Customers using VITROS TSH Lot 5040 will potentially experience an unexpected increase in the unsuccessful calibrations potentially causing a delay in testing and reporting results for VITROS TSH. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter on July 28, 2016, via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received VITROS Immunodiagnostic Product TSH, Lot 5040 to inform them of the issue and the required actions. Instructions were provided to customers. Customers with questions were instructed to contact Ortho Care" Technical Solutions Center at 1-800-421-3311. Foreign affiliates were informed by e-mail on July 28, 2016, of the issue and instructed to notify their consignees of the issue and the required actions.For questions regarding this recall call 908-218-8776. |  | 4144 | Domestic: 2883, Foreign: 1261 | Worldwide Distribution - US to VA, WV, GA, TN, KY, MN, SD, AR, OK and Internationally to Canada | Open, Classified |  |  | 4144 | Not_Computer
Z-2584-2016 | 74419 | No | MQN | Bone plate. | Dental | Dental | 510(k) | TI Single Vector Distractor Body with Left Foot/30mm; Catalog ID 487.965 |  |  | May 23, 2016 | 2016 | Synthes (USA) Products LLC | DePuy Synthes is voluntarily initiating a recall due to a possibility the fastener on the Titanium Single Vector Distractor Body may become prematurely separated from the Proximal Foot Plate. | Under Investigation by firm | The firm, DePuy Synthes, mailed out an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 5/23/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately take the following actions: If you DO have any of the identified devices call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, complete the Verification Section (page 3 of the letter) indicating the number of devices that were found and return the Verification Section (page 3 of the letter) with the product to the following: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132, by fax 844-449-8498 or scan/email: Synthes8096@stericycle.com., and complete the Verification section even if you DO NOT have the identified product.Note: If the Verification Section is answered on behalf of more than one facility and/or individual clearly indicate the name and address of the facility and/or individual.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 259 | 259 units | Worldwide Distribution-US (nationwide) distribution to states of: AZ, CA, CO, DC, FL, IL, MA, MI, NJ, NY, PA, TX, UT, WI, and WV; and countries of: Australia, Canada and Switzerland. | Open, Classified |  |  | 1817 | Not_Computer
Z-2403-2016 | 74843 | [6 empty] | OraQuick Ebola Rapid Antigen Test, Product Code 3001-2807; Intended for the qualitative detection of antigens for Ebola viruses. |  |  | July 20, 2016 | 2016 | OraSure Technologies, Inc. | Failed stability testing at 8 months. This lot produced intermittent false negative results for the positive kit control with devices stored at 30C conditions. | Process control | OraSure contacted the sole customer via email and mail (letter dated July 21, 2016) to request return of all the product from the affected lot. |  | 3075 | 3,075 units | Distributed to one customer only in US state of Georgia. | Terminated | May 10, 2017 | 294 | 3075 | Not_Computer
Z-2552-2016 | 74781 | No | LTI |  |  | Gastroenterology/Urology | PMA | ReShape Integrated Dual Balloon System, Gen 1, Catalog No. RSM110, Model No. 01-0013-001 |  |  | May 04, 2016 | 2016 | Reshape Medical Inc | Due to a suspected misalignment of the connection between the proximal balloon cap of the implant and the delivery catheter. The possible misalignment is the result of an assembly error during manufacturing and may lead to leakage of saline solution during the balloon inflation procedure. | Process control | A Dear Doctor letter dated May 4, 2016, was sent to all the customers to inform them that ReShape Medical is recalling ReShape Integrated Dual Balloon System, Gen 1 due to a suspected misalignment of the connection between the proximal balloon cap and the delivery catheter. The letter informs the customers that the possible misalignment is the result of an assembly error during manufacturing and may lead to leakage of saline solution during the balloon inflation procedure. Customers are informed of the planned field correction and if customers have any questions then they're instructed to contact ReShape Medical Customer Service at (844) 937-7374.For further questions please call (949) 429-6680 ext. 106 |  | 56 | 56 units | US: Distribution to the states of : NY, OH and CA. | Terminated | August 17, 2016 | 105 | 56 | Not_Computer
Z-2565-2016 | 74602 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Valleylab Laparoscopic Wire J-Hook Electrode 36CM |  |  | July 08, 2016 | 2016 | Medtronic | Potential for compromise of the package resulting in a breach of the sterile barrier | Process control | Medtronic/Covidien sent an Urgent Medical Device Recall letter dated July 8, 2016, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. The letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Customers with questions were instructed to contact their Medtronic representative or Customer Service at 800-882-5878. |  | 349629 | 349629 | Worldwide Distribution - US (nationwide) and Internationally toAustralia, Canada, Japan, Korea, Hong Kong, Thailand, | Open, Classified |  |  | 995222 | Not_Computer
Z-2553-2016 | 74403 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | METS SMILES Total Knee Replacement |  |  | May 31, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | The manufacturer has identified that the stated Femoral Plateau Plates may not fit into the intended Femoral Knee mating component and thus sit flush upon the Femoral Knee component prior to implantation. | Device Design | The recalling firm sent recall notification letters to their consignees on 05/31/2016. Follow up communciation was sent out on 7/26/2016 after a technical and medical assessment was completed. |  | 3 | 3 units | Directly distributed to TX, further distributed to AL and MS | Open, Classified |  |  | 3 | Not_Computer
Z-2593-2016 | 66448 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Hospira RECEPTAL CANISTERS: 1L(1000mL), 1.5L(1500mL), 2L(2000mL) |  |  | September 10, 2013 | 2013 | Hospira Inc. | Hospira has become aware of customers using the incorrect size Receptal liners with their Receptal canisters, e.g., 1.5 L liner, 2L liner or 3L liner with a 1L canister. It is possible that if an incorrectly sized liner is used with a different size canister the lid may not fit and the vacuum/suction will not be sufficient and/or consistent. | Process control | URGENT MEDICAL DEVICE RECALL letters dated September 10, 2013 were distributed to direct accounts notifying them of the recall. The letter requests that customers to ensure that they only use liners and canisters of corresponding sizes (e.g. a 1000mL liner with a 1000mL canister) and to only use Hospira components. The letter further advises customers of an ongoing recall of the 1000mL Receptal systems. Customers were requested to complete and return the attached reply form and to notify other healthcare providers within their organization.Customers who further distributed the Receptal systems were requested to forward the letter to their customers and to request that the contact Stericycle at 1-888-943-5175 (Monday - Friday, 8:00 AM - 5:00 PM, CST) to obtain a reply form. |  | 4544224180761321 | 67,410,861 liners and 67,410,861 lids/canisters in total | *** US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY), District of Columbia, Puerto Rico; *** FOREIGN: Australia, Canada, Costa Rica, Singapore, South Africa | Open, Classified |  |  | 4544224180761321 | Not_Computer
Z-2591-2016 | 73465 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | PEDIATRIC KIDNEY BIOPSY TRAY |  |  | October 05, 2015 | 2015 | Avid Medical, Inc. | Avid Medical Inc. states a latex-containing Telfa" Non-adherent Dressing 4"x3" (10.2cm x 7.6 em) was included in a limited number of convenience kits (custom procedure trays) that were not labeled as containing latex. | Employee error | Avid Medical Inc. states a latex-containing Telfa" Non-adherent Dressing 4"x3" (10.2cm x 7.6 cm) was included in a limited number of convenience kits (custom procedure trays) that were not labeled as containing latex. |  | 20 | 20 | Recalled product was distributed to 2 states: Virginia and Massachusetts. | Terminated | January 10, 2017 | 463 | 700 | Not_Computer
Z-2587-2016 | 74595 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | OPTICAL BIOMETER AL-Scan: Software Version: 1.09 and earlier. |  |  | July 01, 2016 | 2016 | Nidek Inc | Nidek Inc. received information from our manufacturer NIDEK CO. Japan that the shown axiallengths were much shorter at AL-Scan measurement. | Software in the Use Environment | Dear Doctor/Recall letters were sent to affected customers on July 7, 2016. The letters describe the issue and provide instruction should the issue occur. The letters inform users that Nidek will reach out in the next few weeks to schedule installation of new software that will solve the problem.Questions or concerns may be directed to:Customer Service 1.800.722.0219; Monday- Friday 8 am to 5pm PSTE-mail: AL-Scan_support@noritsuservice.com. |  | 17 | 17 | US Nationwide Distribution in the states of: NY, PA. GA. AZ, VA, DE, ME,. MO, TN, and NM. | Open, Classified |  |  | 17 | Computer
Z-2588-2016 | 74752 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROFEMUR(R) MODULAR NECK, HIP FEMORAL NECK MicroPort total hip systems intended for use in total hip arthroplasty for reduction or relief of pain and/or improved hip function in skeletally mature patients. |  |  | June 30, 2016 | 2016 | MicroPort Orthopedics Inc. | Profemur(R) neck, Item Number PHA01222 (Short AR VV2 Ti Neck) was labeled as Item Item Number PHA01212 (Short AR VV1 Ti Neck). | Error in labeling | The firm initiated their recall by email on 06/30/2016. |  | 14 | 14 units | Belgium, Italy, Spain, and Poland | Open, Classified |  |  | 14 | Not_Computer
Z-2606-2016 | 74800 | No | LOD | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Cobalt HV Bone Cement |  |  | July 26, 2016 | 2016 | Encore Medical, Lp | The outer packaging was mislabeled on the box indicating "Cobalt HV with Gentamicin". The bone cement does not contain antibiotics. | Labeling mix-ups | The recalling firm sent an Urgent Field Safety Notice letter dated July 26 2016 to Sales Agents. The letter identified the affected product, problem and actions to be taken. The sales agents were instructed to pass on the notice to any organization where the potential affected product has been transferred. Agents were instructed to identify hospitals in their territory, complete a visually inspection of the affected product in the hospital's inventory, complete an Acknowledgement and Receipt form. If affected products is found contact Customer Service at 1-800-456-8696 for replacement. |  | 866 | 866 units | US Nationwide Distribution | Terminated | January 13, 2017 | 171 | 866 | Not_Computer
Z-2603-2016 | 74801 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TIGR Matrix Surgical Mesh, Model No. NSTM1015, NSTM1520, NSTM2030 |  |  | July 19, 2016 | 2016 | Novus Scientific Ab | Novus Scientific is implementing additions to the product's instructions for use concerning certain inguinal hernia repairs. | Other | An Urgent Safety Alert letter dated 7/19/16 was sent to all their customers to inform them that Novus Scientific is implementing additions to the product's instructions for use (IFU) with a new contraindication and warning related to the clinical study result. The letter informs the customers that until the new IFU is made available, surgeons using the product for hernia repair should follow the instructions in the urgent safety alert in addition to those in the current IFU.Customers are informed of the actions to be taken:" Additional contraindication: Not suitable for the repair of direct inguinal hernias." Additional warning: Because TIGR Matrix Surgical Mesh is fully resorbable, it should not be used in repairs where permanent support from the mesh is required.Customers with any questions or concerns are instructed to contact Novus Scientific AB ( +46 (0)18 700 11 50. |  | 3458 | 3458 units total (657 units in US) | US : IL, CA, FL, KY, MO, NH, VA, TX, NC, OH, NY, TN, Worldwide: Sweden, Italy, Denmark, UK, Singapore, Germany, Belgium, Spain, Austria, Israel, Malta, Switzerland, Ireland, Norway, Netherlands, Turkey, New Zealand, Bulgaria, France, Iran, Finland, Poland, South Africa | Terminated | December 23, 2016 | 157 | 3458 | Not_Computer
Z-2604-2016 | 74884 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE VC150 Vital Signs Monitor; Intended to monitor a single patient's vital signs at the site of care. |  |  | July 27, 2016 | 2016 | INNOKAS MEDICAL OY | A software error on released software versions 1.6.12, 1.6.12F and 1.6.16 may give wrong time data to measurements. | Software design | Innokas Medical has notified the one US distributor on 7/27/16 via letter of the deficiency and correction of the device. |  | 1715 | 1715 units | AR, AZ, FL, IN, LA, MA, MI, NC, NJ, NM, NY, OH, PA, TN, WI | Terminated | April 19, 2017 | 266 | 1715 | Computer
Z-2595-2016 | 73607 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Adult ThermoGardÂż Dual Dispersive Electrodes (for patients &gt;15 kg.), 10' (3.05m) Cable, Catalog Number 51-7310. |  |  | May 18, 2016 | 2016 | Conmed Corporation | A version of the dual dispersive electrodes may not be compatible with some electrosurgical generators, placing patients at risk for undetected pad lift and a potential burn. | Device Design | On May 17, 2016 CONMED Corporation distributed Urgent Device Recall notification letters and Business Reply Forms to their customers via courier service. CONMED Corporation has decided to recall these ThermoGardÂż Dual Dispersive Electrodes for all customers who do not use an electrosurgical generator that employs the Contact Quality Monitoring systems listed. If your electrosurgical generator is not listed, your generator should not be used with the lot codes of the ThermoGard Dual Dispersive Electrodes identified. Customers are advised to inspect inventory for any of the recalled devices. If you use the ThermoGardÂż Dual Dispersive Electrodes in conjunction with an electrosurgical generator listed, no further action is required. However, if your electrosurgical generator and CQM system is not listed above, do not use the affected lot codes of ThermoGardÂż Dual Dispersive Electrodes. Please note that ThermoGardÂż Dual Dispersive Electrodes produced after July 5, 2015, are not subject to this recall. Customers should contact all of those departments or organizations within their facility and any other facilities that you may have supplied or given these affected products to. It is imperative that all end users of these devices receive this notice and respond immediately. If you have questions, please contact Patricia Cotter, CONMED Recall Coordinator 315-624-3237 (8AM-5PM EST, M-F) or fax to 315-624-3225 or email thermog@CONMED.com. |  | 307684 | 307,684 units | Worldwide Distribution -- USA, including AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; Puerto Rico; and, the countries of Australia, Canada, Belgium, Brunei Darussalam, China, France, Guatemala, Hong Kong, Israel, Italy, Japan, South Korea, Malaysia, Mexico, New Zealand, Norway, Portugal, Saudi Arabia, Spain, Sweden, and United Arab Emirates. | Open, Classified |  |  | 447697 | Computer
Z-2360-2016 | 74665 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Alere INRatio2 PT/INR Professional Monitoring System |  |  | July 11, 2016 | 2016 | Alere San Diego, Inc. | Alere San Diego is initiating a voluntary recall for the Alere INRatioÂż/INRatioÂż2 PT/INR Monitoring Systems (Test Strips and Meters). This action is being initiated as a result of patients who had a therapeutic or near-therapeutic INR with the Alere INRatioÂż/INRatioÂż2 PT/INR Monitoring System but a significantly higher INR (outside of therapeutic range) when performed by a central laboratory. | Other | Alere issued a press release on July 11, 2016. An Urgent Medical Device Letter dated July 26, 2016 was sent to affected customers on July 29, 2016. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to consult with their healthcare provider as soon as possible to transition to an alternate method of PT/INR testing and to contact the firm regarding return or disposal of unused product and obtaining device replacement from a different manufacturer. For questions contact the Alere INRatid1 Recall Hotline at 1-866-723-2535. |  |  |  | Worldwide Distribution - US Nationwide and foreign distribution to the countries of: Argentina, Australia, Azerbaijan, Bangladesh, Belarus, Belgium, Brazil, Brunei Daruss, Canada, Chile, China, Colombia, Egypt, France, Hong Kong, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Malaysia, Maldives, Netherlands, Russia, Rwanda, Singapore, South Africa, South Korea, Spain, Taiwan, Thailand, Trinidad, Ukraine, Uruguay, and Vietnam. | Terminated | July 20, 2017 | 374 |  | Not_Computer
Z-2605-2016 | 74737 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Panacea Air Pro Mattress Control Unit, Model: Platinum 5000, US Model Number 5000. |  |  | July 11, 2016 | 2016 | American National Mfg Inc | The Panacea Air Pro Mattress Control Unit is being recalled due to a firmware problem, which may result in the device not having the firmness range required. | Software Manufacturing/Software Deployment | An Urgent Medical Device Recall letter, dated 7/11/16, was sent to customers to inform them of a recall involving the Panacea Air Pro Mattress Control Unit also labeled as Platinum 5000, US Model Number 5000, Serial numbers 9500051 through 9500300. The letter informs the customers that the recall has been identified due to firmware problem. The letter informs the customers of the problem identified and the actions to be taken. Customers were asked to immediately examine their inventory and quarantine the affected device. A shipping label would be provided for return and update of the affected devices. All customers that the device was further distributed to would need to be notified of the recall. |  | 31 | 31 units | Distributed to the US -- WI, OH, IL, TX, AZ, MD, IA, PA, CO, WA, ID, MN, LA, WY, and MS. | Open, Classified |  |  | 31 | Not_Computer
Z-2608-2016 | 74656 | No | HHI | Vacuum abortion system. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Conical portio adapter (acorn) small &amp; large are used for examination and diagnosis in conjunction with endoscopic accessories in gynecological laparoscopy, to determine patency of the fallopian tubes. |  |  | May 19, 2016 | 2016 | Richard Wolf Medical Instruments Corp. | The distal end of the Conical portio adapter (acorn) is composed of a small cap (head) that partially sits inside a hollow shaft (tube). These two parts are welded together to create one piece. In the impacted batches of product, the weld was improperly performed. As a result, the two pieces are not properly joined. This potential failure can result in the separation of the two pieces. The possible separation of these two components presents a potential risk to patients. | Nonconforming Material/Component | Customers impacted were notified via Urgent Medical Device Recall Letter dated 5/31/2016 via certified mail. The letter describes the reason for the recall. Richard Wolf Identified the following measures to be taken by the consignees:1. Immediately examine your inventory and quarantine product subject to recall. Richard Wolf Medical Instruments Corporation strongly recommends not using the potentially defective products. Please make sure that all users of the aforementioned products become aware of this urgent safety information. 2. If you have further distributed this product, please identify your customers and immediately notify them about this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter, but must include instructions on what customers should do with the recalled product. 3. PLEASE COMPLETE ENCLOSED ANSWER FORM {even if the product is no longer in your inventory) and return the record to Richard Wolf Medical Instruments Corporation before June 17, 2016. Your return of the answer form will: " Confirm receipt of this safety information" Trigger a replacement order. Richard Wolf Medical Instruments Corporation will provide your facility with a return authorization number and shipping instructions and will bear the cost for shipping (returns to RWMIC and replacements to facility). 4. Please immediately inform Richard Wolf Medical Instruments Corporation of any adverse events.If the consignees have any questions they are asked to contact Lisa Williams, Regulatory Affairs Assistant at Richard Wolf Medical Instruments Corporation Richard Wolf Medical Instruments Corporation.Telephone: (847) 913-1113 Ext: 256 Telephone: (847) 913-1113 Ext: 225Fax: (847) 913-0924 Fax: (847) 913-0924Email : lwilliams@richardwolfusa.comMonday through Friday, 7:30AM to 4:00 PM CST. Monday through Friday, 9:00AM to 5:30 PM |  | 16883.81 | Product code 8378.90- 94 pieces &amp; Product code 8378.91 - 32 pieces | U.S. Consignees: AK, AL, AR, AZ, CA, CT, DC, FL, GA, IA, IL, KY, MA, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, TX, UT and VA. | Open, Classified |  |  | 16883.81 | Not_Computer
Z-2602-2016 | 74545 | No | MHY |  |  | Neurology | PMA | Medtronic, Lead Kit for DBS Stimulation. Model Numbers 3387S-40 (p/n 3387S0007V) &amp; 3389S-40 (p/n 3389S0008V). |  |  | June 10, 2016 | 2016 | Medtronic Neuromodulation | Medtronic has voluntarily decided to retrieve unused Model 3387S-40 and 3389S-40 DBS leads from three manufacturing lot numbers.During the manufacturing process of a DBS lead component, Medtronic identified the potential for lead insulation damage. | Process design | Consignees were delivered in person by Medtronic Field Representatives a "Medical Device Removal" letter dated June 2016. The letter described the problem and the product involved in the recall. Advised consignees to return the unused product and to complete the Customer Confirmation Form. For questions contact Medtronic Technical Services at 800-707-09033 weekdays 7 am - 6 pm, or contact your Medtronic representative. |  | 70 | 70 | US: AR, CA, FL, IL MD, MO, NE, NY, OH, TN WI. | Open, Classified |  |  | 70 | Not_Computer
Z-2607-2016 | 74252 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Radiesse (+) Lidocaine, Injectable Implant 1.5cc, Rx only, P/N: 8063M0K1, Merz Aesthetics. |  |  | May 20, 2016 | 2016 | Merz North America, Inc. | Complaints reported of Radiesse (+) unable to be expelled from the syringe and/or needle. | Nonconforming Material/Component | Merz sent an Urgent Medical Device Recall Letter dated May 20, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete the enclosed response form and return via fax to (262) 835-3330 or email to complaints2@merz.com. For questions contact Merz at (262) 835-3300 ext. 2627. |  | 1591 | 1,591 units | US Nationwide Distribution in the states of AR, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MN, MO, NC, ND, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, VT and WA. | Terminated | May 08, 2017 | 353 | 1591 | Not_Computer
Z-2600-2016 | 74855 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Large QWIX Compression screw, dia. 7.5 mm L 90 mm, Lot No. FH86, Rx only, |  |  | July 21, 2016 | 2016 | Integra LifeSciences Corp. | Labeling on the Qwix Screw product states that it is a "compression screw" instead of a "positioning screw". | Labeling mix-ups | Integra sent an Urgent Voluntary Medical Device Recall letter and acknowledgement return response form dated July 21, 2016 to their affected customers. The recall letter identifies the product and lots involved, problem and actions to be taken. All customers are asked to return the completed acknowledgement and return response form by email to FCA1@integralife.com or fax to 1-609-275-9445. Should any customer has any questions regarding the instructions, they can contact Customer Service at 1-800-654-2873. |  | 120 | 120 units from multiple lots | US Nationwide Distribution in the states of AL, GA, ID, MA, MD, MN, NC, NJ, NY, OH, OK, PA, PR, TX, VA and WA | Open, Classified |  |  | 120 | Not_Computer
Z-2601-2016 | 74564 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDPointÂż 500 v2.2.2A Software Upgrade Kit; SMN 11066719, Software Version 2.2A |  |  | July 05, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Some v2.2.2 upgrade kits include a dialysate mode which not cleared/approved for shipment in the United States | Under Investigation by firm | Siemens sent a Customer Notification dated July 2016 to all affected domestic (US) customers via Federal Express on July 5, 2016. Customers were asked to take the following actions: Please do not use the dialysate mode. If you have enabled this feature, please disable it by following these instructions:1. At the System screen, select Setup. 2. Select Secured Options &gt; Analysis Options. 3. Unselect Dialysate. 4. Select the Continue button 3 times.Please review this letter with your Medical Director. Complete and return the Field Correction Effectiveness Check attached to this letter within 7 days.Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens Technical Support representative.For questions regarding this recall call 781-269-3000. |  | 76 | 76 | Nationwide Distribution to AL, CA, MA, NJ, OH, TX, WA, PA | Terminated | August 02, 2017 | 393 | 76 | Not_Computer
Z-2609-2016 | 66606 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 1) The HomeChoice automated peritoneal dialysis system. Dataplate located on the device is labeled in part: Baxter. |  |  | October 07, 2013 | 2013 | Baxter Healthcare Corp. | In systems with version 10.4 software Initial Drain logic, the device will attempt to drain the patient to empty at the beginning of therapy (initial drain). The operator can stop, but cannot bypass, the active Initial Drain. This is to mitigate against Unintended Increased Intraperitoneal Volume (IIVP). This can cause serious problems in patients with unrelated a co-morbid condition of ascites. | Labeling design | An IMPORTANT PRODUCT INFORMATION letter dated October 7, 2013, was sent to all affected customers. The letter included instructions for clinicians to: 1) be aware that HomeChoice will attempt to drain the patient to empty at the beginning of therapy (Initial Drain); 2) review HomeChoice instructions for use for additional details on HomeChoice 10.4 Initial Drain Logic; 3) acknowledge your receipt of the notification by completing the attached Customer Reply Form and returning by fax to either 224-270-5457 or scanning and e-mailing it to fca@baxter.com; 3) if the recipient of the letter is a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, to forward the communication to all customers who may have received these products. Clinical questions should be addressed to Baxter's Renal Division Clinical Helpline at 888-736-2543, option 2 (Monday - Friday, 8:00 AM - 5:00PM CST). General questions about this recall should be addressed to The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 65590 | *** 1) Product Codes 5C4471 and 5C4471R: Approximately 48,600 units; *** 2) Product Codes 5C8310 and 5C8310R: Approxiamtely 16,990 units | Worldwide distribution. US nationwide including Puerto Rico; Australia, Canada, New Zealand, and Japan | Completed |  |  | 65590 | Computer
Z-2610-2016 | 74010 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | BeaconÂż Tip Torcon NBÂż Advantage Catheter |  |  | April 15, 2016 | 2016 | Cook Inc. | Increase in reports of polymer degradation of the catheter tip, resulting in tip fracture and/or separation.degradation could include loss of device function, separation of a device segment leading to medical intervention, or complications resulting from a separated segment. Such complications include device fragments in the vascular system, genitourinary system, or other soft tissues. Fragments within the vascular system could result in embolization to the heart or lungs, or occluding blood flow to end organs. | Unknown/Undetermined by firm | Cook Inc. sent an URGENT: MEDICAL DEVICE RECALL NOTIFICATION letter dated April 15, 2016, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers with any medical questions or concerns, were instructed to contact Cook Medical Customer Relations at 800.457.4500 or 812.339.2235. For information regarding the recall, please contact Stericycle Expert Solutions at 866.201.9067. |  | 2662505 | 2,662,505 | Worldwide Distribution - US (Nationwide) and Internationally to Algeria, Andorra, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Barbados, Belgium, Bermuda, Brazil, Brunei Darussalam, Bulgaria, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Djibouti, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guatemala, Guernsey, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jersey, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macau, Malaysia, Malta, Martinique, Mauritius, Mexico, Monaco, Mongolia, Morocco, Myanmar, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, "Palestinian Territory, Occupied", Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad And Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, U.S. Virgin Islands, Yugoslavia. | Terminated | August 04, 2017 | 476 | 4146309 | Not_Computer
Z-2627-2016 | 74919 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 31, 2016 | 2016 | Merge Healthcare, Inc. | Potential for RadSuite AV viewer to skip image slices when the user presses page up or page down when the viewer displays multiple viewports and the user has mouse focused on any viewport other than the first viewport. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated January 28, 2016, and via email on January 31, 2013, to all affected customers. the letter identified the product the problem and the action needed to be taken by the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it. If you have any additional questions call Merge Customer Service at (877) 741-5369.For further questions, please call (262) 367-0700. |  | 25 | 25 sites potentially have the affected versions | Distribution was made to AL, AZ, CA, IL, MD, MA, MI, MO, NV, NJ, NC, OH, PA, TN, and VA. Government distribution was made to OK. There was no foreign/military distribution. | Open, Classified |  |  | 25 | Not_Computer
Z-2630-2016 | 74759 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 120 V Neptune 3 Rover, Model Number: 0703-001-000 |  |  | July 01, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker initiated a voluntary recall of the 120 V Neptune 3 Rover on 07/01/2016, due to a potential for the device to experience system errors causing the device to shutdown if high levels of electromagnetic interference are present in the operating room. | Under Investigation by firm | Stryker Instruments initiated a voluntary recall of the Stryker Neptune 3 Rover via certified mail on 07/01/2016, due to the potential for the Neptune 3 Rover to experience system errors if there are high levels of electromagnetic interference in the operating room. 1. Immediately review the Recall Notification.2. Confirm that the Rover serial numbers identified on the Business Reply Form (BRF) are at your facility.3. A Stryker representative will facilitate the repair of your affected Neptune 3 Rover(s). Please coordinate the availability of the product to assure the repair can be implemented.4. Please complete and return the Business Reply Form (BRF).Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.5. Fax the completed, signed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com.6. Stryker will document the repair of each unit by serial number.For questions regarding this recall, please contact Stryker Instruments:Kelly Jo Davis269-389-2921kellyjo.davis@stryker.com |  | 78 | 78 | Domestic:MI, CA, ID, IAVA/DOD: NoneForeign: None | Terminated | June 02, 2017 | 336 | 78 | Not_Computer
Z-2624-2016 | 74840 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS with software version VA48A-SP2; Model # 10430603, computed tomography x-ray system |  |  | July 13, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to an internal communication error between the firmware and the software of the components, the planned CT scan executes properly, but the injector is not started. Therefore, the contrast agent is not injected and the desired examination result is not achieved. This error only affects the automatic mode or coupled mode and does not affect the manual control of the injector. | Software in the Use Environment | Siemens mailed to customers an Important Customer Safety Notice to inform about a possible malfunction when using the affected systems and what actions to take to eliminate the problem. Also, Siemens informed their customers that they were currently developing a new firmware to eliminate the internal communication error, which should be released in August 2016. Customers were asked to contact their service organization at 800-888-7436, if they had any questions. |  | 1 | 1 unit | Distributed to: MI, NY, CA, KY, ND, NE | Open, Classified |  |  | 7 | Computer
Z-2628-2016 | 74808 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC UroView 2800, fluoroscopic x-ray system |  |  | July 22, 2016 | 2016 | GE OEC Medical Systems, Inc | A firmware issue in the supplied LCD workstation monitor(s). Under some circumstances, the monitor(s) may appear dark with no video displayed. | Software design (manufacturing process) | On July 22, 2016 GE Healthcare Surgery began distributing an Urgent Medical Device Correction letter to customers via mail to inform them of the issue and provide instructions for immediate mitigation of the issue. The letter was sent to 3 titles at each facility affected: 1.) Facility Administrator, 2.) Director/Manager of Radiology and Radiologists, 3.) Radiology Department. Consignees instructured that local GE Healthcare Field Engineer will contact them to coordinate replacement of the affected monitor(s), and will ensure that system meets specifications. |  | 47 | 47 total | Distribution US nationwide and Korea. | Open, Classified |  |  | 47 | Not_Computer
Z-2631-2016 | 74850 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | RESPONSIVE, KNEE: FEMORAL COMPONENT CR/CS, SIZE 3R, REF 90-SRK-311300, QTY: 1, STERILE R, |  |  | July 18, 2016 | 2016 | Medtronic Sofamor Danek USA Inc | The alpha total knee replacement components that may have not been passivated in accordance with the requirements of the job traveler. | Employee error | On 07/18/2016, the firm initiated the correction by verbally notifying (by visit and telephone) three of four (3 of 4) implanting surgeons; the fourth implanting surgeon was verbally notified on 07/25/2016. The firm followed with a letter to all four (4) surgeons on 07/28/2016. Medtronic recommended that the condition of the patients receiving the subject part be assessed for any indications of premature failure during normal post-operative follow-up visits. |  | 3 | 3 units | AR, MN,MA | Open, Classified |  |  | 16 | Not_Computer
Z-2645-2016 | 74552 | No | MHY |  |  | Neurology | PMA | DBS Pocket Adaptor, Model 64001 and Model 64002. |  |  | June 23, 2016 | 2016 | Medtronic Neuromodulation | Due to the potential for pocket adaptor conductor wire fractures involving DBS Pocket Adaptor Model 64001 and Model 64002, a safety notification was sent to healthcare professionals to reinforce device specific labeling for the handling of pocket adaptors and system integrity checking during implant procedures. | Nonconforming Material/Component | Consignees were sent on 6/23/2016 a "Medical Device Safety Notification," dated June 2016. The letter was addressed to Healthcare Professional. The letter described "Background" information on the problem and product involved in the recall. It also provided "Recommendations" which included "Implanting the pocket adaptor with the neurostimulator" and "Checking system integrity".For questions contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7 am - 6pm CST. |  | 20541 | 7,542 in the US and 12,999 outside the US. | Worldwide Distribution. | Open, Classified |  |  | 20541 | Not_Computer
Z-2668-2016 | 74607 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeamÂż Radiotherapy Delivery System and EDGE" Radiotherapy Delivery System, K140528. The TrueBeam and Edge Systems are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation therapy is indicated for adults and pediatric patients. |  |  | June 30, 2016 | 2016 | Vision RT Ltd | Vision RT Ltd received three (3) reports from different sites that discovered the following issue, unintended changes were made to planned couch parameters, specifically the couch rotation parameter, during patient set-up for the Optical Surface Monitoring System [OSMS].1. CRM 3194, No Serious Injury, No MDR, Aware date: Dec 11, 2015 2. CRM 3321, No Patient Involved, No MDR, Aware date: Jan 19, 20163. CRM 4570, No Serious Injury, No MDR, Aware date: March 15, 2016No patient harm was reported in any of these cases. | Other | Vision RT Ltd will provide all affected AlignRTÂż Plus Users with an Urgent Medical Device Correction Notification letter, dated 2016-06-24 Subject: Unexpected Overwrite of Plan Couch Parameter. The letter will provide Users with a detailed description of the problem and provide instructions to avoid the issue. Non-responders will be notified again until we achieve 100% for North America and International consignees. Vision RT will make three (3) attempts using at least two different methods of contact to obtain proof of notification.Containment: Vision RT will provide a copy of the Urgent Medical Device Correction letter for any site accepting an AlignRT device. The letter will provide Users with a detailed description of the problem and provide instructions to avoid the issue.Product will not be returned to Vision RT.Varian Medical Systems and Vision RT are investigating development of a technical fix for the affected products. Notification of the intention to correct is provided in the Urgent Medical Device Correction Notification letter, dated 2016-06-24 Subject: Unexpected Overwrite of Plan Couch Parameter. Vision RT Ltd will contact all affected AlignRTÂż Plus users and schedule upgrades to implement the correction when available. |  | 47 | 47 | Product distributed in US.The foreign countries where the product was distributed are as follows: Canada, Germany, Japan, Sweden, Switzerland, Taiwan, and United Kingdom.No product was distributed to US government accounts. | Open, Classified |  |  | 47 | Not_Computer
Z-2639-2016 | 74013 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision RFD 3D Mobile C-arm for Mobile Fluoroscopic Imaging |  |  | April 15, 2016 | 2016 | Ziehm Imaging Inc | Malfunction of the radiation-warning timer and the 3D-specific interval warning. No 5-minute timer dialogs or beeper will occur until the unit is rebooted. | Device Design | Ziehm Imaging, Inc planned action to bring these defects into compliance: 1. You will contact customers and initiate a software update to correct the defects.2. You will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.For further questions, please call (866) 949-4346. |  | 31 | 31 devices | US Distribution including Puerto Rico and to the states of :FL, GA, VA,MA. ME, MI, NY, NJ, PA, RI, IL, TX, AZ, CA, NV, UT, and WA. | Open, Classified |  |  | 62 | Computer
Z-2646-2016 | 74790 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BVX-330BR Bovie Precise 360 Handpiece 33cm, blade |  |  | July 18, 2016 | 2016 | Bovie Medical Corporation | The adhesive may be inadequately applied or the assembly may be manipulated after application of the adhesive resulting in an incomplete bond to the angled tip. | Nonconforming Material/Component | Consignees were notified by written recall notice sent via certified mail or courier on July 18, 2016 to quarantine their inventory of the Precise 360 Handpiece, identify and contact their customers that have received the affected product and provide the a copy of the attached field safety notice. Customers are requested to return affected product, and return all affected product currently in their possession. All customer inventory should be reviewed, discontinue use, and return all Bovie Precise 360 handpieces to: Quality Assurance Department, Bovie Medical Corporation, 5115 Ulmerton Road, Clearwater, Florida 33760, Attn: Precise 360 Recall. Also, please write the authorization number RMA XXXX on the outside of the shipping container and use UPS account number 342562 and return via ground shipping and Fax or email the attached recall return response form to the email or fax indicated on the form as soon as possible and include a copy with the shipment. A completed recall return response form is required even if the customer has no product to return. If you have any questions about this recall or the product affected by the recall, call 727-803-8617. |  | 95 | 95 | Worldwide Distribution - US Nationwide in the states of: MD, GA, MO and FL and the country of Finland. | Open, Classified |  |  | 95 | Not_Computer
Z-2658-2016 | 74820 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core128, Model number 728323 |  |  | July 20, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | After upgrading to 3.6.7 software version via FCO72800643, during reconstruction of gated helical scans, the planned anatomy can be cut off on the CT images. This issue can occur on gated CT helical reconstruction that is not planned at iso-center (0,0). If the region of interest is not visualized in the images, a CT rescan may be performed. | Software Design Change | Customers were sent a letter dated July 28, 2016 informing them of the problem, under what circumstances it can occur and the actions that the customer can take to minimize the effect of the problem. The letter also includes the actions planned by Philips to correct the problem. Philips plans on releasing a field change order (FCO) to correct the affected systems.If you need any further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. ForNorth America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377). |  | 3 | 3 | US Nationwide Distribution in the states of: AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,NC,ND,NE,NJ,NM,NV,NY,OH,OK,OR,PA,PR,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY | Open, Classified |  |  | 615 | Computer
Z-2650-2016 | 74923 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic, Viva Quadripolar Implantable Cardioverter Defibrillators: |  |  | August 12, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | 78 Viva CRT -Ds and Evera ICDs that were manufactured with a specific subset of circuit components may experience rapid battery depletion due to a low resistance path developing within the circuit component. There have been seven (7) confirmed failures (9%) through July 16, 2016. Reported complications have included shortness of breath, pocket heating, low heart rate, and early device explant. | Nonconforming Material/Component | Consignees were sent or hand delivered on August 12, 2016 a Medtronic "Urgent Medical Device Correction" letter dated August 2016. The letter was addressed to Physician. The letter described the problem and the devices involved in the recall. The letter described a number of "Patient Management Recommendations". For questions contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636. |  | 78 | 78 (39 US, 39 OUS) | Worldwide Distribution - US Nationwide in the states of US: AZ, CA, FL, GA, IL, KY, MD, MN, NJ, NY, NC, OH, OK, PA, PR, TX. and countries of: Australia, Finland, France, Germany, Iran, Italy, Japan, Korea, Netherlands, Puerto Rico, Saudi Arabia, Spain, and United Kingdom. | Open, Classified |  |  | 78 | Not_Computer
Z-2660-2016 | 74806 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite F50 instrument. |  |  | July 25, 2016 | 2016 | Tecan US, Inc. | The batteries have the potential to overheating, melt or char. | Nonconforming Material/Component | Tecan contacted all customers by Safety Announcement, via certified mail starting July 25, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Because affected batteries pose a fire and burn hazard, it is extremely important that customers check whether their battery is affected. Customers were advised to go immediately to the HP Battery Recall website at: http://www.HP.com/go/batteryprogram2016 to check if their battery is affected by this recall. Customers were also advised to complete and return the Customer Response Form at QA@tecan.com or fax it to 919-287-2961.Customers with questions should contact HP at 1-888-202-4320 or Tecan Technical Service's Helpdesk at 800-832-2687, option 1. |  | 1 | 1 unit | Worldwide Distribution - US including AK, AL, AR, CA, FL, GA, HI, KY, MA, MD, NC, NY, PA, PR, SC, TN, TX, UT, VA. International account: Canada. | Open, Classified |  |  | 65 | Not_Computer
Z-2659-2016 | 74937 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NovaPACS versions 8.4.7, 8.5.3, and 8.5.6. Distributed by Novarad Corporation. Picture archiving and communication system (PACS) |  |  | August 12, 2016 | 2016 | Novarad Corporation | Novarad Corporation announces a voluntary field action for the NovaPACS versions 8.4.7, 8.5.3, and 8.5.6 due to a "pop-up screen" that was prematurely released. | Software design | Novarad sent an Urgent Medical Device Correction letter dated August 12, 2016, to all affected consignees. As a workaround, customers are advised to use the HTMLS Viewer for review purposes only and not for diagnostic purposes. The issue has been fixed in all future releases of NovaPACS software. You can eliminate the possibility of incorrect information appearing by contacting Novarad Technical Supoort to arrange for an upgrade to a corrected version of NovaPACS software. This correction is available to all affected customers at no cost and will be included din all future upgrades.Customers will need to fax, mail, or email th eincluded Acknowledgement and Receipt Form back to Novarad. Customers with questions or concerens regarding this notification or would like to request an upgrade, should contact Novarad Technical Support at 801-221-5895.For questions regarding this recall call 801-642-1001. |  | 90 | 90 | Worldwide Distribution - US (Nationwide) Canada and other foreign consignees. | Open, Classified |  |  | 90 | Computer
Z-2537-2016 | 74561 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Roadrunner Uniglide Hydrophilic Wire Guide, Catalog Numbers and GPN: HPWA-18-260, G30489; HPWA-18-320, G30490; HPWA-25-260, G30491; HPWA-35-260, G30492; HPWAS-35-260, G30506; HPW-35-150, G56149; HPW-35-180, G56150; HPWS-35-150, G56152; HPWS-35-180, G56153; HPWA-35-80, G56172; HPWA-35-150, G56173; HPWA-35-180, G56174; HPWA-18-180, G56160; HPWAS-35-150, G56176; HPWAS-35-180, G56177. |  |  | June 24, 2016 | 2016 | Cook Inc. | Potential coating contamination with glass particles. | Material/Component Contamination | Cook Inc. initiated a voluntary recall of the Roadrunner UniGlide Hydrophilic Wire Guides on June 24, 2016 via courier due to the potential for the device to be contaminated with glass particles. The firm issued a press release dated August 12, 2016. Customers were asked to take the following actions: 1. Examine inventory immediately to determine if you have affected products and quarantine affected products. 2. Return affected products to Stericycle Expert Solutions (a third-party recall administration service provider). Use the enclosed label and a copy of the Acknowledgement and Receipt Form to receive a product credit.3. Even if you do not have affected products on hand, you must still complete the Acknowledgement and Receipt Form and fax it to 866.796.4780 or email it to cookmedical4502@stericycle.com.4. Report adverse events to Cook Medical Customer Relations at 800.457.4500 or 812.339.2235. Monday through Friday between 7:30 a.m. and 5:00 p.m. (Eastern Time). Or by email at CustomerRelationsNA@cookmedical.com.Should you have any medical questions or concerns, please contact Cook Medical Customer Relations at 800.457.4500 or 812.339.2235. For information regarding the recall, please contact Stericycle Expert Solutions at 866.912.9552.Cook is providing notification to distributors and informing them that this notice must be shared with appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred. |  | 8750 | 8,750 devices were distributed in commerce | Worldwide distribution. US (nationwide), Australia, Austria, Belgium, Canada, China, Czech Republic, France, Germany, Hong Kong, Hungary, Italy, Jordan, Korea, Spain, South Africa, Taiwan, Turkey, and United Kingdom | Open, Classified |  |  | 8750 | Not_Computer
Z-2671-2016 | 64796 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Centricity Laboratory Core Lab System 4.1 |  |  | November 20, 2012 | 2012 | GE Healthcare It | The Centricity Laboratory removes free-text notes on a master panel when the user deletes the slave panel. The deletion of a slave panel will (auto slash) the corresponding test (item) results on the master panel. Additionally, the displayed and HL7 results show neither the appropriate abnormal flag value nor the reference range values, when using large_num.g &gt;7 to report numeric results. | Software design | GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated Nov 20, 2012 to all affected customers. The letter described the product, problems and actions to be taken. The customers were instructed to follow the safety instructions and to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI #86071.Customers were instructed to contact their GE Service Representative with any questions regarding this issue. For other questions, contact GE Helpdesk:North America; (905) 479-1249 or 1-888-778-3375; email: tgg.nahelpdesk@med.ge.comEurope; +44 1603 877595; email: tgg.ukhelpdesk@med.ge.comAsia Pacific; +61 3 5327 4000; email: tgg.ozhelpdesk@med.ge.com.For questions regarding this recall call 847-277-5000. |  | 17 | 17 | Worldwide Distribution - US Distribution to GA, OH, NY, TX, TN ,MN and Internationally to India, Australia, Canada, England, Qatar, and Scotland. | Open, Classified |  |  | 17 | Computer
Z-2678-2016 | 74927 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Geo-Med, TURP/CYSTO PACK PGYBK REF 89-7931.07, Rx Only, STERILE EO, |  |  | August 11, 2016 | 2016 | DeRoyal Industries Inc | The firm is recalling two custom surgical kits containing the BARD(TM) Toomey Irrigation Syringe 70cc with Resectoscope Tip, Catheter Tip, Luer Adapter and Cap. The BARD(TM) component is being recalled for a possible slit defect on the package that may affect the component's sterility. | Nonconforming Material/Component | DeRoyal issued their recall on 08/11/2016 to the end-user level by sending overnight letters to its direct accounts. The customers were asked to identify the affected trays, place them in quarantine and return them to DeRoyal for replacement or credit. |  | 26 | 26 units | Distributed in LA and TN. | Terminated | June 13, 2017 | 306 | 50 | Not_Computer
Z-2670-2016 | 74616 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | UNIQ 1.0: UNIQ FD10, 722026; UNIQ FD10/10, 722027; UNIQ FD20, 722028; UNIQ FD20/10 biplane, 722029; UNIQ FD10 OR Table, 722033; UNIQ FD10/10 OR Table, 722034; UNIQ FD20 OR Table, 722035; UNIQ FD20/20, 722038; |  |  | July 07, 2016 | 2016 | Philips Electronics North America Corporation | Increase in the failure rate of certain low-voltage DC power supplies (DCPS) | Component change control | Philips sent an Urgent - Field Safety Notice dated June 20, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Philips Field Service Engineers will replace the affected DCPS free of charge.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377.For further questions regarding this recall please call (978) 659-3000. |  | 1536 | 1536 devices (includes both Allura and UNIQ) | Nationwide, Canada, AlgeriaArgentinaAustraliaAustriaBahrainBangladeshBelgiumBrazilBrunei DarussalamBulgariaCHILEChinaColombiaCroatiaCubaCyprusCzech RepublicDenmarkDominican RepublicEcuadorEgyptFinlandFranceFrench PolynesiaGeorgiaGermanyGreeceHondurasHong KongHungaryIcelandIndiaIndonesiaIranIrelandIsraelItalyJapanJordanKazakhstanKenyaKorea, Republic ofLebanonLithuaniaMacedoniaMalaysiaMaltaMexicoMoroccoMyanmarNepalNetherlandsNew CaledoniaNorwayOmanPakistanPalestine, State ofPERUPhilippinesPolandPortugalQatarRÂżunionRomaniaRUSSIAN FEDERATRussian FederationSaudi ArabiaSingaporeSlovakiaSouth AfricaSpainSri LankaSwedenSwitzerlandTaiwanTanzaniaThailandTunisiaTurkeyUkraineUnited Arab EmiratesUnited KingdomUruguayUzbekistanViet Nam | Open, Classified |  |  | 1536 | Not_Computer
Z-2673-2016 | 74666 | No | ELC | Ultrasonic scaler. | Dental | Dental | 510(k) | Diamond Coated Tips and Burs are accessories to the air powered Bone drill (VARIOSURG with and without light) |  |  | July 08, 2016 | 2016 | Nakanishi Inc. | Through an error, the labeling failed to include the symbol or other text indicating that the diamond coated tips were single use only. A contributing factor to this error is that in other markets such as Japan, the diamond coated tips are multiple use and are not limited to single use only. The tips are stainless steel with a coating of fine diamond powder. The tips are sold non-sterile and are to be sterilized prior to first use. The tips were cleared as single use only.This problem was discovered during an ongoing, proactive, quality system improvement plan that was initiated by Nakanishi in October 2014. Because the correct instructions for sterilization are not available, it is possible that the devices might not be sterile prior to use. The absence of the single use statement introduces a risk of cross-contamination if the tips are reprocessed and reused. | Error in labeling | Naknishi sent out an URGENT : MEDICAL DEVICE CORRECTION letter on July 18, 2016 describing the potential safety issue regarding their Diamond Coated Products. The firm is requesting customers take the following actions: 1. Examine their inventory immediately to determine if they have any product, both used and unused, bearing any of the Nakanishi Product Codes listed in List of Products enclosed with this letter.2. Immediately dispose of any USED product according to applicable regulations.3. Segregate any UNUSED product for return to Nakanishi.4. In order for the firm to complete their records, they are requesting that customers complete the enclosed Correction Activity Sheet and FAX or Mail to: NSK America, Corp.1800 Global Parkway, Hoffman Estates, IL 60192 US Facsimile: 1-800-838-9328Customers with questions were advised to Telephone: 1-888-675-1675.NSKrecall@NSKDental.comFor questions regarding this recall call 847-843-7664. |  | 2574 | 2574 units | Nationwide Distribution to Illinois, Florida and Georgia. | Open, Classified |  |  | 37577 | Not_Computer
Z-2677-2016 | 74811 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | CURVED-TIP STAPLER 30,IS4000; Model number 470530; |  |  | July 14, 2016 | 2016 | Intuitive Surgical, Inc. | Field failures were reported related to the da Vinci Xi Surgical System EndoWrist Stapler 45 and 30instruments having an interruption in shaft rotation and / or instrument jaws not opening. | Nonconforming Material/Component | Intuitive sent a Stop Use e-mail to all affected sites on July 14, 2016 to inform users of the issue prior to the commencement of a recall. E-mail described the problem and asked that all surgeons who use the stapler be informed to immediately stop use. Formal letters will follow. Medical Device Recall letters were sent on July 21, 2016. Customers were instructed to return quarantined affected product whether used or unused to Intuitive Surgical. Customers should contact customer service to initiate the standard Return Material Authorization (RMA) process by phone or email:North and South America: 800-876-1310, option 3 (6AM to 5PM PST)Email: customersupport..servicesupport@intusurg.comEurope, Middle East, Asia and AfricaPhone: 800-821-2020 or +41 21 821 2020Email: ics@intusurg.comCustomers were also instructed to complete the attached Return Acknowledgement Form and submit via email to Regulatory Compliance at isi.compliance@intusurg.com or by fax to +1 (408) 523-0619.For questions regarding this recall call 408-523-2100.10-31-16 Update: Monthly status update states that an additional 93 instruments have been identified and an additional communication will be sent to these sites in Mid November 2016, |  | 875 | 875 instruments (827 a la carte + 48 in kits); total, all models | Worldwide Distribution - US (nationwide) Internationally to Australia, Belgium, Denmark, France, Germany. Italy. Japan, Spain, Sweden, Switzerland, Taiwan, and Turkey, | Terminated | January 03, 2017 | 173 | 875 | Not_Computer
Z-2680-2016 | 75008 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | NIO-A / NIO-P is an instant IO access device comprises housing with spring load mechanism that injects a trocar needle into the bone marrow of a patient and allows connection of syringe or IV administration set. |  |  | August 18, 2016 | 2016 | Waismed Ltd | WaisMed received a report in which during the operation of the NIO-A, the needle was not released from the device as expected. The device was placed on the floor and after several minutes the needle was released spontaneously. NIO-A and the NIO-P modelshave the same activation mechanism. However, none of the affected lots of the NIO-A are distributed in the US, and thus the company is only recalling a single lot of NIO-P in the US. | Component design/selection | The firm emailed the US Consignees on August 18, 2016 that they were conducting a recall on the NIO/A and NIO/P products. The US distributor shall contact the customers and will collect the products from each of the customers. The recalled products shall be moved to Storage quarantine within one week. The company shall keep track on the process and the recall efficacy. |  | 50 | 50 | WI, TX, CA, LA and the following countries: UK, Spain, Poland, Israel, Czech Republic, France, Belgium, Denmark, Italy, Cyprus, Austria, Greece, Slovenia, Australia, Germany, Estonia, New Zealand, Switzerland, Costa Rica, Guatemala, Canada, Malaysia, Poland, Singapore, Lithuania, Romania | Open, Classified |  |  | 50 | Not_Computer
Z-2687-2016 | 63898 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Non-DEHP Micro-Volume Extension Set, Male Luer Lock Adapter, 36 (91 cm), Volume 0.32 mL; an Rx sterile, single use nonpyrogenic fluid pathway; 50 sets per case; |  |  | December 14, 2012 | 2012 | Baxter Healthcare Corp. | Some of the individual extension set packages were improperly sealed, compromising the sterility of the product. | Unknown/Undetermined by firm | Baxter sent an Urgent Product Recall notification letter dated December 14, 2012, to all affected customers. Baxter instructed customers to locate and remove all affected product from their facility. Customers were instructed to contact Baxter Healthcare Center for Service to arrange for return of product and credit and to complete the attached Customer Reply Form and return it to Baxter or scanned email. Customers with questions should call 1-888-229-0001. For questions regarding this recall call 800-422-9837. |  | 15257 | 15,257 units | Worldwide Distribution - US including FL, GA, IL, KS, KY, MI, MO, NC, ND, OH, PA, SC, TN, TX, WV, Puerto Rico and Internationally to Canada. | Completed |  |  | 15257 | Not_Computer
Z-2685-2016 | 74924 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Interventional Radiology (IR) images are stored as JPEG2k Lossless in Merge Enterprise Archive (EA) and are not displaying correctly in RadSuite. | Software design | The firm, Merge, sent an "URGENT: MEDICAL DEVICE RECALL" letters dated 1/29/2016 and 5/25/2016 via e-mail or via certified mail if they did not have an e-mail address for the customers. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: you must discontinue using the affected version of RadSuite or iConnect Enterprise Archive with GE AW Workstation; use the workaround listed in the letter; complete and return the enclosed form no later than June 17, 2016; ensure that all users of the product are provided with the notification; if you have further distributed this product, please identify your customers and notify them at once of this product recall Merge has a released fix available for this issue.If you have any additional questions, please send an email to recall@merge.com or call 800-724-5970 ext 8027. |  | 96 | 96 sites potentially have the affected versions for both products | US Distribution to: CO and OK. | Open, Classified |  |  | 96 | Not_Computer
Z-2681-2016 | 74920 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Compano Basic Reader using 100-127V as mains voltage. Model numbers: 732027, 732028, and 732040. |  |  | August 12, 2016 | 2016 | Philips Electronics North America Corporation | A defect component of the Compano reader power supply can cause the power supply to burn and some smoke to leave the housing. | Component design/selection | Philips will send a Field Safety Notice to all affected customers on August 12, 2016. The notice informs consignees of the issue and informs them that Philips will replace the power supply unit of the PCR Compano Basic reader system of all affected units free of charge (FCO73200037). |  | 388 | 388 | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC,ND, NH, NJ, NM, NV, NY, OH, OR, PA, TX, VT, WA, WI, WV, and WY; and, the countries of Canada, Brazil, Germany, Iraq, Japan, Korea, Kuwait, Libya, Mexico, Portugal, and Qatar. | Open, Classified |  |  | 388 | Not_Computer
Z-2684-2016 | 74867 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | INTEGRA Spinal Tray, Catalogue No: 3404875, Contents: 10 Each per Case, Sterile. |  |  | August 04, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is recalling Spinal Trays, Cat No. 3404875, because of the affected recalled Hospira 5% Lidocaine HCL that is a component of some lots of theSpinal Trays. | Other | Integra sent an Urgent Medical Device letter dated August 4, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to return unused, affected product. Confirmation of the recall notification delivery via delivery verifiable email; Return of Recall Acknowledgement form; Return &amp; disposition of unused, affected product, if the customer in their response indicates they have affected devices. Any products returned to Integra as a result of the recall will be stored in a segregated, labeled area. Recalled products that are returned will be dispositioned as scrap and destroyed. Any Lidocaine returned with the products will be dispositioned as directed in the Hospira Inc. Drug Recall Notification.If you have questions of a customer service nature or beyond the information in the Hospira Inc. URGENT DRUG PRODUCT RECALL NOTICE, please contact either your local Integra Pain Management Sales Representative or customer service at 1-800-241-2210.For further questions, please call (801) 886-9505. |  | 200 | 200 packages (20 cases) | US Distribution to the states of : OH and CA. No foreign, Govt/VA/military consignee. | Open, Classified |  |  | 200 | Not_Computer
Z-2683-2016 | 74793 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET Servo Humidifier 163, Inmed Mfg. Sdn. Bhd. Lot PT 2577, Jalan Perusahaan, 4, 34600 Kamunting, Perak, Malaysia. |  |  | December 03, 2015 | 2015 | Teleflex Medical | Cracking may occur at the 15mm and 22mm ISO femaleconnector. | Device Design | INMED sent an Urgent - Field Safety Notice dated December 3, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to cease use and distribution of stock of the affected product batch and quarantine immediately. Customers are to return all affected product to MAQUET. An Acknowledgement form was included to be completed and returned to their MAQUET representative. If the product was distributed, they were instructed to provide the field safety notice to all of their customers who have received product in scope of the Field Action. Their customer is then required to complete the Acknowledgement form and return it to them.Should you require any further information or support concerning this issue, please contact your local MAQUET representative at fieldactions@maquet.com.INMED is committed to providing high quality, safe and effective products. W e sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service. For further questions, please call (610) 378-0131. |  | 2700 | 2,700 units | US Distribution to the states of : FL, KS, and WV. | Completed |  |  | 2700 | Not_Computer
Z-2690-2016 | 74525 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | FUJIFILM Digital Mammography System Aspire HD (FDR MS-1000) with Biopsy Positioner (FDR-1000BPY) |  |  | June 09, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | If an exposure is aborted during a stereo biopsy workflow, an error may occur which would prevent the system from continuing to operate properly. | Software design | The firm, Fujifilm, sent a "FIELD SAFETY NOTICE" letter dated June 9, 2016 via Federal Express to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to read and follow the instructions in the letter; call your local service representative, in case an error occurs, and complete and return the Customer Feedback Form via fax (203)-251-7862. Both affected systems will be corrected via on-site visits by FMSU Customer Service Engineers.If you have any questions about this matter, please contact the Product Manager, Women's Health directly at (303)-888-4549 or call your local FUJIFILM office. |  | 2 | 2 Units | US Distribution to: NC and HI. | Open, Classified |  |  | 2 | Not_Computer
Z-2704-2016 | 74830 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry XPT, SMN 10723034, IVD. --- This issue affects only the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results when HbA1c values are reported in International Federation of Clinical Chemistry (IFCC) equivalent units (HbA1cR). |  |  | July 20, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens identified an issue with the ratio equation provided on the ADVIA Chemistry XPT System Software Test Definition (TDef) Version 1.0 disks (SMN: 11127343 and 11222123). This issue affects only the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results when HbA1c values are reported in International Federation of Clinical Chemistry (IFCC) equivalent units (HbA1cR). The ratio equation used to calculate ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment (A1c_3) results in IFCC units contains an error. The error results in falsely depressed HbA1cR results. | Software design | Siemens sent an Urgent Field Safety Notice dated July 20, 2016, to all affected customers via Fed Ex and the Urgent Field Safety Notice (CHSW16-02 July 2016) was e-mailed to regional countries for implementation Outside US (OUS). Customers have been advised to stop reporting results and call Siemens immediately to schedule service. The file with the typographical error is located on ADVIA Chemistry XPT Test Definition (TDef) disk. -- Field service personnel were sent a support bulletin describing the issue and instructing them how to correct the equation. Service engineers have been instructed to remove all version 1.0 TDef disks from customer sites regardless of whether they run the ADVIA Chemistry Hemoglobin A1c_3 Automated Pretreatment assay and report in IFCC units.Customers with questions were advised to contact their Siemens Customer Care Center or their local Siemens technical support representative. |  | 229 | Domestic: 32 units; Foreign: 197 units | Worldwide Distribution - US to MA, NY, MI, IL, TX, CO, AZ, CA, WA and Foreign distribution to the following countries: Australia, Austria, Bahrain, Brazil, China, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Italy, Malaysia, Netherlands, New Zealand, Norway, Portugal, Portugal, Republic Korea, Singapore, Slovakia, Spain, Sweden, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 229 | Computer
Z-2691-2016 | 74967 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | HexaPOD evo RT System |  |  | August 18, 2016 | 2016 | Elekta, Inc. | Potentially unrecognized incorrect position of the treatment couch in 3D workflow, i.e. the HexaPOD has not moved fully to the 3D position. | Nonconforming Material/Component | Elekta sent an Important User Notice 618-03-303-025 to affected customers on 8/18/2016. The notice identified the affected product, problem, clinical impact, solution and actions to be taken. The Notice reminds users that the Interlock Check is not possible during 3D Workflow as iGUIDE is not in use. The notice reminds users that if iGUIDE is not being used - (A) Ensure the HexaPOD is at its pre-defined 3D position before treatment, and (B) Ensure the External Inhibit LED at the EnableSwitch Board is off. For questions contact your local Elekta office. |  | 13 | 13 | Worldwide Distribution - US Nationwide in the states of LA, PA, WA, and the countries of: Australia, Denmark, France, India, Italy, and Japan. | Terminated | May 23, 2017 | 278 | 13 | Not_Computer
Z-2688-2016 | 74064 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Invivo Sentinelle Venguard Breast MRI Auxiliary Table/Tabletop with 8/16 Channel Coil Array for GE 1.5T/3T MRI Systems. |  |  | May 06, 2016 | 2016 | Invivo Corporation | A diode was inadvertently not assembled into the system, which may result in coil overheating | Under Investigation by firm |  |  | 3 | 3 | WI | Completed |  |  | 3 | Not_Computer
Z-2689-2016 | 74758 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | A7 Anesthesia Delivery System, a device used to administer to a patient, continuously or intermittently, a general inhalation anesthetic and to maintain a patient's ventilation. Part numbers 0632F-PA0000X (US domestic) and 0632B-00014 (international) |  |  | June 15, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Potential for a leak to occur on the back-up O2 and air e-size cylinder yokes on the A7 Anesthesia Delivery System. | Nonconforming Material/Component | Mindray sent via certified mail with return receipt a recall letter dated June 15, 2016 to their affected customers. Mindray will replace the gasket on all e-size cylinder yokes on the affected A7 systems. Customer was instructed to contact their local Mindray Service Representative to arrange for this replacement. Customers can continue to use their A7 system while awaiting the replacement of the gasket(s). Customers can contact Ms. Diane Arpino, Director, Quality Operations and Regulatory Affairs via email to: d.arpino@mindray.com or via telephone to: (201)995-8407 |  | 167 | 167 units (165 units - US) and (2 units - International) | US Nationwide Distribution to AL, MA, MD, MN, MO, NE, NJ, OK, OR, PA and VA; and Canada | Open, Classified |  |  | 167 | Not_Computer
Z-2705-2016 | 74190 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Double Pump RF Patient Cassette; Product Number: 72204055. |  |  | May 10, 2016 | 2016 | Medical Vision Ab | Sterility of device may be compromised due to sterile package breakage. | Unknown/Undetermined by firm | The An Urgent Field Safety Notice: Medical Device Field Safety Corrective Action / Recall letter was sent via certified mail on May 10, 2016. The letter informs customers of the issue and asks them to take the following actions: 1. Locate and quarantine affected unused devices immediately; 2. Return quarantined product to your national Smith &amp; Nephew agency/distributor. 3. Complete the return slip and fax it to your national Smith &amp; Nephew agency/distributor. 4. Please make sure this safety information is passed on to all those who need to be aware of it within your organization. 5. Please maintain awareness on this notice and resulting action until the Field Safety Corrective Action is terminated to ensure effectiveness of this action. |  | 3295 | 3,295 devices | Distributed to the US, to the states of NH, OK, AL, PA, IL, MA, TX, MI, NC, CA, MN, CT, GA, NY, WI, KY, OH, VA, AL, NJ, TN, SD, and IN. | Open, Classified |  |  | 3295 | Not_Computer
Z-2692-2016 | 74875 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System, Catalog No. 775222 |  |  | August 04, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System to reduce the potential for flammable liquids coming into contact with electrical components in the Stainer module possibly causing a fire. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device recall letter dated August 4, 2016, to all affected customers to inform them that the UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System to reduce the potential for flammable liquids coming into contact with electrical components in the Stainer module possibly causing a fire. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the recall notice were instructed to contact Beckman Coulter Customer Support Center:http://www.beckmancoulter.com/customersupport/support; 800-526-7694 in the United States and Canada; and outside the United States and Canada, are instructed to contact their local Beckman Coulter Representative. |  | 759 | 759 units total (358 units in US) | Worldwide Distribution to US (nationwide) and Internationally to Australia, Belgium, Canada, China, Croatia, Czech Republic, France, Germany, Hong Kong, India, Israel, Italy, Japan, Kuwait, Malaysia, Mayotte, Mexico, New Zealand, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Slovakia, South Korea, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, and Vietnam. | Open, Classified |  |  | 759 | Not_Computer
Z-2707-2016 | 74827 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SMR allen wrench 5 mm, product code 9013.50.210 |  |  | July 21, 2016 | 2016 | Limacorporate S.p.A | Product found to be dimensionally non-conforming (slightly over dimensioned on the diameter of the Zimmer connection) and non-functional for connecting with some T-handles. | Nonconforming Material/Component | The firm received two complaints from the US field about the inta-operative malfunction with the SMR allen wrench 5mm (product code 9013.50.210) with the T-handle. Both parts involved were manufactured with the same lot #14AA311. No adverse effects for patients reported. |  | 76 | 76 | US, Italy, Australia | Terminated | January 13, 2017 | 176 | 76 | Not_Computer
Z-2706-2016 | 74856 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Terumo Custom Cardiovascular System Cardiovascular Procedure Kits labeled containing Coated Tubing as follows: |  |  | July 25, 2016 | 2016 | Terumo Cardiovascular Systems Corp | Terumo Custom Cardiovascular Procedure Kits are labeled as Xcoating surface coated tubing, but tubing is non-coated | Process control | Terumo Cardiovascular Systems issued the Safety Alert dated July 25, 2016 and advises the user that their kit may include uncoated tubing, and further advises the consignee to continue using their kits with consideration of anticoagulation protocols during a procedure. Each consignee is asked to complete a Customer Response Form, and to return the Customer Response Form to Terumo acknowledging their receipt of the form. Consignees can return their affected product for credit/replacement; however, return of the product from the field is a customer option and is not under recommendation from Terumo Cardiovascular Systems. For further questions, please call (800) 262-3304. |  | 2207 | 2207 kits | NationwideForeign: Canada, Latin America | Open, Classified |  |  | 2207 | Not_Computer
Z-2682-2016 | 74685 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DuraDiagnost stationary X-ray system |  |  | September 17, 2015 | 2015 | Philips Electronics North America Corporation | Buttons or keys on the monitor allow the user to turn the volume setting down to 0", in which case the beep will not be heard at the conclusion of an exposure, as required 21 CFR 1020.31(j), and the Height-Adjustable Tabletop is missing the required labels [21 CFR 1010.2., 21 CFR 1010.3(a), and 21 CFR 1020.30(3)]. | Radiation Control for Health and Safety Act | Philips plans to bring the products into compliance with the Federal standard at no cost to customers by implementing a field correction.1. You will contact customers and initiate the actions to correct the failure to comply issues.2. You will attach the warning label to maintain an audible volume, and issue a supplement to the user manual that instructs the user how to adjust the volume for the local noise environment and include a warning that the user shall maintain an audible volume.3. You will attach the identification label to the affected Height-AdjustableTabletop.4. The customer notification letter which includes a statement that you will without charge, remedy the defect or bring the product into compliance.If you need any further information or support concerning this, please contact your local Philips representative or Philips Customer Service: 1-800-722-9377. |  | 15 | 15 | US Distribution to the states of : ME, NE, OR, CA, TX, NC, WA and OH. | Open, Classified |  |  | 15 | Not_Computer
Z-2711-2016 | 74900 | No | NQE | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | SPECULAR MICROSCOPE CEM-530; Software version 1.08 and 1.09. |  |  | August 08, 2016 | 2016 | Nidek Inc | Software version 1.08 and 1.09 for the Specular Microscope CEM 530 included a change of analysis results feature that was not reviewed and approved by the FDA. | Software change control | The firm, Nidek, sent a Dear Doctor letters entitled "RECALL NOTIFICATION" dated August 15, 2016 to affected sites by Federal Express. The letter informs users of the issue and stated that Nidek field service representatives will visit to install updated software.If you have any questions, please call Customer Service at 1-800-722-0219; Monday-Friday 8am to 5pm PST or email: cem530update@noritsuservice.com. |  | 29 | 29 affected devices | US Distribution to states of: TX, KS, FL, NV, CA, PA, KY, OR, TN, AL, VA, MO, OK, CT, NY. LA, NM, and IN. | Terminated | January 09, 2017 | 154 | 29 | Computer
Z-2712-2016 | 74961 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | August 18, 2016 | 2016 | Elekta, Inc. | In a specific workflow where contours are edited (enlarged, moved, copied, etc.) and then the contour is deleted on some slices and saved in the same session, it is possible that the deleted contours are still present which would not be the intent of the planner. | Unknown/Undetermined by firm | Important User Notice 382-03-MON-003 was sent to customers on August 18, 2016. The notice informs users of the specific product and version numbers affected by the issue, and any work arounds that can be used to avoid the issue. The letter also states that the problem will be in a future Monaco release. The customers were instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta as soon as possible but no longer than 30 days from receipt. Questions can be directed to the local Elekta office. |  | 3012 | 3,012 units | Worldwide Distribution -- USA, to the states of AK, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MO, MT, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, Puerto Rico, Algeria, Antigua and Barbuda, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Laos, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Nambia, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 3012 | Not_Computer
Z-2709-2016 | 75027 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. The firm name on the label is Merge Healthcare, Hartland, WI. Image processing system. |  |  | February 09, 2016 | 2016 | Merge Healthcare, Inc. | A report can be confirmed with the incorrect patient demographics resulting in it being saved under the incorrect patient record in the Electronic Medical Record (EMR). | Software design | MERGE sent an Urgent Medical Device Recall letter dated February 9, 2016, to all affected customers via email or via certified mail if they did not have an e-mail address for the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.For further questions, please call (800) 724-5970 ext. 8027. |  | 116 | 116 customers potentially have the affected versions | Nationwide Distribution | Open, Classified |  |  | 116 | Not_Computer
Z-2710-2016 | 74909 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Push Button Blood Collection Set, Sterile, Rx only, 0.6 x 19 mm x 305 mm, 23G x 3/4" x 12", REF 367342. |  |  | August 08, 2016 | 2016 | Becton Dickinson &amp; Company | Becton Dickinson is recalling various lot numbers of BD Vacutainer Push Button Blood Collection Set because of reports that specific lots exhibit separation of fron and rear barrels upon activation of the safety feature, which retracts the needle. | Under Investigation by firm | BD notified their customers via e-mail and sent an "Urgent Product Recall Letter" and a Recall Response Form dated August 8, 2016 via UPS. |  | 1788150 | 1,788,150 units | US and International | Open, Classified |  |  | 1788150 | Not_Computer
Z-2708-2016 | 74824 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Alcon AcrySof IQ IOL with ULTRASERT Delivery System |  |  | July 28, 2016 | 2016 | Alcon Research, Ltd. | The ULTRASERT Delivery System from certain lots have an inferior surface characteristic that could result in the IOL becoming lodged in the Delivery System | Process control | Alcon Research sent an Urgent Field Safety Notice letter dated July 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check their inventory and segregate any product being recalled. Customers were asked to complete and return the Field Safety Notice Response Form. For questions or concerns contact Alcon. |  | 371 | 371 units | Armenia, Belgium, Bulgaria, Croatia, Finland, France, Germany, Hungary, Italy, Lithuania, Malaysia, Netherlands, Norway, Poland, Portugal, Qatar, Romania, Slovenia, Spain, Sweden, Switzerland, Thailand, United Kingdom (Great Britain and Ireland), and United Arab Emirates. *** While this product is available in the US, these lots were not distributed to US consignees *** | Terminated | June 07, 2017 | 314 | 371 | Not_Computer
Z-2713-2016 | 74706 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | The 9025TRU Arterial Blood Gas (ABG) Kit is used to obtain arterial blood sampling for pH, Blood Gas, Electrolyte and Metabolite Analysis. |  |  | June 27, 2016 | 2016 | Cardinal Health | Arterial Blood Gas (ABG) Kit- Missing Safety Component. The ABG Kit(s) have been identified as missing the locking clasp component. CareFusion/BD has identified a potential customer/user safety risk associated with the missing locking clasp component. | Under Investigation by firm | CareFusion/BD sent out a FIELD SAFETY NOTICE of the field correction in writing on June 27, 2016. The letter described the nature of the recall and provided actions to be performed. Carefusion/BD directed customers to take the following actions:Step #1: Inspect current inventory on-hand to identify the presence of the affected Arterial Blood Gas Sampling Kit(s). Step #2: Complete the Customer Response Form by checking all appropriate boxes and providing all accurate information as required. Step #3: Destroy all affected product(s) in-stock in accordance with their facility's destruction protocol. Step #4: Return the Customer Response Form to GMB-GLB-ALFieldActions@CareFusion.com Step #5: If the customers want to obtain credit and/or replacement of destroyed product(s) they are asked to call their CareFusion/BD Customer Support Department at 800.323.9088 (Option #1 ). For any additional questions and support concerning the voluntary Field Safety Corrective Action, customers are asked to call (847)362.9303. |  | 2709466 | Lot #884993 -16,000 units, Lot #903136 -1,900 units and Lot #903137 - 300 units. | Worldwide Distribution - US (nationwide) and Canada. | Open, Classified |  |  | 2709466 | Not_Computer
Z-2721-2016 | 75052 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Medical Linear Accelerator |  |  | July 08, 2016 | 2016 | Elekta, Inc. | There have been instances in hospitals where the joint between the drive gear and the gantry base drive wheel has become loose, leading to fatigue failure of some of the securing bolts. | Other | Important Field Safety Notice 200-01-103-081 was released July 8, 2016 for distribution to all customers with Digital Accelerators without XVI. It informs users of the specific product and version numbers affected by the issue. The modification involves an inspection and correction procedure. Should the inspection pass, the corrected in not required. All gantry base wheel bolts should be checked for presence and torque and should be replaced if found to be showing signs of fatigue. Elekta service is actively reaching out to affected customers in order to organize a time to implement the inspection and correction. The notice included an acknowledgment form to be returned. |  | 553 | 553 | AL, AZ, CA, FL, GA, KY, MI, MN, MO, NE, NJ, NY, NC, OH, OK, PA, SC, TX, WA, WV, Argentina, Austria, Australia, Bosnia and Herzegovinia, Bangladesh, Belgium, Bulgaria, Brazil, China, Columbia, Cuba, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Iran, Italy, Japan, Jordan, Lebanon, Libya, Morocco, Mexico, Nigeria, Netherlands, Norway, Pakistan, Panama, Peru, Philippines, Portugal, Serbia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Syria, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Vietnam, Yemen | Open, Classified |  |  | 553 | Not_Computer
Z-2731-2016 | 75081 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare SIGNA Creator, SIGNA Explorer. Diagnostic Imaging Device. |  |  | August 31, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare is conducting a recall dealing with an incorrect RF Power Monitor Function involving SIGNA CREATOR and SIGNA EXPLORER which could result in higher than expected thermal dose to the patient and higher than expected localized heating. There have been no failures of this type reported, and no injuries reported as a result of this issue. | Software design | Consignees were sent on 8/31/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 67921 dated August 31, 2016. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, and Radiologists.The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 383 | 383 (US = 24;OUS = 359) | US: AZ, CA, CT, FL, IL, MA, MS, MO, NC, PR, TX, WI.OUS: Algeria, Australia, Bangladesh, Belgium, Brazil, Canada, Colombia, Dominican republic, Egypt, France, Germany, Ghana, Guatemala, HONG KONG, Hungary, India, Indonesia, Iran, Italy, Jamaica, Japan, Korea, Kuwait, LEBANON, Lithuania, Maldives, Martinique, Mexico, MOROCCO, Netherlands, New Zealand, OMAN, Pakistan, Philippines, Poland, Portugal, Romania, SAUDI ARABIA, Singapore, South Africa, Spain,Sweden, Tanzania, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Open, Classified |  |  | 383 | Not_Computer
Z-2719-2016 | 75033 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ReliaMedÂż Silver Alginate/CMC. Each dressing is packed in a pouch (primary packaging), each pouch is packed in a carton (secondary packaging) holding multiple (depending on stock keeping unit) pouches and each carton contains one (1) Instructions For Use (IFU) leaflet. |  |  | January 15, 2015 | 2015 | Advanced Medical Solutions Ltd. | On 08 January 2015 AMS Sole Customer for the products notified AMS of the issue after discovering the misplaced IFUs during its review of packaging in its warehouse | Package design/selection | Advanced Medical Solutions' (AMS) Sole Customer has already received notification via telephone conference call, followed by emails and correspondence. |  | 47569 | 47,569 cartons | OH | Terminated | November 23, 2016 | 678 | 76773 | Not_Computer
Z-2732-2016 | 74912 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis System. |  |  | August 08, 2016 | 2016 | Beckman Coulter Inc. | There is a possibility of specimen misidentification for the stand-alone DxH 800 and DxH 600 where the Primary Identifier is configured to Tube Position ID. | Software design | An Urgent Medical Device Recall letter dated 8/8/16 was sent to all customers who purchased the Unicel DxH 800 and DxH 600 Coulter Cellular Analysis System to inform them that there is a possibility of specimen misidentification for the stand-alone DxH 800 and DxH 600 where the Primary Identifier is configured to Tube Position ID. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the notice are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; 800-526-7694 in the US and Canada; outside the US and Canada, contact local Beckman Coulter representative. |  | 5282 | 5,282 units total (2,877 units in US) | Affected geographies: Andorra, Australia, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Columbia, Croatia, Czech Republic, Egypt, Ecuador, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay and Vietnam. | Terminated | April 29, 2017 | 264 | 5282 | Not_Computer
Z-2715-2016 | 74853 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite software. Radiological image processing system. |  |  | April 09, 2015 | 2015 | Merge Healthcare, Inc. | When RadSuite is used with IPID (Issuer of Patient ID) as a part of the "Patient Identifier," it is possible in some circumstances that the demographics of one patient will be applied to a study or studies for another patient. | Software design | The recalling firm issued letters dated 4/3/2015 via email on 4/9/2015 notifying the customer of the issues. |  | 10 | 10 sites have the affected software that is configured with the specific conditions listed in Code Information | Distribution was made to medical facilities in AL, MI, MO, PA, TN, and TX. | Open, Classified |  |  | 10 | Not_Computer
Z-2717-2016 | 74975 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | NOMAD Pro2 Packaged X-Ray System, Part Number 08500021; NOMAD Pro2 Packaged X-Ray System Black, Part Number 08500022; NOMAD Pro2 Vet Packaged X-Ray System, Part Number 08500023; NOMAD Pro2 Vet Packaged X-Ray System Black, Part Number 08500024; NOMAD Pro2 Packaged X-Ray System, GBL, Part Number 08500025 and Part Number 08500026; NOMAD Pro2 Packaged X-Ray System, GLB, Vet, Part Number 08500027 and Part Number 08500028; NOMAD Pro2 Package X-Ray System, GLB, Black, Part Number 08500029 |  |  | July 18, 2016 | 2016 | Aribex Inc | Firm received customer complaints where end users reported that Nomad Pro and Pro 2 handsets had appeared to "overheat", "melt", "smoke" and/or "ignite." Investigation found that the internal lithium ion batteries had entered into a thermal runaway condition as a result of a short circuit on the battery pack PCB. | Other | The firm sent a "FIELD ACTION NOTICE" to all affected customers. |  | 16871 | 16,871 | AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA. HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY , Puerto Rico, Afghanistan, Argentina, Australia, Bahamas, Bahrain, Belgium, Bermuda, Chile, China, Columbia, Denmark, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Kuwait, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Saudi Arabia, Papua New Guinea, Philippines, Poland, Qatar, Saudi Arabia, Slovenia, South Africa, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, UAE, United Kingdom, Venezuela, Vietnam | Terminated | July 26, 2017 | 373 | 16871 | Not_Computer
Z-2714-2016 | 74862 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Acecide-C High Level Disinfectant &amp; Sterilant 875 mL. Packaging is 6 inner cartons in one master carton;1 inner carton consists of a set of two liquids (1 acid &amp; 1 buffer) banded together. |  |  | August 02, 2016 | 2016 | Best Sanitizers Inc | Certain lots of the acid solution in Acecide-C are degrading faster than normal. | Nonconforming Material/Component | Customer letters dated 8/2/16 will be sent. |  | 3023 | 3,023 cases USA; 666 cases shipped to Canada (OTC drug there) | Distributed to one (1) account in the state of PA.Distributed to one (1) account in Canada | Terminated | August 18, 2017 | 381 | 3023 | Not_Computer
Z-2733-2016 | 75028 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Cranial Access Kits, Catalogue Nos. INSHITH; INSHITHRZN; INSHITHND |  |  | August 24, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences has been notified of a medical device recall by Adaptive Surgical, LLC involving their Riverlon" Nylon 3-0 Sutures that are packaged in Integras Cranial Access Kits. It was found that product seals may become open during extreme distribution conditions and render the device unsterile. | Packaging | Consignees notified on August 12, 2016 via letter to identify and return product. |  | 710 | 710 | Nationwide and VA/govt/military. No foreign consignees. | Terminated | March 30, 2017 | 218 | 710 | Not_Computer
Z-2734-2016 | 74886 | [6 empty] | 2 over 2 Elevating Monitor Suspension - Omega Medical Imaging, LLC Elevating Monitor Suspension Systems. |  |  | June 21, 2016 | 2016 | Omega Medical Imaging, Inc. | Elevating Monitor Suspension separated from the lifting column bracket | Device Design | Consignees were notified by a Field Safety Advisory Notice - FSAN 16-0001 of this immediate Recall on 6/21/2016. |  | 26 | 26 units | CO, MN, GA, NC, IN, OH, KY, IL, NY, AL, MO, LA, PA, and SD Denmark | Open, Classified |  |  | 26 | Not_Computer
Z-2736-2016 | 74894 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Model 882482, gamma camera for Single Photon Emission Computed Tomography (SPECT) |  |  | May 25, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The firm became aware of an issue where during a non-Auto Body Contouring (ABC), non-circular, rel-180 scan, the detector may come in contact with the patient. The issue was caused by a software defect in the workflow setup script. | Software design | The firm is informing the consignee of this issue through Customer Information Letter. |  | 424 | 424 units | Worldwide distribution. US nationwide, Algeria, Argentina, Armenia, Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Finland France, Gabon, Germany, Greece, India, Indonesia, Ireland, Japan, Korea, Republic of Korea, Republic of Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Viet Nam. | Open, Classified |  |  | 544 | Computer
Z-2739-2016 | 64761 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Archive is a software product for receiving, archiving and sending of medical data. Qualified system administrators install, monitor and maintain the system. DICOM devices (e.g. modalities, workstations) communicate with the archive using the DICOM protocol (published by ACR-NEMA). The clinical use is limited to the interaction other systems have with the Centricity Enterprise Archive. Centricity Enterprise Archive is not intended for use in patient vicinity. The intended users for Centricity Enterprise Users service personnel and System administrators, and non clinical users. As of V4.0 XDS enabled systems can communicate with the archive using the XDS and XDS-I profiles (published by IHE). |  |  | November 16, 2012 | 2012 | Ge Healthcare It | There is a potential for data loss associated with empty path names resulting from HL-7 updates when the cache disk is almost full. In an unlikely event that the cache disk is almost full (&lt;100MB) and HL-7 update events are received, the update action fails and an empty path name is written to the EA database. As a result the instance is not deleted from the storage device but no longer referenced. When trying to retrieve the instance an error is returned. | Software design | URGENT MEDICAL DEVICE CORRECTION letters dated November 16, 2012 and April 19, 2013, were sent to all affected customers.The letters requested that customers verify that all Enterprise Archive partitions have enough free space available and that the cache disk has at least 10% free space available. The letters further requested that customers whose disks are not being purged to the configured watermarks and the cache disk is almost full to contact their local GE service person. A permanent product correction (Service Pack 12 and release notes) will be provided by GE Healthcare IT. Customers with questions can contact their GE service representative or the GE Healthcare IT Remote Online Center Customer Care Center at 1-800-437-1171. |  | 1379 | 1,379 units | Nationwide Distribution | Terminated | January 11, 2017 | 1517 | 1379 | Not_Computer
Z-2723-2016 | 64261 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Infusion Pump. List Number 11971. The Plum A+ Infusion Systems are a multi-purpose family of devices that provides precise delivery of multiple therapies across the general spectrum of clinical care. The Plum A+ family infusion pumps are self-contained, microprocessor based infusion devices. |  |  | February 06, 2013 | 2013 | Hospira Inc. | The fluid shield diaphragm for the Plum A+ Infusion pumps was undersized, possibly resulting in alarms and delay or interruption of therapy. | Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter, dated February 6, 2013, to customers. Customers were requested to return a reply form to Stericycle indicating that they have received the letter, and notified their users and/or customers. The customers were asked to supply the number of Plum A+ infusers at their facility. Product will be corrected by Hospira in the field. |  | 357778 | 357,778 total | United States nationwide distribution including PR, and USVI. Costa Rica, Turks and Caicos, and Dominican Republic. | Terminated | November 14, 2016 | 1377 | 357778 | Not_Computer
Z-2738-2016 | 73833 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PINNACLEÂż Introducer Sheath |  |  | March 16, 2016 | 2016 | Terumo Medical Corporation | Terumo Medical is initiating this recall due to complaints of valve leakage that have been related to a manufacturing issue. If affected introducer sheaths are used, there is possible risk of incremental blood loss through the valve. | Device Design | Terumo sent an Urgent Product Recall Bulletin dated March 7, 2016, to all affected customers. All affected customers will be notified regarding the required recall actions (stop using/quarantine the affected products). TMC will strive for 100% verification to confirm that all customers have acknowledged the recall. Follow-up communication is being sent specifying the mechanisms to be used for labeling and returning the involved product from the customer level.We regret any inconvenience this action may cause your operations and we appreciate your understanding. Please contact us if you need assistance procuring additional inventory. We are committed to continuing to offer products that meet the highest quality standards that you expect from Terumo Medical Corporation. If you have questions or need further information, please call Customer Care at 1-800-888-3786 or contact your local sales representative.For further questions regarding this recall please call (410) 392-7226. |  | 1370995 | 1,370,995 eaches have been distributed to the field | Nationwide Distribution | Open, Classified |  |  | 1370995 | Not_Computer
Z-2743-2016 | 74613 | No | NOV | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Achillon Minimally Invasive Device for Achilles Tendon Rupture Treatment |  |  | June 17, 2016 | 2016 | Integra LifeSciences Corp. | Data from internal studies suggests there may be a potential that the secondary package seal integrity does not provide the level of sterility intended for Minimally Invasive Achilles Tendon Suture System Achillon. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter and recall acknowledgment and return form dated June 17, 2016 to their customers via traceable courier. Customers were asked to review their inventory and complete the attached form. Customers with questions should contact Customer Service at 1-800-654-2873.For questions regarding this recall call 717-840-3431. |  | 6729 | 6,729 units | Nationwide distribution | Open, Classified |  |  | 6729 | Not_Computer
Z-2744-2016 | 74710 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microtek Franz Angiography Drape with Radiation Shield, Item # RBA202, and Microtek Angiography Drape with Radiation Shield and Pouches, Item # RBA203. Sterile EO. Microtek Medical, Inc., Columbus, MS 39702. |  |  | July 14, 2016 | 2016 | Ecolab Inc | Ecolab is recalling 2 models of Angiography Drapes with Radiation Shields because the sterile packaging may contain small holes in the film of the pouch. | Packaging process control | Consignees were sent on 7/14/2016 an Ecolab "Urgent Medical Device Recall" letter dated July 14, 2016. The letter described the problem and the product involved in the recall. Advised consignees to immediately quarantine the product and to notify their customers. For questions contact Customer Service at 1-800-824-3027. |  | 6750 | 6750 eaches (675 cases) | US: OH, NY, CA, MN. | Open, Classified |  |  | 6750 | Not_Computer
Z-2747-2016 | 74895 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | MENTOR 2" Tall Height TE With Sutures 350cc PN: 354-9312 |  |  | August 09, 2016 | 2016 | Mentor Texas, LP. | Tissue expander product mix found: CPX4 350cc was placed in a carton for the CPX4 450cc | Labeling mix-ups | The firm, Mentor, sent an "MEDICAL DEVICE REMOVAL" letter dated 8/11/2016 to affected consignees. The letter will be followed up with a respective sales representative visit. The letter described the product, problem and actions to be taken. The consignees were instructed to immediately identify all unused units and return them per the normal return process; forward the letter to anyone your facility that needs to be informed; if you have sent these units to another facility, forward the letter to that facility; and maintain awareness of the letter until all units in your inventory have been returned.If you have any questions, please contact your local Mentor sales representative or Mentor Customer Service at 1-800-235-5731. |  | 65 | 65 units | Worldwide Distribution: US (nationwide) and countries of: South Korea, Argentina, Mexico, Canada, Australia and New Zealand. | Open, Classified |  |  | 135 | Not_Computer
Z-2741-2016 | 74986 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Radial Artery Catheterization Kit with Sharps Safety Features |  |  | August 08, 2016 | 2016 | Arrow International Inc | Shipping carton labeling error. | Process control | Teleflex sent an Urgent Medical Device Recall Notification letter dated August 23, 2016, via FED-Ex 2 Day Air. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine all affected products. Instructions for the products return were provided in the firm's recall notification letter. Customers were asked to complete the enclosed Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. Customers with questions were instructed to call 1-866-246-6990. |  | 690 | 690 kits | Nationwide Distribution to OH, IN, TN, NC, and KY. | Open, Classified |  |  | 690 | Not_Computer
Z-2746-2016 | 74915 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Connecting Bracket (Fixture) Device (1002.65A0) is designed as a connecting element to support ancillary head rests on a Maquet operating room table immediately before, during and after surgical interventions as well as for examination and treatment. |  |  | August 12, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | Post Market Surveillance cases reported to Maquet showing instances in which the welded joint between the square mount and the first fixable swivel joint of Maquet's Connecting Bracket (Fixture) Device is broken. | Use error | Maquet sent an "Urgent Product Recall Medical Device Correction Letter and Medical Device Correction Response Form" dated August 12, 2016 via UPS to their affected customers. Customers were advised to inspect all devices that include the fixture for damage before every use. Customers were asked to fill out and sign the enclosed Medical Device Correction Response Form and return it via email to FieldAction@maquet.com or fax to 1-973-909-9927. Customers with questions were advised to contact their MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383.For questions regarding this recall call 973-709-7487. |  | 19 | 19 units (US) and 1975 units (OUS) | Nationwide Distribution to CA, FL, IA, IN, MO, UT and WI | Open, Classified |  |  | 19 | Not_Computer
Z-2745-2016 | 74771 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Whipknot Soft Tissue Cinch (USP#5) suture. PN: 7211015 Used during ligament reconstruction. |  |  | July 19, 2016 | 2016 | Smith &amp; Nephew, Inc. | Sterility compromised due to breach in sterile packaging | Package design/selection | Smith &amp; Nephew issued recall letters via Federal Express on July 19, 2016 to the customers. Customers asked to inspect inventory and locate any unused devices, and quarantine immediately. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return. |  | 34065 | 34,065 units | Worldwide distribution. US nationwide, and the countries of: AU, AT, BE, CA, CH, CR, DB, DE, ES, FI, FR, GB, GR, IC, IL, IN, IT, MY, NL, NO, PL, PR, PT,SE, SG, TH, and ZA. | Open, Classified |  |  | 34065 | Not_Computer
Z-2749-2016 | 74812 | Yes | LWP |  |  | Cardiovascular | PMA | Pacemakers: Accolade MRI DR Model L311; Accolade MRI DR EL Model L331; and Essentio MRI DR Model L111. |  |  | June 15, 2016 | 2016 | Boston Scientific Corporation | Five Accolate MRI and Essentio MRI DR pulse generators with an alternate MICS module configuration were shipped to US customers and implanted. This alternate configuration did not receive FDA approval yet. | Process control | As we considered the value of a direct physician communication we have not identified any information that would be of value to the physician because device performance and behavior remain unchanged. |  | 5 | 5 | Distributed to the states of TX, NH, MA, PA, and MN. | Open, Classified |  |  | 5 | Not_Computer
Z-2742-2016 | 74917 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA CentaurÂż PSA Assay; Cat No. 06574155/SMN 10310292 (100 test)/02676506/SMN 10310293 (500 test) |  |  | August 09, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Communication provided to emphasize that the PSA values should be interpreted in accordance with current clinical guidelines for defining biochemical recurrence following radical prostatectomy. | Under Investigation by firm | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on July 29, 2016 to be delivered to customers on August 1, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on July 28, 2016. These emphasize that the prostate specific antigen (PSA) values should be interpreted in accordance with current clinical guidelines for defining biochemical recurrence following radical prostatectomy (e.g., the 2013 American Urological Association (AUA) Guidelines or the 2015 European Association of Urology (EAU)). These guidelines define biochemical recurrence of prostate cancer as a detectable or rising PSA value post-radical prostatectomy that is &gt; or = 0.2 ng/mL (ug/L) with a second confirmatory level of &gt; or = 0.2 ng/mL (ug/L). The communications also provide functional sensitivity from a recent study to give an expectation of precision at the low end of the assay. |  |  |  | Nationwide AlbaniaAlgeriaAngolaArmeniaAustriaBahrainBelgiumBulgariaBurkina FasoCroatiaCyprusCzech RepublicDenmarkEgyptEstoniaFinlandFranceGeorgiaGermanyGreat BritainGreeceHungaryIcelandIraqItalyJordanKazakhstanLatviaLithuaniaLuxembourgMacedoniaMaltaNetherlandsNorwayPakistanPolandPortugalQatarRomaniaRussian Fed.SenegalSerbiaSlovakiaSloveniaSpainSwedenSwitzerlandTunisiaTurkeyUnit.Arab Emir.UzbekistanVatican | Open, Classified |  |  |  | Not_Computer
Z-2760-2016 | 73196 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The system provides the ability to transmit patient data files for storage, viewing and analysis at distributed locations via the intranet or Internet, or may function as a stand-alone device. |  |  | October 19, 2015 | 2015 | Invivo Corporation | Display of SpO2 and Non-Invasive Blood Pressure (NIBP) Measurements May Freeze | Unknown/Undetermined by firm | Phillips Medical Systems sent out a letter to their customers, dated October 19, 2015. The letter notified customers and distributors of the voluntary recall of the Xper Flex Cardio Patient Monitoring System, and identified the product issue as "Display o fSpO2 and NIBP measurements may freeze". Phillips assured customers that the recalled device can be used if each monitored patient is closely observed and not left unattended. These directions are directed within the instructions of the manual for regular use. Continued use is advised until Phillips contacts the individual customer to correct the issue with a software update. Some hardware is more dated, so software updates will require special handling. If customers have any additional questions, they are encouraged to contact their local Phillips representative. For further questions, please call (877) 468- 4861 option 1 then option 2. |  | 2929 | 2,929 | Nationwide Distribution to the states of :AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TX, TN, UT, VA, WA, WI, WV, Alaska, Hawaii, &amp; Washington, DC. Worldwide Distribution to the countries of : Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bulgaria, Canada, Cayman Islands, Chile, Columbia, Croatia, Democratic Republic of Congo, Denmark, Ecuador, Egypt, Estonia, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Islamic Republic of Iran, Ireland, Israel, Italy, Jordan, Kazakhstan, Republic of Korea, Lebanon, Libya, Malaysia, Mauritius, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, State of Palestine, Panama, Papua New Guinea, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Taiwan Province, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, United Republic of Tanzania, Uzbekistan, &amp; Vietnam. | Open, Classified |  |  | 2929 | Computer
Z-2758-2016 | 74395 | No | IYX | Scintillation (gamma) camera. | Radiology | Radiology | 510(k) | e.cam and Symbia E gamma camera systems with patient beds |  |  | August 05, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | "We have received reports of patients hair being caught in the e.cam and Symbia E patient bed up/down drive. The purpose of this letter is to remind users of the patient positioning instructions for your e.cam or Symbia E system and to inform you of an update that will be performed by your local Siemens service representative." | Device Design | The firm, Siemens, sent a "Customer Advisory Notice" letter on 8/5/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the "safety" sections of the user manual with your staff; continue using your system under the guidelines set forth in the letter; place advisory in the system's instructions for use; and if you have sold this equipment and it is no longer in your possession, forward the notice to the new owner and inform us about the new owner of the equipment.Siemens has designed an enhancement kit to further reduce the gap in the patient bed up/down drive. Your local Siemens service organization will be contacting you to schedule this update.If you have any questions regarding this important notice, contact your local Service representative at the following numbers: America: 1-800-888-7436; Europe, Middle East, and Africa: +49 9131 940 4000 or Asia and Australia: +86 (21) 3811 2121. |  | 5451 | 5451 | Worldwide Distribution: US (Nationwide) including states of: District of Columbia and Puerto Rico and countries of: Algeria, Argentina, Austria, Australia, Bangladesh, Bahrain, Belarus, Belgium, Brazil, Cameroon, Canada, Chile, China, Colombia, Congo-The Democratic Republic of the, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Germany, Great Britain, India, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Lebanon, Mauritius, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, Nigeria, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Spain, Serbia (Republic of Serbia), South Africa, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Uganda, United Arab Emirates, Uzbekistan, Viet nam, and Yemen. | Open, Classified |  |  | 5451 | Not_Computer
Z-2752-2016 | 74930 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small. This product is sold sterile. The parts are placed into an inner cavity and are sealed using a nitrogen process. The inner cavity is then placed into an outer cavity and heat sealed with outer Tyvek lid on the outer cavity. The patient record label is placed on the Tyvek outerlid. The sealed cavities and package insert is placed into a folding carton, sealed and shrink wrapped. |  |  | August 16, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a voluntary medical device recall of a single lot of the Coonrad/ Morrey Elbow Pin/ Bushing Replacement size Kit Extra Small because the lot was packaged with a missing humeral bushing. The missing bushing could result in a surgical delay or additional surgery if a replacement is not available. | Process design | Zimmer Biomet sent an Urgent Medical Device Recall Removal letter dated August 16, 2016, for their C/M Elbow Pin/ Bushing Replacement Kit via FedEx. All distributors were notified via electronic mail. Hospital risk managers, as well as distributors with product, were notified via courier. Customers were instructed to complete the following items including locating and removing the product in their territory, as well as identifying hospitals that currently have the affected product. Distributors will return on-hand product to Zimmer Biomet and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers were provided with a letter identifying the issue and their responsibilities. These responsibilities include:-Assisting the Zimmer Biomet sales representative with the quarantine of the product and completion and return of the Attachment 1 form.-Complete and return Attachment 1 (Acknowledgement of Responsibility) to the fax number or email address provided.If there are questions or concerns please call the customer call center at 1-800-348-9500 ext 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com.For questions regarding this recall call 877-946-2761. |  | 25 | 25 units | Worldwide Distribution - US including AL, AZ, CA, CT, GA, MN, NC, OH, OR and Internationally to Australia, Germany, Japan, Korea, and Taiwan. | Terminated | June 12, 2017 | 300 | 25 | Not_Computer
Z-2767-2016 | 75006 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Sterile Drape, SK6000 PAQ drape pouch (for use with the SPY Fluorescent System (SPY 2000/2001)) |  |  | August 18, 2016 | 2016 | NOVADAQ TECHNOLOGIES INC. | Novadaq is informing SK6000 PAQ users that a pinhole through the drape pouch film material was identified in some SK6000 PAQs during shipping validation testing. | Packaging | The firm, Novadaq, sent an "URGENT DEVICE CORRECTION: SK6000 PAQ Drape Pouch" letter dated 8/18/2016 to affected consignees. The Firm instructed consignee to cease distribution of affected product; accept the shipping-validated replacement PAQ; complete and sign the enclosed "Acknowledgment of Receipt" form, scan it and email to quality@novadaq.com. If you have any questions, contact Director of Regulatory Affairs at 1-905-629-3822 ext 205 or email: jpendlebury@novadaq.com.. |  | 20 | 20 units | US Distribution in states of: FL, MI, and VA. | Open, Classified |  |  | 20 | Not_Computer
Z-2768-2016 | 74974 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Certegra Workstation used in conjunction with Medrad Stellant CT Injection System |  |  | August 02, 2016 | 2016 | Bayer Healthcare | CertegraWorkstation when used in conjunction with the Medrad Stellant CT Injection System may exhibit thermal events, such as: smoke and sparks of fire. | Under Investigation by firm | Bayer Healthcare mailed an Urgent Field Safety Notification letter on August 03, 2016 to customers who were affected by the recall informing them of the issue and what actions to take if they were impacted. Customers were advised to use the enclosed Interim Instructions for Use for further information. For questions customers should call 1-877-229-3767. |  | 165 | 165 systems | Worldwide Distribution - US (nationwide) and Internationally to United Arab Emirates, Australia, Canada, Germany, Spain, and United Kingdom. | Open, Classified |  |  | 165 | Not_Computer
Z-2703-2016 | 74785 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE(R) II MODULAR TIBIAL BASE W/SCREWHOLES, Knee Tibial Base, REF KTTIHA60, SIZE 6 STD; SURFACE POROUS, HA COATED; FIT CEMENTLESS; USE W/ SIZE 6 FEMUR &amp; SIZE 6 TIBIAL INSERT; CONTENTS, 1 EACH, STERILE R, Rx ONLY, MicroPort Knee Systems Total Knee Systems MicroPort Total Knee Systems are indicated for use in knee arthroplasty in skeletally mature patients with the following conditions: 1) noninflammatory degenerative joint disease including osteoarthritis, traumatic arthritis, or avascular necrosis; 2) inflammatory degenerative joint disease including rheumatoid arthritis; 3) correction of functional deformity; 4) revision procedures where other treatments or devices have failed; and treatment of fractures that are unmanageable using other techniques. |  |  | June 21, 2016 | 2016 | MicroPort Orthopedics Inc. | ADVANCE(R) HA Coated tibial bases are being recalled due to a higher than expected revision rate as a result of component loosening. | Use error | MicroPort sent an Urgent Field Safety Notice dated June 22, 2016, to all affected customers. The firm requested customers complete the enclosed form and return by fax 1-901-451-6032 or by email to Postmarket@ortho.microport.com. Customers were also requested to return the affected product. Non-responding consignees will be notified after 30 days.For questions regarding this recall call 901-867-4771. |  | 38 | 38 units | Worldwide Distribution only to Australia, Austria, France, Italy, Spain, and United Kingdom | Terminated | July 26, 2017 | 400 | 889 | Not_Computer
Z-2755-2016 | 74584 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Intermediate carton label reads in part as: "Bard-Parker***Stainless Steel Surgical Blades***#12B***QTY 150 (3 x 50)***REF 371712***LOT 0088464." |  |  | June 27, 2016 | 2016 | Aspen Surgical Puerto Rico Corp. | After receiving a complaint the firm disclosed that Conventional Stainless Steel Surgical Blade catalog number 371712 lot 0088464 was incorrectly identified in the single unit foil pack as 12 instead of 12B. | Mixed-up of materials/components | An "Urgent: Medical Device Correction" letter was sent on June 27, 2016. The letter provided background on the issue, as well as actions to be taken by the user. Customers have the alternative to either keep the product or return it for replacement. A response form was included for customers to complete and return. Customers are to maintain awareness of the incorrect identification prior to use. Questions concerning the correction can be directed to Pedro.Garcia@WelchAllyn.Com. |  | 6000 | 6,000 pieces | US distribution only. | Terminated | May 15, 2017 | 322 | 6000 | Not_Computer
Z-2759-2016 | 75020 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | WECKÂż Horizon" Manual-Load Ligating Clip Applier, For Medium-Large Size Horizon Ligation Clips Curved Jaw, 11 (28cm), Rx Only, REF 337111, |  |  | August 24, 2016 | 2016 | Teleflex Medical | Misbranded: Incorrect etching on the device. | Error in labeling | Teleflex sent an Urgent Medical Device Notification letter dated August 24, 2016 to customers via FedEx. The letter identified the affected product, problem and actions to be taken. For questionscontact your local sales representative or Customer Service at 1-866-246-6990. |  | 49 | 49 units | Worldwide Distribution US Nationwide in the states of FL, IL, KY, MI, NY and OK and countries of Mexico and Canada. | Open, Classified |  |  | 49 | Not_Computer
Z-2761-2016 | 74891 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Nylon Surgical Suture for human use and vet use, sold under brand name Riverlon, Securolon, Weblon. Product is sterile. Part numbers: |  |  | August 05, 2016 | 2016 | Riverpoint Medical, LLC | Nylon Surgical Sutures are recalled because the seals may become open during distribution or subsequent packaging. | Nonconforming Material/Component | Riverpoint Medical sent an Urgent Medical Device letter, dated August 5. 2016, and the Customer Response Form to their customers via email. The letter identified the affected product, problem and actions to be taken. Customers are instructed to check inventory, discontinue selling affected sutures and return affected devices to the firm. Distributors are instructed to notify their downstream customers. Consignees should submit the Customer Response Forms. For questions call Riverpoint Medical at 1-866-445-4929. |  | 2182 | 2110 boxes distributed in the US and 72 boxes distributed outside US | US Nationwide distribution in the states of ID, MA, OR, NY and in the country of Australia. | Open, Classified |  |  | 2243 | Not_Computer
Z-2764-2016 | 74893 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Patient Cable. Part Numbers: 4460, 6735. Sterile within Chipboard Boxes within corrugated cartons. |  |  | August 10, 2016 | 2016 | Blackrock Microsystems, LLC | Blackrock Microsystems, LLC announces a voluntary field action for the Blackrock NeuroPort Biopotential Signal Processing System because the Neural Signal Amplifier and the Patient Cable are incorrectly labeled as Type CF Applied Parts. | Labeling False and Misleading | Blackrock Microsystems sent an Urgent Medical Device letter dated August 12, 2016 via email. The letter identified the affected product, problem and actions to be taken. Customers were instructed not use the amplifier in application requiring a type CF applied part; perform a physical count of the inventory of amplifiers and patient cables with the part numbers listed. For questions regarding this Correction call 801-582- 5533, ext. 240. |  | 120 | 120 total cables | Worldwide Distribution - US Nationwide and Canada, | Open, Classified |  |  | 120 | Not_Computer
Z-2753-2016 | 74722 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Axiom REG implants Sterile |  |  | June 20, 2016 | 2016 | Anthogyr | Following a FDA inspection showing that the firm failed to validate the design and process of its implants. | Device Design | The firm, Anthogyr, contacted physicians by phone and sent an " URGENT: MEDICAL DEVICE RECALL" notification by mail, dated 06/20/2016, informing users of a recall due to inadequate validation of design and process of implants by the FDA. The notice described the product, problem and actions to be taken. The physicians were instructed to isolate the parts and return any Anthogyr implants from lots subject to recall to Business Area Manager, 247 SouthWest 28th road, Miami, FL 33129;and complete and return the RETURN RESPONSE form by fax to: +33 (0)4 50 93 78 60 or e.mail: j.duthil@anthogyr.com or mail: Anthogyr- Director of Quality &amp; Regulatory Affairs- 2237 avenue- Andre Lasquin-France. The firm stated that they will not market their products in the United States until they obtain validation through 21 CFR Part 820. If you have any questions, call 404-718-9689 or email: jduthil@anthogyr.com or mfoley@anthogyr.com. |  | 111 | 111 units | US Distribution to sates of: NC, NY, and OR. | Terminated | March 28, 2017 | 281 | 121 | Not_Computer
Z-2756-2016 | 74934 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Aortic Arch Cannula and Coronary Artery Perfusion Cannula with Ballon (part number NA-55X6). The Aortic Arch Cannulae are designed to be used in the extracorporeal circuit during cardiopulmonary bypass surgery. The cannula consists of varying lengths of non-wire reinforced (flexible) polyvinyl chloride tubing that terminates in an angled tip. These devices come in direct contact with the central circulatory system but they are not intended to control, diagnose, monitor or correct a defect. |  |  | August 12, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field for the Aortic Arch Cannula and Coronary Artery Perfusion Cannula With Ballon because the tip of the cannula should be bent at either a 45 or 90 degree angle. Samples have been returned where the angle of the tip is less than the requirement. | Process control | LivaNova issued an Urgent Safety Alert letter via certified mail on August 12, 2016 to all affected customers. The letter instructed customers to check their inventory to determine if the Cannulae belongs to the product code and lots listed. If the customer does not have any of the products described in the communication, please complete the response form below and return it per the instructions indicated on the form; otherwise, please check the tip angle to determine if it is conforming to specifications. If specification is not met, please file a complaint through your normal process whether you decide to keep it and use it or return the product to Sorin Group; if you are not willing to use it, return the affected parts to Sorin Group USA. Please contact Customer Service 1-S00-650-2623 or email CustomerService@livanova.com to have a Return Material Authorization (RMA) issued. Complete the Customer Response Form attached to this letter to acknowledge receipt of this letter. Please return this form no later than September 15, 2016.For questions regarding this recall call 303-467-6306. |  | 200 | 200 units worldwide | Worldwide Distribution - US including IL, MO, TX. Internationally to Canada and Italy. | Terminated | March 27, 2017 | 227 | 1405 | Not_Computer
Z-2750-2016 | 74760 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Affinity Four Birthing Bed Model P3700 |  |  | July 15, 2016 | 2016 | Hill-Rom, Inc. |  | Nonconforming Material/Component | Hill-Rom Inc. initiated a voluntary recall of the Affinity Four Birthing Beds on July 15, 2016, with letters to customers mailed certified via FedEx. Customers were asked to do the following: During use, if you encounter a situation where the CPR mechanism fails to activate, the head of the bed can still be lowered, electrically, using the siderail articulation controls. A Hill-Rom representative will be contacting each facility to make arrangements for replacing the affected head drive.Please pass this notice on to all those who need to be aware within your organization and/or to any organization where the affected devices have been transferred.If you have any questions concerning this correction, please contact Hill-Rom Technical Support at 800-445-3720. |  | 94 | 94 beds | US Distribution to the states of : CA, FL, IN, KS, LA, MI, MO, NC, NY, OK, OR, PA, TN, TX, WA, WYForeign:BULGARIA, INDIA, UNITED ARAB EMIRATES | Terminated | March 23, 2017 | 251 | 94 | Not_Computer
Z-2751-2016 | 74798 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hard Tissue Replacement (HTR) Implant, Rx only, for use in single procedure only. Custom made. |  |  | June 13, 2016 | 2016 | Biomet Microfixation, LLC | Potential sterility deficiency. | Packaging | Zimmer Biomet Customers were sent an Urgent Product Recall letter dated June 13, 2016, via via FedEx mail. The letter notified customers of a recall of Hard Tissue Replacement Polymer Implants due to voids in sterile packaging. The firm instructed customers to locate and discontinue use of the implants, complete the Fax Response form and fax it to 904-741-9425, send back any quarantined product using the FedEx labels provided along with the recall notification, and to notify facilities where the implant was further distributed by forwarding them a copy of the provided notice. The firm instructed distributors to forward the notice to officials of their customers who are in charge of receiving recall notices. Additionally, the firm informed distributors that they are responsible for locating and returning distributed product. Questions are to be directed to 1-800-874-7711 or 904-741-4400 x9574, Monday thru Friday from 8am to 4pm ET. Her after hours phone is 812-614-9499.For questions regarding this recall call 904-741-9225. |  | 54 | 54 units | Worldwide Distribution - US including AZ, CA, FL, MO, NH, PA, TX, &amp; VA, and Internationally to Austria, Canada, England, Germany, Greece, Italy, Netherlands, &amp; South Africa. | Open, Classified |  |  | 54 | Not_Computer
Z-2769-2016 | 74649 | No | BZC | Pulmonary-function data calculator. | Anesthesiology | Anesthesiology | 510(k) | Innocor Cardiopulmonary Exercise and Lung Function Testing System |  |  | May 13, 2016 | 2016 | Innovision A/S | The purpose of the recall is to replace data transmission cable (USB-cable) connecting the lnnocor device with an external computer.The replacement USB cable will provide galvanic isolation of the lnnocor from the external computer via opto-couplers, as opposed to the grounding protection currently in use.The replacement is intended to prevent risk of electric shock. | Device Design | Recall notifications described the issue and actions to be taken by the firm. The notice also stated that replacement costs would be carried by Innovision ApS. Acknowledgment of the notice can be made by contacting service manager Knud Pedersen at kp@innovision.dk or (+45) 65 95 91 00) to schedule the service activity and to answer any questions. |  | 8 | 8 | Worldwide Distribution -- US, to Massachusetts and Washington, DC; and, the countries of Canada, Germany, and United Kingdom. | Open, Classified |  |  | 8 | Not_Computer
Z-2766-2016 | 74839 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SureSight Obturator Introducer Localization Systems (ILS) for ATEC in MRI. |  |  | August 02, 2016 | 2016 | Hologic, Inc | Tip Breakage at the glue joint between the obturator body and tip | Process control | Hologic sent an Important Product Recall letter to customers in the United States through Certified letter. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory and cease use of all affected recalled product. Customers were asked to complete the enclosed Recall Response Card indicating whether you have any of the recalled product on hand, as well as the quantity of product that you have in stock. |  | 535 | 535 units | US Nationwide Distribution | Terminated | March 30, 2017 | 240 | 1975 | Not_Computer
Z-2777-2016 | 73689 | No | MDD |  |  | General &amp; Plastic Surgery | PMA | Integra Meshed Dermal Regeneration Template 5 cmx 5 cm (2 in x 2in) Rx Only |  |  | March 11, 2016 | 2016 | Integra LifeSciences Corporation | Integra's post QA release review of historical product release test results for Meshed IDRT products identified the Peel Strength test average result was incorrectly calculated for a single Lot (Lot 105A00324750). | Under Investigation by firm | Integra LifeSciences Inc. sent an urgent voluntary medical device recall letter/recall acknowledgement and return form dated March 11. 2016. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for the affected product and immediately stop using and remove from service. Customers were asked to complete the attached form and return by email or fax as indicated on the form. Customers with questions were advised to contact Customer Service at 1-888-601-0203.For questions regarding this recall call 609-275-2700. |  | 14 | 14 packages | Nationwide Distribution to PA | Open, Classified |  |  | 14 | Not_Computer
Z-2770-2016 | 74654 | stains, hematology | KQC | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Fluorocell PLT reagent, Catalog #CD994563 |  |  | July 01, 2016 | 2016 | Sysmex America Inc | There is a potential that fluorescent platelet (PLT-F) and immature platelet fraction (IPF%) results performed using the recalled lots may exhibit lower than expected PLT-F and IPF values. The issue may be recognized by a large discrepancy between the Impedance platelet (PLT-I) value and a false low PLT-F value. In most cases, the falsely decreased PLT-F results displayed a "PLT Abn Scattergram" message with asterisks beside the results indicating the data is unreliable. In rare instances, platelet flags may be absent. The PLT-I counts are unaffected.The PLT-F is a reflex test used to confirm PLT-I results due to flagging or low values. As such, the likely impact to patient results is low. Affected results display a discrepancy between the PLT-I and PLT-F, with lower than expected PLT-F results that do not match the patient's clinical picture. | Under Investigation by firm | Sysmex sent a Product Notification Letter dated August 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. . Customers were instructed to cease use and discard the affected lots, and reply to Sysmex America, Inc. using the provided fax-back form.An updated Product Notification (Attachment 1) was mailed to all XN-Series customers on August 12, 2016 informing them of the additional recalled lots. This notification included the phone number for the Technical Assistance Center at 1-888-879-7639 in the U.S., and 1-888-679-7639 in Canada. |  | 3868 | 3868 units | Worldwide Distribution - US (Nationwide, Washington DC) Canada; Cuba; Chile; and Dominican Republic. | Open, Classified |  |  | 3868 | Not_Computer
Z-2771-2016 | 74898 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Catheter, Intravascular, Therapeutic, Long Term Greater than 30 days. |  |  | February 17, 2016 | 2016 | Arrow International Inc | Notification of component molding abnormality by supplier used to package certain Arrow Select Kit (ASK) products. | Unknown/Undetermined by firm | The firm, Arrow International, sent an "Urgent Medical Device Recall Notification" letter dated 02/19/2016 via FedEx 2-day delivery to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any of the affected products found; return product-a customer service representative will contact you with a RGA number and instructions; and complete and return the enclosed Recall Acknowledgement Form via fax to 1-855-419-8507 Attn: Customer Service or email to recalls@teleflex.com, even if you do not have no affected product.If you have any questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 6121 | 6,121 units in total | US Distribution to states of: NC, OH, FL, and MD. | Open, Classified |  |  | 6121 | Not_Computer
Z-2776-2016 | 75012 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 EP Navigation System and Accessories, Product Codes: FG540000 and FG540000C, Software Versions: v 3.2 , v 4.3 and v 2.3. |  |  | August 15, 2016 | 2016 | Biosense Webster, Inc. | Biosense Webster, Inc. is initiating a voluntary correction to re-install identified CARTO 3 Systems that were moved and installed in China without testing and approval by Biosense Webster Technical Services, as described in the Instructions for Use (IFU). | Under Investigation by firm | Biosense Webster sent a Customer Notification letter dated August 23, 2016, to all affected customers to inform them that Biosense Webster, Inc. recently became aware that some systems were improperly transported and installed at hospitals in China by Biosense Webster, Inc. Clinical Specialists. The letter informs the customers that according to the CARTO 3 System Instructions for Use (IFU), only the Biosense Webster, Inc. Technical Services team members are authorized to properly install the systems. Customers are informed that each impacted CARTO 3 System must be fully tested and approved by the Biosense Webster, Inc.Technical Services Team. The letter informs the customers of the actions to be taken. Actions Requested on Your Part:1. Read this Medical Device Correction notification carefully.2. If a CARTOÂż 3 System currently resides within your hospital, the system has been corrected, and you should ensure that you have received a Field Service Report that indicates the system is in order .For questions regarding this recall call 909-839-8500.3. Review, complete, sign, and return the Acknowledgement Form in accordance with the instructions listed on the form.4. In the future, if a CARTOÂż 3 System is installed in your facility, you should only conduct a clinical procedure after you have received a Field Safety Report dated after the installation of the device indicating that the system is in order .5. Maintain awareness of this notice and pass this notice on to anyone in your facility that needs to be informed.Customers with any questions are instructed to contact their Biosense Webster, Inc. sales representative. |  | 338 | 338 units | International Distribution to China, Hong Kong and Brazil. | Open, Classified |  |  | 338 | Not_Computer
Z-2775-2016 | 74837 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | Endura MR Mass Spectrometer u using software versions Endura MD Software 1.0 and 1.0 SP1. |  |  | July 19, 2016 | 2016 | Thermo Finnigan LLC | Thermo Fisher has determined that the Endura MD mass spectrometer instrument control software versions 1.0 and 1.0 SP! have a software defect which affect data accuracy. | Software design | ThermoFisher Scientific sent an Urgent Medical Device Field Correction letter dated July 19, 2016, to all affected customers via FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of the affected product until a company authorized Field Service Engineer installs corrected software to the instrument system. Customers were also instructed to communicate to all who operate the instrument and to complete the provided Acknowledgment Form, sign, and return it via email FieldActionTeam@thermofisher.com. Customers with questions should contact the Field Action Team at FieldAction@thermofisher.com. |  | 6 | 6 (5 in Us distribuion and 1 in Japan) | Worldwide distribution - US including FL, NY and Internationally to Japan. | Open, Classified |  |  | 6 | Computer
Z-2779-2016 | 73989 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VIDAS NT-pro-BNP |  |  | April 15, 2016 | 2016 | Biomerieux France Chemin De L'Or | Due to a packaging error; The kit does not contain the correct number of C1 or C2 control vials. | Packaging | bioMerieux. sent an Urgent Product Removal Notice letter dated April 26, 2016 to customers via FedEx. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check for affected products in their inventory and discard the incomplete products. Customers were asked to complete the attached Acknowledgement Form and return it via fax to bioMerieux. For questions contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666. |  | 229 | 229 kits | US Nationwide Distribution in the states of: AL, AZ, CA, FL, GA, IA, KS, KY, MA, MI, MN, NC, NE, NJ, NY, OH, SC, TX, VA and WV. | Terminated | July 21, 2017 | 462 | 229 | Not_Computer
Z-2778-2016 | 74940 | No | LBS | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LDX Optical Check Cassettes (OCC) |  |  | August 02, 2016 | 2016 | Alere San Diego, Inc. | Incorrect expiration date on the magnetic strip that does not match the labeled expiration date. | Packaging | Alere issued a recall notification to customers via telephone and email. Customers was informed of the affected product, problem and actions to be taken. Customers were instructed if they have the affected lot contact the Alere technical service department at 877-308-8289 to report this lot and receive a free replacement. For questions call 877-308-8289 and reference their specific case number. |  | 12951 | 12,951 | Worldwide Distribution - U.S Nationwide in the states of: TN, MA, GA, OK, VA, FL, TN, MO, AL, TX, IL, NY, VT, RI, KY, WI, AL, NC, LA, SD, TX, VA, MI, MN, IN, IA, CA, AR, PA, HI, WV, KS, CO, CT, UT and the countries of Japan, England, Canada | Open, Classified |  |  | 12951 | Not_Computer
Z-2781-2016 | 74899 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | CPX4 Breast Tissue Expander 650 cc |  |  | August 09, 2016 | 2016 | Mentor Texas, LP. | The CPX4 Tissue Expander may have been packaged in the Artoura Tissue Expander box. | Employee error | Mentor Texas sent an Urgent Field Safety Notice dated August 11, 2016 to customers.via letter and a follow-up with respective Sales Representative visit. The letter identified the affected product, problems and actions to be taken. Customers were informed the firm will credit/replace returned affected products. For questions contact your Sales Representative . |  | 30 | 30 units | Worldwide Distribution - US Nationwide and the countries of: Germany, UK, and Belgium | Open, Classified |  |  | 78 | Not_Computer
Z-2785-2016 | 74851 | No | LDH |  |  | General Hospital | 510(k) | GREER 2 ML Sterile Empty Vials 2-13 vial 25 ea - Silver Cap |  |  | February 03, 2015 | 2015 | Allergy Laboratories, Inc. | Potential for glass particles within the vials | Process control | Allergy Laboratories notified the sole consignee via phone call on February 3, 2015. For further information, please call (405) 235-1451. |  | 81625 | 81,625 vials (3,265 packs) | US Distribution to the state of : NC | Terminated | March 13, 2017 | 769 | 81625 | Not_Computer
Z-2786-2016 | 74931 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | AccuLIF PL 6-9mm x 11mm x 27mm x 8 Cage, Rx only, Sterile R, Legal The AccuLIF PL Cages are indicated for intervertebral body fusion with autograft and/or allogeneic bone graft comprised of cancellous and/or corticocancellous bone graft when the subject device is used as an adjunct to fusion in patients with degenerative disc disease (DDD) at one level or two contiguous levels from L2 to S1. DDD is defined as back pain of discogenic origin with degeneration of the disc confirmed by history and radiographic studies. Additionally, the AccuLIF PL Cages can be used as in adjunct to fusion in patients diagnosed with degenerative scoliosis. The AccuLIF Pl Cages are always to be used with supplemental internal spinal fixation. Additionally, the AccuLIF PL Cages are to be used with autograft and/or allogeneic bone graft comprised of cancellous and/or corticocancellous bone graft when the subject device is used as an adjunct to fusion. |  |  | July 26, 2016 | 2016 | Howmedica Osteonics Corp. | Since Stryker acquired the AccuLIF product in March of 2014, there have been Product Inquires (PIs) related to the PL implant confirming to have a reduction of height &gt; 1 mm. The current occurrence rate for reported post-operative reduction of cage height (&gt; 1mm) of AccuLIF PL Expandable Cages is approximately 0.47% and has demonstrated a slight upward trend since September 2015. | Unknown/Undetermined by firm | Stryker Spine sent an "Urgent Recall Letter and Product Accountability Form" dated July 26, 2016 to all affected customers. For patients who have had an AccuLIF Posterior Lumbar (PL) Expandable Interbody implant, Stryker Spine is recommending routine clinical and radiographic post-operated evaluation. Should the patient report any change in or develop near-onset symptoms, more urgent clinical and radiographic evaluation should be completed. Customers with questions should call 201-749-8389 or 201-749-8346. |  |  |  | NationwideDistribution | Open, Classified |  |  |  | Not_Computer
Z-2782-2016 | 74932 | No | PCL | Rubeola (measles) virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Trinity Biotech Captia Measles IgM, REF/Product Code 2326060,Test Kit for 96 Determinations, IVD |  |  | July 11, 2016 | 2016 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | The product lot failed to meet QC acceptance criteria during stability testing. The Calibrator OD was dropping causing the positive control to be invalid (running high out of range). | Nonconforming Material/Component | Recall notification letters (dated 7/11/2016) were sent by FedEx on 7/11/2016. Customers were instructed to discontinue use of any remaining inventory (discard on site if any remaining). Customers were also advised that if the test resulted in a valid run, such that all Controls and Calibrator met the Quality Control Criteria as stated in the package insert, the patient results can be reported and a re-evaluation of patient results is not necessary. For questions, please contact Trinity Biotech's Help Desk at 1-800-325-3424 Option #3. |  | 56 | US: 6 kits, Foreign: 50 kits | Distribution to US nationwide including California, and China (Hong Kong). | Terminated | May 09, 2017 | 302 | 56 | Not_Computer
Z-2791-2016 | 74727 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Supracondylar Nail T2 SCN 09x300 mm |  |  | June 30, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker discovered that a Tibial Nail and femoral Nail were incorrectly labeled and distributed to the field. | Labeling mix-ups | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 30, 2016 by e-mail to its customer requesting that they quarantine affected devices and sent the letter via UPS (with return receipt) on July 1, 2016. The letter described the product, problem and actions to be taken. The customer was instructed to inform users of the Medical Device Removal and forward the notice to all who need to be aware; inspect your inventory and return affected products to Stryker Orthopaedics/PFA Product Returns, Attn: Distribution Inventory Team; 325 Corporate Drive; Dock M-East , Mahwah, NJ 07431 REF: PFA #2016-094; complete and return the Business Reply Form via email to SO_M_PRODUCT_FIELD_ACTION_RESPONSE@stryker.com within 5 days; and keep a copy of the completed and executed Business Reply Form for your records. If you have any questions, feel free to contact the Mgr., Divisional Regulatory Compliance, Marketing Quality at (201) 831-6693. |  | 1 | 1 unit in total | US distribution to state of: CA. | Terminated | June 21, 2017 | 356 | 1 | Not_Computer
Z-2789-2016 | 75015 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Patella Assembly Instrument, REF 3182-1000, Non-Sterile, NDC (01) 07613327215854 (10) |  |  | August 01, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Upon failure of the press-fit, the clamping sub components and pins could disassociate from the Scorpio Patella Assembly Instrument. | Nonconforming Material/Component | Stryker sent an "Urgent Medical Device Recall Notification" Letter and Recall Business Reply Form dated August 1, 2016 to their affected customers via email and also via UPS on August 1, 2016.Customers are asked to quarantine the product subject to this recall. Branches/Agencies are asked to return all affected products available at their location to: Stryker C/O Stericycle, Attn: RA 2016-098 - Event 8296. Hospital/Branches: Customers are asked to complete and sign the Business Reply Form and fax a copy within 5 business days to: 877-272-6141 or email to Stericycle strykerortho8296@stericycle.com. Any questions, customers can call (201) 831-6693. |  | 2039 | 2,039 units | US Nationwide and Internationally | Terminated | June 08, 2017 | 311 | 2039 | Not_Computer
Z-2790-2016 | 74903 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira MedNet Medication Management Suite software, version 6.1, List Numbers 16037-64-02 and 16037-64-03, in combination with the SapphirePlus 13.1x infusion pump |  |  | July 28, 2016 | 2016 | Hospira Inc., A Pfizer Company | Hospira MedNet 6.1 software, in combination with the SapphirePlus 13.1x infusion pump, may result in incorrect bolus amount calculations for drugs whose concentrations are listed in Million Units. | Software design | Hospira sent an Urgent Medical Device Correction letter dated July 28, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter provided instructions for step-by-step process to mitigate the issue. Customers with questions were instructed to contact Hospira Global Complaint Management at 1-800-441-4100 or Hospira Technical Service Center at 1-800-241-4002, option 4. |  | 20 | 20 | Worldwide Distribution - US including ND, TX, MI, MN, WI, IA, ME, HI, PA, MA, NJ, NY, CT, and Internationally to Canada. | Open, Classified |  |  | 20 | Computer
Z-2793-2016 | 75125 | No | LLG | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Syngo Plaza picture archiving and communication system (PAC) (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. |  |  | August 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a Customer Advisory Notice to explain meaning of a suffix o.p. that shows on some DICOM images on area, distance or perimeter values. This suffix stands for on projection and indicates that a shown measurement value may be different from the real value. | Error in labeling | Siemens mailed to customers on August 04, 2016 a Customer Advisory Notice that provided interpretations about the meaning of measurement values like area, distance and perimeter shown as "o.p." by the syngo.plaza. |  | 101 | 101 Systems | Nationwide Distribution | Terminated | March 20, 2017 | 230 | 101 | Not_Computer
Z-2795-2016 | 74732 | No | PII |  | Microbiology | Microbiology | 510(k) | T2CandidaÂżCartridges; Part Number: 90-01794 (box of 12) and Part Number: 91-01534 (individual cartridge base). |  |  | July 07, 2016 | 2016 | T2 Biosystems Inc | Elevated levels of invalid results. | Under Investigation by firm | Customers were notified of the recall vie e-mail and certified mail on July 7, 2016, with an Urgent Medical Device Recall letter. Customers were asked to take the following actions:1. Please examine your current inventory and determine whether you have any of the recalled lots listed. (a.) If you DO have cartridges from the recalled lots, please note the quantity of cartridges you have of each lot number on the attached form entitled T2Candida Cartridge Recall Customer Response Form (CRF). Please also immediately quarantine any cartridges from the indicated lots and do not use them. (b.) If you DO NOT have any cartridges from the recalled lots, please also note this in the space provided on the form2. Send the completed CRF to T2Biosystems via e-mail to orders@t2biosystems.com or via fax to 781-240-0541. Your primary contact at T2 for all recall-related correspondence or questions is Georgette Yannios (gyannios@t2biosystems.com or 877-504-8282 Option 3). All product will be replaced at no charge to you as soon as possible.3. Once T2 receives the CRF, we will issue a Return Material Authorization (RMA) number to you along with a FedEx shipping label. After you receive this RMA number and label, please follow the instructions provided with the RMA for shipment of cartridges back to T2 (please do not ship reagent trays as they are not subject to this action). |  | 6864 | 6,864 | Distributed to the states of UT, FL, NH, PA, GA, MI, AL, IN, NJ, CA, NY, LA, RI, TN, TX, WI, and CT. | Open, Classified |  |  | 6864 | Not_Computer
Z-2796-2016 | 74815 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | May 06, 2013 | 2013 | Merge Healthcare, Inc. | Multiple studies were missing at a customer location that never made it to the archive before watermarking | Software design | The recalling firm first issued a letter via email dated May 6, 2013. A letter dated January 28, 2016, was also issued via email. Both letters informed the customer a software upgrade was available and of the issue and makes recommendations about using the software. They are also informed an upgrade is available to correct the issue. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.For further questions please call (800) 724-5970 Ext. 8027. |  | 45 | 45 sites potentially have the affected versions | Distribution was nationwide to medical facilities. Government distribution was made to LA and OK. | Open, Classified |  |  | 45 | Computer
Z-2802-2016 | 74910 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RAPIDLab 1260 Blood Gas Analyzer |  |  | August 11, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | There is a potential for the first and/or last name of one patient to be printed with Patient ID and result data from a different patient, even though those fields have been turned Off for the system | Software design | Siemens Healthcare issued Urgent Field Safety Notice to (US) customers via Federal Express. . All affected customers outside the US provided a copy of an Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures.This notice informs customers of the issue, Actions to be Taken by the Customer and to return the response form.Questions contact your Siemens Customer Care Center or your local Siemens technical support representative |  | 114 | 114 units | NationwideForeign:AlbaniaAlgeriaAngolaArgentinaArmeniaAustraliaAustriaBahrainBangladeshBelarusBelgiumBosnia Herzeg.BotswanaBrazilBruneiBulgariaCanadaChileColombiaCroatiaCzech RepublicDenmarkEcuadorEgyptEstoniaFinlandFranceFren.PolynesiaGeorgiaGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIraqIrelandIsraelItalyIvory CoastJapanKazakhstanKenyaKuwaitLatviaLebanonLesothoLithuaniaLuxembourgMacedoniaMalaysiaMexicoNetherlandsNew CaledoniaNew ZealandNorwayP.R. ChinaPeruPhilippinesPolandPortugalQatarRep. of YemenRepublic KoreaRomaniaRussian Fed.Saudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanThailandTunisiaTurkeyTurkmenistanU.A.E.United KingdomUruguayVatikancityVietnam | Open, Classified |  |  | 13245 | Not_Computer
Z-2830-2016 | 74712 | No | EZL | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Option-vf Urinary Catheter. Female use only. Latex Free. Sterile, Rx only. |  |  | July 11, 2016 | 2016 | C.R. Bard, Inc. | During an FDA inspection it was found out that the Practical Foley Catheters to be potentially nonsterile. | Device Design | C.R. Bard sent an Urgent - Medical Device Product Recall letter dated July 8, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm notifies customers of the recall, provides a clinical risk statement, and gives instructions regarding product disposition. Facilities are instructed to examine their inventory and quarantine any recalled product. The firm requested that customers complete the Recall &amp; Effectiveness Check Form if product is or is not in inventory. If product was further distributed, customers should be forwarded the recall notification letter and Recall&amp; Effectiveness Check Form.If you or the patient using these catheters has had an adverse event related to the recalled catheters, please contact Bard Medical Division Field Assurance at 1-800-526-4455 (option 5, then option 4) or via email at BMD.FieldAssurance@crbard. com. |  | 274 | 274 units | US Distribution to the states of : AL, CA, CO, FL, IL, IN, MA, MD, NJ, NY, OR, PA, TN, WA, and WI. | Terminated | March 22, 2017 | 254 | 274 | Not_Computer
Z-2804-2016 | 74009 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The VITROSÂż Chemistry Products Calibrator Kit 31, Lot 3155 used in conjunction with: VITROSÂż Chemistry Products HbA1c Reagent Kit, Generation 6 (GEN 6) and Assay Data Disk (ADD), Data Release Versions (DRV) 5873, 5874, 5875, or 5876, is used to calibrate the VITROS 5,1 FS Chemistry System, VITROS 4600 Chemistry System and the VITROS 5600 Integrated System for the determination of percent glycated hemoglobin (HbA1c) in human whole blood. |  |  | April 19, 2016 | 2016 | Ortho-Clinical Diagnostics | There is an incorrect value (data/calibration mathematics) on ADDs. This incorrect value will prevent a successful calibration of the assay. | Employee error | On 19 April 2016, Ortho Clinical Diagnostics distributed Correction Notification letters (Ref. CL2016-087) and ADD DRV 5877 notices to their US consignees via courier service and to their foreign affiliates via email. Customers were advised to install the Assay Data Diskette DRV 5877 or above on their VITROS system. Following a successful calibration and quality control assessment, it is acceptable to use VITROS HbA1c Reagent Kit, GEN 06. Customers who further distributed the product should also provide a copy of the enclosed letter to all customers who received the product. It is suggested that customers complete and return the Confirmation of Receipt form to indicate that you have been informed of this issue. Customers with questions can contact the Technical Solutions Center at 1-800-421-3311. |  | 520 | 520 Units | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CT, DC, FL, GA, IL, IN, KY, LA, MD, MI, MO, MS, NC, NY, OH, OK, OR, PA, SC, TN, TX, WA, WV &amp; Puerto Rico;** Foreign: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy &amp; Spain. | Open, Classified |  |  | 520 | Not_Computer
Z-2832-2016 | 74991 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Bigger-Better-Bladder, TM 2 With 3/8"ID tubing, 72" long tubing, ITEM BBB38-72 |  |  | July 01, 2016 | 2016 | Circulatory Technology Inc | The seal between the balloon and the housing may leak. | Device Design | The firm, Circulatory Technology Inc., sent an amended "Urgent: Expanded Device Recall" letter dated July 1, 2016, and response form to their customers via email. The letter described the product, problem and actions to be taken. The customers were instructed to immediately inspect their inventory, quarantine and prepare to ship the product back to Circulatory Technology Inc. Customers were also instructed to complete and return the response form via email to ytamari@cirtec.com. Customers with questions can call 516-624-2424, Monday through Friday 9:00 AM - 7:00 PM, EST. |  | 20 | 20 units | US Distribution to states of: LA, MI, NY, OH, OR &amp; TX. | Open, Classified |  |  | 20 | Not_Computer
Z-2834-2016 | 74825 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | LAUREATE World Phaco System |  |  | July 28, 2016 | 2016 | Alcon Research, Ltd. | Steps 16-25 are missing from the LAUREATE Multipak FMS Directions for Use (DFU) | Labeling Change Control | All remaining inventory of the identified affected products within Alcon DCs will be supplied corrected directions for use material prior to any further shipments. |  | 8646 | 8,646 UNITS | Burundi, China, Iran, South Korea, Taiwan, Thailand, USA. *** The recalled product is not sold in the USA or its territories. However, donated units are in use by a reduced number of USA clients for Medical Missions outside the country. | Terminated | June 07, 2017 | 314 | 8646 | Not_Computer
Z-2809-2016 | 74966 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Transpac IV Monitoring Kit with 2 SafeSet Reservoir and 2 Blood Sampling Ports, 84" Tubing, Disposable Transducer, 3 mL Squeeze Flush, Macrodrip (Pole Mount), Item No. 42646-06 |  |  | August 10, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kitscontaining the Transpac IV transducers. | Process control | An Urgent Medical Device Recall Notification letter dated 8/10/16 was sent to all their customers who purchased the Transpac IV with SafeSet reservoir kit and without SafeSet Reservoir, and with Transpac IV with marvelous valve without SafeSet reservoir. The letter informs the customers that ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. The recall letter list the affected products and lot numbers. Customers are informed of the actions to be taken. Customers with questions are instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or e-mail to recall@icumed.com. |  | 4180 | 4180 units | US Nationwide Distribution | Terminated | July 10, 2017 | 334 | 9205 | Not_Computer
Z-2829-2016 | 74963 | No | LQK | Home uterine activity monitor. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | AirStrip Sense4Baby Model C Fetal Monitor |  |  | August 04, 2016 | 2016 | AirStrip Technologies, Inc. | Belt clip may become detached causing electronic components to be exposed. | Nonconforming Material/Component | The firm sent out customer notification letters on August 4, 2016. Customers were instructed to cease distribution and use of the device and notify any customers in possession of it. Customers were informed that Airstrip would make arrangements to return the affected Central Units to AirStrip for correction at AirStrips expense. If customers have not been contacted by AirStrip to make these arrangements call AirStrip tollfree at 1 (877) 258-5869.For further questions, please call (210) 805-0444. |  | 42 | 42 | U.S. distribution to the following; AR, LA.Foreign distribution to the following; South Africa, Netherlands. | Terminated | November 30, 2016 | 118 | 42 | Not_Computer
Z-2841-2016 | 73995 | No | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IG1, IGF-I reagent, Catalog # LKGF1; For the quantitative measurement of insulin-like growth factor I (IGF-I) in serum or heparinized plasma. |  |  | April 15, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | 24 minutes of incubation is required prior to processing patient samples and for the pretreatment solution to reach full equilibration when using pretreatment solution when using the pretreatment solution (LGFA) contained in the kit. | Unknown/Undetermined by firm | On April 15, 2016, Siemens Healthcare Diagnostics Inc. distributed Urgent Medical Device Correction notices to their domestic customers and Urgent Field Safety Notices to their internationally customers via courier service. Customers are advised to take the following four actions: 1) To continue using IGF-I kit lot 411 and the associated pretreatment solution lot 055 to report patient results, ensure the IGF-I pretreatment solution is allowed to incubate with the patient sample for a minimum of 24 minutes prior to processing on the IMMULITE/IMMULITE 1000 systems. 2) Complete and return the Effectiveness Check Form attached to the letter within 30 days. 3) Please review the letter with your Medical Director. 4) If you have received any complaints of illness or adverse events associated with the product, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative. If customers have further distributed the product, please retain the letter with your laboratory records, and forward to those who may have received this product. |  | 525 | 525 units | Domestic: IL, FL, NY, OH, LA, MO, MA; International: Canada, others will be inserted | Terminated | April 20, 2017 | 370 | 525 | Not_Computer
Z-2836-2016 | 75114 | No | KIF | Root canal filling resin. | Dental | Dental | 510(k) | Sealapex Xpress, Part No. 33639 |  |  | August 31, 2016 | 2016 | Ormco/Sybronendo | SybronEndo is recalling the Sealapex Xpress because it was labeled with an incorrect expiration date of 2017-11 while the syringe was labeled with the correct expiration date of 2016-11. | Employee error | The firm, Kerr, sent an "Urgent: Medical Device Recall" letter dated 8/31/16 to its customers. The letter describes the product, problem and actions to be taken. The customer were instructed to contact SybronEndo Customer Care at (800) 537-7123 to receive a RMA number if they have any of the affected product; and to complete and email back the enclosed acknowledgement form to Sealapex.Response@kerrdental.com , even if they do not have any of the affected product.If you have any questions, call 714-516-7802 or email to: mark.dzendzel@syndrondental.com. |  | 179 | 179 units | Worldwide Distribution: US (nationwide) and countries of: Canada, Australia, Mexico, Belgium, Denmark, Germany, Great Britain, and Sweden. | Open, Classified |  |  | 179 | Not_Computer
Z-2835-2016 | 74849 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | FPS 35mm x 3.5 mm non-locking screw |  |  | July 21, 2016 | 2016 | Ortho Solutions Inc | Mislabeling of a single batch of FPS 3.5 x 35 mm Locking Screw as being Non-Locking Screws. | Labeling mix-ups | Ortho Solutions sent a "Field Safety Notice" letter dated July 21, 2016 to their affected customers.The letter identify the Solutions Business Development Manager will contact the customer to arrange for the isolation and removal of the unused product from their inventory. Replacement product will be provided. The customer was advised to passed on the "Field Safety Notice" to all who need to be aware within their organization and to any organization to which their have further distributed or transferred the affected screw. Upon removal of the affected device, customers are asked to please sign and return the "Field Safety Corrective Action Declaration Form", by which to confirm they have received the "Field Safety Notice" and that they declare that the required recall has been completed at their hospital. For further questions, call (201) 258-7305. |  | 15 | 15 | US Distribution | Open, Classified |  |  | 15 | Not_Computer
Z-2839-2016 | 75121 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | G1 Dissection Tool - 3mm FLUTED BALL; For cutting and shaping bone including spine and cranium. |  |  | August 26, 2016 | 2016 | The Anspach Effort, Inc. | The firm discovered on March 4, 2016 that several product codes were inadvertently left off of Hold 1503. | Device Design | On April 26, 2016 DePuy Synthes sent a letter informing their customers of the recall. ACTION REQUIREDOur records show that your facility has received one or more of product(s) subject to the recall (correction). The AnspachEffort, Inc. asks that you:1) There is no requirement for retaining your inventory of Gl Dissection Tools.2) To ensure Gl Cutting Tool &amp; Attachment compatibility, please reference Table 1 to review which Attachments arecompatible. The Cutting TooVAttachment combination noted with a check mark in Table I will function as intended.The Cutting Tool/Attachment combinations noted with an "X" in Table 1 will not secure effectively and should not beused.3) Review, complete, sign and return the attached reply form confirming your receipt of this letter by fax or email detailsprovided on the form. Returning the form promptly will prevent you from receiving repeat notices.1. Share this letter with others in your facility that needs to be made aware of this recall (correction).2. If you distribute any of the product(s) subject to the recall (correction) to other services or facilities, please forwardthis letter as appropriate immediately.3. Keep a copy of this notice with the product(s) subject to the recall (correction).If you have any questions regarding this correction, please contact the Complaint Handling Unit Manager at 561-494-3673or contact your Anspach Sales Representative. |  | 63 | 63 | AZ, DC, CA, MD, MI, and ORSwitzerland | Open, Classified |  |  | 397 | Not_Computer
Z-2843-2016 | 74928 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm Bowed Locking Compression Femur Plates, 14 hole |  |  | July 27, 2016 | 2016 | OrthoPediatrics Corp | OrthoPediatrics Corp.is voluntarily recalling multiple lots of the 3.5mm Bowed Locking Compression Femur 14-hole Plate and 18-hole Plate due to the devices being mislabeled. | Labeling mix-ups | Orthopediatrics initiated the recall of 3.5mm Bowed Locking Compression Femur Plates by distributing e-mails on 0July 27, 2016, followed by telephone calls to all customers on July 28, 2016. All distributors and direct representatives were sent an e-mail to provide them with a written notification of the voluntary recall and instruct them to quarantine all identified devices from the affected lot numbers immediately. Customers were instructed to quarantine recalled product and response via a hard copy Reply Form that was included in a hard copy of the Recall Notice. Each sales representative was also asked to contact Orthopediatrics Logistics Department to receive a Return Authorization Number for returning recalled product.in order to track their affected plates. If customers have any questions about this recall, they were instructed to reach out to the firm via phone (office: 574-267-0865) or email (acargill@orthopediatrics.com). |  | 41 | 41 | Nationwide Distribution to AZ, CA, CO, DE, FL, GA, KY, MI, MN, MO, KY, RI, NY, OH, OR, RI, TN, TX, VA, WAForeign: NoneVA/DOD: None | Terminated | March 24, 2017 | 240 | 41 | Not_Computer
Z-2842-2016 | 75041 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator |  |  | August 26, 2016 | 2016 | Philips Electronics North America Corporation | The handle can separate from the MRx housing due to breakage of mounts on the rear case. | Device Design | Philips Healthcare issued customer an Urgent Medical Device Correction notification/Field Safety Notice on 11/19/14. The Urgent Medical Device Correction notification/Field Safety Notice FSN86100158A informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer, and identifies what action Philips plans to take to remedy the issue.You can continue to use your MRx prior to receiving the hardware upgrade. If one or both sides of the handle are found to have separated from the MRx housing, continue to treat your patient per existing protocols. Do not use the handle to carry the MRx. The MRx can be carried in another manner. For example, the MRx shoulder strap or bed rail hook can be used if available on the unitA hardware upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Contact your local Philips representative or call 1-800-722-9377. |  | 394 | 394 units | Worldwide distribution. US Nationwide, Canada, AUSTRALIA, FRANCE, GERMANY, INDIA, JAPAN, NETHERLANDS, SOUTH AFRICA, SPAIN, SWEDEN, THAILAND, and UNITED KINGDOM. | Open, Classified |  |  | 394 | Not_Computer
Z-2845-2016 | 74896 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | 3I OSSEOTITE CERTAIN DENTAL IMPLANTS, 4x11.5 mm, Sterile, Rx only. |  |  | August 08, 2016 | 2016 | Biomet 3i, LLC | Product was incorrectly packaged. | Packaging process control | An Urgent Medical Device Removal Notice is being sent to Customers August 12, 2016, notifying them of the packaging of OSSEOTITE Certain Implant (4x11.5mm) in T3 Implant boxes. This recall is specific to Lot No. 2016031461. The material shipped within the box was described as being correct ,but only the box it was shipped in was incorrect. Customers were notified that there are no immediate or long term health consequences. In the notification letter, customers are instructed to complete the Certificate of Acknowledgement, and to return it via email (postmarket@zimmerbiomet.com) if no product is on hand or to return the certificate with product by shipping materials to: Biomet 3i - Post-Market Returns. Replacement product will be sent to customers upon return. Customers were also advised to keep a copy of their Certificate of Acknowledgement. Further questions should be directed to 1-800 443-8166 between the hours of 8:00 AM and 6:00 PM, Monday thru Friday. For further questions, please call 1-800 342-5454. |  | 92 | 92 units | US Distributions to the states of : CA and FL., and Worldwide to : Colombia, France, Germany, Italy, and Spain. | Open, Classified |  |  | 92 | Not_Computer
Z-2849-2016 | 75080 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TriMed Semi-Tubular Plate, 10-Hole, Model No. SMTP-10 |  |  | August 29, 2016 | 2016 | Trimed Inc | TriMed is recalling the Semi-Tubular Plates with 10 holes (SMTP-10) because a non-locking screw passed through a hole in a SMTP-10 plate being implanted. | Device Design | The firm, TriMed, sent a "Medical Device Recall Notification" Letter dated 8/29/16 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product subject to this recall; notify your customers at once of this product recall; discontinue use; return product, and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form as soon as possible by Fax to: 1-661-254-8485 or email to: davidmedoff@trimedortho.com. Customers with questions are instructed to contact Director of Operations at (800) 633-7221, davidmedoff@trimedortho.com, M-F, 9am-5pm, Pacific Standard. |  | 30 | 30 units | Worldwide Distribution-US (nationwide) including states of: Alabama, Florida, Indiana, Connecticut, California, Wisconsin, Missouri, Florida, Hawaii; and countries of: Sweden and South Africa. | Open, Classified |  |  | 30 | Not_Computer
Z-2850-2016 | 75094 | No | MQL | Reproductive media and supplements. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sage In-Vitro Fertilization- CSC (Choline Substituted Cryopreservation) Freezing Medium . Product Number: ART -8017 |  |  | August 25, 2016 | 2016 | CooperSurgical, Inc. | Products contain international package insert and notthe correct USA package insert | Packaging change control | CooperSurgical sent an Advisory Notice dated August 25, 2016, to affected customers via UPS with confirmed delivery receipt. The letter identified the affected product and detailed the necessary actions to be taken by customers who carried the affected lots in their inventory. The advisory notification asked customers to replace the previously inserted incorrect Instructions for use (IFU) with the correct IFU sent with the advisory notification.For questions regarding this recall call 203-601-5200. |  | 1 | 1 unit | Nationwide Distribution to CA, GA, and NY | Open, Classified |  |  | 5 | Not_Computer
Z-2846-2016 | 74647 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | Broselow/Hinkle, PEDIATRIC EMERGENCY SYSTEM, PINK/RED Oxygen Delivery Module 2/PACK, Product Code 7700RAW2 |  |  | June 29, 2016 | 2016 | Carefusion 2200 Inc | The Oral Airway pouch within the PINK/RED Oxygen Delivery Module, 7700RAW, was incorrectly packaged with the 8 cm Berman Oral Airway. The pouch contained within the module is correctly labeled as 5 cm but actually contains an 8 cm Berman Oral Airway component. Therefore, the component within the pouch is incorrect.The PINK/RED Oxygen Delivery Module, 7700RAW, is intended to be used within the PINK/RED Pouch of the BROSELOW pediatric safety system.Three (3) lots (lot# 914873, lot# 914876 and lot #914878) of the PINK/RED Oxygen Delivery Module, 7700RAW, are affected by this voluntary Field Safety Corrective Action. | Under Investigation by firm | Recall notices were sent 6/29/2016. |  | 544 | 544 (2/PACKS) | Illinois | Open, Classified |  |  | 594 | Not_Computer
Z-2848-2016 | 74885 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS Sterile Transparent Film and all the RENASYS Dressing kits that contains the RENASYS Sterile Transparent Film |  |  | July 21, 2016 | 2016 | Smith &amp; Nephew, Inc. | Sterility seal integrity failures on the Renasys Sterile Transparent Film | Unknown/Undetermined by firm | Smith &amp; Nephew sent an Urgent Medical Device Removal Notice dated July 21, 2016, to all consignees informing them of a voluntary medical device recall involving Sterile RENASYS Transparent Film and the RENASYS Dressing Kits that contain the folded RENASYS Transparent Film. Smith &amp; Nephew have identified that the process of folding the Sterile RENASYS" Transparent Film pouch during its packaging has the potential to damage the pouch seals resulting in a loss of sterility of the product. The loss of sterility has the potential to contribute to a local infection. Actions to be taken by customers and distributors:All affected Sterile RENASYS Transparent Film and the RENASYS Dressing Kits that contain the folded RENASYS Transparent Film must be returned to Smith &amp; Nephew.For questions regarding this recall call 727-642-5849.Users must complete the Acknowledgment, Receipt and Confirmation of Return Form attached with this notice and return to fieldactions@smith-nephew.com. This will trigger the return authorization process if you have product to return. An Acknowledgement, Receipt and Confirmation of Return Form must be returned even if you do not have any affected product in your inventory. Device users are reminded that reports of adverse events experienced with medical devices should be reported to FDAs MedWatch Adverse Event Reporting program. Reports may be made online, by fax or by mail. For questions, or additional information please contact Customer.Large@smith-nephew.com. |  | 45455 | 45,455 devices (in the USA) | Nationwide Distribution to AK, AL, AR, AZ, CA, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TX, UT, VA, WA, WI, Hawaii, and Puerto RicoArgentina, Australia, Bermuda, Brazil, Brunai, Canada, Curacao, Germany, Great Britain, Hong Kong, Ireland, Jordan, Malaysia, New Zeland, Peru, Saudi Arabia, Serbia, Singapore, South Africa, SriLanka, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 45455 | Not_Computer
Z-2869-2016 | 74989 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sertera Biopsy Device, Sterile; Used primarily in ultrasound-guided breast biopsies. |  |  | August 22, 2016 | 2016 | Hologic, Inc | Incorrect label on the Sertera Biopsy Kit | Labeling Change Control | Hologic notified customers in the United States through Certified letter on 8/22/16. The letter contains instructions of returning the product, provides contact information, and includes a postage-paid return card for the consignee to complete and return to Hologic. Contact technical support at 508-263-2900 or l-800-442:9892 Press Option 6 ). |  | 570 | 570 units | Nationwide. | Open, Classified |  |  | 570 | Not_Computer
Z-2867-2016 | 74296 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression Information Portal (IP5) Model 865471. The IP5 is intended to be used as a secondary display control unit for the Expression MRI Patient Monitoring System and the Expression MR200 MRI Patient Monitoring System. |  |  | September 14, 2014 | 2014 | Invivo Corporation | Frozen Display Numerics and Disabled Menu Keys after extended run time. This customer notification was sent August 22, 2014. | Software in the Use Environment | A recall notification letter was sent to hospitals who purchased the recalled product via UPS March 2014. The letter instructed users to power cycle the IP5 unit every day. This is done by pressing the power button on the side of the unit. Customers were instructed to communicate the recall to all those subject and to attach an addendum to the IP5 Instructions for Use. The firm advised customers to call their Philips representative, option 1, then option 2, if they had any questions or concerns. |  | 639 | 639 units | Worldwide distribution. US nationwide, Australia, Austria, Denmark, France, Germany, Iceland, Ireland, Italy, Netherlands, Norway, Poland, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, United Arab Emirates, and England | Completed |  |  | 639 | Not_Computer
Z-2854-2016 | 75046 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic Endovascular Devices |  |  | August 09, 2016 | 2016 | EKOS Corporation | The recalling firm identified ten devices which were distributed and could potentially be non-compliant. These devices potentially underwent a routine manufacturing test with non-compliant parameters, which may result in a small blister in the epoxy over transducers in the treatment zone. It is not known at the time the recall was initiated whether the non-compliant test parameters, or presence of the blisters, adversely affect operation of the device. | Storage | All affected consignees were contacted by e-mail on August 9, 2016, and asked to sequester implicated devices. |  | 10 | 10 units | Three international consignees - Germany, New Zealand and Taiwan | Completed |  |  | 10 | Not_Computer
Z-2868-2016 | 75029 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | PHS Spinal Tray 24G Gertie Marx, Catalogue No. 899379. |  |  | August 24, 2016 | 2016 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences has recently identified that one lot of Spinal Tray 24 G Gertie Marx Catalogue # 899379 was shipped containing the 24 G Gertie Marx needle without an introducer. | Employee error | Consignees notified on 8/24/16 via an Urgent - Field Corrective Action letter e-mail. The letter identified the affected device and the reason for the recall. Customers were asked to review their inventory of the Spinal Tray 24 G Gertie Marx to confirm the quantity received. The attached form should be completed and returned by e-mail or fax as indicated. Once the form is received, Customer Service will place an order for the customer to receive replacement devices. Questions should be directed to Customer Service at 1-800-241-2210. |  | 320 | 320 packages | Distributed to CA. | Open, Classified |  |  | 320 | Not_Computer
Z-2866-2016 | 75036 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | StÂżckert S5 System, S5 Heart-lung machine, Cardiopulmonary bypass heart-lung machine console |  |  | August 25, 2016 | 2016 | Sorin Group USA, Inc. | Sorin/LivaNova is initiating a field correction on the S5 Heart-lung machine because of the potential for failure of the Uninterruptible Power Supply. | Employee error | Sorin Group Deutschland GmbH created Field Correction Order (FCO) 2016-003 on August 11, 2016 to replace the affected power supply with a new unit. The customers will be contacted by the Field Service Representatives to schedule the replacement at the customer facilities. No direct action is required by the customer. |  | 5264 | 5264 Units Worldwide | Worldwide Distribution - US Nationwide in the states of AZ, CA, CO, FL, HI, KS, MA, NC, TX, UT and the countries of: Great Brittan, Peru, Russia, India, Hungary, Thailand, Germany, Azerbaijan, Austria, Italy, Slovakia, Bahrain, and Malaysia. | Terminated | May 26, 2017 | 274 | 5264 | Not_Computer
Z-2852-2016 | 75143 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.x picture archiving and communication system |  |  | August 05, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software error. Incorrect values for the volume calculation from a freehand VOI at the customer site. In volume calculations of prostate as well as in volume calculations of liver were too high. | Software design | Siemens sent an Customer Letter dated August 5, 2016, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer. The letter informed customers of a software change VB10B_HF02 that they have implemented to address several issues.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions please call (610) 219-4834. |  | 338 | 338 systems | Distributed Nationwide | Open, Classified |  |  | 338 | Computer
Z-2865-2016 | 75034 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Patient information in the header is only found on the first page of the report and not on the subsequent pages. | Software design | Merge sent an Urgent Medical Device Recall letter dated January 28, 2016 to affected customers via e-mail or via certified mail if they did not have an e-mail address. The letter identified the affected product, problem and actions to be taken. For questions call Merge Customer Service at (877)741-5369 or support@merge.com. |  | 42 | 42 sites potentially have the affected version | US Nationwide Distribution | Open, Classified |  |  | 42 | Not_Computer
Z-2857-2016 | 75051 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | TECNIS Symfony Extended Range of Vision IOL, Model No. ZXR00 |  |  | August 26, 2016 | 2016 | Abbott Medical Optics Inc. (AMO) | Abbott Medical Optics is recalling Intraocular Lenses (IOLs) because they were not being checked prior to release and this may result in the release of mislabeled IOLs. | Process control | Abbott Medical Optics, Inc. sent an Urgent Field Safety Notice to all affected customers to inform them that Abbott Medical Optics (AMO) is recalling 737 IOLs due to detection of an inspection equipment malfunction that resulted in a total of 737 distributed lenses not being fully checked prior to release. The letter informs customers that this malfunction may result in the release of mislabeled IOLs. The letter informs the customers of the actions to be taken. Customers are instructed to complete Customer Reply Form and fax it to AMO Quality Assurance at 714-247-4510 or email to RegCompliOne@abbott.com. Customers with any questions related to the IOL return process are instructed to contact an AMO Customer Service Representative at 877-266-4543.Customers with product complaints or adverse events are instructed to inform AMO by calling 877-266-4543. |  | 10 | 10 units | Worldwide Distribution - US nationwide, PR, and Internationally to Canada, Latin America, Europe, Africa, Middle East, and Asia Pacific. | Terminated | December 23, 2016 | 119 | 737 | Not_Computer
Z-2873-2016 | 64230 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeCare PCA Infusion System with Hospira MedNet Software, PCA Serial List Number 20709 |  |  | March 08, 2013 | 2013 | Hospira Inc. | Screw rotation error occurring when the pump detects a mismatch between the number of steps given to the motor driving delivery and the number of steps the motor actually took. When this malfunction occurs, the pump will alarm and infusion stops. | Nonconforming Material/Component | Hospira will contract with Stericycle, Inc. in Indianapolis, IN to send letter and a reply form to Hospira, Inc. Receipt of the letters will be confirmed by UPS tracking. Customers will be requested to return a reply form to Stericycle indicating the number of pumps on hand and that they have received the letter and notified their users.In addition, letters will be sent by email to ECRI, NRAC, Noblis and ISMP for notification of their members and posting on their websites. The letter will also be posted to the Hospira website. Letters will be sent to Hospira customers of record who purchased the devices directly from Hospira. Customers will be directed to notify their consignees if they further distributed the product and confirm they have done so.Telephone attempts will be made to contact customers that do not respond in order to confirm receipt of the letter. Final effectiveness will be done as the devices are upgraded with new plastic motors. |  | 8794 | 8,794 | United States (Nationwide) Distribution including Puerto Rico and American Virgin Islands. | Open, Classified |  |  | 8863 | Not_Computer
Z-2871-2016 | 75158 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software. |  |  | November 01, 2012 | 2012 | Merge Healthcare, Inc. | An incorrect biopsy or missed target could result if the incorrect grid is selected within the application. | Error in labeling | Merge sent Urgent Medical Device Recall dated January 28, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 1349 | 1,349 sites potentially have the affected versions | US Nationwide Distribution | Open, Classified |  |  | 1349 | Not_Computer
Z-2875-2016 | 74954 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Pentax Video Bronchoscope |  |  | May 31, 2016 | 2016 | Pentax of America Inc | Pentax America Inc. is recalling various bronchoscopes which contain an incorrect Operation and Reprocessing IFUS. | Packaging process control | Pentax Medical sent a U.S. Urgent Field Correction notice to all affected customers with a recall letter/return response form on May 31, 2016 via Fed Ex.Customers with questions were instructed to call 800-431-5880, ext. 2064. |  | 1719 | 1719 units | Nationwide Distribution | Open, Classified |  |  | 1719 | Not_Computer
Z-2870-2016 | 74952 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | MED 5223 - VacuFlow Safe, Multi-Sample Blood Collection Set With Luer Adapter Tube 12", Sterile, Rx only, Distributed by Tanner Scientific, Sarasota, FGL 34243  Made in China. |  |  | August 09, 2016 | 2016 | Global Healthcare Inc | The multi-sample needle retractable sheath may not fully retract over the needle during the changes of the blood collection tubes. | Under Investigation by firm | The distributors were first notified on August 9, 2016 via phone to place product on hold. They were notified of the recall on August 10, 20l6 by email. |  | 282 | 282 cartons (1000 units per carton) | FL | Open, Classified |  |  | 282 | Not_Computer
Z-2872-2016 | 75194 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology, versions 12.1.1 and 12.2 |  |  | July 12, 2016 | 2016 | Mckesson Medical Imaging Group | The STAT priority icon was not displayed in the Study List when the study was opened. | Software design | McKesson sent an Urgent Field Safety Notice on July 12, 2016 to customers with potentially affected software versions. The letter requested that they contact McKesson to acknowledge that they read and understood the Field Safety Notice and to arrange installation of software update/upgrade. Customers with questions were instructed call the McKesson Support department and ask to speak to their Support Manager.For questions regarding this recall call 404-338-3556. |  | 164 | 164 devices | Worldwide Distribution - US including AL, AK, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, VA, WA WV, WI, WY, and Internationally to Canada, France, Great Britain and Ireland | Open, Classified |  |  | 164 | Computer
Z-2878-2016 | 75113 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Therapist, Accelerator, Linear, Medical |  |  | July 26, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software patch installation to address several safety issues. | Software design | Siemens sent a Field Safety Notice to all affected customers informing them of a new software patch that they plan to install to address software issues. Customers were advised to notify anyone at their organization who should be aware of the information contained in the Notice. Customers were also advised to include the Field Safety Notice in their System Owner Manual, chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 610-219-4834. |  | 103 | 103 accelerators | Nationwide distribution | Open, Classified |  |  | 103 | Computer
Z-2881-2016 | 74805 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | TRUEread Blood Glucose Test Strips (Glucose Oxidase) are used with TRUEresult and TRUE2go meters. TRUEread Blood Glucose Test Strips are available in vials with 10 ct., 25 ct., 50 ct., and sold as 10 ct. with meters. |  |  | June 28, 2016 | 2016 | Trividia Health | Product gives incorrect low blood glucose levels. | Nonconforming Material/Component | The firm, Trividia, issued a press release on 06/28/2016 that instructed users to locate the test strip lot number on product on hand in order to determine if they possess affected product. Customers were instructed to contact Stericycle at 1-888-835-2723 or check online at www.trividiahealth.com/product-notice to see if their inventory was affected by the recall. Any strips under recall will be returned for disposal and replaced with no charge. In addition, the firm sent an "Urgent Medical Device Voluntary Recall" notice to customers addressing the recall and steps to follow. The customers were instructed to identify affected product using the same method: contacting Stericycle at 1-888-835-2723 or checking www.trividiahealth.com/product-notice to see if the product in possession was from affected lots; to notify patients that they may continue use of TRUEread, TRUEtrack, TRUEbalance, and TRUEtest strips not under recall, and to use them with the appropriate meter. Pharmacists/DME providers were instructed to call Stericycle if health care professionals contacted them regarding the recall. Distributors were requested to check their meter and strip inventories to see if they had any affected lots. Distributors were instructed to notify their own customers of the recall and collect product that is affected by the recall. Once all affected product is received by customers, distributors need to call Stericycle to obtain prepaid return labels. All returned product will be replaced at no charge.If you have any questions, contact Vice President of Marketing, at (800) 342-7226 ext. 2183 or email: LNardi@trividiahealth.com. |  | 22416 | 5,527,921 In total =4,621,639 strips &amp; 393,808 kits domestically; 500,058 strips &amp; 22,416 kits internationally | Worldwide Distribution-US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV; and Puerto Rico and countries of: Austria, Bangladesh, Cambodia, Canada, Pakistan, Sri Lanka, Thailand, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 22416 | Not_Computer
Z-2879-2016 | 74508 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Syringe Module Model 8110 with software version 9.15 |  |  | August 08, 2016 | 2016 | CareFusion 303, Inc. | A software anomaly with the Alaris Syringe module software version 9.15 may cause an infusion to unexpectedly stop when the Syringe module is transitioning from one rate to another. | Software design | The firm, BD, sent an "Urgent: Medical Device Recall Notification" letter dated 8/8/16 to customers to inform them that BD has identified a software anomaly associated with the Alaris Syringe module software version 9.15. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the instructions and promptly complete and return the Customer Response Card to expedite the corrective action process.Customers with recall related questions are instructed to contact the BD Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229. |  | 12000 | 12,000 units | Worldwide Distribution: US(nationwide) and countries of: Canada, Australia and United Arab Emirates. | Open, Classified |  |  | 12000 | Computer
Z-2884-2016 | 74861 | No | DKJ | Salicylate test system. | Toxicology | Toxicology | 510(k) | DRI Salicylate Serum Tox Assay, Catalog number 0977. |  |  | July 25, 2016 | 2016 | Microgenics Corporation | Some lots of DRI Salicylate Serum Tox Assay negative patient samples are recovering higher than the product sensitivity claim of 4.4 mg/dL. | Under Investigation by firm | The firm, Thermo Fisher Scientific, sent an "URGENT MEDICAL DEVICE FIELD ACTION" letter dated July 25, 2016 by regular mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to determine if they are using or have inventory of any of the affected lots; discontinue use and destroy any remaining inventory of affected lots per local wastes ordinances; retain a copy of the letter for laboratory records; if you have forwarded kits of the affected lots, provide a copy of letter to them; and complete and return attached Medical Device Field Action Response Form within 5 days to Thermo Fisher Scientific Technical Service via fax at 1-888-527-8001 (USA toll free), 510-979-5420 (outside of USA) or email to techservice.mgc@thermofisher.com.If you have any questions contact Technical Services at 1-800-232-3342 (USA only) or 510-969-5000 (outside the USA) press option 2 then option 3. |  | 1708 | 1708 kits | Worldwide Distribution: US (nationwide) to states of: AZ, CA, MO, TX, and PR and countries of: Canada, Germany, Hong Kong, Italy, Malaysia, and Taiwan. | Terminated | November 21, 2016 | 119 | 1708 | Not_Computer
Z-2887-2016 | 75016 | No | JNM | Transferrin immunological test system. | Immunology | Immunology | 510(k) | R&D; Systems, QuantikineÂż IVDÂżsTfR Immunoassay Kit, for the quantitative determination of soluble transferrin receptor (sTfR) concentrations in human serum and plasma For in Vitro Diagnostic Use, REF DTFR1. |  |  | August 24, 2016 | 2016 | R &amp; D Systems, Inc. | R&D; Systems, Inc. received two customer complaints that results of the Controls provided with the QuantikineÂż IVDÂż Human sTfR Immunoassay kit were out of the range high.Internal testing within R&D; Systems confirmed a problem with Control results being out of range high. | Nonconforming Material/Component | Consignees were sent via US mail on 8/24/2016 a R&D; systems "Urgent Medical Device Recall" letter dated August 23, 2016. The letter described the problem and the product involved in the recall. Advised consignees to stop using the kits and to discard all unused kits and complete and return the enclosed Acknowledgement and Receipt Form. For questions contact Bio-Technes Technical Service Department at 1-612-379-2956 or techserv@bio-techne.com. |  | 725 | 725 kits (502 US, 223 OUS) | US: CA, CO, KY, LA, MD, MA, NJ, OH, RI, UT, VA, WA.OUS: AUSTRALIA, China, INDIA, INDONESIA, ISRAEL, JAPAN, MEXICO, SOUTH KOREA, UNITED KINGDOM. | Terminated | July 26, 2017 | 336 | 725 | Not_Computer
Z-2886-2016 | 75078 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Xhibit Telemetry Receiver (XTR), Model 96280. |  |  | August 25, 2016 | 2016 | Spacelabs Healthcare Inc | The firm received multiple reports of Xhibit Telemetry System going offline or locking up. In addition, short-duration asystole alarms may not display an audio and visual alarm as reported from a single Canadian hospital. | Software design | Spacelabs sent a letter dated August 25, 2016 to affected US consignees informing them of the going offline or locking up issue when Xhibit Central Station is used as part of an Xhibit Telemetry System. On September 1, 2016, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record.Spacelabs sent the second Urgent - Medical Device Correction letter, dated September 15, 2016, to affected US consignees informing them that they have received reports of certain short-duration asystole alarms generated by the Xhibit Telemetry Receiver (XTR) where there may be no audible or visual alarm. On September 22, 2016, a second letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record.Spacelabs Healthcare will contact consignees to schedule a convenient time for Spacelabs to update the recalled Xhibit Telemetry System at no cost. Customers with Xhibit Telemetry Receivers will get all Xhibit Central Stations updated regardless if connected to an Xhibit Telemetry Receiver.Consignees are directed to technical support for additional information or technical assistance:Technical SupportSpacelabs Healthcare, Inc35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 316 | 316 XTR units total (248 units in the US and 68 units outside US) | Worldwide distribution. US nationwide including Puerto Rico, Australia, Canada, China, France, Germany, Ireland, Italy, Netherlands, Philippines, Poland, Romania, Switzerland, and United Kingdom. | Terminated | March 13, 2017 | 200 | 316 | Not_Computer
Z-2888-2016 | 75160 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Spiegelberg Brain Pressure Monitoring System |  |  | December 29, 2009 | 2009 | Spiegelberg Gmbh &amp; Co. KG | In the vast majority of products specified, a shrinkage of the air chamber during sterilization has occurred. Because of these false high readings were displayed in the lower measuring range from 0 to 20 mmHg.The observed measurement error was particularly at the lowest end of the range (to 0 mmHg) up to 8 mmHg. | Packaging | Spiegelberg Gmbh &amp; Co. KG sent an Urgent Medical Device Recall letter dated December 29, 2009, to all affected customers. Customers were instructed to quarantine and return stored products.Furthermore, they were asked to inform all their customers which have received the affected products and recall the products accordingly. Furthermore, they were asked to tell us about the actions taken by them. |  | 22 | 22 units | Nationwide Distribution to PA only | Completed |  |  | 22 | Not_Computer
Z-2897-2016 | 75082 | No | EIH | Porcelain powder for clinical use. | Dental | Dental | 510(k) | BruxZir Shaded, Model No. 70-1167-NOW0001, 70-1167-NOW0002, 70-1167-NOW0003, 70-1167-NOW0005, 70-1167-NOW0006 |  |  | August 15, 2016 | 2016 | Prismatik Dentalcraft, Inc | Prismatik Dentalcraft is recalling BruxZir Shaded because the new design was manufactured at risk and were to be quarantined and held until the validation was completed. Prismatik failed to properly identify the lots as quarantined product. | Release of Material/Component prior to receiving test results | Prismatik DentalCraft, Inc. sent an Urgent Medical Device Recall letter dated August 15, 2016, to all their customers to inform them that Prismatik Dentalcraft is recalling BruxZir NOW product because the new design was manufactured at risk and was supposed to be quarantined and held until validation was completed. Prismatik failed to properly identify the implicated lots as quarantined product. Customers are informed that the BruxZir Now with the thicker sprue design was packaged and shipped to customers in error. Customers are informed of that there is no risk to the patient and of the actions to be taken for the customers and from Prismatic Dentalcraft Inc. For questions regarding this recall call 949-440-2600. |  | 150 | 150 units | Nationwide distribution to SC, MN, NC, IN, AZ, OR, WA, CA, WY | Open, Classified |  |  | 150 | Not_Computer
Z-2890-2016 | 75100 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | The EKOS cart has the catalog number 700-60101. |  |  | August 22, 2016 | 2016 | EKOS Corporation | Casters on the EKOS carts may be missing washers which can cause premature wear and dislodgement of the casters/wheels leading to failure. | Device Design | The firm, EKOS, sent an "URGENT FIELD SAFETY ACTION" letter dated August 22, 2016 to its customers. The letters were sent to customers informing them of the field safety action, how to identify the affected carts at their facility and instructing customers locate the cart, quarantine it and do not use the cart for any procedures. A representative from EKOS Corporation will be contacting customers in the near future to arrange a date and time when the cart at a facility can be inspected and any affected caster(s) replaced. As the cart is inspected, it will be tagged to easily identify it and confirm it has been inspected and repaired as necessary. Customers will be provided with a copy of the inspection/repair form for the cart they inspect. Please retain that form for records.If customers have any questions, please contact EKOS Customer Service at 1-888-400-3567 or your local EKOS representative. |  | 455 | 455 units | US Nationwide Distribution in states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WY. | Terminated | December 30, 2016 | 130 | 455 | Not_Computer
Z-2895-2016 | 74892 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Zyno Medical Administration Set, Sterile. RX |  |  | July 29, 2016 | 2016 | Zyno Medical LLC | Administration Set potential filter leaking | Process control | Zyno issued Recall letter dated July 26th, 2016 to affected customers. Customers who received affected products were contacted between 7/29/2016 and 8/1/2016. Customers were asked to return unused product for replacement.For questions contact Customer Service Representative, at 508-907-7796, or laryloring@zynomed.com.Expanded Recall: Two recall letters issued on 11/3/2016 to address the newly identified lots. One letter for customers who received the initial recall letter and the other letter for newly affected customers who did not receive the initial recall letter. |  | 78063 | 78063 sets | US Nationwide Distribution | Terminated | July 11, 2017 | 347 | 318202 | Not_Computer
Z-2783-2016 | 74835 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Light Adaptor for Small Battery Drive and Small Battery Drive II. |  |  | June 24, 2016 | 2016 | The Anspach Effort, Inc. | There was a potential for the Adaptor and Light Adaptor for Small Battery Drive and Small Battery Drive II to generate excessive internal pressure that may cause the products to burst. | Device Design | An Urgent Notice for this Recall was provided to the consignees on 6/24/2016. They were to review their inventory and return the products with instructions. IF YOU HAVE ANY PRODUCT: To immediately quarantine product so it's not used; ensure all in their facility read the Urgent Notice for the Recall; Call the DePuy Synthes Customer Support to obtain a Return Materials Authorization (RMA) Number; to complete the Verification Section indicating the product subject to recall was located; to return the Verification Section with the product to Product Returns, The Anspach Effort Inc., 4500 Riverside Drive, Palm Beach Gardens, FL 33410; Send a copy of the completed Verification Section to DePuy Synthes, Customer Quality Dept by: FAX (561) 627-2682 or Scan/email: DPYUS-PowerToolsFieldActions@its.jnj.com;Keep the notice visibly posted for awareness until all products have been returned; maintain a copy of the notice subject to this action and keep a copy for your records. The return documentation acknowledges your receipt of the medical device removal information.IF YOU DO NOT HAVE the identified product: complete the Verification Section of the recall, by checking the box indicating no affected product located; Include your name title, address, telephone #, signature and date; Send a copy of the completed Verification Section to DePuy Synthes Customer Quality Dept by: Fax: (561) 627-2682, or Scan/e-mail: DPYUS-PowerToolsFieldActions@its.jnj.com; This return document acknowledges your receipt of medical device removal information;ADVISORY: For Alternative products: This product does not have any inventory for replacement. Since the recall of this device may have an impact on your surgical schedule, DePuy Synthes Power Tools would like to propose alternative products to assist you with a replacement for the Recalled Adaptor (05.100.024) or Light Adaptor (05.001.108) . The products are the Small Battery Drive II Hand-piece (532.110) can also be operated |  | 1833 | 1,833 devices (05.001.024 and 05.001.108) | Worldwide Distribution - US Nationwide and including the countries of Canada, Austria, Australia, Belgium, Brazil, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, Slovakia, Spain, South Africa, South Korea, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam, and areas of Middle East, Africa, Asia &amp; Pacific. | Open, Classified |  |  | 1833 | Not_Computer
Z-2889-2016 | 74901 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ETHICON MERSILENE" Tape |  |  | August 01, 2016 | 2016 | Ethicon, Inc. | ETHICON MERSILENE" Tape (also referred to as MERSILENETM Polyester Fiber Strip or MERSILENETM Fiber Ligature) products are being recalled since a group of MERSILENE" Tape product codes are being supplied with an incorrect Instructions for Use (IFU) insert. | Other | Ethicon Inc. sent an Urgent - Medical Device Correction Notification dated August 1, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ACTION REQUIRED:1. Examine your inventory immediately to determine if you have product subject to this correction on hand and locate such product(s). Refer to Attachment 2 for the Product Identification Tool to identify the products that are subject to this correction by using package labels.2. Please read Attachment 1 included within this letter thoroughly - MERSILENE" Tape Instructions for Use (IFU) .3. Ensure all users are aware of the issue by distributing this notice to all users of the Ethicon MERSILENE" Tape devices in your facility.4. Replace the existing incorrect MERSILENE" Suture Instructions for Use (IFU) with the MERSILENE"Tape Instructions for Use (IFU) provided in Attachment 1 and dispose of the incorrect MERSILENE" Suture Instructions for Use (IFU) per your internal procedures.5. Complete the Business Reply Form (BRF) (Attachment 3) confirming receipt of this notice and fax or email it to Stericycle at 1- (866) 723-2836 or Ethicon8698@stericycle.com within three (3) business days. Please return the BRF even if you do not have product subject to this correction.6. Keep this notice visibly posted with product subject to this correction in your facility for awareness.For further questions, please call (908) 218-0707. |  | 396024 | 396,024 eaches | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico. | Open, Classified |  |  | 396024 | Not_Computer
Z-2892-2016 | 75171 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza VB10A, Picture Archiving and Communication System |  |  | September 06, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software upgrade to eliminate several issues | Software design | Siemens notified customers via mail on 09/07/2016 informing them of the latest syngo.plaza software change that was available to address several issues. |  | 47 | 47 systems | Distributed Nationwide | Terminated | May 08, 2017 | 244 | 47 | Computer
Z-2898-2016 | 74955 | No | PHV |  | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CareLink iPro Version 1.10, Catalog No. MMT-7340 |  |  | August 05, 2016 | 2016 | Medtronic Inc. | Medtronic MiniMed is recalling the CareLink iPro Therapy Management Software due to a time stamp error. | Software design | Medtronic sent an Urgent Field Corrective Action letter dated August 5, 2016, to affected customers who received the CareLink iPro Therapy Management Software. The letter informed the customers that the recall affects the users of CareLink iPro software who have downloaded reports that include data that was uploaded from certain models of Roche BG meters. Customers were informed that the recall is due to time stamp error and data from certain Roche meters uploaded to CareLink iPro may resulted in incorrect reports. Customers were informed that if they manually enter Logbook data from the Roche ACCU-CHEK meters or upload data from a blood glucose meter other than the Roche meters are not affected by the recall situation. Customers with any questions regarding the recall letter are instructed to contact the Medtronic 24-Hour HelpLine at 1.800.646.4633.For questions regarding this recall call 818-576-4700. |  | 33 | 33 units | Nationwide Distribution to TN, MN, and WA | Open, Classified |  |  | 33 | Computer
Z-2899-2016 | 74951 | No | DDR | Myoglobin immunological test system. | Immunology | Immunology | 510(k) | Siemens ADVIA CentaurÂż Calibrator U, |  |  | August 08, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Failure of analytical sensitivity with Calibrator U kit lots ending in 63 and 64 and a drift from the internal standardization | Nonconforming Material/Component | Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction (UMDC) to customers in the US on August 8, 2016, to be delivered to customers on August 9, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on August 8, 2016. The communications notify customers of the failure of analytical sensitivity with Calibrator U kit lots ending in 63 and 64 and a drift from the internal standardization. The communications also inform customers that analytical sensitivity and alignment to the internal standardization will be restored with the release of Calibrator U kit lots ending in 65. Customers may continue to use Calibrator U lots CU63 and CU64 with any indate reagent kit lot, current QC ranges and current Master Curve Material lot 45719. Revised QC targets and ranges are located on the Bio-Rad website at QCnet.com for use with CU65 and higher. For further questions, please call (508) 668-5000. |  | 2248 | 2,248 kits | Worldwide Distribution - Nationwide Distribution including the countries of:Argentina Australia Brazil Canada Chile China Ecuador India Japan Korea Mexico Malaysia Peru Philippines Saudi Arabia Singapore Taiwan Vietnam South Africa | Open, Classified |  |  | 2248 | Not_Computer
Z-2900-2016 | 75111 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS CL Flow Cytometer, Catalog No. B30166; Intended for use with in vitro diagnostic flow cytometry applications using up to four florescent detection channels. |  |  | September 06, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS CL System because the use of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. | Software Design Change | An Urgent Medical Device Recall letter dated 9/6/16 was sent to customers to inform them that the use of Positions 9 and 10 of the reagent carousel can lead to misalignment of the carousel and damage to the prep probe. The letter informs the customers of the impact of the recall, action to be taken, and the resolution by Beckman Coulter. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding this notice, are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support; call 1-800-369-0333 in the US and Canada; outside the US and Canada, contact their local Beckman Coulter Representative. |  | 148 | 148 units total (12 units in US) | US: IL, PA, NY, GA, WA, FL, SC.Affected geographies: Angola, Australia, Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Israel, Italy, Malaysia, Nambia, Netherlands, Panama, Poland, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Uganda, United Kingdom and Zambia | Terminated | August 07, 2017 | 335 | 148 | Not_Computer
Z-2917-2016 | 75019 | No | OEA | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | BLUE IMPERVIOUS GOWN WITH THUMBHOOKS, X-LARGE. Non-Sterile, Single Use Only, Plastic Non-Surgical Isolation Gown Non-Sterile, Single Use Only, Plastic Non-Surgical Isolation Gown. Serves as physical barrier to the transfer of microorganisms, body fluids and particulate material. Level of barrier protection is nonspecific. |  |  | July 20, 2016 | 2016 | Cardinal Health | Blue Impervious Gown with thumbhooks, X-large. All lots do not meet design specifications, specifically 16 CFR 1611 - Standard for the Flammability of Vinyl Plastic Film. The standard defines a specific test method that quantifies the rate of flame propagation against a minimum value using an engineered test fixture. The method requires that the test be performed on different orientations of the film (machine direction and cross direction). The machine direction is defined by the manufacturing process generating the film. The formed film exits the machine moving in the machine direction. The cross direction is at 90 degrees to the machine direction. The more rapid burning direction defines the results of the test (either a pass or fail). The products being recalled failed when tested in the machine direction with a result of 1.9-2.0 inches/second versus a limit of 1.2 inches/second. They passed by self-extinguishing when tested in the cross direction. | Process control | The firm sent an URGENT: PRODUCT RECALL dated 7/20/2016 describing the reason for the recall. The letter also provides instructions for the user on how to address this issue.Cardinal Health anticipates replacement inventory of 5213PG will be available by August 22, 2016 and V5213PG by August 29. Until that time, the recommended substitute product is 5210PG, which shares all the design characteristics of 5213PG (blue polyethylene film material, over-the-head donning/tear-away removal, half-back, thumb hook wrists) and comes in similar packaging (15 individually packed gowns per box, 5 boxes per case for a total of 75 gowns per case). The firm advises that the 5210PG is "Universal" size, which is smaller than 5213PG (XL size), so it may not work for all users.The firm is requesting all customers to take the following actions:1. Immediately stop distributing affected product.2. Return the enclosed acknowledgment form via fax to 847.689.9101, whether ornot they have affected product, as Cardinal Health is required to confirm receipt ofthis notification.3. Contact the Distributor Management Customer Service team at 800.635.6021 toarrange for credit and return of any affected product they may have.4. Customers do not need to notify any customers to whom they may have distributed product affected by this recall.If they have any questions relating to this recall, please contact Cardinal Health Customer Advocacy at 800.292.9332. |  | 4062511 | 4,062,511 eaches | Nationwide to U.S. consignees located in the following states:AL ,AR ,AZ ,CA ,CT ,FL ,GA ,IL ,IN ,KY ,LA ,MA ,MD ,MI ,MO ,NC ,NJ ,NM ,NV ,NY ,OH ,PA ,SC ,TX ,UT ,VA ,WA and WV. | Open, Classified |  |  | 4062511 | Not_Computer
Z-2906-2016 | 75065 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CUSA Dissectron Fingertip Laparoscopic Cannula, REF F75221, Sterile EO, Rx only, packaged in a Tyvek pouch and secondary packaged in Tyvek pouch, Five (5) sealed pouches are packaged in a cardboard box. |  |  | August 23, 2016 | 2016 | Integra LifeSciences Corp. | During final sampling inspection, one of the lots was discovered with a total incomplete packaging seal of the sterile product. | Process control | Integra sent an "Urgent Voluntary Medical Device Recall" Letter and Recall Acknowledgement and Return Form dated August 23, 2016 to their affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was advised if they have distributed the product to identify those customers and forward a copy of the enclosed Integra Medical Device Voluntary Recall Notification to those customers. The customer was also asked to review their inventory for the affected product and complete the attached form indicating whether or not they have any of the affected product in their inventory. Customers with questions were instructed to contact Customer Service at 1-800-654-2873. |  | 9 | 9 | Nationwide Distribution to KS only | Open, Classified |  |  | 53 | Not_Computer
Z-2914-2016 | 75048 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CT3Plus (a.k.a. CT3PZ); |  |  | September 19, 2016 | 2016 | New Star Lasers, Inc. | FDA inspection found that customers who had received devices with promotional material with outside use indications were not notified when labeling was revised to remove these indications. | Labeling False and Misleading | Letters sent on September 19, 2016 via US mail advised customers of the unapproved claims that may be in their literature and advised of the approved uses for each device. |  | 97 | 44 systems CT3Plus; 47 systems CT3PZ | US distribution only. | Open, Classified |  |  | 119 | Not_Computer
Z-2911-2016 | 75014 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Walrus Arterial Embolectomy Catheter |  |  | July 07, 2016 | 2016 | Arrow International Inc | Arrow is recalling due to incorrect labeling of products. | Component design/selection | The firm, Arrow International, Inc., sent an "Urgent Medical Device Recall Notification" letter dated July 07, 2016 via FedEx to its customers. The letter described the product, problems and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any products they may have on hand; return product, and complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com, even if you have no affected stock.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2425 | 2,425 catheters | US Nationwide Distribution | Open, Classified |  |  | 2425 | Not_Computer
Z-2918-2016 | 75050 | No | HTN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Arthrex Suture Washer |  |  | August 12, 2016 | 2016 | Arthrex, Inc. | The device was assembled incorrectly. Suture and washer were not assembled; components were placed loose in the package. | Packaging process control | Consignees were notified 8/12/2016 of the potentially affected devices; Product/Batch Numbers: AR-7000-818T Batch #10028617. For Healthcare Professionals: immediately discontinue use of all affected products with stated numbers in their possession. Call (855) 552-5056 or email returns@arthrex.com with the product number, lot, and quantities so they may ship replacement products and expedite return of recalled product. To please complete the "Recall Acknowledgement of Receipt: postal card and return to Arthrex, Inc.Distributing Agencies: Immediately identify and hold all stated products in your inventory. Call 888-519-7255 ext. 72067 or e-mail AIM@arthrex.com Attn: Christi Zimmer. Have product number, lot number and quantities you may have so replacement product may be shipped and expedite the return of the recalled product. Please complete the "Recall Acknowledgement of Receipt" postal card and return to Arthrex. Inc. Any questions about the return of products to call 866-267-9138 and ask for Ellen Barry. Also by email to complaints@arthrez.com. |  | 45 | 45 devices | OR, NY, MN, KY, CA, ID, CO, MA, and FL.Thailand. | Open, Classified |  |  | 45 | Not_Computer
Z-2833-2016 | 74838 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TriFoillmaging Triumph/Triumph II Research CT scanners. |  |  | May 24, 2016 | 2016 | Northridge Tri-Modality Imaging, | The Firm has determined that several Triumph/Triumph II Research CT scanners produced by Northridge Trimodality Imaging, inc dba TriFoillmaging (or by Gamma Medica-Ideas, Inc and field upgraded with a replacement x-ray tube) may be out of compliance with 21 CFR 1020.40 (Cabinet X-ray standards) | Radiation Control for Health and Safety Act | TriFoil lmaging planned action to bring defect into Compliance:1. All repairs were performed at customer facilities. Customers were contacted by phone to arrange a convenient repair time.2. TriFoil service engineers installed a collimator on each of the affected cabinet x-ray systems to bring their emission rates into compliance with the emission limit in the performance standard.3. Full radiation surveys were performed, at maximum exposure settings, to verify each repair was effective.4. All costs associated with the repair and follow-up testing were paid directly by TriFoil Imaging.CDRH approves the CAP and understands that the CAP was completed onAugust 29, 2016.For further questions please call (818) 709-2468. |  | 3 | 3 | USA Distribution | Open, Classified |  |  | 3 | Not_Computer
Z-2913-2016 | 75108 | No | JFS | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens RAPIDPoint 405 Measurement Cartridge |  |  | September 02, 2016 | 2016 | Siemens Healthcare Diagnostics Inc | Error in the value assignment for the Chloride calibrators in the RAPIDPoint 400 and RAPIDPoint 405 Measurement Cartridge | Under Investigation by firm | Siemens Healthcare issued An Urgent Field Safety Notice via e-mail to affected Siemens Healthcare Diagnostics Regional offices on September 2, 2016 both in the United States and outside of the United States, for communication with affected customers e: the bias of the chloride values. US accounts sent hard copy via Federal Express and requested to:Complete and return the Field Correction Effectiveness Check form attached to this letter within 7 days. Review this letter with your Medical Director.Questions:Contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 6638 | 6638 cartridges | Nationwide:Foreign:Canada, Bahamas,Germany, Japan, Mexico.Belgium, Italy,Spain, Korea, France, Austria, Korea, France, Poland, Australia, china, AE, Bangladesh, Columbia, Czechkoslovakia, Bosnia | Terminated | August 07, 2017 | 339 | 7327 | Not_Computer
Z-0008-2017 | 74982 | No | OUF | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Thermo Scientific QMS Everolimus Assay; |  |  | August 15, 2016 | 2016 | Microgenics Corporation | A reagent stability issue (still under investigation) causes less separation in the calibration curve, which affects results across the entire range. For controls, maximum bias of +/- 35% for low levels, +/- 33% for mid-range levels, and +/- 47% for high levels is observed. Bias in patient samples ranges from -15.7% to +5.2% with an average bias of -5.6% when comparing the QMS everolimus assay to an everolimus LC-MS reference method. Patient samples across the measuring range  sub-therapeutic, therapeutic, and supra-therapeutic  could be affected. | Under Investigation by firm | Thermo Fisher sent an Urgent Medical Device Field Action letter dated on August 15, 2016 to customers via regular mail. The letter identified the affected product, problem and actions to be taken. Customers are advised to Discontinue use and destroy any remaining inventory of the affected lots per your local waste ordinances. Customers were instructed to cmplete the attached Medical Device Field Action Response Form and return theform to Thermo Fisher Scientific Technical Service. For questions contact Technical Service at 1-800-232-3342 (USA only) or 510-979-5000 (outside the USA) then press option 2, then option 3. |  | 1663 | 163 kits in US, 669 kits worldwide - toatal 831. | Worldwide Distribution - US Nationwide in the states of KY, CA, NY, DC, NC, TX, IL, OH, IL, FL, IK, IA, MS and countries of Mexico, Taiwan, Guatemala, Malaysia, Argentina, Dominican Republic, Philippines, Germany, Australia. | Terminated | January 04, 2017 | 142 | 1663 | Not_Computer
Z-0003-2017 | 75235 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard 360 Revision Knee System |  |  | August 08, 2016 | 2016 | Zimmer Biomet, Inc. | The drill bit used to drill the hex pre-drilled hole in the screw head of the Tibial Augment Bolts included in the Vanguard 360 Revision Knee System was larger than allowable per the process specification. As a result, the hex geometry can still be visible, but there are radial cuts into the hex flat. Therefore, when the bolt is inserted into the augment, the screwdriver is unable to catch on the hex geometry, or the screwdriver strips the remaining hex geometry, and the bolt is unable to be fully tightened.The bolt stripping or the bolt not being fully tightened will be noticed by the surgeon intraoperatively when he or she is unable to screw the augment bolts into the augment. This possible event can result in an intraoperative delay resulting in increased time under anesthesia. | Process control | Zimmer Biomet sent an URGENT MEDICAL DEVICE RECALL REMOVAL notifications dated August 8, 2016, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of business hours will receive a prompt to record a voicemail and will be returned upon receipt. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. For further questions, please call (800) 348-2759. |  | 454 | 454 | Worldwide Distribution - US Distribution to the states of : VA, UT, TX. OR, MN, AL, PA, CA and KY., and to the countries of : AUSTRALIA, China, Italy, Netherlands and Japan. | Terminated | June 05, 2017 | 301 | 454 | Not_Computer
Z-0002-2017 | 75127 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Merit Inflation Syringe Kit. Catalog Number KOS-02857 |  |  | August 25, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for Merit Inflation Syring Kit due to unit labels missing required device labeling information. | Labeling design | Merit Medical Inc. notified all affected customers on August 25, 2016, via verbal and written notification to return the affected units. The returned units will be re-labeled.For further questions, please call (801) 316-3690. |  | 1297 | 1,297 total units | One consignee located in CA. No Canadian, other foreign or VA/govt/military consignees. | Open, Classified |  |  | 1297 | Not_Computer
Z-0001-2017 | 75118 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDIÂż H/S Cuvette |  |  | December 11, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received complaints of the CDI System 500 monitor displaying the H/S DISCONNECT AT CUVETTE error message when the CDI H/S Cuvette does not make a proper connection to the CDI H/S Probe. In this situation, blood parameter values for HCT, Hgb, and SO2 do not display. | Other | Terumo (Planned Action) plans to execute the following steps as part of the correction activity:Notify all customers of the potential that a cuvette may disengage from the circuit (probe) leading to a "disconnect" message on the monitor. This notification will be performed via a Customer Notification Letter. For further questions, please call (410) 392-7277. |  | 140000 | 140,000 units | Worldwide Distribution - USA (nationwide) Distribution. | Completed |  |  | 140000 | Not_Computer
Z-0017-2017 | 74320 | No | EIH | Porcelain powder for clinical use. | Dental | Dental | 510(k) | Zenostar MT Color Impulse blue, Art. no. 681165 |  |  | May 27, 2016 | 2016 | Ivoclar Vivadent, Inc. | The primary packaging of these products may show leaks in individual cases. As a consequence, the liquid may escape inadvertently and direct contact with this liquid may lead to skin irritation. | Process design | Wieland sent an Urgent Medical Device Recall Notification letters dated May 27, 2016, and Response Forms to their customers via courier service to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Segregate any affected materials immediately. Customers were asked to complete the attached Acknowledgment form and return it to Ivoclar Vivadent, Inc. by mail, fax or email as indicated on the form. For questions contact Ivoclar Vivadent Technical Customer Service at 800-533-6825. |  | 79 | 79 units | US Nationwide Distribution in the states of: AK, AL, AZ, CA, CO, CT, FL, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NE, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WI. | Open, Classified |  |  | 654 | Not_Computer
Z-0038-2017 | 75173 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 2.5, 3.0, 3.5, 4.0, 4.5, 4.7, 5.0 and 4.3 (InverseArc 1.0) -- Radiation Therapy Treatment Planning System |  |  | September 01, 2016 | 2016 | RAYSEARCH LABORATORIES AB | RaySearch became aware of the problem as it was discovered in cooperation with a customer experimenting with the dose calculation on a phantom. The correction concerns two issues found with the dose calculation when using a region of interest (ROI) of type Fixation or Support with material override within the patient outline (External ROI) in RayStation 2.5, 3.0, 3.5, 4.0, 4.5, 4.7, 5.0 and 4.3 (InverseArc 1.0). | Under Investigation by firm | RaySearch sent a Field Safety Notice, on September 1, 2016, (RSL-D-61-301, dated 9/5/2016), was sent by e-mail to all affected customers. ACTIONS TO BE TAKEN BY THE USERDo not use ROIs of type Fixation or Support within the patient outline. Always use other ROI types formaterial override within the External ROI. A small overlap with the External ROI is acceptable if the Fixationor Support ROI is mainly outside the External ROI.Please educate planning staff and all users about this workaround.Inspect your product and identify all installed units with the above software version number, then confirmyou have read and understood this notice (contact information below).SOLUTIONA correction of these issues will be available in the next version of RayStation, scheduled for market releasein December 2016. In the meantime, this field safety notice is distributed to all customers. Until a correctedversion has been installed, all affected users must maintain awareness of this field safety notice.TRANSMISSION OF THIS FIELD SAFETY NOTICEThis notice needs to be passed on to all those who need to be aware within your organization. Pleasemaintain awareness of this notice as long as this version of RayStation is in use to ensure effectiveness ofthe workaround. |  | 1368 | USA: 522 units, OUS 846 units | Nationwide distribution | Open, Classified |  |  | 1368 | Computer
Z-0039-2017 | 74987 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension VistaÂż Calcium (CA) FlexÂż reagent cartridge |  |  | August 03, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics determined that the Dimension Vista Calcium (CA) Flex reagent cartridge lot # 16060BB may produce erroneously low or high results from specific well sets. | Under Investigation by firm | Siemens mailed on August 09, 2016, an Urgent Medical Device Recall letter to all customers who purchased the CA lot 16060BB. The notification letter requested that customers discontinue use and discard the affected kit. |  | 5991 | 5,991 cartridges | Distributed Nationwide and the countries of : Austria, Bahamas, Belgium, Canada, Czech Republic, Denmark, Germany, Italy, Japan, Netherlands, Portugal, Saudi Arabia, South Korea, Slovenia, and Spain. | Terminated | May 11, 2017 | 281 | 5991 | Not_Computer
Z-0016-2017 | 75017 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Panoramic X-ray Model PC-1000, Laser-1000 packed in a crate |  |  | August 19, 2016 | 2016 | Panoramic Rental Corp. | Panoramic Corporation is performing a voluntary Medical Device Correction on certain PC-1000 X-ray systems due to the potential for a safety nut to be ineffective, which could result in blunt force trauma that may impact the head, neck or shoulders. | Device Design | Panoramic Corporation sent an Urgent Medical Device Correction letter dated August 19, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form and return to Panoramic. For questions contact Panoramic at correction@pancorp.com or 888-381-2577 Monday through Thursday 8:00 AM to 6:00 PM and Friday 8:00 AM to 5:00 PM EST in order to verify the effectiveness of the safety component. |  | 13340 | 13,340 | US Nationwide Distribution including Puerto Rico | Open, Classified |  |  | 13340 | Not_Computer
Z-0010-2017 | 75157 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Model Number 728321 |  |  | August 24, 2016 | 2016 | Philips Medical Systems | The infant performance test did not meet the test specification for the Infant Head Uniformity test with High Resolution. As a result, when the user performs High Resolution head scans, the reconstructed images may exhibit degraded image quality manifested as non-uniformity of the image, CT number shift greater than 5 Hounsfield units, and reduction in gray/white matter differentiation. The issue is most sever in the High Resolution mode, and the factory reference pediatric scan protocols are set to High Resolution default. Adults head scans have a factory reference protocol that defaults to standard resolution which doesn't experience the issue. | Software design | Philips sent an Urgent Field Safety Notice (FSN) CLE 16-045 letter dated August 26, 2016, via a certified letter to customers. The letter identified the affected product, problem, actions to be taken by customer/user and actions planned by Philips to correct the problem. The letter informed customers a Philips field service representative will install a software update on the affected systems. For questions contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, follow the prompts).. |  | 358 | 358 | Worldwide Distribution - US Nationwide in the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, GA, IA, IL, IN, FL, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NH, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY. and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Bolivia, Brazil, Cambodia, China, COlumbia, Cuba, Denmark, Dijbouti, Dominican REpublic, Egypt, Estonia, Finland, France, Germany, Gilbraltar, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lebanon, Libya, Lithuania, Malaysia, Martinique, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Seitzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uniter Arab Emirates, United Kingdom, Uzbekistan, and Vietnam. | Open, Classified |  |  | 1302 | Computer
Z-0013-2017 | 75126 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Multix Fusion Stationary X-Ray System |  |  | August 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential issue with the patient table for Ysio Max and Multix Fusion systems. There is an unlikely risk that an internal part of the table may detach causing the table top to tilt to the left side which could result in an injury to a patient on the table. | Process control | Siemens sent a Customer Safety Advisory Notice letter dated August 4, 2016 to their customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to discontinue use of the tables until one of their service technicians can inspect the affected part of each table. |  | 37 | 37 Systems | US Nationwide Distribution | Open, Classified |  |  | 37 | Not_Computer
Z-0045-2017 | 75068 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ROI-A Anterior Delivery Device. Orthopedic manual surgical instrument. Model number: SI-ROIA-0023 |  |  | August 29, 2016 | 2016 | LDR Spine USA, Inc. | The ROI-A Anterior Delivery Device T-Handle, part number SI-ROIA-0023 has experienced binding of the rotation of the t-handle. There have been no reports of patient injury | Device Design | The firm, LDR, contacted affected consignees via phone and then sent a follow up email dated 9/8/16. The phone script and email described the product, problem and actions to be taken. The consignees were instructed to promptly return the instrument using the RMA#1180. The firm will perform a rework that prevents the binding to occur.If you have any question, contact the Quality Engineering Manager by email: ron.musselman@zimmerbiomet.com, phone: 512-344-3436 or Field Inventory Manager - US Supply Chain or phone: 512-344-3436. |  | 24 | 24 units | US Distribution to states of: GA, OR, FL, KY, NC, AR and PA. | Terminated | January 13, 2017 | 137 | 24 | Not_Computer
Z-0043-2017 | 75115 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Operating Table TruSystem 7000 U, Material Number 1604788; Operating Table TruSystgem 7000 U (MB), Material Number 1604786 |  |  | July 10, 2016 | 2016 | Trumpf Medical Systems, Inc. | Three complaints describing TruSystem tables unexpectedly descending approximately 20 centimeters. | Nonconforming Material/Component | Trumpf Medical sent an Urgent Field Safety Notice dated July 8, 2015, to all affected consignees. The letter indicated that the tables should not be used until an authorized service engineer has exchanged the spindles of the affected operating tables. The letter included an acknowledgement form which is to be returned. The letter also requested a sub-recall if the product had been further distributed. For questions regarding this recall call 843-329-0543. |  | 84 | 84 units | Worldwide Distribution - US including CA, CT, GA, IL, NC, OR, SC, VA, WI, WY, and Internationally to Australia, Azerbaijan, Brazil, Finland, France, Germany, Hungary, Italy, Luxembourg, Norway, Poland, Russia, Slovakia, and Switzerland | Open, Classified |  |  | 84 | Not_Computer
Z-0046-2017 | 74871 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems, Inc Aquilion Lightning TSX-035A CT Scanner. |  |  | April 25, 2016 | 2016 | Toshiba American Medical Systems Inc | It has been found that if the scan range extends beyond the maximum Field of View (FOV), a reconstruction operation error may occur. As a result, scanning may be interrupted and rebooting of the system may be necessary. Please note that this issue has not occurred in the U.S. | Software design | Toshiba America Medical System, Inc. planned action to bring the defect into Compliance:1. A notification will be sent to customers with affected systems in compliance with 21 CFR 1003.21 upon approval by the Secretary. The notification letter includes a statement by the correction will be made at no cost to the customer.2. To correct this issue, modified software will be installed on the systems to prevent this occurrence from happening. The consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.3. To avoid this problem in scan planning, the consignee will be advised to do the following until the modified software is installed:a. Confirm that the scan range does not extend beyond the maximumFOV.b. If the ROI for the scan range (indicated by the yellow solid lines) is set as the same size as the maximum FOV (indicated by green dashed lines), the ROI position must be changed. Instructions for changing ROI will be provided as attachments to the notification letter.4. The modification to the software will be completed within 90 days from the date that the letters are submitted to Toshibas customers.5. Progress letters will be provided to the Secretary, including the number of electronic products repaired. For further questions, please call (714) 730-5000. |  | 165 | 165 | US Distribution to the states of: PA, IA, MA, SD, MI, LA, KY, NC, NY, AR, IN, WI, OH, TY, FL, WV, WA, CT, MS, AL, ND, CO, HI, NO, MN, ID, OH, GA, MT, LA, NM, VT, TN, CA, NH, NJ, NE, SC and KS. | Terminated | July 13, 2017 | 444 | 165 | Computer
Z-0050-2017 | 75208 | No | DGC | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | IgE Calibrator Series |  |  | September 07, 2016 | 2016 | Randox Laboratories Ltd. | According to the firm, the IgE Calibrator Series IE2492 (lot 387-392IE) value assignment has now been aligned to the 3rd International Standard for human serum IgE 11/234.Turbidimetric (Non-IFCC Cal.) Quality Control Targets and Ranges now apply. The calibrator value sheet is to be updated with the new values. | Under Investigation by firm | The firm plans on contacting each consignee about the recall. |  | 4 | 4 | All consignees are in the US. | Open, Classified |  |  | 4 | Not_Computer
Z-0007-2017 | 75031 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 0.2 Micron Filter, 50 mm |  |  | August 24, 2016 | 2016 | Baxter Corporation Englewood | Baxter Healthcare Corporation is issuing a voluntary product recall for all unexpired lots of the 50mm 0.2 Micron Filter (Product Code: H93835) due to the potential for the filter membrane layer to be missing and for particulate matter to be present. | Nonconforming Material/Component | The firm, Baxter, sent an "URGENT PRODUCT RECALL" letter dated 8/24/2016 to its Consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) locate and remove all affected product; 2) contact Baxter Health Care Center for Service to arrange for return of product at 888-229-0001 between the hours of 7am-6pm Central Time, Monday-Friday; 3) Complete and return the Customer Reply form to Baxter via fax to 224-270-5457 or scan and email to fca@baxter.com, and 4) if you distributed this product to others facilities or department within your institutions forward copy of the communication.For general questions, contact Baxter Product Surveillance at 800-437-5176 between the hours of 8am-5pm Central Time, Monday through Friday. |  | 202647 | 202,647 | Worldwide Distribution - US (Nationwide) and countries of: Austria, Belgium, Czech Republic, Denmark, Finland, Germany, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | July 26, 2017 | 336 | 202647 | Not_Computer
Z-0057-2017 | 74614 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Better-Bladder With 14" ID tubing, ITEM BB14, STERILE, |  |  | July 01, 2016 | 2016 | Circulatory Technology Inc | Twisted or wrinkled balloon components which are considered a cosmetic defect by the firm. | Use error | The firm, Circulatory Technology Inc., sent an "Urgent: Device Recall" letter dated July 1, 2016 and response forms to their customers via email. The letter described the product, problem and actions to be taken. The letter instructed customers to immediately inspect their inventory visually for units with wrinkled or twisted balloons and quarantine. Customers were also advised to complete and return the response form via email to ytamari@cirtec.com. Customers with questions are asked to contact Yehuda Tamari at (516) 624-2424, Monday to Friday, 9:00 AM - 7:00 PM, EST. |  | 958 | 958 units | US Nationwide Distribution to states of: AL, AR, CA, Washington, D.C., FL, GA, IA, IL, IN, KY, LA, MI, MS, NC, NJ, NY, OH, OK, PA, SC, TN, TX &amp; VA. | Terminated | April 14, 2017 | 287 | 958 | Not_Computer
Z-0051-2017 | 75213 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | MicroStic Plus w/SS Wire 5F and MicroStic Plus w/NiTi Wire &amp; Echo Needle 4F. |  |  | September 08, 2016 | 2016 | Argon Medical Devices, Inc | The supplier initiated a recall because affected lots of the micro-introducer have an outer diameter of the dilator hub that is slightly smaller than is capable of securing the dilator to the sheath | Nonconforming Material/Component | An Urgent - Product Recall Notice was sent to consignees identifying the affected device and the reason for the recall. Customers were asked to return all affected devices. The firm will send replacement devices upon receipt of the returned affected product and once assembly of a cleared lot that does not contain recalled devices has been completed. If there are questions regarding the letter or recall action, customers can contact Beckie.Ellis@argonmedical.com. |  | 200 | 200 units | Worldwide Distribution -- US, to the states of TN and MS; and the country of Canada. | Terminated | June 07, 2017 | 272 | 200 | Not_Computer
Z-0058-2017 | 75130 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Teleflex Percuaneous Solutions, Gripper Grasper, Ref ML003EE, Rx Only Sterile Manufactured for: Teleflex Medical, Park, NC 27709 |  |  | September 09, 2016 | 2016 | Teleflex Medical | The product is being recalled due to a labeling error. The Gripper Graspers are labeled as Johans Graspers and the Johans Graspers are labeled as Gripper Graspers. | Process control | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" Letter dated 9/9/2016 to consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any products listed; to return product- complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com even if you do not have the affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 160 | 160 ea. in total | US Distribution to states of: AZ, CA, FL, NY, NC, and TX. | Open, Classified |  |  | 160 | Not_Computer
Z-0056-2017 | 74627 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | The Medix NatalCare LX Intensive Care Incubator. Neonatal incubator. |  |  | August 25, 2016 | 2016 | Natus Medical Incorporated | NatalCare LX incubators may have a defective base where the weld holding the wheel bushing to the tube is inadequate. There were several complaints of wheels bending out of alignment. | Unknown/Undetermined by firm | Natus sent an Urgent Notification Field Safety Notice on August 25, 2016 to all affected customers. Customers are asked to inspect the quality of the welds and wheel alignment for the Natal Care LX Incubators at their facility per attached Inspection Instructions for Assessing the Natal Care LX Wheeled-Based Joint and Weld.Customers are provided with the Natal Care LX Rolling Base Verification Form (DOC-019727) and are requested to verify the weld and gap status for each of the four wheels on the incubator and to note the results on the verification form. After the Verification form is completed and returned to Natus, Natus will arrange to ship a repair kit for each affected wheel, along with instructions and simple tools to facilitate the repair of the Natal Care LX base.Natus Technical Service will follow-up with phone calls to all non-responding customers and again to verify that the repairs have been performed.Consignees who have any questions, please call Technical Service at 888-496-2887 or send email to technical_service@natus.com |  | 104 | 104 units | US Distribution to the states of : AL, CA, FL, IL, IN, KY, LA, MA, MS, NJ, NY, OK, PA, SC and VA, and Internationally to Pakistan. | Open, Classified |  |  | 104 | Not_Computer
Z-0060-2017 | 75077 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Universal Acetabular Cup Positioner/Impactor, REF 2101-0200, Non-Sterile, Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430, USA, A Subsidiary of Stryker Corp. |  |  | August 18, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received reports of the thread length protruding past the dome of the acetabular trial or implant for the Trident Universal Impactor/Positioner. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Recall Notification dated August 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions Needed: 1. Please inform users of this Urgent Medical Device Recall and forward this notice to all those individuals who need to be aware within your organization.2. Branches/Agencies: Return all affected products available at your location toStryker C/0 Stericycle3. Hospitals/Branches: Complete and sign the enclosed Business Reply Form and fax a copy to 877-546-0444 or email to Stericycle strvkerortho6536@stericycle.com. Our records indicate that you have received the above referenced instrument. It is our responsibility to ensure that customers who may have received this affected instrument also receive this important communication. Please assist us in meeting our regulatory obligation by faxing back the attached Business Reply Form within 5 days.We regret any inconvenience this action may cause you and if you have any questions, feel free to contact me at (201) 831-6693.For further questions regarding this recall, please call (201) 831-5000. |  | 3768 | 3,768 units | US and Internationally | Open, Classified |  |  | 3768 | Not_Computer
Z-0062-2017 | 66544 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia T Series SPECT + CT |  |  | August 16, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The detectors on the Symbia S and Symbia T camera systems are supported and moved horizontally and vertically by four linear bearing cars. There is a potential for the linear bearing car to fail allowing the ball bearings to fall out. | Component design/selection | Customers were contacted by a service organization that performed product inspections at their sites. |  | 1109 | 1109 units | Worldwide Distribution - US: Nationwide, including Washington D.C., and Puerto Rico. and the countries of: Poland, France, Austria, Czech Republic, Ireland, Germany, Italy, Japan, Canada, United Arab Emirates, Turkey, Netherlands, United Kingdom, Brazil, South Africa, Kuwait, Belgium, Russia, Taiwan, Australia, Bangladesh, Bulgaria, Brazil, Switzerland, Chile, China, Colombia, Denmark, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, Hong Kong, Croatia, Hungary, Indonesia, Iran, Israel, India, South Korea, Kuwait, Kazakhstan, Lithuania, Latvia, Morocco, Mexico, Norway, New Zealand, Peru, Phillipines, Pakistan, Portugal, Qatar, Romania, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Thailand, Tunisia, Taiwan, Venezuela, and Vietnam. | Completed |  |  | 1841 | Not_Computer
Z-0063-2017 | 75177 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Blood Pressure Monitor UA-767PBT-Ci for small cuff. Designed to be used by end users who are eighteen (18) years and older at home to monitor their blood pressure (systolic and diastolic) and pulse rate. |  |  | September 12, 2016 | 2016 | A &amp; D Enginering Inc | A&D; may have shipped blood pressure monitors that were not labeled for use with small cuff. | Labeling Change Control | Notification letters were sent on September 12, 2016 to all affected customers. Letters identify the product and reason for the recall. Letters include a serial number list that is customer specific and ask that inventory be reviewed and that other organizations that may be impacted by this issue be notified. |  | 378 | 378 units | to distributors in NY, MA, VA. TX, | Open, Classified |  |  | 378 | Not_Computer
Z-0064-2017 | 74922 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | TransMotion Medical, TMM3 Series Video-Fluoroscopy Swallow Study Stretcher-Chair. |  |  | June 29, 2016 | 2016 | Winco Mfg., LLC | Non-conforming component caused arms of TMM4 Stretcher chair to no longer meet design specifications. | Nonconforming Material/Component | Customers were sent an Urgent: Medical Device Recall notification on 7/12/2016 for the specific models of the TransMotion (TMM) Stretcher chairs. The letter explained the reason for the recall and asked customers to take action. Customers are asked to follow the instruction in the Operator Manual for proper egress. Patients are not allowed to egress from the chair without assistance nor should they be left unattended. The chairs will be repaired on site; and, removal of the chair from service is not required with proper supervision and use. Customers are asked to complete the instructions for the acknowledgement form and return. The firm will contact the customers to arrange for repair. Questions regarding the field correction should be directed to the Customer Care department at 800-237-3377 or 352-854-2929 or e-mail QA-RA@wincomfg.com |  | 264 | 264 Stretcher Chairs | Worldwide Distribution -- USA, including AR, CA, CO, CT, FL, GA, IL, KS, LA, MD, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, PA, RI, TN, TX, VA, WA, WI, and Puerto Rico; and the countries of Canada, Australia, Italy, Panama, Switzerland and Taiwan. | Open, Classified |  |  | 264 | Not_Computer
Z-0069-2017 | 75297 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | V. Mueller Neuro/Spine, SHADOW-LINE, SPINAL DISTRACTION SCREW, STERILE, various sizes (12mm, 14mm, 16mm and 18mm) 5 PER BOX. |  |  | September 23, 2016 | 2016 | Cardinal Health 200, LLC | Potential risk associated with corrosion demonstrated on the distraction screw surface. | Package design/selection | On September 23, 2016, Becton Dickinson sent a Field Safety Notice to all affected customers describing the details of the recall and necessary action to take. ACTIONS TO BE TAKEN BY BD:A credit will be applied to your account for the affected product upon receipt of the affected unit(s) or receipt of certification of destruction documented on the customer response form.Complete the Recall Response form certifying that all affected product in your inventory has been destroyed and mail/email a copy of the response form.: The firm requests customers to promptly return the enclosed response form to expedite the correction process and acknowledge receipt of this notification. Or email: GMB-GLBÂżVMFieldActions@carefusion.com For further questions, please call (847) 473-1500. |  | 72076 | 72,076 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Germany, New Zealand, Saudi Arabia, South Africa, Switzerland and Thailand. | Open, Classified |  |  | 72076 | Not_Computer
Z-0067-2017 | 75148 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | The Introes Pocket Bougie is an endotracheal tube introducer. The device is an endotracheal tube introducer and it is used as a guide during intubation. Anesthesiology. |  |  | August 19, 2016 | 2016 | Brathwaites Oliver Medical | The Introes Pocket Bougie an endotracheal tube introducer size 14 Fr (4.7mm) x 60 cm may break during intubation. | Nonconforming Material/Component | On August 19, 2016 the firm sent a MEDICAL DEVICE RECALL NOTICE &amp; CORRECTIVE ACTION letter to affected customers. The letter described the reason for the recall and necessary actions to be taken by the consignees. These included:BOMImed is requesting bougies with lot# 101615 to be returned for evaluation due a potential manufacturing concern.Complete the included Medical Device Recall Customer Letter and follow the instructions to obtain a Return Material Authorization number andapplicable shipping instructions to return the affected products.BOMImed Regulatory department also requires the below CustomerConfirmation section to be completed and return toregulatory@bomimed.com to confirm this notification was received. |  | 760 | 760 units | U.S. Consignee located in the following states: IL, WI, WE, OR, NY, UT, IN, NE, PA, MO, OH, AZ, MA, CT and NC.*** Foreign Consignee: Canada. | Open, Classified |  |  | 760 | Not_Computer
Z-0071-2017 | 75330 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Brivo 715 Prime, OEC Brivo 785 Essential, OEC Brivo 865 Advance |  |  | September 23, 2016 | 2016 | GE OEC Medical Systems, Inc | GE Healthcare Surgery announces a voluntary field action for the OEC Brivo 715 Prime, OEC Brivo 685 Essential, and OEC Brivo 865 Advance because the circuit board that controls vertical motion of the C-arm may fail, resulting in the C-Arm moving up or down without command. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter dated September 23, 2016 to customers. The letter informed customers of the affected product, safety issue, and safety instructions for immediate mitigation of the issue. GE Healthcare will correct all affected products at no cost to customers. A GE Healthcare representative will the customers to arrange for the correction. Please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 148 | 148 total | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, FL, KS, MS, NJ, NV, NY, PA, TX, UT, and WV. No Candaian or Va/govt/military consignees. There are 133 other foreign consignees. | Terminated | April 24, 2017 | 213 | 148 | Not_Computer
Z-0068-2017 | 75273 | No | OMT | Absorbable lung biopsy plug. | General &amp; Plastic Surgery | Anesthesiology | 510(k) | BioSentry Tract Sealant System, Product Code: 768022019S; Indicated to provide accuracy in marking a biopsy location for visualization during surgical resection. |  |  | September 24, 2016 | 2016 | Surgical Specialties Mexico S DE RL DE CV | Surgical Specialties Mexico is recalling the BioSentry Track Sealant System due to the potential for an open seal on the product package. | Process control | An Urgent Product Recall Notification Letter dated 9/24/16 was sent to customers to inform them that Surgical Specialties Mexico is recalling the BioSentry Track Sealant System due to the potential for an open seal on the product package. The letter informs the customers to complete the customer acknowledgement form and return via fax to their quality assurance department at (610) 404-4010 or return via email to kknappenberger@surgicalspecialties.com. Customers with questions are instructed to contact Kelly Knappenberger at (484) 220-2402 or Kris Liberacki at (262) 994-8083. |  | 512 | 512 boxes (2560 units) | US: OH, PA, CO, MA, DE, CA, NY, VA, WA, NJ, CT, FL, TX, NH, ID, OK, AZ, GA, PA, NV, VT.Worldwide: Spain, Italy, United Kingdom, Portugal, Turkey, Germany, Denmark, | Open, Classified |  |  | 512 | Not_Computer
Z-0070-2017 | 75279 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-arm O2 Surgical Imaging System Product Catalog Number: BI-700-02000 |  |  | September 22, 2016 | 2016 | Medtronic Navigation, Inc.-Littleton | O-arm O2 Surgical Imaging System Spatial calibration may be erroneous in Stealth Station navigated images | Software design | Medtronic Navigation contacted customers via telephone beginning September 23, 2016. Customers were informed of the affected product, problem and actions to be taken. Medtronic requests that users immediately discontinue the use of the O-Arm O2 Surgical Imaging system for automatic registration in Stealth Station navigated surgical procedures until further notice.Medtronic Service personnel have been scheduled to begin the necessary inspections and re-calibration activities on customer systems beginning on September 23, 2016. |  | 14 | 14 units | Worldwide Distribution - US Nationwide in the states of CO, KA, NE, MI, MO, OH, WA and country of Switzerland | Terminated | May 12, 2017 | 232 | 14 | Not_Computer
Z-0073-2017 | 75192 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fujifilm Synapse PACS, Software versions: 4.0.xxx, 4.1.xxx, version 4.2.xxx, version 4.3.xxx |  |  | September 15, 2016 | 2016 | Fujifilm Medical Systems U.S.A., Inc. | Image data for a patients image may not be correct | Software design | The firm, FMSU, sent a "Medical Device Correction" letter dated September 15, 2016 to its customers. The letter describes the product, problem and actions to be taken. The firm will install a patch that corrects the issue (Synapse IPSS v1.3.1.0 Patch).Customers with Version 4.0 may upgrade to version 4.4 (will not require the patch) or upgrade to 4.1, 4.2, or 4.3 and install the patch. FMSU will install the patch on the server remotely if your site uses Active Monitoring. Your notification that the patch was installed will be your service ticket. For sites not using Active Monitoring, FMSU will schedule a time to install the patch on your server.Questions or concerns, contact the FUJIFILM Customer Support Center at 1-888-FUJIMED (1-800-385-4633) Option #1. International callers can access the FUJIFILM Customer Support Center by calling 1-203-602-3580. |  | 251 | 251 units | US Nationwide Distribution. | Open, Classified |  |  | 251 | Not_Computer
Z-0072-2017 | 74162 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur XPT Immunoassay System |  |  | April 28, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | The ADVIA CentaurÂż XPT default setting for the Daily Maintenance Task (Daily Cleaning Procedure) frequency may have the Daily Maintenance Task frequency set to, As needed instead of Daily. Not performing the daily maintenance may impact any assay and can be detected through monitoring of quality control and calibration results. Systems running in a language other than English are not impacted. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated April 28, 2016, to all affected customers. The letter advised customers not to release results generated by the system if quality control results are not within acceptable ranges. Customers were advised to perform the following steps to check the Setting for the Daily Maintenance Task frequency on your instrument: *Open the Maintenance Tab, * Select Daily Maintenance task and * Verify that the Frequency Column is set to "Daily at..". The Daily Maintenance Task frequency setting can only be reset by a Siemens Customer Service engineer. If your instrument is set to "As Needed": Daily Maintenance Task Frequency Setting May be Incorrect. Customers should contact their local Technical Support Center to schedule a visit by your Siemens Customer Service engineer. Customers should also perform the daily maintenance by selecting the Daily Maintenance Task and then selecting Perform on the right bar. In addition, all customers were asked to complete and return the Field Correction Effectiveness Check within 30 days.For questions regarding this recall call 914-631-8000. |  | 292 | 292 units Total (3 domestically &amp; 289 internationally) | Worldwide Distribution - US including MD, MN, &amp; MS; **Internationally** AE, AT, AU, BE, BR, CA, CN, DE, DK, EG, ES, FI, FR, GB, GP, HU, IN, IT, JO, JP, KR, MY, NO, NZ, PH, PT, QA, SA, SE, SG, TR &amp; VN | Open, Classified |  |  | 292 | Not_Computer
Z-0078-2017 | 75350 | No | NAW |  |  | Radiology | PMA | TheraSphere(R) YTTRIUM-90 CLASS MICROSPHERE INJECTION |  |  | September 21, 2015 | 2015 | Biocompatibles U.K., Ltd. | One mislabeled dose of TheraSphereÂż Y-90 Glass Microspheres, which was shipped to Hong Kong (total radioactivity incorrect). The product ws labeled as 10 GBq and the actual does was 5 GBq. | Packaging | Biocompatibles UK Ltd sent an Urgent Field Safety Notice dated September 21, 2015, to all affected customers. The firm did not request return or destruction of the product. Customers with questions were instructed to call 613-410-1870.For questions regarding this recall call 613-801-1848. |  | 64 | 64 doses | Worldwide Distribution - US to CA, FL, IL, WI, IN, OR, PA, MA, NY, KY, TX, MI, MD, NV, WA, OH, MN, and Internationally to Germany, Turkey, Hong Kong, Italy, Portugal | Completed |  |  | 64 | Not_Computer
Z-0076-2017 | 75303 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | September 29, 2016 | 2016 | Elekta, Inc. | When creating 3D plans using either MU or Dose weighting modes, if the user changes the Physician's Intent Rx Dose and/or the number of fractions, and then modifies the wedge angle, the MU value is scaled incorrectly. | Device Design | Elekta sent an Important Field Safety Notice 382-01-MON-005 was sent to all affected customers on 9/29/2016. The notice informs users of the specific product and version numbers affected by the issue,and any work arounds that can be used to avoid the issue. The notice included an acknowledgment form which is to be returned.For questions regarding this recall call 314-993-0003 or 800-878-4287. |  | 365 | 365 | Worldwide Distribution - US to CA, FL, KY, MD, MO, MT, NJ, NC, PA, WA, WI. Internationally to Algeria. Antiqua &amp; Barbuda, Argentina, Australia, Austria, Belarus, Botswana, Brazil, Bulgaria, China, Columbia, Croatia, Czech Republic, Ecuador Estonia, France, Germany, India, Iran, Israel, Italy, Japan, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Portugal, Romania, Russia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Zimbabwe. | Open, Classified |  |  | 365 | Not_Computer
Z-0077-2017 | 75296 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Head Frame - Model number HHF -ALL Head Frame Posts (PN1008016 ) and Post Kits (PN1008016). |  |  | September 13, 2016 | 2016 | Varian Medical Systems, Inc. | Varian Medical Systems has received a report that a user was able to easily rotate the head frame posts when attached to the mounting cam on the metal head ring. There was no report of serious injury due to this issue | Device Design | Varian Medical Systems sent an Urgent Medical Device Removal letter dated September 14, 2016, to all affected sites. Each site was also contacted by telephone to advice users of the issue and request of return of the head frame posts on September 14 and 15th. Calls enabled Varian to prioritize replacement parts for those with patients scheduled in the immediate future.Recommended User Actions:1.CEASE USE of the Varian Head Frame Assembly and QUARANTINE the Varian Head Frame Assembly and components.2.Collect, Clean, Disinfect all Head Frame Posts [PN1008016] if used previously.3.Contact Varian 1.888.827.4265 for Return Material Authorization number, and return all Head Frame Posts in accordance with instructions provided with the enclosed Recall Return Material Form.We sincerely apologize for any inconvenience and thank you in advance for your cooperation. If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager. This notice has been provided to the appropriate Regulatory Authorities. For further questions, please call (650) 424-6833. |  | 75 | 75 devices | US Distribution and Internationally to 1 site in Canada, India and Morocco | Open, Classified |  |  | 75 | Not_Computer
Z-0082-2017 | 75037 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | SIEMENS IMMULITE/IMMULITE 2000 Systems Rubella IgM (RUM), REF/Catalog Number for US: L2KRM2(D), Siemens Material Number (SMN) for US: 10374021, REF/Catalog # for OUS: L2KRM2, SMN for OUS: 10381327; Unique Device Identification Number (UDI) 00630414961996; IVD --- |  |  | August 24, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens received multiple customer complaints indicating an increase in false positive and indeterminate patient sample results and increased imprecision for several lots of Rubella IgM assay on the IMMMULITE 2000/IMMULITE 2000 XPi platforms. | Under Investigation by firm | Siemens sent an Urgent Medical Device Recall Letter IMC16-20.A.US , dated August 24, 2016, to the US consignees via FedEx, and an Urgent Field Safety Notice IMC16-20.A.OUS (dated August 2016) was sent to foreign representatives for distribution to foreign consignees. Customers were instructed to discontinue use of and discard the affected kit lots. Customers were also requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within 30 days. Field service personnel will be sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For questions, customers were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative. |  | 3612 | Domestic: 177 Kits; Foreign: 3435 Kits | Worldwide Distribution n- US (Nationwide) and Foreign distribution to the following countries: Argentina, Bangladesh, Brazil, Bulgaria, China, Colombia, Costa Rica, Cyprus, Egypt, France, Germany, Greece, Guatemala, India, Indonesia, Israel, Italy, Kazakhstan, Macedonia, Mexico, Paraguay, Poland, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, South Africa, Syria, United Arab Emirates, and Uruguay. | Open, Classified |  |  | 3971 | Not_Computer
Z-0006-2017 | 74999 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD), US product: Model No.: 1103; and International Product: Model No.: 1104XX. |  |  | August 17, 2016 | 2016 | HeartWare, Inc | Foreign material found within the driveline connector of the HVAD Pump that could cause increased impedance resulting in electrical faults. | Component design/selection | HeartWare notified their customers with an Urgent Medical Device Recall notice on 8/17/2016 of the HVAD Pump, listing the US Product: Model No. 1103 and the International Product: Model No.: 1104. The letter discussed the reason for the recall as well as the risk to health. Customers were to identify affected product in their inventory; and, complete and return the acknowledgement form. Acknowledgement should be provided no later than 30 days from date of the letter either to the HeartWare representative; or, via e-mail to FSCA @Heartware.com; or, FAX the signed form to (305) 364-2665. Affected product should be returned upon receipt of replacement product; and, the completion form should also be completed and returned. Questions or concerns should be directed to a local HeartWare representative. |  | 344 | 344 pumps | Worldwide Distribution - US, in the states of AL, AR, AZ, CA, DC, FL, IL, IN, KY, LA, MA, MI, MN, MO, NC, NE, NJ, NY, OH, OK, PA, SC, TX, UT, VA, and WI; and, the countries of Argentina, Belarus, Bosnia, Herzegovina, Brazil, Canada, Chile, Colombia, Hong Kong, India, Israel, Kazakhstan, Kuwait, Lebanon, Malaysia, New Zealand, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Taiwan, Turkey, and Vietnam. | Open, Classified |  |  | 344 | Not_Computer
Z-0075-2017 | 75274 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Banyan Medical 5 x 55mm Pyramidal Tip Trocar w/ 2X Cannula w/ |  |  | September 07, 2016 | 2016 | Banyan Medical, LLC | Potential lack of assurance of sterility | Process control | Banyan Medical issued Urgent Voluntary Recall Notification on 9/7/16 to Distributors to remove product from inventory and contact hospital to examine and return product. Questions contact your local Banyan Sales Representative |  | 320 | 64 boxes of 5 | IN, TN | Open, Classified |  |  | 415 | Not_Computer
Z-0080-2017 | 75240 | No | EAJ | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Zero-Gravity Radiation Protection System |  |  | September 09, 2016 | 2016 | Contour Fabricators Inc | Two customer complaintssuggest that excessive rotation of the cable in one direction (clockwise or counter-clockwise)without allowing the cable to unwind to its normal condition may cause the cable to breaksuddenly and the body shield to fall. While no injuries or deaths occurred in either failure, theZero-GravityTM body shield weighs 55 pounds, and in the event of a cable failure, the bodyshield will fall immediately. A fall of this nature places the user, associated staff and the patientat risk for a potential injury. | Labeling design | On 9/9/2016, URGENT: MEDICAL DEVICE CORRECTION letters dated 9/8/2016 were sent to the affected consignees via certified mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. TIDI Products is taking the following action to correct this issue:1. Revising instructions for use to include the step to allow the cable to unwind after useand prior to storing the body shield in a locked or secured position.2. Modifying user and staff training protocols to include the requirement to allow the cableto unwind after each use.3. Adding additional warning labels to the system.4. Replacing all cables in the field to ensure that any potentially stressed cable is removed.A representative from TIDI Products will be contacting you to schedule a visit from a technicianto implement this corrective action.Customer ActionsCables should be carefully inspected for signs of wear or rotational tension, which will create abend or buckle in the cable. Practitioners and medical institutions should review their routineuse of the Zero-Gravity Radiation Protection System to determine if the body shield undergoesrepeated rotation without allowing the cable to relax back to its normal condition. If the cablehas been consistently rotated in one direction without being allowed to relax back to its normalcondition or if there is any doubt about excessive rotation in one direction, TIDI recommendsthat practitioners and medical institutions discontinue use until such time as corrective actioncan be implemented by TIDI Products.If facilities continue to use their Zero-Gravity" Radiation Protection System until correction andinstall of the replacement cable can be performed, they should make sure that they are allowingthe unit to go back to its relaxed state.Please complete and return the enclosed response for |  | 447 | 447 | Nationwide. Canada.AUSTRALIAAUSTRIABELGIUMBRUNEICHINAFRANCEGERMANYINDAINDONESIAISRAELITALYKSAMALAYSIANETHERLANDSPOLANDQATARQUEBECSAUDI ARABIA SINGAPORESOUTH AFRICATAIWAN (R.O.C.) | Open, Classified |  |  | 447 | Not_Computer
Z-0079-2017 | 74750 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System, Model 5.0 |  |  | May 30, 2016 | 2016 | RAYSEARCH LABORATORIES AB | For a treatment plan consisting of multiple beam sets, the table for ROI plan dose statistics in the report may show the statistics for a beam set dose. The error can only be triggered when using a report template where statistics for a beam set dose is included as the last dose statistics prior to the plan dose statistics. | Software design | On May 19th, 2016, RaySearch Laboratories distributed Field Safety Notice, Medical Device Correction notices and reply forms to their customers via email. Actions to be taken by the customer include: 1) Inspect all report templates that are intended to be used for plans with multiple beam sets. 2) Make sure that POI Dose statistics (Plan) is included immediately before ROI Dose statistics [Plan dose] in the Report content area. 3) Please educate planning staff and all users about this workaround. 4) Inspect your product and identify all installed units with the above software version number. Finally, complete and return the reply form via email to: freddie.cardel@raysearchlabs.com or by fax to: 888-501-7195 to confirm that you have read and understand the information. Customers with questions can contact David Hedfors via email: david.hedfors@raysearchlabs.com |  | 491 | 491 units (187 domestically &amp; 304 internationally) | AZ, CA, FL, MT, NC NY, TX, OH &amp; WA | Open, Classified |  |  | 491 | Computer
Z-0087-2017 | 75229 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Uniflex Humeral Nail: 467265 G-PIN 2.8BL 3.2PL 60CM 2PK |  |  | August 02, 2016 | 2016 | Zimmer Biomet, Inc. | The G-PIN 2.8BL 3.2PL 60 CM 2PK is intended to contain one (1) 467265-01 Ball Tip Guide Pin and one (1) 467255-02 Plain Tip Guide Pin. The G-PIN 2.8BL 3.2PL 60 CM 2PK may have been packaged incorrectly either with two (2) 467265-01 Ball Tip Guide Pins and zero (0) 467255-02 Plain Tip Guide Pins or with zero (0) 467265-01 Ball Tip Guide Pins and two (2) 467255-02 Plain Tip Guide Pins.If the G-PIN 2.8BL 3.2PL 60 CM 2PK contains the incorrect combination of Plain Tip Guide Pins and Ball Tip Guide Pins, the surgeon would either have to implant the nail without the Plain Tip Guide Pin as a guide, ream with a Plain Tip Guide Pin, or choose to not ream the humeral canal. The risk associated with not having the correct combination can include a delay in surgery of less than 30 minutes, the reamer tip remaining in the distal humeral canal, or an iatrogenic fracture. | Mixed-up of materials/components | The firm, Zimmer Biomet, sent an " URGENT MEDICAL DEVICE RECALL REMOVAL" notification dated 8/2/2016 affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The consignees were instructed to review the notification and ensure affected team members are aware of the contents; immediately locate and quarantine affected product in your inventory; complete and return the Certification of Acknowledgment portion of Attachment-1 Inventory Return Certification Form via digital copy within three (3) days via email to CPWARFieldAction@zimmerbiomet.com or via fax to 574-372-1683;Immediately return all affected product from your distributorship and affected hospitals within your territory; and retain a copy of your recall acknowledgement and product return form.If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of business hours will receive a prompt to record a voicemail and will be returned upon receipt. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 22 | 22 units | US Distribution to states of: MD,TX, FL, GA, VA and PA. | Terminated | March 24, 2017 | 234 | 22 | Not_Computer
Z-0004-2017 | 75181 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Skintact Electrodes for Defibrillation, DF29N. |  |  | September 01, 2016 | 2016 | Leonhard Lang Medizintechnik GmbH | There is a risk that defibrillation electrodes model DF29N will be connected with the defibrillator Welch Allyn AED 10 only with delay or not at all. This may cause a situation in which a patient, who is in a life threatening condition and requires adefibrillation shock, cannot be treated in good time. | Device Design | Consignees was sent via e-mail a Leonhard Lang "Important Safety Notice" dated September 1, 2016. The letter described the product being recalled, Description of the defect, Actions and time frame of the recall, &amp; Compensation for the recalled electrodes. Advised consignees to inform all users within their organizations of the recall; and, secure and destroy the unused electrodes in their inventory. Confirmation of destruction can be done by completing and returning the "Confirmation of Destruction / Consumption" form by October 14, 2016. If the product was further distributed, they were to forward a copy of the Safety Notice. For questions contact sales staff. |  | 11110 | 11,110 ( US 8,040; OUS 3,070) | Worldwide Distribution - US, to the state of Florida; and, the countries of South Africa, Germany, Slovenia, France, United Arab Emirates, Great Britain, Italy, Israel, Serbia, Thailand, Thailand, South Africa, Lebanon, and Poland. | Open, Classified |  |  | 11110 | Not_Computer
Z-0084-2017 | 75234 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Phoenix Recon Targeting Arm |  |  | August 09, 2016 | 2016 | Zimmer Biomet, Inc. | Old revisions and new revisions of the soft tissue sleeve and recon targeting arm are not interchangeable, as the new revision tissue sleeve will not fit into the old revision targeting arm. The old revision tissue sleeve has too much clearance with new revision targeting arm.Risks of encountering the affected product include: the correct instrument may not be readily available for the procedure causing the surgeon to complete the procedure by hand; or a delay in surgery greater than 30 minutes may occur while another Recon Targeting Arm or Recon Soft tissue Sleeve is located or while the surgeon completes the procedure by hand. | Finished device change control | On 8/9/2016, URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions or concerns call the customer call center at 1-800-348-9500 ext 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 31 | 31 | US Nationwide Distribution in the states of: MI, UT, SD, OH, WA, KS, MO, NJ, GA, and TX. | Terminated | March 23, 2017 | 226 | 31 | Not_Computer
Z-0085-2017 | 75056 | No | JOS | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CONCEPT HEATWAVE Electrode 130 mm x 2.3 mm diameter, REF/Catalog Number CS-023, STERILE, Rx Only |  |  | August 29, 2016 | 2016 | ConMed Corporation | CONMED received complaints of damage to the CS-023 electrodes shaft insulation which could result in burn marks on the shaft. If the shaft insulation is damaged in any way, it could increase the risk of patient burn. | Process change control | Conmed sent an URGENT DEVICE RECALL letter dated August 31, 2016 and Business Reply Form were sent to the domestic consignees on August 29, 2016 via USPS First Class Mail. The letter identified the affected product, problem and actions to be taken. For questions or requests contact the Field Action Support Team at 1-800-535-8536, fax to 315-624-3225, or email largheatwv@conmed.com. |  | 4770 | Domestic: 1740; Foreign: 4770 | Worldwide Distribution - US Nationwide in the states of: CA, HI, IL, IN, KS, LA, MA, NC, NM, NY, OH, OK, PA, TN, TX, UT, WA and WI; plus Foreign distribution to Australia, Canada, France, Germany, Italy, Spain, Czech Republic, Lithuania, Netherlands, Poland, Korea,Morocco, Malaysia, Mexico, Taiwan, and South Africa. | Open, Classified |  |  | 4770 | Not_Computer
Z-0005-2017 | 72001 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System Controller |  |  | June 08, 2016 | 2016 | HeartWare Inc | Loose connector ports in controller body. | Under Investigation by firm | HeartWare sent an Urgent Medical Device Correction letter dated June 8, 2016, to all affected customers, HeartWare advised clinicians to remind patients to follow unit instructions and water avoidance, and to regularly schedule appointments to check the integrity of the unit and to be cautious with connectors when inspecting connectors. Clinicians were also instructed to send back the enclosed "Acknowledgement Form" within 30 days and to forward the message to necessary individuals within their organization.Customers with questions were instructed to contact their local HeartWare representative.For questions regarding this recall call 305-364-1575. |  | 13363 | 8799 Worldwide, 4,564 in US | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and countries of Australia, Austria, Belgium, Canada, France, Germany, Lebanon, Malaysia, Sweden, Switzerland, &amp; United Kingdom. | Open, Classified |  |  | 13363 | Not_Computer
Z-0090-2017 | 74950 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products LAC Slides, REF/Product Codes 843 3880 (300 slides per sales unit) and 815 0112 (90 slides per sales unit), Unique Device Identifier Nos. 10758750004911 and 10758750004454, Rx ONLY, IVD --- |  |  | August 11, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics confirmed that N-Acetylcysteine (NAC), a drug used for acetaminophen overdose, mucolytic therapy and is used as nutritional supplement, can interfere with VITROS GLU, LAC, TRIG, and URIC Slides assays. | Under Investigation by firm | Ortho-Clinical Diagnostics sent an URGENT PRODUCT CORRECTION Letter (Ref. CL2016-152) dated August 11, 2016 via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to all customers who could order and use affected product to inform them that new interferant were identified for GLU, LAC, TRIG and URIC Slides which may cause biased results. Foreign affiliates were informed by e-mail on August 11, 2016, of the issue and instructed to notify their consignees of the issue. For questions, please contact our Ortho CareTM Technical Solutions Center at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 62013 | US: 32,081 sales units; Foreign: 29,932 sales units | Worldwide Distribution - US (Nationwide) Puerto Rico and Bermuda), Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Poland, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 805842 | Not_Computer
Z-0093-2017 | 75163 | No | ORZ | Eyelid thermal pulsation system. | Ophthalmic | Ophthalmic | 510(k) | LipiFlow thermal Pulsation System, consisting of Console and Activator, Model LFD-1100 |  |  | September 13, 2016 | 2016 | Tearscience, Inc | Complaints related to an error condition occurring during patient treatment where the FipiFlow System automatically stopped treatment after detecting a problem with the Activator. | Device Design | The firm, TearScience, sent a "Medical Device Recall" letter dated September 9, 2016 to its consignees/customers on 9/13/2016. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the enclosed Acknowledgement and Receipt form by fax to (919) 882-9603, email to: Customerservice@tearscience.com or mail to TearScience, ATTN: CUSTOMER SERVICE RCL001, 5151 McCrimmon Parkway, Ste 250, Morrisville, NC 27560; to discontinue use of the product; to remove all a product with lot #201615220020 from your inventory; and call TearScience Customer Service at 919-459-4891 to return the product for above mentioned lot# and request new product for replacement.For any questions regarding this letter, please call (919) 459-4891 Monday through Friday, 8:30AM to 5:30PM, Easter Time. |  | 12000 | 1,200 Activators (120 boxes of 10 units) | US Distribution including states of: AK, AR, CA, CO, FL, GA, ID, IL, IN, LA, MD, MA, MN, MO, NV, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA , WA, WI and Puerto Rico. | Open, Classified |  |  | 12000 | Not_Computer
Z-0088-2017 | 75257 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SDMS (SenTec Digital Monitoring System) |  |  | September 15, 2016 | 2016 | SenTec AG | Sentec AG found that with one batch of an internal SenTec Monitor component (Docking Station Module) there is a risk that a fastening screw could become loose under certaincircumstances. A loose screw within the monitor can create uncontrolled electrical connections, including short circuits, with the risk of smolder or fire and electrical shock (patient and/or operator). | Component change control | Distributors were sent via e-mail a SenTec 'Mandatory initiated field correction action" letter dated September 15, 2016 . The letter provided a "Description of the issue, Affected products, Identify and localize affected products, Removal from use, Replacement of affected products at end customer &amp; distributor / sub-distributor level". It also provided "action information for sub-distributors and affiliates, and end customers". For questions contact SenTec Inc. by e-mail (customersevice.usa@sentec.com) or phone (636-343-0000). |  | 21 | 75 (54 US; 21 =OUS) | Worldwide Distribution - US Distribution to the states of : LA, MI, MO, OH, TN. and to the countries of : United Kingdom, The Netherlands and Australia. | Terminated | June 23, 2017 | 281 | 21 | Not_Computer
Z-0118-2017 | 75161 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software |  |  | October 19, 2011 | 2011 | Merge Healthcare, Inc. | Customers may experience an issue with the software study preferences when changes are made to the study protocol, resulting in incorrect patient follow-up. | Software design | Merge sent an Urgent Medical Device Recall dated January 28, 2016 to affected customers via email and certified mail if no email address was available. The letter identified the affected product, problem and actions to be taken. For questions Customer Service at 877-741-5369 or support@merge.com. |  | 844 | 844 sites potentially have the affected versions | US Nationwide Distribution | Open, Classified |  |  | 844 | Computer
Z-0113-2017 | 75119 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Cytoplast Titanium-Reinforced Posterior Singles Membrane |  |  | September 07, 2016 | 2016 | Osteogenics Biomedical, Inc. | The device has the potential to be packaged in the wrong product box. | Labeling mix-ups | Osteogenics Biomedical sent an Urgent Customer Notification letter dated September 7, 2016, to all affected customers. This recall has been extended to the user level (clinician). All units sold domestically were to the end users (clinicians). End users were instructed to inspect inventory and return any affected product.For further questions, please call 1-888- 796-1923. |  | 32 | 32 boxes | Worldwide Distribution - US Distribution and to the countries of : Colombia, Venezuela/Ecuador, Japan, India, Canada, and UK. | Open, Classified |  |  | 32 | Not_Computer
Z-0136-2017 | 75255 | No | JEC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Agee-WristJackÂż Fracture Reduction System; |  |  | September 14, 2016 | 2016 | Hand Biomechanics Lab Inc | A number of Agee WristJackÂż Fracture Reduction Systems have been manufactured with a molded beam slightly larger than optimal. This may result in the surgeon experiencing difficulty fine-tuning the length (red) and radial/ulnar alignment (blue) adjustments on the device. | Under Investigation by firm | Hand Biomechanics Lab, Inc. sent a Voluntary Medical Device Recall letter to all affected customers beginning on September 14, 2016 and sent by phone/fax numbers - or e-mail. Letters request return of devices which have not yet been used. Customers with questions were instructed to contact Customer Service at (800) 522-5778.For questions regarding this recall call 916-923-5073. |  | 65 | 65 units | Nationwide Distribution | Open, Classified |  |  | 65 | Not_Computer
Z-0135-2017 | 74774 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | COREÂż Trumpet Handpiece only, Catalog Number CD8450 |  |  | June 27, 2016 | 2016 | ConMed Corporation | The packaging seal may contain a crease. The crease may result in an open channel. These channels may compromise sterility of the product. | Under Investigation by firm | ConMed sent an Urgent Device Recall letter dated June 27, 2016, to all affected customers with reply forms via USP First Class Mail. International customers received a notice dated July 11, 2016 by courier service. Customers were advised to inspect their inventory for any of the devices subject to recall. Customers were urged to contact all departments within their facility and any other facilities within their organization that may have received affected products. It is imperative that all end users of these devices receive the notification. Customers who have received any of the affected product are asked to complete and return the reply form along with the device to: CONMED Corporation 525 French Road Utica, NY 13502 USA Attn: Ed Kovac. Please do not return used devices. Replacement product will be issued unless otherwise requested. If customers do not have any affected devices to return, complete and return the business reply form indicating they have no devices and return by one of the means listed:1. Email to: coresihp@conmed.com2. Fax to: Field Action Support Team at 315-624-3225. 3. US mail, using enclosed postage-paid envelope, to: CONMED, 525 French Road, Utica, NY 13502; Attn: Field Action Support Team. Customers with questions or requests, should contact the Field Action Support Team at 315-624-3237, fax to 315-624-3225, or email coresihp@conmed.com. |  |  |  | Worldwide Distribution - US (Nationwide) Internationally: to Australia, Austria, Uruguay, Denmark, Germany, Ireland, Italy, United Kingdom, Guatemala, Czech Republic &amp; Canada. | Open, Classified |  |  |  | Not_Computer
Z-0123-2017 | 75207 | No | CDT | Triglyceride test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Randox Lipase (Colorimetric) |  |  | August 02, 2016 | 2016 | Randox Laboratories Ltd. | According to the firm carry over was observed when the Lipase test is run directly before or after the LDL and Triglycerides assays. Customers are instructed to amend the running order for all instruments using these reagents to ensure Lipase is not place directly before or after LDL or Triglycerides. A correction was made to the IFU for the Lipase assay and all customers who purchased the Lipase kit in the USA were contacted with the updated instruction. | Nonconforming Material/Component | Randex Laboratories sent an Urgent Field Safety Notice letter dated June 17, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the response section of this form technical.services@randox.com. For questions contact Randex Technical Service. |  | 18 | 18 | US Nationwide Distribution | Open, Classified |  |  | 18 | Not_Computer
Z-0125-2017 | 74397 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Calix P PEEK Lumbar System, PLIF Trial |  |  | May 17, 2016 | 2016 | X Spine Systems Inc | The trials and rasps used to prepare the surgical site for placement of the implant may become detached from the inserter assembly part and become lodged in the intervertebral space. | Component design/selection | Consignees will be notified via email and courier and will be instructed to return the affected product. Upon receipt of the recalled trays, the trials and rasps will be removed from the tray and quarantined for evenutal destruction. |  | 1504 | 1504 units | Product was shipped to the following states: AZ, CA, CO, FL, GA, ID, IN, KY, LA, MD, MI, MO, NC, NV, NY, PA, PR, TX &amp; WV.Product was also shipped to the following countries: Australia, England, Germany, Italy, Mexico, Portugal &amp; United Kingdom. | Open, Classified |  |  | 3067 | Not_Computer
Z-0112-2017 | 75018 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Sagittal Blade 18.0X0.97X90MM, Model Number: 6118-097-090 |  |  | August 17, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the Sagittal Blade 18.0X.97X90MM due to the potential for a sterile barrier breach in the product's packaging. | Packaging | Stryker Instruments is voluntarily recalling the Sagittal Blade 18.0X.97X90MM, product number 6118-097-090, lot number 16029017, due to a breach in the packaging causing a breach in sterility. There is a potential for a bone or soft tissue infection which may require medical or surgical intervention. Customer notices were issued via certified mail on 08/17/2016. Actions to be taken by the Customer/User:1.Immediately review this Recall Notification. 2.Check all stock areas and/or operating room storage to determine how many Sagittal Blades from the affected lot (listed above) are at your facility. Quarantine and discontinue use of the recalled Sagittal Blades.3.Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if customers dont have any affected product on hand. Note: Customer signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4.If customers have further distributed this product, please forward the recall letters and the attached BRF to all affected locations. Please indicate each location on the BRF. 5.Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com.6.If the BRF for customer facilities indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. 7.Upon receipt of the recalled product, replacement product will be shipped to customer accounts.For questions regarding this recall, please contact Stryker Instruments: Kelly Jo Davis, 269-389-2921, kellyjo.davis@stryker.com |  | 225 | 225 | US Nationwide in the states of RI, WA, CA, CO, WI,FL, ID, ND, MD,TXForeign: Australia, New Zealand, Netherlands, France, Italy | Terminated | March 29, 2017 | 224 | 225 | Not_Computer
Z-0122-2017 | 74918 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | 25 gauge VersaVit 2.0 Vitrectomy Cutter, Product/Part Number 70935, packaged in a Tyvek/Mylar double pouch configuration. |  |  | August 11, 2016 | 2016 | Synergetics Inc | Vitrectomy cutters have been found not to be operating within the correct pressure pulse per specification. | Other | A customer notification letter was sent out on 8/11/2016 to the hospital/surgical center level. Customers were asked to fill out and return a recall acknowledgment form and send back any affected devices. |  | 9 | 9 units | Devices were distributed to KY, VA, TN, OH, IN, CA, NJ, NY, TX, SC, AL, MS, MI, and UT. Foreign distribution to the UK, France, Mexico, and India. | Terminated | April 27, 2017 | 259 | 408 | Not_Computer
Z-0111-2017 | 75237 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Distal Femoral Augment 75 X 5 RL/LM |  |  | August 09, 2016 | 2016 | Zimmer Biomet, Inc. | Contents of the package do not match the product labeling.Product is labeled: Vanguard Distal Femoral Augment 75 X 5 RL/LMCarton contains: Vanguard Distal Femoral Augment 75 X 5 LL/RMRisks:Adverse health outcomes that could occur as a result of this event are a delay less than 30 minutes and a different size augment being used to complete the procedure.As the augment will not assemble with femoral component due to its shape, it would be easily identified during surgery. | Labeling mix-ups | On 8/9/2016, URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notice you have further questions or concerns please call 1-800-348-9500 ext. 1251 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of operating hours will receive a call back the next business day. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. |  | 5 | 5 | Worldwide Distribution - US Nationwide in the states of NJ, MI, FL and the countries of Sweden and Canada. | Terminated | May 17, 2017 | 281 | 5 | Not_Computer
Z-0128-2017 | 75250 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | CD79b PE (ASR) , Catalog number 335816 ; |  |  | September 21, 2016 | 2016 | Becton, Dickinson and Company, BD Biosciences | CD79b (SN8) PE is contaminated with a FITC conjugate of unknown antigenic specificity | Under Investigation by firm | BD Life Sciences sent an Urgent Product recall letter dated September 2016 on September 21, 2016 ,to all affected customers via UPS and/or e-mail. Customers were instructed to inspect their inventory to determine if they have any of the affected lots and to discard the affected lots, and BD will issue replacement product. Customers were also instructed to complete the attached Customer Response Form whether or not they have any affected product. Customers needing further assistance should contact BD Customer Support at 855-236-2772 (prompt 3). For customers outside the US, contact their local BD Biosciences representative or distributor.For questions regarding this recall call 408-954-6307. |  | 14 | 14 units | Nationwide Distribution to CA, MN, OH and TX. | Open, Classified |  |  | 14 | Not_Computer
Z-0044-2017 | 75053 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Willy Rusch GmbH Tracheostomy Tube Set |  |  | August 26, 2016 | 2016 | Teleflex Medical | The connector may disconnect from the tracheostomy tube during use. | Use error | Teleflex sent an Urgent Field Safety Notice dated August 26, 2016, to all affected consignees. The letter requested that consignees cease use and distribution of stock, quarantine immediately, and return the product. Also, the letter requested a sub-recall if the product had been further distributed. The letter included an Acknowledgement Form which is to be returned. Customers with questions were instructed to contact their local sales representative or Customer Service.For questions regarding this recall call 610-378-0131. |  | 1095 | 1095 units | Worldwide Distribution: US (nationwide) to CA and countries of: Austria, Belgium, France, Germany, Italy, Japan, Poland, and Turkey. | Terminated | March 21, 2017 | 207 | 1095 | Not_Computer
Z-0138-2017 | 74542 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OER-Pro Endoscope Reprocessor, Endoscope washer/disinfector |  |  | June 23, 2016 | 2016 | Olympus Corporation of the Americas | New warning requiring Users to conduct duodenoscope precleaning and manual cleaning even when using an Automated Endoscope Reprocessor (AER) that has instructions that may indicate a user could forego certain steps in precleaning and manual cleaning of the endoscopes. | Under Investigation by firm | Olympus America Inc. (OAI) mailed a letter to customers informing them that they are issuing an updated Operation Manual for the OER-Pro Automated Endoscope Reprocessor.The following actions were asked to be taken:-Olympus has discontinued previously distributed copies of the OER-Pro Operation Manual. - Inspect your inventory of Operation Manuals and discard any existing inventory of OER-Pro Operation Manuals. - Implement use of the enclosed OER-Pro Operation Manual, which contains the new Warning on pages 73 and 80.- Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the requirement to perform precleaning and manual cleaning of Olympus duodenoscopes prior to placing the duodenoscopes in the OER-Pro for high level disinfection. - Additional copies of the new OER-Pro Operation Manual can be obtained by contacting our Technical Assistance Center at 1-800-848-9024, option 1, or by indicating on the enclosed questionnaire. Additional Operation Manuals will be mailed to your facility - Please indicate on the enclosed questionnaire that you have received this notification. Fax the completed form to (484) 896-7128. |  | 2686 | 2, 686 units | Distributed to: AK,AL,AR,AZ,CA,CO,CT,DC,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NC,ND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,TN,TX,UT,VA,VT,WA,WI,WV,WY | Open, Classified |  |  | 2686 | Not_Computer
Z-0289-2017 | 74590 | No | FBK | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Vari-Safe Injection Needle US endoscopy 5976 Heisley Road Mentor, Ohio 44060 |  |  | May 18, 2016 | 2016 | US Endoscopy Group Inc | US Endoscopy is conducting a voluntary product recall of one lot of the Vari-Safe Injection Needle as it was identified that the lot was incorrectly distributed. The lot was assembled with internal catheter components which differ from the original Vari-Safe needle internal catheter components. All components are dimensionally similar to the original Vari-Safe needle internal catheter components and the components have history of use on other US Endoscopy needle products. | Other | US Endoscopy sent an Urgent- (Vari Safe) injection needle recall letter dated July 6, 2016.US Endoscopy requests that you:1.Quarantine the affected Vari-Safe Injection devices in your stock.2. Destroy each affected device by opening the device package, cutting the device catheter with a scissors, and then discarding the device and package.3.Complete the attached Customer Response Card and return to the attention of US Endoscopy Customer Service, via fax 1-440-639-4495 or email orders@us endos copy.com .A no charge replacement will be provided for each unused Vari-Safe Injection Needle device from the affected lot. We apologize for any inconvenience this matter may cause you, and as always, US Endoscopy is dedicated to supporting our products and valued Customers. For further questions please call (800) 769-8226. |  | 36 | 36 units | US Distribution to the states of : OH, OR, WA and Internationally to the country of Italy. | Terminated | May 10, 2017 | 357 | 36 | Not_Computer
Z-0137-2017 | 75239 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CADstream software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | The software was not presenting the PURE filter user interface message on DICOM transfer studies. If a site starts using PURE filter, the user interface message will not display. This may result in incorrect patient follow-up. | Software design | Merge Healthcare Inc. sent an Urgent Medical Device Recall letter dated January 28, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. In addition, the customer is requested to call Customer Service or send them an email to schedule the upgrade. If the customer chooses not to upgrade, they are requested to complete the attached form or go online and complete it.Merge Healthcare is committed to improve efficiencies and enhancing the quality of healthcare worldwide.If you have any additional questions, call Merge Customer Service at (877) 741-5369 or support@merge.com. |  | 844 | 844 sites potentially have the affected versions | Worldwide Distribution - USA (nationwide) Distribution. Foreign/military/government distribution was also made. | Open, Classified |  |  | 844 | Computer
Z-0262-2017 | 71968 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Plastic Kit, Kit number CMP1084(B |  |  | April 22, 2015 | 2015 | Windstone Medical Packaging, Inc. | The Medtronic COVER LIGHT HANDLE DISPOSABLE GLOVES, Part # 13140, of the custom AMS kit were recalled because gloves may contain splits or holes compromising sterility. | Unknown/Undetermined by firm | Aligned Medical Solutions (AMS) sent the Medtronic Urgent Medical Device for Devon Light Glove letter, dated April 23, 2015 to all affected consignees.Consignees were informed that the sterile kits assembled and distributed by AMS contain the recalled Devon COVER LIGHT HANDLE DISPOSABLE Glove (AMS part #13140 or Medtronic part # 571711) due to splits or holes. Consignees were advised to identify any remaining affected kits and follow the over-labeling instruction. At the time the kit is opened for use, any light gloves should be identified and disposed of pursuant to the medical waste policies in effect at your institution. AMS can provide replacement product if needed or you can contact your local Medtronic sales rep for sterile replacements as well.Due to the manufacturing process and finished sterilized packs in inventory, or at your facility, over labeling of inventory and/or future shipment affected by this recall will be labeled with a sticker attached to each pack. This sticker will read:RECALL NOTICEMedtronic announced the recall l of Devon Light Glove contained in this kitafter it was sealed and sterilized. All other components in the kit are notaffected by this recall.1. At the time the kit is opened for use any Light Glove manufactured by Medtronic should be identified and set aside2. The recalled light glove should be rendered unusable to protect against inadvertent use and disposed of pursuant to the medical waste policies in effect at your institution.Consignees should complete the attached Reply Form and fax to AMS at 406-256-9875.For any questions about this recall, please call 321-527-7714.For questions regarding this recall call 404-259-6387. |  | 1164 | 1164 Kits | Nationwide Distribution to California, Florida, Georgia, Illinois, Kansas, Louisiana, Michigan, New Jersey, Nevada, Texas, Virginia, and Wyoming. | Completed |  |  | 104187 | Not_Computer
Z-0290-2017 | 75242 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | RTI Surgical Streamline OCT Occipito-Cervico-Thoracic Spinal Rod, CoCr, 3.2mm (D) x 240mm (L) |  |  | September 16, 2016 | 2016 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | RTI Surgical Inc. (RTI) is conducting a voluntary recall on two specific lots of the StreamlineÂż OCT StraightRod, CoCr. 3.2mm (D) x 240mm (L) due to the potential for small cracks to be present on the rods. The small cracks on the rods used in implants could cause reduced strength of these affected implants, and lead to patient injury. | Process control | RTI issued e-mail cutsomer notices for their voluntary recall of lots 234829 and 234829 of the StreamlineÂż OCT StraightRod, CoCr, 3.2mm (D) x 240mm (L) on 09/16/2016. RTI Surgical has produced replacement implants that have been verified as not containing any cracks. These implants will be sent to customers once they have identified how many replacementsare needed. Please follow the below instructions.1. Review inventory located in your facility and any other point of use locations that you may be responsible for. Immediately quarantine any implants listed in this recall notice. Please refer to the pictures included in customer notices for the etch location of the lot numbers for these implants:2. Contact the RTI Surgical Field Inventory representative Jennifer Mannor so that replacements can besent and an RGA and return shipping labels can be created if needed.3. Review, complete and return the attached recall reply form to Dan Nelson (dnelson@rtix.com) in accordance with the directions on the business reply form within 5 business days of receipt of this notification.Should customers have any questions please do not hesitate to contact RTI Surgical Field Inventory representative Jennifer Mannor (Email: jmannor@rtix.com Phone: 906-225-5872. |  | 66 | 66 | US Nationwide Distribution in the states of: AL, CA, CO, FL, IN, IL, MI, MO, MS, NC, NY, OK, TN, TX, and VA,. | Terminated | June 01, 2017 | 258 | 66 | Not_Computer
Z-0296-2017 | 74701 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Endoscopic Articulating Linear Cutter, 38ATW35. VASCULAR/THIN 35MM |  |  | July 15, 2016 | 2016 | Ethicon Endo-Surgery Inc | Quality control identified a component issue where the pinion gear in the device could fail under extreme use cases. If this condition occurs, staples will be formed past the cut line and the device can be opened and removed from the patient. The firing stroke may be interrupted and the knife may not fully return to the home position potentially exposing the healthcare professional to a sharps injury. | Device Design | Ethicon sent an Urgent Medical Device Recall letter dated July 15, 2016, to all affected consignees via UPS next day mail. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to examine their inventory immediately and quarantine any affected product and return to ENDOPATH. If product subject to the recall was further distributed consignees were instructed to notify the other facility and arrange for return of the affected product. Customer were instructed to complete the Business Reply Form confirming receipt of this letter and fax to Stericycle at 1-888-228-5719 or email to ethicon5182@stericycle.com.Customers with questions were instructed to contact the Ethicon Customer Support Center at 1-877-384-4266.For questions regarding this recall call 513-337-3582. |  | 213 | 213 | Worldwide Distribution - US including AL, AR, AZ, CA, CO,CT, DC, FL,GA, IA,IL,KS,KY,LA,MA,MD, ME, MI,MN,MO, MS, MT,NC, ND, NH,NJ, NM, NV, NY, OH, RI, TN, TX, VA, VT, WA, WI and Internationally to Brazil and Canada. | Open, Classified |  |  | 318 | Not_Computer
Z-0055-2017 | 75243 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions Twin-Pass .023" Dual Access Catheter, model 5230. Sterilized with Ethylene Oxide. |  |  | September 16, 2016 | 2016 | Vascular Solutions, Inc. | Vascular Solutions is recalling Twin-Pass dual access catheters because there is a potential for excess manufacturing material to remain at the tip of the catheter or within the distal portion of the rapid exchange lumen. | Process control | Vascular Solutions sent an Urgent Medical Device Recall letter via first class mail beginning 9/16/2016. The letter identified affected product, stated the issue, and instructed that product should not be used and to be returned to Vascular Solutions. A response form was requested to be returned. Questions can be directed to Customer service by e-mail at customerservice@vasc.com. |  | 11670 | 11,670 units total | Worldwide Distribution: US (nationwide) including Puerto Rico and country of Italy. | Open, Classified |  |  | 11670 | Not_Computer
Z-0291-2017 | 75209 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue MX40 Patient Monitor: IntelliVue MX40 WMTS Smart-hopping (1.4GHz) |  |  | September 19, 2016 | 2016 | Philips Electronics North America Corporation | Philips IntelliVue MX40 Patient Wearable Monitor Configuration Setting Disables Generation and Delivery of ECG Alarms to PIIC iX | Software design | The firm, Philips Healthcare, sent a "Field Safety Notice (FSN)" dated 9/14/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to make sure that the MX40 Patient Wearable Monitor local configuration setting 'Alarms On' is configured to its factory default until your software is upgraded; and follow the instructions listed in the letter. A Philips Healthcare representative will contact customers with affected devices to arrange for the installation of the software upgrade.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377. |  | 9804 | 9804 units | Worldwide Distribution-US (Nationwide) and countries of: Aruba, Australia, Austria, Belgium, Bulgaria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Ghana, Greece, Hong Kong, Iceland, India, Iraq, Ireland, Italy,Japan, Korea, Republic of, Malaysia, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Singapore, Slovenia, South Africa, SpainSweden, Switzerland, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 13840 | Computer
Z-0294-2017 | 75425 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Fusion Workstation.; Indicated for the transmission and review of radiological images. |  |  | October 28, 2011 | 2011 | Merge Healthcare, Inc. | After a period of time running Fusion Workstation, the Hounsfield measurement tool will report incorrect values. | Software design | The recalling firm initiated phone calls to affected sites beginning 1/30/2016. |  | 13 | 13 sites potentially hae the affected version | Distribution was made to AZ, CA, ID, IL, IN, MI, MN, and NJ. | Open, Classified |  |  | 13 | Not_Computer
Z-0298-2017 | 74809 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Neptune Plus (Hemostatic Pad), Part Numbers: 8870-03 |  |  | July 07, 2016 | 2016 | TZ Medical, Inc. | Neptune Plus was mis-labeled and included the statement WITH ANTIMICROBIAL BARRIER . This statement should not appear on label. | Labeling Change Control | TZ Medical notified their consignees by telephone on July 7, 2016.Customers with the boxes with the identified statement, "With Antimicrobial barrier, the firm plans to overlabel the statement on-site. If customer does not have the mis-labeled box upon contact, no action is planned. The firm also determined to overlabel devices on-hand.For questions regarding this recall call 503-639-0282. |  | 1410 | 1410 (141 boxes of 10) | Nationwide Distribution to AL, CA, FL, GA, IL, IN, KY, MO, MS, NC, NV, NY, PA, SC, TN, TX, and WV. | Open, Classified |  |  | 1410 | Not_Computer
Z-0297-2017 | 75376 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | The PAXWIRE Occlusion Balloon System consists of an occlusion catheter, inflation handle and syringes. It is designed for use in the iliofemoral artery to provide temporary occlusion as well as a path for dilator and/or sheath introduction, removal or repositioning during catheterization procedures. |  |  | October 06, 2016 | 2016 | AccessClosure, Inc., A Cardinal Health Company | AccessClosure, Inc. (ACI) is voluntarily recalling the PAXWIRE Occlusion Balloon System device because it was made aware from customers that the device may be removed in the absence of vascular sheath across the arteriotomy. This is a deviation from the firm's standard catheterization technique. | Device Design | Cardinal Health sent an Urgent Medical Device Field Safety Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed complete the attached Field Action Return Response Form as soon as possible indicating the number of devices that have been used as their location and the number that will be returned. Customers should discuss the reimbursement for their purchased and return product with their sales representative, who will coordinate the process. Customers with questions were instructed to call 408-610-6570.For questions regarding this recall call 408-610-6500. |  | 6 | 6 in USA | US distribution to FL and NJ | Open, Classified |  |  | 6 | Not_Computer
Z-0117-2017 | 75393 | Yes | NIK |  |  | Cardiovascular | PMA | Quadra Assura, Model No. CD3265-40, CD3265-40Q, CD3365-40C, CD3365-40Q. |  |  | October 10, 2016 | 2016 | St Jude Medical Inc. | St Jude Medical is recalling implantable cardiac defibrillator (ICD) and cardiac resynchronizationtherapy defibrillator (CRT-D) devices because they are subject to premature battery depletion due to lithium cluster with the battery causing a short circuit. | Device Design | St. Jude Medical, Inc., sent a customer letter to all affected customers to inform them that St. Jude Medical recently announced a global medical device advisory for a subset of their company's Fortify, Fortify Assura, Quadra Assura, Unify, Unify Assura and Unify Quadra implantable cardiac defibrillator (ICD) and cardiac resynchronization therapy defibrillator (CRT-D) devices. The letter informs customers that the cause for Premature Battery Depletion is attributed to the formation of lithium clusters within the battery causing a short circuit. There is a possibility that affected devices may lose battery power within days. Thus far, there have been two (2) deaths and ten (10) serious events (fainting) that may have been associated with premature battery depletion. Customers with questions were instructed to contact www.sjm.com/batteryadvisory, (866) 915-5065. |  | 193954 | 193,954 units | Worldwide Distribution - US (nationwide) and OUS | Open, Classified |  |  | 603524 | Not_Computer
Z-0299-2017 | 75224 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Keystone Dental PrimaConnexÂż Straight Implant TC, RD Internal Connection 4.0 x 11.5mm. |  |  | September 15, 2016 | 2016 | Keystone Dental Inc | Small diameter (SD) cover screw was packaged with the regular diameter implant and do not match the implant diameter (RD) | Packaging process control | Keystone Dental initially called all customers on September 15, 2016. Forms were completed to document these calls. In addition a letter has been sent via Federal Express/or hand delivered on September 19, 2016, to all distributors and customers. Customers with questions were instructed to contact Customer Relations.For questions regarding this recall call 781-328-3390. |  | 55 | 55 units | Worldwide distribution - US (Nationwide Distribution to CA, CO, IL, MA, MN, ND, NY, NV, PA, and VA.) and Italy | Terminated | January 03, 2017 | 110 | 55 | Not_Computer
Z-0300-2017 | 75167 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System (Product Code ENGEN) configured with Thermo Scientific Centrifuge Module and TCAutomation Software Version 3.6.1 and Below; IVD. --- Thermo Scientific Centrifuge Module has Product Code |  |  | September 07, 2016 | 2016 | Ortho-Clinical Diagnostics | Ortho confirmed that it is possible for the mis-association of sample identification and assay results due to an error with Thermo Fishers centrifuge module. It is possible for mis-associated assay results to be reported from the laboratory prior to the identification of the Cross Check error on the enGen Laboratory Automation System. Assay results associated with the wrong patient and reported out of the laboratory could lead to inappropriate intervention with the potential for serious injury to the patient. | Software design | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" Letter (Ref. CL2016-183, dated 9/7/2016) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all enGen customers with centrifuge modules to inform them of the issue and advise them to check samples associated with cross check errors as instructed in the operator manual, and perform the preventative maintenance and cleaning procedures for that centrifuge module as provided in the maintenance manual and user guide. The customers were also instructed to Complete and return the Confirmation of Receipt form by 16 September 2016 via Fax:1.888.557.3759 or 1.585.453.4110 or email:RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM. For questions regarding this notification, please call the Ortho Care Solutions Center any time at 1-800-421-3311. Foreign affiliates were informed by e-mail on 9/7/2016 of the issue and instructed to notify their consignees of the issue and appropriate actions to take. |  | 60 | 60 (US: 24; Foreign: 36) | Worldwide Distribution-(Nationwide) and foreign distribution to the following countries: Belgium, Bermuda, Canada, Chile, Denmark, France, Italy, Mexico, Norway, Portugal, Sweden, Taiwan and United Kingdom. | Open, Classified |  |  | 60 | Not_Computer
Z-0301-2017 | 75384 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FemTouch delivery system includes the FemTouch Kit (AC-1007415) and FemTouch" Tip (SPSA-20002590) |  |  | September 25, 2016 | 2016 | Lumenis Limited | The cleaning instructions provided in FemTouch Operator's Manual might be ineffective for a fully assembled device and may lead to a potential for cross-contamination. | Software Manufacturing/Software Deployment | On 9/25/16 consignees were notified via an Urgent - Field Safety Corrective Action letter with a FemTouch disassembly tool kit as well as FemTouch disassembly instructions. The instructions describe the simple disassembly process to be performed as part of the Pre-cleaning phase and the process for reassembling the device prior to re-use. It is recommended that customers clean and sterilize all of their handpieces according to the pre-cleaning disassembly process before each use. Consignees were asked to complete and return a verification form. Questions about the FSCA, should be directed to the Lumenis field correction administrator, Rafi Barel at Rafi.Barel@lumenis.com. |  | 69 | 69 units | Worldwide Distribution -- USA, Italy, Russian Fed., Australia, United Arab Emirates, Spain, United Kingdom, India, Kuwait, and Belarus. | Terminated | April 11, 2017 | 198 | 69 | Not_Computer
Z-0302-2017 | 75265 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ex-TEM Assay for ROTEM delta Thromboelastometry System. Art. no.: 503-05 US, ex-tem. |  |  | September 13, 2016 | 2016 | TEM Systems Inc | Customers complained of low/empty reagent vials. | Nonconforming Material/Component | Customers were notified of the recall by an Urgent Important Customer Information letter on 9/13/2016. The letter described the device and the problem. Customers were told to stop using the product and to call for replacement product. The letter included a Confirmation of Compliance Letter which was to be completed and returned. |  | 647 | 637 boxes x 10 vials | Nationwide Distribution -- AR, CA, DC, GA, HI, ID, IL, IN, LA, ME, MD, MA, MI, MN, MO, NE, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, and WV. | Open, Classified |  |  | 647 | Not_Computer
Z-0305-2017 | 63994 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids |  |  | February 04, 2013 | 2013 | Hospira Inc. | Symbiq Infusers have the potential to experience a white screen during titration of a Keep Vein Open (KVO) delivery. This can allow the clinician to select the "NEXT" button prior to completion of the programming. If this error occurs, a delay in therapy could occur and potentially result in significant injury. | Software design | Hospira sent an Urgent Device Field Correction notice dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the steps provided in the letter while Hospira is developing a software change. Customers were asked to complete the attached reply form and return it via fax to the number on the form. Customers with questions were instructed to contact Stericycle at 1-888-201-9332.For questions regarding this recall call 224-212-4892. |  | 35596 | 35,596 devices in total | Worldwide Distribution - USA (nationwide) and Canada | Open, Classified |  |  | 35596 | Computer
Z-0304-2017 | 75408 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | 4205 Journey Rolling Walker |  |  | October 01, 2016 | 2016 | Nova Ortho-Med Inc | Nova Medical is recalling the Journey Rolling Walker because the screws used to hold the black side support pieces to the frame can become loose over time. This can result in the locking nut coming off and the screw falling out. | Component design/selection | A customer notification letter dated 10/1/16 was sent to customers to inform them that the screws used to hold the black side support pieces to the frame can become loose over time. This can result in the locking nut coming off and the screw falling out. This allows the frame support piece to detach from the walker and causes instability resulting in possible injury. The letter informs the customers of the root cause and the solution. Customers are instructed to record the serial #s of the product that have been corrected and email them to NOVA Customer Service (patricias@novajoy.com). Customers can also 1-800-557-6682 x131. |  | 1380 | 1,380 units | US Nationwide Distribution in the states of: CA, IN, HI, CT, UT, WA, FL, VA, WI, OH, MN, KS, MD, IL, MO, WY, MT, ME, PA, TN, MI, OR, NY, NJ, ID, NC, IA, OK, AR, VT, SC, TX | Open, Classified |  |  | 1380 | Not_Computer
Z-0303-2017 | 75404 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | ReTurn 7500/7500 |  |  | April 15, 2014 | 2014 | Handicare USA, Inc. | During the period of April 2009  May 2011, Handicare AB used a wing handle in which the screw and handle, after repeated tightening, may slide apart. This causes a small amount of play in the attachment of the ladder, which in turn results in high load at the attachment holes of the ladder. Eventually, this may cause mechanical fatigue in the material of the ladder. | Nonconforming Material/Component | Handicare sent a Field Safety notice dated April 21, 2014, to all affected customers. The notification instructed customers to examine the affected ReTurn units for signs of wear and to complete the response form that was mailed with the customer notification letter indicating whether additional parts were required to complete the field correction. For further questions, please call (717) 733-1444. |  | 76 | 29 units , plus 47 wing handles were shipped as spare parts | US Distribution to the states of : NY, NJ, NV, NE, IL, CT, PA and ME | Completed |  |  | 76 | Not_Computer
Z-0094-2017 | 75186 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi MODULAR Analytics Combination Systems |  |  | July 15, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche has received four reports of injury caused by the sharp edge of the center guide rail of ananalyzers sample rack tray. The cleaning procedure for the sample rack tray has been revised to clarify the cleaning process. | Device Design | On 7/15/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, an updated cleaning procedure, and instructions for responding to the formal recall notification. Please contact the Roche Support Network Customer Support Center at 1-800-428-2336 if you have questions about the information contained in this UMDC.On 10/5/2016, URGENT MEDICAL DEVICE CORRECTION UPDATE notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The notification updated the 7/15/16 notification to include additional affected systems and to provide updated cleaning instructions for all systems via the enclosed Operators Manual addendum. |  | 9014 | 9014 in total | US Nationwide Distribution | Open, Classified |  |  | 9014 | Not_Computer
Z-0311-2017 | 75375 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | PhoenixSpec Calibrator &amp; Kit is used to calibrate the PhoenixSpec &#38; Phoenix AP Instruments. The instruments are then used to measure the turbidity of microbial suspensions equivalent to McFarland standards 0.10 to 4.50. |  |  | June 01, 2016 | 2016 | Becton Dickinson &amp; Co. | BD has determined that the BBL PhoenixSpec Calibrator kits referenced in this recall from the following lots (5364776, 5345796 ) are labeled with an incorrect expiration date (year). The expiration date printed on the individual calibration tubes within the kit are correct but difficult to read on a portion of the tubes. | Process control | All US customers were contacted via letters sent via UPS. Ex-US BD sites will contact customers in their local regions. The firm will use an acknowledgement form included with the follow up letter that is mailed or faxed back to BD will be reconciled with the customer list. BD will contact 100% of our customers that have not responded to the communication to verify acknowledgement |  | 20 | 20 | BD has contacted all 31 US customers and 22 Ex-US sites which received product from the associated catalog numbers.US: AL, CA, CO, KY, MD, MI, MS, NE, NJ, OH, PA, SC, SD, TX, VA, WA and WI.Ex-US: Argentina, Europe, Canada, China, Ecuador, India, Japan, New Zealand, Peru, Singapore, Thailand and Vietnam4 US government sites were contacted | Completed |  |  | 114 | Not_Computer
Z-0313-2017 | 74772 | No | MJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Microbiology | 510(K) Exempt | VITROS Immunodiagnostic Products HBsAg Controls, Product Code 6800598 |  |  | June 22, 2016 | 2016 | Ortho-Clinical Diagnostics | The signal/cutoff (s/c) results may increase throughout the open-reconstituted timeframe of the allowed 5 days. Increased results may occur on both the negative and positive Controls. The positive Control, although showing elevated results, was still within acceptable performance. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, distributed "Important Product Correction Notification" letters (Ref. CL2016-129) dated 6/22/2016 and Confirmation of Receipts to their customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only). Foreign affiliates were informed by email (on 22 June 2016) of the issue and instructed to notify their consignees of the issue and required actions. All customers who received the affected product were instructed to immediately discontinue using and discard the affected lots. Forward a copy of the notification if you have provided the affected lot(s) outside of your facility. Post a copy of the notification by each VITROS System that processes VITROS HBsAg Controls and Complete and return the Confirmation of Receipt-Response Required form via fax to: 1.888.557.3759 or 1.585.453.4110 or scan to PDF and email to: RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM no later 7/1/2016. (Note: If you DO NOT have an alternate lot, it is acceptable to continue using the affected lot(s) following these instructions: 1) Control results must be within acceptable limits on the day of reconstitution (Day 1). 2) Store open-reconstituted vials frozen for up to 4 weeks. The firm recommends that customers aliquot the open-reconstituted material prior to freezing so that multiple aliquots of the same vial can be used. 3) Upon receipt of your replacement order, discard all remaining inventory. Customers with questions please call the Ortho Technical Solutions Center at 1-800-421-3311. |  | 1621 | 1,621 units | Worldwide Distribution-USA (nationwide) and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 1621 | Not_Computer
Z-0316-2017 | 75142 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Safil QUI + UNDY 5/0 (1) 45CM DSMP13 Suture |  |  | August 01, 2016 | 2016 | Aesculap Implant Systems LLC | There is a potential for the needle to detach from the suture thread during surgery. If the needle becomes detached the defective suture will need to be replaced causing the wound to be sutured a second time. | Process control | An Urgent Medical Device Recall Notification Letter was mailed to customers on August 11, 2016 asking them to remove and return the affected product immediately. |  | 180 | 180 units | Distributed in Arizona | Open, Classified |  |  | 180 | Not_Computer
Z-0317-2017 | 75371 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Gender Solutions" Patellofemoral Joint Prosthesis |  |  | September 28, 2016 | 2016 | Zimmer Biomet, Inc. | Complaints have been reported that the handpiece is inoperable. The investigation found that the complaints were related to the lack of preventative maintenance contrary to what is recommended in the instructions for use included with the handpiece. An inoperable device may lead to a delay in surgery | Labeling design | On 9/28/2016, URGENT MEDICAL DEVICE FIELD REMOVAL notifications were sent to the affected distributors and Risk Managers / Surgeons via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Risk Managers and SurgeonsRisk Manager Responsibilities: 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative with quarantining all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility for the initial preventative maintenance and arrange for a loaner device. It is estimated that your handpieces will undergo preventative maintenance sometime between November 2016 and March 2017. 4. Complete Attachment 3  Certificate of Acknowledgement. a. Return a digital copy via email to CPWARFieldAction@zimmerbiomet.com or via fax to 574-372-1683 within three (3) days. b. Retain a copy of the Acknowledgement Form with your field removal records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to CPWARFieldAction@zimmerbiomet.com. Surgeon Responsibilities: 1. Review this notification for awareness of the contents. 2. There are no specific patient monitoring instructions related to this field removal that are recommended beyond your existing follow up schedule. 3. Complete Attachment 3  Certificate of Acknowledgement. a. Return a digital copy via email to |  | 703 | 703 | Nationwide.CANADAHONG KONGKOREASINGAPORECHINAVENEZUELAAUSTRALIAJAPANNETHERLANDSGERMANYINDIATAIWAN | Open, Classified |  |  | 703 | Not_Computer
Z-0315-2017 | 75473 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da VinciÂż Xi" Surgical System, model number IS4000, A70_P5x with P5 Software; |  |  | October 05, 2016 | 2016 | Intuitive Surgical, Inc. | Intuitive Surgical has identified a software anomaly in the da Vinci Xi P5 software that can result inunexpected master movement and potential instrument tip movement under certain circumstances | Software design | Urgent Medical Device Correction letters were sent on October 12, 2016 by traceable method to all affected customers. Letters notify customers of the issue and provided instructions to prevent it from happening until ISI software can install new software upgrade in 1st Q 2017. |  | 677 | 677 devices | Australia, Austria, Belgium, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Monaco, Netherlands, Norway, Portugal, Puerto Rico, Qatar, Romania, SouthKorea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United States | Terminated | May 01, 2017 | 208 | 677 | Computer
Z-0318-2017 | 75332 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Gastroscope is intended to be used with a Pentax video processor (including light source), documentation equipment, monitor, EndoTherapy Device such as a Biopsy Forceps, and other ancillary equipment for endoscopy and endoscopic surgery within the upper digestive tract including the esophagus, stomach and duodenum. |  |  | August 15, 2016 | 2016 | Pentax of America Inc | PENTAX Medical is initiating this field action to provide customers that have purchased the affected endoscopes with the most recent Operation and Reprocessing IFUs. | Labeling Change Control | Pentax sent as "Urgent Field Correction" and Field Correction Response Form dated August 15, 2016 to their affected customers. |  | 3436 | 3,436 (US) and 14,488 (OUS) | US Nationwide and Japan | Open, Classified |  |  | 3436 | Not_Computer
Z-0319-2017 | 75090 | No | LKQ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | IMMULITE/ IMMULITE 1000 Systems CMM CMV IgM Assay, REF/Catalog Number LKCM1(D) - 100 tests (US), Siemens Material Number (SMN) 10370301 (US), Unique Device Identification Number (UDI) 00630414985718 (US); and REF/Catalog Number LKCM1 - 100 tests (OUS); SMN 10381296 (OUS), UDI 00630414963747 (OUS); IVD --- Made in UK Siemens Healthcare Diagnostics Products Ltd., Glyn Rhonwy, Llanberis, Gwynedd, LL55 4EL UK |  |  | August 30, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Multiple customers reported seeing an increase in imprecision with some patients resulting in high % CVs with the IMMULITE 2000/2000XPi CMV (IgM) assay starting with reagent lots 259 and above. | Under Investigation by firm | The firm, Siemens, issued an "Urgent Medical Device Recall" Letter (IMC16-22.A.US, dated 8/30/2016) in the US and an "Urgent Field Safety Notice" (IMC16-22.A.OUS, dated August 2016) Outside the US (OUS) starting on 8/30/2016 The notices instructed customers to discontinue use of and discard the affected kit lots listed in the customer notice; review your inventory of the products to determine your laboratory's replacement needs; review the UMDR and or the UFSN with their Medical Director, In the US, the method of distribution is FedEx and Outside the US the distribution is determined at the country level; complete and return the Field Correction Effectiveness Check Form via Fax to: Customer Care Center at (312) 275-7795 within 30 days, and retain this letter with your laboratory records, and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 377 | US: 5 kits; Foreign: 372 kits | Worldwide distribution- US (Nationwide) and countries of: United Arab Emirates, Afghanistan, Angola, Argentina, Australia, Bangladesh, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Belarus, Switzerland, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Dominican Republic, Algeria, Ecuador, Egypt, Spain, France, Great Britain, Greece, Guatemala, Honduras, Hungary, India, Iraq, Italy, Jordan, Canary Islands, Kyrgyzstan, Rep. Korea (S), Kazakhstan, Lebanon, Lithuania, Luxembourg, Latvia, Macedonia, Mexico, Malaysia, Nicaragua, Netherlands, Norway, Nepal, Peru, Pakistan, Poland, Portugal, Paraguay, Qatar, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, El Salvador, Syria, Turkmenistan, Tunisia, Turkey, Uruguay, Uzbekistan, Vietnam, Serbia,and South Africa. | Open, Classified |  |  | 10336 | Not_Computer
Z-0314-2017 | 74786 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | SimpliRED D-Dimer, Product Code DSRK4 |  |  | June 01, 2016 | 2016 | BBI SOLUTIONS OEM LTD | A component of the kit is showing a sensitivity performance detected in stability testing at 15 months of an 18-month shelf life. Qualitative results may be affected. False negative results are possible which may incorrectly influence the clinical diagnostic decision. | Nonconforming Material/Component | BBI Solutions sent an Urgent Medical Device Recall letter dated June 30, 2016, to all affected customers. Customers were advised to immediately discontinue use of the product, consider the requirement for review of reported test results, complete and return the customer acknowledgement form to BBI Solutions and to complete and return the customer reconciliation form to the distributor. Customers who have forwarded product to other sites, should inform those customers of the product recall and provide them with a copy of the notice. Customers with questions were instructed to call Tel: +44(0)1382 564093. |  | 420 | 420 units (183 units quarantined) | US distribution to CT and IN | Open, Classified |  |  | 420 | Not_Computer
Z-0114-2017 | 75211 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | VIA 27 Microcatheter, Part No. FG27-154-01, Model No. VIA-27-154-01 |  |  | September 23, 2016 | 2016 | Sequent Medical Inc | Sequent Medical is recalling the VIA 27 Microcatheter because it may not adequately re-sheath the interventional device into the catheter as designed. | Under Investigation by firm | Sequent Medical sent an Urgent Medical Device Recall letter dated September 23, 2016, to all affected customers. The letter informed customers that Sequent Medical is recalling the VIA 27 Microcatheter because it may not adequately re-sheath the interventional device into the catheter as designed. The letter informs the customers of the risk to health and actions to be taken. Customers were instructed to send all confirmation forms and questions to www.customerserviceus@sequentmedical.com. Customers with questions should call (949) 830-9600x130.For questions regarding this recall call 949-923-9400. |  | 618 | 618 units | Worldwide Distribution - US including NY, MN, FL, MO, MD, TN, NC, OH, CA, KS, TN, WV and Internationally to Czech Republic, Germany, UK, France, Netherlands, Italy, Spain, Finland, Switzerland, Slovenia, Hungary, Northern Ireland, Turkey, and Denmark, | Terminated | May 01, 2017 | 220 | 618 | Not_Computer
Z-0321-2017 | 75062 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TD60 display Telemetry Transmitter, SHENZHEN Mindray Bio-Medical Electronics Co., LTD, Mindray Building, Keji 12th Road South, Hi-Tech Industrial Partk, Nanshan, Shenzhen, 518057 P.R. China |  |  | September 08, 2016 | 2016 | Mindray DS USA, Inc. dba Mindray North America | Mindray has become aware that the TD60 Display Telemetry Transmitter used with the BeneVision Central Station may crack on the bezel surrounding the display and also the transmitter may not power on or will power on but display an error message. | Device Design | Mindray sent a "Medical Device Action" letter dated September 8, 2016 via UPS to their affected customers. |  | 1278 | 1,278 units | US and Canada | Open, Classified |  |  | 1278 | Not_Computer
Z-0322-2017 | 75124 | No | GAH | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Visualase Visualization Stylets PN 020-2301. The individual components within the Visualase Body Accessory Kit are packaged in Tyvek pouches. The packaged devices are then placed into a cardboard box which is labeled with the Body Accessory Kit label. Labeled sterile. |  |  | September 01, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. announces a voluntary field action for the Visualase Visualization Stylets and Visualase Body Accessory Kit because when the Visualization Stylet is placed in the magnetic field of a magnetic resonance imaging system it pulls and/or bends toward the field. | Mixed-up of materials/components | Medtronic sent an Urgent Medical Device Recall letter dated September 2, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their stock for the affected product and if any are found contact Medtronic to arrange for a no charge replacement. Customers were asked to complete the Customer Confirmation form attached to the letter and return it to Medtronic at RS.NavFCA@medtronic.com or by Fax to 651-367-7075. Customers with questions were instructed to contact Medtronic Technical Services at 1-800-595-9709.For questions regarding this recall call 720-890-3200. |  | 100 | 100 total units | Worldwide Distribution - US including CA, GA, FL, and TX. Internationally to Australia. | Terminated | March 27, 2017 | 207 | 100 | Not_Computer
Z-0325-2017 | 75214 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr PS Knee System Box Cut Guide; Used to make the housing resection for the post of the posterior stabilized implant. |  |  | September 19, 2016 | 2016 | Encore Medical, Lp | It has been discovered that during Empowr PS Total Knee replacement, while chiseling for bone removal in the box area, femoral chipping has occurred in surgery. The current design of the captured box cut guides can create an instance where the surgeon is not fully completing the box cut on the side. | Under Investigation by firm | The recalling firm sent a recall notification letter to affected consignees on 9/19/16. The field safety notice provided a newly revised Surgical Technique to educate users on proper use of the device. |  | 436 | 436 units | US | Open, Classified |  |  | 436 | Not_Computer
Z-0324-2017 | 75221 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reprocessed Pressure Tourniquet Cuff (PTC) |  |  | September 19, 2016 | 2016 | Stryker Sustainability Solutions | Issue with the process for detection of leaking Pressure Tourniquet Cuffs (PTC). Leaking PTC devices may result in excessive blood loss requiring surgical or medical intervention or systemic leakage of local anesthetic during Bier Block requiring medical intervention. | Device Design | Stryker Sustainability Solutions sent an Customer Notification letter September 19, 2016. Customers were instructed to discontinue use of the affected products. The letter stated to follow the link below to complete the Recall Effectiveness Check Form and indicate if any devices containing expiration dates from May 2016-May 2018 remain in inventory. This form must be completed even if no affected product is found.Use the following reference number as you complete the form: [Ship to Account #]If the form indicates that any affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Credit for all affected devices returned will be issued to the facility..Adverse reactions or quality problems experienced with the use of this product may be reported to:o Stryker Sustainability Solutions Complaint Hotline: +1(888) 888-3433 X5555o http://www.stryker.com/productexperience/o The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax. |  | 1695760 | 1,695,760 | U.S. distribution nationwide.No foreign distribution | Open, Classified |  |  | 1695760 | Not_Computer
Z-0326-2017 | 75197 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Product Number 3102071 Ceiling Bracket, Arch 71 White with use on ceiling |  |  | August 09, 2016 | 2016 | Hill-Rom, Inc. | Ceiling Bracket 71 wasshipped with a larger center hole. A potentiallyhazardous situation exists that if the Ceiling Bracket 71with a larger center hole is used for a concrete ceilingUltra System installation with the M8 screw and the safeworking load is exceeded, the system could fall. | Mixed-up of materials/components | On 8/9/2016, URGENT: MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any questions regarding this Field Safety Correction, please contact Hill-Rom Technical Support at 1-800-445-3720, your distributor or your Hill-Rom representative. |  | 188 | 188 | WANCOHPAMIAustraliaSweden | Terminated | March 24, 2017 | 227 | 188 | Not_Computer
Z-0086-2017 | 74018 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | 15 g R.K. Epidural Needle (TW), Catalog #/REF #100-1415 |  |  | April 22, 2016 | 2016 | Epimed International, Inc. | There is a possibility that the needles may have the potential to skive a catheters outer coating. If this were to occur, there exists the risk that a portion of the catheter could remain in the patient, which may require otherwise unnecessary medical invention. | Employee error | On April 22, 2016, Epimed sent an Urgent Medical Device Recall notices and Recall Return Response Forms to their customers via courier service. Customers were advised to examine all inventory and quarantine product subject to recall immediately. Once the product has been quarantined, please complete and return the enclosed response form as soon as possible by fax to: 518-725-0207, Attn: Quality Assurance. In addition, if you find any product that corresponds to the lot numbers given, please contact the Customer Service Department so that Epimed can arrange for replacement/credit (518-725-0209 or 518-514-1891).Further questions, please call (518) 725-0209. |  | 576 | 576 units | US Distribution to the states of : AL, AK, FL, IN, TN &amp; TX | Open, Classified |  |  | 576 | Not_Computer
Z-0327-2017 | 74879 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Affected Pain Control Tray Component (5% Lidocaine Hydrochloride and 7.5% DExtrose Injection, USP) |  |  | July 11, 2016 | 2016 | B. Braun Medical, Inc. | B. Braun Medical Inc. is voluntarily recalling specific lots of their Custom Spinal Anesthesia Tray which contains a drug component, 5% Lidocaine Hydrochloride and 7.5 % Dextrose Injection, USP that Hospira Inc. has recalled because the drug product does not meet the specification for color throughout shelf life. | Under Investigation by firm | The firm, B.Braun, mailed to customers a "VOLUNTARY DRUG RECALL NOTIFICATION" letter dated July 12, 2016. The letter described the product, problem and actions to be taken and to inform the customers of a drug recall issued by Hospira, Inc. The Customers were instructed to determine their current inventory of the affected lots Do not destroy any affected product), complete and return "Product Removal Acknowledgement" form via fax to: B.Braun Medical Inc., Quality Assurance department at (610) 849-1197 or email to PA_QualityAssurance.BBMUS_Service@bbraun.com within two weeks of receipt, even if you have no inventory. A BBMI Customer Service Representative will contact you to provide instructions for handling the affected product and arrange for return to BBMI.Should you have any questions or concerns regarding the attached information, please contact our Customer Support Department at (800) 227-2862. |  | 1704 | 1704 units | US Distribution to: AL, IL, OK, MN, MA and WI. | Terminated | March 15, 2017 | 247 | 1704 | Not_Computer
Z-0329-2017 | 75066 | No | MRV | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Glucose Tolerance Beverage- Orange 50 gram Concentration, Liquid 4/Pack of 6 bottles, 24/case |  |  | August 24, 2016 | 2016 | Azer Scientific, Inc. | Out of Specification Microbiological test results | Unknown/Undetermined by firm | Azer Scientific mailed an Urgent Medical Device Recall letter dated August 30, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to discontinue use of product and discard to avoid further distribution or use and to notify their customers/end users of the recall. Customers were asked to acknowledge receipt of the letter by completing the attached form and faxing it to (610) 901-3046. Customers with questions were instructed to call (610) 524-5810. |  | 426 | 426 cases | Nationwide Distribution | Open, Classified |  |  | 1567 | Not_Computer
Z-0332-2017 | 75477 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Xhibit Central Station, Model 96102. |  |  | October 18, 2016 | 2016 | Spacelabs Healthcare Inc | The firm received reports of telemetry SpO2 numerics dropping off the Xhibit Central display. Desaturation, high, and low limit alarms work normally. | Software design | The firm sent the Urgent - Medical Device Correction letter, dated 18 October 2016, via priority service, return receipt requested, to all U.S. customers on 10/18/16. The firm will e-mail a customer letter (translated as necessary) to all international subsidiaries and distributors of record on October 25, 2016.Xhibit Central Stations used without Xhibit Telemetry Receivers, are not affected.Customers were asked to circulate the letter to all person affected and place a copy with the appropriate monitor user manual until corrections have been implemented. Spacelabs Healthcare will contact affected consignees to schedule a convenient time for Spacelabs to update the affected Xhibit Central Station(s) with a new software version 1.1.6 at no cost.Customers with questions can call Technical Support at 1-800-522-7025 and select 2. |  | 348 | 348 units (292 units in the US and 56 units outside the US) | Worldwide Distribution -- USA, to AL, CA, CT, FL, GA, IA, ID, IL, IN, KY, MI, MN, MS, NC, NE, NJ, NY, OH, OK, PA, RI, SC, SD, TX, VA, WA, WY, and Puerto Rico; and, to the countries of AUSTRALIA, CANADA, CHILE, CZECH REPUBLIC, FRANCE, MEXICO, NETHERLANDS, PANAMA, SAUDI ARABIA, SWITZERLAND, TAIWAN, and UNITED KINGDOM. | Terminated | April 03, 2017 | 167 | 348 | Computer
Z-0331-2017 | 75372 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | HANDS-OFF Infusion Port Thermodilution Catheter consists of the Arrow IPTD thermodilution catheter enclosed in a contamination shield (Arrow Cath-Gard) with integral flushing/balloon test chamber, enabling the practitioner to prepare, test, and insert the catheter without exposing it to external contamination. |  |  | September 02, 2016 | 2016 | Arrow International Inc | Labeling inconsistency | Labeling mix-ups | Arrow sent an Urgent Medical Device Recall Notification letter dated September 20, 2016, to all affected customers via FedEx 2-day air. The letter identified the problem and provided instructions to immediately discontinue use and quarantine any products with the associated lot numbers indicated in the letter. If product was found, customers were asked to complete the Recall Acknowledgement Form and a Customer Service Rep will issue a Return Goods Authorization (RGA) Number for the product's return. Disposition of recalled product will be scrapped. For further questions, please call (610) 378-0131 |  | 1361 | 330 units in US and 1,031 units OUS | Nationwide distribution | Open, Classified |  |  | 1361 | Not_Computer
Z-0333-2017 | 75447 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Guider Softip XF; |  |  | October 05, 2016 | 2016 | Stryker Neurovascular | Firm became aware of potentially defective Guider Guide Catheters due to a manufacturing non-conformance that could lead to hub leaks in the resulting product. | Employee error | Stryker Neurovascular sent an Urgent Medical Device Voluntary Recall letter dated October 5, 2016, to all affected customers by Federal Express. Letters advised of the reason for recall and models and devices affected and requested that all affected devices in inventory be returned to Stryker. Customers were advised to circulate the Field Safety Notice internally to all interested/affected parties and to complete the attached customer response form and return the form to their nominated Stryker Representative or to NVFieldActions@stryker.com.For questions regarding this recall call 510-413-2106. |  | 1478 | 1478 devices, all models and lot numbers. | Worldwide Distribution - US including AK, IL, OH and Internationally to:ARGENTINAAUSTRALIAAUSTRIABELGIUMBRAZILCANADACHILECZECH REPUBLICECUADOREGYPTFRANCEGERMANYGREECEHUNGARYINDIAIRANISRAELITALIAJAPANMEXICONETHERLANDSNORWAYPERUPHILIPPINESPOLANDRUSSIASERBIASINGAPORESLOVAKIASOUTH AFRICASOUTH KOREASPAINSWEDENSWITZERLANDTHAILANDTUNISIATURKEYUKURUGUAYVIETNAM. | Terminated | March 23, 2017 | 169 | 1478 | Not_Computer
Z-0354-2017 | 63993 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids; Product List Number 16027 |  |  | February 05, 2013 | 2013 | Hospira Inc. | Hospira has received reports of Malfunction S205 Backup Battery Failure alarms in Symbiq v3.13 pumps. This issue to an incorrect installation of components on some of the Power Supply Controller (PSC) Boards distributed prior to September 2012. On pumps where incorrect installation of components has occurred, the backup battery power discharges and subsequently, the battery cannot be charged. Because of the battery depletion, or lack of charge, a visual and audible warning alarm, S205, is displayed. An S205 alarm is intended to notify the customer that the coin call battery, powering the backup buzzer, may need to be charged or replaced. | Process control | Hospira sent an Urgent Field Device Correction letter dated Feburay 5, 2013, to all affected customers.Hospira has implemented a correction to the manufacturing process to address this issue. Symbiq v3.13 pumps that have been shipped from Hospira since September 2012 and pumps that were upgraded to v3.13 in the Field after September 2012 are not impacted by this issue. If you are unsure of when you were shipped Symbiq v3.13 pumps or when your devices were upgraded, contact Hospira.Please complete the attached reply form and return it via fax number or e-mail address on the form, even if you do not currently have the impacted infuser. lease contact Stericycle at 1-866-201-9072 (M-F, Bam - 5pm EST) to obtain additional copies of the reply form.If you have further distributed these devices, please notify your accounts who may have received these infusers from you and ask them to contact Stericycle at 1-866-201-9072 (M-F, 8am- 5pm EST) to receive a reply form. |  | 15285 | 15,285 | Nationwide Distribution and Canada | Open, Classified |  |  | 35596 | Not_Computer
Z-0352-2017 | 75128 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Rigid Suction Canister 3000 CC and 2000 CC. |  |  | August 10, 2016 | 2016 | Medline Industries Inc | Risk of implosion when the Rigid Suction Canister is subjected to excessive negative pressure during use. | Process control | Medline Industries, Inc. sent an Urgent Recall Immediate Action Required letter dated August, 23, 2016, and destruction form via first class mail to affected customers on August 17, 2016 and August 23, 2016. The firm recommended the following actions: REQUIRED ACTION:1. Immediately check your stock for item number OR220 and OR230 with lot numbers beginning with 9961601xxxx through 9961604xxxx (listed on the original case packaging). Destroy all affected products.If the rigid canisters are NOT in the original case packaging, we are instructing you to destroy them and we will send replacements.2.Please return the completed enclosed destruction form listing the quantity of product(s) destroyed. Even if you do not have any of this product please complete and return the form, as indicated on the destruction form.When we receive your completed destruction form your account will receive replacements.3.If you are a distributor, promptly notify any of your customers that may have received affected product.If you have any questions, please contact 866-359-1704. |  | 7680 | 7,680 units | US Distribution to the following states: CA , CO , FL , GA, GA , IL , MA, MA , MD , MI, MO , NC , NE , NJ , NY, NY , OH, OK , PA, TX and WA. | Open, Classified |  |  | 7680 | Not_Computer
Z-0349-2017 | 73286 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | Model 055: Thermophore Classic, Professional Quality, Deep-Heat Therapy, Model 055 Large 14' x 27", |  |  | September 15, 2016 | 2016 | Battle Creek Equipment Co. | Products manufactured with the SPT cord have a known failure mode of sparking,causing smoke, fire and burns. | Device Design | All consumers for which Battle Creek has contact information and that purchased a recalled pad will be contacted via first class mail. Consumers will be asked to contact us by phone after verifying that the model number and lot code is part of the recall. Battle Creek plans to add information about the recall on its web site including a pdf of the consumer letter. All customer consignees (non-consumer) will be informed of the recall. Email will be the most effective method , and mail and telephone calls will be used as needed. |  | 99060 | 99,060 | Distributed within the United States, Canada, Mexico, Puerto Rico and Columbia | Open, Classified |  |  | 99060 | Not_Computer
Z-0360-2017 | 74242 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SureT G29 6mm, 18", Paradigm; SureT G29 6mm 23"; SureT G29 6mm, 23", Paradigm; SureT G29 6mm 32"' SureT G29 6mm, 32", Paradigm; SureT G29 8mm, 23"; SureT G29 8mm, 23", Paradigm; SureT G29 8mm, 32"; SureT G29 8mm, 32", Paradigm; SureT G29 10mm, 32", Paradigm |  |  | November 07, 2014 | 2014 | Unomedical As | Tubing detachment | Device Design | ConvaTec sent an Urgent Medical Device Correction letter to all affected consignees on November 7, 2014. The letter notified the consignees of what action they need to take if they experience a detached tubing. You can continue to use your infusion sets by following the advice below.1. When changing your infusion set, closely follow the instructions for use included in the product box. Check the tubing at the connect/disconnect location identified in the drawings above to make sure it is not loose.2. As always, it is essential to monitor your blood sugar levels frequently using your blood glucose meter. Proactively check your tubing connections occasionally throughout the day to ensure tubing is secure. It is especially important to check your blood sugar and tubing connections at bedtime to confirm insulin delivery is occurring.3. If you experience a high blood sugar, check your tubing connections and infusion site closely to ensure your tubing is secure.If you discover the tubing is detached:1. Do not attempt to reattach the tubing. Replace the infusion set immediately.2. Treat any high blood sugar based on guidelines provided by your healthcare professional.3. Call your distributor's Helpline to report the issue. They will give you instructions on how to return the affected infusion set.The letter included a reply card to be returned. |  | 4748000 | 4,748,000 units | Worldwide Distribution - US including CA, IN, PA, and Internationally to Canada, Korea, and Japan. | Completed |  |  | 17165789 | Not_Computer
Z-0348-2017 | 75271 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Heater Cooler Unit |  |  | September 27, 2016 | 2016 | Maquet Cardiovascular Us Sales, Llc | There is the potential for contaminated water to enter parts of the Heater-Cooler Unit circuits. There is a potential risk of bacterial contamination including Mycobacteria species. | Device Design | Maquet issued an Urgent Medical Device Recall letter/Return Response form dated September 27, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customer were instructed to complete the return response form via fax to 1-973-629-1518. For questions contact Maquet Customer Service at 1-888-627-8383 option 2 followed by option 2 Monday through Friday between 8 amd and 6 pm. EST. |  | 305 | 305 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 305 | Not_Computer
Z-0350-2017 | 75283 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TANDEM (TM) COCR SHELL UHMWPE LINER BIPOLAR, |  |  | August 29, 2016 | 2016 | Smith &amp; Nephew, Inc. | Some Bipolar shells were manufactured with an out of specification retainer groove. | Process control | The firm initiated their recall by FedEx and Email on 08/29/2016 to hospitals and distributors and requested the return of any in stock product. International consignees were notified by Email on 08/29/2016. Additionally, the firm notified the US physician of the issue by overnight letter on 09/08/2016. |  | 2352 | 2352 units | Nationwide and Australia, Belgium, Canada, China, Cyprus, France, Germany, India, Italy, Japan, Korea, Mexico, Netherlands, Poland, Portugal, Puerto Rico, Singapore, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates | Open, Classified |  |  | 8158 | Not_Computer
Z-0361-2017 | 64005 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids. List number 16026. Note that Hospira retired this pump June 30, 2015. |  |  | March 16, 2011 | 2011 | Hospira Inc. | May go to a white screen due to a problem with the SOM2 module. A white screen is a diagnostic screen displayed to the user as part of the safety shutdown procedure of the device. The infusion and the processor stop, sound a 3 minute audible alarm from the backup buzzer and deactivate the couch screen and the automated loading/unloading of cassettes. Delay or interruption of therapy. | Device Design | Hospira sent an Urgent Device Recall Letters, dated February 5, 2013, to all affected customers. Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form. Healthcare professionals were to contact Stericycle. Hospira discontinued support of these devices as of June 30, 2015. Customers with questions were instructed to call 1-800-201-9064.For questions regarding this recall call 224-212-4892. |  | 10337 | 10,337 | Worldwide Distribution - US (nationwide) and Canada | Open, Classified |  |  | 21234 | Not_Computer
Z-0363-2017 | 75452 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | X-Force Nephrostomy Balloon Dilation Catheter Kit with Inflation Device, PTFE Sheath, 8 mm x 15 cm, Product Code 996081 |  |  | October 12, 2016 | 2016 | C.R. Bard, Inc. | Contract manufacturer packaged an incorrect sheath with the product. Packaged 30Fr instead of 24Fr. | Unknown/Undetermined by firm | The firm, Bard Medical, sent an "URGENT: MEDICAL DEVICE PRODUCT RECALL" letter dated 10/13/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine product subject to recall; do not use or further distribute any affected product; complete and return the attached Recall &amp; Effectiveness Check Form regardless of whether or not you have any of the affect product by fax to 1-770-784-6469 or email a scanned copy to BMD.FieldAction@crbard.com by October 24, 2016, and if you have further distributed any units of the affected lot, identify your customers and notify them at once of the product recall.If you have any questions, contact BMD Recall Coordinator at 1-800-793-8110. |  | 192 | 192 units | Worldwide distribution: US distribution to: CA, KS, MO, NY, VA, WA, and to countries of: China, Italy, Sweden and United Kingdom. | Open, Classified |  |  | 192 | Not_Computer
Z-0368-2017 | 75483 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ZimmerÂż Air Dermatome II Handpiece Loaner, item number 01-8851-001-00. Device is packaged inside of a padded-foam corrugate box. A product label is applied to the lid of the box. The width plates, when sold individually, are placed in a labeled poly zipper bag, which is then placed into a corrugated shipper carton or padded envelope. |  |  | July 04, 2016 | 2016 | Zimmer Surgical Inc | Zimmer Biomet Surgical has received product complaints reporting that the surface coating applied to the device is blistering, peeling and discoloring after usage over time (see Image 1). There have been no injuries associated with the reported complaints. In addition, an independent assessment determined that the blistering, peeling and potentially flaking of the surface coating is not expected to result in any long range health consequences. Given that the frequency of this surface coating condition can worsen with age, Zimmer Biomet Surgical is currently making preparations for replacement activities to follow in order to prevent future complaints. This document is provided to alert all users of the potential for the surface coating to exhibit blistering, peeling, and discoloring after usage over time. A separate recall removal notice will be issued to facilities, in phases, with detailed instructions beginning in August 2016. Your sales representatives will be notified when a replacement is ready. Do not return product at this time as part of this action.Image 1. | Process design | The firm, Zimmer Biomet Surgical, sent an "URGENT MEDICAL DEVICE RECALL CORRECTION NOTICE" dated 7/6/2016 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:1. Carefully review this letter and ensure all users of the Zimmer Air Dermatome II at your facility have been informed of this notice.2. Inspect the Zimmer Air Dermatome II devices within your facility and indicate on the attached Certificate of Acknowledgment the condition of the device. Complete and sign the attached Certificate of Acknowledgment and return it to corporatequality.postmarket@zimmerbiomet.com.3. Before each use, continue to examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are notpresent, you may continue to use. Immediately after each use, re-examine the surface coating for blistering/peeling. If the blistering/peeling characteristics are present, discontinue use. In the event that alternate devices or therapies are not available, surgeons may, at their discretion, determine the use of the device as a medical necessity and continue to use. After use, follow your facilitys wound irrigation protocol or Zimmer Biomet recommends copious irrigation with a normal saline solution or immersion of the graft in a bowl of normal saline. Zimmer Biomet recommends irrigating the donor site with saline as well. Carefully examine the skin graft and donor site for gray coating flakes and continue to irrigate as necessary. Report this as an adverse event to Zimmer Biomet Surgical at SurgicalRegulatoryReporting@zimmerbiomet.com.If after reviewing this notification you have further questions or concerns please call 1-330-364-0989 between 8:00 am and 5:00 pm EST, Monday through Friday. |  | 1030 | 1030 units | US Nationwide Distribution including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, and WI. | Terminated | June 21, 2017 | 352 | 1030 | Not_Computer
Z-0369-2017 | 75510 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | TBS iNsight Version v.3.0.1 |  |  | October 03, 2016 | 2016 | Medimaps Group | The FRAX adjusted for TBS values are not correct when: The FRAX feature is activated in TBS iNsight; and TBS has been computed from a spine scan where some vertebrae were excluded. | Under Investigation by firm | Medimaps sent an Urgent Field Safety Notice letter dated October 3, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. |  | 15 | 15 units in US | Worldwide Distribution - US Nationwide - in the states of: IL, NC, CA, DE, NY, MD, WI, MN. And the following countries: France,Thailand, Croatia, Canada, Italy, Spain, Russian Federation, Ukraine, Lithuania, South Africa, Brazil, Lebanon, Saudi Arabia, Korea, China Portugalm Iraq, Syrian Arab Republic, Czech Republic, Australia, Belgium, Austria, Romania, New Zealand, Switzerland, Hong Kong, Sweden, Mexico, Islamic Republic of Iran, Serbia, Netherlands, Denmark, and UK. | Terminated | June 07, 2017 | 247 | 15 | Not_Computer
Z-0365-2017 | 75302 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Covidien Devon Light Glove sub-assembly non-sterile |  |  | September 29, 2016 | 2016 | Covidien LLC | Added Instructions for Use: After application, inspect the Light Glove for barrier integrity | Process control | Medtronic issued a Correction Notice on September 29, 2016. Customers were notified via Federal Express and the letter informed customers of the addition of the following statement to the Instructions for Use (IFU): After application, inspect the Light Glove for barrier integrity. Questions or concerns, contact your Medtronic representative or Customer Service at (800) 962-9888 |  | 118800 | 118,800 | Worldwide Distribution - US NationwideForeign:Canada:AndorraAustriaBelgiumCzech RepublicDenmarkFinlandFranceGermanyGreeceIrelandIsraelItalyLatviaLuxembourgMonacoNetherlandsNorwayPolandPortugalRussian Federation Saudi ArabiaSloveniaSouth AfricaSpainSwedenSwitzerlandUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 123750 | Not_Computer
Z-0374-2017 | 75509 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash |  |  | October 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Hardware issue that may expose the user to a potential risk for serious injury due to exposure to rotating or electrical parts. | Nonconforming Material/Component | Siemens mailed a Customer Safety Advisory Notice to affected customers on October 19, 2016, to inform them about the hardware issue. The notice provided guidance on risk prevention and actions they are taking to develop a solution to correct the issue. For further questions, please call (610) 448-6471. |  | 57 | 57 systems distributed in USA | Nationwide Distribution | Terminated | March 13, 2017 | 145 | 57 | Not_Computer
Z-0370-2017 | 75414 | No | KDL | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | GAMBRO Capillary Dialyzer, Polyflux(R) Revaclear MAX, STERILE, Product code 110634 |  |  | October 12, 2016 | 2016 | Baxter Healthcare Corp. | Potential presence of particulate matter on the blood side of the dialyzer | Process control | Baxter sent an Urgent Product Recall letter dated October 12, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:I. Locate and remove all affected product lots from their facility.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.3. Complete the enclosed Baxter Customer Reply Form and return it to Baxter by either fax or scanned email.For questions contact Baxter Product Surveillance at 800-437-5176. |  | 22656 | 22,656 units | Worldwide Distribution - US Nationwide and the countries of Canada and Mexico. | Terminated | August 19, 2017 | 311 | 22656 | Not_Computer
Z-0372-2017 | 75288 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Replacement floor locking cylinders for the Schaerer model 7300 surgical tables. Replacement cylinders are part number 85118. |  |  | September 27, 2016 | 2016 | Schaerer Medical USA Inc | The floor locking cylinders are contained in the base of the Schaerer 7300 model surgical tables. The cylinders were manufactured with the incorrect "O" ring which may cause the cylinder to leak. | Nonconforming Material/Component | Schaerer Medical sent a Field action letters dated September 27, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. The customer is required to complete and return a response form with the number of affected tables and cylinders. All affected cylinders will be replaced in the field and returned to the manufacturer. For questions Quality Assurance at 513-272 -9116. |  | 10 | 10 replacement cylinders. | Worldwide Distribution - US Nationwide in the states of: CA, CT, FL, IA, IN, MI, MN, MO, OH, PA, and VAReplcement cylinders wre distributed | Open, Classified |  |  | 10 | Not_Computer
Z-0309-2017 | 75358 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Marathon Flow Directed Micro Catheter, UltraFlow HPC Flow Directed Micro Catheter |  |  | October 05, 2016 | 2016 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Delamination and detachment of the PTFE (polytetrafluoroethylene) coating material may lead to PTFE coating in the blood stream. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated October 5, 2016, to all affected customers to inform them that Medtronic has identified the potential for an issue with a specific set of lot numbers of Medtronic Neurovascular products as noted where the PTFE (polytetrafluoroethylene) coating could delaminate and separate from the delivery wire. Customers are instructed to complete the attached customer confirmation certificate and fax it to Medtronic at (949) 434-5020 to the attention of Neurovascular Quality.On October 14, 2016, Medtronic announced that it has notified customers of a voluntary recall of certain lots of its Pipeline embolization device, Alligator retrieval device and X-Celerator hydrophilic guidewire. The recall also includes the stylet containing UltraFlow flow directed micro catheters and Marathon flow directed micro catheters. This voluntary recall is being conducted due to the potential separation and detachment of the polytetrafluoroethylene (PTFE) coating on parts of these devices. Should the PTFE separate from the delivery wire or stylets, PTFE particulate could enter the blood stream of the patient. PTFE in the blood stream, based on the size and quantity, could lead to a thromboembolic event.Customers were asked to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 1-949-434-5020 to the attention of Neurovascular Quality. Customers with questions were instructed to contact their Medtronic representative. |  | 84278 | 84,278 units total (17,811 in US) | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lithuania, Malaysia, Mexico, Nepal, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 84278 | Not_Computer
Z-0378-2017 | 75246 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LFIT Anatomic V40 Femoral Head, Low Friction Ion Treatment, Sterile, 36 mm, REF 6260-9-236; Modular components designed to be locked onto a femoral hip stem trunnion during surgery for total hip replacement. |  |  | August 29, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Stryker received several complaints describing incidence of harm secondary to taper lock failure for specific lots of numerous catalog numbers of LFIT Anatomic CoCr V40 Femoral Heads. | Under Investigation by firm | Stryker notified their Branches/Agencies of this recall by e-mail on August 29, 2016 and they were asked to quarantine the affected devices. A Recall Notification Letter and Product Accountability Form was also sent on August 29, 2016 via UPS (with return receipt) to their Branches/Agencies/Hospital Risk Management and Surgeons. On October 11, 2016, Stryker sent an updated recall notification via UPS with return receipt to their affected customers because additional customers and lot numbers were identified. |  | 42519 | 42,519 units (total Catalog numbers) | US Nationwide and Internationally | Open, Classified |  |  | 42519 | Not_Computer
Z-0328-2017 | 75395 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | FindrWIRZ Guidewire System Catalog# 40-05 |  |  | September 29, 2016 | 2016 | SentreHeart Inc | A complaint was received on September 27th, 2016 at a medical center in Germany for delamination and flaking of the PTFE lubricious coating on threeFindrWIRZ Guidewire System .025 guidewires removed from their packaging hoops. | Under Investigation by firm | SentreHeart sent an Urgent Medical Device Voluntary Recall letter dated September 29, 2016, to all affected customers by traceable delivery method. Letters explained the issue and requested that any devices from the two affected lots be returned to SentreHEART immediately. Customers were asked to complete and return the enclosed response form as soon as possible via fax at 650-354-1204 or email to customerservice@sentreheart.com. Customers with questions were instructed to call 1-855-256-7330.For questions regarding this recall call 650-241-6008. |  | 132 | 132 worldwide distribution (98) of those units in US | Worldwide Distribution - US including MO, CA, OR, GA. NH, NY, MI, TX, MD, PA, IL, KS, NE, CT, AZ and Internationally to United Kingdom and Germany | Terminated | March 01, 2017 | 153 | 132 | Not_Computer
Z-0379-2017 | 75448 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Lactic Acid Part # 09D891T21 |  |  | October 11, 2016 | 2016 | Abbott Laboratories | Abbott has identified negative interference from the drug N-Acetyl Cysteine (NAC) with the ARCHITECT Lactic Acid reagent (LN 9D89-21). | Under Investigation by firm | Abbott Laboratories sent an Product Correction Immediate Action Required letter dated October 10, 2016, to all affected customers. Please review this letter with your Medical DirectorTaking the above information into consideration, you may continue to use ARCHITECT Lactic Acid reagent (LN 9D89-21)Please complete and return the Customer Reply FormIf you have forwarded the product listed above to other laboratories, please inform them of this Product Correction and provide to them a copy of this letterPlease retain this letter for your laboratory records.If you or any of the health care providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 1-877-4ABBOTT (available 24 hours a day, 7 days a week). Customers outside the U.S., please contact your local area Customer Service.For further questions please call (847) 937-2550. |  | 26837 | 26, 837 kits | Nationwide and Worldwide Distribution. | Terminated | May 17, 2017 | 218 | 26837 | Not_Computer
Z-0380-2017 | 75409 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Daig Steerable and Daig Supreme Fixed Curve Diagnostic Electrophysiology Catheters |  |  | October 06, 2016 | 2016 | Stryker Sustainability Solutions | The EP Catheters may be mislabeled for French size during reprocessing. | Under Investigation by firm | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/6/2016 to its customers. The letter described the product, problem and actions to be taken.Customers were instructed to discontinue use of the recalled product. Other EP Catheters reprocessed by Stryker Sustainability Solutions should be considered as alternatives. To complete and Return the Recall Effectiveness Check Form, even if no affected product is found, to your local Stryker Sustainability Sales Representative or email to SSSPFA@stryker.com or mail to:Stryker Sustainability Solutions1810 West Drake DriveTempe, AZ, 85283Attn: Jodie RueckertIf you have any questions, contact Senior Director RA at 1(888)888-3433. |  | 4 | 4 devices | U.S. distribution to the following; AZ, MN and MO. | Open, Classified |  |  | 4 | Not_Computer
Z-0386-2017 | 75554 | No | CCQ | Medicinal nonventilatory nebulizer (atomizer). | Anesthesiology | Anesthesiology | 510(K) Exempt | LMA MADdy, Rx only, |  |  | October 27, 2016 | 2016 | Teleflex Medical | Product Defect; These products may produce a straight stream instead of a fully atomized plume of medication. | Process control | Teleflex sent an Urgent Medical Device Recall Notification letter on/about October 27, 2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and quarantine affected products. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to Teleflex Customer Service or email to recalls@teleflex.com. For questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. Teleflex sent an Amended Urgent Medical Device Recall Notification letter dated November 22, 2016, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to continue to return affected product for a full refund per the procedure set forth in the letter of October 27, 2016. Some customers have indicated that, due to medical necessity, they plan to continue using the affected lots rather than return them. Teleflex is advising all such customers to follow the supplemental instructions for Use included as Attachment B. These supplemental instructions allow non-destructive testing of each unit prior to the procedure to determine if it is defective. Institutions that intend to continue using affected lots with this supplemental testing, please send the acknowledgment form to recall@teleflex.com with their name and contact information and the lots that they will continue to use, or call 866-246-6990. Customers with questions should contact their local sales representative or Customer Service at 1-866-246-6990. |  | 2674650 | 2,674,650 units in total | Worldwide distribution. U.S. (Nationwide) and Internationally to Canada, Belgium, Singapore and Thailand. | Open, Classified |  |  | 2674650 | Not_Computer
Z-0377-2017 | 75178 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AtriClip LAA Exclusion System with preloaded Gillinov-Cosgrove Clip |  |  | September 22, 2016 | 2016 | AtriCure, Inc. | There is a potential for a component failure of the device which could necessitate medical intervention. | Device Design | AtriCure sent an Urgent Advisory Notice dated September 21, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and determine if they have any affected product. If so customers should ensure that all users of the affected devices are aware of the Advisory Notice. If customers further distributed the product, they should identify their customers and notify them at once of the Advisory Notice. Any units from the affected lots should be returned to AtriCure, Inc. Customers should contact AtriCure customer service at 1-866-349-2342, and select option 6 to return any affected product(s). Customers were asked to complete and return the enclosed Device Notification Acknowledgement Form as soon as possible. Customers with questions should call 513-755-4563. |  | 298 | 298 | Nationwide Distribution to the follwoing states: AL, AR, AZ, CA, CT, FL, GA, IL, KS, MD, MI, MN, MO, MS, NC, NM, NY, OH, OK, PA, SC, TN, TX, VA, VT &amp; WI.Product was also shipped to the following countries: Poland and Switzerland | Open, Classified |  |  | 298 | Not_Computer
Z-0394-2017 | 75150 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | catalog # 405612 |  |  | August 02, 2016 | 2016 | Becton Dickinson &amp; Company | A recall was initiated by Hospira, Inc. for several lots of 5% Lidocaine HCI and 7.5% Dextrose Injection, USP, 2mL Ampul, NDC # 0409-4712-01. Some BD" Spinal Trays contain the 5% Lidocaine Hydrochloride and 7.5% Dextrose Injection, USP, 2mL Ampul (NDC Number 0409-4712-01) being recalled by Hospira. As a result, BD is initiating a recall for the affected BD Spinal Trays. | Nonconforming Material/Component | BD sent recall letters/business reply card was sent via UPS/email on 8/2/2016.Stericycle is handling returns. |  | 39130 | 39,130 trays | Nationwide. | Open, Classified |  |  | 39130 | Not_Computer
Z-0387-2017 | 75357 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA c111 |  |  | September 27, 2016 | 2016 | Roche Diagnostics Operations, Inc. | cobas c 111 analyzers (catalog numbers 04777433001 and04528778001) with software versions up to and including 4.20 may encounter the following alarm: 7002: 108000572, A software error occurred. This alarm is generated due to a measurement timingerror. Under very rare conditions, the instrument may process two tests in the same cuvette if the runrestarts. If a used cuvette is used again result of the test(s) will be erroneous.These erroneous results may not be flagged.Falsely low or high patient results may lead to incorrect diagnostic measures and medical therapeuticdecisions. The medical risk depends on the parameter. | Software design | Roche sent an URGENT MEDICAL DEVICE CORRECTION letter dated September 27, 2016 to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were asked to complete the attached fax form and fax it to 1-866-808-1159. For questions contact the Roche Support Network Customer Support Center at 1-800-428-2336. |  | 139 | 139 | US Nationwide Distribution | Terminated | July 21, 2017 | 297 | 139 | Computer
Z-0388-2017 | 75388 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SB Knife -Electrosurgical Knife, Sterile |  |  | September 09, 2016 | 2016 | Akita Sumitomo Bakelite Co., Ltd. | Mislabeled product: Sterilization pouch and carton box labeled as Standard type instead of Jr type-Long | Error in labeling | Sumitomo Bakelite on September 9, 2016 initiated removal of all SB Knife from US accounts.The information was emailed to the accounts and Sumitomo Bakelite requested all customers not to use the mislabeled SB Knife and return product for replacement. |  | 15 | 15 units | US in the states of MA | Terminated | November 10, 2016 | 62 | 23 | Not_Computer
Z-0392-2017 | 75475 | No | LTI |  |  | Gastroenterology/Urology | PMA | ReShape Integrated Dual Balloon System, Gen 1, Model No. 01-0013-001, Catalog No. RSM110 |  |  | October 14, 2016 | 2016 | Reshape Medical Inc | ReShape Medical is recalling the Integrated Dual Balloon System Gen 1 due to the potential leak of distillation fluid during balloon inflation. | Device Design | ReShape sent a customer notification letter dated October 14, 2016 was sent to customers to inform them that ReShape Medical is recalling the Integrated Dual Balloon System due to customer reports of instillation fluid leaking during the device inflation procedure. The letter informs the customers that any instillation fluid leakage during the inflation process will cause an incomplete balloon fill and require immediate removal of the device. Customers are instructed to contact ReShape Medical Customer Service at (844) 9377374 for any questions related to the product returns. |  | 114 | 114 units | US Nationwide Distribution in the states of: MO, TX, AZ, CA, NY, TN, GA, LA, FL, MA, OH, IN, IL | Open, Classified |  |  | 114 | Not_Computer
Z-0390-2017 | 75386 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Accessory Adapter Part number 600525A0 |  |  | October 05, 2016 | 2016 | Maquet Medical Systems USA | Maquet Inc. is initiating a voluntary field action on the Accessory Adapter part number 600525A0 due to a complaint in which a head rest mounted to the accessory adapter slipped out of the mountain rail. | Nonconforming Material/Component | Maquet Medical Systems USA sent an Urgent Product Recall Medical Device Removal letter dated October 5, 2016 to the affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:1. Read this Medical Device Removal Notice and disseminate the informationprovided to all the users of the device. 2. Examine your inventory immediately to determine if you have the Accessory Adapter(s). If so, remove the affected products, quarantine them and keep them in a secure location for shipment back to Maquet. 3. Please fill out and sign the enclosed Medical Device Removal Response Form on page 4 and return it either via email to FieldActions@maquet.com or fax to 1-973-909-9927. For question contact your local MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383 (Press option 2, then followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 5:00 p.m. EDT. |  | 9 | 9 Units (US) 141 Units (OUS) | Worldwide Distribution - US Nationwide in the states of AZ, CA, DC, MI, MO, and PA | Open, Classified |  |  | 9 | Not_Computer
Z-0391-2017 | 75412 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OR Table, Top U26 H V, Model # 1909793; OR Table Top U26 H V U, Model # 1909794; OR Table Top U24 H V, Model # 1909796; OR Table Top U24 H V U, Model # 1909797; OR Table Top U24 H V W, Model # 1909798; OR Table Top U14 H V, Model # 1909799 |  |  | October 25, 2016 | 2016 | Trumpf Medical Systems, Inc. | Trumpf has identified six incidents of pinched cables. | Process control | Trumpf Medical sent an Urgent Medical Device Correction letter dated October 7, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken The letter indicated that a authorized service provider will inspect the table tops and make necessary corrections. A confirmation notice was included with the letter which is to be returned. For contact Trupmpf Medical Systems, Inc. at 888-Âż474-Âż9359. |  | 4 | 4 (US) &amp; 229 (Foreign) | Worldwide Distribution - US Nationwide in the states of AZ and SC and the countries of Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Iceland, Italy, Norway, Qatar, Slovenia, Spain, Sweden, Switzerland | Open, Classified |  |  | 4 | Not_Computer
Z-0393-2017 | 75474 | No | DXF | Septostomy catheter. | Cardiovascular | Cardiovascular | 510(k) | NRG Transseptal Needle and NRG RF Transseptal Kit. |  |  | October 11, 2016 | 2016 | Baylis Medical Corp * | The firm is voluntarily recalling lots of the NRG Transseptal Needle and NRG RF Transseptal Kit due to the possibility for an integrity defect to the Tyvek pouch of the NRGTransseptal Needle. | Packaging process control | A recall notification, marked URGENT Voluntary Medical Device Recall was sent by Baylis Medical Company via overnight courier to all customers affected by the recall with an Acknowledgement Form. Customers were asked to examine their inventories, quarantine any affected lots, and complete the attached Acknowledgement Form. The form should be returned to Baylis Medical Company via the instructions provided. Upon receipt, the firm will contact customers with information on how to return affected product. Questions regarding the letter can be directed to the Customer Service Department at 1-800-276-4416. |  | 28093 | 28,093 units | Worldwide Distribution -- US, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Mexico, Netherlands, Spain, UAE, and UK. | Terminated | July 10, 2017 | 272 | 28093 | Not_Computer
Z-0399-2017 | 75419 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Merge LIS software. |  |  | March 05, 2015 | 2015 | Merge Healthcare, Inc. | There is a potential for duplicate container numbers to be created for patients. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016. If no email address was available, the letters were issued via certified mail. Letters to nonresponders were issued 5/26/2016. |  | 413 | 413 sites potentially have the affected software | Distribution was made nationwide and to the Virgin Islands. Foreign distribution was made to the Bahamas. | Open, Classified |  |  | 413 | Not_Computer
Z-0401-2017 | 75256 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista IRON Flex reagent cartridges |  |  | August 25, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has become aware of deferoxamine interference with Dimension and Dimension Vista IRON assays at a different concentration than listed in the Instructions for Use (IFU). | Error in labeling | Siemens mailed an Urgent Medical Device Correction letter to all customers who received the Dimension and Dimension Vista IRON Flex reagent with an update to the Limitation of Procedure section for the Dimension and Dimension Visa IRON assays Instructions For Use (IFU) to indicate that: patients treated with metal-binding drugs (e.g. deferoxamine) may have depressed iron values, as chelated iron may not properly react in the iron assay. Customers were told that the information provided in the letter, related to deferoxamine, supersedes the information related to deferoxamine int he current IRON IFU's for Dimension and Dimension Vista products until the IFU's are updated. |  | 30247 | 30,247 units | Distributed Nationwide | Open, Classified |  |  | 68477 | Not_Computer
Z-0398-2017 | 75394 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA CentaurÂż Vitamin D Total Assay |  |  | October 05, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | Vitamin D Total- Change in Correlation Between Serum and Plasma Specimen Tubes- negative bias with plasma vs. serum samples | Device Design | Siemens issued an Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on October 5, 2016 to be delivered to customers on October 6th, 201 6. An Urgent Field Safety Notice ~UFSN) was sent out for distribution to all customers outside the United States on October 5, 2016. These notices inform the customer ofthe change in correlation between serum and plasma specimen tube results generated on the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT Vitamin D Total assay. A revised letter was issued on November 7, 2016 by Fedx for delivery on November 8, 2016 and advises users to: In the interim, Siemens requires that customers discontinue use of plasma tubes for specimen collection for the ADVIA Centaur Vitamin D Total assay until further notice.Customers can continue to use the ADVIA Centaur Vitamin D Total assay with serum specimen collection tubes.Accounts to Complete and return the Field Correction Effectiveness Check Form attached within 30 days. Review the letter with your Medical Director.and retain this letter with your laboratory records, and forward this letter to those who may have received this product.Question contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 37437 | 37437 kits | NationwideForeign:Afghanistan, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil,Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt,Finland, France, French Polynesia, Georgia, Germany, Great Britain, Greece, Guadeloupe,Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Jordan, Latvia, Lebanon,Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Nepal, Netherlands,New Zealand, Norway, Oman, Pakistan, Paraguay, Philippines, Poland, Portugal, Qatar,Romania, Russian Federation, Saudi Arabia, Singapore, Serbia, Slovakia, South Africa, SouthKorea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates,Uruguay, Vietnam | Open, Classified |  |  | 107218 | Not_Computer
Z-0400-2017 | 75387 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | PET Infusion Pump |  |  | September 30, 2016 | 2016 | Bayer Healthcare | The Source Administration Sets (SAS) used with the Medrad Intego PET Infusion System do not meet pressure testing specifications in an occlusion situation when the product is more than one year past manufacturing date. | Process control | Bayer mailed an Urgent Medical Device Customer Notification letter to the affected consignees on September 30, 2016 asking that they review their current SAS inventory and determine if you have any product that is more than one year past its manufacturing date, Quarantine unused units and call Bayer Customer Care line to receive a Return Goods Authorization (RGA) number for the product's return. |  | 165127 | 165,127 (79,714 units in U.S.) | Distributed Nationwide and internationally to : Taiwan, Ukraine, Turkey, Thailand, Slovania, Singapore, Sweden, Saudi Arabia, Serbia, Romania, Qatar, Portugal, Poland, Panama, Oman, Norway, The Netherlands, Mexico, Malta, Morraco, Libyan, Lithuania, Lebanon, Kuwait, Italy, Israel, Hungray, United Kingdom, France, Finland, Spain, Denmark, Germany, Czech Republic, Chile, Switzerland, Belguim, Bosnia, Austria, Australia, | Terminated | March 24, 2017 | 175 | 165127 | Not_Computer
Z-0395-2017 | 75366 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Tina-Quant Hemoglobin A1c Gen. 2 |  |  | October 04, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Roche Diagnostics has confirmed elevated QC and patient sample recovery for individualTina-quantÂż Hemoglobin A1c Gen.2 and Tina-quant Hemoglobin A1cDx Gen.2 cassettes/bottles.This issue occurred over multiple lots. Single cassettes/bottles of the reagents are affected.The issue can lead to erroneously high HbA1c results. If a patient with known diabetes is affected,an erroneously high HbA1c result might lead to therapeutic consequences, such as therapy escalation(introduction of another oral antidiabetic medication or insulin) or increase of the dosage. This canfurther lead to an increased risk of hypoglycemia. Considering the unreliable detectability of the issue,relevant medical risk for the patient cannot entirely be excluded. | Storage | On 10/4/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected Direct Consignees and Distributors via courier. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Direct ConsigneeActions Required" Follow the instructions in the Customer Actions for Prevention/Detection of the Issue section of this Urgent Medical Device Correction (UMDC)." Complete the attached fax form (7184-00-1016) and fax it to 1-844-294-7181." File this Urgent Medical Device Correction (UMDC) for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a weekat 1-800-428-2336 if you have questions about the information contained in this UMDC.DistributorActions Required" Ensure proper storage conditions (2-8ÂżC) are maintained for the reagents listed in the table above,per their respective method sheet." If you have not directly shipped any of the products in the table above, you do not need to distributethe enclosed UMDC to your customers." If you have directly shipped any of the products in the table above, please distribute the enclosedUMDC 16-157 to those customers." Complete the attached fax form (7185-00-1016) and fax it to 1-844-294-7181.Questions Please contact the Roche Support Network Customer Support Center at 1-800-428-2336 if you havequestions about the information contained in this distributor letter or the enclosed UMDC 16-157. |  | 4976 | 4976 | Nationwide. | Terminated | June 05, 2017 | 244 | 4976 | Not_Computer
Z-0403-2017 | 75492 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Fluoro-4 Silicone Ureteral Stent, BardÂż Fluoro-4", Silicone Ureteral Coil Stent; BardÂż Silicone Uretera l Coil Stent with Figure Four, End, Manufacturer: C.R. Bard, Inc., Covington, GA 30014. |  |  | October 13, 2016 | 2016 | C.R. Bard, Inc. | Labeling: Statement was added to the Warnings/Precautions section requesting end users to consider the presence of a knot if significant resistance is encountered during attempts at removal. | Labeling Change Control | Bard Medical sent an Urgent: Medical Device Communication Letter dated October 13, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action required:Examine your inventory and identify any product subject to this notification. If you have any remaining inventory, include a copy of this communication with the unit(s) for future use.If you may have further distributed or transferred this product, please identify the respective organizations and notify them at once of this communication. Your notification to these organizations may be enhanced by including a copy of this product communication letter.This notification should be carried out to the user level.This product notification is being made with the knowledge of the Food and Drug administration.We appreciate your cooperation and assistance in dealing with this matter and sincerely apologize for any inconvenience that may result from this action.For further questions, please call (770) 784-6220. |  | 75089 | 75,089 units | U.S. Nationwide. | Open, Classified |  |  | 75089 | Not_Computer
Z-0414-2017 | 75491 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | 4 Port "ON" Manifold (600 psi) w/4 ft. TranspacÂż IV, Item No. 46300-14NS |  |  | September 27, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified a potential for leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. | Device Design | ICU Medical sent an Urgent Medical Device Recall Notification letter dated September 27, 2016, to all affected customers to inform them of the expanded recall of the monitoring kits with Transpac IV transducers. Customers were informed of updated communications from August 10, 2016, when ICU Medical sent a communication indicating a potential leak with certain SafeSet Reservoir and blood sampling kits containing the Transpac IV transducers. Based on continued evaluation, ICU Medical is expanding the notification to now include certain other monitoring kits that contain the Transpac IV transducers. A small number of these devices may have the potential to leak or potentially introduce small amounts of air into the line. Customers with questions were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 2, or e-mail the following address: productreturns@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or e-mail to recall@icumed.com. Customers with questions were instructed to call 866-829-9025. |  | 20 | 20 units | Nationwide Distribution | Terminated | July 11, 2017 | 287 | 210737 | Not_Computer
Z-0411-2017 | 75494 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 20, 2016 | 2016 | Mevion Medical Systems, Inc. | Software defect:2D projection of contours in Verity. The defect causes structures to appear off from their true position in the DRR and the radiograph with an error that increases with the distance of the structure from isocenter | Software design | Mevion issued an Important Safety Notice letter on October 14, 2016 to their customers. The letter identified the affected product, problem, containment, mitigation and actions taken. Customers were informed that Mevion will be updating software when it is released. For further Information, contact: srosenthal@mevion.com. |  | 6 | 6 units | US Nationwide Distribution in the states of: FL, MO, NJ, OH, OK | Terminated | May 15, 2017 | 207 | 6 | Computer
Z-0409-2017 | 75496 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 31, 2016 | 2016 | Mevion Medical Systems, Inc. | Completing Setup Without Program Move Complete Can Cause An Isocenter Error | Software design | Mevion issued an Important Safety Notice letter on October 31, 2016 to customers. The identified the affected product, problem, mitigation and advisory information. The letter informed customers that Mevion will release and install a software patch that removes the reliance on tracking isocenter changes for the correct reporting of the Couch position to the user displays and to the OIS. For further information mailto:srosenthal@mevion.com |  | 6 | 6 units | US Nationwide in the states of FL, MO, NJ and OK | Terminated | June 29, 2017 | 241 | 6 | Computer
Z-0410-2017 | 75498 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250 |  |  | October 21, 2016 | 2016 | Mevion Medical Systems, Inc. | Recovery of Incorrect Isocenter -operational change | Software Design Change | Mevion issued an User Notice letter on October 24, 2016 to their customers. The letter identified the affected product, problem and actions taken. Customers were informed that Mevion will be updating software when it is released. For further clarification or assistance, please contact Rob Cessac by email at rcessac@mevion.com or by phone at (978) 540-1501. |  | 6 | 6 units | US Nationwide Distribution in the states of FL, MO, NJ, OH and OK, | Terminated | May 10, 2017 | 201 | 6 | Computer
Z-0407-2017 | 75526 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Ostase Calibrators, Catalog No. 37305, B83876 |  |  | October 24, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has determined that the Access Ostase Calibrator and Access Ostase QC lots do not meet their expiration date claims within the 10% criteria. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated October 24, 2016 was sent to all their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions regarding the notice are instructed to contact Customer Technical Support Center at http://www.beckmancoulter.com, (800) 854-3633 in US and Canada, and to contact local Beckman Coulter representative if outside of US and Canada. |  | 357 | 357 kits total (94 kits in US) | US Nationwide Distribution in the states of: UT, MI, MA, LA, IN, NC, CA, TX, NJ, IL, WI, MN, FL, KA, NY, CO, WA, KY, VA, and AL.Affected geographies: Belgium, Canada, China, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Italy, Japan, Mexico, Philippines, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Taiwan. | Open, Classified |  |  | 357 | Not_Computer
Z-0413-2017 | 75456 | No | KRG | Pacemaker programmers. | Cardiovascular | Cardiovascular | PMA | The ORCHESTRA PLUS Programmer |  |  | September 29, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the ORCHESTRA Programmer and ORCHESTRA PLUSProgrammer because the residual longevity estimate may be less accurate. | Software Design Change | The firm, LivaNova, sent an "Urgent Field Safety Notice" dated 9/29/2016 to its customers. The notice described the product, problem, actions to be taken. The notice states that a Sorin representative will contact customers to upgrade the software in the facility's programmer.If customers have any questions, they are urged to call their LivaNova representative or contact LivaNova Technical Services at 877.663.7674. |  | 586 | 586 programmers | Programmers were distributed nationwide, to VA/govt/military, Canadian, and other foreign consignees. | Open, Classified |  |  | 1433 | Computer
Z-0589-2017 | 75487 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Advia Chemistry Calibrator |  |  | October 07, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the DBIL_2 Assigned Calibrator Values (ACV) currently assigned to the affected lots may cause, on average, a positive proportional bias of approximately 41% on patient samples compared to the reference method for Direct Bilirubin measurement. | Under Investigation by firm | Siemens issued an Urgent Medical Device Correction in the US and an Urgent Field Safety Notice Outside the US beginning on October 6, 2016. Customers were notified of the issue and instructed to use the reassigned DBIL_2 calibrator value for Chemistry Calibrator Lot Numbers 680726, 680726A, 680726B, 680726C and 680726D. Customers were also requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days by faxing to the Customer Care Center at 312-275-7795.For questions regarding this recall call 914-631-8000. |  | 10614 | 10,614 units | Worldwide Distribution - US (Nationwide) Internationally to Canada, Mexico, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Chile, Cyprus, Denmark, Egypt, Ecuador, Finland, France, Germany, Great Britain, Hong Kong, Hungary, Italy, Japan, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Paraguay, Peru, Poland, Portugal, Russian Fed. Slovakia, Slovenia, Spain, Sweden, Switzerland, Thailand, Turkey, Vatican City and Vietnam | Open, Classified |  |  | 10614 | Not_Computer
Z-0584-2017 | 75485 | No | GXD | Radiofrequency lesion generator. | Neurology | Neurology | 510(k) | RF MultiGenÂż, 0406-900-000; Refurb RF MultiGenÂż, 0406-900-000U |  |  | October 11, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the Care Instructions/Instructions for Use packaged with the Radiofrequency (RF) MultiGenÂż Generator, because the instructions for use incorrectly contained instructions supporting the sterilization of the MultiGenÂż cables. | Labeling Change Control | Stryker Instruments initiated a voluntary recall of the RF MultiGen, and Refurb RF MultiGen because the instructions for use contained incorrect instructions on sterilizing the device. Stryker distributed customer notices via certified mail on 10/11/2016. Customer's were instructed to do the following: Immediately review this recall notification.2. Enclosed was a revised version of the IVS Care Instructions. Please distribute the revised version to the appropriate person or department.3. Remove and discard any previously distributed IVS Care Instructions or Instructions for Use.4. Please complete, sign, and return the enclosed Business Reply Form (BRF) acknowledging your receipt and understanding of this notification.5. Fax the completed BRF to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com. Customers with questions can contact: 269 323 7700 . |  | 2675 | 2,675 | Domestic:AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WYForeign:France, Germany, Canada, UK, Lebanon, South Africa, Chile, Hong Kong, Poland, Argentina, Spain, Greece, Italy, India, Brazil, Switzerland, Netherlands, SwedenVA/DOD: | Terminated | July 25, 2017 | 287 | 2675 | Not_Computer
Z-0585-2017 | 75428 | No | FAS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | C-Max" Cutting Loop Electrodes, Storz One-Stem, 24 Fr, .014in., .35 mm, Sterile, Rx only, Manufactured by: C.R. Bard, Inc., Covington, GA 30014. |  |  | October 12, 2016 | 2016 | C.R. Bard, Inc. | Defect; The contact part could possibly separate from the electrode affecting to functionality of the device. | Other | Consignees were notified via letter on 10/12/2016. |  | 70 | 70 units | FL, GA, LA, OH, VA and WA. | Open, Classified |  |  | 70 | Not_Computer
Z-0592-2017 | 75337 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Arsenal Spinal Fixation System, Set Screw, Part No. 47127 |  |  | October 05, 2016 | 2016 | Alphatec Spine, Inc. | Alphatec Spine is recalling the Arsenal Spinal Fixation System Set Screw due to a trend in set screw postoperative disengagement from the screw body. | Under Investigation by firm | Alphatae Spine sent an Urgent Medical Device Recall notification letter dated September 29, 2016 to customers. The letter identified the affected product, problem and the actions to be taken. Customers are instructed to fill out and return the last page of recall letter using one of the methods identified in the letter. |  | 16972 | 16,972 units | Worldwide Distribution - US Nationwide in the countries of Japan, UK, NSW, Spain. | Open, Classified |  |  | 16972 | Not_Computer
Z-0591-2017 | 75417 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | FlexLab Automation Modules |  |  | September 29, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. |  | Under Investigation by firm | Siemens mailed to customers an Urgent Field Safety Notification letter warning them of the problem and to inform them that Siemens Customer Service Engineers shall be visiting customer sites to replace the power cords. |  | 448 | 448 units | Distributed Nationwide and to the following countries:Australia, Azerbaijan, Bahrain, Bangladesh, Belgium, Brazil, Canada, Czech Republic, Denmark, Egypt, Germany, Greece, Indonesia, Italy, Kingdom Saudi Arabia, Lebanon, Malaysia, P.R. China, Philippines, Portugal, Qatar, Republic Korea, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., United Kingdom, Vietnam | Open, Classified |  |  | 448 | Not_Computer
Z-0582-2017 | 75347 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium Medical 28 Fr Trocar catheter, Sterile |  |  | October 26, 2016 | 2016 | Atrium Medical Corporation | Trocar Catheters are packaged with a label containing an icon that shows more holes (eyelets) than the product actually contains. | Labeling design | Atrium Medical issued an "Urgent Medical Device Recall (Removal)" communication to customers on 10/26/16 via FedEx 2 day delivery. The recall identifies the problem, product, and risk factors. In addition, it provides instructions on how to handle product that may be affected. Accounts are to read the letter completely and share with all users and staff. They have been requested to complete the enclosed Recall Reply Form, acknowledging receipt of the notice; and, remove and quarantine any affected product for return. Additionally, obtain product from an alternate supplier because replacement product is not available as all Trocar catheters have been discontinued. Contact Atrium Medical Customer Service at 1-800-370-7899, Monday through Friday between 9:00 am to 5:00 pm, for a Return Goods Authorization to return the product and receive credit. |  | 7043 | 7043 cases (10/cs) | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Costa Rica, Dominican Republic, Ecuador, Finland, France, Germany, Honduras, Hong Kong, Ireland, Jamaica, Kuwait, Mexico, Netherlands, New Zealand, Panama, Peru, Portugal, Puerto Rico, Qatar, Reunion, Romania, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, and Venezuela. | Open, Classified |  |  | 31247 | Not_Computer
Z-0587-2017 | 75513 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FlexLab, Accelerator a3600 and Aptio Automation are modular systems designed to automate Pre-Analytical and Post-Analytical processing, sample handling in order to automate sample processing in the Laboratory. |  |  | October 04, 2016 | 2016 | Inpeco S.A. | Inpeco received several notifications from the field about burnt schuko sockets. The customers reported that some smoke was seen coming from the damaged sockets. | Employee error | Inpeco sent an Urgent Field Safety Notice dated October 2016 for the Aptio Automation, the Accelerator a3600 and the FlexLab Shucko power cord cables to all affected US and foreign customers. The Urgent Field Safety Notice described the reason for the recalls and actions that should be taken.The firm is requesting all customers to complete and return the Field Correction Effectiveness Check form attached to the Urgent Field Safety Notice letter within 7 days. The Urgent Field Safety Notice letter should be forwarded to affected departments or consignees.To mitigate any possible risks to the laboratory users' health conditions due to the defective power cords Inpeco released a Technical Service Bulletin (TSB) addressed only to Field Service Engineer with the procedure which has to be performed in order to prevent other cases of damaged sockets. The FSE has to check visually the appearance of the Schuko female socket. If it is intact, without any detectable deformation nor melting the FSE has to verify that the terminals are correctly crimped into each single wire. If the female socket presents alterations, the socket has to be replaced with a new one. The affected sockets will be destroyed.Customers with questions were instructed to call (+41) 91 9118 258. |  | 188 | 188 units | Worldwide Distribution - US including Texas and New York. Foreign Consignees: Germany, Italy, France, Austria and Spain. | Open, Classified |  |  | 188 | Not_Computer
Z-0586-2017 | 75241 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS) released on July 11, 2016 |  |  | September 14, 2016 | 2016 | Roche Diabetes Care, Inc. | Roche Diabetes Care, Inc. initiated a voluntary recall for the version 1.2.0 of Accu-Chek Connect Diabetes Management App (iOS), released on July 11, 2016, due to the software containing a programming error (bug) in the Bolus Advisor feature, which could lead to an incorrect insulin bolus recommendation. | Software design (manufacturing process) | Roche Diabetes Care, Inc., initiated a voluntary recall on September 14, 2016, of the version 1.2.0 of the Accu-Chek Connect app on iOS products (iPhone, iPad). The Prescription Activation Service (PAS) which allows the users to activate the bolus advice feature was placed on product hold immediately. This measure prevented any new bolus advice activations. An updated version (V 1.2.2) of the Accu-Chek Connect Diabetes Management App for iOS that contains the fix for the issue was released on August 29, 2016. Patients were instructed to:Âż If you use the Accu-Chek Connect App on an IOS product (iPhone, iPad), please verify that you have installed the newest version of the app.Âż Contact Accu-Chek Customer Care at 1-800-628-3346 if you have additional questions.HCPs were instructed to provide a copy of the UMDC to those patients to whom they have prescribed the Bolus Advisor on the Accu-Chek Connect App. No product returns are expected. Customers with questions may contact Accu-Chek Customer Care at 1-800-628-3346. For further questions please call (317) 521-4312, |  | 7909 | 7,909 downloads of iOS version 1.2.0 of the Accu-Chek Connect Diabetes Management App in the US and 8,775 downloads Internationally | US Distribution to the states of :Alabama, Arizona, Arkansas, California, Colorado, Connecticut, DC, Delaware, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Virginia, Washington, West Virginia, WisconsinForeign: None provided | Terminated | June 06, 2017 | 265 | 7909 | Computer
Z-0588-2017 | 75443 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Triage Total 5 Control Level 1, Model No. 88753 |  |  | October 12, 2016 | 2016 | Alere San Diego, Inc. | Encoded with incorrect ranges that are not consistent with the ranges printed on the Expected Value Card. | Employee error | The firm, Alere, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated 10/2016 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of and discard the CODE Chip contained within the Alere Triage Total 5 Control Level 1 PN 88753 Lot C3233A; install the replacement QC Sample CODE Chip for Lot C3233AR; share this information with your laboratory staff and retain this notification as part of your laboratory QS documentation; if you have forwarded the product to another laboratory, provide copy of this notification; and complete and return the Customer/Distributor Verification Form via Fax to Alere Technical Services at +1 858 805 8457 within 10 business days. Customer with any questions about the information contained in the notification are instructed to contact:Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121, U.S.A.Phone: 844-320-5124FAX: 858-805-8457E-mail: Responses.ts@alere.comOn 11/10/16 a revised Urgent Medical Device Correction letter was sent to their customers to inform them of the recall. |  | 513 | 513 kits | US Nationwide Distribution | Terminated | July 24, 2017 | 285 | 513 | Not_Computer
Z-0597-2017 | 75516 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core Computed Tomography X-ray system, Model 728321 |  |  | October 13, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The product label does not include the correct current rating. | Labeling Change Control | The firm will send a customer letter via certified mail and will have the receipts returned for tracking purposes. Also, a Field Service Engineer will implement the Field Change Order on affected systems to resolve the issue. Implementation of the Field Change Order will be monitored. |  | 2 | 2 | Worldwide distribution. The products were shipped to the following states: AL, AZ, CO, GA, IN, KY, LA, MA, MN, NY, OH, PA, SC, TX, and UT.There are five (5) government accounts for this recall: DEPARTMENT OF VETERANS AF, VETERANS WAY AND LAMONT ST, Mountain Home, TN 37684; US ARMY EVANS COMMUNITY H, 1650 Cochrane Cir, Colorado Springs, CO 80913; US VETERANS ADMINISTRATION, 3710 SW US VETERANS HOSPITAL, PORTLAND, OR 97239; VA MEDICAL CENTER, 4100 W 3RD ST, DAYTON, OH 45428; VAMC - PHILADELPHIA, UNIVERSITY &amp; WOODLAND AVE, PHILADELPHIA, PA 19104.There are five (5) Canadian accounts for this recall: FIVE HILLS HEALTH REGION, 55 DIEFENBAKER DR, MOOSE JAW, SASKATCHEWAN S6J 0C2; INSTITUT UNIV CARDIOLOGIE, 2725 CH STE FOY, SAINTE FOY, QUEBEC G1V 4G5; LAKERIDGE HEALTH, 1 HOSPITAL CRT, OSHAWA, ONTARIO L1G 2B9; LAKERIDGE HEALTH BOWMANVI, 47 LIBERTY ST S, BOWMANVILLE, ONTARIO L1C 2N4; MCGILL UNIVERSITY HEALTH, 1051 BOUL DÂżCARIE, MONTRÂżAL, QUEBEC H4A 0A5.There were zero (0) Mexican accounts for this recall.The recalled products were shipped to the following countries: Argentina, Australia, Belgium, China, Egypt, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Republic of Korea, Lebanon, Netherlands, New Zealand, Norway, Russian Federation, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 117 | Not_Computer
Z-0611-2017 | 75497 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iConnect Access software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | May 25, 2016 | 2016 | Merge Healthcare, Inc. | Software displayed incorrect prior reports in the viewport area, only when more than one prior study (2 or more) was viewed. | Software design | The recalling firm issued a notification letter dated 1/27/2016 via email or via certified mail if the firm did not have an email address for the customer. Nonresponders were issued a second letter dated 5/25/2016 on 5/26/2016. |  | 60 | 60 sites potentially have the affected versions | Distribution was made to medical facilities nationwide. Foreign distribution was also made. | Open, Classified |  |  | 60 | Computer
Z-0604-2017 | 75555 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS, Computed tomography system |  |  | October 19, 2016 | 2016 | Siemens Medical Solutions USA, Inc | A problem can occur with the small light marker windows, used for the positioning laser and are integrated in the front cover of the SOMATOM Definition AS and Definition Flash systems, by either becoming loose or potentially dropping out. There is a risk that a person could touch rotating or electrical parts of the gantry if they were to reach through the window opening. | Nonconforming Material/Component | Siemens' sent a Customer Safety Advisory Notice letter dated October 19, 2016 to affected customers. The letter identified the affected product, problem, potential risks, how to avoid the potential risk related to this issue and actions to be taken. |  | 762 | 762 systems distrtibuted in U.S. | US Nationwide Distribution | Terminated | March 13, 2017 | 145 | 4599 | Not_Computer
Z-0610-2017 | 75499 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | OMNIBotics Tracker Kit -Model: RM-10000 |  |  | October 14, 2016 | 2016 | OMNI LIFE SCIENCE | Incorrect lot number on outer kit | Error in labeling | The firm, OMNIlife Science, Inc., issued an "URGENT MEDICAL DEVICE ADVISORY NOTICE" dated 10/14/16 by email to its consignees/customers. U.S. and foreign agents were also sent the notification packet via overnight mail. Acknowledgement of receipt is requested. The notice described the product, problem and actions to be taken. The customers were instructed to immediately quarantine lot 19-09-021 for OMNIBotics Tracker Kit RM-1000; overlabel the outside packaging with the label provided following the provided training paperwork and rework instructions; note advisory notice only includes listed product, it does not affect other lots; and complete and return the Acknowledgement and Receipt Form and accompanying FN-102 by Fax to 508-519-3390.Accounts are to hold item RM-10000 with the lot # 19-09-021 until a corrected label is provided for the outer box. The patient labels inside the box are correct.For further information , Regulatory Affairs Specialist at 774-226-1847 or 800-448-6664 Monday through Friday, 8:00 AM to 4:30 PM, Easter Time. |  | 15 | 15 units | Worldwide Distribution: US distribution to OK, MD, TX and country of: Australia. | Open, Classified |  |  | 15 | Not_Computer
Z-0600-2017 | 75365 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 Antibacterial Hydrophilic Intermittent Catheter, Male 14 Fr. (30/Box), Sterile |  |  | October 05, 2016 | 2016 | C.R. Bard, Inc. | Misbranding; the product labeled as an Antibacterial Hydrophilic Intermittent Catheter, is instead an Antibacterial Intermittent Catheter (no Hydrophilic Coating). | Mixed-up of materials/components | Bard Medical sent an Urgent Medical Device Product Recall letter dated October 5, 2016 to customers The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine and quarantine and return affected inventory. Customers were asked to complete and return the accompanying Recall &amp; Effectiveness Check Form attached. |  | 10290 | 10,290 units | US Nationwide Distribution in the states of: CA, FL, IL, IN, MN, MO, NC, NJ, NY, OH, OK, PA, TX and UT. | Open, Classified |  |  | 10290 | Not_Computer
Z-0608-2017 | 75677 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash, Computed tomography x-ray system |  |  | November 02, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is releasing a software update that provides bug-fixes to improve system performance for customers with SOMATOM Definition AS, SOMATOM Definition Edge, SOMATOM Definition Flash Systems with software version VA48A_SP2 and Care Contrast license correction of potential safety issues. | Software design | Siemens mailed a Customer Safety Notice on November 9, 2016, to customers who were affected by the recall to inform them of a possible malfunction with the systems. The notification explained the underlying issue, the effect if has on the operation of the system, potential risks involved; and the actions they plan to implement to correct the issues. Customers with questions were instructed to call 1-800-888-7436.For questions regarding this recall call 610-448-6471. |  | 7 | 7 systems | Nationwide distribution to MI, NY, CA, KY, ND, and NE. | Open, Classified |  |  | 7 | Computer
Z-0609-2017 | 75187 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Model 208-70-110, Long Thread Tiger Large Cannulated Headless Screw, 7 x 110 mm labeled as "7.0 x 110mm LTTHCS REF 208-70-110 QTY 30 LOT TSL003721 2016-03-29 Trilliant Surgical, LTD" |  |  | September 02, 2016 | 2016 | Trilliant Surgical Ltd. | Screws were identified as having the incorrect part number and lot number marked upon them. The part number incorrectly indicated a length of 115 mm rather than 110 mm. The lot number incorrectly indicated lot number of TSL003765 rather than TSL003721. | Nonconforming Material/Component | The firm sent out written notifcation of the recall on 9/2/16; the letter instructed the consignees to discontinue use of the affected lot, quarantine, and return product of the affected lot in accordance with an included acknowledgement sheet. |  | 11 | 11 | CA, AZ, WI, FL, NJ, OH | Terminated | May 19, 2017 | 259 | 11 | Not_Computer
Z-0599-2017 | 75629 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | The software did not show unviewed images when the last view was skipped. The physician may potentially misdiagnose when not all images are available for viewing. | Software design | The recalling firm issued a letter dated March 28, 2016 via e-mail or via certified mail to customers. The letter identified the affected product, problem and actions to be taken. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. |  | 534 | 534 sites potentially have the affected versions | Worldwide Distribution - US Nationwide including Canada. | Open, Classified |  |  | 534 | Computer
Z-0603-2017 | 75606 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis, Angiographic x-ray system |  |  | November 01, 2016 | 2016 | Siemens Medical Solutions USA, Inc | The possibility exists that the semi-automatic focus switch may not work as intended when pressing the footswitch multiple times. | Software in the Use Environment | Siemens mailed a Customer Safety Advisory Notice to affected customers to inform them of a corrective action they plan to perform to prevent a possible hazard to persons and equipment. The notice identified the associated risks, underlying issues, impact on system operation and actions to be taken. |  | 4657 | 4,657 systems Worldwide | US Nationwide Distribution. | Terminated | June 21, 2017 | 232 | 4657 | Not_Computer
Z-0612-2017 | 75523 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA 19-9 Assay (50 test kit)-for use, with the ADVIA Centaur and ADVIA Centaur XP, ADVIA Centaur XPT and ADVIA Centaur CP systems. |  |  | October 19, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc | CA 19-9 assay on the ADVIA Centaur CP does not meet the correlation to the CA 19-9 assay on the ADVIA Centaur/XP/XPTsystems as stated in the Instructions for Use. | Process design | Siemens Healthcare issued Urgent Medical Device Correction (UMDC) to Siemens Healthcare Diagnostics customers in the United States on October 19, 2016 to be delivered to customers on October 20, 2016. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on October 18, 2016. The notice informs the customer of the potential to observe a bias between CA 19-9 results generated on the ADVIA Centaur CP and the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT systems. Customers may continueto use the ADVIA Centaur CA 19-9 assay to report patient results on the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT and ADVIA Centaur CP systems. Customers should not use the ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT systems interchangeably with the ADVIA Centaur CP when generating CA 19-9 results.For customers using the ADVIA Centaur CP and ADVIA Centaur/XP/XPT for CA 19-9 interchangeably, a look back of patient results over the previous three months is recommended.A Field Correction Effectiveness Check is to be completed, signed and returned Siemens. Questions, contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 18733 | 18,733 kits | NationwideForeign:Albania, Algeria, Angola, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Cambodia, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Germany, Great Britain, Greece, Guadeloupe, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy,Japan, Jordan, Kazakhstan, Latvia, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, Nepal, Netherlands, New Zealand, Nicaragua, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uruguay, Vatican City, Vietnam | Open, Classified |  |  | 48955 | Not_Computer
Z-0622-2017 | 75661 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion Prime TSX-303A |  |  | October 25, 2016 | 2016 | Toshiba American Medical Systems Inc | Toshiba American Medical Systems (TAMS) is recalling the Aquilion CT system because ithas been reported that the cooling oil in the x-ray tube heat exchanger may leak during system operation. | Device Design | Toshiba American Medical Systems, Inc. sent an Urgent Medical Device Correction letter dated October 25, 2016, to all affected customers who received the Aquilion CT system to inform them that Toshiba America Medical System has received reports that the cooling oil in the x-ray tube heat exchanger may leak during system operation. The letter informs the customers of the description found and corrective action. Customers were instructed to complete the attached customer reply form and fax it to (877) 349-3054 or send it to raffairs@tams.com. Customers with questions are instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 291 | 291 units | Nationwide Distribution | Terminated | August 19, 2017 | 298 | 291 | Not_Computer
Z-0623-2017 | 75145 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Greatbatch Stiffer Coaxial Micro-Introducer. Model #/ Name: 10904-001 Stiffer Coaxial Micro-Introducer 4FR; 10904-002 Stiffer Coaxial Micro-Introducer 5FR. |  |  | September 08, 2016 | 2016 | Greatbatch Medical | Greatbatch Medical has identified an issue with the Stiffer Coaxial Micro-Introducer 4FR and 5FR after a field complaint reported that the rotating luer detached from the dilator hub during use. | Process design | Greatbatch sent a "Urgent Medical Device Recall" letter dated September 7, 2016 to all affected customers .The letter described the product, "Reason for Field Correction Action", "Risk to Health" , "Replacement Devices", &amp; "Instructions to Customers" . Advised customers to quarantine the product and contact end users and share the recall information with them. Requested consignees to complete &amp; return the Field Recall Response Form. For questions contact your local Greatbatch Medical representative, contact us at 1-763-951-8235, or e-mail us at FieldActionCenter@Greatbatch.com. |  | 23075 | 23,075 | US: NY, TX. | Open, Classified |  |  | 23075 | Not_Computer
Z-0614-2017 | 75267 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite,f/k/a: Emageon Advanced Visualization , RadSuite versions 8.30.7.8, 8.30.7.9, 8.30.8, and 8.31. |  |  | July 14, 2016 | 2016 | Merge Healthcare, Inc. | It was reported by a customer that RadSuite images are not appearing as readable images. This product malfunction has the potential to result in delay in care and a possible re-exposure of the patient to radiation (x-ray). | Software design | Merge sent an Urgent Medical Device Recall letter dated October 19, 2016,(Recall #2016-075) to all affected customers. Initially, before it was considered a recall, the firm provided customers with a software upgrade to fix the issue. The letter described the Product, Issue, Potential Harm, Containment by the Customer/User, Action by Merge and Actions by Customers. Requested consignees to complete and return the response form. For questions email recall@merge.com. Customers with questions were instructed to send an email to recall@merge.com.For questions regarding this recall call 262-367-0700. |  | 29 | 29 | Nationwide Distribution including AL, CA, FL,IN, MI, MO, NY, PA, TN, and WI. | Open, Classified |  |  | 29 | Computer
Z-0633-2017 | 75367 | No | OJV | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Waste Management Kit. Catalog Number K10-04381AP |  |  | September 23, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for various Merit Medical Procedure Kits due to a compromised sterile barrier. | Process control | Consignees notified via letter on 09/23/16 to quarantine any devices and discontinue use, ensure all personnel to whom the devices were distributed are made aware of the field action, and return affected devices to Merit Medical. |  | 30 | 30 units | Nationwide to VA, UT, and NC. Other foreign consignee Japan. No Canadian and VA/govt/military consignees. | Terminated | March 27, 2017 | 185 | 431 | Not_Computer
Z-0636-2017 | 75306 | No | FDF | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pentax Video Colonoscope is intended to be used with a PENTAX video processor (including light source), documentation equipment, Monitor, EndoTherapy Device such as a Biopsy Forceps, and other ancillary equipment for endoscopy and endoscopic surgery within the lower digestive tract including the anus, rectum, sigmoid colon, colon and ileocecal valve. |  |  | August 15, 2016 | 2016 | Pentax of America Inc | Pentax is initiating the product correction of several model numbers of Pentax's Colonoscope that were shipped between the clearance date of April 2014 and September 2015 because they contain the operation and reprocessing IFU's that were in use prior to the 2014 clearances and therefore did not include the cleared operation IFU, the updated reprocessing IFU or the dedicated reprocessing adaptor (for applicable models). | Other | Pentax sent as "Urgent Field Correction" and Field Correction Response Form dated August 15, 2016 to their affected customers. |  | 5090 | 5,090 (US) and 14,209 (OUS) | US Nationwide and Japan | Open, Classified |  |  | 5090 | Not_Computer
Z-0627-2017 | 75454 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | DePuy Spine SKYLINE ANTERIOR CERVICAL PLATE SYSTEM VARIABLE SELF-TAPPING SCREW 4.0 x 26mm |  |  | October 12, 2016 | 2016 | DePuy Spine, Inc. | SKYLINE Variable Ti screws may have contained SKYLINE Constrained Ti screws within the packaging. | Under Investigation by firm | The firm, DePuy Synthes Spine, sent an "URGENT! VOLUNTARY PRODUCT CORRECTION NOTIFICATION" letter dated 10/12/2016 to its Customers. Sales consultants supporting these customers received an email notification with required reply forms. The letter described the product, problem and actions to be taken. A DePuy Synthes Spine Sales consultant has been notified and will be reviewing inventory to ensure correct screw placement within your set. The customers were instructed to review, complete, sign and return the attached business reply form to DePuy Synthes Spine via email to: DPYUS-SpineFieldActions@its.jnj.com within 5 business days of receipt of this notification; forward the notice to anyone in your facility that needs to be informed; if devices have been forwarded to another facility, contact that facility and provide them with notice, and keep a copy of the notice.Questions about this recall, please contact Recall Coordinator, DePuy Synthes Spine (via telephone: 508.828.3647 or via email: DPYUS- SpineFieldActions@its.jnj.com). |  | 6 | 6 units | US Distribution to states of: CA, ID, IL, IN, KS, KY, LA, MO, NC, OR, PA, TX, VA, WA, and WV. | Open, Classified |  |  | 72 | Not_Computer
Z-0637-2017 | 75502 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 2.5 mm Reaming Rod with Ball Tip &amp; Extension/950 mm, Sterile |  |  | October 03, 2016 | 2016 | Synthes (USA) Products LLC | The peel pouches for the affected reaming rods and extraction hooks are delaminating. | Nonconforming Material/Component | DePuy Synthes mailed a Customer Safety Advisory Notice on October 19, 2016 to affected customers advising them of the malfunction, potential risks and the actions needed to be taken to avoid injury, as they work on a solution to correct the problem. |  | 1951 | 1,951 units | Distributed Nationwide | Terminated | July 06, 2017 | 276 | 11305 | Not_Computer
Z-0650-2017 | 75540 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ACTIS Flex Reamers SZ 0/1; 2/3; 4/5; 6/7; 8/9; 10/11; 12, Product is sold non-sterile, these instruments are packed non-sterile, within a rigid tube with foam protectors, with label and IFU |  |  | October 25, 2016 | 2016 | Depuy Orthopaedics Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary recall for all lots of the ACTIS Flex Reamers due to the instruments breaking and potentially leaving pieces in the patient. | Under Investigation by firm | Depuy Orthopaedics initiated a recall on October 25, 2016. Notices were sent to the US Distributors via email on October 25, 2016. Customers were instructed to do the following: 1. Immediately Inspect Inventory, 2. Immediately Return US Distributor Inventory: If any affected instruments are found in a US Distributors inventory, return to:DePuy Synthes Joint Reconstruction, a division of DePuy Orthopaedics, Inc.3. To expedite the return and credit process: Be sure to reference H16-14 on all return paperwork and/or online return forms and on the outside of the box when returning recalled lots of the instrument.4. Upon receipt of the affected instruments, affected instruments will be credited against the US Distributors B&R; budget.5. Within 5-Business Days: Complete the Distributor Card or for the Reconciliation Form: Within 5-business days of initiation, the completed Reconciliation Form should be returned to DePuy Orthopaedics, Inc.6. Retain copies of all field action documents in customer filesFor further questions please call (574) 371-4917. |  | 137 | 137 | US Distribution to the states of: CA, KY, LA, MD, ME, NC, NH, NY, TX, VA, WAForeign: Austria, Ireland, and JapanVA/DOD: | Terminated | June 06, 2017 | 224 | 137 | Not_Computer
Z-0649-2017 | 75011 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Custom Fluid Management Set, Sterile, EO, Rx Only |  |  | August 12, 2016 | 2016 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for Fluid Administration Sets due to a molding defect in the drip chamber. | Component design/selection | Consignees were notified via an Urgent Product Recall Notice on 8/12/16 to: 1. Please immediately determine if any of the devices are within your facility, quarantine them, and discontinue use. 2. Ensure that all personnel to whom the devices were distributed are made aware of this field action. 3. Complete the Customer Response Form and return it to Merit Medical Systems with any affected products. If you have no affected products, please complete the Customer Response Form. 4. E-mail the completed Customer Response Form to Customer Service at RESPONSE@merit.com, or fax to (801) 316-4880. 5. If you have any questions concerning this communication or for assistance arranging product return, please dont hesitate to contact your Merit Sales Representative or Vicki Biehn at vbiehn@merit.com or at (801) 208-4260. |  | 107341 | 107,341 total devices | Nationwide Distribution. | Terminated | May 09, 2017 | 270 | 107341 | Not_Computer
Z-0648-2017 | 75503 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS Lyse Reagent Kit, Catalog No. B23538 |  |  | October 19, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter is recalling the AQUIOS Lyse Reagent Kit because it may lead to an increased generation of notifications and flags when used on the AQUIOS CL when running the AQUIOS Tetra application. The issue could lead to an increase in the number of samples requiring review and re-runs with a potential delay to the reporting of results. Depending on the specimen age, a new blood specimen may be required. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated October 19, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the notice are instructed to contact our Customer Support Centerhttp://www.beckmancoulter.com/customersupport/support; 1-800-369-0333 in the US and Canada; outside the US and Canada contact local Beckman Coulter Representative. |  | 2367 | 2367 units total (207 units in US) | Worldwide Distribution - US Nationwide and the countries of Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Italy, Malaysia, Netherlands, Poland, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, United Kingdom and Zambia | Terminated | February 24, 2017 | 128 | 2367 | Not_Computer
Z-0645-2017 | 75478 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LIAISON Estradiol II Gen, 310680, 310680-CN |  |  | July 13, 2016 | 2016 | Diasorin Inc. | DiaSorin is issuing an urgent medical device correction for the LIAISONÂż Estradiol II Gen assay. Due to the risk of a recently identified cross reactivity, the LIAISONÂż Estradiol II Gena assay should not be used to test patients being treated with fulvestrant (Faslodex). | Device Design | DiaSorin Inc. sent an Customer Notification letter dated July 13, 2016, to all affected customers. The letter described the problem and the product involved in the recall. The letter listed the "Recommended Actions" and requested customers to complete and return the response portion. For questions contact Product Support at 1-800-328-1482. |  | 10739 | 10,739 | USA (nationwide) Distribution and Worldwide Distribution including the states of : AR, AZ, CA, FL, ID, NJ, NV, NY and PA.., and to the countries of : Saudi Arabia, Colombia, Mauritius, Iraq, Mongolia, Greece, Albania, Austria, Bosnia, Bosnia Herzegovina, Belarus, Bulgaria, Macedonia, Poland, Serbia, Dominican Rep., Tunisia, Paraguay, Hungary, Jordan, Egypt, Belgium, Caribbean, Germany, Brazil, Mexico, Italy, India, South Korea, Cyprus, Bangladesh, Morocco, Chile, Bahrain, Palestine, Utd.Arab.Emir., Kosovo, Uruguay, Venezuela, Hungary, Australia, Vietnam, Sri Lanka, Turkey, Vietnam, Romania, Mauritius, Ecuador, Lebanon, Nepal, Argentina, Hong Kong, China and Canada. | Open, Classified |  |  | 10739 | Not_Computer
Z-0646-2017 | 75511 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Pathfast D-Dimer, Reference number: PF1051-KUS |  |  | August 16, 2016 | 2016 | Lsi Medience Corporation | Two complaints were received about decrease in sensitivity (control value and/or calibration value) of the product. | Nonconforming Material/Component | Customers were notified by letter sent via email on/about August 1, 2016.They were instructed to stop using and destroy the affected kits. They were to perform a retrospective analysis of the patients tested on the batch and investigate further. A Destruction Certificate was attached to be completed and returned to Verbatim Americas, LLC.For further questions please call (704) 547-6517. |  | 2520 | 42 boxes (each containing 60 test cartridges/one kit contains 10 reagent cartridge trays and one tray has 6 cartridges). | US Distribution to the states of : AL, FL, ID and MA. | Open, Classified |  |  | 2520 | Not_Computer
Z-0651-2017 | 75324 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | PeroxiClear Peroxide Solution, (3oz. and 12 oz., 12 oz. Twin Pack, 2x12 oz. Twin Pack for Target stores, 12 oz. for Walmart Vision Center and 3 oz. Starter Kit US), Manufactured by Bausch + Lomb, Rochester, NY 14609. |  |  | September 19, 2016 | 2016 | Bausch &amp; Lomb Inc Irb | Product was found to contain excess residual levels of hydrogen peroxide. | Under Investigation by firm | The firm, Bausch &amp; Lomb, sent an "URGENT MEDICAL DEVICE VOLUNTARY RECALL" letter dated 9/29/2016 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter: review your inventory including floor displays; contact Stericycle, the firm conducting the product recall on behalf of Bausch &amp; Lomb, to arrange for return of the product; and complete and return the Recall Acknowledgement Form to Stericycle Inc,, even if you do not have any of the product in your inventory, via Fax to 1-877-779-7814, email: BauschandLomb5833@Stericycle.com or phone: 1-877-877-0481. If you have any questions about this voluntary recall process or would like additional information, please contact Stericycle, Inc., the firm conducting the recall on behalf of Bausch &amp; Lomb at 1-877-877-0481. |  | 4652822 | 4,652,822 units | Worldwide Distribution-U.S. Nationwide and country of: Canada. | Open, Classified |  |  | 4652822 | Not_Computer
Z-0375-2017 | 75551 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare System X-ray Tomography Computed : |  |  | March 15, 2016 | 2016 | Philips Electronics North America Corporation | It was discovered the patient support tabletop may fail to move in the horizontal direction when it is commanded to do so during CT scans due to a potential mechanical linkage failure. | Radiation Control for Health and Safety Act | Philips planned action: 1. Philips issued a customer notification letter on October 13, 2016. The letter includes a statement that the corrections will be made free of charge.2. Philips will initiate a field corrective action to address the issue.3. Philips will contact the customer and schedule an inspection to determine if the unit requires repair.4. If repair is required, the couch horizontal linkage assembly will be replaced.For further questions please call (978) 659-3000. |  | 1363 | 1363 | USA(nationwide) Distribution to the states of : AZ, NY, GA, UT, TN, TX, AR, NV, Il, FL, MS, LA, MD, MO, MA, OK, KY, WY, SC, ND, NE, DE, MI, WA, CT, RI, KS, OR, IN, PA, CA, NJ, NC, MN, AL and IA. | Open, Classified |  |  | 1363 | Not_Computer
Z-0653-2017 | 75373 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | COBRA Fusion 50 Ablation System, COBRA Fusion 150 Ablation System, and COBRA Fusion Magnetic Retriever System. Model Numbers: 700-001, 700-001S, 700-001MI, 700-002, and 700-003. |  |  | September 22, 2016 | 2016 | AtriCure, Inc. | Due to product design and process control, part of the device can separate and remain behind when the device is removed that can potentially cause permanent impairment or damage to body function or structure. | Component design/selection | An Urgent: COBRA Fusion Ablation System Advisory Notification Letter with Acknowledgement Form was sent to all users. The letter described the issue and identified the affected device. Customers were asked to examine their inventory and determine if any affected product was on hand. All users, including those that the device may have been further distributed should be notified with a copy of the letter. Customers are to contact AtriCure customer service at 1-866-349-2342, and select option 6 to return any affected devices. The Device Notification Acknowledgement Form should be completed and returned as soon as possible. |  | 5263 | 5,263 | Worldwide Distribution -- USA, in the following states AL, AR, AZ, CA, CO, CT, GA, FL, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, WA, WI, &amp; VA; and, to the countries of Albania, Australia, Belgium, France, Germany, Ireland, Italy, Norway, Netherlands, Poland, Sweden, Turkey, United Arab Emirates, &amp; United Kingdom. | Open, Classified |  |  | 5263 | Not_Computer
Z-0654-2017 | 75641 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | The Hemotherm 400CE 115v, containing power board p/n 39945 |  |  | November 07, 2016 | 2016 | Cincinnati Sub-Zero Products Inc | The device may not maintain water flow or temperature control. | Material/Component Contamination | Cincinnati Sub-Zero sent an Urgent Medical Device Correction letter date November 7, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm has notified all customers to discontinue use and contact CSZ. Upon receipt of the new board, immediately remove and discard the affected power board and replace it with the newly provided power board. The complete and return the customer response form. For further questions, please call (513) 772-8810 |  | 50 | 50 | US Distribution to the states of : CA, FL, GA, MA, MN, NC, NJ, OH, PA and SC. | Open, Classified |  |  | 50 | Not_Computer
Z-0652-2017 | 75484 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Chemistry 1 Calibrator |  |  | September 15, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens determined that discrepant low calcium recovery of QC and patient samples can occur with Dimension VistaÂż CHEM 1 CAL lots 5GM081 and 5GM082. | Under Investigation by firm | Siemens mailed an Urgent Medical Device recall notification letter on September 19, 2016 to notify affected customers who purchased the product and to make them aware of the potential risk to health and the actions to be taken. For questions regarding tis recall. For questions regarding this recall call 1-800-441-9250. |  | 9935 | 9,935 units | Worldwide Distribution - US (Nationwide) Internationally to Australia, Austria, Bahamas, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Italy, Japan, Netherlands, New Zealand, Norway, Portugal, Qatar, Saudi Arabia, South Korea, Slovakia, Slovenia, Spain, and Switzerland. | Open, Classified |  |  | 9935 | Not_Computer
Z-0655-2017 | 75558 | No | HHM | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | SURESOUND22US Multipack (US Distribution) containing 2 packs of six devices. |  |  | October 28, 2016 | 2016 | Hologic, Inc | Breach in sterile barrier compromising product sterility | Package design/selection | Hologic notified customers in the United States through Certified letter on 10/28/16. The letter contains instructions with the option of returning or destroying the product, provides contact information, and includes a postage paid return card for the consignee to complete and return to Hologic.Questions or require clarification of the response process please contact 1- 800-442-9892 and selection option 6 and then option 2 for surgical product technical support.For questions regarding this recall call 508-263-8920. |  | 600 | 600 units (US) ; 24 units (OUS) | Worldwide Distribution - US (Nationwide) and Internationally to Finland and Austria. | Terminated | August 18, 2017 | 294 | 600 | Not_Computer
Z-0656-2017 | 75580 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Flexible Shaft, Orthopedic manual surgical instrument |  |  | October 31, 2016 | 2016 | Synthes (USA) Products LLC | Non-conforming material used. | Under Investigation by firm | The firm, DePuy Synthes, mailed an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 11/2016 to customers who were affected by the recall. The letter described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory to identify and quarantine all affected products; if you have product to return call DePuy Synthes at 1-800-479-6329 to obtain an RA; complete and return the Verification Section with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132 within 30 business days; send a copy of the completed Verification section via fax to 888-257-7905 or scan/email: Synthes3503@stericycle.com, even if you have no product to return; forward notice to anyone in your facility that needs to be informed; if any of the affected product has been forwarded to another facility, contact that facility to arrange return and keep a copy of the notice.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 198 | 198 units | US Nationwide Distribution. | Terminated | June 12, 2017 | 224 | 198 | Not_Computer
Z-0660-2017 | 75664 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System; Used to make treatment plans for patients with prescriptions for external beam radiation therapy. |  |  | November 03, 2016 | 2016 | Elekta, Inc. | Incorrect Dose when using the reset function. | Software design | Important Field Safety Notice 382-01-MON-006 was sent to all affected customers on 11/3/2016. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. An acknowledgement form was included which is to be returned to Elekta. |  | 671 | 671 | AK, AZ, CA, CO, FL, GA, IL, IN, MA, MI, MN, MO, NV, NH, NY, NC, OH, OK, OR, PA, SD, TX, UT, VA, WA, WI, Algeria, Australia, Austria, Brazil, Canada, China, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Laos, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Venezuela, | Open, Classified |  |  | 671 | Not_Computer
Z-0658-2017 | 75553 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConforMIS iTotal Cruciate Retaining (CR) Knee Replacement System, Sterile |  |  | October 18, 2016 | 2016 | ConforMIS, Inc. | Mislabeled: incorrect serial number on the F4 A/P cut guide iJig. | Labeling Change Control | ConforMIS notified the sales representative on 10/18/16 after it became aware of the device mix-up. The sales representative was requested to return any remaining components from iTotal CR kit S/N: 0381042. |  | 2 | 2 units | Distributed to the states of CO and NV. | Completed |  |  | 2 | Not_Computer
Z-0663-2017 | 75617 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Bridge Occlusion Balloon. Model: 590-001 |  |  | October 31, 2016 | 2016 | Spectranetics Corporation | Spectranetics Corporation announces a voluntary field action for the Turbo Elite Atherectomy Catheter, ELCA Coronary Atherectomy Catheter, and Bridge Occlusion Balloon due to potentially compromised integrity of the outer sterile packaging. | Packaging | Spectranectics sent an Urgent Medical Device Recall letter dated October 31, 2016, to all affected consignees. Consignees were instructed to discontinue use of the potentially impacted lots and return the product for replacement. Consignees were also instructed to complete the attached Acknowledgement and Receipt Form to facilitate the product return of any remaining product in inventory. Customers with questions were instructed to call 1-800-231-0978, Option 2. For questions regarding this recall call 719-651-8517. |  | 873 | 873 units | Worldwide Distribution - US (nationwide and PR) Internationally to Austria, Belgium, Cyprus, Denmark, France, Germany, Hungary, Italy, Netherlands, Norway, Poland, Spain, Sweden, Switzerland, United Kingdom, Australia, Bahrain, Brazil, Canada, Chile, Colombia, Dominican Republic, Israel, Japan, Kuwait, Taiwan, Thailand. | Terminated | April 18, 2017 | 169 | 37276 | Not_Computer
Z-0668-2017 | 75718 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco TRP System |  |  | November 17, 2016 | 2016 | Elekta, Inc. | Contours are not handled correctly if the CT dataset contains slices with DICOM Z coordinates of xx.x50mm. | Software design | Elekta sent an Important Field Safety Notice 382-01-MON-007 was sent to all affected customers on November 17, 2016, to all affected customers. The notice informs users of the specific product version numbers affected by the issue, and any work arounds that can be used to avoid the issue. The notice included a reply form which is to be returned. |  | 2148 | 2148 | Worldwide Distribution - US including AK, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MD, MA, MI, MN, MO, MT, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, SD, TX, UT, VA, WA, WI, Puerto Rico and Internationally to Algeria, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Bahrain, Belarus, Botswana, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Laos, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Tunisia, Turkey, Unite Kingdom, Ukraine, Venezuela, Vietnam, and Zimbabwe. | Open, Classified |  |  | 2148 | Not_Computer
Z-0666-2017 | 75336 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-MR1 Ventilator: Catalog# 161010 |  |  | September 26, 2016 | 2016 | Hamilton Medical, Inc. | Oxygen tubing and the oxygen connector of the Hamilton-MR1 could potentially become loose during the preparation for ventilation. | Device Design | Hamilton Medical sent a letter dated September 6, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Hamilton Medical will update each affected Hamilton-MR1 ventilators by installing a new oxygen mixer mounting plate with the correct screws at no cost to each facility. A Hamilton Medical Field Service Technician will contact customers soon to schedule a convenient time to install the hardware upgrade. Customers may also call 1-800-426-6331, option #2 to be issued an RGA for return to the service center for the upgrade kit to be installed. Customers with questions should call 817-909-0308. |  | 44 | 44 units | Nationwide Distribution | Open, Classified |  |  | 44 | Not_Computer
Z-0667-2017 | 75639 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Z6Ms transesophageal transducers used on ACUSON SC2000 ultrasound systems, Z6Ms model 10436113. |  |  | November 07, 2016 | 2016 | Siemens Medical Solutions USA, Inc. | Firm received reports about "acquisition errors" and the deterioration of material covering the articulating section of the Z6Ms transesophageal transducer. | Device Design | Siemens sent a Customer Safety Advisory Notification letter to all affected customers on November 7, 2016 by certified mail advising them of the issue and identified steps that can be taken to avoid potential risk of this issue. This letter has been released first in English while translations are being completed. Once the translations are completecustomers in the countries who require translations will receive the letter.Prior to this mailing, ECO #641310 was released on July 5, 2016 and implemented into forward production on July 28, 2016 (SN 63028009) to correct the susceptibility to Gastro-flex circuit corrosion and electro migration. The ECO changed standard copper traces and traditionally applied soldermask on the circuit to fully gold plated traces with a dry film solder mask.ECO #645821 was released on September 8, 2016 date and implemented into forward production on September 8, 2016 (SN 63530013) to correct susceptibility to articulation sleeve damage. The ECO changed the articulation sleeve material from C-Flex to Viton.For questions regarding this recall call 650-694-5398. |  | 33 | 33 | Worldwide Distribution - US (nationwide) Internationally to AUSTRIA, GERMANY, SOUTH AFRICA, NETHERLANDS, MAURITIUS, MEXICO, and ITALY,JAPAN, SWEDEN | Terminated | May 26, 2017 | 200 | 33 | Not_Computer
Z-0664-2017 | 75334 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Merge FlexConnect software, a component of Merge LIS. |  |  | January 30, 2016 | 2016 | Merge Healthcare, Inc. | Communication protocols interfacing with the affected software version with some select instruments were not properly handled, preventing data captured by these instructions to be processed by the rest of the LIS software. | Software design | The recalling firm issued a letter dated 1/27/2016 via e-mail on 1/30/2016. Nonresponders were issued a second letter dated 5/25/2016 via email on 5/26/2016. |  | 324 | 324 sites potentially have the affected version | Nationwide Distribution | Open, Classified |  |  | 324 | Computer
Z-0665-2017 | 75564 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. |  |  | April 30, 2014 | 2014 | Merge Healthcare, Inc. | The application may crash during the cath lab procedure. | Software design | Merge sent an Important Alert letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact Merge Support at 866-261-5761 or support@merge.com. For questions regarding this recall call 877-741-5369. |  | 192 | 192 sites potentially have the affected software | Nationwide Distribution | Open, Classified |  |  | 192 | Not_Computer
Z-0673-2017 | 75569 | No | CDS | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MSS CASSETTE GLU/LAC/UREA (BUN) as a part of the |  |  | October 31, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Incorrect Urea results may be generated sporadically on cobas b 221&lt;6&gt;Roche OMNI S6 system and cannot be detected by QC. QC measurement results can be affected as well. This issue is only detectable if comparison result of a second measurement is available. The cause for this issue is an error in the software measurement procedure.In children, decreased urea results can point at inborn disorders or urea cycle. Therefore, false normal urea values could lead to delayed/incorrect diagnosis of such a disorder.Besides the patients most at risk, the risk for the overall patient population with falsely low or high values is considered remote. No severe adverse health effects are expected; but, unnecessary diagnostic measures may be initiated.Falsely normal results may result in further necessary examination not being carried out, with the possibility of a delay in the diagnosis and treatment. A medical risk cannot entirely be excluded. | Software design | Roche sent an Urgent Medical Correction letter dated October 31, 2016, to all affected customers. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Actions Required" Use the workaround outlined in this Urgent Medical Device Correction (UMDC) for Ureameasurement until the updated software version is available." If your facility has distributed the affected product to another site, please ensure this UMDC isprovided to that site." Complete the attached fax form and fax it to 1-877-766-7471." File this UMDC for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272 if you have questions about the information contained in this UMDC. |  | 97 | 97 | Nationwide Distribution including ALCA CO CT FL IL IN KS MD MI MO NC NE NJ NM NY OH OK PA PR SC TX VA and WI | Open, Classified |  |  | 97 | Computer
Z-0669-2017 | 75370 | Yes | NIP |  |  | Cardiovascular | PMA | MISAGO RX Self Expanding Peripheral Stem |  |  | August 04, 2016 | 2016 | Terumo Medical Corp | During recent routine testing, the MISAGO Peripheral Stent was found not to conform to product specification related to stent diameter at the proximal/distal ends and/or the shape of the stent. | Device Design | Terumo Medical Corporation issued an Urgent Product Recall Bulletin dated August 4, 2016, via Fed Ex.Customers must immediately discontinue use and quarantine any available stock of the Misago peripheral stent to prevent further use. The customers should complete the enclosed Inventory Confirmation form as soon as possible and email or fax the completed formto misago.rc@terumomedical.com or (734) 330-2034, Attn: MISAGO Recall.Customers are encourage you to contact Terumo Medical Corporation with any questions or concerns-Terumo Medical Corporation Quality Department: 1.800.283.7866, Press 4 for list of departments, Press 7 for Quality Assurance; Monday  Friday, 8 a.m.  5 p.m. ET- Email Quality Assurance: misago.rc@terumomedical.com- Customer Care: 1.800.888.3786- Contact your local Terumo representative |  | 7965 | 7,965 units | Nationwide | Open, Classified |  |  | 7965 | Not_Computer
Z-0671-2017 | 75311 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter SIGMA Spectrum INFUSION PUMP with Master Drug Library, V8 Spectrum Infusion Pump, Product Code 35700BAX2, Unique Device Identifier GTIN 00085412091570, Rx ONLY -- Baxter Healthcare Corporation, Medina, NY 14103 --- NOTE: The speakers are a part of the rear case assembly located in the back of the infusion pump. The speaker is intended to generate audible tones during each keystroke, and during alarm conditions. However, some speakers have a variation in their internal wiring that increases the potential for the speaker to fail prematurely. Spectrum Infusion Pumps with non-functional speakers will continue to display visual alarm indications, and will otherwise continue to operate as intended. |  |  | September 28, 2016 | 2016 | Baxter Healthcare Corp. | A correction was issued for the SIGMA Spectrum Infusion Pump with Master Drug Library in order to replace speakers installed in pumps or included in spare part rear case assemblies that have an increased potential to fail prematurely. | Process control | Baxter sent an Urgent Medical Device Correction letter dated September 28, 2016, to all affected customers. On 10/27/2016, Baxter Healthcare sent a f/u UMDC Letter (dated 10/27/2016) to customers. The f/u letter contains the device correction plan which states that Baxter will be replacing at no charge the rear case assembly on all affected Sigma Spectrum Infusion Pumps, as well as affected stand-alone rear case assemblies shipped to customers as spare parts.For questions, please contact your Baxter sales representative, or Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday, between 7:00 am and 7:00 pm ET. |  | 15517 | 15,517 units (US: 9560; Canada: 5957) | Nationwide Distribution, Puerto Rico and Canada | Open, Classified |  |  | 21932 | Not_Computer
Z-0367-2017 | 75043 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Greatbatch, Offset Cup Impactor with the following Parts and Catalogue #: |  |  | August 02, 2016 | 2016 | Greatbatch Medical | On December 17, 2015, Greatbatch discovered that the Standard Offset Cup Impactor with a POM-C handle grip failed sterility testing when housed in a dedicated instrument case using the sterilization parameters identified in the Standard Offset Cup Impactors Instructions for Use (IFU). | Under Investigation by firm | Greatbatch sent an Urgent Medical Device letter dated August 2, 2016, via FedEx to all affected consignees. The letter described the Reason for Field Correction Action, Risk to Health, and Instructions to Customers. Advised consignees to" immediately inform those responsible for sterilization of the devices to utilize the updated sterilization instructions, and post the attached notice where it is clearly visible to personnel who perform sterilization activities". Also instructed consignees to notify their customers and to complete and return the Field Correction Action Response Form. For questions contact your local Greatbatch Medical representative or 1-763-951-8235 or e-mail FieldActionCenter @Greatbatch.com. |  | 4478 | 4478 (2906 US, 1572 OUS) | Worldwide Distribution - US including CA, IN, KS,NJ, NC, TN, TX, and UT and Internationally to ITALY, JAPAN, FRANCE, SPAIN, GERMANY, NETHERLANDS, GREECE, LUXEMBOURG, SOUTH KOREA, IRELAND, UNITED KINGDOM, AUSTRALIA, AUSTRIA, BELGIUM, SWITZERLAND, NORWAY, SWEDEN, and TAIWAN. | Terminated | July 27, 2017 | 359 | 4478 | Not_Computer
Z-0676-2017 | 75637 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ARROW OnControl Bone Marrow Biopsy System Comprehensive Tray, Sterile, The Arrow OnControl Bone Marrow Aspiration System is intended for bone marrow aspiration of the iliac crest of adult and pediatric patients. The ArrowÂż OnControl Bone Marrow Biopsy System is intended for bone marrow core biopsy of the anterior or posterior iliac crest of adult patients |  |  | November 08, 2016 | 2016 | Teleflex Medical | Sterility: Due to a potential incomplete seal on the outer sterile package. | Device Design | Teleflex sent an Urgent Medical Device Recall Notification letter dated November 8, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product, complete the enclosed Recall Acknowledgement Form and fax to 1-855-419-8507, Attn: Customer Service or email to recalls@teleflex.com. Customers with questions were instructed to call their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 919-433-4940. |  |  |  | Worldwide Distribution - US (nationwide) Internationally to Brazil, Canada, Ecuador, Guatemala, Mexico and Peru. | Open, Classified |  |  |  | Not_Computer
Z-0674-2017 | 75589 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Adiprep Adipose Transfer System. Product Codes: 51431; 51432; 51433; 514001-347. |  |  | November 03, 2016 | 2016 | Terumo BCT, Inc. | Terumo BCT announces a voluntary field action for the Harvest : Adiprep Adipose Transfer System because the affected product is incorrectly labeled and has the potential for operator confusion. | Error in labeling | Consignees will be notified on 11/07/16 via letter to contact the North America Terumo BCT support center at +1.877.842.7837 or +1.303.542.5050 if they have any affected product to issue a Return Goods Authorization for product return. |  | 660 | 660 | Nationwide, Canadian, and other foreign consignees. No VA/govt/military consignees. | Open, Classified |  |  | 660 | Not_Computer
Z-0683-2017 | 75063 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Lineage Liners: |  |  | June 21, 2016 | 2016 | MicroPort Orthopedics Inc. | Revision rate trends from within the MicroPort Orthopedics Inc. database evaluated for the modular head Metal on Metal THA systems show an increasing overall trend from 2009 to present and it was found that there was a specific hazard/harm for suspected tissue reaction to metal debris over one percent for modular head CONSERVE(R) systems and DYNASTY(R)metal acetabular liner systems. | Device Design | MicroPort sent an Urgent Field Safety Notice dated June 22, 2016, to all affected customers notifying them of the change in the instructions for use on July 7th, 2016 by email. The firm also notified US hospitals where the implantation surgeries occurred within the last twelve (12) months by letter on July 13th, 2016 delivered by FedEx Priority. Customers were asked to complete and return the enclosed form even if they do not have any affected product to return by fax to 1-901-451-6032 or PostMarket@ortho.microport.com.For questions regarding this recall call 901-867-4771. |  | 5009 | 5,009 units | Worldwide Distribution - US (nationwide) Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Cyprus, Czech Republic, Denmark, Finland, France, Georgia, Germany, Greece, Iran, Italy, Jamaica, Japan, Lithuania, Malaysia, Philippines, Poland, Portugal, Russia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Netherlands, Turkey, Arab Emirates, United Kingdom, and Vietnam. | Terminated | July 01, 2017 | 375 | 513126 | Not_Computer
Z-0690-2017 | 75627 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FMsealer Open Shears. Model FM3001 |  |  | October 24, 2016 | 2016 | Domain Surgical, Inc. | Domain Surgical, Inc. announces a voluntary field action for the FMsealer Open Shears due to fixed jaw fractures during clincal usage. | Other | Consignees were notified on 10/24/2016 via letter to return affected products to Domain Surgical, Inc. |  | 245 | 70 units US; 175 units OUS | Nationwide and other foreign consignees. No Canadian or VA/govt/military consignees. | Terminated | April 11, 2017 | 169 | 245 | Not_Computer
Z-0682-2017 | 75573 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | STÂżCKERT HeaterCooler System 3T; Item Number 160281, HeaterCooler 3T, 240V/60Hz; Item Number 160282, HeaterCooler 3T, 208V/60Hz; Item Number 160285, HeaterCooler 3T, 120V/ 60Hz |  |  | August 01, 2014 | 2014 | Sorin Group USA, Inc. | On August 1, 2014 Sorin Group USA, Inc. informed consignees via letter to review the disinfection and monitoring practices and to strictly adhere to the instructions in the Instructions for Use ( IFU ). | Material/Component Contamination | Sorin Group sent an Important Information letter dated August 1, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by then consignee. Consignees were instructed to review the disinfection and monitoring practices and to strictly adhere to the instructions in the Instructions for Use ( IFU ). Consignees with questions were instructed to call 303-467-6527 or 303-467-6503.For questions regarding this recall call 281-228-7260. |  | 208722 | 1,618 units in the US; 129 units outside of US | Worldwide Distribution - US (Nationwide and Puerto Rico) and Internationally to Canada, Colombia, Ecuador, El Salvador, Germany, Honduras, Mexico, Saudi Arabia, Trinidad, and Tobago. | Terminated | April 20, 2017 | 993 | 208722 | Not_Computer
Z-0681-2017 | 75694 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes 0.76 mm Drill Bit/Stryker J-Latch/with 5mm Stop/44.5 mm |  |  | November 01, 2016 | 2016 | Synthes (USA) Products LLC | Package may be incorrectly labeled as containing 0.76 mm Drill Bit, but in fact contain a 1.1 mm diameter Drill Bit. | Under Investigation by firm | DePuy Synthes sent an Urgent Notice to customers affected by the recall to inform them of the potential impacts, and actions to be taken for the return of the identified units. |  | 84 | 84 units | Distributed Nationwide | Terminated | June 01, 2017 | 212 | 84 | Not_Computer
Z-0647-2017 | 75536 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Med Single Lumen Catheters sold in Centurion convenience kits. |  |  | October 21, 2016 | 2016 | Centurion Medical Products Corporation | The kits contain Multi-Med Single Lumen Catheters thathave a potential for excess material to remain at the tip of the catheter from the manufacturing process. Itis possible that the excess material may separate from the catheter during use and pose a potential riskof embolism, which could result in serious injury or death. | Process control | Centurion sent an Urgent Recall Notice dated October 21, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of this product immediately, and forward a copy of this notice to other users to whom this product may have been further distributed. In addition, customers were instructed to complete the attached form at their earliest convenience and fax to 517-546-3356.For questions regarding this recall call 517-546-5400. |  | 1000 | 1000 kits | Nationwide Distribution | Terminated | July 20, 2017 | 272 | 1000 | Not_Computer
Z-0692-2017 | 75594 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force, System x-ray, tomography, computed |  |  | October 20, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is providing software update versionVA50A_SP3 to address the software bugs thatwere identified through normal field monitoringand the Global Complaint Handling Process.Correction for the problems are as follows:1. Correction to volumetricmisrepresentations of high contrastobjects when using ADMIRE.2. Correction to highly sporadic scanaborts due to temporarily tubecurrents at 0mA.3. Correction to missing MicrosoftHotfixes (MS16-001, MS15-135,MS15-088, MS15-048).4. Correction to highly sporadic imageartifacts caused by error signals fromhigh voltage chain at extremely hightube currents.5. Correction to highly sporadicincomplete image reconstruction dueto temporary slow hard disk writespeed.6. Correction to highly sporadic scanaborts due to temporarily high diskload of system disk.7. Correction to highly sporadic imageartifacts generated by scattercorrection in case of tube arcings. | Under Investigation by firm | Siemens sent a Customer Information Letter dated November 2, 2016, to customers affected by the recall to inform them of a release of a software update that will provide bug-fixes to improve system performance. Customers will be contacted to arrange a date for onsite installation. For further questions, please call (610) 448-6471. |  | 48 | 48 units distributed in U.S. | Nationwide Distribution. | Terminated | May 30, 2017 | 222 | 48 | Computer
Z-0696-2017 | 75756 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Core 128 728323 |  |  | November 12, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Software error due to the filament on timer | Software design | Philips sent a Customer Information Letter dated November 14, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377). |  | 509 | 509 | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and Internationally to Algeria, Argentina, Australia, Austria, Bangladeshm Belgium, Bolivia, Brazil, Cambodia, Canada, Chile, China, Colombia, Croatia, Cuba, Czech Republic, Denmark, Djbouti, Dominican Republic, Egypt, Estonia, Finland, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic of Korea, Latvia, Lebanon, Lithuania, Malaysia, Martinique, Mexico, Monaco, Mongolia, Mozambique, Myanmar, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestinian territory, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Rwanda, Saudi Aarabia, Senegal, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Vietnam. | Open, Classified |  |  | 1419 | Computer
Z-0698-2017 | 75632 | No | PAX |  |  | Neurology | HDE - Humanitarian Device Exemption | Elana Surgical KitHUD |  |  | November 09, 2016 | 2016 | Elana, Inc. | Elana has issued this Field Action to voluntarily remove specific serial numbers of our Elana Surgical KitHUD due to potentially compromised integrity of the outer sterile packaging of the Elana Catheter 2.0. | Packaging | Elana sent an Urgent Medical Device Recall letter dated November 7, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected device and to return the product for replacement. Customers were also instructed to complete the attached Acknowledgement and Receipt Form to facilitate the product return and exchange. Customers with questions were instructed to call 301-215-7010. |  | 4 | 4 | Worldwide Distribution - US including NY and Internationally to Switzerland and the Netherlands. | Open, Classified |  |  | 4 | Not_Computer
Z-0702-2017 | 75630 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MANOS EX (CAT#TY-12-101) |  |  | November 04, 2016 | 2016 | Thayer Intellectual Property, Inc. | The firm failed to notify current customers of IFU update/warning. During a recent FDA inspection, it was discovered that Thayer Intellectual Property, Inc. had revised their IFU to include a warning not to bend, deform or otherwise modify any part of the device. The warning was added as a result of a MAUDE report filed in December 2014 where a piece of the MANOS EX broke during use and was left in the patient's hand. The patient suffered a tendon injury and an additional surgery was required to remove the broken piece of the device and repair the tendon. FDA inspection found that while the firm had updated their IFU, they had not notified customers who had been shipped the device with the older version that did not include the warning. Risk: It is important for MANOS EX users to be aware that bending, deforming, or modifying the device could result in device breakage which could impact patient safety and device performance. | Packaging change control | Thayer sent an Urgent Medical Device Correction letters dated November 4, 2016 by Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were asked to check if any of the affected lots were still available, and if so, to change out the IFU with version 5. For questions call 415-923-9375. |  | 339 | 339 devices | US Nationwide Distribution in the states of CA, TX. MD, MI,VA, DC, UT, CT, ID, SD, IN, and GA. | Open, Classified |  |  | 339 | Computer
Z-0699-2017 | 75444 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Aequalis Fx2 implant parts and instrument trays. |  |  | September 19, 2016 | 2016 | Tornier, Inc | Tornier is conducting a recall on Aequalis Fx2 ( implant parts and instrument trays) due to reports of dislocations of the poly insert and the stem. | Component design/selection | Wright sent an "Urgent Medical Device Recall" letter dated September 19, 2016. The letter was addressed to Healthcare Professional. The letter identified the affected product, problem and actions to be taken. It advised consignees to stop further distribution or use of product immediately. Requires consignees to notify all accounts that may be end-users, or contact Wright with your request for us to contact them. Advised consignee to complete and e-mail the attached questionnaire to FieldAction@wright.com. For questions contact the same. |  | 30 | 30 | US Distribution TX, CA, and MI | Open, Classified |  |  | 30 | Not_Computer
Z-0704-2017 | 75796 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | UltraSeal XT Plus Clear Refill. The package contains 4x 1.2 ml UltraSeal XT Plus. |  |  | November 17, 2016 | 2016 | Ultradent Products, Inc. | There was a labeling error with product, UltraSeal XT Plus Clear Refill, Part Number (PN) 565, Lot Number (Lot) BD79T. The package was mislabeled as PQ1, PN 641; while the barcode, IFU and products were labeled for UltraSeal XT Plus Clear Refill, PN 565, Lot BD79T. | Labeling mix-ups | Consignees notified via phone on 17Nov2016 and by letter on 22Nov2016 to identify and return products. Consignees with any questions regarding this field action can be sent to 801-553-4858 or 801-553-4220. |  | 31 | 31 | Worldwide Distribution - US Nationwide | Terminated | June 05, 2017 | 200 | 31 | Not_Computer
Z-0701-2017 | 75644 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 3M Universal Electrosurgical Pad, REF 9130 (100 pads / case) and |  |  | October 26, 2016 | 2016 | 3M Company - Health Care Business | 3M is recalling Universal Electrosurgical Pads because a report of an unintended material (process liner) in the product. This may prevent or impede the safe return of electrosurgical current following Electrosurgical Unit (ESU) activation. This defect has the potential to increase the risk of a patient burn directly under the pad or at an alternate site of the body. | Equipment maintenance | 3M sent an Urgent Medical Device Recall letter dated October 26, 2016 to affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to complete and return the attached Customer Product Return and Recall Form. Customers with questions can contact 3M Health Care Customer Helpline at 1-800-228-3957, 8:30 am - 7:00 pm EST. |  | 32000 | 32,000 pads | Worldwide Distribution - US Nationwide in the states of AR, AZ, CA, FL, GA, ID, IL, KY, NC, NY, TN, TX, VA and the countries of Canada, Chile, Costa Rica, Germany, Japan, Mexico and South Africa. | Terminated | August 01, 2017 | 279 | 32000 | Not_Computer
Z-0703-2017 | 75608 | No | NJL |  |  | Orthopedic | PMA | LCS Complete Knee Revision System VVC Insert |  |  | November 08, 2016 | 2016 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is voluntarily recalling one lot of the LCS Complete Knee Revision System VVC Insert Small 22.5mm, due to the insert being manufactured outside of specifications, which prevents the insert from properly seating in the mating tibial tray. | Nonconforming Material/Component | DePuy Orthopaedics, Inc. initiated a voluntary recall on 11/8/2016 via letter mailed to the customer. Only one device was distributed in the U.S., and this was removed from the hospital on 21-Oct-2016. This recall does not affect any other lots of the LCS Complete Knee Revision System VVC Insert Small 22.5mm devices. The recall notice provides instructions for notifying the medical facility that received, used, or purchased the affected lot of the LCS Complete Knee Revision System VVC Insert Small 22.5mm. The purpose of this device recall is to notify medical professionals of the possible effects of using the affected device and to communicate that there are no recalled devices distributed at this time. The customer was asked to do the following: Complete the Reconciliation Form and return to your DePuy Orthopaedics Sales Consultant or fax to 574-371-4939 or email to DPYUS-JointReconFieldActions@its.jnj.com within five (5) business-days of this notice." Retain a copy of the completed Reconciliation Form and this notice in your files." Forward this notice to others in your facility that need to be informed.Effectiveness will be determined by the return of one US Medical Facility Reconciliation Form.For questions about device recall information provided, please contact Kim Earle, Senior Recall Coordinator, at 574-371-4917 or kearle@its.jnj.com or DPYUS-JointReconFieldActions@its.jnj.com. |  | 1 | 1 | US Distribution in the state of TN | Terminated | June 05, 2017 | 209 | 1 | Not_Computer
Z-0700-2017 | 75458 | No | MHY |  |  | Neurology | PMA | Model 37751 Recharger |  |  | October 03, 2016 | 2016 | Medtronic Neuromodulation | Medtronic has identified an increased number of complaints from customers involving reports of Rechargers (Medtronic Model 37751 Recharger, which is included in Models 37754 and 97754 Charging Systems,Spinal Cord Stimulation)that are in an unresponsive error state, where the Recharger is non-functional with a blank display screen and is beeping every 5 seconds. Medtronic has determined all Rechargers manufactured starting in November 2014 (indicated by serial numbers beginning with NKA4 or NKU4 ) are more susceptible to this error state. | Software change control | Consignees were sent a series of documents which included an introductory Medtronic and "Urgent Medical Device Correction" letters dated September 2016. The urgent letters included one for Deep Brain Stimulation (DBS) and one for Spinal Cord Stimulation(SCS) . Each letter was addressed to Healthcare Professional. Only physicians associated with both SCS and DBS therapy received the cover letter. The letters provided information on "Background, Issues Mitigation, Recommendation, Additional Information and Patient Notification (DBS).Requested consignees to complete and return the Physician /HCP Reply Form.For questions contact Medtronic Technical Services at 1-800-707-0933 weekdays 7 am to 6pm, CT. |  | 42887 | 42,887 | Worldwide Distribution - US Nationwide including DCOUS: Canada and others. | Open, Classified |  |  | 42887 | Not_Computer
Z-0708-2017 | 75322 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Xcela Power Injectable PICC Maximal Barrier Nursing Kit |  |  | August 29, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | Medline Industries, Inc. supplies Navilyst with sterile package gloves included in the tandem packing of the Maximal Barrier Nursing Kits. Medline has determined that specific lots of the gloves supplied to Navilyst contain an incorrect expiration date. Navilyst has confirmed that Medlines affected sterile gloves (Medline Part Number MSG1075/MSG3075) have been included in a limited number of kits. | Under Investigation by firm | Navilyst sent an Urgent Voluntary Medical Device Recall notices, reply forms and a copy of Medline's Recall Notice to the their customers. The letter identified the product, the problem, and the action to be taken by the customer. Navilyst Medical, Inc. (an AngioDynamics Company), the manufacturer of the Xcela" PICC w/PASV, the BioFlo PICC w/ Endexo, and the BioFlo PICC w/ Endexo and PASV is conducting a medical device recall based on information received from Medline. Medline has determined that the sterile gloves listed in their Recall Notification dated August 3, 2016 may contain an incorrect expiration date. Actions to be taken: - Customers should immediately remove the affected product from inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location). - Segregate this product in a secure location for return to Navilyst Medical. - Immediately forward a copy of this recall notification to all sites to which customers may have further distributed affected product. Customers should also promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) via email to recall@angiodynamics.com. Customers with questions can call 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday  Friday: Eastern Standard Time).For questions regarding this recall 518-792-4112. |  | 75 | 75 packages (boxes of 5) | Nationwide Distribution including AZ, CA, FL, GA, HI, IL&lt; IN&lt; KS, LA, MI, MN, MO, ND, NM, OH, OR, PA, TN, TX, WA &amp; the District of Columbia. | Open, Classified |  |  | 428 | Not_Computer
Z-0707-2017 | 75270 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge, eFilm Workstation and eFilm Lite |  |  | October 20, 2016 | 2016 | Merge Healthcare, Inc. | A product issues happens for RF projection images. If user measures on RF image, .cal (Calibration) is not displayed and measurement is not correct. The .cal label is not displayed on the projection images after calibration is performed which should be present if the measurement is presented based on Pixel Spacing. | Software design | Merge Healthcare sent an Urgent Medical Device Recall letter dated October 12, 2016, to all affected consignees. The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Actions by Merge, and Actions by Customers. Advised consignees to discontinue using the product, and identify &amp; notify all customers. Requested consignees for a response to the notification. Customers with questions were instructed to send an email to recall@merge.com. |  | 2163 | 2163 | Worldwide Distribution - US: Nationwide including DC, GU, PR and Internationally to Canada, Argentina, Armenia, Australia, Austria, Belgium, Brazil, Chile, China, Colombia, Congo, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Guatemala, hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, New Caledonia, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Open, Classified |  |  | 2163 | Not_Computer
Z-0712-2017 | 75665 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | TENNIS ELBOW SUPPORT, PneuGel(TM), UNIVERSAL, REF 200A10, QTY: 1 Each, Rx Only, |  |  | November 07, 2016 | 2016 | DeRoyal Industries Inc | DeRoyal received complaints of mold on Tennis Elbow Straps. | Under Investigation by firm | DeRoyal issued their recall on 11/7/16 to the physician / clinician level by sending recall letters to its direct accounts and sales representatives. The foreign customer was notified by email on 11/7/2016. The customers were asked to identify the affected products, destroy them and return the provided notice of destruction form to DeRoyal for credit. |  | 56 | 56 units | US Nationwide Distribution | Open, Classified |  |  | 847 | Not_Computer
Z-0711-2017 | 75715 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr Knee System |  |  | November 15, 2016 | 2016 | Encore Medical, Lp | A size 4 Empowr baseplate was mislabeled as a size 6. The affected lot was laser marked as M6, marked with the lot number for the size 6 minus lot, and distributed as size 6 minus. | Mixed-up of materials/components | The recalling firm sent an Urgent Field Safety Notice letter dated November 15, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 760-734-3551. |  | 7 | 7 units | US Nationwide Distribution in the states of CO, IL, IN, FL, CA, AZ | Open, Classified |  |  | 7 | Not_Computer
Z-0716-2017 | 75570 | No | LWR |  |  | Cardiovascular | PMA | Perceval Sutureless Aortic Heart Valve size S, Item # ICV1208, REF PVS21; Perceval Sutureless Aortic Heart Valve size M, Item # ICV1209, REF PVS23; Perceval Sutureless Aortic Heart Valve size L, Item # ICV1210, REF PVS25; Perceval Sutureless Aortic Heart Valve size XL, Item # ICV1211, REF PVS27. |  |  | October 31, 2016 | 2016 | Sorin Group USA, Inc. | Sorin Group USA, Inc. announces a voluntary field action for the Perceval Sutureless Heart Valve due to updated steps associated with the implantation procedure. | Use error | The firm, LivaNova, sent a "Customer Letter" dated 10/31/2016 to its Consignees with clarifications about the implant of the Perceval Sutureless Aortic Valve. The letter described the product, problem and actions to be taken. The consignees were instructed to reinforce the steps that may influence procedural success and mitigate potential complications and complete and return the Customer Response Form via fax to: LivaNova Canada Corp, Canada at +1 303 895 2335 or email to: HV.FEEDBACK@LIVANOVA.COM.If you have additional questions or request of clarifications, please contact your LivaNova representative or Customer Service.For questions regarding this customer letter, please contact your LivaNova sales representative or email: hv.feedback@livanova.com. |  | 845 | 845 units | Worldwide Distribution-US (Nationwide) including states of: MD, ME, MA, NH, NY, NJ, AL, MS, VA, FL, CA, HI, TN, GA, WI, TX, MI, OH, KS, MN, NC, IL, IN, WA MO, PA, NM, and UT and countries of: Austria, Australia, Belgium, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Jordan, India, Italy, Iran, Islamic Republic of, Ireland, Israel, Korea, Republic of, Kuwait, Luxembourg, Malaysia, Malta, Mexico, Netherland, New Zealand, Norway, Panama, Poland, Portugal, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South African, Spain, Switzerland, Romania, Russian Federation,Taiwan, Province of China, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Open, Classified |  |  | 845 | Not_Computer
Z-0715-2017 | 74948 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Peripherally Inserted Central Catheters (PICC), UPN H965750191, Catalog # 75-019 |  |  | June 08, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | There is a potential for stylet stiffening that is not accurately represented on the labeling. The label states that product includes a stylet with an outside diameter (OD) measuring 0.014, however, the product may contain a stylet with an outside diameter (OD) measuring 0.016. | Mixed-up of materials/components | Angiodynamics, Inc. (Navilyst Medical, Inc.) sent an Urgent Voluntary Medical Device Recall letter dated June 8, 2016, with reply forms to all affected customers. Customers were advised to immediately remove any affected product from inventory, segregate and secure in a location for return. Immediately forward a copy of the recall notification to all sites to which they have distributed affected product. If affected product is located in your institution, please call Navilyst Medical Customer Service at 1-800-772-6446 8am-7pm, EST Monday - Friday to obtain a replacement or credit for your returned product. Promptly complete, sign and return the Reply Verification Tracking Form by email to recall@angiodynamics.com or by Fax to 1-800-782-1357.For questions regarding this recall call 518-792-4112. |  | 50 | 50 cartons (boxes of 5) | Nationwide Distribution to CA, FL, GA, HI, LA, MI, MN, MO, NV, OH &amp; TX | Open, Classified |  |  | 98 | Not_Computer
Z-0719-2017 | 75620 | No | ELW | Impression material. | Dental | Dental | 510(k) | Virtual XD Test Pack Heavy/Light Fast Set, 2 x 50 ml, REF #/Product Code 646469, Rx ONLY -- |  |  | November 09, 2016 | 2016 | Ivoclar Vivadent, Inc. | The firm received complaints claiming the dental material failed to set up. As the dental material ages, the set time may increase. | Under Investigation by firm | Ivoclar Vivadent sent and URGENT - MEDICAL DEVICE RECALL Letters (dated 11/07/2016) and Recall Response Forms to customers via Certified Mail-Return Receipt Requested. The letter identified the affected product, problem and actions to be taken. Customers were advised to return all affected products in stock. For questions contact Ivoclar Vivadent Customer Service at 800-533-6825. |  | 2886 | US: 2090 units, Canada: 331 units, Australia: 465 units | Worldwide Distribution - US Nationwide and the countries of Canada and Australia | Open, Classified |  |  | 10606 | Not_Computer
Z-0680-2017 | 75575 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Model: VANTA Analytical X-Ray System |  |  | October 27, 2016 | 2016 | Olympus Scientific Solutions Americas | Possible unintentional x-ray emission after users attempt early termination of the sequence. | Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 41 units located in the United States will consist of the following:1. Shipments were stopped until a software correction was designed and validated. The software correction was designed by October 27, 2016, and was implemented on November 7, 2016 for all units that had not left the place of manufacture and for all new products.2. The software correction will be made available to all distributors, sales personnel and users. Issue 1 is corrected by altering the behavior of the stop button and trigger such that they stop the current test and terminate the test sequence. To correct issue 2 the software will be modified to grey-out the Stop button during this very small period making it consistent with rest of the VANTA Start/Stop test UI paradigm. Both during user Start and Stop actions, users are shown a greyed button to acknowledge user Click and the button is un-greyed when the system completes the user operation. Both user operations typically take 100  400 milliseconds to complete.3. All VANTA units in the United States are currently in the possession of Olympus or your distributors. No units are in the possession of users.4. Olympus or its distributors will apply the software correction prior to delivering a VANTA system to any purchasers in the United States.For further questions, please call (781) 419-3900. |  | 41 | 41 units | US Distribution to the states of : TX, WA, GA, MA, SC, AZ, PA, VA, CA, OH, NJ, MD, CO, IL, LA, and MN. | Open, Classified |  |  | 41 | Computer
Z-0720-2017 | 75633 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System |  |  | October 11, 2016 | 2016 | RAYSEARCH LABORATORIES AB | An error may occur with the display of dose computed on images other than the planning CT (auxiliary CT) when using multiple patient cases in RayStation 5. If a CT image set with the same Frame of Reference as the displayed auxiliary CT set exists in another case, the dose display may be incorrect. When the error occurs, the display of dose in patient views, including the maximum dose position, the dynamic isodose lines and the dose grid may be incorrect. The dose value normally displayed in the upper left corner when pointing in the 2D view may be incorrect or missing. | Software design | On October 11, 2016 RaySearch Laboratories distributed a Field Safety Notice, Medical Device Correction #13529 to customers via email. Customers were instructed to take caution when a patient has multiple cases. The following may be incorrect if there are image sets that shre the same Frame of Reference residing in different cases: *Display of evaluation dose computed on another CT image set and *Dose tracking fraction dose and deformed dose. Until a corrected version has been installed, all affected users must maintain awareness of this field safety notice. In these cases, the display of the dose values in the 2D view, including the dose value shown when pointing in the view, maximum dose position, dynamic isodose lines and dose grid display may be incorrect. Customers are encouraged to educate staff and all users about this. For regulatory information, please contact David Hedfors, at +46 8 510 530 12 or david.hedfors@raysearchlabs.com. |  | 540 | 540 units | US Nationwide Distribution in the states of AZ, AR, CA, CT, DE, FL, GA, HI, IL, LA, MA, MD, MN, MI, MS, MO, NJ, NY, NC, OH, PA, RI, TX &amp; WA | Open, Classified |  |  | 540 | Not_Computer
Z-0727-2017 | 74368 | No | NAC | Hydrophilic wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AMD-Ritmed Transparent Dressing - Window with label, 4 " x 4 3/4" (10 cm x 12 cm), Reorder No./Code A60044-W, Sterile |  |  | June 08, 2016 | 2016 | Amd-Ritmed, Inc. | The lot of transparent dressing was found to be contaminated (not sterile) based on FDA sampling and analysis. | Material/Component Contamination | URGENT: MEDICAL DEVICE RECALL Letter (dated 6/8/2016) was e-mailed to the consignees on 6/8/2016. The recall is being conducted to the retail level. For questions. please contact AMD Customer Service by phone at 1-800-445-0340 or by e-mail at contactus@amdritmed.com. |  | 216 | US: 216 cases (200 dressings per case) | NY, MA, VA, TN | Open, Classified |  |  | 216 | Not_Computer
Z-0726-2017 | 75713 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | Cut lines on the image may present horizontally rather than vertically. | Component design/selection | MERGE sent an Urgent Medical Device Recall dated March 28, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Merge Healthcare is working on correcting the issues and will inform customers once a fix is made available. Customers were asked to reply using the enclosed form and the return addressed envelope. If customers have further distributed the affected product they need to identify their customers and notify them at once of the product recall. Customers with questions were instructed to send an email to recall@merge.com. |  | 537 | 537 sites potentially have the affected versions | Worldwide Distribution - US (nationwide) Canada and foreign distribution | Open, Classified |  |  | 537 | Not_Computer
Z-0724-2017 | 75683 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Halyard C-Section Drape with Clear Screen and Full Incise (Non-Sterile), Product Code 44965 NS; C-Section Drape with Clear Screen and Full Incise (Sterile), Product Code 44966 00; C-Section Drape with Clear Screen (Non-Sterile), Product Code 44967 NS; C-Section Drape with Clear Screen (Sterile), Product Code 44968 00; C-Section Drape with Pouch and Fenestration (Non-Sterile), Product Code 44977 NS; C-Section Drape with Pouch and Fenestration (Sterile), Product Code 44978 00 |  |  | November 09, 2016 | 2016 | Halyard Health, Inc | Drapes within the affected lots may include a manufacturing variation that prevents convenient removal of the coated release paper, which inhibits easy access to the surgical site. | Nonconforming Material/Component | Accounts were notified of the recall by letter on 11/9/2016. The consignees were instructed to quarantine and discontinue use of the recalled lots and to discard/destroy per the facility's standard disposal procedure for nonwoven materials. A response form was included which was to be sent back. |  | 132 | 132 cases | US Nationwide Distribution in the states of AL, FL, GA, IL, KS, MN, MO, NE, NJ, OH, PH, TN, VA, WI, WY | Open, Classified |  |  | 132 | Not_Computer
Z-0722-2017 | 75681 | No | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | Barbiturates Intercept Micro-Plate EIA (enzyme immunoassay, barbiturate) |  |  | November 10, 2016 | 2016 | OraSure Technologies, Inc. | Shipping error. Cocaine metabolite Microplate distributed instead of Barbiturate Microplate due to incorrect selection of Microplate during manufacturing. | Mixed-up of materials/components | Due to the small number of customers that were affected by the recall, OraSure notified customers on November 10, 2016, via telephone with a follow-up email. Customers were asked to quarantine the units on hand and they will schedule a pickup by Fed EX to retrieve the affected units.For further questions, please call (610) 882-1820, ext. 1657 |  | 4 | 4 kits | Distributed to KS, and VA | Terminated | April 17, 2017 | 158 | 4 | Not_Computer
Z-0723-2017 | 75652 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 560 Hematology Systems, Siemens Material Number (SMN) 11170842, IVD. |  |  | October 21, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens is investigating an issue which may cause an incorrect result to be reported. Siemens received two reports of multiple discordant records for the same Sample ID in the ADVIA 560 Hematology System database that occurred during the installation of the systems. The database should only contain one record of a Sample ID number for any given time and date. If there are multiple records for the same Sample ID, it is possible that multiple results may be manually or automatically sent to the Laboratory Information System (LIS), printed or displayed on the results report screen. | Software design | An Urgent Medical Device Correction Letter HI17-02.A.US (dated October 21, 2016) and Response Form were sent to US Customers via Federal Express. An Urgent Field Safety Notice was sent to customers outside the US on October 21, 2016, (distribution is determined at the country level). The letters instruct customers to contact their local Siemens technical support representative if they obtain multiple results for one Sample ID. It also informs the customer what to do if the ADVIA 560 Hematology System generates any of the six error messages listed in the letter. For further questions, please call (312) 275-7795. |  | 164 | US: 23 systems; Foreign: 141 systems | Worldwide Distribution - US Distribution to the states of including Puerto Rico and : CA, CO, FL, IL, MI, NJ, NY, OH, TX, WA, WI, WI and WY., and to the countries of : Angola, Australia, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, France, Germany, Greece, India, Italy, Japan, Kenya, Malaysia, Mexico, Myanmar, Netherlands, Norway, Pakistan, Portugal, Saudi Arabia, Slovenia, Spain, Switzerland, Turkey, Uganda, United Kingdom, Vietnam. | Open, Classified |  |  | 164 | Not_Computer
Z-0721-2017 | 75731 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 4.0, 4.5, 4.7 and 5.0; Radiation Therapy Treatment Planning System |  |  | November 09, 2016 | 2016 | RAYSEARCH LABORATORIES AB | Software anomaly; an issue was found with the proton Pencil Beam Scanning (PBS) dose calculation in RayStation 4.0, 4.5, 4.7 and 5.0. For treatment plans with a combination of range shifter, large air gap and beams that enter the patient surface at an oblique angle, the dose calculation accuracy may be less than expected. | Software design | Ray Search sent an Urgent Field Safety Notice Medical Device Correction letter (dated 11/8/2016) via e-mailed to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Inspect their product and identify all installed units with the above software version number, then confirm you have read and understood this notice (contact information included). |  | 66 | US: 66 units | US Nationwide Distribution in the states of California, Florida, Illinois, Louisiana, Michigan, New Jersey, Tennessee, Texas, Washington | Open, Classified |  |  | 66 | Computer
Z-0725-2017 | 75263 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Magnesium on RX Imola analyser |  |  | April 14, 2016 | 2016 | Randox Laboratories, Limited | According to the firm, Carry over was observed when the amylase or pancreatic amylase test is run directly before or after the Magnesium assay. An update is required to the RX Imola analyser running order and an acid wash recommended when setting up the magnesium assay. A correction was made to the IFU for the Magnesium assay and all RX Imola customers in the USA were contacted with the updated instruction. | Device Design | Randox sent an Urgent Field Safety Notice to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review and update Rx Imola running order where required. Rework any remaining stock with updated IFU and Important Notice. Complete and return the vigilance response section of the form to technicalservikces@randox.com. |  | 15 | 15 | Nationwide Distribution The product was only distributed to US Consignees. | Open, Classified |  |  | 15 | Not_Computer
Z-0659-2017 | 75600 | No | LOZ |  |  | Cardiovascular | PMA | 5000 Series Freedom Drivers. Freedom Driver System for Temporary Total Artificial Heart (TAH-t). Part number 595000-001 |  |  | October 21, 2016 | 2016 | SynCardia Systems Inc. | The Main Printed Circuit Board Assembly (PCBA) of the affected Freedom Drivers may fail and cause the Freedom Driver to stop functioning without visual or audible alarms, resulting in the loss of life-sustaining function. | Process control | On 10/21/16 customers were notified by phone. Customers were informed that patients who were being supported by affected Drivers must be switched to unaffected Drivers.The firm will provide unaffected Freedom Drivers to replace the affected Drivers. Affected Drivers must be return to the firm. Instructions for return provided, and Acknowledgment Form must be completed.Contact Information For QuestionsMary Pat SloanSenior Vice PresidentGlobal Certification and LogisticsSynCardia Systems, LLC1992 E Silverlake RdTucson, AZ 85713Direct 520-547-7470Toll-free 866-480-1122 Ext 1308Monday-Friday 7:30 AM to 5:00 PM Pacific/Arizona timeHotline 24/7 866-771-9437www.syncardia.comOn 10/27/16 customer notification letters were sent. The letter provided instructions on how to recognize the device could fail. The patient and caregiver will recognize that an affected Driver is failing because the Freedom Driver will stop pumping without generating any sound or any visual or audible alarms.When functioning, the Freedom Driver exhibits a distinctive pumping sound; when it stops pumping the lack of sound is noticeable.Generally, if the Freedom Driver function is interrupted, a visual and/or audible alarm will occur. However, as a result of this component failure, when affected Freedom Drivers stop functioning, there will be no visual or audible alarms, and no pumping sound.Actions to Be Taken by The Customer/UserSynCardia has provided unaffected Freedom Drivers to the two hospitals to replace the affected 5000 series Drivers. All patients with affected Freedom Drivers have undergone the exchange and received unaffected Drivers.The hospitals have returned all the affected Freedom Drivers to SynCardia.The hospitals are requested to complete the Recall Acknowledgement and Receipt Form below and return to SynCardia. |  | 5 | 5 | US distribution to Virginia and Arizona. | Open, Classified |  |  | 5 | Not_Computer
Z-0731-2017 | 75853 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | December 02, 2016 | 2016 | Elekta, Inc. | Edits to Particle field definition parameters may not be saved when the field definition window is saved. | Software design | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTIFICATION" letter dated 11/2016 was sent to affected customers on 12/2/2016. The letter describes the product, problem and actions to be taken. The customers were instructed to save and re-open the field definition form when manual edits have been performed to confirm that the edits were save; post this notice in a place accessible to all users; advise the appropriate personnel working with the product the content of the letter; and complete and return the Acknowledgement Form via fax to: 408 830 8003 immediately upon receipt but no later than 30 days.Elekta will resolve the problem with the MOSAIQ 2.60 SP7F, MOSAIQ 2.60 SP14 and 2.64 SP3. Customers are encouraged to contact Elekta support to get the corrected version.If you have any questions, contact RA Manager Americas at 1-770-300-9725 or email to: Elekta_Global_PMS@elekta.com. |  | 47 | 47 | Worldwide Distribution-US (nationwide) to states of: FL, IL, MO, NJ, OH, OK, TN, TX, VA, and countries of: Italy, Japan, South Korea, Sweden and United Kingdom. | Open, Classified |  |  | 47 | Not_Computer
Z-0729-2017 | 75825 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | If images are sent without an order in the system, they will be matched with the latest order on the current patient/modality matching potentially resulting in the matching of the report to the incorrect accession number (but still associated to the correct patient). | Software design | MERE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers via email or via certified mail if they did not have an email address for the customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to respond to the notification using the enclosed form and the return addressed envelope. Customers were also instructed to ensure that all users of the product are provided with the notification and if product was further distributed customers need to identify their customers and notify them at once of the product recall. Customers with questions should send an email to recall@merge.com. For questions regarding this recall call 877-741-5369. |  | 198 | 198 sites potentially have the effected versions | Nationwide Distribution | Open, Classified |  |  | 198 | Not_Computer
Z-0730-2017 | 75714 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | October 01, 2015 | 2015 | Merge Healthcare, Inc. | Reporting feature times out after inactivity for more than an hour sending the user back to the study list, which causes all reporting data being entered to be lost. | Software design | MERGE sent an Urgent Medical Device Recall letter dated March 28, 2016, to all affected customers. Non-responders were issued a second letter dated July 11, 2016, The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to reply using the enclosed form and the return addressed envelope and also to ensure that all users of the product were provided with the notification. If product was further distributed, customer were asked to identify their customers and notify them at once of the product recall. Customers with questions were advised to email recall@merge.com.For questions regarding this recall call 877-741-5369. |  | 116 | 116 sites potentially have the effected versions | Nationwide Distribution | Open, Classified |  |  | 116 | Not_Computer
Z-0728-2017 | 74488 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Carry bar which is an accessory to a lift used for patients in hospitals and other facilities. Recall impacts the following model numbers: 360741, 360750, 360751, 360755, 360755 Rev B, 360756, and 360757. |  |  | August 30, 2016 | 2016 | ErgoSafe Products, LLC (DBA) Prism Medical | The black plastic puck on the carry bar is breaking resulting in potential for patient harm. | Device Design | The recalling firm sent field safety notices to their customers by mail on August 30, 2016. The notice instructed customers to inspect their carry bar to see if it used a black plastic puck. If it did, they were instructed to contact the recalling firm to receive a metal insert to install. If the puck is cracked or damaged, they can request a new, replacement carry bar at a discounted price. For further questions, please call (866) 891-6502. |  | 17964 | not greater than 17,964 units | Distributed US nationwide and to Canada, Japan, Australia, Thailand, Great Britain, and South Korea | Open, Classified |  |  | 17964 | Not_Computer
Z-0746-2017 | 75385 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Harmony Insight Diagnostic Monitor Support System - 100 Spring Arm (EMS) |  |  | October 05, 2016 | 2016 | Steris Corporation | The Drager Pendula spring arm device has a set of screws located on the spring arm that were not properly secured with Loctite and may become loose. | Component design/selection | Steris sent an Urgent Voluntary Field Correction Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remind staff personnel that Harmony Insight Diagnostic Monitor Support Systems, including the affected Drager Pendula spring arms, are not designed or intended to be placed over hospital personnel or patients and should be moved with care. Customers with questions were instructed to call 440-392-7358 or Steris Customer Service at 1-800-548-4873. For questions regarding this recall call 440-392-7231. |  | 63 | 63 | Nationwide Distribution including AK, AZ, CA, CO, FL, IL, KY, MI, MO, NB, NJ, NY, OH, PA, RI, SD, TN, TX, UT, WA &amp; WV. | Terminated | April 14, 2017 | 191 | 63 | Not_Computer
Z-0732-2017 | 75460 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Lumex, Patriot Full Electric Homecare Bed, &amp; Patriot LX Full Electric Homecare Bed containing the Okin Dumat 3 Head/Foot Motors. |  |  | September 30, 2016 | 2016 | GF Health Products, Inc | GF Health Products, Inc., in conjunction with OKIN America, has initiated a voluntary field correction affecting the OKIN America "DUOMAT 3" Head/Foot motor used in the manufacture of our Lumex Patriot Full Electric Bed Model US0458/US0468 and Patriot LX Full Electric Bed Model US6000.The failure mode is directly related to contaminated material in the plastic housing used to produce the DUOMAT 3 motor. Contaminated motor housings may crack resulting in a motor that will no longer actuate the head section per the intended design.If the head section were to fall it is unlikely but possible an injury could occur. No injuries have been reported. | Process control | Graham Field sent an "Urgent Correction" letter dated September 30, 2016. The letter described the problem and the product involved in the recall. The letter described the " Action Required: Motor Replacement" . Informed consignees to contact GF Health Products at 1-888-974-4347 or fieldcorrection@grahamfield.com within ten (10) business days from receipt of this notification to schedule motor replacement order. Requested consignees to return the completed Certification of Replacement and Return Agreement to: fieldcorrection@grahamfield.com or Fax: 1-920-929-8213. For additional information, please call 1-888-974-4347; or at fieldcorrection@grahamfield.com. |  | 3547 | 3,547 (3063 beds, 484 motors) | Worldwide Distribution - USA ( nationwide) Distribution including Puerto Rico and to the countries of : Bahamas, Canada, Chile, El Salvador and Santo Domingo. | Open, Classified |  |  | 3547 | Not_Computer
Z-0747-2017 | 75508 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer RoboRack 25ul, Clear, Non-Conductive, Filter Tips, Pre-sterilized |  |  | June 24, 2016 | 2016 | PerkinElmer, Inc | Filter Tips molded incorrectly and may not seal to the Varispan tip adaptor properly causing inaccurate aspiration and dispensing | Process control | PerkinElmer initiated recall on June 24, 2016 to the Product Manager and was delivered to the customer by email facilitated through PerkinElmer Sales Team.Customers were requested to take the following actions until PerkinElmer's corrective measures are completed:Inspect customer inventory of the 6000689 25 Clear Filter Pipette Tip and identify ifthe affected lots are present. Complete and return the attached Response Form with the required information. Destroy all quantities of tips from the affected lots.Upon receipt of the return Response Form, a shipment of replacement tips will be shipped. |  | 595 | 595 trays (571,200 tips) | Worldwide Distribution - US including CA, MA, NY, IA, OR , PA, WA and Internationally to Belgium, Canada, Finland, France, Germany, Japan, Netherlands, Singapore, and UK | Open, Classified |  |  | 595 | Not_Computer
Z-0741-2017 | 75638 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | McKesson Powder Free Vinyl Exam Gloves Catalog numbers 14-116, 14-118, 14-120, 14-136, 14-138 and 14-140 |  |  | October 12, 2016 | 2016 | CYPRESS MEDICAL PRODUCTS | A container of McKesson vinyl gloves that failed the FDA leak test was inadvertently distributed. | Process control | A letter, dated October 19, 2016, on Cypress Medical Products letterhead, was sent to customers. Customers were requested to stop use and quarantine the gloves. |  | 2005.2 | 2005.2 cases have been distributed | US Distribution. This recall is limited to one distribution center in Auburn, Washington, and the customers serviced from that distribution center. | Terminated | March 27, 2017 | 166 | 2005.2 | Not_Computer
Z-0733-2017 | 75793 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Cook Spectrum Minocycline/Rifampin Impregnated Double Lumen Central Venous Catheter Tray |  |  | November 28, 2016 | 2016 | Cook Inc. | Cook Medical has received reports of the trays containing the incorrect needle, which in some cases could result in the inability to pass a wire guide through the needle.Potential adverse events that may occur as a result of utilizing needles of incorrect size include delay in treatment, which could be critical in an emergent situation. There have been no reports of illness or injury associated with this issue. | Mixed-up of materials/components | Cook Medical sent an Urgent - Medical Device Recall letter dated November 23, 2016. to all affected customers. The letter identified the product the problem and the action needed to be taken. Action to be taken: 1. Please review the attached list of affected products and lot numbers that were shipped to your account and quarantine any affected products that remain unused. 2. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit. 3. Please complete the attached Acknowledgement and Receipt Form and return it to Cook Medical. 4. Please report any adverse event to Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235, Monday through Friday between 7:30 am and 5:00 pm Eastern Standard Time or e-mail at fieldactionsna@cookmedical.com. Transmission of this notice: This notice must be passed on to appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at (800) 457-4500 or 1(812) 339-2235. |  | 99 | 99 units | AZCADCGAILLAMNMSNCNVNYOKPATNTXWI | Open, Classified |  |  | 99 | Not_Computer
Z-0740-2017 | 75802 | No | OFT | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | Enforcement Discretion | BD Spinal Anesthesia Tray |  |  | November 07, 2016 | 2016 | Becton Dickinson &amp; Company | BD is conducting a voluntary product recall of the BD Spinal Anesthesia Tray, Cat (Ref) 405671, lotB01K089D since a very low number of trays are incorrectly labeled as BD Single Shot Epidural Tray,Cat (Ref) 406069. | Error in labeling | BD sent an Urgent Product Recall letter dated November 7, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately review their inventory for the affected product. Return the affected products with the Business Response Card form following the instruction on the enclosed packing instruction. The completed form should be faxed to 1-866-240-5212 or email the completed form to bd6549@stericycle.com. Customers with questions were instructed to call 1-866-792-5451. For questions regarding this recall call 866-792-5451. |  | 4500 | 4500 | Nationwide Distribution | Open, Classified |  |  | 4500 | Not_Computer
Z-0739-2017 | 75820 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Performance-LOAD Ambulance Cot Fastener |  |  | November 08, 2016 | 2016 | Stryker Medical Division of Stryker Corporation | Stryker Medical initiated a voluntary recall of Stryker Performance-LOAD Cot Fasteners due to complaints that fastening system may not have been securely fastened or had an inability to fasten into the Performance-LOAD and bounced back during loading which could cause injury to consumers. | Under Investigation by firm | Stryker Medical initiated a voluntary recall of Stryker Performance-LOAD Cot Fasteners MODEL: 6392, due to complaints and service reports that some cots compatible with the Performance-LOAD cot fastening system may not have been securely fastened or had an inability to fasten into the Performance-LOAD (bounced back during loading) which could lead to injury of consumers. Customer notification letters were sent via Fedex on November 8, 2016. Customers were asked to quarantine the impacted devices and submit completed business reply forms by fax (269)488-8691 or email productfieldaction@stryker.com. Upon receipt of the completed business reply form, Stryker will contact you toarrange for the modification/repair of your Performance-LOAD unit to ensure that the unit securely fastens the cot within the ambulance.4. In the interim, until the service can be performed, there are additional actions that users can and should take to ensure that the cot is securely fastened into the Performance-LOAD system.a. Per the PerformanceLOAD Operations Manual (P/N 6392-009-001, pages 26-27. If customers have loaned or sold any of the products listed in this letter, please forward a copy of the recall notice to the new users and advise Stryker of their new location in the space provided on the business reply form. If you have any questions or concerns, please contact Stryker Customer Service (1-800-327-0770). Our normalbusiness hours are Monday-Friday 8 a.m-6 p.m. (EST). |  | 749 | 749 units | Worldwide Distribution to AK, AL , AR , AZ , CA , CT , DC , DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA , RI, SC, TN, TX, UT,VA, VT ,WA ,WV ,WY and Internationally to Canada and the Netherlands | Open, Classified |  |  | 749 | Not_Computer
Z-0742-2017 | 75812 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition AS System, x-ray, tomography, computed |  |  | November 15, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Software update that provides software and firmware bug-fixes to improve system performance | Software Design Change | Siemens mailed a Customer Information Letter to affected customers on November 21, 2016 informing them of software updates that would provide bug-fixes to improve system performance and stability for customers with SOMATOM Definition AS, Definition Edge and Definition Flash systems using Software versions VA48A_SP1 and VA48A_SP2. The update will remotely be pushed to their systems or customers will be contacted by one of Siemens service organization to arrange a date for installation of the update on site. Customers with questions were instructed to call 1-800-888-7436. |  | 104 | 104 systems | Nationwide Distribution | Open, Classified |  |  | 104 | Computer
Z-0734-2017 | 75863 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System using software version 10.6; |  |  | November 22, 2016 | 2016 | Accuray Incorporated | Accuray has become aware of a potential safety issue involving unexpected treatment robot motion when removing an accessory. This issue is limited to certain CyberKnifeÂż Systems with software version 1 0.6. | Software design | Accuray sent an Urgent Safety Advisory Notification letter dated November 22, 2016, to all affected customers. Letters alerted users to the issue and the sequence of events necessary for it to occur. Accuray reps will visit all sites to perform a software upgrade to correct the issue.For more information or any questions about this notification, please contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Technical Solution Center at http://www.accuray.com/services-support.For further questions, please call (408) 716-4600. |  | 40 | 40 affected devices | Worldwide Distribution - US Distribution to the states of : PA, WA, MT, NY, SC, CA, MN, AK, NJ, FL, KY, WI., and to the countries of : Bulgaria, Canada, France, Germany, India, Japan, Luxembourg, Netherlands, Portugal, Switzerland and Turkey | Terminated | May 22, 2017 | 181 | 40 | Computer
Z-0738-2017 | 75469 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | BARDIAÂż 60cc Piston Irrigation Syringe, Sterile, Rx Only, The BARDÂż Syringe is available in a single use plastic design which features a 60cc or 70cc capacity graduated barrel. It has a thumb ring plunger for one-handed use. A catheter adapter tip and Luer adapter tip are included. BARDÂż Syringe allows one-hand use with a 60cc or 70cc graduated barrel and thumb-ring plunger and is indicated for catheter irrigation use. BARDÂż Syringe is intended for Cytoscope or Resectoscope irrigation, catheter irrigation and Foley catheter balloon inflation. |  |  | August 03, 2016 | 2016 | C.R. Bard, Inc. | Package Defect: Products may be at risk for having a slit defect on the package and that may affect the product sterility. | Process control | BARD Medical sent an Urgent Medical Device Product Recall letter dated August 3, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately examine their inventory and quarantine product subject to the recall. Complete and return the accompanying Recall &amp; Effectiveness Check Form attached to the letter. If product was further distributed please identify those customers and notify them at once of the product recall by forwarding a copy of the letter. For questions regarding this recall call 770-784-6220. |  |  |  | Worldwide Distribution - US (nationwide) and Internationally to Australia, New Zealand, Canada, Mexico, France and Italy. | Open, Classified |  |  |  | Not_Computer
Z-0745-2017 | 75252 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD 23 G x .75 in. BD VacutainerÂż Push Button Blood |  |  | August 08, 2016 | 2016 | Becton Dickinson &amp; Company | BD hasreceived multiple reports of units associated with the 23 G x .75 in. BD VacutainerÂż Push Button BloodCollection Set with 12 in. Tubing whose lots are exhibiting separation offront and rear barrels upon activation of the safety feature, which retracts theneedle. | Device Design | The firm, BD, sent a "URGENT PRODUCT RECALL" letter dated 8/8/2016 to its customer with a return response form via UPS 2nd Day mail on 8/8/2016.The letter described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory for the specific lot numbers, quarantine the product and discontinue the shipment of the affected product; complete and return the Recall Response Form by fax to BD at 1-866-913-0661 or email to bd6835@stericycle.com (even if you do not have any product); return all affected products with completed response form per packing instructions; if you have distributed this product, identify your customers and notify them immediately of this product recall; if you would like BD to conduct the notification to your customers, email your customer list within 3 business days to Matthew.Kelleher@bd.com.if you have any questions or require assistance with the return of the recalled product and/or availability of credit product, contact 1-888-628-0732 between 8Am and 5 PM ET Monday through Friday. |  | 1788150 | 1,788,150 units | Worldwide distribution- US (nationwide) and countries of: Canada, Belgium, Brazil. | Open, Classified |  |  | 1788150 | Not_Computer
Z-0750-2017 | 75786 | No | POA |  |  | Microbiology | EUA - Emergency Use Authorization | LightMixÂż Zika rRT-PCR Test |  |  | November 16, 2016 | 2016 | Roche Molecular Systems, Inc. | The LightCyclerÂż 480 algorithm used for the LightMixÂż Zika rRT-PCR Test, EUA (catalog number07987897001) occasionally (~1% of the time) calls a positive result for a negative sample if there isan inflection point in the fluorescence signal. | Software design | Roche Molecular Diagnostics Inc. sent an Urgent Medical Device Correction Letter dated November 16, 2016, to all affected customers via UPS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until Roche provides updated Instructions for Use to assist in result interpretation. If product was further distributed, ensure the Urgent Medical Device Correction letter is provided to that site. Also, customers were instructed to completed the attached fax form and fax it to 1-888-912-2188. For questions customers should call 1-800-526-1247. |  | 24 | 24 pieces | Nationwide Distribution | Open, Classified |  |  | 24 | Not_Computer
Z-0748-2017 | 75351 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ellman A CYNOSURE COMPANY Minor Surgery Bipolar Cable, Contents: Five (5) Disposable Bipolar Cables,STERILE |  |  | September 07, 2016 | 2016 | Ellman International, Inc. | Mislabeling of the "expiration date". The expiration date on the outer box and the individual pouches do not match. The expiration date on the box is correct with a date of 2018-02. | Employee error | Cynosure, Inc. dba Ellman distributed Urgent Medical Device Recall Notices dated September 6, 2016, and response forms to their customers via UPS overnight. Cynosure has identified an error in the Expiration Date on the label of the Minor Surgery Bipolar Cable pouch. The Expiration Date on the individual pouches that are provided in the 5 pack boxes do not match the Expiration Date on the box. The Expiration Date on the box is correct with a date of 2018-02, but the label on the individual pouches has a date of 2021-02. Actions To Be Taken - Return any cables subject to recall. You will be contacted by a Cynosure Customer Service Representative who will provide you with a RMA number for any cables that are being returned. Complete and return the response form via email to choy@cynosure.com. Cynosure will replace affected products at no charge upon receipt of the returned cables. Customers with questions should call 978-256-4200 or 800-886-2966.For questions regarding this recall call 516-267-6502. |  | 370 | 370 boxes (US - 159 boxes, OUS - 211 boxes) | Worldwide Distribution - US including CA, FL, MA, MD, MI, MN, MS, OH, PA, VA, TX &amp; WV; **International**Australia, Bangladesh, Greece, Italy, Japan, Lebanon, Russian Federation, Singapore, Spain &amp; Vietman. | Open, Classified |  |  | 370 | Not_Computer
Z-0752-2017 | 75710 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Metzenbaum Scissors 7 , Curved, Sterile |  |  | October 28, 2016 | 2016 | Integra York PA, Inc | The Chevron seals of 6x10 Tyvek peel pouch used to package the sterile instruments may separate causing the packages to open and compromise sterility of the instrument. | Process control | The firm, Integra, mailed an "URGENT MEDICAL DEVICE RECALL NOTICE" dated October 26, 2016 to customers affected by the recall to notify them of the issue. The notice described the product, problem and actions to be taken. The customers were instructed to immediately review their inventory to determine if they had any of the recalled product on hand and to discontinue distributing the product; review your customer traceability records for shipment of the product and forward a copy of this Recall Notice to any of your customers that purchased the affected lots; complete and return the Acknowledgement Form that accompanied the notice (even if you do not have any product(s)) by email to: FCA3@integralife.com or fax to: 1-609-750-4220.Should you have any questions regarding these instructions, please contact Customer Service at 1-866-854-8300. |  | 50 | 50 units | US Distribution in NC and TX. | Terminated | April 17, 2017 | 171 | 200 | Not_Computer
Z-0749-2017 | 75682 | No | OUP | Powered surgical instrument for improvement in the appearance of cellulite. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Cellfina Prep Pack, Part No. CP1 |  |  | November 14, 2016 | 2016 | Ulthera Inc | It was discovered that a non-sterile vacuum tube is supplied within the pack, although the Instructions for Use (IFU) describe the use of a sterile vacuum tube as part of the procedure setup. | Other | Ulthera sent an Urgent Medical Device Recall letter dated November 15, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It was discovered that a non-sterile vacuum tube is supplied within the pack, although the Instructions for Use (IFU) describe the use of a sterile vacuum tube as part of the procedure setup. Customers are informed of the actions to be taken. Customers with any technical or logistical questions are instructed to contact technical support at Ulthera, Inc. at (480) 619-4069. The letter informs the customers with questions related to their Cellfina business are instructed to contact Merz North America at 888-751-7546 or at mymerzsolutions@merz.com. |  | 6562 | 6562 units | Worldwide Distribution - US Distribution and to the countries of :Europe (Austria, France, Germany, Italy, Netherlands, Switzerland, Spain, UK) | Open, Classified |  |  | 6562 | Not_Computer
Z-0755-2017 | 75584 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Avance, Avance CS2, Amingo. |  |  | November 03, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where certain Avance CS2, Avance and Amingo anesthesia devices can transition to a System Malfunction state ifthe lower storage drawer containing the optional large tray insert accessory closed with an abnormally high amount of force. | Device Design | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 34079 dated November 3, 2016, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 23098 | 23,098 (US = 8,711; OUS = 14,387) | USA ( nationwide) Distribution including Puerto Rico and DC. and to the countries ofAustraliaAUSTRIAAzerbaijanBahrainBangladeshBarbadosBelarusBelgiumBOLIVIABOSNIA-HERZEGOVINABrazilBRUNEIBULGARIACanadaCHILEChinaCOLOMBIACOSTA RICACroatiaCyprusCzech RepublicDenmarkECUADOREgyptEstoniaEthiopiaFinlandFRANCEGeorgiaGERMANYGhanaGibraltarGREECEHONG KONGHUNGARIAIcelandIndiaIndonesiaIRAQIrelandIsraelITALYJAPANJORDANKAZAKHSTANKenyaKOREA, REPUBLIC OF KuwaitLatviaLebanonLibyan Arab JamahiriyaLITHUANIAMACEDONIAMALAYSIAMaltaMexicoMOLDOVAMOROCCONAMIBIANetherlandsNew ZealandNicaraguaNigeriaNORWAYOmanPakistanPanamaPARAGUAYPERUPHILIPPINESPolandPORTUGALQatarROMANIARussian FederationSaudi ArabiaSerbiaSERBIA-MONTENESingaporeSlovakiaSLOVENIASOUTH AFRICASPAINSWEDENSWITZERLANDSYRIATAIWANThailandTunisiaTURKEYUkraineUnited Arab EmiratesUNITED KINGDOMUruguayUZBEKISTANVenezuelaVietNamYemen | Open, Classified |  |  | 23098 | Not_Computer
Z-0759-2017 | 75693 | No | HCF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Weck Facial Closure Systems have application in laparoscopic procedures for approximation of tissues and percutaneous suturing for closing incision sites. |  |  | November 16, 2016 | 2016 | Teleflex Medical | The devices wings may become partially detached from the EFx Shield during use. | Other | Consignees were notified via letter on 11/16/2016. |  | 2265 | 2265 units | Domestic; US Nationwide; International: Belgium, Canada. | Open, Classified |  |  | 2265 | Not_Computer
Z-0757-2017 | 75249 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Ultrasound Video Bronchoscope |  |  | May 31, 2016 | 2016 | Pentax of America Inc | The EB-1970UK Ultrasound Video Bronchoscope shipped between theclearance date of April 2014 and September 2015 contained the Operation andReprocessing IFUs that were in use prior to the 2014 clearance, and therefore did notinclude the cleared Operation and Reprocessing IFUs. | Process control | Pentax sent a US Urgent Field Correction letter to all affected customers with a field correction/field safety response form via USPS on May 31, 2016. The letter identified the product, the problem, and the action to be taken by the customer. Revised instructions were as follows:The instructions have been revised in the following manner:Âż Operation IFUo Integrated all models of each family into the IFUo Modified the indications for use language to be consistent with similar productso Enhanced the explanation of the operation of the deviceo Revised and added cautions and warningsÂż Reprocessing IFUo Included more specificity regarding reprocessing activities to aid in the proper method for reprocessing theendoscopes, including:Âż number of brush and rinse steps,Âż quantity of fluid volumes,Âż use of additional text, figures, cautions, and warningsIdentification of Affected Endoscope ModelsTable 1 provides an overview of the Affected Endoscope models, and the part number and revision of the historical andrevised IFUs. The accompanying USB flash drive contains all the revised IFUs identified below. Please distribute only thoseIFUs which support the products owned by your facility.Table 1Affected Models Historical Version of IFUs Revised IFUsEB-1970UKUltrasound Video BronchoscopeZ870 Owners IFUS016_R02 Operator IFUS021_R04 Reprocessing IFUNo.388_R01 AddendumNo.392_R00 AddendumEB-70K seriesVideo BronchoscopesZ705 Operation IFU S015_R04 Operator IFUS020_R04 Reprocessing IFUNo.388_R01 AddendumNo.392_R00 AddendumZ706 Reprocessing IFUEB-75K seriesVideo BronchoscopesS003 Owners IFUCustomers were instructed to ensure that user departments discard the prior versions of the Instructions for Use and ensure that all users read and carefully follow the revised Operation and Reprocessing IFUs to ensure the proper use and cleaning of the affected PENTAX endoscopes. Customers were advised to contact their PEN |  | 113 | 113 units | Nationwide Distribution | Open, Classified |  |  | 113 | Not_Computer
Z-0762-2017 | 75262 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Randox Liquid Cardiac Controls Level 1, Level 2, and Level 3 are liquid controls containing BNP, CK MB Mass, Digoxin, D-Dimer, Homocysteine, hsCRP, Myoglobin, NT-ProBNP, Troponin I, and Troponin T. They have been developed for use in the quality control of BNP, CK MB Mass, Digoxin, D-Dimer, Homocysteine, hsCRP, Myoglobin, NT-ProBNP, Troponin I, and Troponin T assays on various clinical chemistry and immunoassay systems. This in vitro diagnostic device is intended for prescription use only. |  |  | September 16, 2016 | 2016 | Randox Laboratories | An internal real-time stability monitoring program revealed decreased recovery of myoglobin (-20% bias) for Randox Immunoturbidimetric and Siemens Centaur Classic methods and CKMB (-21% bias) for Roche e411, BioMerieux Vidas, and Siemens Centaur Classic methods. | Nonconforming Material/Component | Randox sent an Urgent Field Safety Notice dated September 2, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please note the control value sheets have been updated with the new values. A FSN has been completed to advise customers on the appropriate action to take.For further questions, please call (304) 728-2890. |  | 2817 | 2817 | Worldwide Distribution - US Distribution. | Open, Classified |  |  | 2817 | Not_Computer
Z-0760-2017 | 75312 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | INOVO Purefill Oxygen Compressor Model PF1100 |  |  | July 21, 2016 | 2016 | Inovo, Inc | Device can fail during operation. | Under Investigation by firm | The firm, Inovo, Inc, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated Sept 2, 2016 to all their consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and isolate the devices in their inventory and return them to Inovo Inc. Also, if the consignee sold or distributed the device, the consignee needs to retrieve and return the units to Inovo Inc. as soon as possible and complete and return the Medical Device Return Response form in the self-addressed stamped envelope provided or email to: regulatory@inovoin.com or fax to Inovo, Inc. Any questions please contact at 814-443-7602 or via e-mail at regulatory@inovoinc.com. |  | 117 | 117 units | US Nationwide Distribution to states of: AL, AR, CT, FL, IA, KS, KY, MD, MI, MN, MO, NE, NJ, NV, NY, OH, RI, SC, TN, TX, UT, and WA. | Open, Classified |  |  | 117 | Not_Computer
Z-0756-2017 | 75481 | No | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Covidien Precision Specimen Container- 4 oz (118 mL), Plastic Screw Top Cap, OR Packaged, Sterile |  |  | October 14, 2016 | 2016 | Covidien LLC | Sterility of the outer surface of the container compromised due to packaging defect. | Packaging process control | Covidien initiated recall by letter on October 14, 2016. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory of the affected lots to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department.Accounts requested to completed Recalled Product Return From to obtain an RGA.Questions: Contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 5187100 | 5,187,100 Units | NationwideForeign: Japan | Open, Classified |  |  | 5187100 | Not_Computer
Z-0754-2017 | 75601 | No | FAS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HF Cable WA00014A, Endoscopic electrosurgical unit and accessories |  |  | November 03, 2016 | 2016 | Olympus Corporation of the Americas | Software malfunction that results in incorrect generation or display of error codes. | Under Investigation by firm | Olympus mailed on 11/03/2016 an Urgent Communication Letter to customers affected by the recall. The letter requested that consignees inspect their inventory for the affected lots and call Customer Service to obtain a Returned Material Authorization (RMA) to return the product. |  | 494 | 494 units | Distributed Nationwide | Open, Classified |  |  | 494 | Computer
Z-0763-2017 | 75770 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Ceramax Skin Barrier Cream: |  |  | November 21, 2016 | 2016 | Oculus Innovative Sciences, Inc. | Firm became aware that professional samples of Ceramax are not meeting stability standards set for the product. | Packaging process control | Oculus issued Medical Device Recall letters dated November 21, 2016, to Sales reps notifying them to stop all further distribution and immediately remove the affected lots of Recall product from sample inventory. Customers were also instructed to sign the Response form verifying receipt of the letter. Customers with questions were instructed to call Oculus Innovative Science-Regulatory Affairs at 707-559-7181. For questions regarding this recall call 559-7191. |  | 68970 | 68,970 units | Nationwide Distribution | Terminated | June 13, 2017 | 204 | 68970 | Not_Computer
Z-0753-2017 | 75826 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ecolab, Microtek Rolled Probe Cover PC0907NB, Rx Only. |  |  | November 30, 2016 | 2016 | Ecolab Inc | The reason for the recall is that the inner carton box of the Rolled Probe Cover PC0907NB has the incorrect expiration date (2015-10-01) on the label while the individual pouch labels and outer shipper have the correct expiration date (2018-10-01). Themanufacturing date was inadvertently used as the expiration date on the inner carton box labels. | Incorrect or no expiration date | Ecolab sent an "Urgent Medical Device Recall" letter dated November 30, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised to quarantine the product for return. If the product was further distributed notify them of the recall. Requested customers to complete and return the Recall Response Form. For questions contact Customer Service by phone at 1-800-824-3027 or e-mail customerservice@microtekmed.com.For further questions please call (651) 250-4385. |  | 40 | 40 cases (144 eaches per box) | US Distribution to the states of : CA, FL, MI, NJ, TX, UT and VA.. | Open, Classified |  |  | 40 | Not_Computer
Z-0766-2017 | 75645 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr PS Knee System Box Cut Guide |  |  | September 19, 2016 | 2016 | Encore Medical, Lp | It has been discovered that during Empowr PS Total Knee replacement, while chiseling for bone removal in the box area, femoral chipping has occurred in surgery. | Device Design | DJO Global sent an Urgent Field Safety Notice dated September 19, 2016, to all affected customers. The field safety notice provided a newly revised Surgical Technique to educate users on proper use of the device. Customers were instructed to become familiar with the new surgical techniques and notify all of their applicable users of the revised methods. Customers were also instructed to sign the Acknowledgement and Receipt Form. Customers with questions were advised to call 760-734-3551.For questions regarding this recall call 512-832-6302. |  | 162 | 162 units | Nationwide Distribution | Open, Classified |  |  | 162 | Not_Computer
Z-0758-2017 | 75663 | No | LZA | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Sterling Nitrile Sterile Powder Free Exam Gloves, S, Sterile. |  |  | November 10, 2016 | 2016 | Halyard Health | Defects: Product may be discolored and brittle resulting in the inability of the user to don the gloves without tearing. | Under Investigation by firm | Consignees were notified by letter via Fed Ex on/about 11/10/2016. |  | 59 | 59 cases | CA, FL, GA, IA, IN, LA, MA, MN, MT, NM, OR, TX and UT. | Terminated | May 04, 2017 | 175 | 59 | Not_Computer
Z-0765-2017 | 75675 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reprocessed Xcel Bladeless Trocar (H12LP) |  |  | November 11, 2016 | 2016 | Stryker Sustainability Solutions | An increase in reports indicating that the reprocessed H12LP Trocar adjustable plug may break in the packaging or during use. | No Marketing Application | Stryker sent an Urgent Medical Device Recall letter to all affected customers on November 11, 2016.Customers were instructed to discontinue use of the affected lots and were asked to follow the link below to complete the Recall Effectiveness Check Form and indicate if any devices from the lots referenced in Attachment 1 remain in their inventory. This form must be completed even if no affected product is found.www.novasyte.com/sss142Use the following reference number as you complete the form: [Ship to Account #].If the form indicates that any affected devices remain in your inventory, a prepaid shipping label will be issued for the return of the product. Credit for all affected devices returned will be issued to your facility.Address any questions regarding the recall to Novasyte by calling (760) 279-3180 or emailing sss142@novasyte.com. |  | 40476 | 40,476 devices | Worldwide Distribution - US (nationwide) and Internationally to Israel and CanadaForeign distribution to the following; IL and CA | Terminated | May 05, 2017 | 175 | 40476 | Not_Computer
Z-0764-2017 | 75735 | No | GDT | Removable skin staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | WECK Visistat 35W, Disposable Skin Stapler, Rx Only, Sterile, Teleflex Medical, NC 27709. |  |  | November 18, 2016 | 2016 | Teleflex Medical | Sterility Deficiency: The packages are not sealed properly and thus the sterility is not assured. | Packaging process control | The firm, Telexflex Medical, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 11/18/2016 to its Consignees/Customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any affected products; complete and return the Recall Acknowledgement Form by fax to 1-855-419-8507, Attn: Customer Service or email to: recalls@teleflex.com (even if you do not have any affected product) and return the product with the completed Recall Acknowledgement Form. A customer service representative will contact you with a RGA # and provide instructions for the return of the product.If you have any other questions, feel free to contact your local salesrepresentative or Customer Service at 1-866-246-6990. |  | 96402 | 96,402 units (9,120 ea US, 87,282 ea OUS) | Worldwide Distribution- U.S. (Nationwide) and countries of: Australia, Belgium, Canada, China, Germany, Hong Kong, India, Japan, Korea, Malaysia, New Zealand, Singapore and Thailand. | Open, Classified |  |  | 96402 | Not_Computer
Z-0761-2017 | 75743 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 15 monitor/defibrillator is intended for use by trained medical |  |  | March 11, 2010 | 2010 | Physio-Control, Inc. | Customer complaint was received relating to a LifePak15 device that would power on then power down unexpectedly after a short period of use. When the LP15 powered down, power could be immediately restored by activation of the power button. The device was connected to a cellular modem from MultiTech systems for use with LIFE NET services. The corrective action was completed by Velocitor Solutions in December 2010. | Environmental control | Velocitor initiated this field action and notified affected customers on March 11, 2010. This action was reported to Physio-Control as completed on 12/21/2010. |  | 1399 | 1399 units (866 in the US and 533 outside US) | worldwide | Completed |  |  | 1399 | Not_Computer
Z-0770-2017 | 74587 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE blanket LED Phototherapy system, neonatal phototherapy. Used to provide treatment for neonatal Hyperbilirubinemia. |  |  | October 28, 2016 | 2016 | Natus Medical Incorporated | neoBLUE blanket LED Phototherapy System distributed since release of the product in 2011 is recalled due to customer complaints regarding discoloration/ degradation of fiberoptic bundle at connection of pad and box. | Device Design | Natus began sending the URGENT RECALL NOTIFICATION / FIELD SAFETY NOTICE neoBLUE blanket system performance letter dated October 27, 2016 and the Technical Bulletin to US customers on October 28, 2016.Natus plans to notify their international customers on December 2, 2016. Customers were informed in March 2015 of a potential issue of discoloration/degradation/melting of the fiber optic bundle at the pad connector which is inserted in the neoBLUE blanket light box. Included with that letter was a Technical Bulletin which described the issue and addressed many of your potential questions.The purpose of this current letter (dated October 27, 2016) is to communicate Natus' plan for a follow-on field correction for neoBLUE blanket systems shipped between 2011 through March 21, 2016 (Serial Numbers up to and including xxxx004282), and to remind customers of the instructions contained in the neoBLUE blanket Technical Bulletin.Natus is in the process of confirming an updated neoBLUE blanket configuration which will not be susceptible to the degradation described above. However, the date of project completion is uncertain at this time. Natus will contact customers when the new configuration is available, at which time Natus intend to initiate a field corrective action to replace or upgrade your device(s).Please contact Natus Quality_Programs@natus.com if you have any further questions. Please call (800) 426-0353. |  | 2619 | 2619 units (1925 units in the US and 694 units OUTSIDE the US) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : including American Samoa and Guam, and in the following countries:Albania, Argentina, Australia, Austria, Bangladesh, Canada, Chile, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, Indonesia, Iraq, Israel, Italy, Korea, Kuwait, Latvia, Malaysia, Mongolia, Namibia, Netherlands, Nigeria, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, and Virgin Islands (British). | Open, Classified |  |  | 2619 | Not_Computer
Z-0768-2017 | 75576 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Data Management System. |  |  | October 25, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has received and confirmed reports of an intermittent and unexpected behavior when loading new panels into a WalkAway instrument usingLabPro Data Management System version 4.42. The issue could cause workflow interruption with a potential of delaying reporting results due to the inability to begin processing new panels. | Device Design | An Urgent Medical Device Recall letter dated 10/25/16 was sent to all their customers to inform them that Beckman Coulter has received and confirmed reports of an intermittent and unexpected behavior when loading new panels into a WalkAway instrument using LabPro Data Management System version 4.42. Customers are informed that the issue could cause WalkAway workflow interruption with a potential of delaying reporting results due to the inability to begin processing new panels. The letter informs the customers of the actions to be taken and the resolution for the product on recall. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the notice are instructed to contact their Customer Support representative: http://www.beckmancoulter.comor call 1-800-677-7226, option 1 in the United States. |  | 131 | 131 units total (130 units in US) | Worldwide Distribution -- United States and Mexico. | Terminated | May 01, 2017 | 188 | 131 | Computer
Z-0767-2017 | 75571 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | IVEA 500A |  |  | November 01, 2016 | 2016 | Firefly Medical, Inc. | The dowel pin could potentially slip out of its press fit, resulting in the product potentially slumping to one side which could cause the patient to stumble. | Process control | Firm notified customers on **11/18/16** via certified "Urgent: Medical Device Recall (Field Correction) IVEA Model 500A" letter to correct product in the field. Customers were asked to thoroughly read the letter and to immediately follow the actions in the enclosed Product Correction Kit. The instructions include steps for cleaning the correction area and installing the permanent adhesive caps, as well as, completing and returning the product correction form to Firefly Medical. Customers may continue to use their units until the Product Correction actions are complete, as long as pins are not protruding. If the pins are protruding, then customers are to stop using the device immediately and complete the product correction instructions. Customers can call 970-472-5323 or e-mail info@iveamobility.com if they have any questions. |  | 239 | 239 | Worldwide Distribution -- USA and Canada. | Terminated | August 01, 2017 | 273 | 239 | Not_Computer
Z-0769-2017 | 75228 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Updated Natus neoBLUE blanket LED Phototherapy Systems. It provides phototherapy underneath the infant and can be used in a bassinet, open bed, radiant warmer, incubator, or while holding the infant. |  |  | September 16, 2016 | 2016 | Natus Medical Incorporated | The firm sent the Technical Bulletin to consignees who received the reliability updated neoBLUE blanket LED Phototherapy Systems between March 2016 to August 2016. The reliability update to prevent the early failure of the fiberoptic pads experienced by some customers on older systems. These failures on previous systems involved discoloration and melting of the fiberoptic bundle at the connector that is inserted in the neoBLUE blanket light box. | Device Design | Natus sent the Urgent Recall Notification/Field Safety Notice and Customer Reply Form (DOC-021042A), dated September 16, 2016 and the Technical Bulletin (DOC-021016) to US consignees via email.The letters and technical bulletin for foreign consignees was sent via email to the appropriate Natus Distribution Partner (DP) on November 7, 2016.US Customers can email the customer reply form to Natus_Quality_Progams@natus.com or fax to (847-949-6031). US Customers with questions can call Natus Technical Service at 888-496-2887. International customers should contact their local distributor with any questions. |  | 713 | 713 units total (403 units in US and 310 units outside US) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, NC, ND, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI, WV and Puerto Rico. , and to the countries of : Argentina, Austria, Canada, Chile, China, Denmark, Dominican Republic, Finland, France, Germany, Hong Kong, Hungary, Indonesia, Italy, Latvia, Malaysia, Namibia, Paraguay, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Taiwan, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, Viet Nam, and Virgin Islands (British). | Open, Classified |  |  | 713 | Not_Computer
Z-0779-2017 | 75708 | No | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DELFIA hFSH kit. In Vitro Diagnostic for the determination of hFSH in serum |  |  | November 16, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | A DELFIA hFSH kit lot failed real time stability testing. The real time stability data demonstrated a decrease in hFSH concentrations between -4.8% to -11.3% for control sera. The -11.3% bias was observed at an hFSH target concentration of 101 IU/L. The change observed in the serum samples within the range 1.76  67.0 IU/L hFSH averaged -2.9%. The reference range of the assay spans 0.60 to 98.0 IU/L. | Other | PerkinElmer initiated recall on November 16, 2016, delivered to the end customer by email facilitated through the PerkinElmer Technical Support group located in Akron Ohio. Technical Support will be asked to make sure the customers receive the Recall Letter and understand the recommended actions.Users asked to inspect inventory, discontinue use destroy the kits and complete the Response Forms and Destruction.Replacement kits will shipped. |  | 79 | 79 kits | CA, CO, NY, WAForeign:Austria, Bolivia, Bosnia , Sweden, Norway, Netherlands, Japan, Great Britain, | Terminated | July 05, 2017 | 231 | 79 | Not_Computer
Z-0772-2017 | 75420 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS Percutaneous Introducer Set under the following labels: 1) 5 F Stiffened Micro-Introducer Kit, 45CM SS/SS ECHO B PG, Item Number H787065970395, Catalog No.REF 06597039, 2) 5 F Stiffened Micro-Introducer Kit, STF 5F M.I. KIT 45CM SS/SS ECHO S PG, Item Number H787065970415, Catalog No. REF 06597041 |  |  | September 28, 2016 | 2016 | Angiodynamics, Inc. | Greatbatch Medical the manufacturer of the 4F and 5F sheath/dilator components has determined that the products listed in their September 7, 2016 Recall Notification have the potential for the rotating luer to detach from the dilator hub during use. AngioDynamics has confirmed that affected sheath/dilators (Greatbatch Model Numbers 10904-001 and 10904-002) have been included in packaged AngioDynamics Stiffened Micro-Introducer Kits. | Component design/selection | Angiiodynamics sent an Urgent Voluntary Medical Device Recall letter dated September 28, 2016, to all affected customers. Recall notifications were delivered by Federal Express. Customers were instructed to segregate and return all affected devices to AngioDynamics and to complete and return the Reply Verification Tracking Form, provided in the recall notification by fax to 1-800-782-1357. Customers were encouraged to forward a copy of the recall notification to all sites to which have distributed affected product. If affected product is located in your institution, please call AngioDynamics Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday  Friday: Eastern Standard Time) to obtain a replacement or credit for your returned product. Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on this page and the Reply Verification Tracking Form via to recall@angiodynamics.com. Customers with questions should contact customer service at 1-800-772-6446. |  | 162 | 162 boxes | Nationwide Distribution to AL, AR, CA, FL, GA, IL, KY, LA, MI, MN, MO, NC, OH, OK, SC, NY, ME, PA, MD, VA, WA, WI, WV, TN, and TX | Open, Classified |  |  | 2145 | Not_Computer
Z-0780-2017 | 75588 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PerkinElmer 20ul Clear RoboRack Tip, Pipette (Qty 960) |  |  | October 13, 2016 | 2016 | PerkinElmer, Inc | Defective tips have the potential to pipette/dispense volumes 50-100% lower than is specified for the tips. | Process control | PerkinElmer initiated the recall on Sept 1. 2016 by letter dated . Customers were told of the defect, potential for inaccurate measurements, and requested to inspect for the lots stop using and destroy. Replacement product to be provided. Accounts asked to compelte Response form. Questions call your PerkinElmer representative. |  | 33312 | 33312 | IA, MD, NY, WAForeign: France, Germany, The Netherlands, Hungary, Singapore | Open, Classified |  |  | 33312 | Not_Computer
Z-0771-2017 | 75781 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH, Pocket PAC IC Intermittent Cath Kit, 8 FR, 12 FR, 14 FR, 10 FR and 16 FR, Rx Only, Sterile, |  |  | November 22, 2016 | 2016 | Teleflex Medical | Labeling Inconsistency: The products have beenlabeled with the incorrect expiration date which exceeds the actual, validated 35-month shelf life. | Error in labeling | Teleflex sent an Urgent Medical Device Recall Notification letter dated November 22, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers with affected stock, immediately discontinue use and quarantine any affected products. To return product, customers were advised to complete the enclosed Recall Acknowledgment Form and fax it to 1-855-419-8507, Attn: Customer Service or email to recall@teleflex.com. Customers with no affected stock should also complete the enclosed Recall Acknowledgment Form to confirm receipt of the letter. Customers with questions should contact Customer Service at 1-866-246-6990. |  | 102843 | 102,843 units | Nationwide Distribution to CA, CO, FL, GA, IL, KS, KY, MA, MO, MS, NC, NJ, NY, OH, OK, OR, PA and TX. | Open, Classified |  |  | 102843 | Not_Computer
Z-0786-2017 | 75794 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AbbVie PEG Kit 20 FR |  |  | November 22, 2016 | 2016 | AbbVie Inc. | AbbVie is issuing an urgent field correction notification due to an increased number of reports of disconnections of the AbbVie 20FR PEG tube from the Y-connectors. | Device Design | The firm sent an URGENT: MEDICAL DEVICE FIELD CORRECTION NOTIFICATION AbbVie PEG Kit 20 FR to affected customers describing the reason for the notification. The letter asks affected customers to forward this letter to all Movement Disorder Specialists and physicians (e.g. GIs, IRs, and surgeons) that place PEG- J Tubes within your facility.The letter describes actions to be taken for patients still using the AbbVie PEG 20FR Tube. The letter requests user facilities to continue to monitor the performance of the Tube. The firm states that there is no need to replace the AbbVie PEG 20FR Tube if it is functioning as expected and maintaining a viable connection with the Y-Connector. The letter provides the following instructions when establishing a connection between Abbvie PEG 20FR Tube with the 20FR Y-connector, always refer to AbbVie J Intestinal Tube 9 FR to PEG 15 and 20 FR Instructions for Use for 62918-001. The letter notes the following:1. Ensure that the set-up process clearly follows the sequence of events as described in the instructions for use and the following clarifications. 2. Ensure that the AbbVie PEG 20FR tube is connected to the corresponding 20 Y- connector (see photograph below which shows the difference between the 15FR and 20FR connection diameter). 3. DO NOT USE ALCOHOL CONTAINING PRODUCTS with the AbbVie PEG 20 FR tube as this may result in damage to the tube.4. Ensure that the Tube Clamp and Tube Clip on Fixation Plate are both closed. 5. Cut off the proximal cone and Fixation Loop of the AbbVie PEG Tube (15FR or 20FR). Note: Ensure that the cut is perpendicular (90 degree angle) to the tube and not on a diagonal as this ensures proper seating of the tube when engaged into the pin of the Y-connector fitting. 6. Slide the colored Fixation Screw (including the outer ring) over the Abbvie PEG Tube:a. For 15 FR, use the Blue/White screwb. For 20 FR, use the Violet/White screw7. Grasp near the |  | 541 | 541 units | U.S. consignees located throughout the country and Puerto Rico. Foreign Consignees: Worldwide. | Open, Classified |  |  | 541 | Not_Computer
Z-0782-2017 | 75055 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultroid Hemorrhoid Management Systems 110/220 VAC and procedures kit including the Ultroid Disposable Sterile Probes |  |  | October 19, 2016 | 2016 | Ultroid Technologies, Inc. | Products are not designed, qualified, manufactured, and/or managed under a state of control per internal quality system procedures and external laws, regulations, directives, standards, and/or guidance (such as 21 CFR 820 Quality System Regulation) | Process control | On October 19, 2016 Ultrid sent a recall letter to all their consignees informing them of the recall and to cease further distribution of this product, respond the self-addressed response form and returned for a voucher. For more information contact Melanie Rosario at 727-898-0717 |  | 1129 | 1,129 | AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, KS, LA, MD, ME, MI, MN, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA and Puerto RicoAsturias, Australia, Chile, Greece, Hong Kong, India, Indonesia, Iraq, Japan, Kuwait, Malaysia, Mexico, Nigeria, Qatar, Russia, Slovenia, Spain, Syria, Trinidad, Turkey, UAE, UK, and Vietnam | Open, Classified |  |  | 61059 | Not_Computer
Z-0781-2017 | 75459 | No | LGW |  |  | Neurology | PMA | Algovita Spinal Cord Stimulation System, Percutaneous Leads, 8-Electrode Percutaneous Lead Kit; |  |  | October 11, 2016 | 2016 | Nuvectra | Nuvectra is conducting a recall due to two clinical risks that are associated with certain percutaneous leads currently being used in the Algovita Spinal Cord Stimulation System.It has been reported that in some percutaneous leads, the stylet protruded beyond the distal lead tip. | Nonconforming Material/Component | Nuvectra field sales associate or contracted distributor were sent a Nuvectra Urgent Field Safety Notice (Field Action Number: 2016-03) on 10/11/2016 via Fed-Ex. The letter provided background information, described the problem, &amp; the product involved in the recall. The letter listed the potential clinical implications, rate of occurrence, recommendations and mitigations and action needed. Provided consignees with notices for them to deliver to physicians as well as Response forms for them and the physician. For questions contact Kathy Jo Fahery +01.763.951.8219. |  | 1391 | 406 US, 985 OUS | US: Nationwide (CA, TX, KY, FL, IL,PA,OH, GA, NJ, NC, KS,AR, CO, NE, NY, VA, AZ, OK).OUS: Germany | Open, Classified |  |  | 1391 | Not_Computer
Z-0787-2017 | 75816 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-651, 120" Microbore extension set with slide clamp; Product Code: AMS-651 |  |  | November 18, 2016 | 2016 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | Incorrect priming volume is printed on the device package. | Device Design | Vygon sent an Urgent Medical Device Recall letter dated November 18, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to take the following actions: 1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately; 2. Count and document your affected inventory on the "Recall Acknowledgement and Inventory Return Form; 3. Fax or email the form to Vygon including customer contact information; 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number; 5. Return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up; 6. All inventory will be credited or replaced at no cost to you. For questions contact local sales representative or Vygon's Customer Service Department at 1-800-473-5414.or by email at customerservice@vygonus.com. |  | 31500 | 31,500 | US distribution in the state of CA | Open, Classified |  |  | 31500 | Not_Computer
Z-0788-2017 | 75486 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II Implantable Drug Infusion System. This system includes: Model 8870 Software Application Card, Model 8840 N'Vision Clinician Programmer, and Model 8637 SynchroMed II Implantable Drug Infusion Pump (supplied in 20 ml or 40 ml reservoir size.) |  |  | October 03, 2016 | 2016 | Medtronic Neuromodulation | Medtronic is following up to a May 2013 communication regarding the Priming Bolus function for the SynchroMed Infusion System. Medtronic is updating the Model 8870 software application card (to version AAU01) and the SynchroMed pump labeling to address the priming bolus issue. | Under Investigation by firm | Medtronic Representatives began visiting consignees on 10/3/2016 to replace previous version Model 8870 software application cards, contained within the model 8840 NVision Clinician Programmer, with the updated version (new version is AAU01). Beginning October 20, 2016, Implanting and Following physicians were sent an Urgent Medical Device Correction letter dated September 2016, . This letter to physicians is informing them of the updated Model 8870 software application card, and changes to SynchroMed infusion system labeling. A response form is requested to be returned to Medtronic.Customers with questions were instructed to contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 22298 | 22,298 software cards | Worldwide Distribution - (US Nationwide) Internationally to Germany, Greece, Australia, United Kingdom, Switzerland, Italy, Norway, France, Canada, and Malta. | Open, Classified |  |  | 22298 | Computer
Z-0825-2017 | 75844 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia MR-OR; 781437, 781438 |  |  | December 01, 2016 | 2016 | Philips Electronics North America Corporation | Possibility that a patient or user may be exposed to a very small level of touch current. | Unknown/Undetermined by firm | Philips sent an Urgent Field Safety Notice Medical Device Correction letter dated November 2016 to all affected customers. The letter explains the issue and actions planned by Philips. There is no need for action to be taken by customer or user. Customers with questions should contact their local Philips representative Technical Support Line 1-800-722-9377. |  | 9 | 9 | Worldwide Distribution -US including CA, NY and Internationally to United Kingdom, Australia, Norway, and Thailand | Open, Classified |  |  | 9 | Not_Computer
Z-0789-2017 | 75742 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | AcuvanceÂż2 Safety Catheter, Product No. 1720-AI |  |  | November 21, 2016 | 2016 | Smiths Medical ASD Inc. | Smiths Medical became aware that three (3) lots of Acuvance2 Safety Catheter, product number 1720-AI, contain devices with missing or incorrectly positioned side port valves. | Process control | Consignees were sent on 11/21/2016 a Smiths Medical "Urgent Medical Device Field Safety Notice: dated November 18, 2016. The letter was addressed to "Clinician who uses the AcuvanceÂż2 Safety Catheter". The letter described the Reason for Recall, Risk to Health, and Instructions to Customers. Requested consignees to complete and return the "Urgent Medical Device Field Safety Notice Response Form" within 10 days of receipt. For questions contact Stericycle via email at the smithsmedicalrecall@stericycle.com. |  | 26750 | 26,750 | Foreign Distribution in the Italy, France, Switzerland, Germany, Sweden | Terminated | August 01, 2017 | 253 | 26750 | Not_Computer
Z-0821-2017 | 75892 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tibial Component Metal Cased Small Long; Tibial Component Metal Cased Small Short; Tibial Component Metal Cased Std Long; Tibial Component Metal Cased Std Short |  |  | December 07, 2016 | 2016 | Stanmore Implants Worldwide Ltd. | The recalling firm has received 5 customer inquiries reporting that at least one pouch of the double pouch packaging configuration has been punctured by the tip of the device's stem. | Packaging | Stanmore Implants sent an Urgent Field Safety Notice dated November 25, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately check their internal inventory and quarantine all subject devices, circulate the Field Safety Notice internally to all interested/affected parties, maintain awareness of the notice internally until all required actions have been completed within their facility, inform Stanmore Implants if any of the subject devices have been distributed to other organizations, complete the attached customer response form and list of affected batches form, return the completed form and any affected devices to their S/W Representative. Customers with questions were instructed to contact their local sales representative. |  | 51 | 51 units | Nationwide Distribution to OH, NE, FL, LA, NY, TX, MS, MD, CT, AL, PA | Terminated | July 17, 2017 | 222 | 105 | Not_Computer
Z-0817-2017 | 75785 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | MEDTRONIC ARRIVE BRAIDED TRANSSEPTAL SHEATH 10F |  |  | October 05, 2016 | 2016 | Oscor, Inc. | Complaint that during insertion of the dilator through the sheath a fragment of inner liner came out of the sheath prior to use. | Nonconforming Material/Component | Oscor sent two(2) Product Recall notifications:-HOSPITALS: Immediately check your inventory to confirm that you do or do not have units from these lot(s) in your possession. Identify and set aside any unit from the identified lot in a manner that ensures the affected product will not be used. Check all storage and usage. Review, complete, sign and return the enclosed Acknowledgement Form attached in the letter, directly to Oscor Inc. at the fax number or e-mail on the form, for immediate replacement (as applicable).Who to call: Please call Oscor Inc. Customer Service Department at 727-937-2511 or email RGA@Oscor.com Hours of operations: Monday to Friday from 8:30AM to 5:30PM Eastern Time.- DISTRIBUTORS: Immediately check your inventory to confirm that you do or do not have units from these lot(s) in your possession. Identify and set aside any unit from the identified lot in a manner that ensures the affected product will not be used. Check all storage and usage. Please pull a list of the hospitals (end user) impacted by the affected lots and retrieve inventory for replacement, if available. Review, complete, sign and return the enclosed Acknowledgement Form (Exhibit A) attached in this letter, directly to Oscor Inc. at the fax number or e-mail on the form, for immediate replacement (as applicable). Who to call: Please call Oscor Inc. Customer Service Department at 727-937-2511 or email RGA@Oscor.com Hours of operation: Monday to Friday from 8:30AM to 5:30PM Eastern Standard Time. |  | 991 | 991 | Worldwide Distribution - US including CA, NJ, MD, TX, FL, NC, MO, and Puerto Rico. Internationally to AUSTRIA, BOLIVIA, CANADA, CHILE, CYPRUS, NETHERLANS, EGYPT, GERMANY, HONG KONG, ISRAEL, ITALY, SWITZERLAND, and VIETNAM | Open, Classified |  |  | 7361 | Not_Computer
Z-0828-2017 | 75761 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i CERTAIN LOW PROFILE 17Âż ABUTMENT 3.4MM(D) X 2MM(H) |  |  | September 14, 2016 | 2016 | Biomet 3i, LLC | A limited number of product pouches may not have been sealed during packaging. | Employee error | BIOMET 3i will conduct this recall with Dentists and Health Care Professionals. Notices will be sent via a carrier with a traceable method of delivery by BIOMET 3i team members." All dentists and health care professionals will be notified via courier: - Dentists and health care professionals will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the product in their control. - Dentists and health care professionals will return on-hand product to BIOMET 3i via a courier agency and ensure all of their products are accounted for using the form provided in the letter. - Dentists and health care professionals will complete and return the Inventory Return Certification Form to the fax number or email address provided.International distributors will be notified via email. Further distribution of the international customer notices will be sent through the local and regional distributors following local language translation. Communications will occur as translations are completed and are anticipated to occur during the week of September 12, 2016.1233. Customers with questions were instructed to call the customer call center at +44 (0) 800 652-1233. |  | 127 | 127 | Worldwide Distribution to Austria, Belgium, Switzerland, Germany, Spain, France, Great Britain, Ireland, Italy, Lithuania, Portugal and Sweden | Open, Classified |  |  | 9363 | Not_Computer
Z-0826-2017 | 75795 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Affinity Fusion" Oxygenator with Integrated Arterial Filter, |  |  | November 16, 2016 | 2016 | Medtronic Inc | Medtronic is initiating a voluntary product recall for specific lot numbers of the Affinity FusionÂż Oxygenators with BalanceÂż1 Biosurface. These are distributed as stand-alone devices, or as a combination unit with the Affinity Fusion Cardiotomy/Venous Reservoir, or as part of tubing packs. Medtronic has identified an out-of-specification condition exhibiting additional plastic (flash) in the arterial sampling port adjacent to the arterial outlet port of the oxygenator. | Process control | The firm, Medtronic, sent an "Urgent Medical Device Recall" letter dated November 2016 to its Consignees on 11/16/2016. The letter was addressed to Risk Manager or Healthcare Professional. The letter described the problem and the product involved in the recall and actions to be taken. The consignees were instructed to quarantine all unused affected product in your inventory, return the affected product (contact Medtronic Customer Service at 800-848-9300 to initiate a product return) and complete &amp; return the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or scan and email to RS.CFQFCA@medtronic.com. For questions contact your Medtronic Representative. If you require assistance in contacting your representative, please contact Medtronics National Answering Service at 800-633-8766. |  | 5609 | 5609 (663 US; 4946 OUS) | Worldwide Distribution-US (nationwide) in states of: AK, AZ, CA, CO, CT, FL, IA, IL, IN, KS, KY, LA, MN, ND, NV, NY, SD, TN, TX, WI, WY and countries of: Canada, Chile, Croatia, Germany, Greece, Hungary, Iran, Italy, Latvia, Lebanon, Poland, Portugal, Romania, Spain, Tanzania, and United Kingdom. | Terminated | August 01, 2017 | 258 | 5609 | Not_Computer
Z-0827-2017 | 75855 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Zeus-P Lumbar Interbody Device |  |  | November 22, 2016 | 2016 | Amendia, Inc | Labeling error where three boxes in inventory had conflicting patient and primary labels. Evaluation indicated that the primary label and patient label did not match. The primary label matched the label on the inner pouch, but the patient label was incorrect. | Error in labeling | Amendia Inc. sent an Urgent - Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand.For further questions please call (404) 610-7215. |  | 25 | 25 units | US Distribution to the states of : GA. MI, OK, SC, TX, VA | Terminated | June 05, 2017 | 195 | 25 | Not_Computer
Z-0864-2017 | 75692 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | VITROS CRP Slides: 1) VITROSÂż Chemistry Products CRP Slides (250 slides per sales unit), Catalog # 1926740 and 2) , VITROSÂż Chemistry Products CRP Slides (90 slides per sales unit), Catalog # 1926740 |  |  | October 26, 2016 | 2016 | Ortho-Clinical Diagnostics | CRP samples diluted using VITROS Chemistry Products Specialty Diluent, VITROS Chemistry Products FS Diluent Pack 3, or a sample with a low concentration of CRP had positively biased results if the endogenous level of CRP in the diluent is not subtracted from the final result. The IFUs do not provide specific instructions for how to account for endogenous CRP in the diluent when calculating the sample result. Orthos diluents used to dilute samples containing CRP are manufactured using human plasma based fluids that contain endogenous CRP. Without accounting for the CRP in the diluent and removing its contribution from the result, the reported result will include the CRP present in both the sample and diluent, possibly leading to positively biased results. | Process control | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated October 26, 2016., via FedEx overnight courier and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who do NOT use ORTHO PLUS e-Communications and received CRP slides within the last 12 months to inform them of the possibility of positively biased CRP results when using VITROS CRP Slides with diluted samples and to provide them with original instructions for how to perform dilutions when processing CRP samples with concentrations above the measuring range. On 1 November 2016, Ortho Clinical Diagnostics issued customer letter (Ref. CL2016-209) sent via ORTHO PLUS e-Communications to all customers who use ORTHO PLUS e-Communications and received CRP slides within the last 12 months to inform them of the possibility of positively biased CRP results when using VITROS CRP Slides with diluted samples and to provide them with two revised instructions on how to perform dilutions when processing CRP samples with concentrations above the measuring range. Foreign affiliates were informed by email on October 26, 2016, of the issue and instructed to notify their consignees of the issue. This included the original instructions for how to perform dilutions when processing CRP samples with concentrations above the measuring range. On 1 November 2016, two revised Instructions for CRP Dilutions enclosures (Ref. CL2016-209) were sent via FedEx overnight courier and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received CL2016-209 via FedEx overnight courier or US Postal Service Priority Mail that provided additional detailed instructions on how to perform a dilution for samples exceeding the measuring range. If you have further questions, please contact our Ortho Care Technical Solutions Center at 1-800-421-3311. |  | 124063 | 124,063 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico, and to the countries of : Bermuda, Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 124063 | Not_Computer
Z-0865-2017 | 75891 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery Systems, models: G3, G4, VSI. |  |  | December 06, 2016 | 2016 | Accuray Incorporated | Accuray has become aware of a potential safety issue involving the possibility of electric shock duringmaintenance activities from the lon Pump HV Power Supply at the back of the modulator cabinet. This issue is limited to certain G3, G4, and VSI models of the CyberKnifeÂż Robotic Radiosurgery System. | Device Design | Accuray Incorporated sent an Urgent Safety Advisory Notification letter to all affected customers on December 1, 2016 by traceable method.. Letters identified the product, the problem, and the action to be taken by the customer. provided users with affected product and described the issue. The final resolution will result in a modulator cabinet inspection and installation of the Plexiglas cover - as necessary. Customers will be contacted by Accuray to schedule the inspection. Prior to this modulator cabinet inspection, inform all applicable personnel of the potential danger and need to closely follow electrical safety procedures. Customers with questions were instructed to contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Tehnical Solution Center at http://www.accuray.com/services-support or call:North America +1-877-666-8667, Europe +800-4141-9595, Hong Kong +852-2247-8688, Japan +81-3-6265-1530. Customers were also asked to complete the attached acknowledgement and return to Accuracy.For questions regarding this recall call 608-824-3362. |  | 115 | 115 affected devices. | Worldwide Distribution - US including CA. MN, OK. FL. MO, TN, KS, TX, AZ, MS, IA, MA. NC, NY. IN, NJ, VA. LA, WA, WI, PA. CT, G,. MT, MD, LA., WV, OH, and DC.Foreign distribution to Taiwan, Japan, Korea, Turkey. Malaysia, Vietnam, Greece, China, Spain, France, Italy, Pakistan, and Georgia, | Terminated | June 09, 2017 | 185 | 115 | Not_Computer
Z-0877-2017 | 75886 | No | MKQ |  |  | Pathology | PMA | BD SurePathTM Preservative Collection Vials, catalog numbers 490505, 490527, 491095, 491096, and 491097. |  |  | February 12, 2016 | 2016 | Becton Dickinson &amp; Co. | BD has confirmed that a portion of BD SurePath" collection vials associated with specific production lots may contain caps that are cracked which can lead to leaking of the preservative fluid. | Nonconforming Material/Component | BD Life Sciences sent the initial recall notification on February 12, 2016. The distributors were provided with a written communication via email and customers were provided with a written communication via UPS.For further questions, please call (410) 316-4000. |  |  | 16,682.396 | Worldwide Distribution - USA (nationwide) Distribution including BB(Barbados) and PR (Puerto Rico) and to the states of : AK, AL, AR, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and to the countries of: China, Europe, Hong Kong, India, Korea, Mexico, New Zealand, Panama, Singapore, and Taiwan.480 US customers and 3 US distributors. Product was also shipped to 10 ex-US regions. 15 US government sites received affected product | Completed |  |  |  | Not_Computer
Z-0878-2017 | 75909 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge OrthoCase software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | October 08, 2015 | 2015 | Merge Healthcare, Inc. | Measurements in the software are changing after saving a plan with a measurement result, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge issued a letter dated March 28, 2016, via email on April 4, 2016, or via certified mail if they did not have an email address for the customer. The letter notifies the customer of the issue, informs them of a workaround and that a fix is not yet available but they will be notified once one is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016.For further questions, please call (877) 741-5369. |  | 30 | 30 sites potentially have the effected versions | Distribution was made to medical facilities in CA, CO, CT, MA, MD, MN, MT, NE, NV, NY NJ, PA, SC, WA, and WI. Foreign distribution was made to Canada, as well as other countries. | Open, Classified |  |  | 30 | Computer
Z-0867-2017 | 75828 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes VA Implant Rack for the Compact Distal Radius System |  |  | November 29, 2016 | 2016 | Synthes (USA) Products LLC | The screw rack plate was assembled incorrectly which could cause the screws in the screw rack to sit low and touch the surface beneath the screw rack. | Process control | DePuy Synthes mailed an Urgent Notice Medical Device Recall letter to all affected customers informing them of the potential impacts, actions to be taken and instructions for the product return. Customers were instructed to immediately review their inventory to identify and quarantine all affected products, complete the attached Verification Section and send a copy to Depuy Synthes by scan/email: Synthes5748@stericycle.com or fax to 1-888-943-5169. If any of the affected product has been forwarded to another facility, contact the facility to arrange return. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. |  | 31 | 31 units | Nationwide Distribution | Terminated | August 03, 2017 | 247 | 31 | Not_Computer
Z-0866-2017 | 75716 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Insufflator, laparoscopic |  |  | November 21, 2016 | 2016 | Northgate Technologies, Inc. | Device could cause a patient overpressure situation without any visual or auditory indication or warning | Software design | Northgate Technologies sent an Urgent Medical Device Recall Notification letter dated November 17, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to examine their records and ascertain the location of affected devices, cease distribution or use and quarantine those products appropriately. Ship all affected products in their possession to Northgate Technologies, Inc., 1591 Scottsdale Court, Elgin, IL 60123. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 847-608-9405. If product was further distributed, those customers should be notified immediately. Customers with questions should call 224-856-2253. For questions regarding this recall call 800-348-0424. |  | 49 | 49 devices | Worldwide Distribution - US including FL and MA. Internationally to Canada | Open, Classified |  |  | 49 | Not_Computer
Z-0873-2017 | 75819 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | 8450-009-030 3.0MM Precision Round, 13CM, single use, Sterile |  |  | November 30, 2016 | 2016 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling specific lots of sterile MIS Toolsteel Burs because the burs may exhibit signs of corrosion. | Storage | The firm, Stryker Instruments, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated November 2016 to its Customers Overnight service, FedEx on November 30, 2016. The letter described the product, problems and actions to be taken. Stryker Sales Representatives and Stryker International will be notified via email. Customers are instructed to do the following: 1. Immediately review the Recall Notification.2. Check all stock areas and/or operating room storage to determine how many affected Burs from the affected lots are at your facility. Quarantine and discontinue use of the recalled Burs.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of the notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you dont have anyaffected product on hand. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to StrykerInstrumentsRecalls@stryker.com.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification.4. If you have further distributed this product, please forward this letter and the attached BRF to all affected locations. Please indicate each location on the BRF.5. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to StrykerInstrumentsRecalls@stryker.com.6. If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. Upon receipt of the recalled product, a credit will beissued to your account.Recalled products will be destroyed, packaging will be defaced, opened and discarded. Product will be discarded.For questions regarding this recall, please contact Stryker Instruments: StrykerInstrumentsRecalls@stryker.com, or |  | 1257 | 1,257 in total | Worldwide Distribution - US (nationwide) including states of: AK, AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI and countries of: Canada, Japan, Venlo (Netherlands), France, and UK (United Kingdom). | Terminated | June 01, 2017 | 183 | 1257 | Not_Computer
Z-0897-2017 | 75862 | No | PJI |  | Immunology | Immunology | 510(k) | EnLite Neonatal TREC Kit;an in vitro diagnostic device intended for the semi-quantitative determination of TREC (T-cell receptor excision circle) DNA in blood specimens dried on filter paper. The test is for use on the VICTOR" EnLite instrument. The test is indicated for use as an aid in screening newborns for severe combined immunodeficiency disorder (SCID). Product Number: 3401-001U |  |  | December 05, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | Unique Device Identifier (UDI) is missing from outer kit box label. | Employee error | Customers were notified of the recall via email on December 5, 2016. Technical Support will ensure the customers receive the Recall Letter and understand the recommended actions. Customers were asked to retain this Medical Device Correction in your laboratory documentation until the affected lot inventory is either depleted or reaches expiration and to complete the enclosed response form and return it by fax to number 1-330-825-8520 (United States) /+358 2 2678 357 (outside United States) or as scanned by e-mail to TurkuQMresponse@perkinelmer.com as soon as possible, but not later than December 29, 2016. |  | 73 | 73 kits | MA, AK | Open, Classified |  |  | 73 | Not_Computer
Z-0881-2017 | 75858 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | K-Wire 1.4mm x 12in Non-Threaded, Part Number 9080-12U; K-Wire Threaded 1.4mm x 18, Part Number 9080-18T; K-Wire Non-Threaded 1.4mm x 18, Part Number 9080-18U; 1.4mm x 24 Threaded Guide Wire, Part Number 9080-24T; K-Wire Non-Threaded 24, Part Number 9080-24U; Blunt K-Wire, 18 inches, Threaded, Part Number 9080B-18T; Blunt K-Wire, 18 inches, Unthreaded, Part Number 9080B-18U; Blunt K-Wire, 24 inches, Unthreaded, Part Number 9080B-24U; 1.4mm x 18 Blunt Threaded Nitinol Guide Wire, Part Number 9080B-N-18T; 1.4mm x 18" Blunt Non-Threaded Nitinol Guide Wire, Part Number 9080B-N-18U; Lodestar K-wire Non-Threaded, 18, Part Number 9080L-18U; 1.4mm x 18 Threaded Nitinol Guide Wire, Part Number 9080-N-18T; 1.4mm x 18 Non-Threaded Nitinol Guide Wire, Part Number 9080-N-18U |  |  | November 22, 2016 | 2016 | Amendia, Inc | The rods are packaged in packages with unacceptable seals. In some cases, the integrity of the seal is compromised by channels or voids. | Employee error | Amendia sent an Urgent Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter requested that customers examine their inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand. Customers with questions were instructed to call the following numbers:770-575-5224770-575-5236770-575-5369For questions regarding this recall call 404-610-7215. |  | 5833 | 5833 units | Nationwide Distribution to AL, AZ, AR, CA, CO, FL, GA, IA, KS, KY, LA, MA, MI, MN, MO, NV, NJ, NC, SC, TN, TX, VA, WA, WV, WI and Puerto Rico | Open, Classified |  |  | 7151 | Not_Computer
Z-0891-2017 | 75856 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Optimus Fixed Awl |  |  | November 22, 2016 | 2016 | Amendia, Inc | Complaints that the tip of the awl broke after impaction during surgery. | Device Design | Amendia Inc. sent an Urgent : Medical Device Recall letter dated November 22, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand.For further questions, please call (404) 610-7215. |  | 9 | 9 devices | US Distribution to the states of : FL, GA. IL, MO, TX | Open, Classified |  |  | 9 | Not_Computer
Z-0879-2017 | 75758 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics PediNail Intramedullary Nailing System |  |  | November 18, 2016 | 2016 | OrthoPediatrics Corp | Labeled as a 48mm Cortical Screw but measuring at 50mm | Labeling mix-ups | OrthoPediatrics sent an URGENT MEDICAL DEVICE RECALL notification on November 18, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. . It is important that you carefully follow these recall instructions:" Please immediately take possession of any of these screws your customer may have or that you have, and quarantine them. Contact logistics@orthopediatrics.com to receive a Return Authorization Number. It is very important that we contain these devices rapidly and maintain them in quarantine. Replacement devices can be ordered from Customer Service." The FedEx package you receive will contain a pre-addressed FedEx return label for the return of the recalled screw(s)." Included in this FedEx package is a Mandatory Reply Form for you to indicate the total number of affected screw(s) in your sets.We will be providing any additional follow up instructions as needed. If you have any questions call (574) 267-0865.For further questions regarding this recall , please call (574) 268-6379 |  | 22 | 22 units | Worldwide Distribution - US Distribution to the states of : TX, MO, FL, GA, NC, AZ, NY., and to the countries of France and Ireland. | Terminated | July 27, 2017 | 251 | 22 | Not_Computer
Z-0887-2017 | 75869 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Black Widow 6.6mm x 20mm Screw, Part Number TBS020; Black Widow 6.6mm x 25mm Screw, Part Number TBS025; Black Widow 6.6mm x 30mm Screw, Part Number TBS030 |  |  | November 22, 2016 | 2016 | Amendia, Inc | There is a problem with the mating feature between the plate and screw which could cause the plate opening not to interfere enough with the plate and head of the screw. | Process change control | Amendia sent an Urgent Medical Device Recall letter dated November 22, 2016, to all affected consignees.. The letter requested that they examine inventory and quarantine the product. The letter also requested that they contact Amendia to confirm receipt of the letter and to obtain a Return Material Authorization for any product on hand. Customers with questions should call 1-855-755-3329.For questions regarding this recall call 404-610-7215. |  | 96 | 96 units | Nationwide Distribution to IN, LA, and TX | Terminated | July 31, 2017 | 251 | 144 | Not_Computer
Z-0899-2017 | 75836 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AQURE System |  |  | November 23, 2016 | 2016 | Radiometer America Inc | Design error when displaying additional information in the patient view window; error may result in misreading a parameter and its value. | Software design | Radiomet4er sent a letter dated November 30, 2016, to all affected customers via email with a request for customers to complete a return form. Customers/Users were advised to not use the "more" button to expand patient results when viewing them in the AQURE patient view. Customers/Users Radiometer Representative will contact them to schedule a visit of a remote session. During the visit or remote session the representative will run a script on the AQURE system, which serve to remove the misalighnment possibility.It is noted in the letter to customers that expanding the results by using the More button will not affect the stored results.If customers are not the end-user of the affected product ensure that the notification letter is distributed to the final end-user.If any questions regarding the notification letter contact Radiometer America Technical Support at 1-800-736-0600 opt.1.For questions regarding this recall call 440-871-8900. |  | 7 | 7 total (2 in US) | Worldwide Distribution - US including MI and NY and Internationally to England, Switzerland, and United Arab Emirates. | Open, Classified |  |  | 7 | Not_Computer
Z-0895-2017 | 75746 | No | FCH | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BREAST PACK, Surgical Instrument Tray, REF/Catalog No. 5406, STERILE, Rx Only -- busse Hospital Disposables Hauppauge, NY --- Device Listing # E193298. |  |  | November 15, 2016 | 2016 | Robert Busse &amp; Co. Inc. d.b.a. Busse Hospital Disposables | Covidien issued a labeling update for the Covidien Devon Light Gloves. Busse had incorporated the affected gloves in several kits and distributed them prior to receipt of the recall notification. | Other | Busse Hospital Disposables sent a Medical Device Correction letter (dated November 11, 2016) and response form to the three customers via FedEx on 11/15/2016 to inform them of the recall of the kits containing the affected gloves. This was done without notifying NYK-DO. |  | 35 | US: 35 cases | PA, FL, VA | Open, Classified |  |  | 131 | Not_Computer
Z-0900-2017 | 75740 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe Reverse Shoulder Fixed Angle Torque Defining Screw Kit |  |  | October 25, 2016 | 2016 | Exactech, Inc. | Disengaged screw head could potentially not be retained in the Torque Defining Screwdriver. | Component design/selection | Exactech sent an Advisory Notice on October 25, 2016, to Exactech field agents in the U.S. Agents outside of the U.S. were notified on November 2, 2016. Field agents were notified of the recall in the Advisory Notice and asked to notify their accounts that may have the recalled product in stock. Agents were required to complete and return the Field Advisory Notice Response Form within 5 working days. For further questions, please call (352) 377-1140. |  | 2278 | 2,278 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MD, ME, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, SC, TN, VA, WA and WI., and to the countries of :Austria, Australia, Canada, Germany, France, Spain, India, Switzerland, Italy, South Korea, United Kingdom, Brazil, Luxembourg, Greece, Guatemala, Netherlands, South Africa, and Sweden | Open, Classified |  |  | 2278 | Not_Computer
Z-0888-2017 | 75649 | sling, arm | ILI | Arm sling. | Physical Medicine | Physical Medicine | 510(K) Exempt | Medi Choice Arm Slings |  |  | October 31, 2016 | 2016 | Owens &amp; Minor Distribution, Inc. | Owens &amp; Minor is the initial importer and private label distributor of Medi Choice Arm Slings (Product). The Product is manufactured in China by Tops Medical Articles Nantong Inc.Owens &amp; Minor first learned of the product issue that prompted the product removal action via customer complaints received in September 2016. Two customer complaints received on September 1"1, 2016 reported surface contamination on the arm slings. Two additional complaints reported on September 19, 2016 and October 7, 2016 reported the same issue.Upon receipt of the complaints on September 1, 2016, Owens &amp; Minor immediately initiated a CAPA investigation for root cause analysis and determination of corrective/preventive actions. A quality inspection of inventory confirmed that the surface contamination was confined to lots manufactured in May, June, and July of 2016. The surface contamination for these lots was found to be frequent. Subsequent August product lots were inspected and no surface contamination was found. Limited inventory from lots prior to May 2016 were available and inspected as well. No surface contamination was found on these prior lots.Owens &amp; Minor has quarantined the affected product lots. The product is on hold at all Owens &amp; Minor Distribution Centers (DC). Samples of contaminated product have been sent to the manufacturer for evaluation and to a certified testing lab for identification of contaminate. These results are pending at this time.Based on the frequency of the product defect as well as evaluation of potential adverse health event risk, Owens &amp; Minor initiated a voluntary recall for the above stated affected product lots. | Under Investigation by firm | The firm, Owens&Minor;, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 11/9/2016 to its consignees/customers including a Recall Response Form. The letter describe the product, problem and actions to be taken. The customers were instructed to examine your inventory and identify all affected product; immediately stop using all affected product; quarantine all affected product per your facility's Standard Operating Procedures; if you have further distributed this product, identify patients/customers who received product and notify them at once; and complete and return Recall Response Form via email: www.owens-minor.com, fax to 804-723-7100, or mail to the O&M; Distribution Center (DC) Attn: Recall Officer (even if you do not have any affected product). Note: upon receipt of the Recall Response Form, the DC will contact you with the appropriate RGA and instructions for product return.If you have any questions concerning credit and/or replacement of affected product or this recall, please call 804-723-7000 or contact your O&M; Sales Representative or Customer Service Representative at the local O&M; Distribution Center. |  | 568 | 568 cases | US Nationwide Distribution. | Open, Classified |  |  | 568 | Not_Computer
Z-0890-2017 | 75301 | No | FEI | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Best Practices kits containing certain production lots of Clearify Visualization System |  |  | October 04, 2016 | 2016 | Medtronic | Product sterility is compromised due to breach of the sterile barrier | Employee error | Covidien issued an Urgent Medical Device Recall letter on October 4, 2016 and a corrected letter on October 10, 2016 to US Customers via Federal Express. The letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Disposition of product within Medtronic control and customer returned devices will be reworked.Questions or concerns, contact your Medtronic representative or customer service at 800-882-5878. |  | 230 | 230 systems | Worldwide Distribution - US (Nationwide)Foreign:AustriaBelgiumBulgariaBrazilChileCyprusDenmarkEgyptFinlandFranceFranceGermanyGreeceHungaryIranIrelandIsraelItalyJapanKenyaKuwaitLatviaLuxembourgMexicoNamibiaNetherlandsNigeriaNorwayPolandPortugalRomaniaSaudi ArabiaSerbiaSlovakiaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited Arab EmiratesUnited Kingdom | Open, Classified |  |  | 1293399 | Not_Computer
Z-0885-2017 | 75798 | No | MTF |  |  | Immunology | PMA | IMMULITE 2000/IMMULITE 2000 XPi Systems Prostate-Specific Antigen (PSA) Assay, REF/Catalog Number L2KPTS2, SMN 10706281 and REF/Catalog Number L2KPTS6, SMN 10706282 (US); REF/Catalog Number L2KPS2, SMN 10380986 and REF/Catalog Number L2KPS6, SMN 10380996 (OUS); IVD --- Origin: UK Siemens Healthcare Diagnostics Products Ltd. Llanberis, Gwynedd, LL554EL UK |  |  | November 18, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics confirmed that the IMMULITE 2000/ IMMULITE 2000 XPi Systems PSA assays are not meeting the current High-dose Hook Effect expectation in the Instructions for Use. | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Correction Letter" (IMC 17-07.A.US, dated 11/2016) was sent via FedEx on 11/18/2016 to the US customers and an "Urgent Field Safety Notice" dated 11/2016 to customers outside the US (distribution is determined at the country level). The letter/notice described the product, problem and actions to be taken. Additionally, Field service personnel were sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. Customers were instructed to review letter with your medical doctor; review products and follow instructions; complete and return the attached Field Correction Effectiveness Check Form within 30 days by fax to Customer Care Center at (312) 275-7795; retain this letter with your laboratory records and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 6697 | 6697 kits (US: 1646 kits; OUS: 5081 kits) | Worldwide Distribution-US (nationwide)and the following countries: Afghanistan, Albania, Algeria, Argentina, Aruba, Austria, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg, Brazil, British Guyana, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Rep., Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Greece, Great Britain, Grenada, Guatemala, Honduras, Hungary, Iceland, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Jordan, Kazakhstan, Kosovo, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Mexico, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Spain, South Africa, South Korea, Suriname, Switzerland, Syria, Taiwan, Tadjikistan, Thailand, Trinidad, Tobago, Tunisia, Turkey, Turkmenistan, United Arab Emirates, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 11402 | Not_Computer
Z-0880-2017 | 75792 | No | NDE | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAM Kirschner Wire, Sterile R, T2 K-Wire, Sterile R and T2 K-Wire Recon, Stryker Trauma GmbH, Prof.-Kuntscher-Strasse 1-5, Schonkirchen, Germany 24232; T2 K-Wire, Sterile R |  |  | November 09, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | A review of packaging revealed the seal integrity of the pouch may be compromised. More specifically, there is a potential that the sterile pouch is not sealed at one end due to a manufacturing error. | Packaging process control | Stryker sent an Urgent Medical Device Removal letter dated November 9, 2016, to all affected customers via UPS (with return receipt). Stericycle will be handling the returns. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform users of the Medical Device Recall and forward the notice to all those individuals who need to be aware within their organization. Return all affected products available at their location to:Stryker C/O Stericycle2670 Executive Dr., Suite AIndianapolis, IN 46241Attn: RA 2016-169 - Event XXXCustomers were also instructed to complete and sign the enclosed Business Reply Form and return to Stericycle.Customers with questions should call +1-201-851-6683.For questions regarding this recall call 201-831-5000. |  | 2084 | 2,084 units | Worldwide Distribution - US (nationwide) Internationally to AU, CA, CL, CH, ES, FR, GB, SE, JP, FR, ZA, and NL. | Open, Classified |  |  | 2084 | Not_Computer
Z-0892-2017 | 75780 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3, PowerCross .018 OTW PTA Dilatation Catheter, Sterile EO. |  |  | November 10, 2016 | 2016 | Medtronic Inc | Medtronic identified an issue with specific models and lots of the PowerCross .018 Over the Wire (OTW) Percutaneous Transluminal Angioplasty (PTA) Dilatation Catheter. During an unrelated manufacturing inspection of scrapped products, an out of specification condition was observed. The catheter outer shafts were pinched and deformed at the proximal balloon bond. Bench testing of affected devices showed that this resulted in either the inability to inflate the device, a slow deflation, or the inability to deflate the device post expansion. | Process control | Medtronic verbally communicated about the issue on November 20. 2016, to affected customers. Additionally, consignees were hand delivered by Field Sales Representatives a Medtronic "Urgent Medical Device Recall" letter dated November 2016. The letter was addressed to Risk Manger or Health Care Professional. The letter described the problem and the product involved in the recall. Advised consignees to immediately identify, quarantine the affected product, and return it to Medtronic. Requested consignees to complete and email the Customer Confirmation Certificate to RS.CFQFCA@medtronic.com or fax it to Medtronic at 651-367-0612, attention: Customer Focused Quality.For questions, contact your Medtronic representative. If you require assistance in contacting your representative, please contact Medtronics National Answering Service at 800-633-8766. |  | 315 | 315 | NY, NC, WA | Terminated | August 02, 2017 | 265 | 315 | Not_Computer
Z-0893-2017 | 75737 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gammex Non-Latex Sensitive Neoprene Surgical Gloves |  |  | October 24, 2016 | 2016 | Ansell Healthcare Products LLC | Ansell Healthcare Products is initiating a voluntary recall involving 15 cases of Gammex Non Latex Sensitive Surgical Gloves because the packaged contents incorrectly included a natural rubber Latex surgical glove. | Packaging change control | Ansell Healthcare Products, LLC. sent an Urgent: Voluntary Surgical Glove Recall letter dated October 28, 2016, to all affected customers. The letter identified the problem the product and the action the customer needs to take.Please check your product stock at the warehouse and shelf level in order to advise us directly if any gloves having the recalled lot numbers are found. This information, along with your tracking number, must be communicated to our Customer Service Department by calling 1-800-952-9916, or by returning the attached Response Form to Ansell by fax at1-800-722-8155, or via email customerserviceus@ansell.com . Returned product will either be replaced or credited without delay. Your Ansell representative will be available to follow up with you regarding any concerns you may have. |  | 15 | 15 Cartons | US Distribution to the states of : PA, FL, NJ, IL and MA. | Open, Classified |  |  | 15 | Not_Computer
Z-0902-2017 | 75817 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Laboratory 5002-0180, Specimen Gate eReports 5002-0370, Software Version 1.0 |  |  | December 01, 2016 | 2016 | PerkinElmer Health Sciences, Inc. | Issue may cause eReports to consume an excessive amount of disk space causing the software to stop operating. | Software design | PerkinElmer Health Sciences issued a Medical Device Correction letter dated December 1, 2016, to all affected customers email facilitated through our PerkinElmer Software Services group located in Akron Ohio. Software Services will be asked to make sure the customers receive the Recall Letter and implement the recommended corrective measures.Actions to be takenA field correction of the issue is recommended. PerkinElmer will correct the issue with an update installed through remote connection. You will be contacted by your PerkinElmer representative to make the necessary arrangements for implementing the update.In the interim period, please do not attempt to login multiple times in quick succession. Allow the software time to respond to first login command.Please contact your local PerkinElmer representative or Specimen Gate Support for further information.Other InformationPlease inform those affected in your organization accordingly.To comply with regulatory requirements we request that you complete the enclosed response form and return it by fax to number 1-330-825-8520 (United States) /+358 2 2678 357 (outside United States) or as scanned by e-mail to TurkuQMresponse@perkinelmer.com as soon as possible, but not later than December 22, 2016.For questions regarding this recall call 781-663-6900. |  | 21 | 21 | Worldwide Distribution - US including CO, FL, GA, IL, MA, MI, NV, NJ, OH, SC, AZ, KY and Internationally to Italy, Ireland, UK, Spain, and Peru | Terminated | May 01, 2017 | 151 | 21 | Computer
Z-0904-2017 | 75851 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Innova 2000 |  |  | November 21, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare is conducting a recall due to the potential fall of aged CRT monitors - Prestige 2, Prestilix, Prestige VH, Prestige SI, RFX/SFX, Legacy, Precision 500D, ADVANTX and Innova 2000 systems. | Nonconforming Material/Component | Consignees were sent on 11/21/2016 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# FMI 10913 dated November 21, 2016. The letter was addressed to Director of Radiology, Risk Manager/Hospital Administrator, and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 1598 | Total for all products 1,598 (USA 824, OUS 774) | US: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT,NC, NE NH, NJ, NM, NY, NV, OH, OK, PA, RI, SC, SD, TN, TX,, VA, WA, WO, WY, WV. OUS: AlgeriaArmeniaAustraliaAustriaBrazilCanadaChileChinaDenmarkEgyptFranceGermanyGuatemalaHondurasHong KongHungaryIcelandIndiaIrelandItalyJapanKoreaLibyaMexicoNetherlandsNetherlandsPeruPolandPortugalRussiaSaudi ArabiaSerbiaSierra LeoneSlovakiaSpainTaiwanThailandTurkeyUnited Arab EmiratesUnited KingdomVenezuelaVietnam | Open, Classified |  |  | 1598 | Not_Computer
Z-0876-2017 | 75383 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 128 Multi-slice CT Scanner System with |  |  | August 31, 2016 | 2016 | Neusoft Medical Systems Co., Ltd. | The Firm discovered during contrast agent tracking scan, when the expected concentration of contrast agent is reached, later planned scan may be interrupted. The system must be powered down and restarted to clear the interruption error before scanning can resume. | Software design | Neusoft Medical Systems Co. Ltd. planned action to bring into compliance.1.Affected Customers will be notified by letter to provide them a description of the affected device, instructions for using the device prior to correction, hazards associated with the defect, and a statement of corrective actions to be taken at no cost to the owner.2.Field Service Engineers will visit owner sites to install a software update which resolves the defect.3.Corrections will be made at no cost to the owner4.Corrections will be completed by May 11, 2017. CDRH approves the CAP subject to the following conditions:.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.Further questions please call (281) 453-1205. |  | 23 | 23 units | US Distribution | Open, Classified |  |  | 23 | Computer
Z-0901-2017 | 74834 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | NIM EMG Endotracheal Tube, 6.0 mm I.D. x 8.8 mm O.D. (27 FR), Reinforced, Sterile, Rx only. Individually packaged. |  |  | June 28, 2016 | 2016 | Medtronic Xomed, Inc. | Wire in tubing can become exposed, posing potential harm to the intubated patient. | Use error | Medtronic sent an Urgent Product Safety Advisory letter dated June 28, 2016, to all affected customers informing them of the correction and providing instruction to avoid causing injury when using the EMG Endotracheal Tube. Recommended Actions included: inspecting the product thoroughly before use, testing the cuff by inflating it then deflating it before intubation, using a non-malleable stylet that closely matches the natural curve of the tube, not bending and flexing the tube excessively before use, and lubricating the cuff with a non-paralyzing, aqueous lubricant and a lubricated stylet. If customers found an EMG tube that appeared to have an electrode wire protruding from the distal end, the tube is not to be used and practitioners are urged to call Medtronic Product Quality Experience Department at 1-866-849-4003. For further questions , please call (904) 296-9600. |  | 3752 | 3752 | Worldwide Distribution - Nationwide Distribution including US Virgin Island and Puerto Rico , and to the countries of :Japan, Korea, China, Canada, Dominican Republic, Colombia, Chile, Costa Rica, Bahamas, Peru, Panama, Mexico, Argentina, Ecuador, Brazil, Trinidad and Tobago, Hong Kong, Taiwan, United Kingdom, Netherlands, Sri Lanka, Austria, Italy, Germany, United Arab Emirates, Saudi Arabia, Kuwait, France, Belgium, Oman, Ireland, Azerbaijan, India, Spain, Portugal, Luxembourg, Switzerland, Russian Federation, Austria, Greece, Cyprus, Portugal, India, Algeria, Iceland, Iran, Denmark, Hungary, Czech Republic, Ukraine, Turkey, Croatia, Israel, Norway, Oman, Sweden, Libya, Finland, Spain, Portugal, Algeria, Morocco, Poland, Canada, Malaysia, Singapore, Australia, and New Zealand | Open, Classified |  |  | 3752 | Not_Computer
Z-0918-2017 | 75834 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | Shimadzu X-ray High Voltage Generator Model No. UD150B-40 on SONIALVISION safire II/17 systems with 0.7/1.2JG326D X-Ray Tube |  |  | November 20, 2016 | 2016 | Shimadzu Medical Systems | Shimadzu Medical Systems is recalling the Shimadzu X-ray High Voltage Generator due to the possibility of oil leaking from the x-ray tube housing under certain conditions. | Under Investigation by firm | The firm, Shimadzu Medical Systems USA, sent an "Urgent Medical Device Recall notice" and Field Safety Notice dated 11/20/16 to customers to inform them that Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray High Voltage Generator UD150B-40. Customers are informed of how to identify the affected products, corrective action plan by Shimadzu, and the actions to be taken by customers. Customers that have not received a call or visit from Shimadzu are instructed to call Shimadzu Medical Systems National Technical Support at (800) 228-1429.Shimadzu sent an updated "URGENT Voluntary Medical Device Recall Notice" letter dated 11/30/16 to its customers on 11/30/16 to inform them that Shimadzu Corp. has identified a potential issue in the Shimadzu X-ray High Voltage Generator UD150B-40 on SONIALVISION safire II/17 systems with 0.7/1.2JG326D X-Ray Tube. The letter informs the customers that when operated for prolonged periods of time at low rate serial exposure, the heat load capacity of the X-ray anode disc (target) may become damaged and leak insulating oil. Customers with questions are instructed to contact Shimadzu National Technical Support at (800) 228-1429. |  | 13 | 13 units | US Distribution to states of: AZ, WA, CA, LA, FL, SC, and NJ. | Open, Classified |  |  | 13 | Not_Computer
Z-0903-2017 | 75476 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products Intact PTH Reagent Pack: 1) UPN 10758750006287 / Product Code 6802892 - Reagent Pack, 2) UPN 10758750006250 / Product Code 6802893 - Calibrators |  |  | October 05, 2016 | 2016 | ORTHO-CLINICAL DIAGNOSTICS | The device defect is a positive bias of up to 20% relative to the currently marketed Roche device. In addition, the firm states that during the investigation they discovered that the cleared reference interval (RI) was not set optimally at product launch in 2010 and that the upper limit in the IFU of 53.5pg/mL is too low and should be closer to 80 pg/mL. They are in the process of submitting the data to support the revised RI in a new 510K submission and will be changing the IFU once the 510K is cleared. | Under Investigation by firm | Ortho Clinical sent an Urgent Product Correction Notification letter dated October 5, 2016, to all affected customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received VITROS Immunodiagnostic Products iPTH to inform them of the issue and the required actions. Foreign affiliates were informed by email on October 5, 2016, of the issue and instructed to notify their customers of the issue and the required actions. Required Actions - Until further notice, be aware of positively biased results when using the device. Discuss any concerns regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action. Post the notification by your VITROS System or with your user documentation. In accordance with regulatory requirements, complete and return the Confirmation of Receipt form via email to: RA-OCDUS-CONFIRMAD@ORTHOCLINICALDIAGNOSTICS.COM. Customers with questions can call the Ortho Care Technical Solutions Center at 1-800-421-3311.For further questions regarding this recall, please call (585) 453-3452. |  | 7027 | 7,027 Total units | Worldwide Distribution - Nationwide Distribution : and to the countries of : Bermuda, Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 7027 | Not_Computer
Z-0917-2017 | 75833 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu X-ray TV System, Catalog No. SONIALVISION G4 |  |  | November 20, 2016 | 2016 | Shimadzu Medical Systems | Shimadzu Medical Systems is recalling Shimadzu X-ray TV System SONIALVISION due to possibility of oil leaking from the x-ray tube housing under certain conditions. | Under Investigation by firm | An urgent medical device recall notice and a field safety notice dated 11/20/16 was sent to customers to inform them that Shimadzu Corporation has identified a potential issue in the Shimadzu X-ray TV System SONIALVISION G4. Shimadzu informed their customers that they have received a report that oil had leaked from X-ray tube device of SONIALVISION G4 during examination. Customers are informed of the description of the issue, the actions planned by Shimadzu to correct the issue, and the actions that should be taken by the customer. Customers are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 39 | 39 units | US: TX, NY, SD, PA, MD, OH, IA, IN, MT, IL, FL, SC, OR, CA, AR, GA, CT, NV, PA, NE, WA | Open, Classified |  |  | 39 | Not_Computer
Z-0928-2017 | 75824 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Arthex S3 Arm Sleeve Connector; Catalog Number AR-1650-02; manufactured and distributed between March and November 2016. |  |  | November 23, 2016 | 2016 | Allen Medical Systems, Inc. | During use, the unit could separate un-expectantly. This failure would allow the patient's arm to fall out of suspension during a procedure, possibly causing moderate injury to the patient. This potential defect is not apparent to the user. | Under Investigation by firm | The firm's customer was notified of the recall via letter on November 23, 2016. They were asked to take the following actions: 1. Immediately remove from service your Arthrex S3 Arm Sleeve Connector (AR-1650-02). All serial numbers are impacted. 2. Please follow the included instructions for the return of all S3 Arm Sleeve Connectors (AR-1650-02). Allen Medical will inspect, repair or replace all units as necessary at no cost to you. Please take steps to retrieve the units from your customers and return them to Allen Medical via the standard Return Material Authorization (RMA) process. In order to minimize this inconvenience, we will repair the Arthrex S3 Arm Sleeve Connector and return the device to you as soon as possible. For further questions, please call (978) 266-4280. |  | 420 | 420 devices | US Distribution to the state of : FL | Open, Classified |  |  | 420 | Not_Computer
Z-0930-2017 | 75591 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sterling Diagnostics, Inc., HDL Cholesterol (PEF Method) for IN VITRO DIAGNOSTIC USE in the quantitative determination of analytes. These are CONSUMABLES - a customer buys another when kit is consumed. |  |  | July 20, 2016 | 2016 | Sterling Diagnostics, Inc. | Sterling Diagnostics, Inc. initiated a recall on 7/20/2016 for multiple diagnostic kits due to a recision of the previously extended expiration dates for the kits. | Process change control | Sterling Diagnostics, Inc. initiated their voluntary recall on 07/20/2016 by calling customers by phone to determine what kits remianed at customers' and compare withthe firm's sales records. Sterling Diagnostics followed up by sending a letter to each customer, and instructed them to take note of the corrected expiration dates for the dignostic kits, apply the revised dates to the kits in their possession, and return a customer acknowledgement to the Office Manager Luz Santa Mariaat Sterling Diagnostics, Inc. at 36645 Metro Court, Sterling Heights, MI 48312. Customers with questions can call 586-979-2141. |  | 245 | 245 | Domestic: MIForeign: PhillipinesVA/DOD: None | Terminated | March 22, 2017 | 245 | 1135 | Not_Computer
Z-0939-2017 | 75876 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS software. |  |  | October 15, 2015 | 2015 | Merge Healthcare, Inc. | The software fails to associate to the correct MG image if there are two images for the same view. | Software design | Merge sent an Urgent : Medical Device Recall letter dated July 11, 2016, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Be aware of the following workaround: Customer I User:Ensure both" tome" images are reviewed when more than one is captured for the same view. Merge has a released fix available for this issue.YOUR RESPONSE TO THIS NOTIFICATION IS REQUIREDPlease reply using the enclosed form and the return addressed envelope. Your response is required no later than August 2,2016. For further questions regarding this recall please call (877) 741-5369 |  | 9 | 9 sites potentially have the effected software | Distribution was made to medical facilities located in MT, CA, PA, and TX. | Open, Classified |  |  | 9 | Computer
Z-0938-2017 | 75813 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific, INGEVITY" MRI Leads, Sterile EO, RxOnly. |  |  | November 30, 2016 | 2016 | Boston Scientific Corporation | Boston Scientific is conducting a recall on Boston Scientific INGEVITY" MRI endocardial pacing leads. There is a possibility that the polyurethane boot at the terminal end of the lead was not securely connected to the lead body. | Process control | Boston Scientific sent an "Important Medical Device Information" letter dated November 30, 2016. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall. The letter also described the "Clinical Considerations, Recommendations, Affected Population &amp; Further Information". For questions contact your local Boston Scientific representative or Technical Services. For further questions, please call ( 651) 582-4000. |  | 15 | 15 (2 US; 13 OUS) | Worldwide Distribution - US Distribution to the states of : IA and NY., and to the countries of : United Kingdom, Netherlands, Italy, Germany and Belgium. | Open, Classified |  |  | 15 | Not_Computer
Z-0919-2017 | 75860 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Scope, Computed tomography x-ray system |  |  | November 22, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens is providing a new installation of the Residual Current Monitor (RCM) in the Line Connection Box (LCB) to resolve a potential defect that could lead to a power cutout during use of the SOMATOM Scope, Scope Power and Perspective systems. | Nonconforming Material/Component | Siemens mailed on December 01, 2016 a Customer Safety Notice to affected customers to inform them of a corrective action that is being performed to prevent a possible hazard to persons and equipment. The notice explained the underlying issue, the effect the issue has on the system, the associated risks and actions they will take to correct the issue.In addition, customers will receive a Customer Safety Notice via Hand Delivery from the Siemens Customer Service Engineer at the scheduled appointment for the corrective installation. |  | 13 | 13 systems | Distributed to: ILILGASCSDOKNYMIINMDAR | Terminated | May 30, 2017 | 189 | 13 | Not_Computer
Z-0926-2017 | 75348 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Normal Saline Flush 0.9% USP Sodium Chloride Injection Syringe, 3 ml fill, 5 ml fill and 10 ml fill. Product codes: 1203, 1205, 1210, 1210-BP |  |  | October 04, 2016 | 2016 | Nurse Assist, Inc | Potential contamination with B. cepacia. | Unknown/Undetermined by firm | Nurse Assist, Inc. sent an Urgent Medical Device Recall letter dated October 4, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately discontinue use, a Returned Goods Authorization would be provided for return of the affected product. If product was further distributed consignees were instructed to make every effort to retrieve any of the unused product. For questions consignees should call 800-649-6800 or email productremovalinfo@nurseassist.com.For questions regarding this recall call 817-231-1300. |  | 386175 | 386,175 syringes | Worldwide Distribution - US (Nationwide), St. Thomas, US Virgin Islands, and Internationally to Canada. | Terminated | August 16, 2017 | 316 | 386175 | Not_Computer
Z-0942-2017 | 75881 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TRU-Balance 3 Power Positioning Systems, TB3 Lift/Tilt/Recline (Gen 1). |  |  | December 01, 2016 | 2016 | Pride Mobility Products Corp | The interface between the TB3 back and tilt base can separate when tilted beyond the specified angle. | Use error | Pride Mobility Products Corp. mailed a letter to affected customers to inform them of the issue. Systems were not required to be returned. Customers were informed that a TB3 Seat Depth Adjustment Reinformcement Kit which contains (2) reinforcement plates and associated hardware to strengthen the interface will be installed to address the issue. Direct accounts were asked to contact patients to inform them of the product correction. Customers were also requested to complete and return Pride's completion card once the correction has been made. Pride will issue a $65.00 credit once the completion card is received. Customers with questions were instructed to call Pride's Upgrade Department at 866-222-6212.For questions regarding this recall call 800-800-8586. |  | 12343 | 12,343 chairs | Nationwide Distribution | Open, Classified |  |  | 12343 | Not_Computer
Z-0940-2017 | 76014 | No | OGZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | TVC Imaging System model TVC-MC9 , for the near-infrared examination of coronary arteries in patients undergoing invasive coronary angiography |  |  | December 15, 2016 | 2016 | Infraredx, Inc. | Near-infrared spectroscopic image (Chemogram) generated by the TVC Imaging System TVC-MC9 in some cases demonstrated an unusual appearance in that the yellow in the image (indicating the presence of lipid core plaques) was horizontally-oriented | Device Design | Infraredx, Inc. sent an Urgent Product Recall Medical Device Field Correction letter dated December 15, 2016, describing the problem and replacing the defective lasers. A Field Service Representative will contact customers to schedule a replacement of the system's laser rack module. If customers receive this letter and are not contacted within 48 hours, please call Infraredx Technical Support@ 1-800-596-3104.For questions regarding this recall call 781-221-0053. |  | 7 | 7 units | Worldwide Distribution - US including MI and Internationally to South Korea | Open, Classified |  |  | 7 | Not_Computer
Z-0943-2017 | 75888 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROFEMUR(R) R RASP HANDLE, REF PPW38078, 1 EACH, NON-STERILE, MicroPort Orthopedics, Arlington, TN |  |  | November 17, 2016 | 2016 | MicroPort Orthopedics Inc. | The recall is being initiated because all broach handles in the affected lot had the same non-conformity with the connection feature. The broach handles may have machining issues that do not allow the handle to properly mate with an associated broach. | Process control | MicroPort Orthopedics sent an Voluntary Device Product Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the enclosed form even if they do not have any affected product to return. Return the completed form by fax to 1-901-451-6032 or by email to PostMarket@ortho.microport.com. Attempt #1 for acknowledgment will be made by email. For questions regarding this recall call 901-867-4318. |  | 20 | 20 units | Worldwide Distribution to Spain, China, and Italy | Open, Classified |  |  | 20 | Not_Computer
Z-0941-2017 | 75882 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | G3 Opticage Expandable Interbody Fusion Device, Model No. 9180-00, 9180-01, 9180-06, 9180-07, 9180-17, 9101-00, 9101-06 |  |  | November 22, 2016 | 2016 | Interventional Spine Inc | During implantation, the proximal end of the device could separate if the angle between the device and insertion instrument is too steep. This separation will prevent the device from expanding. | Device Design | Interventional Spince, Inc. sent an An Urgent Medical Device Recall letter dated November 22, 2016, to all affected customers to inform them that during implantation, the proximal end of the device could separate if the angle between the device and insertion instrument is too steep. This separation will prevent the device from expanding. The letter informs the customers of the risk to health and actions to be taken by the customer/user. Customers were instructed to discontinue use and return device per the enclosed instructions. Complete the enclosed acknowledgment letter and fax, email or mail it to Interventional Spine. Customers with questions were instructed to call Monday through Friday 8am to 5pm (Pacific Time), 949-472-0006 or email cs@i-spineinc.com. |  | 239 | 239 units | Nationwide Distribution to NY, NC, SC, PA, TX, OH, GA, NJ, DE, AZ, DC, CA, and VA. | Completed |  |  | 239 | Not_Computer
Z-0944-2017 | 75873 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular REDAPT(TM) Hip Systems |  |  | November 15, 2016 | 2016 | Smith &amp; Nephew, Inc. | All lots of modular neck hip prostheses are being recalled due to a higher than anticipated complaint and adverse event trend. | Device Design | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated November 15, 2016, to all affected customers. Customers were instructed to inspect their inventory and locate any unused devices and quarantine immediately. Complete the Inventory Return Certification Form and obtain a return authorization (RA) number if there was product to return.For questions regarding this recall call 978-749-1440. |  | 6266 | 6266 units | Worldwide Distribution - US (nationwide) and Internationally to Austria, Australia, Belgium, Canada, Switzerland, China, Cyprus, Czech Republic, Dubai, Germany, Denmark, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Mexico, Netherlands, Poland, Puerto Rico, Portugal, Russia, Turkey, and South Africa. | Open, Classified |  |  | 31936 | Not_Computer
Z-0949-2017 | 75999 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PreludeEASE Hydrophilic Sheath Introducer Dilator |  |  | December 01, 2016 | 2016 | Merit Medical Systems, Inc. | Recalling dilators included with the 6F PreludeEASE Hydrophilic Sheath Introducer because the dilator tips were manufactured with an undersized inner diameter and therefore will not allow passage of an 0.021 (0.53 mm) guide wire. | Process change control | Merit Medical sent an Urgent Product Recall Notice dated December 1, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to identify any affected devices within their facility, quarantine them, and discontinue use. Ensure that all personnel to whom the devices were distributed are made aware of this field action. Consignees were asked to work with their Merit Sales Representative to arrange product return. Consignees with questions were instructed to contact their Merit Sales Representative or Merit Customer Service at response@merit.com or at 801-208-4381.For questions regarding this recall call 801-208-4623. |  | 335 | 335 | Nationwide Distribution | Terminated | April 18, 2017 | 138 | 335 | Not_Computer
Z-0947-2017 | 75937 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gomco Circumcision Clamps with separate O-Ring Component: The device is packaged non-sterile and has no shelf-life |  |  | October 28, 2016 | 2016 | Aesculap Implant Systems LLC | Aesculap has received complaints of excessive bleeding after use of Gomco Circumcision Clamps. | Under Investigation by firm | Aesculap sent an Urgent Medical Device Recall Notification letter dated October 28, 2016, to all affected customers were sent an Urgent Medical Device Recall Notification letter via Fed-Ex overnight.. Customers were asked to immediately remove the affect product from their inventory and return to Aesculap. A return label was provided to efficiently remove and return the affected products. Customers with questions should call 610-984-9414.For questions regarding this recall call 800-258-1946. |  | 437 | 437 units | Nationwide Distribution | Open, Classified |  |  | 437 | Not_Computer
Z-0948-2017 | 75827 | Yes | OOD | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GalaSHAPE 3D; product code SH3D03, a medium size oval |  |  | November 21, 2016 | 2016 | Tepha Incorporated | Endotoxin values from testing of retains slightly exceeded the established specification. | Process design | Tepha Medical Devices sent a Voluntary Product Recall Notification letter dated November 22, 2016, to all affected customers. Beginning on 11/21/2016, all impacted customers were notified of the recall via phone or in person by Galatea Surgical, a wholly owned subsidiary of Tepha, responsible for the sales and marketing of GalaSHAPE 3D. That notification was followed up by the Tepha recall letter dated November 22, 2016, delivered by certified mail. Customers were instructed to do the following:1) Immediately review your inventory for the affected lot number listed above.2) If you do not have any of the affected lot in your inventory, complete the enclosed RecallAcknowledgment Form, indicate that you have zero (O) inventory and fax the completed form back to (781) 357-1701 or email it to contact@galateasurgical.com.3) If you do have inventory from the affected lot, please complete the enclosed Recall Acknowledgement Form and fax back to Tepha, Inc. at (781) 357-1701 or email it to contact@galateasurgical.com. A Galatea representative will contact you to arrange for the product return.If you have any questions or require assistance with the return of the recalled product, please call Galatea Surgical, at (781) 357-1752.For questions regarding this recall call 781-357-1750. |  | 29 | 29 | Nationwide Distribution to IN, CA, FL, MI, NV, and MT | Terminated | January 05, 2017 | 45 | 29 | Not_Computer
Z-0951-2017 | 75846 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Natus Warmette Blanket Cabinet and Warmette Dual Cabinet |  |  | October 15, 2016 | 2016 | Natus Manufacturing Limited | The fuse blows during initial power up or after a period of time. This results in a complete loss of power to the Cabinet and in some cases has resulted in the fuse cover for the power inlet melting slightly. | Mixed-up of materials/components | Natus Manufacturing Limited sent an "Urgent Field Safety Notice" letter dated April, 2016, via email. The letter described the problem and the product involved in the recall. Informed the consignees that a repair kit to correct the issue is available for repair of affected units. Requested consignees to complete the attached form and Faxto 847-949-6031 or e-mail it to Natus Quality Programs. In turn the firm will send the repair kit. For questions contact Technical Service at 888-496-2887.For questions regarding this recall call 888-496-2887. |  | 31 | 31 | Worldwide Distribution - US including NJ, NC, TX, IL, VA, NY, FL, CA, MS, WA, IN, and MA. and Internationally to Canada and Costa Rica. | Completed |  |  | 31 | Not_Computer
Z-0960-2017 | 76029 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Tiger Screw Sterilization Tray Base Assembly Model 210-00-001 Revision N, O |  |  | December 08, 2016 | 2016 | Trilliant Surgical Ltd. | Four lots of Tiger Screw Sterilization Tray Base Assemblies were released to the field with a printing error where k-wire description markings indicate diameters of 0.45 and 0.35 instead of 0.045" and 0.035", respectively. | Under Investigation by firm | Trilliant Surgical Ltd. sent an Urgent Medical Device Recall letter dated December 8, 2016, to all affected consignees. The letter instructed the consignees to 1) identify any/all affected product; 2) document any/all identified product via a provided response form and return to Trilliant Surgical; 3) exchange product upon receipt of replacements; and 4) return any/all affected product as originally identified and reported on the return response form. Customers with questions were instructed to call 800-495-2919. |  | 50 | 50 units | Nationwide Distribution | Terminated | May 17, 2017 | 160 | 50 | Not_Computer
Z-0952-2017 | 76104 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | SCREW 02.7-12MM LOCKING, REF W27 ST112: a) DRP, b) NEO, In2Bones, 28 chemin du petit bois, 69130 Ecully - France |  |  | December 19, 2016 | 2016 | In2Bones, SAS | Dimensional non-conformity of the thread on the screw head. | Component design/selection | IN2BONES en a Field Safety Corrective Action letter dated December 15, 2016, to the sole US distributor. The letter identified the product, the problem, and the action to be taken by the customer. We therefore recommend you to follow the instructions here below:1. Identify all the devices with the batches subject of this recall that might still be in your inventoryand quarantine them2. Inform and distribute this Recall Notification to all relevant persons within your organization3. To distributors only: Identify all the devices with the batches subject of this recall that weredelivered to your customers, and if relevant, instruct them to also follow these instructions(identification and quarantine)4. Fill in and return the fax back form enclosed. With this form, you will certify that you havereceived this Recall Notification and intend to comply with the recommendations listed. This faxback form will enable In2Bones to conduct effectiveness checks.Customers with questions were instructed to contact the Quality and Regulatory Affairs team at +33 4 72 29 26 26 / +33 7 61 88 44 30 or by email:qualte@in2bones.com.For questions regarding this recall call 844-602-6637. |  | 19 | 19 units | Worldwide Distribution - US TN only and Internationally to France and Switzerland | Open, Classified |  |  | 85 | Not_Computer
Z-0955-2017 | 75782 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | VOCO Futurabond M+ adhesive, RF 1515, Manufactured by VOCO GmbH, Futurabond M+ is a dental adhesive for use with methacrylate-based restorative, core build-up or luting materials. It can be applied with phosphoric acid etching (selective enamel etching or total-etch) of the dental hard tissue as well as without (self-etch). A1] Adhesive for direct restorations of all classes of cavities using light-curing composite /compomer / Ormocer materials with a methacrylate base. [A2] Adhesive for core build-ups made of light-curing composite materials. [A3] Adhesive for indirect restorations with light-curing luting composites. [B] Intraoral repair of composite, compomer or Ormocer restorations, ceramic veneers and allceramic restorations without an additional primer. [C] Treatment of hypersensitive tooth necks. [D] Protective varnish for glass ionomer cement restorations. [E1] Sealing of cavities prior to amalgam restorations. [E2] Sealing of cavities and core preparations prior to the temporary luting of indirect restorations. In combination with Futurabond M+ DCA (Dual Cure Activator): [F1] Adhesive for direct self-curing or dual-curing composite restorations of all classes of cavities and for core build-ups. |  |  | November 03, 2016 | 2016 | Voco GmbH | Labeling mix-up: Futurabond M+ DCA (Dual Cure Activator) was marked labeled as Futurabond M+ Universal Adhesive. | Labeling mix-ups | Voco Gmbh notified consignees via email on/about November 3, 2016. The consignees were informed about the incorrect labeling and requested not to use this affected lot as well as to return it. They were also requested to pass this information to their customers |  | 50 | 50 units | Nationwide Distribution to FL, IA, NH, PA, SC and TX | Open, Classified |  |  | 50 | Not_Computer
Z-0956-2017 | 75901 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Oxford Fixed Lateral Bearing Size F Right, Part Number 154375 Lot 744260, Sterile, The product is placed into an Aluminum Oxide coated PET pouch and vacuum sealed with an Argon flush. The vacuum sealed part is then placed into a coated 1073B Tyvek/film pouch and heat sealed. This pouch acts as the sterile barrier. The pouch is then labeled and placed into a carton with an IFU. The carton is then labeled and shrink wrapped. The pouch and carton labels also contain sterilization indicator dots. |  |  | December 06, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Biomet is initiating a removal of a single lot of Oxford Fixed Lateral Bearings due to mislabeling. Zimmer Biomet received product complaints reporting that when the package labeled as a Right was opened it actually contained a Left Bearing. | Employee error | Zimmer Biomet sent an Urgent Medical Device Recall letter dated December 6, 2016. to all affected customers for a single lot of Oxford Fixed Lateral Bearings via e-mailed letter. Zimmer Biomet received two product complaints reporting that when the package labeled as a Right was opened it actually contained a Left Bearing.Risks: If another bearing is not available, a delay greater than 30 minutes may occur. If the product of the incorrect orientation is implanted, premature revision may be required. Customers were instructed to review the recall notification and ensure affected team members are aware of the contents.2. Immediately locate and quarantine affected product in your inventory.3. Complete the Certification of Acknowledgement portion of the Inventory Return Certification Forma. Return a digital copy within three (3) days via email to corporatequality.postmarket@zimmerbiomet.com.4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Inventory Return Certification Form to Zimmer Biomet.Request a Return Authorization Number via email at rgarequest@zimmerbiomet.com . Be sure to specify RECALL as the RGA type.b. For each return, send a copy of Inventory Return Certifictaion Form via email to corporatequality.postmarket@zimmerbiomet.com.c. Include a hardcopy of Inventory Return Certifictaion Form to Zimmer Biomet with your shipment for immediate processing.d. Mark the outside of the returns box(es) clearly with RECALL. 5. If after reviewing this notice you have further questions or concerns please call 1-877-946-2761 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of business hours will receive a prompt to record a voicemail and will be returned upon receipt. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com.Customer ServiceZimmer Biomet, Inc.P.O. Box 708Warsaw, I |  | 12 | 12 units distributed | Domestic: NoneFroeign: Netherlands &amp; SwitzerlandVA/DOD: None | Terminated | June 01, 2017 | 177 | 12 | Not_Computer
Z-0958-2017 | 75983 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Twist Drill PRO Âż4.2mm,long,41mm,SS-Non sterile. Dental implant accessory |  |  | November 23, 2016 | 2016 | Straumann Manufacturing, Inc. | Drills cannot be inserted into handpiece due to tolerance specification not met | Under Investigation by firm | Straumann sent an Urgent Field Safety Notice to all affected customers via UPS on November 23, 2016 for next day delivery. Customers were requested to check inventory for the listed article and lot numbers and return all unused, packaged parts. Replacement drills will be sent upon receipt. Complete the Confirmation Form and fax the form to the number listed on the form. For questions regarding this recall call 800-448-8168. |  | 120 | 120 units | Nationwide Distribution | Open, Classified |  |  | 285 | Not_Computer
Z-0962-2017 | 75943 | Yes | NIK |  |  | Cardiovascular | PMA | Amplia MRI" CRT -D SureScan", |  |  | December 09, 2016 | 2016 | Medtronic Inc., Cardiac Rhythm and Heart Failure | During internal testing conducted as part of next generation product development, it was discovered that the AdaptivCRT (aCRT) feature of Amplia MRI and Claria MRI CRT -D devices may appear to be enabled when it is not active. | Software Design Change | Consignees were hand delivered beginning on 12/9/2016 a Medtronic "Urgent Medical Device Correction" letter dated December 2016. The letter was addressed to "Dear Physician". The letter described the issue and the product involved in the recall. The letter informed consignees that a software update is being developed to address the issue and it will also address other issue not covered in this recall. The letter also included Appendix A which listed Patient Management Recommendations (Until the software update has been approved and the device models listed in Table A receive the update, follow the programming recommendations found in Appendix A. These recommendations also apply to any new device implants).Additionally, two other "Urgent Medical Device Correction" letters were hand delivered to the consignees, "Principal Investigator Contact Record" and a supplement to the advisory notification.For questions contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636. |  | 21111 | 21,111 ( 13,198 US; 7,913 OUS). | US: All states in USA including DC &amp; PR. OUS: Australia, Finland, Japan, Saudi Arabia, Austria, France, Korea,Singapore,Bangladesh, Germany,Luxembourg, South Africa, Belgium,Greece, Malaysia, Spain, Brunei Darussalam, Guadeloupe, Netherlands, Sweden, Cayman Islands, Hong Kong,New Zealand, Switzerland, Chile, India, Norway, Thailand, Cyprus, Iran, Poland, Turkey, Denmark, Ireland, Portugal, United Arab Emirates,Dominican Republic, Italy, Reunion,United Kingdom. | Open, Classified |  |  | 21111 | Computer
Z-0963-2017 | 75941 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Ablation Catheter Cable |  |  | December 06, 2016 | 2016 | Stryker Sustainability Solutions | A breach in the sterile packaging may have occurred resulting in potentially non-sterile product. | Under Investigation by firm | Customer notification letters were sent 12/06/16. Customers were asked to use the criteria of Item Number and Expiration Date to determine if any affected product remains in inventory. Use of the products from the affected item number and expiration date should be discontinued.Customers are instructed to follow the web link below to complete the Recall Effectiveness Check Form and indicate if any devices with the affected item number and expiration date listed above remain in inventory. ONLINE form must be completed even if no affected product is found.Use Stryker Sustainability Solutions Ship To Account Number as completethe form. If do not know your Ship To Account Number, contact local salesrepresentative or call Stryker Sustainability Solutions Customer Care at 1-888-763-8803 to obtain this number.If the form indicates that any affected devices remain in inventory, a prepaid shipping label will be issued for the return of the product. Credit will be issued to facility for all affected devices in scope of this recall that are returned.Address any questions regarding the recall to Novasyte by calling (760) 279-3126 oremailing sss151@novasyte.com.Adverse reactions or quality problems experienced with the use of this product may be reported to:- Stryker Sustainability Solutions Complaint Hotline: +1-888-763-8803 http://www.stryker.com/productexperience/- The FDAs MedWatch Adverse Event Reporting program either online, byregular mail, or by fax. |  |  |  | U.S. distribution to the following: MI, TX, TN, IN, AL, FL, PA, AK, WI, MT, SC, CO, CA, NH, NC, NJ, DE, AZ, MN, MS, GA, DC, IA, NY, NM, KS, OH, OK, SD, OR, NV, VA, MD, WA, CTForeign distribution to the following: CA. | Terminated | June 08, 2017 | 184 |  | Not_Computer
Z-0978-2017 | 75049 | No | MSS | Intraocular lens guide. | Ophthalmic | Ophthalmic | 510(k) | Alcon Monarch III (1) "C" Cartridge |  |  | September 05, 2016 | 2016 | Alcon Research, Ltd. | The intraocular lens could become lodged within the cartridge | Device Design | Alcon sent an Urgent Field Safety Action letter dated August 2016 to all affected consignees who received the affected product. Customers were instructed to check their inventory and segregate any of the affected cartridges. Customers were also asked to complete and return the Product Recall Response Form. , requesting the return of any units in their inventory. If customers had transferred product to other organizations, they were asked to provide them with a copy of the notice. Customers with questions were instructed to contact Alcon.For questions regarding this recall call 817-551-3058. |  | 1130 | 1,130 units | Distributed to Mexico only | Open, Classified |  |  | 1130 | Not_Computer
Z-0972-2017 | 76016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q/Q.zen, Angiographic x-ray system |  |  | December 08, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Siemens initiated a corrective action to address two possible, mutually independent causes of a system defect related to the following:- In Artis Systems with A100 Plus or A100G generators, an attempt to resume operation following detection of a fault can result in the failure of a module in the high-voltage generator.- For biplane systems delivered with software version VD11B since April of this year, software problems in conjunction with graphics card may, in rare cases, result in the loss of image display in the examination room. | Software in the Use Environment | Siemens mailed a Customer Safety Advisory Notice (CSAN) dated December 8, 2016, to all affected customers with Customer Safety Advisory Notice AX002/16/S. Additionally, Siemens has initiated a software update via Update Instruction AX001/16/S. Their Service Organization will contact each customer to schedule an appointment to perform the corrective action; however, customers can make contact for an earlier appointment. Customers with questions should call 1-888-826-9702.For questions regarding this recall call 610-448-6461. |  | 1500 | 1,500 distributed Worldwide | Nationwide Distribution | Open, Classified |  |  | 1500 | Computer
Z-0965-2017 | 75290 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner and Reprocessor System 208V, Product Code: 50004 |  |  | September 27, 2016 | 2016 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the EVOTECH Endoscope Cleaner and Reprocessor System because the hose may come loose and cause a leak. | Nonconforming Material/Component | ASP has sent out Service Representatives to all impacted customers to inspect the hoses for the EVOTECH Endoscope Cleaner and Reprocessor System and will replace the hose if required.The firm issued a Technical Bulletin dated 8/22/2016. Customers were instructed to complete online training prior to performing the actions described in the bulletin and follow instructions.For orders and questions call 888-783-7723 or email: RA-ASPUS-ServPartOrd@aspus.jnj.com. |  | 107 | 107 units | US Nationwide Distribution. | Terminated | April 17, 2017 | 202 | 107 | Not_Computer
Z-0979-2017 | 75837 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SIEMENS IMMULITE 2000/2000 Xpi Systems Intact PTH (Parathyroid Hormone) |  |  | November 22, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting this recall for the IMMULITEÂż 2000/IMMULITEÂż 2000 XPi Intact PTH (Intact Parathyroid Hormone) (iPTH) assay kit lot 320. Siemens has confirmed that IMMULITEÂż 2000/IMMULITEÂż 2000 XPi Intact PTH kit lot 320 can exhibit an average negative bias of up to -39% at iPTH concentrations &lt;20 pg/mL with serum and EDTA patient samples vs. a reference kit lot. | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Recall" letter IMC16-27.A.US (dated 11/22/2016) to its customers on 11/22/2016 via FedEx. Customers Outside the US were sent an "Urgent Field Safety Notice" dated 11/2016. The letter/notice describes the product, problem and actions to be taken. Customers were instructed to discontinue use of and discard the affected kit; review your inventory of these products to determine your laboratorys replacement needs; Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days by fax to the Customer Care Center at (312) 275-7795; and retain this letter with your laboratory records, and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 2728 | 2728 kits (US: 565 kits; OUS: 2163 kits) | Worldwide Distribution: US (Nationwide) and countries of: Canada, China, Chile, Mexico, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg, Brazil, Bulgaria, Peru, Colombia, Ecuador, Bolivia, Costa Rica, Honduras, Croatia, Czech Republic, Denmark, Egypt, Estonia, France, Algeria, Greece, Guatemala, Hungary, India, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Netherlands, Norway, Pakistan, Panama, Paraguay, Poland, Portugal, Republic Korea, Romania, Russian Fed., Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, Jordan, U.A.E., United Kingdom and Uruguay. | Open, Classified |  |  | 2728 | Not_Computer
Z-0975-2017 | 76033 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | The OEC 9900 Elite mobile fluoroscopy system is designed to provide fluoroscopic and spot-film images of the patient during diagnostic surgical and interventional procedure. Examples of clinical application may include cholangiography, endoscopy, urologic, orthopedic, neurologic, vascular, cardiac, critical care and emergency room procedure. The system may be used for other imaging applications at the physician's discretionOEC 9900 Elite. MDL Numbers: D148942, D155043 |  |  | December 19, 2016 | 2016 | GE OEC Medical Systems, Inc | GE Healthcare Surgery announces a voluntary field action for the OEC MiniView 6800, OEC 9900 Elite, and OEC 9800 because the workstation power cable assembly can result in intermittent power cord failures. | Use error | GE Healthcare sent an Urgent Medical Device Correction letter dated December 19, 2016, to all affected consignees. GE informed consignees that GE will correct all affected products. If consignnes have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or their local Service Representative.For questions regarding this recall call 801-536-4952. |  |  |  | Worldwide Distribution - US including U.S. States: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS,NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PUERTO RICO, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Internationally to Canada, Mexico, and Costa Rica | Open, Classified |  |  |  | Not_Computer
Z-0967-2017 | 75952 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SURESTEP Tray, BARDEX I.C. COMPLETE CARE, Urine Meter, Temperature STATLOCK Stabilization Device |  |  | July 11, 2016 | 2016 | C.R. Bard, Inc. | The product contained a latex catheter instead of a silicone catheter. | Labeling mix-ups | Bard Medical sent an Urgent : Medical Device Product Recall Letter dated July 11, 2016, to all affected customers. The letter requested that they check inventory and quarantine product, and if the product had been further distributed, they need to conduct a sub-recall. The letter included a reply form to be returned. For further questions, please call (770) 784-6220. |  | 2690 | 2690 units | US Distribution to the states of : CA, CT, FL, GA, IL, IN, KS, KY, MD, MA, MI, NE, NV, NJ, NM, NY, OR, SC, TN, TX, UT, VA, WA | Open, Classified |  |  | 2690 | Not_Computer
Z-0966-2017 | 75857 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ceres-C Standalone Cervical Interbody |  |  | November 22, 2016 | 2016 | Amendia, Inc | Absent tantalum market pin. | Process control | The firm, Amendia, contacted the consignee and issued a Return Material Authorization for the affected product on 8/18/2016.Please contact Regulatory and Quality Consultant for Amendia at 404-610-7215 or by email to bbrosseau@securebme.com with any questions or concerns regarding this recall notification. |  | 6 | 6 devices | US Distribution to GA only. | Open, Classified |  |  | 6 | Not_Computer
Z-0961-2017 | 75838 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | GoDirect Screw-Retaining Top |  |  | November 28, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | Consignees received product 6060-62, Replant Overdenture Abutment with the label for 4500-61 GoDirect Screw-Retaining Top. | Labeling mix-ups | On 11/28/16 customers were called and instructed to return the product. |  | 14 | 14 | U.S. distribution to the following; WA.Foreign distribution to the following; Netherlands. | Open, Classified |  |  | 14 | Not_Computer
Z-0980-2017 | 75895 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 133" 60 Drop 24 Hour Primary Set w/2 Bravo24Âż, 2 Gang 1o2Âż Manifold, Rotating Luer, 1 Ext, Item No. B59006 |  |  | November 30, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified an isolated event involving a 15 drop drip chamber assembled on a device which should contain a 60 drop drip chamber. | Mixed-up of materials/components | An Urgent Medical Device Recall Notification letter dated 11/30/16 was sent to their customers to inform them that ICU Medical Inc. has identified an isolated event involving a 15 drop drip chamber assembled on a device which should contain a 60 drop drip chamber. The letter instructs the customers of the actions required and to return completed recall response form to ICU Medical via fax (801) 264-1755 or by email recall@icumed.com. Customers with questions are instructed to contact ICU Customer Service Monday through Friday between 8:30AM and 4PM Pacific Time, (866) 829-9025 and select option 2 or email productreturns@icumed.com. |  | 225 | 225 units | US: NC | Terminated | June 23, 2017 | 205 | 225 | Not_Computer
Z-0984-2017 | 75948 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Medline E-Z Lubricating Jelly; Bacteriostatic. Water Soluble. Sterile. 2 FL OZ (59 ml). |  |  | December 12, 2016 | 2016 | MEDLINE IND | Product was not sterilized. Product was shipped to distribution centers instead of the sterilization facility due to an operations error by a 3rd party shipping company. | Process control | Medline Industries sent a recall notification letters dated December 12, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to quarantine any affected product on hand and return product to Medline Industries for credit. Distributors were instructed to notify their customers and have them return affected product as well. For questions contact 866-359-1704. |  | 12024 | 12,024 tubes | Worldwide Distribution - US Nationwide in the states of AL, FL, GA, IA, IL, IN LA, MI, PA, TX, WI, and the country of Canada | Open, Classified |  |  | 12024 | Not_Computer
Z-0982-2017 | 75908 | No | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StatLockÂż CV Plus w/Pigtail. Product Code: CV0220CE &amp; CV0220. Packaged as a single unit with a skin protectant pad and foam strips in a Tyvek pouch. Depending on configuration, there are either 25 or 50 pouches per carton and 10 cartons per case. Other configurations may contain additional components such as a foam pad with a pigtail clip. |  |  | December 09, 2016 | 2016 | Bard Access Systems | Bard Access Systems announces a voluntary field action for the StatLockÂż Catheter Stabilization Devices because some single unit pouches have been released without unique product identifiers (product code, lot number, expiration date, etc.). | Process design | Bard sent an Urgent Customer Notification letter dated December 12, 2016. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory and identify any product subject to the notification. Please direct any return requests and/or questions to Shelly Gilbert at 1-800-290-1689. |  | 14500 | 14,500 units total | Worldwide Distribution - US Nationwide | Terminated | May 26, 2017 | 168 | 14500 | Not_Computer
Z-0989-2017 | 76102 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | i.v. Station and i.v. Station ONCO Automated Drug Compounding System, Pharmacy Compounding System |  |  | November 23, 2016 | 2016 | HEALTH ROBOTICS S.R.L. | The firm has initiated a product correction due to a bag carousel motor error causing misalignment and resulting in a mislabeling of final IV bag preparations. | Software design | Omincell mailed an Urgent Medical Device Field Correction Notification letter to affected customers on 11/21/2016. The letter provided information of the event and instructions on how to continue operating the device to prevent risks. The issue will be corrected with a software update. |  | 121 | 121 systems | Shipped throughout United States. | Terminated | March 15, 2017 | 112 | 121 | Computer
Z-0988-2017 | 75093 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Tego Connector, Item No. D1000, NM1000 |  |  | August 24, 2016 | 2016 | ICU Medical, Inc. | ICU Medical Inc. has identified a potential risk of leaking with certain Tego Connector devices. | Under Investigation by firm | The firm, icumedical, sent an "URGENT: Medical Device Recall Notification" letter dated 8/24/16 to all their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to quarantine any affected devices; remove from use, return to ICU medical, Inc.; and complete and return the recall response form to ICU Medical via fax (801) 264-1755 or at recall@icumed.com (even if you do not have any affected product). Customers with questions are instructed to contact ICU customer service Monday through Friday between 8:30AM and 4:00PM Pacific Time, (866) 829-9025 and select option 8 or email the following: customerservice@icumed.com. |  | 476600 | 476,600 units | Worldwide Distribution-US: (Nationwide) and countries of: Saudi Arabia, Germany, Austria, Belgium, Estonia, Poland, Switzerland, Canada, Luxembourg, Netherlands, Bulgaria, Malaysia, Italy, Ireland, Great Britain, France, Brazil, Sweden, Slovenia, and Portugal. | Terminated | July 07, 2017 | 317 | 476600 | Not_Computer
Z-0987-2017 | 76085 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam Oncentra Brachy |  |  | December 23, 2016 | 2016 | Elekta, Inc. | Cross profile for Varian 60 degree wedge shows "horns." | Software design | Elekta sent an Urgent Important Safety Notification letter dated December 23, 2016 to affected customers. The notice informs users of the specific product and version numbers affected by the issue and any work arounds that can be used to avoid the issue. The letter included an acknowledgement form to be returned to Elekta. |  | 433 | 433 | Worldwide Distribution - US Nationwide in the states of AL, CA, DC, NE, NY, ND, OR, PA, RI, SC, TX, and the countries of Angola, Argentina, Armenia, Australia, Austria, Barbados, Belarus, Brazil, Bulgaria, Canada, China Columbia, Curacao, Czech Republic, Denmark, Finland, France, Germany, Greece, Honduras, Hungary, Iceland, India, Indonesia, Ireland, Italy, Japan, Kazakhstan, Kenya, Malaysia, Netherlands, Norway, Oman, Peru, Poland, Portugal, Romania, Russia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Yemen, Zambia. | Open, Classified |  |  | 433 | Not_Computer
Z-0991-2017 | 75969 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConFORMIS ITOTAL PS-IPOLYIMPLANT KIT-RIGHT |  |  | October 31, 2016 | 2016 | ConforMIS, Inc. | iView report displayed femoral cut measurements in the incorrect unit of measure (inches vs. mm). | Software Design Change | ConForMIS intiated the recall by telephone on October 31, 2016, to the Sales Rep to request return of the kit.For questions regarding this recall call 781-345-9001. |  | 1 | 1 unit | Distribution to TX only | Terminated | January 11, 2017 | 72 | 1 | Not_Computer
Z-0986-2017 | 76008 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SafeStepÂż Huber Needle Set with Y-Injection Site. Product Code: LH-0029YN. Packaged as single units in thermoformed trays. There are 25 SafeStepÂż trays per case and 100 per carton. |  |  | December 19, 2016 | 2016 | Bard Access Systems | Bard Access Systems announces a voluntary field action for the SafeStepÂż Huber Needle Set with Y-Injection Site and MiniLocÂż Safety Infusion Set Winged without Y-Injection Site devices because sets may contain one or more units where the needle length and gage differs from the length and gage described on the label. | Mixed-up of materials/components | Bard sent an Urgent Notification letter dated December 16, 2016 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to identify and return affected product. For questions call 1-800-290-1689. |  | 8000 | 8000 units | Worldwide Distribution - US Nationwide in the states of CA, CO, FL, GA, HI, IL, MA, MD, NJ, OH, OR, TX. | Terminated | April 17, 2017 | 119 | 15200 | Not_Computer
Z-0983-2017 | 75915 | No | OHW | Powered corneal burr. | Ophthalmic | Ophthalmic | 510(K) Exempt | Algerbrush-II |  |  | October 12, 2016 | 2016 | Bausch &amp; Lomb, Inc. | According to firm ( Bausch &amp; Lomb, Inc. ) on June 17, 2016 a packaging operator notified his supervisor he had not placed the operating instruction into the last repackaging of SKU 0816 C (Pterygium Burr Medium Diamond). Further investigation found that repackaged Algerbrush-II and its accessories from April 30, 2012 to present, contained out of date operating instruction. Therefore Bausch &amp; Lomb is proposing a Field Correction to provide customer purchasing product from April 30, 2012 to June 20, 2016 with the current operating instruction. | Error in labeling | Bausch + Lomb sent an Urgent Medical Device Voluntary Field Correction letter dated October 12, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Bausch + Lomb customer service at 1-800-338-2020 |  | 2487 | 2487 units Domestic (USA), 1,211 units Foreign (outside USA) | Worldwide Distribution - US Nationwide | Terminated | July 14, 2017 | 275 | 2487 | Not_Computer
Z-0990-2017 | 76018 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System version 13.MR2 [13.06.31 with Smart Segmentation Knowledge Based Contouring version 2.3 [2.3.12] |  |  | December 02, 2016 | 2016 | Varian Medical Systems, Inc. | Modifications in version 13.6MR2 for Contouring, SmartAdapt, and SmartSegmentation workspaces resulted in contours not being saved consistently in Eclipse. Treatment Planning System. The issue only occurs if certain conditions are fulfilled. | Device Design | Varian Medical Systems sent an Urgent Medical Device Correction Letter dated December 5, 2016 , that were sent to all 5 affected sites on December 5, 2016 by a trackable method.Letters advised users of the issue and the specific scenario that must take place for the issue to occur. Varian is investigating the issue and will provide a correction once it is available.We sincerely apologize for any inconvenience and thank you in advance for your cooperation. If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager. This notice has been provided to the appropriate Regulatory Authorities. For further questions please call (650) 424-6833.Varian Oncology Help Desk Contact Information:Phone:+1.888.827.4265Email:support @varian.com |  | 42 | 42 devices are affected | US Distribution to the states of : NJ, TN. OR and FL. | Terminated | May 12, 2017 | 161 | 42 | Computer
Z-0995-2017 | 76127 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT S(40)-4R, Material Number 10248671 |  |  | December 13, 2016 | 2016 | Siemens Medical Solutions USA Inc. | Siemens Medical Solutions, Molecular Imaging has become aware, through our supplier that some laser cover windows may fall off. | Device Design | Siemens Healthineers sent a Customer Safety Advisory Notice letter on December 13, 2016. The letter states that the light marker windows used for the positioning laser and integrated into the front cover of the Biograph mCT or Biograph mCT Flow system may loosen and potentially drop out and warning the consignee not to use the system of the front cover is loose or missing. Customers with questions were instructed to contact their local Siemens representative at the following numbers:America 1-800-888-7436Europe, Middle East and Africa +49 9131 940 4000Asia and Australia +86 (21) 3811 2121For questions regarding this recall call 865-218-2000. |  | 88 | 88 units | Worldwide Distribution - US (nationwide) and Internationally to U.A.E., AT, AU, AZ, BR, CA, ,CH, CL, CN, CO, CZ, DE, DK, EG,FI, FR, GB, HK, HR, IL, IN, IR, IT, JP, KR, MX, NL, NO, NZ, PK, PL, QA, RO RU, SE, SG, SI, SK, TH, TR, AND TW. | Terminated | May 08, 2017 | 146 | 409 | Not_Computer
Z-1013-2017 | 76044 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Catalog No. 629029, B24465, B24802, B68304 |  |  | December 19, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter has determined that additional clarification for the Blast Suspect messages is necessary. In rare situations, the UniCel DxH 800 and DxH 600 Coulter Cellular Analysis System may not flag or detect blasts in some blood samples. This is due to limitations in the available technology as well as sample limitations. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated December 19, 2016, to all affected customers to inform them that Beckman Coulter has determined that additional clarification for the Blast Suspect messages is necessary. In rare situations, the UniCel DxH 800 and DxH 600 Coulter Cellular Analysis System may not flag or detect blasts in some blood samples. The letter informs the customers of the impact, actions to be taken, and resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Center: http://www.beckmancoulter.com/customersupport/support, 800-526-7694 in the United States and Canada, and outside the United States and Canada, to contact their local Beckman Coulter Representative. |  | 4657 | 4,657 units total (2,067 units in US) | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the countries of : Australia, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Columbia, Croatia, Czech Republic, France, Germany, Gibraltar, Ecuador Hong Kong, India, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, New Zealand, Oman, Panama, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay and Vietnam. | Terminated | May 12, 2017 | 144 | 4657 | Not_Computer
Z-1016-2017 | 64220 | No | FPK | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Minibore Extension Set, 59 Inch with Spin Lock Collar, Non-DEHP, Latex-Free |  |  | February 19, 2013 | 2013 | Hospira Inc. | Hospira, Inc. is voluntarily recalling one lot of list 14699-28, Minibore Extension Set, 59 inch with Spin Lock Collar, Non-DEHP, identified above due to a confirmed customer report of an occlusion of solvent in the male luer lock. This occlusion has the potential to prevent the set from being primed for use and cause a delay in therapy. | Process control | Hospira sent an Urgent Device Recall letter dated February 19, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please check your inventory and immediately quarantine any affected product. Complete the attached reply form and return it to the fax number or e-mail address on the form, even if you do not have the affected product. Inform healthcare professionals in your organization of this recall.Return affected product to Stericycle using the label provided with this letter. If you have not received a return label or require additional assistance contact Stericycle at 1-888-965-5821 between the hours of Bam to 5pm EST, Monday through Friday. To ensure proper and timely credit, follow the instructions on the return label for returning the product. Please visit http://expertezlabel.com to request additional labels for returning affected product.If you have distributed the product further, notify your accounts that received the product identified above of this recall and ask them to contact Stericycle to receive a reply form and return labels for returning the product.Please contact Hospira Customer Care at 1-877-946-7747, between the hours of 6am to 6pm CST, Monday through Friday, or your Hospira representative regarding replacement product availability. For further questions please call (224) 212-2000. |  | 48050 | 48,050 units | Nationwide Distribution | Terminated | April 24, 2017 | 1525 | 48050 | Not_Computer
Z-1015-2017 | 75210 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Infant Warmer System (IWS) |  |  | August 04, 2016 | 2016 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related to loose screws in the Heater Head of the Infant Warmer System (IWS). Hot screws fromthe Heater Head of the IWS could fall onto the bed if the Heater Head assembly has been improperly serviced. This situation can be clinically hazardous because thermal injury to a patient could result. Two injuries have been reported as a result of this issue. | Use error | This recall strategy includes the depth of the recall, plan for public warning and effectiveness check. The strategy outlined below addresses these elements for this particular recall. The plan for public warning is to contact all customers with potentially affected product by means of an Urgent Medical Device Correction letter. The letter will instruct customers that they can continue to use the device by following instructions on the Addendum to the Service Manual.The Service Manual Addendum provides instructions for checking and tightening the screws. In the US the letter will be sent by a traceable means [either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX]. This letter will be sent to: Nurse Manager, Labor &amp; Delivery/NICU; Manager, Bio-Medical Engineering; andDirector of Risk Management. |  | 32073 | 32,073 (US = 24,361; OUS = 7,712) | This product has been distributed domestically and internationally . | Open, Classified |  |  | 32073 | Not_Computer
Z-1008-2017 | 75751 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REFLECTION(R) DEAD BLOW MALLET, REF 71362106, QTY: (1), NON-STERILE, Smith &amp; Nephew, Inc,. Memphis, TN 38116 |  |  | November 16, 2016 | 2016 | Smith &amp; Nephew, Inc. | The firm received complaints of cracks in the weld on the head of the mallet. In the reported cases, some of the lead beads and particles escaped from the mallet into the surgical wound. | Process control | Smith and Nephew sent an Urgent - Product Recall 1 st Notification dated November 16, 2016, to all affected customers.Actions for Hospital Representatives:1. Please inspect your inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately.2. Complete the last two columns in the Inventory Return Certification Form on the following page, indicating the quantitiesthat need to be returned. If you do not have product to return, please place an x in the column No Product to Return. Include your contact information in the spaces provided at the bottom of the form.3. Please contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1- 901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return.4. Return any affected product to the address listed on the Inventory Return Certification Form. Please indicate the RA number on your return shipment.Actions for Smith &amp; Nephew Affiliates/Distributors1.Send this letter to your customers and ask them to return the recalled product.2. Please inspect your inventory and locate any unused devices from the above listed product and batch numbers, cease distribution and use of those products, and quarantine them immediately.3. Complete the last two columns in the Inventory Return Certification Form above, indicating the quantities that need to be returned. If you do not have product to return, please place an x in the column No Product to Return . Include your contact information in the spaces provided at the bottom of the form.4. Please contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1- 901-566-7975 to return the completed Inventory Return Certification Form and obtain a return authorization (RA) number if you have product to return.5. Return any affected product |  | 4444 | 4444 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the countries of : Argentina, Austria, Australia, Belgium, Bahrain, Brazil, Canada, China, Colombia, Costa Rica, Dubai, Denmark, Germany, Finland, France, Great Britain, Greece, Hong Kong, Ireland, Israel, India, Italy, Japan, Mexico, Malaysia, Netherlands, Norway, Poland, Portugal, Puerto Rico, South Africa, Sweden, Singapore, Switzerland, Thailand, Turkey and Venezuela. | Open, Classified |  |  | 4444 | Not_Computer
Z-1003-2017 | 75810 | No | ONB | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Halo Closed System Transfer Device, closed syringe adaptor, REF CSA100, Rx only, Sterile, packaged as individual units, packed and shipped 60 units/box and 4 boxes/shipping carton Corvida Medical, Coralville, IA 52241. |  |  | November 22, 2016 | 2016 | J&J; Solutions, Inc. DBA Corvida Medical | Product malfunction: Needle does not retract. | Under Investigation by firm | Corvida Medical sent an Urgent Notice: Medical Device Recall letter dated November 22, 2016, to all affected customers.After completing a Health Hazard Evaluation, recalling firm drafted a Recall Notification by Phone script. on November 21 2016, utilizing the script, the Vice President of Sales contacted all parties (customers) who received the recalled lot. He confirmed each customers understanding of the recall notification and forwarded the contact records to the Recall Coordinator. If you have further questions , please contact your Sales Representative or Corvida Customer Service at (800) 651-6832. |  | 5520 | 5520 units | US Distribution to the states of : MD, TX, and WI. There was no foreign/military/government distribution. | Open, Classified |  |  | 5520 | Not_Computer
Z-1010-2017 | 75850 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | symmetry surgical ULTRA(R) Steam container, REF numbers: C131310, C13134, C13135, C13136, C13139, C13854, C13855, C13856, C181310, C18134, C18135, C18136, C18138, C22854, C22855, C22856, C22858, C24134, C24135, C24136 |  |  | November 29, 2016 | 2016 | Symmetry Surgical, Inc. | Sealant on the base trim of sterilizer container may be deteriorated or absent which could provide a path for contaminants to enter the container under certain conditions post-sterilization. | Nonconforming Material/Component | Siemens sent a Customer Safety Notice letter to customers. The letter identified the affected product problem and actions to be taken. Siemens Customer Service Engineer will contact customers to setup an appointment for installation of the new residual current monitor (RCM). |  | 7613 | 7613 units | Worldwide Distribution | Terminated | August 20, 2017 | 264 | 7800 | Not_Computer
Z-1005-2017 | 75994 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | FV-RF (Clinix RF Plus) |  |  | November 18, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | There's a potential for liquid penetration into the back side of the front control panel of the X-ray diagnostic table, which can cause a short circuit of the terminals of the tilt operation and cause an unintentional tilt of the table. | Component design/selection | Philips sent an Urgent Field Safety Notice Medical Device Correction letter dated December 2, 2016 to customers via USPS certified letter with returned receipt. The letter identified the affected product, problem and actions to be taken. Additionally, your local Philips representative has been provided details regarding this Device Correction and can be contacted if you need any further information or support concerning this issue: Technical Support Line: 1-800-722-9377 |  | 41 | 41 | US Nationwide distribution in the following states: AL, CA, FL, GA, HI, IL, IN, KS, LA, MA, MI, MS, NJ, NY, OH, PA, SC, VA, WI, WV, &amp; WY. | Open, Classified |  |  | 41 | Not_Computer
Z-1004-2017 | 76021 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT |  |  | November 25, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Failure to correctly document the installation of four M12 Bolts into the system rotor. | Employee error | Philips sent a Customer Information Letter to their customer. The letter identified the affected product, problem description and actions to be taken by customers. For further information or support concerning this issue, contact your local Philips representative or local Philips Healthcare office. For SUPPORT North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377). |  | 1 | 1 | The one system was distributed in Canada. There were no government or US distribution. | Open, Classified |  |  | 1 | Not_Computer
Z-1006-2017 | 76024 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IQon Spectral CT with software version 4.7.0 |  |  | December 16, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | Multiple issues have caused the device to result in CT rescans or incorrect scan location or misrepresentation of image results. | Software design | Philips sent a Customer Information Letter dated December 13, 2016 to customers via certified mail. The letter identified the affected product, problem and actions to be taken. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377). |  | 35 | 35 | Worldwide Distribution - US Nationwide in the states of AR, IN, LA, MA, MN, OR, TN &amp; TX. and the countries: Australia, Belgium, Denmark, France, Germany, Israel, Japan, Republic of Korea, Poland, Portugal, Slovenia, Switzerland, Thailand and United Kingdom. | Open, Classified |  |  | 35 | Not_Computer
Z-1009-2017 | 76028 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | December 21, 2016 | 2016 | Elekta, Inc. | Incorrect Enhanced Dynamic Wedge (EDW) or Virtual Wedge (VW) Calculations. | Nonconforming Material/Component | Elekta sent an Urgent Important Field Safety Notification letter dated December 2016 to customers. The notice informs users of the specific product and version numbers affected by the issue and any work arounds that can be used to avoid the issue. The notice included a reply form to be returned to Elekta. |  | 279 | 279 | Worldwide Distribution - US Nationwide in the state of CA, IN, KY, MA, NJ, NY, NC, PA, TX, WI, and the countries of Brazil, Canada, China, Croatia, Estonia, Germany, India, Japan, Netherlands, New Zealand, Norway, Poland, Serbia, South Africa, Spain, Russia, Turkey | Open, Classified |  |  | 279 | Not_Computer
Z-1007-2017 | 76060 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | AxIEM Touch N Go Pointer Probe. Model Number: 9735318. Placed in a 30mil PETG tray sealed with a Tyvek lid and then placed in a shelf carton. |  |  | December 12, 2016 | 2016 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. announces a voluntary field action for the AxIEM Touch N Go Pointer Probe because the patient registration accuracy may compromise navigation accuracy when using the StealthStationÂż System with Touch N Go Pointer Probes for patient registration. | Process change control | The firm, Medtronic, sent an "URGENT: MEDICAL DEVICE RECALL" letter on 12/12/16 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine inventory for any affected products, quarantine them for return to Medtronic; complete and return the CUSTOMER RESPONSE FORM by email at RS.NavFCA@medtronic.com or fax to Medtronic Technical Services at 651-367-7075 (even if you do not have affected product); and contact Medtronic Technical Services at 1-800-595-9709 to get a return material authorization (RMA) and arrange for their return and no charge replacements.If you have any questions regarding this communication, please contact Medtronic Technical Services at (800) 595-9709. |  | 111 | 111 devices | US Nationwide Distribution. | Terminated | May 19, 2017 | 158 | 111 | Not_Computer
Z-1020-2017 | 75889 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reunion TSA Peg Alignment Sound |  |  | December 02, 2016 | 2016 | Stryker Howmedica Osteonics Corp. | Reunion TSA Peg Alignment Sound broke during surgery. | Device Design | Stryker Orthopaedics sent an Urgent Medical Device Recall Notification letter dated December 2, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to inform users of the Urgent Medical Device Recall and forward the notice to all those individuals who need to be aware within their organization. Return all affected products available to: Stryker Orthopaedics/PFA Product ReturnsAttn: Distribution Inventory Team325 Corporate DriveDock M-EastMahwah, NJ 074311360184/RA2016-112Complete and sign the enclosed Business Reply Form and fax a copy to 855-261-3635 or email to SO_M_PRODUCT_FIELD_ACTION_RESPONSE@stryker.com.Customers with questions were instructed to call 201-831-6693. |  | 131 | 131 units | Nationwide Distribution | Open, Classified |  |  | 131 | Not_Computer
Z-1017-2017 | 75602 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Merge Eye Station |  |  | December 09, 2016 | 2016 | Merge Healthcare, Inc. | This recall has been initiated due to an issue related to the potential accidental deletion of record(s) by an Eye Station user. | Software design | Consignees were sent on 12/9/2016 a Merge "Urgent Medical Device Recall" letter. The letter was addressed to Office Manager and Recall # 216-078. The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Action by Merge and Actions by Customer. Requested consignees to provide response to the "Response to this Notification is Required". If the product was further distributed advised consignees to notify them at once. For questions send an e-mail to recall@merge.com. E-mails to be sent to consignees on 12/19/2016. |  | 1597 | 1597 (1451 US; 146 OUS) | US: Alabama, Alaska, Arizona,Arkansas, California,Colorado,Connecticut,Delaware,District of Columbia, Georgia,Hawaii, Idaho, Illinois,Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland,Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, Rhode Island, South Carolina, South Dakota, Texas, Utah, Vermont, Virginia, Washington, West Virginia, Wisconsin, Florida, Tennessee, Wyoming.OUS:Czech RepublicAlgeriaBelgiumBrazilCanadaDenmarkFranceGermanyGreeceIndiaIsraelItalyKorea, Republic ofLuxembourgNetherlandsPolandRussian FederationSingaporeUnited Kingdom | Open, Classified |  |  | 1597 | Not_Computer
Z-1028-2017 | 76037 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008K2 Series : Hemodialysis System |  |  | December 21, 2016 | 2016 | Fresenius Medical Care Renal Therapies Group, LLC | When the UF Rate, Goal or Time is adjusted using the up and down arrow keys, and the change is cancelled by using the esc key, the cancelled UF Rate is actually being executed rather than rate displayed on the machine | Device Design | Fresenius issued Customer Notification via Certified Mail on 12/21/2016. Users provided with problem and recommendation, when adjusting the UF Rate, UF Goal, or Tx Time, is to use the numeric data entry keys instead of the up or down arrow keys. Change the values by using data entry keys, then save by pressing CONFIRM. Fresenius will make arrangements to install software on the hemodialysis machine. Confirmation of the noticed by the consignee will be required via a Reply Form. |  | 60151 | 60,151 units (US) 1773 (OUS) | NationwideForeign: Canada, Mexico, Jamaica, Germany, Guam | Open, Classified |  |  | 205393 | Computer
Z-1023-2017 | 76034 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Medacta International SA, M.U.S.T. Rod, REF 03.50.433, Size 5.5mm, Length 480mm, Type Straight, Mat Co-Cr-Mo |  |  | December 07, 2016 | 2016 | Medacta Usa Inc | Damaged packaging, lack of sterility assurance | Packaging | The firm, Medacta International, immediately notified all customers by phone on 12/07/16 and further notified via email on 12/9/16. The firm also issued an "URGENT- Field Safety Notice" dated 7/12/2016 to all its customers. The notice described the product, problem and actions to be taken. The customers were instructed to return all units to Medacta International SA and forward the notice to all people and organizations potentially affected.If you have any questions contact Regulatory Affairs Manager at +41 (0) 91 696 60 60. |  | 82 | 82 units | Worldwide Distribution- US (nationwide) distribution to states of: TX and CO; and countries of: Australia, Belgium, Spain, German, Italy and Switzerland. | Open, Classified |  |  | 82 | Not_Computer
Z-0976-2017 | 76026 | general purpose reagent | PPM | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | MagSIL (NucliSENS easyMAG Magnetic Silica), |  |  | November 23, 2016 | 2016 | BioMerieux SA | A drift of the performance with the BK test (BK virus, worst case application) was observed for some silica batches of Magnetic Silica (MagSil). This is the same issue as in recall Z-2342-2016. | Nonconforming Material/Component | Consignees were notified via, URGENT PRODUCT SAFETY CORRECTION NOTICE on/about 11/23/2016. This is the same issue as in recall Z-2342-2016, notification letter sent on/about 7/8/2016. |  | 3723 | 3723 units in total | Worldwide distribution. US Nationwide; BioMerieux Subsidiaries in: Argentina, Austria, Australia, Brazil, Canada, Switzerland, Chile, Colombia, Czech republic, Germany, Spain, France United Kingdom, Greece, China, Hungary, India, Italy, Japan, South Korea, Netherlands, Poland, Portugal, Russia Fed., Sweden, Singapore, Thailand, Turkey, South Africa; Distributors in: Peru, Kenya, Morocco, Serbia, Romania, Uruguay, Latvia, Dutch Antilles, Israel, Georgia, Armenia, Malaysia, Belarus and Direct Customers in: Togo, Djibouti, Burkina Faso, Mongolia, Zimbabwe, Uganda and Iraq. | Open, Classified |  |  | 3723 | Not_Computer
Z-1025-2017 | 75970 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Active Articulation ArComXL Polyethylene Bearings |  |  | December 19, 2016 | 2016 | Zimmer Biomet, Inc. | Product complaints indicated that the part and lot numbers on the packaging differed from the physical product inside the packaging. Risk of delay in surgery. | Under Investigation by firm | Zimmer Biomet sent an Urgent Medical Device Recall (REMOVAL)  LOT SPECIFIC notifications dated December 19, 2016, to the affected distributors and Risk Managers and Distributors via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Risk Manager Responsibilities: 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative quarantine all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility. 4. Complete Attachment 1  Certificate of Acknowledgement. a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com. b. Retain a copy of the Acknowledgement Form with your recall records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. Distributor Responsibilities 1. Review this notification and ensure that affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification |  | 19 | 19 | USA (nationwide) and Internationally to ARGENTINA | Open, Classified |  |  | 55 | Not_Computer
Z-1018-2017 | 75898 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Optilite Analyser, Product Code: IE700 |  |  | May 13, 2015 | 2015 | The Binding Site Group, Ltd. | Potential risk of the Optilite lid/cover falling suddenly and causing injury when not placed in the fully open position. | Component design/selection | Binding Site sent an Urgent Field Safety Notice letter dated May 13, 2015 to their customers The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the Eback Form that this notice has been received and the actions to be taken have been completed to technical.services@bindingsite.co.uk |  | 129 | 129 units | Worldwide Distribution - US Nationwide in the states of: CA, VA, GA, PA, NJ, KY, NY, AR, NM, NC, AL, GA, MI, and the countries of Canada, China, Australia, Israel, Ireland, UK, Germany, Belgium, Italy, France, Czech republic, Spain, Portugal. | Completed |  |  | 129 | Not_Computer
Z-1033-2017 | 75651 | No | GNA | Escherichia coli serological reagents. | Microbiology | Microbiology | 510(K) Exempt | PREMIER EHEC and PREMIER EHEC Bulk. |  |  | November 08, 2016 | 2016 | Meridian Bioscience Inc | Internal testing has demonstrated that the indicated kit lots produce false positive results when tested with some broth specimens. | Under Investigation by firm | On November 7, 2016 the firm sent Voluntary Medical Device Recall Letters and reply forms to their customers. |  | 4927 | 4,747 EHEC + 180 EHEC Bulk | Product was shipped to the following states: AL. AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WY.Product was also sent to the following countries: CANADA, GERMANY, HONG KONG, ITALY, JAPAN &amp; NEW ZEALAND. | Terminated | August 16, 2017 | 281 | 4927 | Not_Computer
Z-1030-2017 | 75989 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Novation Cemented Plus Femoral Stem, 12/14, Collared, Standard Offset, Matte, Cemented, Size 15. |  |  | November 07, 2016 | 2016 | Exactech, Inc. | Products were labeled with incorrect Global Trade Item Numbers. The GTIN on the label identifies the devices as another. | Under Investigation by firm | In the recall notice sent to agents, dated 11/01/2016, recipients are instructed to cease product distribution, extend knowledge of the recall to their accounts, identify and quarantine any product in inventory, and complete and return the Recall Inventory Response Form. Consignees are instructed to contact their Exactech Representative within 5 days of receipt of the notice to confirm product quantities at each location. Consignees with any questions are instructed to contact Kaya Davis at kaya@exac.com or 1-800-392-2832. |  | 79 | 79 units | AZ, CA, CO, FL, IL, KS, MD, NJ, NY, OH, OK, TN, &amp; VA; Germany, Italy, Japan, Spain, &amp; Switzerland. | Open, Classified |  |  | 189 | Not_Computer
Z-1034-2017 | 76039 | No | EID | Dental hand instrument. | Dental | Dental | 510(K) Exempt | NanoTips 30g 20pk, 100pk PN 1240, 1426 |  |  | December 14, 2016 | 2016 | Ultradent Products, Inc. | Ultradent Products, Inc. announces a voluntary field action for the NanoTips 30g 20pk, 100pk because the cannula attached to the tip hub may become loose or detach from the blue base that connects to the syringe. | Device Design | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated December 14, 2016, to all affected consignees. The letter instructed consignees to discard remaining inventory and return the response form. For any questions, please call at 801-553-4858. |  | 14300 | 143 packs of 100 tips; 104 packs of 20 tips | Worldwide Distribution - US (Nationwide) - Internationally to Belgium, Switzerland, Germany, Italy, South Korea, Netherlands, and Japan. | Terminated | June 05, 2017 | 173 | 14300 | Not_Computer
Z-0950-2017 | 75466 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module, Model No. 8100 and AIL sensor kits, P/N 147083-102 and P/N 49000221. |  |  | December 02, 2016 | 2016 | CareFusion 303, Inc. | During an infusion, a false Air-in-Line (AIL) alarm would cause the infusion to be interrupted. | Component change control | Carefusion (BD) sent a Medical Device Safety Notification letter dated December 2, 2016, to all affected customers. The letter informed customers that BD has identified an issue with the Alaris Pump module. The letter informs the customers that the firm has received reports of Air-In-Line (AIL) alarms that have occurred when no air is observed in the line. Customers were informed of the issue and of the potential risk involved. The letter informs the customers of the actions for users and additional actions for biomedical engineering. Customers with questions regarding recall related questions were instructed to contact BD Support Center at (888) 562-6018, 7-4pm PT, M-F or email supportcenter@carefusion.com. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266, 24 hrs/day, 7days/wk. or email customerfeedback@carefusion.com.For questions regarding this recall call 858-617-2000. |  | 382635 | 382,635 units | Nationwide Distribution | Open, Classified |  |  | 382635 | Not_Computer
Z-1038-2017 | 76110 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type IIIB; |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX EndovascularAAA System (AFX System). | Other | A Dear Physician letter dated 6/2/15 was sent to customers to inform them that the 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The letter provides update reports for the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System).On January 2016, a Dear Physician letter was sent to customers to inform them of the 2015 Clinical Update covering the Endologix Endovascular Systems for AAA Repair. The letter provided update reports on the cumulative clinical data and commercial experience (i.e., post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safeand effective use of these devices.A Dear Physician letter dated 12/27/16 will be sent to inform customers of the important information related to the AFX Endovascular AAA System (AFX System). The letter provides a series of updates to the Instructions for Use (IFU) and modifications to the product were implemented, including introduction of a graft material processing improvement known as Duraply.Customers with any questions are instructed to contact their Endologix representative. |  | 40967 | 40967 | US and worldwide: Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Terminated | April 28, 2017 | 122 | 40967 | Not_Computer
Z-1037-2017 | 76086 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type IIIA |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX Endovascular AAA System (AFX System). | Device Design | Endologix sent a dear physician letter dated June 2, 2015, to inform them of the enclosed 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The update reports the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System). The letter informs the Physicians of the year's update on the clinical programs and commercial experience (i.e. post-market surveillance).A dear physician letter dated January 2016 was seen to inform them of the enclosed 2015 Clinical Update covering the Endologix Endovascular Systems for AAA Repair. The update reports the cumulative clinical data and commercial experience (i.e. post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safe and effective use of the devices.A dear physician letter dated December 27, 2016, will be sent to provide important information related to the AFX Endovascular AAA System (AFX System). Endologix is informing the customers of a series of updates to the Instructions for Use (IFU) and modifications to the product. Customers with any questions are instructed to contact their local Endologix representative. |  |  |  | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | April 28, 2017 | 122 |  | Not_Computer
Z-1035-2017 | 76075 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type IIIA |  |  | December 27, 2016 | 2016 | Endologix | Endologix initiated a correction by updating reports for the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System). | Other | Endologix sent an Dear Physician letter dated June 2, 2015, were sent to customers to inform them that Endologix has enclosed the 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The update reports the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System). The customer letter includes update on the clinical programs and commercial experience (ie. post-market surveillance). On January 2016, Dear Physician letters were sent to customers to inform them that Endologix has enclosed the 2015 Clinical Update covering the Endologix Endovascular Systems for AAA Repair. The update reports the cumulative clinical data and commercial experience (ie. post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safe and effective use of the devices. The customer letter includes updates on the Clinical Programs and Commercial experience.A Dear Physician letter, "Important Safety Update AFX Endovascular AAA System: Type III Endoleaks", was sent to provide important information related to the AFX Endovascular AAA System (AFX System). The letter informs the customers that Endologix has an active post-market surveillance program that has been monitoring and evaluating the performance of the AFX System since its introduction to the market in 2011. In January of 2013, Endologix conducted an investigation into reports of Type IIIa endoleaks (separation of bifurcated and extension stent grafts at the point of overlap), which was followed by an investigation into Type IIIb endoleaks (disruption of the stent graft material) in September of 2013. For further questions please call (949) 595-7200. |  | 40967 | 40,967 | Worldwide Distribution - US Distribution and to the countries of : Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | April 28, 2017 | 122 | 40967 | Not_Computer
Z-1044-2017 | 76198 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System. |  |  | January 10, 2017 | 2017 | Elekta, Inc. | Incorrect dose after editing beam number an wedge angle. | Other | Elekta sent Important Field Safety Notice 382-01-MON-009 to all affected customers on 1/10/2017. The notice informs users of the specific product and version numbers by the issue, and any work arounds that can be used to avoid the issue. The notice informs customers that Elekta Regional Support Staff will contact them directly. The notice also included an acknowledgment form to be returned. |  | 1999 | 1999 | AZ, AR, CA, CO, CL, GA. IL, IN, LA, MA, MI, MN, MO, NV, NH, NJ, NY, NC, OH, OR, PA, SD, TX, UT, VA, WA, WI, Puerto Rico, Algeria, Antiqua &amp; Barbuda, Argentina, Australia, Austria, Bangladesh, Bahrain, Belarus, Botswana, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Laos, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Napal, Namibia, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Peru, Romania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Venezuela, Vietnam, Zimbabwe | Open, Classified |  |  | 1999 | Not_Computer
Z-1041-2017 | 75966 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew, RENASYS GO Negative Pressure Wound Therapy, Rx only. |  |  | November 11, 2016 | 2016 | Smith &amp; Nephew, Inc. | Lack of 510k clearance for design modification. | No Marketing Application | On letter sent to all their consignees, Smith &amp; Nephew asked to the consignees to removed all existing Renasys Go NPWT Pumps and returned following the instructions on the letter and also removed the affected power supply units identified on the recall letter. Each consignee needs to complete the Acknowledge |  | 15960 | 15,960 units in the US | Nationwide | Open, Classified |  |  | 15960 | Not_Computer
Z-1043-2017 | 76048 | No | DEH | Immunoglobulin (light chain specific) immunological test system. | Immunology | Immunology | 510(k) | Optilite Freelite Lambda Free Kit |  |  | July 05, 2016 | 2016 | The Binding Site Group, Ltd. | Batch numbers listed are reporting Lambda Free Light Chain results &gt;114mg/L that are on average 33% lower than other Optilite Lambda Freelite batches. Therefore users moving to this batch will see results &gt;114mg/L giving lower results and conversely when switching to the next batch results will be seen to shift upwards. | Under Investigation by firm | On 07/05/16 notification letters were sent to customers. The Binding Site Group requests the following:- Immediately Quarantine &amp; Dispose of any LK018.OPT lot 405566 kits within stock;- Remove any affected lots from the Optilite instrument;- Patients with Lambda Freelite results &gt;114mg/L should be reviewed in consultation with clinicians, alongside other clinical findings and retested where applicable;- Notify local Binding Site representative to request replacement kits;- Return completed E-Back Form to local Binding Site Representative;The firm recommended that the notice be circulated to those within the organization who need to be aware. |  | 809 | 809 kits | U.S. distribution to the following; AR, PA, GA, NC, KY, VA, CA.Foreign distribution to the following; Uk, Germany, France, Italy, Spain, Czech Republic, Portugal, Belgium. | Terminated | July 03, 2017 | 363 | 809 | Not_Computer
Z-1045-2017 | 76187 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Ambra PACS |  |  | December 16, 2016 | 2016 | DICOM GRID INC | A software error caused the window/level to become the same in one series regardless if the image had different levels; image results have a washed-out grey appearance. | Software change control | On 12/16/16 a fix is the software issue was conducted via remote cloud system. On 01/11/17 customers were sent Release Notes that addressed the software correction. The issue was resolved and required no further action. |  | 209 | 209 | U.S. distribution to the following; TX, FL, NE, CA, WI, AZ, NH, MO, MA, CO, PA, NV, PR, IL, OH, NY, MD, NJ, MN, MI, CT, VA, GA, UT, KS, WA, MI, HI.Foreign distribution to the following; Canada, Spain, Israel, Singapore, Kosovo, Austria, Brazil, Malaysia, Australia. | Terminated | April 07, 2017 | 112 | 209 | Computer
Z-1040-2017 | 76054 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 2500 Family Catalog No. ODL25120 AutoMate 1200, ODL25125, AutoMate 1250, ODL25250 AutoMate 2500, ODL25255 Automate 2550. |  |  | December 21, 2016 | 2016 | Beckman Coulter Inc. | Beckman Coulter initiated a design change to update the Automate PC image to accommodate the operating system change to Windows 10. | Under Investigation by firm | The firm, Beckman Coulter, initiated a cite visit to correct the 4 systems installed. The firm will perform the correction by Field Service and tracked via Service Modification.If you have any questions, contact the VP, Global Quality Systems &amp; Compliance at 714-961-4933 or email: RegulatoryAgencyInquiries@beckman.com. |  | 4 | 4 units | US Distribution to MD only. | Terminated | May 17, 2017 | 147 | 4 | Not_Computer
Z-0970-2017 | 75875 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Halo One 5F Introducer Sheath 10 cm 5 pack |  |  | December 02, 2016 | 2016 | BARD PERIPHERAL VASCULAR, INC | Complaints of sheath separation, kinking, and/or tip damage during use. | Device Design | Bard sent an Urgent Medical Device Recall Notification letter dated December 2, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to find attached instructions detailing the steps they need to take regarding the product. Customers with questions were instructed to contact their Bard representative. Customers were provided the following instructions:1. Records show that your facility has purchased product codes affected by this voluntary recall. Do not use or further distribute any affected product.2. Check all inventory locations within your institution for affected product code / lot number combination listed in the recall notice. If have further distributed any of the product code / lot numbers, immediately contact that location, advise them of the recall, forward instructions and have them return the affected product to BPV.3. Remove any identified product from shelves.4. If have used the affected product, complete and return the attached Recall and Effectiveness Check Form indicating no product will be returned. Once the product affected by this recall has been removed from inventory:5. Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product that intend to return. It is extremely important that this information is received.6. Please call the firm's Recall Coordinator at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST). Once all information has been verified, the Recall Coordinator will issue you either a Return Authorization (XC) Number or Consignment Recall Number (XH) to facilitate the expedient return of the product. BPV will issue a credit for returned product. Email the completed Recall and Effectiveness Check Form to HaloOneRecall@crbard.com or fax it to BPV at 1-800-994-6772. If cannot email or FAX the form, please call the BPV recall coordinator at 1-8 |  | 101 | 101 | Nationwide Distribution to NY, MO, NY, NY, FL, KS, MI, TX, NV, OH, ME, WA, LA, AZ, UT, NH.No foreign distribution. | Terminated | April 05, 2017 | 124 | 101 | Not_Computer
Z-1051-2017 | 76076 | No | PAM | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Great Basin Staph lD/R Blood Culture Panel |  |  | December 12, 2016 | 2016 | Great Basin Scientific, Inc. | The Great Basin Staph ID/R Panel is being recalled because the S. aureus probe could potentially result in a false positive. | Material/Component Contamination | The firm, Great Basin Scientific, notified customers by phone beginning on December 12, 2016 and by email on December 14 and 15, 2016. The notifications described the product, problem and actions to be taken. The customers were instructed to cease using and immediately discard all affected cartridges, the firm will replace those cartridges as soon as possible; and respond to the email or call Direct: 385.219.4540 or Support Team 888.333.9763 or email: customerssupport@gbscience.com with any questions or concerns. |  | 32 | 32 Kits (320 individual disposable cartridges) | US Nationwide Distribution. | Terminated | March 31, 2017 | 109 | 32 | Not_Computer
Z-1048-2017 | 76256 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System |  |  | January 04, 2017 | 2017 | Endologix | Endologix updated information on the rates of Type III endoleaks and suggestions for patient surveillanceand treatment. | Device Design | Endologix sent a Dear Physician letter on 12/30/16 to inform them information related to the AFX Endovascular AAA System (AFX System). The letter provides updated information on the rates of Type III endoleaks and suggestions for patient surveillance and treatment. Customers were informed of the voluntary recall of the small remaining quantity of original AFX with Strata graft material, and the larger diameter sizes of AFX2. Customers with any questions are instructed to contact their Endologix representative. Customers with questions were instructed to call their Endologix representative.On 1/17/17, Endologix issued a press release to inform consumers that firm is resuming shipments of all sizes of AFX2 Endovascular AAA systems and appointed Laura Nagel as Vice President, Global Quality. Consumers are informed that the large diameter sizes of AFX2 were placed on temporary hold on 12/27/16 to investigate a manufacture issue. |  | 34 | 34 units | Worldwide Distribution - US (nationwide)and Internationally to Austria, Belgium, France, Germany, Greece, Italy, Ireland, Monaco, Poland Portugal, Romania, Spain, Sweden, Turkey, Israel, Argentina, Florida, Mexico, Chile, Brazil, Japan, Singapore, Malaysia, and Thailand. | Terminated | April 28, 2017 | 114 | 866 | Not_Computer
Z-1050-2017 | 76049 | No | DEH | Immunoglobulin (light chain specific) immunological test system. | Immunology | Immunology | 510(k) | Human Lambda Free SPAPlus Kit |  |  | November 28, 2012 | 2012 | The Binding Site Group, Ltd. | Calibration curve activity has increased over time in the kit lots listed. | Unknown/Undetermined by firm | Customer notification letters were sent 11/28/12. Customers were asked to discontinue using the listed kit lots. Discard any unused or partially-used kits of these lots. Let The Binding Site know how many replacement kits are required. |  | 1002 | 1,002 | U.S. distribution to the following; CA, TX, HI, GA, PA, UT, NY, OK, ID, MA, AZ.Foreign distribution to the following; Australia, Bulgaria, Denmark, Hong Kong, Ireland, Israel, Japan, New Zealand, Taiwan, Belgium, France, Italy, Spain. | Open, Classified |  |  | 1002 | Not_Computer
Z-1053-2017 | 76050 | No | DEH | Immunoglobulin (light chain specific) immunological test system. | Immunology | Immunology | 510(k) | Freelite Human Lambda Free Kit for use on SPAPLUS |  |  | October 28, 2016 | 2016 | The Binding Site Group, Ltd. | A change in the calibration curve causing an increase in false prozone flags. | Under Investigation by firm | Binding Site sent an Important Notice dated October 28, 2016, to all affected customers. . Binding Site requests the following:- Immediately quarantine and dispose of any LK018.S lot 400012 / LK018.10S lot 400257 kits held in your organisation;- Remove any affected lots from the SPAPLUS instrument;- Contact Binding Site representative to request replacement kits;- Return completed E-back form to your local Binding Site representative.For questions regarding this recall call 858-453-9177. |  | 3088 | 3,088 | Worldwide Distribution - U.S. (nationwide) and Internationally to the following; Canada, Australia. | Terminated | July 27, 2017 | 272 | 3088 | Not_Computer
Z-1052-2017 | 76213 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Carestream Touch Prime, Catalog # 1738830, and Touch Prime XE, Catalog # 1738822, Ultrasound System |  |  | November 21, 2016 | 2016 | Carestream Health Inc | Software error; Carestream Health Inc, received a complaint stating that when a user accidentally obtains a measurement value of 0 and corrects the value in the report, the resulting measurement unit is not displayed, i.e., centimeters or millimeters. As such, the user expects that the measurement is taken calculated in centimeters, consistent with other values in the report. In actuality, the measurement is taken in millimeters. When this updated measurement is used in an average calculation, the result appears incorrect as two measurements are interpreted as centimeters while the user corrected value is interpreted as millimeters. If the user selects a Calc Result display as Min or Max, the values are also interpreted as millimeters when centimeters were expected. | Software design | Carestream sent an URGENT: MEDICAL DEVICE CORRECTION Letter dated November 17, 2016, with a Consignee Notification Acknowledgement (HRA # MA-2016-019) to all affected US consignees via UPS on November 21, 2016. The letter identified the product, the problem, and the action to be taken by the consigneeAction to be Taken:When reviewing the worksheet, if a zero value is observed and that calculation field needs to be changed, the calculation needs to be deleted from the worksheet. Once deleted, the user needs to acquire a new calculation from a clinical image.Carestream is currently validating a fix for this problem and your local service representative will be contacting you within the next few weeks to arrange to upgrade your software.If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910, available 7 days per week on a 24 hour basis. Outside of the U.S., please call your local Service support number.If you have distributed the device outside your facility, please alert your customer(s) of this field correction and contact Carestream Customer Care Center as listed above.. |  | 30 | US: 14 units; Foreign: 16 units | Worldwide Distribution - US including SC, GA, TX, OH, IA, NY and foreign distribution to Switzerland and Malaysia | Open, Classified |  |  | 30 | Computer
Z-1047-2017 | 76112 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type IIIA |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX EndovascularAAA System (AFX System). | Nonconforming Material/Component | On 6/2/15, a Dear Physician letter was sent to inform their customers of the 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair.The letter provided update reports of the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System).On January 2016, a Dear Physician letter was sent to customers to inform them of the 2015 Clinical Update covering the Endologix Endovascular Systems for AAA Repair. The letter provided update reports of the cumulative clinical data and commercial experience (i.e., post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safeand effective use of these devices.On 12/27/16, a Dear Physician letter will be sent to customers to inform them of the important information related to the AFX Endovascular AAA System (AFX System). The letter provides a series of updates to the Instructions for Use (IFU) and modifications to the product were implemented, including introduction of a graft material processing improvement known as Duraply.Customer are informed of the on-going Endologix investigation that the changes may help prevent the occurrence of Type III endoleaks reported with the AFX device. |  |  |  | US and worldwide: Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Terminated | April 28, 2017 | 122 |  | Not_Computer
Z-1056-2017 | 75998 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | APD Set with Cassette 3-Prong Peritoneal Dialysis Set |  |  | December 16, 2016 | 2016 | Baxter Healthcare Corporation | Customer complaints received for the presence of leaks | Under Investigation by firm | An Urgent Product Recall communication will be sent to affected customers via U.S.P.S., first class mail. |  | 8640 | 8640 eaches | US only | Open, Classified |  |  | 8640 | Not_Computer
Z-1046-2017 | 75957 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | If a reader selects Multi-study review prior to the original study completing its loading process and then immediately returns to image review of the original study, it is possible that not all current images will be present for review, which could result in incorrect treatment/diagnosis. | Software design | The recalling firm issued a letter dated 3/28/2016 via email or via certified mail if they did not have an email address for the customer. A second letter dated 7/11/2016 was issued to nonresponders. |  | 188 | 188 sites potentially have the effected versions | US Nationwide Distribution | Open, Classified |  |  | 188 | Not_Computer
Z-1039-2017 | 76113 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type: IIIB |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX EndovascularAAA System (AFX System). | Device Design | Endologix sent a Dear Phycisicans Letter dated June 2, 2015, to all affected customers to inform them of the enclosed 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The letter informs the customers of the update reports the cumulative clinical data and information on the Powerlink System and the AFX Endovascular AAA System (AFX System).On January 2016, a Dear Physician letter was sent to inform the customers of the 2015 Clinical Update covering the Endologix Endovascular Systemsfor AAA Repair. Customers were informed the update reports the cumulative clinical data and commercial experience (i.e., post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safeand effective use of these devices.On 12/27/16, a Dear Physician letter was sent to inform customers of the important information related to the AFX Endovascular AAA System (AFX System). The letter informs the customers of a series of updates to the Instructions for Use (IFU) and modifications to the product were implemented,including introduction of a graft material processing improvement known as Duraply. The letter informs the customers that at the time of the submission of these modifications to FDA, they were not identified as measures intended to address Type III endoleaks. The on-going Endologix investigation has determined that these changes may help prevent the occurrence of Type III endoleaks reported with the AFX device. Customers with any questions areinstructed to contact their Endologix representative. |  | 4143 | 4,143 devices | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | April 28, 2017 | 122 | 4143 | Not_Computer
Z-1055-2017 | 76043 | No | JZQ | Radial immunodiffusion plate. | Immunology | Immunology | 510(K) Exempt | Digital RID Plate Reader and Software |  |  | October 09, 2012 | 2012 | The Binding Site Group, Ltd. | If a control ring is marked after reading, the software will not flag results that are out of the specified QC range. | Software Design Change | On 10/2012 notification letter was sent to customers requesting they take the following actions:1. If have not already installed software version Vl.0.3 then please continue to work with version V1.0.2 until the above issues are resolved.2. If have already installed software version V1.0.3, when processing results always mark the controlprior to reading the ring.3. If are uncertain about the validity of any assays you have performed using software version Vl.0.3 please review the controlresults obtained and compare them to the specified control range for the RID assay and;"If the control results are acceptable then your assay is valid and results will be unaffected."If the control results are not acceptable your assay should be considered invalid and your patient results should be reviewed and repeated as appropriate.On 02/2013 a second notification was sent in which customers were informed the new software was released.Customers were asked to take the following actions;1. Upon receipt install the new software and uninstall the previou version2. Install following the onscreen instructions. The installation process has not changed from previous versions. The instruction guide received with the RID Plate reader is still the current version.3. After receipt of the new software disc complete and return the Eback form to The Binding Site.For further assistance contact Technical Services at The Binding Site via the email address: technicalservices@bindingsite.co.uk. |  |  |  | U.S. distribution to the following; MA, GA, NY, IA, CA, IL, UT, NJ, NC, MD, MN.Foreign distribution to the following; UK, Germany, Spain, France. | Terminated | April 04, 2017 | 1638 |  | Computer
Z-1054-2017 | 76106 | No | MIH |  |  | Cardiovascular | PMA | AFX Endovascular AAA System, Endoleak Type IIIB |  |  | December 27, 2016 | 2016 | Endologix | Endologix updated the Instructions for Use (IFU) and implemented modifications to the AFX EndovascularAAA System (AFX System). | Under Investigation by firm | A Dear Physician letter dated 6/2/15 was sent to customers to inform them of the 2014 Clinical Update on the Endologix Endovascular Systems for AAA Repair. The letter provided update on the clinical programs and commercial experience (i.e., post-market surveillance).On January 2016, a Dear Physician letter was sent to customers to inform them of the 2015 Clinical Update covering the Endologix Endovascular Systems for AAA Repair. The letter provided update reports for the cumulative clinical data and commercial experience (i.e., post-market surveillance) of the Powerlink System and the AFX Endovascular AAA system and contains important information and recommendations regarding the continued safeand effective use of the devices.On 12/27/2016 a Dear Physician letter provides important information related to the AFX Endovascular AAA System (AFX System). The letter informs the customers that a series of updates to the Instructions for Use (IFU) and modifications to the product were implemented, including introduction of a graft material processing improvement known as Duraply. Customers with any questions are instructed to contact their Endologix representative. |  | 17139 | 17,139 units | US, and worldwide:Austria, Belgium, Czech Republic, France, Germany, Greece, Ireland, Israel, Italy, Latvia, Luxembourg, Monaco, Netherlands, Poland, Portugal, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Terminated | April 28, 2017 | 122 | 17139 | Not_Computer
Z-1064-2017 | 73459 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Apollo: Models: 9784000131, 9784000231, 9784020131, 9784020231, 9784100231, 9784120131, 9784120231, 9784130131 |  |  | February 19, 2016 | 2016 | Villa Radiology Systems LLC | The analysis showed that the incident was caused by a chain of four faults:1.) The injured person overcame a physical barrier (the tabletop support arm. height ca. 60 em) to access a prohibited area not intended as a working position.2) The directions for use (device labeling and instructions for use) did not contain an explicit exclusion of the prohibited area.3.) The movement was activated inadvertently and continuously. and it was not released (the movement requires continuous activation).4.) Delayed activation of the emergency stop. | Device Design | Consignees were notified of the recall via letter starting on 2/19/16. The letter explains the issue and actions to be taken by the consignee and manufacturer. |  | 66 | 66 distributed in US | ALAZCACOCTDCFLIAILINLAMAMDMIMNMOMSNJNYOHPAPuerto RicoTNTXUTWI | Open, Classified |  |  | 66 | Not_Computer
Z-1057-2017 | 75951 | Yes | PAW | Nonabsorbable expanded polytetrafluoroethylene surgical suture. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Chord-X Chordal Sizer is designed to assist in determining the proper length of chordal replacement prostheses to repair the mitral valve and help the surgeon select a preconfigured prosthesis, or configure the correct size using straight suture material, during mitral valve repair. |  |  | December 13, 2016 | 2016 | On-X Life Technologies, Inc. | Four issues were uncovered by the recalling firm's CAPA: 1) Burrs were found to be present on Left Housing, Right Housing and Barrel components, 2) Loose plastic particulate was found within sealed primary packages, 3) Loose hair was discovered within four sealed primary packages and 4) Assembly error, wherein Left and Right Housing components were acceptably press-fit together, was discovered within one sealed primary package. One of the Left Housing press-fit pins was not pressed into the associated hole of the Right Housing, resulting in a small gap in the housing assembly. | Component design/selection | A letter was sent to the implanting surgerons notifyijng them of the recall and attemmpting retrieval. |  | 45 | 45 units | CA, WA, NE, DE, OH, ME, TX | Terminated | April 13, 2017 | 121 | 45 | Not_Computer
Z-0922-2017 | 75845 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD20 722012 is intended for: Cardiovascular and vascular X-ray imaging applications, including diagnostic, interventional procedures (such as PTCA, stent placement and atherectomies), pacemaker implantations and Electro Physiology(EP). |  |  | December 01, 2016 | 2016 | Philips Electronics North America Corporation | Due to a production error, screws were not tightened to specified torque. These screws may come loose from the C-arm shaft causing the C-arm to rotate uncontrolled. | Process control | Customers were notified of the recall via letter sent on 12/1/16. The letter explains the issue and the action that Philips plans to take. |  | 1 | 1 | US: WAOUS: China, Spain, Korea, Republic of, Germany | Terminated | August 10, 2017 | 252 | 6 | Not_Computer
Z-1058-2017 | 76144 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT |  |  | December 23, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | During a retrospective review of production documentation it was observed that the M6 Rotor Counterweight Bolt may have been improperly torqued. | Employee error | Philips sent an Customer Information letter.A document showed an incorrect torque wrench used during the assembly process which could potentially apply 4 Newton-meters less torque than specified to the system rotor counterweight bolts. Out of an abundance of caution, Philips will have a Field Service Engineer ensure the bolts on your system are correctly torqued.If you need any further information or support concerning this issue, please contact your local Philips representative. For further information please call (440) 483-5777. |  | 1 | 1-unit | Product was not distributed in the US.Unit was distributed in Australia. | Open, Classified |  |  | 1 | Not_Computer
Z-1061-2017 | 76080 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Velara Generator with an old Firmware version used with Philips Allura Xper Systems (Allura 15 &amp; 12 (monoplane), Allura FD10C, Allura FD10F, Allura FD 10 F/C, Allura FD20, Allura XPER FD10/10, Allura XPER FD20 Biplane) |  |  | December 09, 2016 | 2016 | Philips Electronics North America Corporation | System may lock up. | Device Design | Consignees were notified of the recall by letter, sent via certified mail on 12/9/16. Philips will upgrade the firmware of the Velara Generator, allowing the Velara Generator to recover in case of intermittent shortcuts. This service will be provided free of charge for all affected systems and will start effective January 2017. A Philips service representative will contact customers with affected devices to arrange for the service. |  | 59 | 59 devices (this number represents all products listed in recall) | US: AL, CA, DC, IL,GA, KS, MD, MO, NC, NY, OH, PA, RI, SC, TX, WAOUS: Brazil, Bulgaria, Costa Rica, France, Germany, Israel, Italy, Japan, Korea, Republic of, Mexico, Russian Federation, Singapore, Spain, Sweden, SwitzerlandSwitzerland, Turkey, United Kingdom | Open, Classified |  |  | 59 | Not_Computer
Z-1060-2017 | 76074 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9900 Elite. MDL Numbers: D148942, D155043 |  |  | December 27, 2016 | 2016 | GE OEC Medical Systems, Inc | GE Healthcare Surgery announces a voluntary field action for the OEC 9800 and OEC 9900 Elite due to additional supplementation to the user documentation to include areas for enhanced planned maintenance and component replacement frequencies. | Packaging change control | The customer communication is expected to begin distribution in January 2017. Consignees will be notified of the availability of the Operator Manual supplement via a letter. No product is being returned. |  | 26256 | 26,256 total | Worldwide including govt/VA/military and foreign consignees. | Open, Classified |  |  | 26256 | Not_Computer
Z-1075-2017 | 76089 | No | EQL | Laryngostroboscope. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Rigid Laryngostroboscopes |  |  | December 12, 2016 | 2016 | Pentax of America Inc | Pentax is initiating a Field Correction to revise the Instructions for Use of the Rigid Laryngostroboscopes Model 9106 and 9108. | Labeling design | The firm, Pentax Medical, sent an "U.S. URGENT FIELD CORRECTION" letter and a Return Response form dated December 12, 2016 to their affected customers via USPS Certified Mail. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all users read and carefully follow the revised IFU in order to properly operate and reprocess the affected product; contact your PENTAX Medical sales representative to arrange for training regarding the new procedures, and complete and return the FIELD CORRECTION RESPONSE FORM via fax to: PENTAX Recall Coordinator at 201-799-4063 (alternate 201-391-4189) or email a pdf copy to customeradvisories@pentaxmedical.com even if you no longer possess the product.If you have any questions regarding this action, please feel free to contact us at: 800-431-5880 (8:30 AM-5:00 PM, Monday-Friday, EST) or email: customeradvisories@pentaxmedical.com. |  | 2056 | 2,056 (1,545 (US) and 511 (OUS)) | Worldwide Distribution-US (nationwide) and countries of: Albania, Australia, Belgium, Brazil, Bulgaria, Canada, Chile, Croatia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Oman, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 2056 | Not_Computer
Z-1074-2017 | 76132 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Unna-Z, Unna Boot with Zinc and Calamine, Compression Wrap, 1 Per Box, 3-4in x 10yds |  |  | November 23, 2016 | 2016 | Medline Industries Inc | Medline Industries, Inc is recalling Unna Z Unna Boot Bandage 3 X10 and Unna Z Unna Boot Bandage 4 X10 due to contamination of the products with common mold (Cladosporium sphaerospermum/cladosporioides/herbarum/phaenocomae/halotolerans and Penicillium corylophilum).. | Environmental control | Medline Industries sent an Immediate Action letter dated November 23, 2016, to all affected customers with response forms to their customers. Customers were instructed to destroy the affected Unna Boot Bandages. Affected lots of product remaining at the firm's distribution centers have been placed on hold and quarantined pending destruction. All affected product will be destroyed. For further questions, please call (847) 643-3245. |  | 117059 | 117,059 units | Worldwide Distribution - USA (nationwide) and to the countries of : Chile, Peru, El Salvador and Taiwan. | Open, Classified |  |  | 117059 | Not_Computer
Z-1077-2017 | 75709 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 123 POC system |  |  | September 03, 2015 | 2015 | Roche Diagnostics Operations, Inc. | under specific settings, an issue may occur during simultaneous Sensor Cartridge and Fluid Pack change on the cobas b 123 &lt;2&gt; POC system and cobas b 123 &lt;4&gt; POC system.The issue occurs when the software function [AutoQC as follow-up] is configured to run all three levels of AutoQC only after a Fluid Pack change, but not after a Sensor Cartridge change. When both are changed simultaneously, starting with the Sensor Cartridge and followed by the Fluid Pack, the analyzer carries out only the follow-up actions associated with the Sensor Cartridge change after completing the change workflow. As a result, no follow-up AutoQC is performed and the three expected AutoQC measurements for the Fluid Pack change are not carried out.Without running quality control, there is a remote possibility that system issues would not be detected and wrong results would not be excluded on all parameters: pH, PO2, PCO2, Na+, K+, Ca++, Cl-, Glu, Lac, Hct, SO2, O2Hb, COHb, MetHb, HHb, and Bili. | Software design | Roche issued an Analyzer Bulletin on September 3, 2015, to all affected Consignees. The recall notifications included a description of the reason for the recall, affected product, and consignee responsibilities; the notification did not include instructions for responding to the notification. Actions Required" Follow the workaround outlined in this Analyzer Bulletin to avoid the described issue." File this Analyzer Bulletin for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center at 1-800-526-2272 if you have questions about the information contained in this Analyzer Bulletin. |  | 30 | 30 | Nationwide Distribution including CA, LA, NE, PR, SC, TX, VA, and WV | Terminated | June 05, 2017 | 641 | 30 | Computer
Z-1067-2017 | 76041 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Human Subclass Kit for use on the Beckman IMMAGE"/IMMAGE"800 Analyser |  |  | September 11, 2014 | 2014 | The Binding Site Group, Ltd. | Customer complaints identified that samples containing extremely high IgG4 levels exhibit antigen excess and will yield spuriously low IgG4 results. | Nonconforming Material/Component | Binding Site sent an Important Information letter dated September 30, 2014, Customers were instructed to retain the notification for records and additional instructions with the firm stating they were updating the instructions for use. Customers with questions were advised to contact their Binding Site representative or technical.service@bindingsite.co.uk. |  |  |  | Worldwide Distribution - US including n FL, OH, TN, MN, NW, MI, CA. and Internationally to Canada, France, UK, Spain, Germany, Czech Republic, Belgium, Italy, Australia, Brazil, Chile, China, Costa Rica, Kuwait, New Zealand, Oman, Pakistan, Portugal, Sweded, Taiwan, and Turkey. | Terminated | May 03, 2017 | 965 |  | Not_Computer
Z-1076-2017 | 75878 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Mallinckrodt Pharmaceuticals INOmax DSIR Plus delivery system. Model 10007. Delivers INOMAX (nitric oxide for inhalation). |  |  | November 21, 2016 | 2016 | INO Therapeutics (dba Ikaria) | Mallinckrodt discovered that some devices may have been manufactured with an out-of-tolerance manometer which can result in in an inaccurate proportional valve calibration table in an affected device. | Equipment maintenance | Recall letter was sent to customers in a e-mail beginning November 21, 2016 with a had copy sent via FedEx overnight with signature receipt beginning December 5, 2016. Devices will be picked up and replaced by Mallinckrodt. Questions can be directed to 1-877-KNOW-INO (1-877-566-9466). |  | 45 | 45 | US Distribution to the states of : CA, CO, DC, FL, GA, HI, IA, IL, IN, MI, MN, NC, NJ, NM, NY, OH, PA, and TX | Open, Classified |  |  | 45 | Not_Computer
Z-1078-2017 | 75818 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sekisui Diagnostics Ammonia L3KÂż Assay; Catalog Number: 293-80-91 and 293-10 |  |  | November 29, 2016 | 2016 | Sekisui Diagnostics P.E.I. Inc. | Certain lots are showing an atypical decrease in optical density (OD) over time, which may result in an impact to performance at the high end of the linear range. | Material/Component Contamination | Sekisui Diagnostics sent an Important Product Correction Notification letter dated November 29, 2016, to all affected customers. Customers were asked to take the following actions:Customer Instructions: Please stop using and discard any of the Ammonia L3KÂż reagent lots listed above in your inventory; Contact your distributor to discuss replacement product. Distributor Instructions: Please send the "IMPORTANT PRODUCT CORRECTION Customer Notification" provided to all customers that received Sekisui Diagnostics Ammonia LJKÂż from the listed lot numbers; Complete, sign and return the attached CONFIRMATION OF NOTIFICATION form by FAX to the number provided. This will: Indicate receipt of this Field Correction, and confirmation of notification of customers that received impacted customers; Act as confirmation of destruction of affected product in inventory; Allow for issuance of replacement product.Customers with questions were instructed to contact Technical Services at 1-800-565-0265 or email PEIDiagnosticTechnical@sekisui-dx.com. |  | 1339 | 774 Liters bulk and 565 kits | Worldwide Distribution - US including DE, KS, FL, TX, CA, MI, TN, MA, ID, NJ, RI, IL and Internationally to Austria, Mexico, Canada, India, South Korea, China, Spain, South Korea, and Sri Lanka | Terminated | July 06, 2017 | 219 | 1339 | Not_Computer
Z-1091-2017 | 76214 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. Merge Hemo monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | There is a potential connection issue when powering up the Merge Hemo Record Station and the Hemo Monitor does not communicate with the Client PC. | Software design | MERGE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to reply using the enclosed form and ensure all users of the product are provided with the notification. Customers with questions were instructed to send an email to recall@merge.com.. |  | 242 | 242 sites potentially have the affected versions | Nationwide Distribution | Open, Classified |  |  | 242 | Not_Computer
Z-1092-2017 | 75711 | No | NQD | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | Tina-Quant (Latex) HS Test System |  |  | February 26, 2016 | 2016 | Roche Diagnostics Operations, Inc. | Tina-quant Cardiac C-reactive Protein (Latex) High Sensitive (CRPHS) lot number 604450 (catalog number 11972855216) shows an under-recovery of up to -25% at concentrations below 5 mg/L on the MODULAR ANALYTICS P module.this issue can lead to erroneous low CRPHS results. CRPHS is used as a marker for the prediction of cardiovascular risk. An erroneous low result could lead to an incorrect cardiovascular risk assessment and subsequent delay of further examination. | Under Investigation by firm | Roche Diagnostics sent a Reagent Bulletin dated February 26, 2016, to all affected consignees. The recall notifications included a description of the reason for the recall, affected product, request for replacement, and consignee responsibilities; the notification did not include instructions for responding to the recall notification. Actions Required" Do not use CRPHS lot number 604450. If you have any affected product in your inventory, discard it per local guidelines." Complete the attached replacement request form to receive replacement product." File this Reagent Bulletin for future reference.QuestionsPlease contact the Roche Support Network Customer Support Center at 1-800-428-2336 if you have questions about the information contained in this Reagent Bulletin. |  | 3204 | 3204 | Nationwide Distribution including AL, AZ, CA, CT, FL, GA, KS, MA, MO, NJ, NY, PA, and WA. | Terminated | June 05, 2017 | 465 | 3204 | Not_Computer
Z-1090-2017 | 76078 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas p 612 pre-analytical system; |  |  | December 27, 2016 | 2016 | Roche Diagnostics Corporation |  | Software design | On 12/27/2016, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected Consignees via courier fax or email. The recall notifications included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The cobas p 612 pre-analytical system (63x) software is being revised to correct the aspiration process. We anticipate this new version (2.0.4) to be available during Quarter 1 2017.For now, discontinue using disposable non-filter tips for Aliquotting on your cobas p 612 pre-analytical system (63x). You will be contacted by a Roche Field Engineering Specialist (FES) about modifying your system to allow for the use of disposable filter tips. Filter tips will be provided at the time of service.Actions Required" Immediately discontinue using disposable non-filter tips for Aliquotting on your cobas p 612 pre-analytical system (63x)." Complete the enclosed faxback form and fax it to number 1-877-906-8982 or email it to Indianapolis.cd_tsc@roche.com." Provide copies of this UMDC to other clinicians who may need to be aware of this software issue." If your facility has distributed the affected product to another site, please ensure this UMDC is provided to that site." File this UMDC for future reference.Please contact the Roche Support Network Customer Support Center at 1-800-428-2336, 24 hours a day/7 days a week, if you have questions about the information contained in this UMDC. |  | 5 | 5 | TX and WA | Terminated | August 04, 2017 | 220 | 5 | Computer
Z-1079-2017 | 75599 | chest drainage kit | PAD | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | Enforcement Discretion | Atrium Medical Ocean Drains as follows: |  |  | November 29, 2016 | 2016 | Atrium Medical Corporation | Outer Packaging is not sterile | Package design/selection | Atrium Medical issued recall letter on 11/29/16 via two-day express mail delivery advising users: of the problem, updated instructions for use; the new Warning Label applied to the packaging; and requesting to post the warning label near all Atrium chest drain product inventory.A Field Notification Reply Form acknowledging receipt of the Field Notification was requested to be completed and returned.Questions contact Atrium/Maquet/Getinge chest drainage representative or Atrium Customer Service at 1-800-528-7486 (603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 964463 | 964,463 cases (US) 521,975 cases (OUS) in total | Worldwide Distribution - US Nationwide and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, China, China, Columbia, Ecuador, Egypt, Finland, France, Germany, Great Britain, Greece, Guatemala, Hong Kong, Ireland, Italy, Kuwait, Lebanon, Libya, Mexico, Netherlands, New Zealand, Pakistan, Panama, Peru, Poland, Portugal, Puerto Rico, Saudi Arabia, South Korea, Spain, Sweden, Thailand, Tunisia, and Venezuela, | Open, Classified |  |  | 964463 | Not_Computer
Z-1094-2017 | 76168 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Maintenance Kit PH1, Cod. 6997266 |  |  | January 04, 2017 | 2017 | Baxter Healthcare Corporation | Baxter Healthcare Corporation is issuing a voluntary product recall for several lots of Phoenix Dialysis Machine Preventive Maintenance Kits due to incorrect silicone straight connectors components packaged in these kits. During Preventive Maintenance, the two silicone connectors must be installed in the Phoenix mass-balance section of hydraulic circuit only if the Phoenix machine has already been modified with"CONN. IMPROVING RETR. KIT" (Code SP00721). Installation of these connectors in machines modified with SP00721 may result in fluid leaks during dialysis treatment. These Preventive Maintenance Kits include the FILTER sub-group, which contains the silicone straight connectors that are larger than expected. | Mixed-up of materials/components | Baxter sent an Urgent Product Recall on January 4, 2016, to all affected customers via U.S.P.S., first class mail. Customers were asked to: 1) Locate and remove all affected product lots from their facility. 2) Check if the Preventative Maintenance Kits were already installed in their Phoenix machines and if the case or if uncertain, contact their Baxter Service Representative at 1-800-525-2623 (Option 2) to schedule an appointment to remediate the machine.3) Contact Baxter Health care Center for Service to arrange for return and credit.4) Complete the enclosed Baxter Customer Reply Form and return it to Baxter by either fax or scanned email.All non-responding customers will be contacted via telephone.All affected product will be destroyed or discarded.Customers with questions were advised to call 800-437-5176. |  | 211 | 211 kits | Nationwide Distribution | Open, Classified |  |  | 467 | Not_Computer
Z-1093-2017 | 76053 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | Tryptic Soy Agar, Acumedia PN 7100 |  |  | April 22, 2016 | 2016 | Acumedia Manufacturers, Inc. | light precipitate may form in prepared plates of Acumedia TrypticSoy Agar (7100) Lot# 108166.While there is no change in growth performance of this medium, the formationof particulate matter may be misinterpreted as growth in poured plates thereby leading to inaccuratecounts. | Component design/selection | Neogen sent a Recall Notification dated March 8, 2016, to all affected customers on April 22, 2016. via e-mail. The recall notification included a description of the reason for the recall, affected product, and offer for replacement. If any Tryptic Soy Agar from this lot remains in your inventory, and you would like to replace this lot, please contact the customer service team at Neogen Corporation by email at foodsafety@neogen.com or by telephone at (800)-234-5333.We would like to assure you that we are committed to providing products that meet our strict quality standards for your testing needs. Neogen will investigate this isolated incident to determine the root cause and implement measures to ensure that this issue does not happen in the future. |  | 482 | 482 | Worldwide Distribution - US Distribution to the states of : AR, CA, CT, FL, GA, ID, IL, IN, MD, ME, MI, NC, ND, NY, OH, OR, TX, TN, UT, VA, WI. and to the countries of : CANADA, UK, ISRAEL, MEXICO, AUSTRALIA, COLOMBIA, TAIWAN, BRAZIL, CHILE and INDONESIA. | Completed |  |  | 482 | Not_Computer
Z-1099-2017 | 76147 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Accu-Chek Connect Diabetes Management App |  |  | December 20, 2016 | 2016 | Roche Diabetes Care, Inc. | iOS and Android: Under certain conditions the affected app versions may disregard historical bolus data potentially leading to an incorrect bolus insulin recommendation being provided to the user.iOS only: Pairing and using multiple meters with the Accu-Chek Connect app can under rare circumstances cause the bolus advisor to fail to offer a correction bolus recommendation within the eligible time window following a blood glucose measurement (10  15 minutes).Depending on the individual metabolic situation potentially incorrect bolus advice could lead to serious health consequences such as hypoglycemia. Both software issues may also cause the amount of active insulin displayed during the bolus calculation process to be incorrect and should not be used to manually calculate a bolus. | Software design | Customers were notified via Google Play Store or the Apple App Store that an update to the Accu-Chek Connect App was available on December 20, 2016.. The description of the new version available in the app stores indicated that the new version resolved issues with the Bolus Advisor. It also includes direction to the Accu-Chek website. On December 21, 2016, information regarding the issues corrected with the new version was placed on the website. The information on the website was replaced on January 10, 2017, with the patient version of the Urgent Medical Device Correction (UMDC) which was sent on January 10, 2017, to US HCPs who have received prescription pads that would allow them to prescribe the bolus advice activation.If you have additional concerns, please contact the Accu-Chek Customer Service Center at 1- 800-858-8072. All other versions of the app as well as other Accu-Chek products using the bolus advice feature are not affected by this issue. |  | 27243 | 27243 | Worldwide Distribution - Nationwide Distribution The Accu-Chek Connect App for iOS is distributed directly to end users via the Apple App Store. The Accu-Chek Connect App for Android is distributed directly to end users via the Google Play App Store. | Open, Classified |  |  | 27243 | Computer
Z-1098-2017 | 76046 | No | NSX |  |  | General Hospital | Contact ODE | Softbank software |  |  | October 06, 2016 | 2016 | Soft Computer Consultants, Inc. | Software error. Potential for incorrect results | Software design (manufacturing process) | Consignees were sent a recall notification, dated 10/06/2016, providing instructions for software correction. The firm indicates that the issue will be resolved with a mandatory hot fix. Hot fix will be released when its effectiveness is confirmed. Consignees are instructed to acknowledge receipt of the recall notification and grant permission for the recalling firm to load the hot fix when available. The firm can be contacted by calling Technical Support at 1-800-763-8522. Until the hot fix is available, consignees can still operate the software by exiting (Esc-Quit) the Test Verify option after each save changes (F12-Accept). |  | 189 | 189 total | Worldwide Distribution - US Nationwide Distribution in the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX, VA, WA, WI, WV, and the countries of Hawaii and Puerto Rico Canada &amp; Jamaica | Open, Classified |  |  | 189 | Computer
Z-1105-2017 | 76190 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | NORMFlow H-110 Irrigation Fluid Warmer, Continental Europe, Model H-1100, Reorder H-1100-INT-230. |  |  | October 25, 2016 | 2016 | Smiths Medical ASD, Inc. | Smiths Medical became aware that that one (1) NORMOFLOÂż H-1100 Irrigation Fluid Warmers was incorrectly labelled with Spanish language labels instead of English labels. Specifically five (5) labels were incorrect; the Operators Manual, Outer Base Box label, Pole Assembly Box label, and two (2) O-ring Lube labels. | Labeling mix-ups | Consignee were hand delivered a Smiths Medical Urgent Medical device Filed Safety Notice letter dated October 25, 2016. The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. Advised consignees to Provide the device to the Smiths Medical Representative who delivered this notice to you. The Smiths Medical Representative will replace the labeling on the device and packaging, and provide you with the correct English language Operators Manual. Complete the Recall Confirmation Form and return it to Smith Medical via e-mail to FCA.Response@Smiths-medical.com. For questions regarding this notification, please ask the Smiths Medical Representativewho delivered the notification or contact Smiths Medical Customer Service Department at +46 859477250. |  | 1 | 1 | Distribution in the foreign country of Sweden | Terminated | July 25, 2017 | 273 | 1 | Not_Computer
Z-1104-2017 | 76101 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | ProSpore 4 mL Ampoule Biological indicator for steam sterilization, contains Geobacillus stearothermophilus, Catalog Number: PS-6-50, contains 50 Ampoules/Box. |  |  | April 04, 2016 | 2016 | Mesa Laboratories Inc | Product resistance has fallen outside of label claim within expiry. | Process control | MesaLabs sent an Urgent Notification letter dated April 4. 2016 to customers. Follow-up emails were sent to verify receipt. The letter identified the affected product, problem and actions to be taken. For questions contact MesaLabs technical support team via email at bi-support@mesalabs.com or call at 303-987-8000 ext. 10042. |  | 5379200 | 328/ 50-count/ boxes; 16,400 individual units | Product was distributed to: AL, AR, CA, FL, IA, MO, NE, NC, NY, OH, and WA. The product was also distributed internationally, including Mexico and Canada. There was no goverment or military distribution. | Terminated | January 26, 2017 | 297 | 5379200 | Not_Computer
Z-1096-2017 | 76056 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Oasis MRI System - C-Spine Coil |  |  | January 06, 2017 | 2017 | Hitachi Medical Systems America Inc | The customer indicated that the coil balun was hot to the touch when removing the coil from the patient table after an exam. | Component design/selection | Hitachi sent an Urgent : Device Correction Notice dated January 3, 2017, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm plans to visit each site and replace the cable balun and also confirm receipt of the notification letter. The firm will quarantine returned baluns in their warehouse until they are reworked. For further questions regarding this recall please call (330) 425-1313 ext. 3720. |  | 180 | 180 units | USA (nationwide) distribution. The product was distributed to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY.There are three (3) government accounts for this recall: VA Medical Center/Lyster Army Health Clinic, Parma VA Medical Center, Outpatient Clinic.There are no Canada accounts for this recall.There are no Mexico accounts for this recall. | Open, Classified |  |  | 180 | Not_Computer
Z-1106-2017 | 76079 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips Cisco 3850 layer 3 switch, 865339 model WS-C3850-12S-E in use with IntelliVue Information Center iX or Classic Information Center |  |  | December 27, 2016 | 2016 | Philips Electronics North America Corporation | May experience a loss of centralized monitoring. | Under Investigation by firm | Philips sent an Field Safety Notice on December 27, 2016, to customers via certified mail. The notice explains the issue and informs the customers that no actions are required. The product may be used until the replacement cable is made available. If you need any further information or support concerning this issue, please call your local Philips Representative at (800) 722-9377. |  | 107 | 107 | Worldwide Distribution - US Distribution to the states of : AK CA CO FL GA KY NC NE NH NJ NY OH PA SC TN TX VA WA and WV., and to the countries of : Canada, Germany, Japan, New Zealand, Poland and United Kingdom | Open, Classified |  |  | 107 | Not_Computer
Z-1109-2017 | 76159 | No | GMD | Cryptococcus neoformans serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Cryptococcal Antigen Lateral Flow Assay (CrAg LFA), Reference # CR2003 |  |  | December 27, 2016 | 2016 | Immuno-Mycologics, Inc | The affected products have approximately a 91% specificity while the non-affected lots have approximately a 99% specificity. The recalling firm is aware of three customers who reported false positive results. | Material/Component Contamination | Immuno sent an Urgent: Medical Device Recall Letter dated December 27, 2016, to all affected customers. All customers were notified via emails to stop using the product on December 27th and December 28th. A recall notification and response form was attached to the email. For further questions, please call (405)364-1058. |  | 630 | 630 units | Nationwide throughout the US | Open, Classified |  |  | 630 | Not_Computer
Z-1107-2017 | 76165 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Genetic Screening Processor (GSP); PerkinElmer Part Number= 2021-0010 |  |  | January 05, 2017 | 2017 | PerkinElmer Life and Analytical Sciences, Wallac, OY | The firm became aware that the ball bushing axels are defective in the GSP instrument. During exposure to humid conditions, there is a risk that the axle material will corrode. Deterioration of the axle can be observed as deposits of corrosion on the defective axles. Ball bushing axles may corrode and corrosion particle may cause contamination in the wells or the corrosion will cause movement issues and possibly spillage and causing contamination of the plate wells. | Component change control | The recall was initiated on January 5, 2017 via email sent through the PerkinElmer Technical Support group located in Akron Ohio.Actions to be taken by user: 1. As instructed in the analyte specific kit inserts, the use of Kit Controls is required to assure the day-to-day validity of results. Patient results should only be reported if control results for the assay meet the laboratorys established criteria for acceptability. 2. If a movement error occurs customers will be directed to contact their local PerkinElmer field service representative. The movement error may indicate deterioration of the axle and therefore the use of the instrument should be avoided until field service personnel have inspected the status of the ball bushing axles or replaced affected ball bushing axles.Actions to be taken by the PerkinElmer: As a corrective action, PerkinElmer will replace the affected ball bushing axles in all the GSP instruments identified in this FSCA/Recall. The replacement may take place either as part of the initial installation activity before the instrument is used, scheduled periodic maintenance, or during a dedicated service visit.This field correction is considered as mandatory by the manufacturer. Customers will be contacted by PerkinElmer to arrange the field service visit. |  | 27 | 24 instruments and 3 spare disk remover modules | MA, SC, TX, ALOUS: China, Mexico, Peru, Morocco, Italy, Finland, Ecuador, Brazil | Open, Classified |  |  | 27 | Not_Computer
Z-1111-2017 | 75978 | No | KDL | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SPS-1, Static Preservation Solution [1 or 2 liter bags], Sterile; Organ Recovery Systems, One Pierce Place, Itasca, IL 60143 |  |  | December 14, 2016 | 2016 | Organ Recovery Systems, Inc. | Organ preservation fluid recalled due to potential for bacterial contamination. | Material/Component Contamination | Firm initiated a voluntary removal of Lot PBR-0060-392 on December 14, 2016. and Lot PBR-0074-330 December 16, 2016. Notification occurred via e-mail to each recipient of the lots. On December 18, 2016, the firm provided subsequent posting notice and FAQs on company website. Follow up telephone conversations were performed post e-mail notifications. Firm worked with FDA to develop and coordinate a unified message that would be noticed on the websites of FDA, UNOS, and CDC. Firm also independently contacted the Association of Organ Procurement Organizations and UNOS to assist with messaging for their constituents. In the recall notification emails, customers were instructed to return affected product to Organ Recovery Systems. All affected product will be destroyed after completion of the firm's internal investigation.Upon learning of the additional lots of potentially contaminated product (PBR-Âż0074-Âż337 and PBR-Âż0060-Âż386), on 1/13/2017 ORS notified their customers via email and phone and updated the posting on their website. The instructions were only to quarantine the product, however. On January 23, 2017, firm began instructing their customers (via email) to return the additional affected lots to ORS. An update to their website was posted January 24, 2017. instructing affected customers to return the lots of product.For further questions regarding this recall, please call (847) 824-2421. |  | 4806 | 4,806 bags (3000 1 Liter bags, 1806 2-Liters bags) | Nationwide | Open, Classified |  |  | 4806 | Not_Computer
Z-1110-2017 | 76319 | No | NPL | Bone grafting material. | Dental | Dental | 510(k) | RENOVIX Guided Healing Collagen Membrane |  |  | December 02, 2016 | 2016 | Biom'up | Lack of sterility assurance | Nonconforming Material/Component | An Urgent Medical Device Recall letter was mailed to affected customer on 12/02/2016 to inform them of the recall. The reason and actions to be taken were explained in the letter. Customer was asked to examine their inventory and contact Biom'Up to arrange for return of product. |  | 1682 | 1,682 units | Distributed to a sole distributor in NC. | Terminated | May 05, 2017 | 154 | 1682 | Not_Computer
Z-1113-2017 | 76156 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cannulated Drill bit 2.0mm and 2.6 mm; Used in conjunction with the Flower Bone Screw Set. |  |  | December 22, 2016 | 2016 | Flower Orthopedics Corporation | The product is being recalled due to incidence and reports of the product breaking during surgery. | Device Design | Flower Orthopedics mailed a letter to customers on December 22, 2016 making them aware of the issue. Customers were asked to return the affected product and to report if any adverse effects resulted from its use. |  |  |  | Distributed throughout the United States | Open, Classified |  |  |  | Not_Computer
Z-1112-2017 | 76111 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | SwishTapered Implant |  |  | December 12, 2016 | 2016 | Implant Direct Sybron Manufacturing, LLC | The extender product code (9037-03) that is packaged with the SwishTapered (Part Number 933308; Lot Number 73241) Implant was packaged in the incorrect position. This discrepancy may lead to the possibility of extender not seating properly on the implant, which may lead to micro-leakage and a possible infection to the patient. | Packaging process control | Customer notification letters were sent on 12/12/16.Customers were requested to review inventory for the affected product. Complete and return the Acknowledgement and Recall Return Form within 48 hours. If customers are an authorized distributor the firm requests that they identify those customers that may have been shipped the affected product lot and contact those customers to inform them of the issue within 48 hours of receipt of the notification in order to provide the customers with replacement product.If customers have any of the affected product the firm will send a replacement part. Any questions contact Implant Direct Sybron Manufacturing LLC Customer Care at 1-888-649-6425. |  | 17 | 17 | U.S. distribution to the following; NM, TX, CA, OR, OH, ME.Foreign distribution to the following; Netherlands. | Terminated | May 22, 2017 | 161 | 17 | Not_Computer
Z-1115-2017 | 75973 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Total Knee System |  |  | January 04, 2017 | 2017 | Zimmer Biomet, Inc. | packaging of some femoral components was incorrectly labeled as either smaller or larger and/or as the incorrect side (right/left). Risk delay in surgery. | Process change control | On 1/4/2017 URGENT MEDICAL DEVICE RECALL REMOVAL- LOT SPECIFIC notifications were sent to the affected distributors and hospital Risk Managers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Distributor responsibilities include: 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet within 2 weeks (January 20, 2017). a. Request a Return Authorization Number via email to rgarequest@zimmerbiomet.com. Be sure to specify RECALL as the RGA type. b. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com. c. Include a hardcopy of Attachment 1 with your shipment for immediate processing. d. Mark the outside of the returns box(es) clearly with RECALL. 5. Note that any hospitals that received direct shipments of this product from Zimmer Biomet or were consigned products will be sent a copy of the Risk Manager Recall Notice directly. It is important that you review the list of hospitals included with the email notification sent to your facility to identify additional accounts Zimmer Biomet has not notified. Using the Additional Accounts Form provided with the email notice sent to your facility, return contact information for any additional hospitals from which you removed affected product. This form will be used to notify each account. If there are no addit |  | 15000 | 15,000 | NATIONWIDE USARGENTINA AUSTRALIAAUSTRIABRAZILCANADACHILECHINACOSTA RICAECUADORGREECEINDIAITALYJAPANKOREA, THE REP OF (SOUTH)MALAYSIAMEXICONETHERLANDSPANAMAPOLANDSINGAPORETHAILAND | Open, Classified |  |  | 15000 | Not_Computer
Z-1114-2017 | 76082 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Orthopedic Salvage System (OSS) 9cm |  |  | December 16, 2016 | 2016 | Zimmer Biomet, Inc. | labeling for specific lots of the Standard OSS Modular Proximal Tibia incorrectly identifies the components as AVL Modular Proximal Tibia.The mislabeling could have the potential risks:A delay in surgery less than 30 minutes.Patient undergoes a revision with increased risk of infection. | Under Investigation by firm | On 12/16/2016 and 12/21/2016, Zimmer Biomet contacted distributors via telephone to notify them of the issue and requested that the product be placed in a quarantine location to prevent use. On 12/22/2016 URGENT MEDICAL DEVICE RECALL REMOVAL- LOT SPECIFIC notifications were sent to the affected distributors and Risk Managers / Surgeons via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Distributor responsibilities include: 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship and affected hospitals and surgeons within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet. a. Request a Return Authorization Number via email to rgarequest@zimmerbiomet.com. Be sure to specify RECALL as the RGA type. b. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com. c. Include a hardcopy of Attachment 1 with your shipment for immediate processing. 5. Note that any hospitals and/or surgeons that received direct shipments of this product from Zimmer Biomet will be sent a copy of the Risk Manager and Surgeon Recall Notice directly. It is important that you review the list of hospitals and/or surgeons included with the email notification sent to your facility to identify additional accounts Zimmer Biomet has not notified. Using the Additional Accounts Form provided with the email notice sent to your facility, return contact information for a |  | 16 | 16 | WINJCAMDWALAMATX | Terminated | July 21, 2017 | 217 | 16 | Not_Computer
Z-1119-2017 | 76124 | No | MTF |  |  | Immunology | PMA | IMMULITE/IMMULITE 1000 Systems, Third Generation PSA, Catalog Number/REF LKUP1, SMN 10380956 (OUS Kit Catalog) |  |  | December 21, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Negative bias with IMMULITEÂż /IMMULITEÂż 1000 Third Generation PSA (prostate-specific antigen) kit. The investigation by Siemens has confirmed that patient samples exhibit an average negative bias of approximately -22% with a range of -16% to -27% for samples recovering from 0.08 to 16.2 ng/mL. | Nonconforming Material/Component | Siemens issued an Urgent Medical Device Recall Notice letter dated December 21, 2016, to their U.S.A customers via Fed Ex and an Urgent Field Safety Notice letter dated December 2016, to customers outside the US (OUS). The notices instruct customers to discontinue use of and discard the affected kit lots listed in Table 1 of the customer communication. Customers are requested to review the UMDR and or the UFSN with their Medical Director. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within 30 days. U.S.A customers are advised to fax the Effectiveness Check questionnaire to (312) 275-7795. Field service personnel will be sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. Customers with questions should contact their local Siemens technical support representative. |  | 1379 | 1379 kits (190 kits in U.S.A and 1189 Kits internationally) in total | Worldwide Distribution - US Nationwide and the countries of Denmark, Brazil, Croatia and Germany | Open, Classified |  |  | 1379 | Not_Computer
Z-1116-2017 | 76278 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza, picture archiving and communications system. |  |  | January 11, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software update for improvements and to resolve several issues | Under Investigation by firm | Siemens mailed a Customer Letter to affected customers on 01/13/2017 to provide information about the latest software change that is now available to address several issues. |  | 62 | 62 systems | Distributed throughout the United States. | Open, Classified |  |  | 62 | Computer
Z-1117-2017 | 76119 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Na+ Slides, (UPN 10758750004812) |  |  | December 20, 2016 | 2016 | Ortho-Clinical Diagnostics | There is a potential for biased results to be generated over the 10 day on-analyzer storage limit when using VITROS Na+ Slide cartridges warmed at room temperature and loaded onto the analyzer after using the minimum warm up requirements stated in the IFU. If Quality Control (QC) results are within acceptable limits, reported results are acceptable and were not affected by this issue. | Under Investigation by firm | Ortho Clinical Diagnostics sent a Urgent Product Correction Notification dated December 20, 2016, to all affected customers via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who have a VITROS analyzer that can process Na+ Slides to inform them of the potential for biased results on Na+ Slides over the 10 day on-analyzer storage limit for Na+ Slides. Customers were instructed to warm up (equilibrate) all Na Slide cartridges for a minimum of eight hours at room temperature prior to loading them onto the analyzer to help ensure the slides have acceptable performance throughout the 10 day on-analyzer storage limit. Foreign affiliates were informed by email on December 20, 2016, of the issue and instructed to notify their Customers of the issue and the appropriate actions to take. For further questions, please call (585) 453-4224. |  | 659957 | 659,957 Total Units (249,570 units domestically &amp; 410,387 units internationally) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Bermuda, Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 659957 | Not_Computer
Z-1118-2017 | 76273 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Mammomat Inspiration full, field digital,system,x-ray,mammographic |  |  | January 11, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software error | Software design | Siemens mailed a Customer Safety Advisory Notice (CSAN) letter dated January 11, 2017 to affected customers. The letter identified the affected product, the potential risks and the steps to take to avoid the risks associated with the issue. The letter also stated Siemens will release a field modification (software update) to resolve the issue as soon as possible. The new software will be provided to affected customers free of charge. |  | 55 | 55 units | United States Nationwide Distribution | Open, Classified |  |  | 55 | Computer
Z-1036-2017 | 76208 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation |  |  | January 07, 2017 | 2017 | Olympus Scientific Solutions Americas | It was discovered under rare circumstances the LED warning light circuit could fail. If this happens, while the instrument is emitting X-rays during at test, the LED warning lights will not illuminate. The test runs normally and will execute as before, but the light will not be on. | Radiation Control for Health and Safety Act | Olympus will bring the defect compliance : 1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office recall coordinator noted below are to be included in the notification.2. Tracking of customer/owner self-installation(s) of the software update should be done to the maximum possible extent via serial-number-tracked downloads, follow-up email reporting, return-mail card or by other similarly effective means. For further questions, please call (781) 419-3500. |  | 74 | US 74 | US Distribution to the states of : PA, MA, CO, OH, NJ, GA, CA, TX, IL, VA, MN, NH, WA, SC, AZ, LA, MD, OK, IN, NY, NC and KS. | Open, Classified |  |  | 74 | Computer
Z-1089-2017 | 76131 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Optilite IgG4 Kit |  |  | December 23, 2016 | 2016 | The Binding Site Group, Ltd. | A change to the lower end of the measuring range and product insert did not reflect the change.The measuring range provided in instrument parameters is correct but contradicts the information provided in the product insert | Under Investigation by firm | Customer notification letters were sent on 12/23/16. The Binding Site recommends to continue testing on the kit as normal and refer to the measuring ranges and limit of quantification stated in the notice and as indicated in the amended insert (INS009.OPT.A, Version: 21st December 2016) provided with this notification.The notice needs to be passed on to all those who need to be aware within the organization or to any Organization where the potentially affected devices have been transferred.Return the completed E-Back Form to your local Binding Site Representative for records and note that The Binding Site intends to report this correction to the FDA recall coordinator.If you have any questions contact Technical Services. |  | 9 | 9 users | U.S. distribution to the following; FL, TX, PA, CA, MA, WA, MI, CT. No foreign distribution. | Terminated | May 03, 2017 | 131 | 9 | Not_Computer
Z-1019-2017 | 75904 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OPT Croutons is a synthetic bone substitute designed for the repair of bony defects and is made of pure beta-tricalcium phosphate. The tricalcium phosphate is a soluble calcium phosphate salt. |  |  | December 05, 2016 | 2016 | Acumed LLC | OPT Croutons, part number 65-0210-S, is recalled due to the result of inadequate documentation of the sterilization dose used in production. Product sterilized at the site did not receive the minimum dose required per the new performance qualification. | Vendor change control | Acumed sent the Urgent Recall Notice letter, dated December 1, 2016, to US customers on December 5, 2016.Customers are advised to follow the instructions provided by Acumed:IMMEDIATELY stop using and/or distributing the product and complete the following actions:1. Perform a physical count of your inventory.2. Remove the affected OPT Croutons, such that the croutons are out of service.3. Record this data on the survey found following this recall letter.4. Fill out the survey and return it via email to RecaiiNotification@acumed.net.5. For any product that needs to be returned:Domestic Customers: please contact Acumed Agent Inventory via email at Agentlnventory@acumed.net or contact your agent directly to be assigned a Return Material Authorization (RMA) number and applicable shipping information.International Customers: please contact Acumed Business Services via email at BusinessServices@acumed.net or via phone at (877) 627-9957 Option 4 (US and Canada) or +001 (503) 627-9957 and choose Option 4 (outside of USand Canada) to be assigned a Return Material Authorization (RMA) number and applicable shipping information.6. Once completed, please return the OPT Croutons that have been removed to Acumed using the Return Label provided by Acumed and included in your replacement shipment.Please contact Acumed Business Services via email at BusinessServices@acumed.net or via phone at (877) 627-9957 Option4 (US and Canada) or +001 (503) 627-9957 and choose Option 4 (outside of US and Canada) with any questions regarding this notification. |  | 83 | 83 units distributed in the US | distributed in FL, IN, KS, MI, MO, NE, OH, PA, TX and Peurto Rico. | Terminated | March 22, 2017 | 107 | 83 | Not_Computer
Z-0788-2017 | 75486 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II Implantable Drug Infusion System. This system includes: Model 8870 Software Application Card, Model 8840 N'Vision Clinician Programmer, and Model 8637 SynchroMed II Implantable Drug Infusion Pump (supplied in 20 ml or 40 ml reservoir size.) |  |  | October 03, 2016 | 2016 | Medtronic Neuromodulation | Medtronic is following up to a May 2013 communication regarding the Priming Bolus function for the SynchroMed Infusion System. Medtronic is updating the Model 8870 software application card (to version AAU01) and the SynchroMed pump labeling to address the priming bolus issue. | Under Investigation by firm | Medtronic Representatives began visiting consignees on 10/3/2016 to replace previous version Model 8870 software application cards, contained within the model 8840 NVision Clinician Programmer, with the updated version (new version is AAU01). Beginning October 20, 2016, Implanting and Following physicians were sent an Urgent Medical Device Correction letter dated September 2016, . This letter to physicians is informing them of the updated Model 8870 software application card, and changes to SynchroMed infusion system labeling. A response form is requested to be returned to Medtronic.Customers with questions were instructed to contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 22298 | 22,298 software cards | Worldwide Distribution - (US Nationwide) Internationally to Germany, Greece, Australia, United Kingdom, Switzerland, Italy, Norway, France, Canada, and Malta. | Open, Classified |  |  | 22298 | Computer
Z-1121-2017 | 76249 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Strata II / Strata NSC Valves (includes Burr Hole and Lumbar Peritoneal. All Catalog Numbers. |  |  | January 17, 2017 | 2017 | Medtronic Neurosurgery | Medtronic Neurosurgery initiated a recall after receiving five (5) Product Experiences between February 15, 2015 and August 23, 2016 reporting of discrepancies between the pressure/performance level (PL) at which the valve was set and the readings determined via an x-ray image. | Other | Medtronic sent an Urgent Product Advisory Notice dated January 19, 2017, was sent to customers to inform them that Medtronic Neurosurgery is initiating a voluntary field corrective action (FCA) for the Medtronic Strata II / Strata NSC valves (includes Burr Hole and Lumboperitoneal. Excludes StrataMR). Medtronic is notifying healthcare providers of an Instructions For Use (IFU) Warning update for the affected products by sending a Product Advisory Notice. For further questions regarding this recall please call (805) 571-8400.. |  | 315997 | 315,997 | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the countries of :AlgeriaIndiaPolandArmeniaIndonesiaPortugalAustraliaIranAustriaIraqQatarBangladeshIrelandRepublic of MontenegroBarbadosIsraelRepublic of SerbiaBelarusItalyRomaniaBelgiumJamaicaRussian FederationBosnia and HerzegovinaJapanSaudi ArabiaBotswanaJordanSingaporeBulgariaKazakhstanSlovakiaCambodiaKenyaSouth AfricaCanadaKuwaitSouth KoreaCanary IslandsLatviaSpainCayman IslandsLebanonSri LankaChinaLibyaSwedenCroatiaLithuaniaSwitzerlandCyprusMacauSyria | Terminated | June 02, 2017 | 136 | 315997 | Not_Computer
Z-1125-2017 | 76062 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Covidien Nellcor SpO2 Forehead Sensor |  |  | December 16, 2016 | 2016 | Covidien LLC | Labeling error: The label indicates the product is latex free, but the headband component included with the forehead sensor does contain latex. | Process control | Medtronic initiated on December 16, 2016. Customers were notified via Federal Express and the letter instructed customers to discontinue use of the product and return all inventory to Medtronic (Covidien)  195 McDermott Road, North Haven, CT 06473 Attn: Field Returns Department. Complete a Recalled Product Return Form (Attachment A, included with letter) and submit it to customer service so a Returned Goods Authorization number (RGA#) can be issued.Questions contact Medtronic representative or customer service at 800-882-5878. |  | 7453 | 7453 Units | Nationwide.Foreign:BelgiumGermanyFranceNetherlands | Open, Classified |  |  | 7453 | Not_Computer
Z-1122-2017 | 76090 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250-Proton Radiation Therapy |  |  | December 16, 2016 | 2016 | Mevion Medical Systems, Inc. | An error can occur causing Delta corrections to be lost when one setup field is closed and another is opened | Software design | Mevion sent an Important Safety Notice to customers on 12/16/16 by email and Salesforce. The letter identified the affected product, problem and actions to be taken. Mevion provided mitigation and will release and install a software patch that removes the possibility that Deltas can be lost between multiple Setup fields. For further information contact rcessac@mevion.com. |  | 2 | 2 units | US Nationwide in the states of OK, NJ | Terminated | April 13, 2017 | 118 | 2 | Computer
Z-1126-2017 | 76130 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Intact Disposable Wands (part of the Intact System) |  |  | December 28, 2016 | 2016 | Medtronic Advanced Energy, LLC | Product sterility compromised due to breach of sterile barrier | Package design/selection | Medtronic Advanced Energy notified customers by letter on December 27, 2016 via Next Day FEDEX. The communication advises users to immediately stop using the affected Intact Disposable Wands and to quarantine them until they can be returned to Medtronic. Medtronic Sales representatives will arrange for the return of all recalled product. |  | 10 | 10 units | Worldwide Distribution - US Nationwide in the countries of China, Greece, Netherlands, Egypt, Turkey, and Switzerland. | Terminated | May 11, 2017 | 134 | 3303 | Not_Computer
Z-1124-2017 | 76182 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Synthes Radial Head Prosthesis System, Surgical instrument motors and accessories |  |  | December 29, 2016 | 2016 | Synthes (USA) Products LLC | There is the possibility that the radial stem may loosen post-operatively at the stem bone interface. | Under Investigation by firm | An Urgent Notice was mailed to consignees explaining the reason for the recall. Customers were asked to immediately review their inventories for the affected product and complete and return the Verification form that was included with the notification and return any of the affected product that was found. |  | 50311 | 50,311 units | United States Nationwide Distribution. | Open, Classified |  |  | 50311 | Not_Computer
Z-1139-2017 | 76206 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Empowr PS Knee Tibial Insert; Model 343-14-709 |  |  | January 10, 2017 | 2017 | Encore Medical, Lp | The Empowr PS Insert was reported as missing the impaction slot feature. | Component design/selection | DJO Global sent an Urgent Field Safety Notice dated January 10, 2017, to all affected customers. The Field Safety Notice instructs the agents to notify all individuals within their organization or to any organization where the potentially affected product has been transferred. Customers with questions should contact customer service at 1-800-456-8696. For questions regarding this recall call 512-832-9500. |  | 24 | 24 units | Nationwide Distribution to CO, NY, and FL | Terminated | May 11, 2017 | 121 | 24 | Not_Computer
Z-1138-2017 | 76134 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-Ray System, Model DRXR-1, Catalog No. 1019397 (all other countries), Catalog No. 1060177 (China only) -- Carestream Health, Inc. 150 Verona Street, Rochester, NY 14608 |  |  | January 12, 2017 | 2017 | Carestream Health Inc | The firm received a complaint alleging that a DRX Revolution System could not be stopped when applying the brake. The Revolution crashed into an elevator. | Device Design | Carestream Health Inc. sent an URGENT: MEDICAL DEVICE CORRECTION Letter dated December 29, 2016, and response forms were sent to the US customers via Certified Mail-Return Receipt Requested on January 12, and January 13, 2017, informing them of the problem. Customers are instructed to read and complete all information and return the response form within 5 working days to the following e-mail address: postmarketra@carestream.com, or fax the completed form to (585) 454-4493. UMDC Letters and Response Forms to Canadian Customers were sent the week of January 9, 2017, via courier. Field Safety Notice (FSN) Letters for the rest of world are in process. FSN Letters for the EMEA should be sent via courier starting later this week.US Carestream Customer Care Center #: 1-800-328-2910 (24/7). Outside of the US, please call your local Service support number. |  | 2773 | US: 1403 units; Foreign: 1370 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico, and to the countries of : Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Denmark, Dominican Republic, Dubai, Finland, France, Germany, Ghana, India, Iran, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Libya, Malaysia, Maldives, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 2773 | Not_Computer
Z-1137-2017 | 75226 | No | LJW | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Celerity ECG Cable Accessory Pack, STERILE, Rx ONLY - distributed under the following labels: (a) angiodynamics UPN/Product No. H787471431, REF/Catalog No. 47-143 -- Legal Manufacturer: AngioDynamics, Inc. 10 Glens Falls Technical Park, Glens Falls, NY 12801; and (b) medCOMP UPN/Product No. MRLCA000, REF/Catalog No. MRLCA000 -- medCOMP, 1499 Delp Drive, Harleysville, PA 19438 --- Description: The Celerity ECG Cable Accessory Pack is comprised of two sterile, single use components: a cable cover and an ECG clip cable (alligator clip). The cable cover is used to sleeve the remote control cable component of the Celerity System so that the user can remotely control the monitor within the sterile field. The stereo jack end of the remote control cable pierces the sterile cable cover and interfaces with the stereo jack connector end of the ECG clip cable. The alligator clip side of the ECG clip cable is clamped to the stylet or guidewire to provide the intravascular cavity ECG reading. |  |  | September 07, 2016 | 2016 | Angiodynamics Inc. (Navilyst Medical Inc.) | The Celerity ECG Cable Accessory Pack has the potential to contain an ECG Clip Cable that was not properly manufactured. The affected product is not soldered per specification, and may result in the device not properly conducting the ECG signal. | Employee error | On 9/7/2016, AngioDynamics, Inc. sent URGENT VOLUNTARY MEDICAL DEVICE RECALL Notifications (dated 9/7/2016) to 47 consignees via Federal Express. Consignees were instructed to segregate and return all affected devices to AngioDynamics and complete and return the Reply Verification Tracking Form, provided in the recall notification. Please call AngioDynamics Customer Service at 1-800-772-6446 Monday - Friday between 8:00 am and 7:00 Pm EST to obtain a replacement or credit for returned product. |  |  |  | AZ, CA, FL, GA, IL, IN, KS, MD, MO, MT, NC, NJ, NV, OH, OR, PA, TN, TX, WI | Terminated | April 06, 2017 | 211 |  | Not_Computer
Z-1142-2017 | 75149 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Merge Eye Station. |  |  | December 09, 2016 | 2016 | Merge Healthcare, Inc. | Merge received reports of Merge Eye Station being unable to capture images of the eye as expected per the intended use of the product. Merge is recalling product from the field to reduce the risk to patient health. | Software design | Merge sent an Urgent Medical Device Recall letter on December 9, 2016. The letter was addressed to Office Manager . The letter described the Product, Issue, Potential Harm, Containment by the Customer / User, Action by Merge and Actions by Customer. Requested customers to provide response to the "Response to this Notification is Required". If the product was further distributed advised consignees to notify them at once. For questions send an e-mail to recall@merge.com. E-mails to be sent to consignees on 12/19/2016. For further questions, please call (262) 367- 0700. |  | 182 | 182 capture stations | Worldwide Distribution - USA (nationwide) including the states of : CA, CO, CT, FL, HI, IL, IN, LA, MD, MA,M I, MN, MS. MK. NY, NC,ND, OH, OK, OR, PA, RI, TN, VA, WA and WV., and the countries of Canada and United Kingdom. | Open, Classified |  |  | 182 | Not_Computer
Z-1143-2017 | 76162 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 12 Defibrillator/Monitor with End-Tidal CO2 option. |  |  | January 16, 2017 | 2017 | Physio-Control, Inc. | The End-Tidal CO2 (EtCO2) reading can intermittently show a value of XXX after start-up or during device operation. | Process control | Physio Control sent the Urgent Medical Device Correction LIFEPAK 15 Monitor/ Defibrillator and LIFEPAK 12 Defibrillator/Monitor with End-Tidal CO2 (EtCO2) option letter, dated January 2017, and confirmation sheets to consignees on January 16, 2017.Physio-Control is contacting customers with affected devices to schedule the correction of their device(s), which will include replacing the EtCO2 module.Required Customer Actions:1. Please forward this information to all affected sites, trainers and users that may have affected LIFEPAK 12 or LIFEPAK 15 device(s).2. Consignees should follow the instructions on the Confirmation Sheet for each serial number listed in their possession. Promptly return the completed Confirmation Sheet to Physio-Control.3. If consignees experience the XXX after power-up or during device operation, please contact Physio-Control to arrange for servicing of their device.Consignees should return completed confirmation sheet:-By fax to: 1-866-448-9567-By email to: rsrecall@physio-control.com-Or by mail to: Physio-Control, Inc. P.O. Box 97006, Dept. 15N Redmond, WA, 98073-9706Consignees with questions can call the firm at 1-800-442-1142, option 7, 6:00 A.M. to 4:00 P.M. (Pacific), Monday  Friday. |  | 7 | 7 units in the US and 13 modules worldwide (5 international; and 8 owned by Physio) | Worldwide distribution. LifePak 12 distributed in US states of AL, GA, NY, SC, and in Australia.LifePak 15 distributed nationwide in the US including Puerto Rico and to ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, CANADA, CHILE, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, FRENCH GUIANA, GERMANY, GREECE, HUNGARY, ICELAND, IRELAND, ITALY, LITHUANIA, LUXEMBOURG, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, SAUDI ARABIA, SENEGAL, SERBIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, TAHITI, UNITED ARAB EMIRATES, AND UNITED KINGDOM. | Open, Classified |  |  | 7 | Not_Computer
Z-1147-2017 | 75338 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Thayer MiniSpacer Dual Spray MDI Adapter Counter Incrementing Actuator with 15mm O.D./I.D. Connections, Reference no. 1543A. |  |  | September 29, 2016 | 2016 | Thayer Medical Corporation | The insert and case label description of the device incorrectly reads Counter Incrementing Actuator with 22mm O.D./I.D. Connections. | Labeling design | Thayer Medical sent an Urgent Medical Device Recall letter dated September 30, 2016, to all affected customers. The letter informed customers that Thayer Medical has identified a labeling error on REF 1543A, Thayer MiniSpacer Dual Spray MDI Adapter with Counter Incrementing Actuator and 15mm O.D./I.D. Connections. Customers were informed to complete and return the attached Customer Reply Form by email to info@thayermedical.com or fax to 1-520-790-5854.For questions regarding this recall call 520-790-5393, ext 2210. |  | 2550 | 2,550 units | Nationwide Distribution | Terminated | February 07, 2017 | 131 | 2550 | Not_Computer
Z-1145-2017 | 75557 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | HOTLINE Blood and Fluid Warmer, Mole l HL-90. Designed for use with the Hotline Fluid Warming Set (L-80, L-70 and L-70N) to warm blood and intravenous (I.V.) fluids and deliver them to the patient's intravenous access site at normothermic temperatures under gravity flow conditions. |  |  | October 31, 2016 | 2016 | Smiths Medical ASD, Inc. | HOTLINE Blood and Fluid Warmers included Greek and Danish language labeling instead of the required Finnish labeling. | Labeling mix-ups | Smiths Medical sent an Urgent-Medical Device Field Safety Notice dated October 28, 2016, to all affected customers. The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. Advised consignees to notify customers and to review and complete the Recall Confirmation Form and return it to Smith Medical via e-mail to FCA.Response@Smiths-medical.com. For further questions, please call (651) 633-2556. |  | 33 | 33 | Internationally to Finland | Open, Classified |  |  | 33 | Not_Computer
Z-1151-2017 | 76275 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InterFuse Tail Traction Tool, Part Number 9116. |  |  | January 13, 2017 | 2017 | Vertebral Technologies, Inc. | Fit issues have been identified with the Tail Traction Tool (PN 9116) when used with the Interfuse T tailed implant. The InterFuse T Tail may not fit into the Thru Hole in the instrument. | Nonconforming Material/Component | Consignees were sent via email a VTI "Advisory Notice" letter. The letter listed the Concern, Impact, Alternate Methods, Affected Lots, and Your Responsibility. Consignees were requested to acknowledge receipt of this notice, Discontinue use of the Tail Traction Tool when implanting the InterFuse T tailed implant, Notify individuals that may potentially use the device and alternate module assembly methods. Advise consignees to contact Customer Service (1-877-912-5401 in US , +-952-912-5400 outside US or customerservice@vti-spine.com to obtain a replacement instrument. |  | 63 | 63 pieces | US: IL, WI, OH, NY, MT, FL, MD, IN, TN, TX, MN.OUS: Italy, Taiwan Spain, Belgium, Canada, Netherlands, Sweden, Switzerland, Portugal, Brazil. Germany, Ireland, | Open, Classified |  |  | 63 | Not_Computer
Z-0927-2017 | 75736 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MEDRAD Intego PET Infusion System - Pump |  |  | November 07, 2016 | 2016 | Bayer Healthcare | Bayer has determined that all current Source Administration Sets (SAS) used with the Medrad Intego Pet Infusion System may produce particulates in radiopharmaceutical (RP) vials when the needle is inserted. | Under Investigation by firm | Bayer mailed an Urgent Medical Device Field Safety Notice dated November 3. 2017 to affected customers to inform them of the issue. Customers were asked to immediately quarantine unused product and resume use once the qualified in-line filter is received or to request a RGA (Return Goods Authorization) number to return the product. Customers were asked to complete the response form indicating if their institution does or does not have affected product and return the form via fax to 1-412-406-0941. Customers with questions were instructed to contact the Customer Care team at 1-800-633-7231, opt 2.For questions regarding this recall call 412-767-2400.On December 5, 2016, the recalling firm sent a new letter indicating the inline filter has been successfully tested and that they have qualified a compatible in-line filter to be used with the Patient Administrative Set (PAS) that will allow you to resume use of your Intego system. |  | 178982 | 178,982 (86,172 units in U.S.) | Nationwide Distribution | Open, Classified |  |  | 178982 | Not_Computer
Z-1150-2017 | 76146 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 9805 |  |  | January 03, 2017 | 2017 | Stryker Leibinger GmbH &amp; Co. KG | An incorrect sterility status on the label, non-sterile products labeled as sterile | Labeling mix-ups | Stryker sent an " URGENT MEDICAL DEVICE RECALL " notifications dated January 3, 2017, to the affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, instructions for returning affected product, and instructions for responding to the formal recall notification.Please report any adverse events or product quality problems to Stryker CMF Customer Service: 1 (800) 962-6558. |  | 619 | 619 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of :Australia, Canada, Chile, Spain, South Africa and Taiwan. | Terminated | July 27, 2017 | 205 | 619 | Not_Computer
Z-1148-2017 | 76017 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AtriClip" LAA Exclusion System with preloaded Gillinov-Cosgrove" Clip - A single PRO2 device is packaged in a conventional thermoformed polyethylene terephthalate glycol (PETG) tray with a TyvekÂż lid. A single-walled non-corrugated box completes the single unit package. Six (6) single unit PRO2 devices are packaged into one (1) corrugate shipper box. |  |  | November 30, 2016 | 2016 | AtriCure, Inc. | A complaint of the PRO2 jaw breaking prior to surgery being performed. | Component design/selection | Recall notification letters were mailed to all Atricure customers. Email notification letters were forwarded to all Atricure field-based personnel, who may also conduct in-person meetings. Telephone calls will be made as needed and effectiveness checks will be conducted. |  | 464 | 464 units | Devices were distributed in the following states: AL, AR, AZ, CA, CO, FL, GA, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO, MS, NC, NE, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, &amp; WI.Devices were distributed in the following countries: Austria, Germany, Netherlands, Poland, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 464 | Not_Computer
Z-1185-2017 | 75939 | No | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Junction Box - 1104343 and 1193560 |  |  | December 02, 2016 | 2016 | Invacare Corporation | A quality issue with the junction box used in the bariatric bed junction that may omit sparks and result into a fire. | Component design/selection | All customers were notified via letter with return receipt. This notification will include a provider letter, provider response card, impacted unit list and a consumer letter. Consignees will be instructed conduct sub-recalls as necessary. Follow-up telephone calls will be made to nonresponders. |  | 161 | 161 units | Products were distributed in the following states: AK, AL, AR, CA, CO, CT, FL, HI, ID, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI.No foreign accounts | Open, Classified |  |  | 864 | Not_Computer
Z-1181-2017 | 76200 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products FS Calibrator 1, REF 680 1873 12 x 3 mL, Rx ONLY, IVD; used in conjunction with VITROS Chemistry Products Calibrator Kits to calibrate VITROS 5,1 FS/4600 Chemistry and integrated Systems. --- Ortho-Clinical Diagnostics, Inc. Rochester, NY 14626 --- Ortho-Clinical Diagnostics Pencoed, CF35 5PZ, UK ---- Background: VITROS assays must be calibrated prior to use in order to be used to produce assay results. The Assay Data Disk (ADD) contains calibrator assigned values/ calculation data for multiple VITROS Calibrators and Lots. The Data Release Version (DRV) number identifies the version of the ADD. Each time the ADD is updated with new ata, the DRV number increments higher. The DRV with the highest number contains the most |  |  | January 05, 2017 | 2017 | Ortho-Clinical Diagnostics | The company received a customer complaint for the inability to calibrate VITROS Chemistry Products dLDL Reagent when using VITROS FS Calibrator 1, Lot 91461 and ADD (DRV 5908). Assay Data Disk (ADD) DRV versions 5904, 5905, 5906, 5907, 5908, 5909, and 5910 are all affected by this issue. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter (Ref. CL2017-002) and ADD DRV 5911dated January 5, 2017, to all affected customers. The letters were sent via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (Ortho Direct, Drop-Ship and US Federal Government) to inform them of the issue and advised them to install Assay Data Diskette, DRV 5911 or above on their VITROS System. Foreign affiliates were informed by e-mail on January 5, 2017, of the issue and instructed to notify their consignees of the issue and required actions. Customers with questions were instructed to contact Ortho Care Technical Solutions Center at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 802 | US: 160 kits, Foreign: 642 kits | Worldwide Distribution - US (Nationwide and, Puerto Rico) and Foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Poland, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 802 | Computer
Z-1176-2017 | 76260 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Potential exists for an incorrect patient image being displayed which could result in the delay in diagnosis or treatment. | Software design | The recalling firm issued a letter dated 3/28/2016 via email on 4/4/2016 or via certified mail if they did not have an email address for the customer. A second letter was issued 7/11/2016 to nonresponders. |  | 29 | 29 sites have the potentially affected software | Distribution was made to medical facilities in AZ, CA, CO, FL, HI, IL, MA, MD, MI, MO, NY, OH, PA, SC, TN, UT, and WI. There was no foreign/government/military distribution. | Open, Classified |  |  | 29 | Not_Computer
Z-1177-2017 | 76133 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Sheridan/CF, Cuffed Tracheal Tube, Catalog # 5-10114 and Sheridan/HVT Cuffed Tracheal Tube, Catalog # 5-10315 |  |  | October 18, 2016 | 2016 | Teleflex Medical | Labeling inconsistency - units labeled with a 7mm tube may contain a size 6.5mm tube, and units labeled 7.5mm tube may contain a size 5.5mm tube. | Unknown/Undetermined by firm | Teleflex sent an Medical Device Recall Notification letter dated October 18, 2016, to all affected customers. Customers were advised to immediately discontinue use and to quarantine product. The letter to distributors requested a sub-recall. Both letters included an acknowledgement form to be returned. Customers with questions were instructed to contact their local representative or Customer Service at 1-800-246-6990. |  | 25274 | 25,274 units | Worldwide Distribution - US including AL, AZ, ARK, CA, CO, CT, FL, GA. HI, KS, KY, IL, IN, IA, LA, MD, MA, MI, MN, MS, MO, MT, NE, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV and Internationally to Canada, Chile, &amp; Japan | Open, Classified |  |  | 25274 | Not_Computer
Z-1178-2017 | 76323 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, EMBLEM S-ICD Programmer Model 3200 Programmer. |  |  | January 12, 2017 | 2017 | Boston Scientific Corporation | There is a potential for radio frequency (RF) interference to alter wireless communication from a programmer, which in rare instances may cause an S-ICD to perform an unintended command. This behavior can only occur during an active, in-clinic interrogation/programming session with the Model 3200 S-ICD programmer. There is no risk of this behavior occurring when the LATITUDE Patient Management System communicates with an S-ICD in an ambulatory setting. | Software design | Boston Scientific sent an "Important Medical Device Information" letter dated on January 12, 2017. The letter was addressed to "Dear Doctor". The letter included information on the Description and Clinical Implications, Rate of Occurrence, Recommendations, Affected Programmer &amp; Additional Information. For additional questions regarding this communication or to report clinical events, contact your Boston Scientific representative or Technical Services at 1 800-227-3422. |  | 4500 | 4,500 | Worldwide Distribution - USA ( nationwide) including PR and to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY., and to the countries of :Canada, Austria, Belgium, Czech Republic, Denmark, Finland, France, French Polynesia, Guadeloupe, Monaco, New Caledonia, Reunion, Germany, Greece, Hungary, Ireland, Israel, Italy, Latvia, Lithuania, Luxembourg, Netherlands, Norway, Poland, Portugal, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Australia, China, Egypt, Hong Kong, India, Iran, Japan, Jordan, Kuwait, Lebanon, Malaysia, New Zealand, Qatar, Saudi Arabia, Singapore, South Africa, Thailand, Tunisia, Turkey, United Arab Emirates, Argentina, Bermuda and Brazil. | Open, Classified |  |  | 4500 | Not_Computer
Z-1180-2017 | 76258 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge RadSuite software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | November 17, 2015 | 2015 | Merge Healthcare, Inc. | The values provided from the Pixel Value tool do not appear to be correct, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge sent an Urgent Medical Device Recall letter dated March 28, 2016 to all affected customers. The letter notifies the customer of the issue and informs them an upgrade is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.A second letter dated July 11, 2016, was issued asking for a response by August 2, 2016, even though all upgrades were made because the effectiveness checks were not yet completed. The letter and response form contained the same information as the first letter.The recalling firm reported emails were tracked for delivery confirmation and the letters were tracked for delivery by USPS certified services. For further questions, please call 877-744-5369. |  | 8 | 8 sites have the potentially affected version | US Distribution was made to medical facilities in AL, IN, MI, PA and WI. | Completed |  |  | 8 | Not_Computer
Z-1179-2017 | 75936 | No | JGP | Total protein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products UPRO Slides, UPN 10758750000593, |  |  | December 08, 2016 | 2016 | Ortho-Clinical Diagnostics | There is a potential for variable negatively and positively biased urine protein results using VITROS UPRO Slides. | Device Design | On December 8, 2016, Ortho Clinical Diagnostics distributed Product Correction Notification letters to their customers via FedEx, ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail. Customers are advised to check their stock for affected product and to immediately discontinue and discard all remaining inventory. Ortho will credit your account for any discarded VITROS UPRO Slides, VITROS Chemistry Products Calibrator Kit 10 and/or VITROS Chemistry Products UPRO Performance Verifier I &amp; II. Please complete the Confirmation of Receipt and return by December 22, 2016. Customers with questions, please contact Ortho CareTM Technical Solutions Center at 1-800-421-3311. |  | 3381390 | 37,571 packs/90 slides totalling 3,381,390 slides | Worldwide distribution - Nationwide U.S.A; International - Bermuda, Australia, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain and Poland. | Open, Classified |  |  | 3381390 | Not_Computer
Z-1189-2017 | 76295 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Buffered Sodium Citrate: (9NC) Blood Collection Tube (4.5 ml x 13 x 75 mm) |  |  | January 16, 2017 | 2017 | Becton Dickinson &amp; Company | A portion of this lot of blood collection tubes was manufactured with approximately twice the expected amount of buffered sodium citrate additive, an anticoagulant. | Process control | BD sent an Urgent Product Recall letter dated January 16, 2017 to customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Customer Recall Response Form and fax the completed form back to BD at 1-844-388-5534 or email the completed form to bd8872@stericycle.com. For questions call 1-844-265-7408. |  | 1488800 | 1,488,800 tubes | Worldwide Distribution - US Nationwide and the countries of Canada, Mexico, Indonesia | Open, Classified |  |  | 1488800 | Not_Computer
Z-1188-2017 | 76378 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Discovery MR450 |  |  | January 20, 2017 | 2017 | GE Healthcare, LLC | Potential safety issue with the patient bore heating on the Discovery MR450 narrow bore 1.5T MRI. In the event of a RF component failure, typically happening after extended system use, a small area on the bore surface may become warmer than normal. If this happens in an area where there is patient contact with the bore, and proper padding is not used, it may cause a serious burn. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter GEHC Ref# 60913 Dated January 20, 2017. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 60 | 60 | Worldwide Distribution- US (nationwide) including states of: CA, DC, IL, MN, MO, NY, SC, TX, WI, WV; and countries of: Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Poland, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, United Kingdom, and Uruguay. | Open, Classified |  |  | 60 | Not_Computer
Z-1186-2017 | 76183 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT |  |  | December 23, 2016 | 2016 | Philips Medical Systems (Cleveland) Inc | The intended use listed in the English and localized language IFUs for BrightView XCT systems did not match the intended use documented in the 510K submission. | Under Investigation by firm | All customers were notified via certified mail with return receipt. International customers were notivied via Philips employees. A field change order was initiated. Non-responders will be followed up on. Lastly, effectivenss checks will be conducted. |  | 425 | 425 | Products were distributed in the following states: AZ, CA, CO, DC, DE, FL, GA, IL, LA, MA, MD, MN, MO, MS, MT, NC, NJ, NY, OH, OK, OR, PA, PR, TX, VT, WA, &amp; WI.Product was distributed internationally to the following countries: Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, France, Gabon, Germany, Greece, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russian Federation, Saudia Arabia, Slovakia, South AFrica, Apain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Viet Nam. | Open, Classified |  |  | 425 | Not_Computer
Z-1193-2017 | 76025 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Stryker Restoris RIO Reamer Handle, Offset, Mako Rx only |  |  | September 29, 2016 | 2016 | Mako Surgical Corporation | Failure of the Offset Cup Reamer Handle not engaging its mating components. | Component change control | Customers were sent a recall notificaiton letter, dated 10/17/2016, addressing the reason for recall and the appropriate actions that should be taken by customers. Consignees are instructed to return any recalled product on hand to Stryker C/O Stericycle, 2670 Executive Dr. Suite A, Indianapolis, IN 46241, Attn RA 2016-083, RA 2016-119 - Event 8811. Consignees are also instructed to inform their users of this recall and forward the notice to affected individuals. A Business Reply Form was provided and is to be completed within 5 business days of receipt by consignees. Completed business reply forms are to be faxed to 855-620-5693 or emailed to strykerortho8811@stericycle.com. |  | 595 | 595 units | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI, WV, &amp; WY and the countries of Italy, Germany, Australia, China, Europe, Hong Kong, South Africa, Germany, Thailand &amp; Singapore. | Open, Classified |  |  | 595 | Not_Computer
Z-1190-2017 | 76236 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Offset Reamer Handle, T5766 |  |  | October 19, 2016 | 2016 | Greatbatch Medical | On June 14, 2016, Greatbatch discovered a significant increase in the number of field complaints for the T5766 Offset Reamer Handle in which the U-joints of the drive chain component seizes during orthopedic procedures rendering the device non-functional. The 2016 failure rate for this failure exceeds the occurrence rate as defined in the DFMEA. | Device Design | Consignee was sent via e-mail on 10/16/2016 a Greatbatch "Urgent Medical Device Recall" letter dated 10/19/2016 The letter described the purpose of the letter, Reason for Notification, Risk to Health, Recommendation, Instruction to Zimmer. Advised consignee to examine their inventory, quarantine the product and provide a list of the affected countries where the product was distributed. Also requested consignee to complete the Customer Notification Response form and return to Greatbatch at FieldActionCenter@Greatbatch.com. For questions contact your local Greatbatch Medical representative, Monday through Friday via telephone at 1-763-951-8235 8:00 AM (CT) to 5:00 PM (CT), or e-mail us at FieldActionCenter@Greatbatch.com . |  | 1441 | 1,441 | Worldwide Distribution - US Nationwide in the states of IN and the countries of Canada, Austria, Germany, Japan. | Open, Classified |  |  | 1441 | Not_Computer
Z-1123-2017 | 76160 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 15 Monitor, Automated External Defibrillator (AED), Manual mode monitoring and therapy functions are intended for use on adult and pediatric patients. Automated external defibrillation mode is intended for use on patients eight years of age and older. |  |  | January 12, 2017 | 2017 | Physio-Control, Inc. | The firm has become aware of a potential device issue where the LIFEPAK 15 device may no longer be able to deliver defibrillation energy due to a potential failure of an internal component (Relay) installed on the Therapy Printed Circuit Board Assembly (PCBA). | Other | Physio plans to send the URGENT MEDICAL DEVICE CORRECTION  ACTION REQUIRED LIFEPAK 15 Monitor/Defibrillator letter, dated January 2017, to their consignees on January 12, 2017.They were informed of the problem and how to identify a malfunction during the User Test of the device as described in the chapter titled Maintaining the Equipment within the LIFEPAK 15 Operating Instructions (or follow page 2 of the notification letter).If the User Test fails, the device will illuminate the Service indicator on the device. Please contact Physio-Control immediately to schedule the device correction. This field corrective action will include replacement of the Therapy PCB assembly on all devices included in the potentially affected population. This will occur as a combination of field and depot service.Customers are required to take the following actions:1. Please forward this letter to all of your sites, trainers and users that have an affected LIFEPAK 15 device as identified on the attached Confirmation Sheet.2. Follow the instructions on the Confirmation Sheet for each serial number listed in your possession. Promptly return the completed Confirmation Sheet to Physio-Control.3. If the device fails the User Test, as described above, contact Physio-Control immediately to arrange for correction of your device.4. Continue to perform the User Test as outlined in the Daily Operators Checklist within the LIFEPAK 15 Operating Instructions.Consignees should return completed confirmation sheet to Physio:-By fax to: 1-866-448-9567-By email to: rsrecall@physio-control.com-Or by mail to: Physio-Control, Inc. P.O. Box 97006, Dept. 15N Redmond, WA, 98073-9706.Customers with questions call the firm at 1-800-442-1142, option 7, 6:00 A.M. to 4:00 P.M. (PST), Monday  Friday. |  | 199 | 338 units total (US =139 units; International = 199 units) | Worldwide Distribution-US ( nationwide) in states of: AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, VA, WA, WI, and WV and countries of: Australia, Brazil, Canada, Czech Republic, Denmark, Finland, France, French Guiana, Germany, India, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Qatar, Saudi Arabia, Slovenia, South Africa, Sweden, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 199 | Not_Computer
Z-1103-2017 | 75971 | No | PAO | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Reverse Shoulder System Humeral Tray |  |  | December 20, 2016 | 2016 | Zimmer Biomet, Inc. | Higher than anticipated rate of fracturing due to design. Risk of revision surgery due to fracturing. | Device Design | On 12/20/2016, URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC notifications were sent to the affected Risk Managers and Distributors via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Risk Manager Responsibilities: 1. Review this notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer Biomet sales representative quarantine all affected product. 3. Your Zimmer Biomet sales representative will remove the affected product from your facility. 4. Complete Attachment 1  Certificate of Acknowledgement. a. Return a digital copy within three (3) days to corporatequality.postmarket@zimmerbiomet.com. b. Retain a copy of the Acknowledgement Form with your recall records in the event of a compliance audit of your facilities documentation. 5. If after reviewing this notice you have further questions or concerns please call 411 Technical Services at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voice mail or be transferred to an on-call representative in the case of an emergency. Alternatively, your questions may be sent by email to corporatequality.postmarket@zimmerbiomet.com. Distributor Responsibilities 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form |  | 268 | 268 | Nationwide. CANADA AUSTRALIA CHILE COLOMBIA KOREANetherlands | Open, Classified |  |  | 268 | Not_Computer
Z-1191-2017 | 76209 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | LIMA Modular Revision Hip Stem Model 428-01-050_110 |  |  | January 10, 2017 | 2017 | Encore Medical, Lp | Lima Proximal Bodies were inadvertently re-sterilized. The safety screw which affixes the distal and proximal bodies of the stem includes a thread-locking plug made from UHMWPE, which is not approved for repeated gamma sterilization. | Nonconforming Material/Component | DJO Global sent an Urgent Field Safety Notice letter dated January 10, 2017 to customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete the Acknowledgement and receipt Form. Customers were instructed to contact Customer Service at 1-800-456-8696 to explore options for a replacement order. |  | 13 | 13 devices | US Nationwide - US Nationwide in the states of MS, OH, RI, NY, OK, TX, CA, HI | Terminated | March 31, 2017 | 80 | 13 | Not_Computer
Z-1192-2017 | 76290 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Integra Cadence Total Ankle System |  |  | January 20, 2017 | 2017 | Integra LifeSciences Corp. | Posterior tibial fractures have been reported. | Use error | The firm intends to update the surgical technique only. No product will be returned. US and non-US consignees were notified by letter delivered by traceable fax or e-mail on 01/20/2017. |  | 944 | 944 units | Nationwide | Terminated | March 30, 2017 | 69 | 944 | Not_Computer
Z-1194-2017 | 76292 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BONE MARROW ASPIRATION NEEDLE 15G x 2.688 MAX and 15G x 4 MAX |  |  | January 18, 2017 | 2017 | Argon Medical Devices, Inc | The recalling firm received a complaint from one of their customers of a potential defect in the forming ofthe primary package of the blister that may compromise the sterility of the product inside the tray. | Equipment maintenance | Argon Medical Devices, Inc. is sending an Urgent Product Recall Notice letter dated January 17, 2017, by registered mail to all affected customers. Customers were asked to determine if their inventory shows the package defect and discontinue use of the device and promptly return the affected product. For questions regarding this recall call 903-403-0749. |  | 497 | 497 boxes | Worldwide Distribution - US (nationwide) and Internationally to Switzerland, Canada, Australia, New Zealand, Ecuador, Colombia, and Japan. | Terminated | May 17, 2017 | 119 | 1284 | Not_Computer
Z-1197-2017 | 76194 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilization System, Product Code: 10033, 10033-002 |  |  | January 09, 2017 | 2017 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has identified that in a rare sequence of events following an external power outage, the system door interlock can be become disengaged during system reboot for a 20-second duration. | Software design | Advanced Sterilization will send an Field Safety Notice dated January 9, 2017,to inform customers that Advanced Sterilization Products (ASP) has identified that in a rare sequence of events following an external power outage, the system door interlock can be become disengaged during system reboot for a 20-second duration. ASP is sending the letter to remind customers to follow the instructions in the STERRAD NX System Users Guide. Specifically, after cycle initiation, the user should only open the chamber door of the STERRAD NX System whenprompted by the Graphical User Interface (GUI) per the Users Guide. Customers are instructed to complete and return the attached Business Reply Card (double-sided) for their acknowledgement within 3 business days of receipt of this notification letter via mail by using the pre-paid postage, by emailat ASP5816@stericycle.com, or by fax to Stericycle: 8662297775. For questions regarding the business reply card, customers are instructed to contact Stericycle at (877) 497- 2425 and reference event # 5816.For further questions regarding this recall please call (949) 453-6400, |  | 7694 | 7,694 units total (5,935 units in US) | Worldwide Distribution - USA (nationwide) including Puerto Rico and to the countries of : Australia, China, Hong Kong, India, Indonesia, Korea (South), Malaysia, New Zealand, Philippines, Singapore, Sri Lanka, Taiwan, Thailand, Vietnam, Canada, Belgium, Czech Republic, Egypt, France, Germany, Great Britain, Greece, Hungary, Ireland, Israel, Italy, Netherlands, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, Japan, Argentina, Bolivia, Brazil, Chile, Colombia, Ecuador, Guatemala, Mexico, Panama, Peru, Uruguay, and Venezuela. | Open, Classified |  |  | 7694 | Not_Computer
Z-1198-2017 | 76257 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | ArthrexÂż Shoulder Suspension Tower; Catalog Number: AR-1650-01 |  |  | January 03, 2017 | 2017 | Allen Medical Systems, Inc. | During use, there is a potential scenario which could allow the fastening screws attaching covers to come loose and potentially land in the sterile field causing a delay in surgery. | Device Design | The distributor was notified of the recall via email on January 3, 2017. The letter identified the affected product, problem and actions to be taken. The letter instructs the consignee to remove the product from service and follow instructions to return the device to Allen Medical for repair. |  | 112 | 112 devices | US state of FL | Open, Classified |  |  | 112 | Not_Computer
Z-1201-2017 | 76344 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Hercules 360 Universal Stabilizing Arm, top mount, 360-degree rotation, reusable stainless steel |  |  | August 20, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) is recalling certain lots of Hercules TM 360 Universal Stabilizing Arms because the two pins that hold the sleeve attached to the distal end of the device may come loose, causing the pins and sleeve to separate from the device. While the arm can still function without the pins and sleeve in place, there is potential for the loose components to fall into the surgical site. | Device Design | Terumo Cardiovascular Systems has contact all affected parties (user facilities and consignees) via a Customer Notification Letter that advises of the recall activity. |  | 260 | The total number of distributed devices stands at 260 | Terumo Cardiovascular has identified 16 domestic consignees and 30 international consignees . | Completed |  |  | 260 | Not_Computer
Z-1199-2017 | 76332 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 6-hole Pubic Symphysis plate (PN 70-0451). The pelvic bone plates for the pelvic ring are a series of plates with varying lengths that function as internal fixation devices for fractures, fusions and osteotomies of the pelvic ring. The plates are strategically pre-contoured where beneficial to the application, and they are secured to the bone with |  |  | January 19, 2017 | 2017 | Acumed LLC | The 6-hole Pubic Symphysis plate (PN 70-0451) is specified to be made of Commercially Pure Titanium (ASTM F67 - CP Grade 4). However, batch 383813 was manufactured out of Titanium Alloy (Ti-6Al- 4V ELI per ASTM F135  Ti alloy) and was distributed. | Mixed-up of materials/components | Acumed sent an Urgent Medical Device Recall letter dated January 16, 2017, to all affected customers on January 19, 2017.Consignees were instructed to follow special instructions and send survey to RecallNotification@acumed.net.Domestic customers with questions or need to return products to Acumed, please contact Acumed Agent Inventory via email AgentInventory@acumed.net to obtain a Return Material Authorization (RMA) number or contact your agent directly.International customers with questions or need to return products to Acumed, please contact Acumed Business Service via email at BusinessServices@acumed.net or via phone 877-627-9957, choose Option 4 (US and Canada), or call 001-503-627-9957, choose option 4 (outside of US and Canada) have to obtain a Return Material Authorization (RMA) number. |  | 15 | 15 units | Worldwide Distribution - Costa Rica, Malaysia, Mexico, South Africa, and Spain | Terminated | May 08, 2017 | 109 | 15 | Not_Computer
Z-1200-2017 | 76281 | No | LFZ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | VIDAS 3 software v. 1.1.4 |  |  | January 11, 2017 | 2017 | BioMerieux SA | During development of the VIDAS 3 software version 1.2, some anomalies have been identified and observed to be already present in the current software version VIDAS 3 version 1.1.4. available in the field. | Software in the Use Environment | Consignees were notified via letter on January 11th, 2017. |  | 1161 | 1161 units | U.S. | Open, Classified |  |  | 1161 | Computer
Z-1204-2017 | 75942 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CentraLinkÂż Data Management System Software Versions: v13x,v14x,v15x |  |  | November 29, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | There is a remote possibility CentraLink may download an order to the ADVIA Automation System without specifying the sample type. This can occur when an order is received from the LIS without a sample type, requiring that the sample type be set in CentraLink based on the sample type of the test in the order. | Device Design | Siemens Healthcare Diagnostics, sent an Urgent Medical Device Correction Notice dated November 29, 2016, to all affected customers within the USA, and and issued notices titled "Urgent Field Safety Notices" to their international customers. Effectiveness Check Forms were also sent to the affected customers. Siemens recommends the LIS specify the sample type in the work order. The CentraLink system is designed to operate in an environment in which the sample type is specified in the work order that is downloaded from the LIS. A Siemens Technical Support Representative will contact affected customers to review your CentraLink configuration. If your LIS does not include sample type information in the sample orders downloaded to CentraLink and Pre-Tests assigned by sample type are configured on the ADVIA Automation System your representative will update your configuration. Complete and return the Field Correction Effectiveness Check Form attached to the letter within 30 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative.For further questions, please call (914) 631-8000. |  | 3893 | 3,893 units | Worldwide Distribution - USA (nationwide) to the states of : AL, AR, AZ, CA, CO, CT FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, WA and WV; and to the country of Canada and others. | Open, Classified |  |  | 3893 | Not_Computer
Z-1203-2017 | 75616 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | The Schiller PB 1000, also known as: Merge Hemo Patient Data Module (PDM), V2 model, a major component of the Merge Hemo software. |  |  | May 31, 2015 | 2015 | Merge Healthcare, Inc. | Non-invasive blood pressure pumps (NIBP) voltage is not sufficient and causes the NIBP pumps to fail prior to reaching adequate initial pressure. | Nonconforming Material/Component | The recalling firm issued letters via email dated 3/28/2016 and 7/11/2016, or via certified mail if no email address was available. A third letter dated 1/23/2017 was also issued.. |  | 341 | 341 sites potentially have the affected product versions | Distribution was nationwide to medical facilities. There was also military and government distribution. | Open, Classified |  |  | 341 | Not_Computer
Z-1202-2017 | 76448 | No | MGB |  |  | Cardiovascular | PMA | StarClose SE Vascular Closure System, Part No. 14679-01, 14679-02. |  |  | February 10, 2017 | 2017 | Abbott Vascular | Abbott Vascular is recalling the StarClose SE Vascular Closure System because it may exhibit difficulty or failure to deploy the StarClose SE Clip. | Other | An urgent field safety notice will be sent to customers on 2/10/17 to inform them that Abbott Vascular has initiated a recall regarding specific lots of the StarClose SE Vascular Closure System. Customers are informed that product from the identified lots may exhibit difficulty or failure to deploy the StarClose SE Clip. Potential risks associated with this event include prolonged procedure times, use of another device or manual compression to achieve hemostasis. Customers are instructed of the actions to be taken and what Abbott Vascular is doing about the recall. Customers with any questions are instructed to contact their local Abbott Vascular Representative or Customer Service Department at (800) 227-9902. |  | 170008 | 170,008 units total (83,029 units in US) | US: Germany, Italy, Puerto Rico, Spain, Sweden, Switzerland, | Open, Classified |  |  | 170008 | Not_Computer
Z-1208-2017 | 75707 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Maquet Express Dry Seal Suction Chest Drain, Single Collection w/AC |  |  | December 23, 2016 | 2016 | Atrium Medical Corporation | Chest drains manual vent valve located at the top of affected chest drains could potentially not re-seal after being depressed. | Process control | Maquet issued a Field action correction notification sent to customers on 12/23/16 via Fed'x for delivery on Tuesday Dec.27, 2016.who have received affected product and describes the failure mode.Accounts to complete the enclosed Field Notification Reply Form to acknowledge receipt of the notification. Return the Field Notification Reply Form to the following e-mail address: drainsmvv.us@getinge.com or you may fax the form to 1-603-386-6590.Questions, contact your local Atrium/Maquet/Getinge chest drainage representative or Atrium Customer Service at 1-800-528-7486 (603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST).UPDATE:On 1/16/17, Maquet issued recall letters that identified correct lot numbers and products in the letter titles: UPDATE-UPDATED AFFECTED DEVICES. Two letters issued one to customers who received the Dec 23 letter and one to new customers who who did not receive the Dec. 23 letter. |  | 681 | 681 cases (4086 units) | Worldwide Distribution - US (nationwide) and Internationally to Canada and Mexico | Open, Classified |  |  | 3891 | Not_Computer
Z-1212-2017 | 76454 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV System, medical charged-particle radiation therapy system |  |  | January 25, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Siemens became aware of Incorrect values for the rotational collimator position displayed due to mechanical problems of the sensor slide. In addition, there are improper weld seam at the overhead suspension. | Device Design | Siemens mailed a Customer Safety Advisory Notice (CSAN) dated February 7, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that Siemens will contact customers to arrange a date for the installation of the software update. |  | 194 | 194 systems | United States distribution | Open, Classified |  |  | 194 | Computer
Z-1218-2017 | 76210 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GN200 Caiman HF generators, electrosurgical, cutting &amp; coagulation &amp; accessories |  |  | January 09, 2017 | 2017 | Aesculap Implant Systems LLC | Generators may have a faulty component which could impact the proper functioning of the device. | Nonconforming Material/Component | Aesculap mailed an Urgent Medical Device Recall Notification letter dated January 9, 2017, to all affected customers. The notification letter advised customers to discontinue use and quarantine the affected unit if found. Customers were asked to complete the enclosed Acknowledgement form, list the serial number in their possession and return via fax at the number that was provided in the notification. For further questions, please call (800) 258-1946. |  | 24 | 24 generators | USA (nationwide Distribution). | Terminated | June 14, 2017 | 156 | 24 | Not_Computer
Z-1216-2017 | 76293 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Enzyme 2 Calibrator is an in vitro diagnostic product for the calibration of Alanine Aminotransferase (ALTI and ALTI) and Aspartate Aminotransferase (AST) methods on the Dimension Vista System. For prescription use only. |  |  | May 05, 2016 | 2016 | Siemens Healthcare Diagnostics, Inc. | Complaints of positive shift of patient and QC results when the ALTI assay is calibrated with the Dimension Vista Enzyme 2 calibrator. Internal investigation has confirmed that patient samples and QC material demonstrate a shift of up to 10% in comparison to results obtained using earlier calibrator lots. | Under Investigation by firm | Siemens mailed an Urgent Medical Device Notification letter to all affected customers on May 10, 2016. The notice provided information on actions to be taken by the customer to work around the issue. For further questions, please call (800) 441-9250. |  | 13147 | 13,147 units | Distributed throughout the United States | Open, Classified |  |  | 13147 | Not_Computer
Z-1219-2017 | 76178 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard Total Knee, PUNCH THRU TRL PLATES, 63 MM Item No: 32-487261, 67 MM Item No: 32-487262, 71 MM Item No: 32-487263, 75 MM Item No: 32-487264, and 79 MM Item No: 32-487265 |  |  | December 15, 2016 | 2016 | Zimmer Biomet, Inc. | Zimmer Biomet is conducting a medical device recall for ARCOS &amp; TPRLC broaches and rasps due to potential alumina inclusions in the raw material batch used to produce the affected products. Inclusions contained within the finished product could lead to the cracking and separation of the instrument. | Material/Component Contamination | Zimmer Biomet is conducting a medical device recall for ARCOS &amp; TPRLC broaches and rasps due to potential alumina inclusions in the raw material batch used to produce the affected products. Inclusions contained within the finished product could lead to the cracking and separation of the instrument |  | 26 | 26 | Domestic: 0Foreign: Singapore &amp; Malaysia | Open, Classified |  |  | 26 | Not_Computer
Z-1220-2017 | 76375 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 1.5mm System, Cranio-maxillofacial Bone Screws 1.5 x 4mm HT SD X-DR Screw, Bone Screw |  |  | January 10, 2017 | 2017 | Zimmer Biomet, Inc. | The identification of the devices was non-conforming with off-centered slots and the no-go gage starting in the pilot hole. | Under Investigation by firm | Zimmer Biomet initiated a voluntary recall of 1.5mm System due to identification of non-conformance with off-centered slots and the no-go gage starting in the pilot hole. One distributor with affected product outside of Zimmer Biomet control was notified via email and asked to return the affected product. This communication was sent on via e-mail onJanuary 10, 2017. All product has been returned.There were 89 units in Zimmer Biomet control at the time the removal was initiated, however all 100 are now in Zimmer Biomet control. Customer ServiceZimmer Biomet, Inc.P.O. Box 708Warsaw, IN 46581-0708Phone: 800-348-2759 |  | 11 | 11 | Domestic: NoneVA/DOD: NoneForeign: Japan | Open, Classified |  |  | 11 | Not_Computer
Z-1680-2017 | 76036 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | JResultNet drivers using ASTM coding language. |  |  | October 23, 2014 | 2014 | Data Innovations, LLC | Software issue resulting in the results of one patient potentially being associated with another. | Software design | Data Innovations, LLC contacted consignees by phone, email, and post. Recall notifications were sent out on November 3, 2014. Consignees were provided instructions on how to reconfigure the drive to work around the issue. If consignees were unable to perform the work around, Customer Support is able to service the units. In order to request this service, consignees were to call 239-931-6004, 802-658-1955, or 800-332-0499. Attached to the recall notification was a response form that is to be faxed to 802-658-2782. Consignees with and without affected drives are to complete this response form. For questions regarding this recall call 239-931-6004. |  | 657 | 657 units | Nationwide Distribution to AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; DC. | Completed |  |  | 657 | Computer
Z-1221-2017 | 76138 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium PVC 8 Fr Straight Thoracic Catheter |  |  | January 20, 2017 | 2017 | Atrium Medical Corporation | Labeling icon depicts a catheter having six (6) eyelets; however the correct number of eyelets on the 8 Fr and 10 Fr thoracic catheters is two (2) | Labeling design | Marquet/Atrium notified consignees by letter dated 01/20/17 informing of the Labeling Icon Clarification. The thoracic catheter labeling is being updated for future production. Users should be made aware of this Field Notification and may continue to use the affected thoracic catheter products with the current labeling.Questions, please contact your local Maquet/Atrium thoracic catheter representative or Maquet Customer Service at 1-800-528-7486 or 603-880-1433, Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 660 | 660 cases (10/case) | Nationwide | Open, Classified |  |  | 2054 | Not_Computer
Z-1225-2017 | 75727 | No | JCM | Abnormal hemoglobin assay. | Hematology | Hematology | 510(K) Exempt | Tosoh Bioscience HLC-723G8 Analyzer for HbA1c testing. |  |  | November 07, 2016 | 2016 | Tosoh Bioscience, Inc. | The device defect associated with this recall may result in HbA1c values with a negative bias in patients with the hemoglobin variants C, D, E and S. The negative bias occurs at the higher levels of HbA1c, e.g. 9%, but not at the diagnostic medical decision point of 6.5%. | Software design | Recall notification letters were distributed to customers on 11/7/16. The letters instructed customers to review the chormatography to determine whether or not a variant peak was detected. If so, restest the sample using an alternative method. |  | 682 | 682 analyzers | Worldwide distribution. The products were shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, AND WY.The recalled products were shipped to the following countries: Brazil, Chile, Columbia, Guatemala, Costa Rica, Dominican Republic, Eduador, Nicaragua, Panama, Peru, Puerto Rico, Uruguay, | Open, Classified |  |  | 682 | Not_Computer
Z-1223-2017 | 76264 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMT Bridle - Nasal Tube Retaining System |  |  | January 19, 2017 | 2017 | Applied Medical Technology Inc | Expired expiration dates are listed on the pouches. | Employee error | Applied Medical Technology Inc. sent an Urgent Medical Device Recall letter dated January 19, 2017, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.ACTIONS TO BE TAKEN BY THE CUSTOMER OR USER:Applied Medical Technology requests that customers IMMEDIATELY TAKE THE ACTION and return all devices in inventory subject to this recall.Do not use any device subject to this recall.Complete and return to Applied Medical Technology the enclosed RECALL NOTIFICATION.For further questions please call (440) 717- 4000. |  | 110 | 110 | US Distribution to the states of : AZ, CA, CO, CT, FL, GA, ID, LA, ME, MO, MN, NC, OH, OK, OR, TN, TX, UT, WA, and WI., and to the United Kingdom. | Open, Classified |  |  | 110 | Not_Computer
Z-1227-2017 | 76251 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | N Latex HCY REAGENT, REF/Catalog OPAX03 /SMN 10445973 , IVD |  |  | January 11, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | The firm confirmed a reduced once-opened and on-board stability for N Latex Homocysteine (HCY) OPAX03 lot 802907632 that may result in erroneously reduced or elevated homocysteine values. The stability issue may lead to higher than expected lot-to-lot variation and an impaired product performance. | Nonconforming Material/Component | URGENT MEDICAL DEVICE RECALL Letter (PP17-005.A.US, dated 1/11/2016) was sent to the sole US consignee via FedEx on 1/11/2017 to notify the customer of the issue. All N Latex HCY OPAX03 lot 802907632 customers are being instructed to discard N Latex HCY OPAX03 lot 802907632. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the UMDR within 30 days. On 1/13/2017, an URGENT MEDICAL DEVICE RECALL letter (PP17-005.B.US, dated 1/13/2017) was sent to the sole US consignee with instructions to discontinue use of and discard the kit lot; complete &amp; return the Field Correction Effectiveness Check/Product Replacement Form attached to this letter within 30 days; Review your inventory of these products to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities. For questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 352 | US: 62 kits; Foreign: 290 kits | Worldwide Distribution - US Nationwide in the state of Colorado and Foreign distribution to the following countries: Austria, Brazil, Canary Islands, Chile, China, Czech Republic, France, Germany, Great Britain (UK), Italy, Kuwait, Latvia, Poland, Saudi Arabia, Spain, Switzerland, Turkey, United Arab Emirates. | Open, Classified |  |  | 352 | Not_Computer
Z-1929-2017 | 76488 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Olympus ENDOEYE HD II Video Telescopes |  |  | February 10, 2017 | 2017 | Olympus Corporation of the Americas | A damaged temperature sensor at the endoscope's tip can cause the endoscope's distal end to become abnormally hot. | Device Design | A letter to consignees was sent February 10, 2017. Olympus will be repairing the device at user sites by disabling the temperature sensor. |  | 863 | 863 units | Nationwide | Open, Classified |  |  | 863 | Not_Computer
Z-1226-2017 | 76329 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | ETEST XM256 (Cefuroxime), Foam: 506958-506918; SPB: 412304-412305 |  |  | January 18, 2017 | 2017 | BioMerieux SA | Potential performance issues. False Susceptible result instead of Intermediate and False Intermediate result with certain methods. | Under Investigation by firm | BioMerieux sent an Urgent Field Safety Notice electronically on 1/18/2017. The notice included instructions and recommendations for users of the product. The notice included an acknowledgment form to be returned. Customers with questions were instructed to contact their BioMerieux representative. For questions regarding this recall call 919-620-3396. |  | 7160 | 7,160 units | Worldwide Distribution - US including AZ, AR, CA, CT, DE, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MS, MO, NE, NH, NJ, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WV, WI, and Internationally to Algeria, United Arab Emirates, Austria, Australia, Brazil, Cambodia, Canada, Chile, Chine, Costa Rica, Croatia, Czech Republic, Ecuador, Estonia, France, Greece, Germany, Great Britain, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, Nigeria, Norway, Oman, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russia, San Marino, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, Uruguay, and Vietnam | Open, Classified |  |  | 7160 | Not_Computer
Z-1245-2017 | 76173 | No | PQQ | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Microbiology | 510(K) Exempt | Keyspan High-High Speed USB to Serial Adapter |  |  | December 22, 2016 | 2016 | Tosoh Smd Inc | Power outages causes reporting software to shutdown. | Software design | Tosoh Bioscience sent a Customer Recall Notification letter dated December 21, 2016 to customers. The letter identified the affected product, problem and action to be taken. For questions call Tosoh Technical Support at 800-248-6764. |  | 34 | 34 | US Nationwide Distribution in the states of CT, MA, ME, MN, MO, NC, NE, NH, PA, SD, TX, VA, &amp; WV. | Open, Classified |  |  | 34 | Computer
Z-1243-2017 | 76414 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew Dyonics Saw Blade, Narrow, Product Number 3705 |  |  | February 02, 2017 | 2017 | Smith &amp; Nephew, Inc. | The single use devices are provided sterile and do not have an expiration date on the label. | Labeling Change Control | Smith &amp; Nephew, Inc. initiated recall by letter on 2/2/17 advising of the problem, requesting inspection of inventory and return of product. Contact Smith &amp; Nephews Field Actions Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to return. |  | 3462 | 3462 | NationwideForeign:Brazil, Canada, Chile , China, Denmark, Hong Kong, India, Mexico, Peru, Singapore, Venezuela | Open, Classified |  |  | 23262 | Not_Computer
Z-1233-2017 | 76368 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Jackson-Pratt Channel Drain Silicone Round, Fully Fluted w/10FR Trocar. Sterile. Rx |  |  | January 31, 2017 | 2017 | Degania Silicone, Ltd. | Product sterility compromised due to breach of sterile barrier | Packaging process control | Degania Silicone Ltd. issued an Urgent Recall letter dated 1/31/17 and requested return of the product to Degania Silicone 14 Thurber Boulevard, Suite A, Smithfield, RI 02917, USA.. For questions regarding this recall call +1-401-349-5373. |  | 3040 | 3040 units | To PR only | Terminated | April 04, 2017 | 63 | 3040 | Not_Computer
Z-1238-2017 | 76362 | No | EGJ | Iontophoresis device. | Physical Medicine | Physical Medicine | 510(k) | Daavlin Aquex (DAAV1000) |  |  | January 30, 2017 | 2017 | Hidrex GmbH | The device was reclassified from a class III device to a class II device and there are new documentation requirements. | Error in labeling | Distributor sent recall notification letter to affected dealers and end users on 2/10/17. Notification letter will include a copy of the addendum to the user manual. The device is not instructed to be returned. |  | 250 | 250 units | Nationwide throughout the US | Terminated | August 04, 2017 | 186 | 494 | Not_Computer
Z-1232-2017 | 76346 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Titan" Stabilizer Attachment and Titan TM 360 Stabilizer |  |  | January 05, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) is voluntarily recalling certain lots of Titan" Stabilizers and Titan" 360 Stabilizers because the Suction Tubing Clip may break when itis being attached to a Terumo CVS Hercules Stabilizing Arm. Although the clip is for convenience purposes and is not required for device operation, there is a potential for broken Suction Tubing Clip fragments to fall into the surgical site. | Other | A recall notification is being forwarded to all users and consignees who have received product. The notification of this voluntary recall will be sent to the following consignee in each facility:+Director of Operating Room ServicesProposed Level of Effectiveness Checks:Terumo CVS will verify that 100% of all letters are delivered to affected consignees |  | 1463 | 1463 | The product was distributed domestically and also to foreign consignees. But not to government agency's or Canada. | Completed |  |  | 1463 | Not_Computer
Z-1235-2017 | 76386 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | The Sarns Centrifugal Pump is a single use device sterilized by ethylene oxide gas. The device is sold single sterile in cartons of 8 units, or may be provided in bulk, non-sterile to another Terumo location for inclusion in Cardiovascular Procedure Kits. Cardiovascular Procedure kits are packaged individually. |  |  | May 27, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems Corporation has received complaints indicating that the Sarns Centrifugal Pump (disposable pump heads) have exhibited leaks. The leaks were detected during priming of the bypass circuit and during CPB procedures.In each of the reported incidents, there were no reports of patient injury as a result of the centrifugal pump leaks.Terumo Cardiovascular System's investigation into the incident has identified the cause of the leaks to be cracks in both the Top Housing and Magnet Back Housing of the pump. When the cracks cause the housing to become compromised, it is possible for fluid to leak from the pump head. | Process control | Customer Notification Letter/Response Form is being sent to all affected consignees. |  | 579 | 579 units have entered into the distribution chain | Terumo CVS has identified 33 domestic consignees and 0 international consignee that are affected by this recall | Terminated | February 23, 2017 | 638 | 579 | Not_Computer
Z-1229-2017 | 76274 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Efficia CMS200 Central Monitoring System; 863352 |  |  | January 17, 2017 | 2017 | Philips Electronics North America Corporation | The monitor may not alarm appropriately for a pediatric or neonatal patient. | Unknown/Undetermined by firm | Philips sent each known affected customer an Urgent Medical Device Correction Notification/Field Safety Notice dated December 2016 to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of providing a replacement MRx unit, free of charge, for all units affected by this issue. When the replacement unit arrives, please use the return shipping label and packaging provided with the replacement MRx unit to return the affected unit to Philips. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Option A:For any neonatal or pediatric patient, if the bedside monitor (without CO2) is connected to CMS200, do not place the bedside monitor in Standby when the patient is transported to a procedure. This will allow the alarms limits to work properly. Contact your service provider to determine the settings of your system and change if necessary.Option B:For any neonatal or pediatric patient, if the bedside monitor (without CO2) connected to CMS200 is placed in Standby when the patient is transported to a procedure. When the same patient returns to unit and is placed back on monitor, take the following steps:If clinician need disable or change alarm limit at CMS200 side, he/she need to discharge this patient at CMS200 side first, and admit same patient ID again. (Reference: Instruction for Use, Document Number 4535 645 61281, Page 89 of 199, section: Discharging a patient) This will allow the alarms limits to work properly. Contact your service provider to determine the settings of your system and change if necessary.For questions regarding this recall call 978-659-3000. |  | 156 | 156 devices, there is no US distribution | Worldwide distribution including AustriaBangladeshBulgariaChileFranceGermanyHong KongIndiaIrelandIsraelItalyKorea, Republic ofMalaysiaMyanmarPhilippinesPolandPortugalSaudi ArabiaThailandTurkeyUnited Arab EmiratesUnited KingdomUnited Republic of Tanzania | Open, Classified |  |  | 156 | Not_Computer
Z-1228-2017 | 76237 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Vascular catheter introduction kit |  |  | January 11, 2017 | 2017 | Arrow International Inc | There have been complaints for peel away sheaths flaring | Device Design | Arrow International mailed an Urgent Medical Device Advisory Notification letter to affected customers on 01/13/2017 to inform them of the issue. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990. |  | 215703 | 215,703 units | Nationwide Distribution | Open, Classified |  |  | 215703 | Not_Computer
Z-1171-2017 | 76179 | No | NRA |  |  | Orthopedic | PMA | Oxford Partial Knee System Phase 3 Tibial Template Right Medial Size D |  |  | January 05, 2016 | 2016 | Biomet U.K., Ltd. | Zimmer Biomet is conducting a medical device field action for various lot-specific Oxford Knee instruments due to potential alumina inclusions in the raw material batch used to produce the affected products. Inclusions contained within the finished product could lead to the cracking and separation of the instrument. | Nonconforming Material/Component | Zimmer Biomet initiated a voluntary recall of the Oxford Knee instruments due to potential alumina inclusions in the raw material batch used to produce the ffected products by mailing letters via FedEx on 01/05/2017. Customers were instructed to Review this notification and ensure affected team members are aware of the contents.2. Immediately locate and quarantine affected product in their inventory.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form.a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form and Attachment 2  Certificate of Decontamination to Zimmer Biomet.a. Request a Return Authorization Number via email to rgarequest@zimmerbiomet.com or through FAST/SMS. Be sure to specify RECALL as the RGA type.b. For each return, send a copy of Attachment 1 and Attachment 3 to corporatequality.postmarket@zimmerbiomet.com.c. Include a hardcopy of Attachment 1 with your shipment for immediate processing.d. Include a hardcopy of Attachment 2 with returned instruments. Mark the outside of the returned boxes clearly with RECALL. 5. Using the Additional Accounts Form provided with the email notice sent to your facility, return contact information for any additional hospitals that may have received or used the affected product. If there are no additional users to notify, please indicate none or NA on the form and return.6. Retain a copy of your recall acknowledgement and product return forms for your records in the event of a compliance audit of your facility.7. If after reviewing the recall notice you have further questions or concerns please call the customer call center at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outs |  | 72 | 72 units in total | Worldwide Distribution in the states of LA, NC, AR, MT, MO, CA, NY, KS, TX, MA, KY, OK, SC, NJ and the countries of Foreign: THAILAND, JAPAN, NETHERLANDS | Open, Classified |  |  | 72 | Not_Computer
Z-1234-2017 | 76276 | No | FCH | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | CMS-2193-R7 and CMS-1911-R3 Custom Biopsy Tray containing a BD 10cc LL Syringe, Sterile, Manufactured for CMS Custom Medical Specialties, Inc. |  |  | September 20, 2016 | 2016 | Custom Medical Specialties, Inc. | Sterility of syringe product could not be assured. | Nonconforming Material/Component | A letter was issued to customers on 9/20/2016 notifying them of the recall of the 10 cc syringe by BD. The letter tells them not to use the syringe in the kit, and to use the syringe on the outside of the kit which is not part of the recall. |  | 80 | 80 syringes | US Distribution to ID and OR. | Open, Classified |  |  | 80 | Not_Computer
Z-1231-2017 | 76337 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Influenza Transport Swab Packages, Reference #20171, |  |  | January 25, 2017 | 2017 | Quidel Corporation | Quidel Corporation is recalling Influenza Transport Swab Packages because it may include an incorrect sample swab mixed in with the proper swabs. | Mixed-up of materials/components | Quidel sent an Urgent Medical Device Recall letter dated 1/25/17 to all affected customers and distributors to inform them that Quidel is recalling the Influenza Transport Swab Packages because they may contain incorrect sample swab mixed in with the proper swabs. The letter informs the customers that the standard influenza swab is packaged individually, one (1) swab per package and uses a polyurethane foam tip. The incorrect swab is packaged as two (2) swabs per individual package and uses a rayon fiber tip. Customers are instructed to complete the enclosed Inventory Assessment/Certificate of Destruction form and return to Quidel. Contact Quidel Technical Support at technicalsupport@quidel.com or by calling 800.874.1517 (in the U.S.), or 858.552.1100 (outside the U.S.), for instructions on how to obtainreplacement product. Our hours of operation are Monday through Friday, 7:00 a.m. to 5:00 p.m. Pacific Time. |  | 1010 | 1,010 packages | Nationwide Distribution to FL, GA, IL, IN, KY, MT, OH, PA, TX | Terminated | June 02, 2017 | 128 | 1010 | Not_Computer
Z-1230-2017 | 76352 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | DIAMONDBACK 360 Peripheral 1.50 Classic, Model DBP-150CLASS145 |  |  | December 07, 2016 | 2016 | Cardiovascular Systems Inc | CSI discovered that a 1.50mm Solid Crown OAD was labeled as a 1.50mm Crown OED. | Employee error | Cardiovascular Systems, Inc. sent an Urgent Medical Device Recall letter, dated 12/07/2016, to all affected customers. The letter identified affected product , stated the reason for recall, asked for product to be returned to CSI. Questions can be directed to Customer Service at 877-274-0901. |  | 8 | 8 units | Nationwide Distribution to LA, MS, SC, TX | Terminated | March 23, 2017 | 106 | 8 | Not_Computer
Z-1248-2017 | 76359 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | SRHS DeRoyal Heart Tracecart REF 53-1913, Rx Only, NON-STERILE PRODUCT CONTAINED WITHIN. |  |  | January 25, 2017 | 2017 | DeRoyal Industries Inc | DeRoyal manufactured custom tracecarts containing a sterile Maquet Getinge Group chest drainage product which was subsequently recalled due to the current packaging configuration of the Atrium Ocean, Oasis, Express and Express Mini chest drainage products. | Other | DeRoyal issued their recall on 01/25/17 to the end-user level by sending overnight letters to its direct accounts and hospitals. Distributors and hospitals were instructed to return the tracecarts to DeRoyal. |  | 1157 | 1157 units | MS &amp; LA | Open, Classified |  |  | 2006 | Not_Computer
Z-1246-2017 | 75914 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Station Import Utility (ESIU) when used with Merge Eye Station and Merge Eye Care PACS. The firm name on the labeling is Merge Healthcare, Hartland, WI. |  |  | March 28, 2016 | 2016 | Merge Healthcare, Inc. | Eye Station images were not importing properly and were imported under "unknown" due to an issue when validating patients using only an Medical Record Number (MRN) | Device Design | The recalling firm issued a letter dated 3/28/2016 via e-mail. If no email address was available, the letters were issued via certified mail. A follow-up letter dated 7/11/2016 was also issued. |  | 1627 | 1,627 sites potentially have the affected versions | Distribution was nationwide to medical facilities. There was also government and foreign distribution, including Canada. | Open, Classified |  |  | 1627 | Not_Computer
Z-1249-2017 | 75829 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Durex Pleasure Pack . |  |  | November 22, 2016 | 2016 | Reckitt Benckiser LLC | Reckitt Benckiser LLC is recalling Durex Condoms Pleasure Packs due to missing information on the package labeling associated with the active ingredient Benzocaine. | Error in labeling | Reckitt Benckiser LLC sent an Urgent Medical Device Recall letter/business response form dated November 22, 2016 to the affected customers. Stericylcle is handling returns. |  | 108342 | 108342 units | Nationwide and Foreign | Open, Classified |  |  | 108342 | Not_Computer
Z-1537-2017 | 76485 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XPT System Software Versions V1.0.1, V1.0.2, V1.0.3, V1.1 and V1.2. The ADVIA Centaur XPT System is a stand-alone, continuous operation, immunochemistry analyzer that performs the following functions: aspirates and dispenses samples; Performs dilutions; Adds reagents; Incubates reaction vessels; Separates solid and liquid wastes; Measures photon emissions; Performs data reduction; Collects and maintains patient demographics and results. |  |  | February 02, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | Multiple software issues, which may affect the operation and workflow of the system. Potential for an apparent delay to testing when these issues occur. The software issues affecting the analyzer may impact all analytes available on the ADVIA Centaur XPT System test menu. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter ( UMDC) dated February 2, 2017. An Urgent Field Safety Notice letter ( UFSN) dated February 2017, was also sent to foreign customers. Customers must follow instructions in the UMDC and UFSN letters. For questions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative. For further questions, please call (312) 275-7795. |  | 729 | US: 17 units; Foreign: 712 units | Worldwide distribution - US Distribution to the states of : NH, MD, MS, MN, IL, TX, CA and WA, and to the countries of : Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Egypt, Finland, France, Georgia, Germany, Greece, Guadeloupe, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Jordan, Latvia, Malaysia, Myanmar, Netherlands, New Zealand, Norway, P.R. China, Peru, Philippines, Poland, Portugal, Qatar, Republic of Korea, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, U.A.E., United Kingdom, Uruguay, Vatican City, and Vietnam. | Open, Classified |  |  | 729 | Computer
Z-1250-2017 | 76363 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | GrebSet Micro-Introducer Kit, model 7941, Sterilized with Ethylene Oxide. |  |  | January 12, 2017 | 2017 | Vascular Solutions, Inc. | Four Lots of GrebSet Micro-introducer Kits, model 7941, contain guidewires that have a shelf life shorter than the kit expiration date. | Error in labeling | Consignees were sent on 1/12/2017 a Vascular Solutions "Urgent Medical Device Recall" letter dated January 11th, 2017. The letter identified the problem and product involved in the recall. The letter provided "Your Immediate Action Is Required". This included to remove the product from inventory and to complete and return the VSI's Account Inventory Form to VSIs Customer Service Dept. |  | 633 | 633 | US Nationwide Distribution in the states of NH, ME, VT, NY, LA, MS, AL CA, IA, WA, OR, MD, DE, PA, NJ, ID, TX, OK, MO, WI, GA, FL, AR, OH, AZ. | Open, Classified |  |  | 633 | Not_Computer
Z-1253-2017 | 76343 | No | NHX | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | EliA Sample Diluent, Article number, 83-1023-01 |  |  | January 30, 2017 | 2017 | Phadia US Inc | Erroneous EliA test results related to problems with liquid level detection and diluent pipetting in Phadia 250 instruments, and deformed bottles. Investigation showed problems with the liquid level detection in diluent bottle batch JJCA resulting in aspiration of no or too low volumes of liquid from the diluent bottle and from the dilution plate. | Under Investigation by firm | Phadia AB, part of Thermo Fisher Scientific, is voluntarily recalling EliA Sample Diluent, Article number, 83-1023-01, lots JM3Y, JS0B, JP87, JK5N, and JU9C, written notices were mailed via FedEx on 01/30/2017. There have been reports of instrument malfunctions for Phadia 250 due to deformed bottles for the below mentioned product and lots. The deformation of the bottle can cause an erroneous volume detection by the instrument which will result in insufficient aspiration of sample diluent, thus generating insufficient dilution of patient samples. The deformation is described as a flange or brim, located on the inside lower part of the vial opening. ACTIONS TO BE TAKEN BY THE CUSTOMER/USER:Âż Stop using the affected lots of the above mentioned product lotsÂż Please, return or attest to scrapping of any unused product lots to the manufacturer and order replacement products free of charge.Âż Review your records to determine if you have had instrument problems due to deformed bottles. If you find the instrument error message 3-145 RIGHT ARM LIQUID DETECTION BELOW LOWER LIMIT Target 2 , please collect all log files available and/or contact customer support, who will further assist in assessing possible impact on the test results in scope.Âż Fill in the Medical Device Recall return response on page 3 and return to the manufacturer by e-mail. If customers have any questions, please contact Phadia US, Inc. at 1-800-346- 4364, option #2. |  | 5957 | 5957 | Domestic: AR, CA, CO, FL, GA, IL, IN, KY, LA, MI, MN, MO, MT, NC, NH, NJ, NM, NY, OH, OR, PA, SD, TN, TX, WI, WVForeign: NoneVA/DOD: None | Terminated | August 07, 2017 | 189 | 5957 | Not_Computer
Z-1252-2017 | 75993 | No | NWB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Uretero-reno videoscope URF-V2 and URF-V2R endoscope and accessories |  |  | December 12, 2016 | 2016 | Olympus Corporation of the Americas | Olympus has received complaints regarding the breakage of the endoscope's insertion tube bending section during surgical procedures. Some of these complaints are associated with tissue trauma, including perforation, and insertion tubes which were stuck inside the patient and had to be surgically removed. | Under Investigation by firm | The firm, Olympus, sent an "Urgent Medical Device Safety Notice" on December 29, 2016, to those that were affected by this issue. The letter addressed the corrective actions Olympus were taking to address the issue. Consignees were asked to inspect their inventory for the specified device, review the enclosed Instructions for Safe Use, which provided instructions to assist with understanding the Warnings and Cautions and complete and return the OLYMPUS URGENT MEDICAL DEVICE SAFETY NOTIFICATION via fax to: Olympus Regulatory Affairs Department at 484-896-7128. |  | 1627 | 1,627 units (769 URF-V2 and 858 URF-V2R) | Worldwide distribution-US Nationwide and country of: Canada. | Open, Classified |  |  | 1627 | Not_Computer
Z-1251-2017 | 75990 | No | FBN | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | URF-P6 Uretero-reno fiberscope and URF-P6R Uretero-reno fiberscope |  |  | December 12, 2016 | 2016 | Olympus Corporation of the Americas | Olympus is recalling the URF-P6/P6R endoscopes due to a similar device (URF-V2/V2R endoscopes) but not identical, where adverse event complaints associated with tissue trauma, perforation and insertion tubes which were stuck inside of patients that had to be surgically removed. In an effort to mitigate a potential risk to patient health, Olympus is undertaking this action to notify users of these complaints and the need for careful inspection of the endoscope prior to use in accordance with the instructions provided int heir notification letter. | Under Investigation by firm | Olympus sent an Urgent Medical Device Safety Notice dated December 29, 2016, to those that were affected by this issue. The letter addressed the corrective actions Olympus that were taken to address the issue. Consignees were asked to inspect their inventory for the specified device, review the enclosed Instructions for Safe Use, which provided instructions to assist with understanding the Warnings and Cautions.You can contact our Technical Assistance Center (TAC) at 1-(800) 848-9024, option 1, to answer any questions. |  | 3461 | 3461 units distributed to US consignees | Distributed Nationwide | Open, Classified |  |  | 3461 | Not_Computer
Z-1255-2017 | 76285 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA Centaur Systems Calibrator A Kits, |  |  | January 10, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc | Siemens Healthcare Diagnostics has confirmed a negative bias for ADVIA Centaur FT4 when used with Calibrator A kit lots ending in 90 on the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur XPT Systems. In addition, Siemens Healthcare Diagnostics confirmed the potential for calibration failures due to above limit calibrator RLU %CVs when using Calibrator A kit lots ending in 90 with the FT4 assay.The performance of the FT4 assay when used with Calibrator A kit lots ending in 90 on the ADVIA Centaur CP System is not affected.Customers may continue to use Calibrator A kit lots ending in 90 for the ADVIA Centaur Systems FT3, T3, T4 and TUp assays. | Unknown/Undetermined by firm | Siemens issued Urgent Medical Device Recall (UMDR) sent to Siemens Healthcare Diagnostics customers in the United States on January 10, 2017 to be delivered to customers on January 11, 2017. An Urgent Field Safety Notice (UFSN) was sent out for distribution to customers outside the United States on January 10, 2017. The customer distribution included all customers thathave received ADVIA Centaur Systems Calibrator A kit lots ending in 90 and ADVIA Centaur FT4 ReadyPacks. These notices inform the customer a negative bias has been confirmed for the ADVIA Centaur FT4 assay when used with Calibrator A kit lots ending in 90 on the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur XPT Systems. In addition, Siemens Healthcare Diagnostics confirmed the potential for calibration failures due to above limit calibrator RLU %CVs when using Calibrator A kit lots ending in 90 with the FT4 assay. ADVIA Centaur FT4 assay customers using the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur XPT systems were directed to discontinue the use of Calibrator A kit lots ending in 90 and instead use Calibrator A kit lots ending in 91 and above to calibrate the FT4 assay. |  | 1937 | 1937 kits | Nationwide.Afghanistan, Albania, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium,Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark,Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Georgia,Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq,Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Latvia, Lithuania, Luxembourg, Macedonia,Malaysia, Malta, Martinique, Mexico, Myanmar, Nepal, Netherlands, New Caledonia, NewZealand, Nicaragua, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal,Qatar, Republic of Korea, Romania, Russian Federation, Saudi Arabia, Senegal, Serbia,Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan,Thailand, Tunisia, Turkey, Turkmenistan, United Arab Emirates, United Kingdom, Uruguay,Vatican, Vietnam | Open, Classified |  |  | 11795 | Not_Computer
Z-1383-2017 | 76487 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PRODISC Orthopedic Manual Surgical Instrument, strut used in spinal surgery. |  |  | January 31, 2017 | 2017 | Synthes (USA) Products LLC | These items are incorrectly etched and labeled. The item etched and labeled as 10mm Strut 03.820.350 lot 5919005 is a 12 mm Strut. The item etched and labeled as 12 mm Strut 03.820.352 lot 5919006 is a 10 mm Strut. | Under Investigation by firm | DePuy Synthes mailed an Urgent Notice to customers affected by the issue on 02/13/2017. The reason for the recall was explained, risks associated with the issue were identified, and actions to be taken were addressed. Customers were asked to complete the enclosed Verification form and mail it back to obtain a Return Authorization Number for the product's return. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. For questions regarding this recall call 800-479-6329. |  | 2 | 2 units | US distribution to TX and LA | Open, Classified |  |  | 2 | Not_Computer
Z-1394-2017 | 76492 | No | MLA | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD AffirM VPIII Microbial Identification Test |  |  | August 27, 2015 | 2015 | Becton Dickinson &amp; Co. | BD has confirmed that a portion of tests associated with the affected lots( 446252 and 446257 ) of product may exhibit blue speckles or fibers on the beads and cards, as shown below. This issue would result in an invalid test, as described in the package insert. | Employee error | Becton Dickinson sent an Urgent Product Recall Letter dated August 2015. The distributors were provided with a written communication via email and customers were provided with a written communication via UPS. Customers were not asked to send product back.If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663.For further questions please call (410) 316-4428. |  | 4502 | Qty Distributed to Field - 4,502 | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WYand to the countries of : AU, BE, IN, KR, SG, | Terminated | March 06, 2017 | 557 | 4502 | Not_Computer
Z-1414-2017 | 76498 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EpiCenter Plus Microbiology Data Management System software version 6.20A, when used in conjunction with the MALDI BioTyper System . |  |  | August 25, 2014 | 2014 | Becton Dickinson &amp; Co. | The information in the EpiCenter interface translation tables was not updated to match the Bruker MALDI BioTyper database identifications for some organisms. This could lead to a possible mis-identification of an organism. | Vendor change control | Customers were initially contacted via phone calls and subsequently provided with a written communication. For questions contact BD Technical Services Department at 1-800-638-8663. |  | 115 | Qty Distr. to Field - 115 | 15 US customers, 18 Ex-US sites, and 0 distributors. Approximately 115 customers totally and 0 distributors | Terminated | March 08, 2017 | 926 | 115 | Not_Computer
Z-1469-2017 | 76568 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC FlexiView 8800, Digital Mobile Imaging System, GE Medical Systems |  |  | February 21, 2017 | 2017 | GE OEC Medical Systems, Inc | OEC FlexiView 8800 workstation power cable assembly issue which can result in a loss of imaging functionality due to intermittent power cord failures. | Process control | On February 21, 2017, an Urgent Medical Device Correction letter was sent to one consignee with an affected unit; a field service representative hand delivered the customer communication. |  | 1 | 1 | One consignee in CA. | Terminated | April 24, 2017 | 62 | 1 | Not_Computer
Z-1461-2017 | 76494 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | L-Varlock Cage/Trial Implant Holder, Reference HD0060, Kisco International, France. |  |  | January 11, 2017 | 2017 | Kiscomedica S.A. | Firm received a complaint of the tip holder breaking during use. | Process control | Manufacturer's Field Safety Corrective Action Report was submitted to ANSM (French Health Products Safety Agency) on January 11, 2017. The Field Safety Notice was sent to U.S. Consignee on January 18, 2017. The letter identified the affected product, problem and actions to be taken. |  | 13 | 13 units | Worldwide Distribution - US Nationwide in the states of MO and country of Switzerland. | Terminated |  |  | 13 | Not_Computer
Z-1405-2017 | 76351 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | A-series lift - A625 Manual w/Load Cell &amp; Wi-Fi, Omni Part No. 324124. Prism Medical. |  |  | January 30, 2017 | 2017 | ErgoSafe Products, LLC (DBA) Prism Medical | During internal cycle test at max load (625 lbs.) 2 units had straps break prematurely. | Device Design | The recalling firm called each customer on Jan 30, 2017 to review the recall and schedule a time for the firm to come to their facility to personally swap the recalled product for a new device. The FSNs were sent on Jan 31, 2017. |  | 21 | 21 | Product was distributed to MT and NV; one VA consignee. | Open, Classified |  |  | 194 | Not_Computer
Z-1382-2017 | 76356 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | OP7543 Design Y Mandible Onlay - Medium - Left and OP7544 Design Y Mandible Onlay - Medium - Right |  |  | January 25, 2017 | 2017 | Matrix Surgical Holdings, LLC | Mislabeling. The packaging labeled as left mandible contained a right mandible. | Other | Matrix Surgical sent an Urgent Product Recall e-mail dated January 26, 2017, to the foreign consignees. The distributors were instructed to return the product in inventory, and to conduct a sub-recall, if needed. For further questions, please call ( 404) 862-3794. |  | 28 | 28 | Worldwide Distribution to the countries of : Brazil, Chile, Colombia, Norway, Saudi Arabia, South Africa, Sweden and Taiwan | Terminated | May 24, 2017 | 119 | 28 | Not_Computer
Z-1464-2017 | 76439 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MaxGuard Extension Set with Injection Site and 0.2 micron filter, Model No. ME3305 |  |  | February 20, 2017 | 2017 | CareFusion 303, Inc. | CareFusion is recalling the MaxGuard Extension Set with Injection Site and 0.2 micron filter because of reports of separation and/or leakages between the Y-connector and tubing. | Process control | An Urgent Medical Device Recall Notification letter will be sent to customers on February 20, 2017 to inform them that CareFusion is recalling the MaxGuard Extension Set with Injection Site and 0.2 micron filtermodel number ME3305 Lot Number 16016790. Customers are informed that CareFusion has received reports of separation and/or leakages between theY-connector and tubing. Customers are informed that leakages from an extension set can cause delay of infusion, interruption of infusion, exposure to medication or hazardous infusates, or underinfusion. The letter informs the customers of the actions to be taken and for recall related questions to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266. |  | 1200 | 1,200 units | US Distribution including MA and MD | Open, Classified |  |  | 1200 | Not_Computer
Z-1318-2017 | 76280 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hystero Pack, part number PSS3535(A |  |  | January 11, 2017 | 2017 | Windstone Medical Packaging, Inc. | The Medtronic COVER LIGHT HANDLE DISPOSABLE GLOVES of the custom Aligned Medical Solutions (AMS) kits were recalled because gloves may contain splits or holes compromising sterility. | Nonconforming Material/Component | AMS notified their consignees by Email, Fax, and Phone on 11/11/2017. AMS sent the notification letter, dated January 11, 2017, and provided yellow stickers "Recall Notice" to customers. They are advised to over-label the affected AMS kits at their facilities. Customers were instructed to place sticker in such a way that it would be immediately noticed but should not cover up or conceal any pertinent information such as Pack Number, Pack Name, Lot Number or Expiration Date on a kit.Customers with questions can call Vicki Davis at 321-527-7714 or send email to vdavis@alignedmedicalsolutions.com. |  | 32 | 32 kits | Distributed in AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NE, NJ, NV, OH, PA, TX, UT, VA, and WA. | Terminated | May 03, 2017 | 112 | 42514 | Not_Computer
Z-1475-2017 | 76497 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SharkCore LG Fine Needle Biopsy System (19GA); Material/Part#: DSL-19-01 |  |  | February 08, 2017 | 2017 | Medtronic | Handle may break during use. | Nonconforming Material/Component | A recall was initiated on February 8, 2017. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 1856 Corporate Dr. Suite 140 Norcross, GA 30093, Attention: Field Returns Department. Disposition of product within Medtronic control and customer returned devices: Will be reworked. |  | 1096 | 1,096 Units (total for all products) | US Distribution to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI; and countries of DK and DE | Open, Classified |  |  | 1096 | Not_Computer
Z-1425-2017 | 76228 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph 16 TruePoint TrueV - 4R, Material Number 10249556 |  |  | February 20, 2017 | 2017 | Siemens Medical Solutions USA Inc. | The mobile-specific instructions was not shipped to a limited population of Biograph Mobile systems. | Labeling Change Control | Siemens sent a Customer Advisory Notice letter on February 20, 2017 to affected customer. The letter identified the affected product, problem and actions to be taken. For questions contact your local Siemens. |  | 4 | 4 units | Worldwide distribution - US Nationwide in the states of CA, CT, FL, GA, IL, KY, MA, MI, ND, NJ, OH, OK, OR, PA, TN, WA, WI, WY and the countries ofSwitzerland, Great Britain, Italy | Open, Classified |  |  | 74 | Not_Computer
Z-1482-2017 | 76440 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MaxGuard Extension Set, Model No. ME2064, MP9009C, MP9209, MP9254C |  |  | February 20, 2017 | 2017 | CareFusion 303, Inc. | CareFusion is recalling the MaxGuard Extension Set due to reports of leaks occurring with the 0.2 micron filter. | Under Investigation by firm | CareFusion sent an Urgent Medical Device Recall Notification letter dated February 2017 to all affected customers inform them that CareFusion is recalling the MaxGuard Extension Set model codes ME20164, MP9009-C, MP9209, MP9254-C with multiple lot numbers. CareFusion has received reports of leaks occurring with the 0.2 micron filter. Customers were informed that leakages on an extension can cause delay of infusion, interruption of infusion, exposure to medication or hazardous infusates, underinfusion or air in line. Customers were instructed to discontinue use of the affected product and contact CareFusion for a replacement lot. The letter informs the customers of the actions to be taken and for recall related questions to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266. Customers were also instructed to immediately complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card. |  | 12850 | 12,850 units | Nationwide Distribution | Open, Classified |  |  | 12850 | Not_Computer
Z-1384-2017 | 76374 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChemVELOCITY Urine Chemistry System, Catalog No. 800-3530, 800-7167, 800-7719, 800-7720, 800-3564, 800-7100, 800-7103, 800-7106, 800-7162, 800-7163, 800-7166, 800-3061. |  |  | January 30, 2017 | 2017 | Beckman Coulter Inc. | The probe misalignment or bending has the potential to lead to delay in reporting results or incorrect patient results. | Device Design | Beckman Coulter/Iris sent an Urgent Medical Device Recall letter dated January 30, 2017 to all affected customers to inform them that Iris International has determined that sample probe misalignment or bending on the iChemVELOCITY may lead to a remote possibility of false negative results due to inadequate strip dosing. Customers were informed that Inadequate dosing may lead to incorrect patient results which may not be detected by quality control, or to a delay in reporting results. Incorrect patient results might be observed as an unexpected discrepancy between instrument results and the patient's clinical picture. The greatest impact could occur when proteinuria is not detected. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers with questions were instructed to contact their local support representative:Âż Via our website, at http://www.beckmancoulter.comÂż By phone, call 800 854-3633 in the United States and Canada |  | 3175 | 3175 units total (1697 units in US) | Worldwide Distribution - US (nationwide) and internationally to Argentina, Australia, Austria, Bangladesh, Belarus, Bolivia, Bosnia and Herzegovina, Brazil, Brunei, Canada, Chile, China, Columbia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Macau, Malaysia, Mexico, Moldova, Montenegro, Netherlands, New Zealand, Nigeria, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Republic and Venezuela. | Terminated | August 19, 2017 | 201 | 3175 | Not_Computer
Z-1385-2017 | 76471 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS. |  |  | April 08, 2016 | 2016 | Merge Healthcare, Inc. | The software is not adding the correct tomo image. | Software design | Merge sent an Urgent Medical Device Recall dated July 11, 2016 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions send an email to recall@merge.com |  | 28 | 28 sites potentially have the affected software | US Nationwide Distribution in the states of AL CA, CO, FL, IL, MD, MN, MT, NV, NY, OR, PA SC, TN, and TX. | Terminated | March 06, 2017 | 332 | 28 | Computer
Z-1386-2017 | 76052 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Hybrid IABP and CARDIOSAVE Rescue IABP affected part number is 0998-:XX-0800-XX |  |  | December 23, 2016 | 2016 | Maquet Datascope Corp - Cardiac Assist Division | Maquet has received 3 confirmed complaints related to the Li-lon Battery used with the Maquet CARDIOSAVE Hybrid IABP and Maquet CARDIOSAVE Rescue IABP of the battery venting after being accidentally dropped. | Device Design | Maquet contacted their customers via letter on 12/23/16 and additional customers identified were notified on 1/10/17 and issued a recall letter/return response form. The letter identified the affected product problem and actions to be taken. Customers were instructed to complete and return the response form via fax to 1-973-629-1518 or email to cardiosavebattery@getinge.com |  | 8000 | 8000 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 8000 | Not_Computer
Z-1692-2017 | 76540 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | UOC Femoral Driver, Product number PE 93045101, orthopedic manual surgical instrument/accessory for use during surgery to implant femoral components and femoral trial, and to remove femoral trial. |  |  | February 10, 2017 | 2017 | United Orthopedic Corporation | The locking lever pin component of the femoral driver may become deformed and break, resulting in malfunction during surgery and causing the device to become inoperable. | Device Design | United Orthopedic Corporation sent an Urgent Field Safety Notice dated January 19, 2017, to all affected consignees via e-mail on February 10. 2017. The Urgent Field Safety Notice instructed consignees to recover any affected units and return them to the manufacturer. A redesigned device will be sent for replacements. Consignees with questions were instructed to call 949-328-3366. |  | 33 | 33 units | Worldwide Distribution - US including State of CA and Internationally to Pakistan, Spain, and Taiwan. | Terminated | June 28, 2017 | 138 | 33 | Not_Computer
Z-1396-2017 | 76500 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology 12.2 - Picture Archive Communication System (PACS) |  |  | December 02, 2016 | 2016 | Mckesson Medical Imaging | Issue for customers that use an EMR login or legacy web URL login or legacy web URL login for McKesson Radiology PACS that may result in missing images in a newly imported study, and/or study imports that remain in an "in-progress" status. | Software design | McKesson sent an Urgent Field Safety Notice on December 2, 2016 to customers with potentially affected software versions. The notice indicated that they were to contact McKesson to install a software update to correct the issue, and that they read and understood the Field Safety Notice. |  | 17 | 17 | AK, CA, FL, HI, KY, MD, MI, MS, NH, OH, PA, TN, TX, VT | Open, Classified |  |  | 17 | Computer
Z-1395-2017 | 76503 | pylon, post surgical | ISM | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | SONNET Mini Battery Pack CableProduct Usage: |  |  | October 25, 2016 | 2016 | MED-EL Elektromedizinische Gereate, Gmbh | Unit Defect: Firm inspection revealed units did not have the compartment lock firmly welded to the housing. | Device Design | MED-EL is notifying all consignees to whom affected products have been directly shipped by Headquarters (based on first delivery shipment records) on 10/25/2016. |  | 538 | 538 units | Worldwide Distribution - US Nationwide in the states of CA, NC and TX; and the countries of Belgium, Belarus, Canada, France, Germany, Hong Kong, Israel, Netherlands, Philippines, Poland, Portugal, Turkey, United Kingdom, and Uruguay. | Terminated | April 04, 2017 | 161 | 538 | Not_Computer
Z-1393-2017 | 76360 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | TriForce Peripheral Crossing Set, Individual product is packaged in a Tyvek-film sterilizable outer pouch. |  |  | February 01, 2017 | 2017 | Cook Inc. | Devices using a specific hemostatic valve design, referred to as a blue valve or polyisoprene valve, have an increase in reports of blood loss associated with the devices which could result in delay in procedure and blood loss from patients. | Under Investigation by firm | Cook Inc. is initiating a voluntary recall of multiple products following the identification of an increase in reports of blood loss associated with devices using a specific hemostatic valve design referred to as the blue valve or polyisoprene valve. Potential adverse events that may occur if these devices were used in the arterial system (Flexor Radial Access Sets) include delay in procedure and blood loss. If devices were used in the central venous system, adverse events that may occur include delay in procedure, blood loss, or air embolism. The firm is distributing notices via mail courier on 02/01/2017. Customers are being instructed to share the recall notice with appropriate personnel, including down to the user level, within the customer's organization or to any organization where the potentially affected devices have been transferred. 1.Examine inventory immediately to determine if customers have affected product and quarantine affected product(s).2.Return affected product(s) to Cook Medical with a copy of the Acknowledgement and Receipt Form to receive a product credit. NOTE: Returned un-affected product will not be credited.3.Even if customers do not have affected product on hand, customers must still complete the Acknowledgement and Receipt Form and return via fax 812.339.7316 or email to Fieldactionsna@cookmedical.com.4.Report adverse events to COOK Medical Customer Relations at 800.457.4500 or 812.339.2235, Monday through Friday between 7:30 am and 5:00 pm (Eastern Time), or by email to: customerrelationsna@cookmedical.com.Adverse events or quality problems experienced with the use of this product may also be reported to the FDA: Online at: http://www.fda.gov/Safety/MedWatch/HowToReport/default.htm (form available to fax or mail) Call the FDA at: 1-800-FDA-1088Effectiveness checks of this recall will be done by response form mailed. All recalled products will be destroyed at the regional distribution warehouses. A cer |  | 1875 | 1,875 | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Internationally to Canada, Armenia, Austria, Belgium, Bulgaria, Cyprus, Czech Republic, Denmark, Spain, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Kenya, Kuwait, Lithuania, Mauritius, Netherlands, Norway, Oman, Poland, Reunion, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Sweden, Switzerland, Turkey, United Arab Emirates, and United KingdomVA/DOD: Yes | Open, Classified |  |  | 47984 | Not_Computer
Z-1400-2017 | 76445 | No | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Remel, Haemophilus Test Medium (Agar) in A) 100 mm REF R01503, and B) 150 mm plates REF R04033, |  |  | February 09, 2017 | 2017 | Remel Inc | Test organisms exhibit poor growth when grown using the test agar. The poor growth can give users incorrect AST zones when performing susceptibility tests. | Nonconforming Material/Component | Thermo Fisher sent an Urgent Medical Device Recall letter dated February 9, 2017, to all affected customers. The recalling firm issued written notifications to their customers. Customers were instructed to destroy any remaining inventory of the recalled lots. Customers were also instructed to contact everyone within their organization or any other organization where the devices have been transferred. An acknowledgement form was included with the recall notice. Customers were asked to complete the acknowledgement form and return it to the recalling firm. If you have any questions, please contact our Technical Service Department at (800) 255-6730 ( US ) or (913) 888-0939 ( International)For further questions, please call (913) 895-4077. |  | 5749 | RO4033 - 163, R01503 - 50. | USA ( nationwide ) Distribution including Puerto Rico and to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SD, TN, TX, VA, VT, WA, WI, WV and WY. | Open, Classified |  |  | 5749 | Not_Computer
Z-1401-2017 | 76400 | No | MCX |  |  | Cardiovascular | PMA | AngioJetÂż Ultra 5000 A Console:Foot Switch |  |  | January 31, 2017 | 2017 | Boston Scientific Corporation | Boston Scientific has received field reports regarding the AngioJet Ultra 5000 A Console foot switch sticking during use resulting in an undesired continued operation of the device. | Under Investigation by firm | Field notification letters were sent on January 31, 2017 by certified mail to affected customers. Letters request that Users follow the operator's manual for service and maintenance and if they notice that the foot switch is sticking to discontinue use and contact BSC Technical support at 1 800-949-6708 or CETechSupportUSA@bsci.com. |  | 6814 | 3402 units; 2626 in US; 786 units OUS | US and Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Cyprus, Denmark, Dominican Rep., Ecuador, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Iraq, Ireland, Israel, Italy, Jordan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, Puerto Rico, Russian Fed. ,Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Tobago, Turkey, United Arab Emir., Venezuela, Virgin Islands, W.Bank/Gaza Strip | Open, Classified |  |  | 6814 | Not_Computer
Z-1407-2017 | 76560 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | MicroAire K-Wires |  |  | February 21, 2017 | 2017 | MicroAire Surgical Instruments, LLC | Mislabeling on three lots of K-Wires. The incorrect product description " dual trocar" instead of "single trocar" was on the product label on the following parts: 1600-9355NS (lot 61005), 1600-9625NS (lot 60297), and 1604-162NS (lot 62856). | Labeling mix-ups | The firm will mail recall letters to the customers who received the mislabeled k-wires on February 21, 2017. Letters will be sent by UPS to expedite the delivery. The firm will contact the non-responding accounts at least two more times. |  | 362 | Total 362 | United States. | Terminated | July 14, 2017 | 143 | 362 | Not_Computer
Z-1403-2017 | 76590 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | When taking measurements from images on the Cardio workstation or from the US cart, numbers are not crossing to the report in the correct unit of measure. | Software design | Merge sent an Urgent : Medical Device Recall letter dated March 28, 2016. If no email address was available, the letters were issued via certified mail. Nonresponders were issued a second letter dated July 11, 2016.For further questions, please call (877) 741-5369. |  | 110 | 110 sites potentially have the affected versions | Distribution was made to medical facilities nationwide. There was no foreign/military/government distribution. | Open, Classified |  |  | 110 | Not_Computer
Z-1257-2017 | 76252 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 1000 defibrillator is intended for use by personnel who are authorized by a physician/medical director and are trained in CPR and the use of the LIFEPAK 1000 defibrillator. |  |  | January 13, 2017 | 2017 | Physio-Control, Inc. | The firm has received complaints that the LIFEPAK 1000 Defibrillator is unexpectedly powering off during device usage. | Device Design | Physio Control sent an Urgent Medical Device Correction letter dated January 2017 to all affected customers. The firm issued the press release on January 13, 2017. The link to press release is http://www.fda.gov/Safety/Recalls/ucm537114.htm.The firm also started mailing letter, confirmation sheet, check list to consignees on January 13, 2017. Information about this notice is available at: www.physio-control.com/lifepak1000-274. Customers were instructed to immediately remove and reinstall the battery from their LIFEPAK 1000 defibrillator. The removal and reinstallation of the battery will clean the contacts of oxidation and will restore power to the device.It is critically important that customers implement a weekly schedule of battery removal and reinstallation for all LIFEPAK 1000 devices. Removing and reinstalling the battery on a weekly basis will help ensure their device is ready for use.Customers were also advised to always carry a spare, fully charged battery. The weekly schedule of battery removal must be performed until their device correction has been completed. Customers were asked to forward this information to all of their sites, trainers and users who have LIFEPAK 1000 devices.The Customer Letter was updated on the web page to clarify the statement "It is also important to always carry a spare, fully charged battery" was provided as a reminder to customers and is consistent with what is stated in the LIFEPAK 1000 Defibrillator Operating Instructions.Customers were asked to complete and return the Confirmation Sheet to Physio-Control.Customers with questions regarding this notification, please contact Physio-Control by calling 1-866-231-1220, 6:00 a.m. to 4:00 p.m. (Pacific) Monday  Friday, or by email to rsrecalls@physio-control.com or fax to 1-866-448-9567. |  | 133330 | total 133,330 units (50,046 units in the US) | Worldwide Distribution - US (nationwide) and in the following countries: Bahrain, Belize, Cambodia, Chile, China, CÂżte d'Ivoire, Ecuador, Germany, Honduras, Hong Kong, India, Indonesia, Jamaica, Myanmar, Netherlands, Niger, Philippines, Saudi Arabia, Senegal, Singapore, Sri Lanka, Thailand, Tunisia, Ukraine, and Zambia. | Open, Classified |  |  | 133330 | Not_Computer
Z-1404-2017 | 76279 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Color contact lenses labeled under the following brands: Diamond, Fierce, Ambition, Circle Collection, Lunatic, Fright, Desire and See Clear. |  |  | January 06, 2017 | 2017 | The See Clear Company | Sterility and Misbranding: Lenses may not be sterile and may be labeled with erroneous lot numbers and expiry dates. | Device Design | The firm, The See Clear Company, sent an "URGENT: DEVICE RECALL" letter dated 12/30/2016 to its Consignees on 01/06/2017. The letter described the product, problem and actions to be taken. The consignees were instructed to immediately examine your stock to determine if you have any of the listed product; immediately discontinue distributing the product and promptly return them via parcel post to Norcross facility, ATTENTION: RETURN GOODS; if you have further distributed any of the listed product, immediately contact your accounts, advise them of the recall, and have them return their outstanding recalled stocks to you; and immediately complete and return RECALL RETURN RESPONSE FORM via mail.If you have any questions, contact us at (678) 313-1399. |  | 119829 | 119,829 units | Worldwide Distribution: US (nationwide) including states of: AZ, CA, DE, FL, GA, HI, LA, NJ, NY, NV, ON, PA and WA; and country of: Canada. | Open, Classified |  |  | 119829 | Not_Computer
Z-1402-2017 | 76318 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviWipes1 Extra Large, Part No. 13-5150 |  |  | January 20, 2017 | 2017 | Metrex Research, LLC. | Metrex is recalling the CaviWipes Extra Large because they may have been contaminated during the packaging process. | Packaging process control | Metrex sent an Urgent Medical Device Recall letter dated January 20, 2017, to all affected customers to inform them that Metrex is recalling one lot of CaviWipes Extra Large because they may have been contaminated during the packaging process. Customers were informed that while handling sealed containers, a production worker sustained a minor finger cut which may have come in contact with the exterior surface of the container. Customers were instructed to contact Metrex Customer care at (800) 537-7123 to receive a RMA number. The RMA will allow for a quick return and replacement or credit. Customers were instructed to identify any customers that may have been shipped the affected product and contact their customers using the customer notification letter that Metrex issued. Metrex requests that those customers contact Metrex Customer Care at (800) 537-7123 and complete the attached acknowledgement form.For questions regarding this recall call 734-947-6700. |  | 960 | 960 cases | Nationwide Distribution to PA, FL, IL, MN, TX, AL, NY | Open, Classified |  |  | 960 | Not_Computer
Z-1399-2017 | 76299 | Yes | CBI | Tracheal/bronchial differential ventilation tube. | Anesthesiology | Anesthesiology | 510(k) | Parker Left Double Lumen Endobronchial Tube with Parker Flex-Tip (Parker Endo-Bronch with Flex Tip). Model Numbers: H-PFTDL-28-1, H-PFTDL-32-1, H-PFTDL-35-1, H-PFTDL-37-1, H-PFTDL-39-1, |  |  | December 16, 2016 | 2016 | Parker Medical | The product was not functioning as intended. The ventilation to the patient was not disrupted, but the single lung isolation and ventilation did not function as intended. | Under Investigation by firm | Parker Medical notified consignees via telephone on December 16, 2016. Consignees were instructed to segregate affected products for immediate pick up. For questions regarding this recall call 239-307-6064. |  | 1060 | 1,060 units | Nationwide Distribution to CO, FL, IN, SC and TX. One VA consignee. | Terminated | April 17, 2017 | 122 | 1060 | Not_Computer
Z-1408-2017 | 76327 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Premier Guard Sterile Drape For Use with Blair-Port Wand (02-0029) and Blair-Port Wand X (01-0046) |  |  | January 23, 2017 | 2017 | Covidien LLC | Sterility compromised due to breach in sterile barrier | Packaging process control | Medtronic issued recall letter on January 23, 2017. Customers were notified via Federal Express and the letter instructed customers to discontinue use of the product and return all inventory to Medtronic (Covidien)  195 McDermott Road, North Haven, CT 06473 Attn: Field Returns Department. |  | 1111600 | 1,111,600 units | Nationwide | Open, Classified |  |  | 1154350 | Not_Computer
Z-1410-2017 | 76408 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The UNI-CP System; Model Number: 330230SND. |  |  | January 25, 2017 | 2017 | Integra LifeSciences Corp. | The firm received one complaint from one non-US sales representative (France) that the label on the UNI-CP plate was incorrect. | Device Design | Integra LifeSciences Corp. issued a recall letter/return response form on 1/25/2017. |  | 39 | 39 units | Nationwide | Open, Classified |  |  | 39 | Not_Computer
Z-1485-2017 | 76571 | bag, urine collection, leg, for external use, non-sterile | NNW | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | PureWick External Catheter for Women-Wick |  |  | February 24, 2016 | 2016 | Purewick Corporation | PureWick is recalling external catheter for women Wicks because the labeling did not contain the phrase "This product contains dry natural rubber" and/or erroneously included the phrase "Latex-free." The white gauze component of the PureWick contains dry natural rubber. | Labeling design | A Urgent Medical Device recall letter dated March 2, 2017 will be sent to customers to inform them that PureWick Corporation is recalling specific lots of PureWick external catheter for women Wicks, Item #625. Customers are informed that PureWick is recalling PureWick Wicks because labeling on product shipped prior to February 22, 2017 either did not contain the phrase This product contains dry natural rubber and/or erroneously included the phrase Latex-free. The white gauze component of the PureWick Wick contains dry natural rubber. PureWick Corporation has received zero complaints or reports of adverse events related to the error in labeling. Customers are informed of the risk to health and the actions to be taken. Customers are instructed to complete the enclosed PureWick Field Correction form and fax to PureWick Corporation at (619) 660-5459 or email to response@purewick.com. Customers with questions or would like assistance completing the PureWick Field Correction Form are instructed to call PureWick toll-free at (844) 584-0734 Monday  Friday from 8:00 a.m. to 4:30 p.m., Pacific Time. |  | 176863 | 176,863 units | US: Nationwide Distribution | Terminated | May 09, 2017 | 440 | 176863 | Not_Computer
Z-1411-2017 | 76158 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | MAKO RIO THA Application User Guides, for total hip arthroplasty, orthopedic surgery. |  |  | December 22, 2016 | 2016 | Mako Surgical Corporation | Five (5) AFMEA Risk Control Measures are missing from the MAKO RIO THA Application User Guides | Labeling design | Letter from Stryker dated 1/28/2017 was sent via USP Next Day Air with return receipt to all their customers. With the following actions; Inform users of this medical device correction and forwar this notice to all those individuals who need to be aware within your organization and return the attached Business Replay Form within 5 business days, to Fax 844-782-5561 or via e-mail to stykerortho4253@stericycle.com |  | 777 | 777 | US nationwide distribution. | Open, Classified |  |  | 777 | Not_Computer
Z-1412-2017 | 76456 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Turon Impaction Fixture |  |  | February 09, 2017 | 2017 | Encore Medical, Lp | During the Turon assembly, the impaction forces caused the polymer, black acetal copolymer from the Impaction Fixture to wear off on the lateral surface of the Humeral Stem titanium plasma spray coating. | Device Design | There are two field safety notices, one for Consignees who have surgeons that use the impaction fixtures (Verson 1) and one for Consignees who had previously indicated their surgeons do NOT use the impaction fixtures (Verson 2). The two recall notification letters were sent out on 2/9/17. |  | 297 | 297 units | US, South Korea, Australia, Canada, United Kingdom/Ireland, Germany | Open, Classified |  |  | 923 | Not_Computer
Z-1536-2017 | 76625 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | 9.6F Plastic Dignity Mid-Sized CT Port w/Pre-Attached Silicone Catheter |  |  | February 23, 2017 | 2017 | Medical Components, Inc dba MedComp | The 9.6F port kits were packaged with the incorrect valved peelable introducer. The label states the kit contains a 10F valved peelable introducer. The kit is packaged with a 9F valved peelable introducer. The port lumen will not fit through the introducer during the insertion procedure. | Unknown/Undetermined by firm | Medcomp mailed a Customer Notification letter dated February 23, 2017, to all affected customers to inform them of the issue. Customers were asked to examine their inventory and return all unused product. Distributors were also asked to locate the end user/facility of where the affected product was sold and communicate the recall to them. Customers were instructed to contact their customer service representative for a Returned Goods Authorization (RGA) number if necessary at 215-256-4201. Customers were also instructed to complete and return page 2 of the letter by fax to 215-256-9191. For questions regarding this recall call 215-256-4201. |  | 71 | 71 units | Worldwide Distribution to PR and Panama | Open, Classified |  |  | 71 | Not_Computer
Z-1415-2017 | 76401 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software using EchoIMS. The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | A situation can occur allowing two physicians to access the same study report in EchoIMS when launched from the Cardio Study List without receiving the read-only notification prompt. | Software design | Merge sent an Urgent: Medical Device Recall letter dated March 28, 2016, via e-mail on April 4, 2016. A letter was issued to nonresponders on July 11, 2016.The letter notifies the customer of the issue, informs them of the workaround, and that no further action was being taken to correct the issue, therefore, the customer needs to be aware of the possible behavior. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they understand the workaround. The response was required no later than April 18, 2016. Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case. For further information please call (877) 741-5369. |  | 17 | 17 sites potentially have the affected versions | US Distribution was made to medical facilities in CA, FL, IL, MD, MO, OH, OK, TX, VT, and WI. Military distribution was also made. | Terminated | March 20, 2017 | 350 | 17 | Not_Computer
Z-1416-2017 | 76153 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Roth Net Polyp Retrieval |  |  | December 13, 2016 | 2016 | US Endoscopy Group Inc | The firm was notified that the expiration date on the outer carton is different than the expiration date on the inner package. The expiration date is shorter than the actual expiration. | Labeling mix-ups | US Endoscopy Group sent an Urgent: Roth Net Polyp Retriever Recall letter dated January 6, 2017, to all affected customers.The firm requested their customers to do the following:1. Quarantine and destroy the affected Roth Net devices in your inventory.2. Contact US Endoscopy Customer Service at 1-800-769-8226 to obtain no cost replacement product.3. Complete and return the attached Response Card.For further information please call (440) 358-6055. |  | 2131 | 2,131 units | US Distribution to the states of : AL, CA, CO, IL, KS, MA, MT, NC , NE , NY, OH, PA, TN, TX and WA.Product was also shipped to the following countries: Australia, Belgium, Canada, Czech Republic, Ireland, Kuwait, Norway, Poland &amp; Sweden. | Terminated | June 13, 2017 | 182 | 2131 | Not_Computer
Z-1705-2017 | 76637 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | The CG+ Arrow PIC Catheters permit venous access to the central circulation through a peripheral vein. |  |  | March 03, 2017 | 2017 | Arrow International Inc | Arrow International, Inc. notified its customers who received the affected Antimicrobial/Antithrombogenic PICC Kits, that the kits were being recalled due to a labeling error. The lidstock on some of the kits reads CDC-45541-HPK1A, while the banner card on the side of the kit reads CDC-45052-HPK1A. Other kits may list CDC-45052-HPK1A on the lidstock and CDC-45541-HPK1A on the banner card. The affected lot numbers are 23F16C0851, 23F16A0727, and 23F16C0844. | Under Investigation by firm | Arrow International mailed an Urgent Medical Device Recall Notification letter to affected customers on March 7, 2017, to inform them of the issue. The notification informs customers to discontinue use and quarantine any products identified with the specific lot numbers. Customers are asked to complete the enclosed Recall Acknowledgement Form for product return. For further questions, please call (866) 246-6990. |  | 1351 | 1,351 catheters | US Distribution | Open, Classified |  |  | 1351 | Not_Computer
Z-1684-2017 | 76501 | No | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cyto-Chex BCT - blood specimen collection device - Immunophenotyping Preservative. IVD. Sterile. 2.0ml only. Packaged 6-tube pack REF 213559 and 100-tube box REF 218980. Streck. |  |  | February 23, 2017 | 2017 | Streck | Complaints of the 2.0ml tube stopper coming loose during phlebotomy when small diameter tube adapters were not utilized. Such adapters aid with centering small diameter tubes to properly pierce the middle of the stoppers. | Error in labeling | On February 23, 2017, Streck electronically notified customers of the updated device correction. The recalling firm used an electronic mail alert to notify all of their customers the IFU for the recalled product had been revised. The specified the revision included phlebotomy guidelines specific to the 2.0 ml blood collections tubes. The revised IFU was provided as a link in the email. The recalling firm is reportedly able to monitor the access of the revised IFU to ensure each customer has received the information. For further questions, please call (402) 333-1982. |  | 128692 | 128,692 tubes | Worldwide Distribution - USA (nationwide) and to the countries of : Belgium, Canada, Germany, Israel, Japan, the Republic of Korea, the Netherlands, and the United Kingdom. There was no military or government distribution. | Open, Classified |  |  | 128692 | Not_Computer
Z-1512-2017 | 76603 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mammotome Revolve Dual Vacuum Assist Biopsy |  |  | February 28, 2017 | 2017 | Devicor Medical Products Inc | The device did not have Revolve Stereo Probe labeling, no product code indicating the size and length of the probe, no sterilization sticker or lot number. | Process control | Devicor sent an Urgent Medical Device Recall letter dated March 10, 2017 to all affected consignees via first class mail. The recall response rate will be tracked and reconciled by consignee's response form, certifications of product inventory, and/or disposal. Consignees were asked to complete and return the enclosed response form as soon as possible. Consignees with any questions were instructed to call Mammotome Customer Service at 1-877-926-2666. For questions regarding this recall call 513-864-9158. |  | 320 | 320 | Worldwide Distribution - US including IA, IL, NY, OH, OK, PA, TN &amp; WV.Internationally to Chile and Brazil | Terminated | August 03, 2017 | 156 | 320 | Not_Computer
Z-1438-2017 | 76481 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Ruby System: |  |  | February 13, 2017 | 2017 | Abbott Laboratories | The Pump Relay Printed Circuit Board Assembly (PCBA) used in the CELL-DYN Ruby instrument may prematurely fail and lead to instrument stoppage | Nonconforming Material/Component | Product correction letters were sent to all affected customers on February 13, 2017 by traceable method. Letters provided information on the reason for the correction with all affected serial numbers listed. Abbott will be contacting each location to schedule replacement of the PCBA Pump Relay Board.Update: five additional serial numbers were added. Correction letters were sent to new consignees - 3 in Brazil and 2 in Mexico. |  | 116 | 116 | US and Afghanistan, Brazil, Egypt, Tanzania, Hungary, Iraq, Kenya, Netherlands, Russia, Romania Vietnam, , United Kingdom, Saudi Arabia, Philippines, Lebanon, Mexico, Japan, Jordan, India, Indonesia, Turkey, UAE, Gaza &amp; Jericho, Germany, Bulgaria, Chile, Argentina, Luxembourg. | Open, Classified |  |  | 116 | Not_Computer
Z-1421-2017 | 76212 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000 Progesterone; Catalog Number: L2KPW2 |  |  | January 04, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc | DHEA-S causes falsely elevated progesterone results. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction (UMDC) letter dated January 4, 2017 to all affected Siemens Healthcare Diagnostics customers in the United States dated . An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on January 4, 2017. These letters identified the affected product problem and actions to be taken. For questions contact your local Siemens Customer Care Center or your local Siemens Technical support. |  | 11963 | 11963 | US Nationwide distribution including Puerto Rico | Open, Classified |  |  | 98294 | Not_Computer
Z-1431-2017 | 76594 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge PACS software. The firm name on the label is Merge Healthcare. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Studies that have QC, PDE, or HL7 ADT updates applied may fail to propagate via telmed protocol to the destination server and may fail to import on the destination server. | Software design | Merge sent an Urgent: Medical Device Recall letter dated March 28, 2016, to all affected customers via email the week of April 4, 2016. If no email address was available, the letters were issued via certified mail. A letter dated February 21, 2017 , was issued was also issued via email or certified mail.The letter notifies the customer of the issue, informs them not receiving updated demographics or image content may result in delay in diagnosis or treatment of the patient or misdiagnosis, and they should be aware of the workaround. The customer is also notified a fix is not yet available, however, the recalling firm is working on correcting the issue and will inform them once a fix is made available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they understand the workaround and are they interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.A second letter was issued dated February 21, 2017,via email or certified mail if no email address was available. The letter notified customers a fix was now available and also contained the same information and response form as the first letter. The letter informed the customer a response was required 15 calendar days after receipt of the letter.The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369. |  | 248 | 248 sites potentially have the affected versions | USA (nationwide ) Distribution was made to medical facilities. Internationally to Canada. There was no government or military distribution. | Open, Classified |  |  | 248 | Not_Computer
Z-1429-2017 | 76367 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Apollo DRF; Model: 9784220831 |  |  | December 27, 2016 | 2016 | Villa Sistemi Medicali S.P.A. | Patient fell from the table because the footrest of the device detached from the table top while the patient was positioned on the table and table was moving from vertical position. | Nonconforming Material/Component | Customers were notified of the recall via DHL or similar mail service. The letter identified the affected product, problem and actions to be taken. |  | 65 | 65 devices in total | US Nationwide distribution in the states of AL, AZ, CA, CO, CT, DC, FL, IA, IL, IN, LA, MA, MD, MI, MN, MO, MS, NJ, NY, PA, Puerto Rico, TN, TX, UT, and W.I | Open, Classified |  |  | 65 | Not_Computer
Z-1426-2017 | 76541 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DR Systems Unity PACS software, now known as Merge Unity PACS |  |  | April 08, 2016 | 2016 | Merge Healthcare, Inc. | Reports were missing demographic header information when they were faxed to referring physicians. | Software design | Merge sent an Urgent : Medical Device Recall letter dated March 28, 2016, to all affected customers via email on April 8, 2016. Nonresponders were issued a second letter dated July 11, 2016.The letter notifies the customer of the issue, makes them aware of a workaround, and informs them a fix is available. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 8, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case.The recalling firm reported emails will be tracked for delivery confirmation.For further questions please call ( 877) 741-5369. |  | 9 | 9 sites potentially have the affected version | US Distribution was made to medical facilities in CA, MT, PA, and TX. There was no foreign/military/government distribution. | Terminated | March 17, 2017 | 343 | 9 | Not_Computer
Z-1434-2017 | 76459 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVATRON M 6730/6740 (MEX) Accelerator, Linear, Medical |  |  | January 25, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Siemens became aware of a defective weld seam at overhead suspensions. Due to this defect it is possible for the weld seam to break and to cause the overhead suspension to fall down. This can lead to a severe injury to a patient or any other person. | Under Investigation by firm | Siemens sent a Customer Safety Advisory Notice (CSAN) dated February 7, 2017, to affected customers to inform them of a planned hardware update regarding the mechanics of the overhead suspension of their Digital Linear Accelerator. The notice also explained the actions Siemens is taking to perform the necessary fix at the affected sites. Customers were advised to notify anyone at their organization that should be aware of this information. Also, the safety notice should be forwarded to the new owner and Siemens also requested they be informed of the identity of the device's new owner where possible. Customers with questions were instructed to call 1-800-888-7436. For questions regarding this recall call 610-448-6478. |  | 35 | 35 systems | Nationwide Distribution | Open, Classified |  |  | 35 | Not_Computer
Z-1441-2017 | 76473 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | RTI Biologics BioSet IC RT Paste 2 cc |  |  | January 23, 2017 | 2017 | RTI Surgical, Inc. | Incorrect instructions for use (IFU) included. | Labeling Change Control | The firm, RTI Surgical, Inc., issued and "IMPORTANT NOTIFICATION Field Action" letter dated January 20th, 2017 to its consignee via e-mail on 01/23/2017. Verification and confirmation of the e-mail was received on the same date. The letter described the product, problem and actions to be taken. The consignee were instructed to immediately quarantine any grafts in your possession; return any unused RTI and include return authorization along with instructions on how to return the product.If you have any questions, Sr. Manager, Regulatory Compliance at 386-418-8888 ext. 4396 or email: amcconnell@rtix.com. |  | 24 | 24 units | International Distribution: Colombia only | Open, Classified |  |  | 24 | Not_Computer
Z-1439-2017 | 76189 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD Solis VIP Ambulatory Infusion Pump, Model 21-21210, Reorder 21-2120-0102-15, |  |  | October 31, 2016 | 2016 | Smiths Medical ASD, Inc. | 20 Pumps sold to the Finnish market contain a message in which one word in the message is mistranslated. When the user follows a specific set of key presses the pump will display the incorrect message. The message indicates that a Patient Controlled Analgesia (PCA) dose is unavailable because the pump is running. It should indicate that the PCA dose is not available because the pump is stopped. The function of the pump is unchanged and no patient injury can occur since no drug is being delivered. | Software design | Smiths Medical sent an Urgent Medical Device Field Safety Notice dated October 26, 2016, via FedEx on October 31, 2016. The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. Advised consignees to contact the Smiths Medical Service &amp; Repair Department for the software to be loaded and to complete the Recall Confirmation Form and return it to Smith Medical via e-mail to FCA.Response@Smiths-medical.com. For further questions please call (651) 633-2556. |  | 20 | 20 | Internationally to Finland | Open, Classified |  |  | 20 | Computer
Z-1470-2017 | 76654 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive when used with RadSuite. The firm name on the label is Merge Healthcare. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The software produced a number of "do not route" exceptions, which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge Healthcare sent an Urgent: Medical Device Recall letter dated March 28, 2016, via e-mail the week of April 4, 2016. If no email address was available, the letters were issued via certified mail. A second letter dated July 11, 2016, was issued to nonresponders.The letter notifies the customer of the issue and informs them a fix has been released for the issue. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016.The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case.The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369. |  | 187 | 187 sites potentially have the affected versions | USA (nationwide ) Distribution to medical facilities. Government distribution was also made. There was no foreign or military distribution. | Open, Classified |  |  | 187 | Computer
Z-1699-2017 | 76633 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Stryker VisumÂż LED Surgical Lighting system is intended to illuminate the operative site during surgical procedures with high intensity light. Affected part numbers are 0682001432 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 9P/FP), 0682001286 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 5P/FP), 0682001298 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 5P), 0682001299 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 5P/5P), 0682001300 (EDS LIGHT SUSPENSION, CENTRAL AXIS, FP/5P/5P), 0682001433 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 9P), 0682001434 (EDS LIGHT SUSPENSION, CENTRAL AXIS, 9P/5P), 0682001435 (EDS LIGHT SUSPENSION, CENTRAL AXIS, FP/9P/5P) |  |  | February 28, 2017 | 2017 | Stryker Communications | The potential hazard that can arise for the combination of the missing screws and the overloading by the end user is insufficient mount force that causes the equipment to fall. | Device Design | Stryker sent an Urgent Medical Device Correction letter dated March 3, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Distributors will notify customers via tracked Customer Notification Letter. A method of written communication will include a confirmation of delivery. For questions contact Stryker Technical Support at 800-243-5135 or comm.techservices@stryker.com |  | 292 | 292 devices (in US) | Worldwide distribution - United States Nationwide - Australia, Canada, Great Britain, Japan, Netherlands and Puerto Rico | Open, Classified |  |  | 292 | Not_Computer
Z-1533-2017 | 76639 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza, Picture Archiving and Communications System (PACS) |  |  | February 20, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software updates | Software design | Siemens mailed a notification letter dated February 21, 2017 to affected customers to inform them of the software updates that were now available to address several issues. Customers were advised to follow the instructions in the notice and where necessary promptly notify and instruct all their staff at their organization who have to be aware of the problem. Customers were also instructed to forward the information to new owners where possible. For questions regarding this recall call 610-219-4834. |  | 13 | 13 systems | Nationwide Distribution to MO,TX, FL, MA, WI, PA, IN, and CA | Terminated | August 07, 2017 | 168 | 13 | Computer
Z-1761-2017 | 76648 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Laboratory; Product Number: 5002-0180 |  |  | March 07, 2017 | 2017 | PerkinElmer Life and Analytical Sciences, Wallac, OY | Potential for errors in patient results generated by the Specimen Gate Laboratory Software. | Software design | PerkinElmer Life and Analytical Sciences sent an Urgent Medical Device Correction letter dated March 7, 2017, to all affected consignees. The letter contains the following instructions: Until PerkinElmer's corrective measures are completed, we request our customers to close the Specimen Gate Laboratory Result Viewer between viewing and especially between manually calculating assay results. Closing and reopening the Result Viewer between processing the results from different assays effectively prevents the described issues from occurring. We also request our customers to follow unreported specimens with available tools to avoid any delay in reporting screening results. The available tools will vary depending on their process or the LIMS system that is used for activities associated with management of patient demographics, patient reports and the overall management of data generated from newborn screening testing. The final corrective action will involve an installation update to the software. PerkinElmer will deploy an installation update that will prevent the described issue when assays are calculated during the same Result Viewer session without closing it between the processing of different assays. Once the installation update is available, Software Services will contact you and arrange for the deployment. The installation update will be provided to you free of charge.Customers were instructed to complete the enclosed response form and return it by fax to 1-330-825-8520 (US) =358 2 2678 357 (OUS).For questions regarding this recall call 317-418-1735. |  | 29 | 29 | Worldwide Distribution - US including AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MD, MI, MO, NV, OH, OK, OR, TN, TX, VA; and Internationally to Canada and Brazil | Open, Classified |  |  | 29 | Computer
Z-1454-2017 | 76587 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bencox Mirabo PE Liner 36/44 Model Number: H1.L61.3644, UDI: 8806373852343 |  |  | February 27, 2017 | 2017 | CORENTEC CO., LTD | Lot may contain the wrong instruction (Instruction for Chinese customer) instead of correct US instructions for use. | Other | Corentec sent an Medical Device Recall letter dated February 23, 2017. The recalling firm notified the consignee via written notification by email and by phone. The recalled products will be returned to the recalling firm. For further questions please call (512) 327-9997 |  | 5 | 5 ea | US Distribution to the state of : PA | Terminated | May 11, 2017 | 73 | 5 | Not_Computer
Z-1495-2017 | 76442 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Orthosize Templating Version 1.2.6 |  |  | February 14, 2017 | 2017 | Zimmer Biomet, Inc. | Digital templates were created with the incorrect files. | Software Manufacturing/Software Deployment | On 2/14/2017 URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected users and 3rd party users via email. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for 3rd party users to respond to the formal recall notification. Orthosize Templating UsersIn approximately 6 weeks you will receive a notice that an update is available for the Orthosize Templating App. Upon selecting the app you will be required to update the app in order to use it.1. Review this notification and ensure affected personnel are aware its contents.2. Immediately update the app when the update becomes available.3. If after reviewing this notice you have further questions or concerns please call 411 TechnicalServices at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday.Alternatively, yourquestions may be sent by email to corporatequality.postmarket@zimmerbiomet.comDigital Template Users (3rd party)1. Review this notification and ensure affected personnel are aware of the contents.2. Immediately update your systems to include the revised digital templates included with notice.3. Complete Attachment 1  Certificate of Acknowledgement.a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com.b. Retain a copy of the Acknowledgement Form with your records.4. If after reviewing this notice you have further questions or concerns please call 411 TechnicalServices at 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Callsreceived outside of the call center operating hours will receive a prompt to record a voicemail orbe transferred to an on-call representative in the case of an emergency. Alternatively, yourquestions may be sent by email to corporatequality.postmarket@zimmerbiomet.com |  | 3232 | 3232 | Worldwide distribution. US Nationwide including Puerto Rico, Canada, Argentina, Australia, Austria, Belgium, Chile, China, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Korea, Lithuania, Malaysia, Mexico, Netherlands, Nicaragua, Norway, Pakistan, Poland, Portugal, Romania, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Open, Classified |  |  | 3232 | Not_Computer
Z-1446-2017 | 76476 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre Vascular Reddick Scoop Tip Cholangiogram Catheter (eIFU) |  |  | February 07, 2017 | 2017 | LeMaitre Vascular, Inc. | The Contains Latex symbol was not printed onto the box label or the sterile (inner) product pouch | Labeling design | LeMaitre Vascular, Inc. issued a Dear Doctor letter, on 2/7/17, to distributors and hospitals via FedEx overnight. The letter contains a form that is requested to be returned to LeMaitre Vascular, Inc. as a record of notification and reconciliation. Any unused product requested to be returned and replaced.Questions concerning this notice contact 781-221-2266 ext. 183. |  | 4673 | 4673 | Worldwide distribution. US Nationwide, AUSTRIA, BELGIUM, FRANCE, GERMANY, ITALY, LUXEMBOURG, Morocco, Saudi Arabia, Singapore, SPAIN, SWITZERLAND, UAE, and UK | Open, Classified |  |  | 21150 | Not_Computer
Z-1681-2017 | 76309 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Care Systems, consisting of Merge Eye Station, Merge Eye Care PACS, and Merge Eye Station Import Utility (ESIU). |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The system will completely stop allowing the downloading or acquisition of new images due to incompatible software application versions. | Software Design Change | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. |  | 1627 | 1,627 sites potentially have the affected software | Distribution was made to medical facilities nationwide. Government distribution was made. Foreign distribution was made to Canada, as well as other countries. There was no military distribution. | Open, Classified |  |  | 1627 | Computer
Z-1459-2017 | 76458 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1-Level Sensor II Gel Pads |  |  | February 20, 2017 | 2017 | Terumo Cardiovascular Systems Corporation | Terumo CVS initiated a voluntary recall for the Level Sensor II Pads and Level Sensor Gel Pads due to non-compliant labeling because the product expiration date is displayed in a format that may not be recognizable to all users. | Under Investigation by firm | Terumo CVS initiated a voluntary recall by issuing a safety advisory for their Terumo Advanced Perfusion System 1-Level Sensor II Pads, and Terumo Advanced Perfusion System 1-Level Sensor II Gel Pads due to non-compliant labeling because the product expiration date is displayed in a format that may not be recognizable to all users. Customers are asked to do the following: 1. Review the Safety Advisory.2. Assure that all users receive notice of this issue.3. Confirm receipt of this Safety Advisory by emailing or faxing the attached Customer Response Form to the email address or fax number indicated on the form.For questions contact Terumo CVS at 1-800-521-2818, Monday  Friday, 8 a.m.  6 p.m. ET. Any adverse events experienced with the use of this product, and/or quality problems should also be reported to the FDAs MedWatch Program:Phone: 1.800.FDA.1088 Fax: 1.800.FDA.0178, Web: www.fda.gov/medwatch/report.htm MedWatch Online Voluntary Reporting Form (mail to address on form): www.fda.gov/Safety/MedWatch/HowtoReport |  | 3794 | 3,794 | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV and the countries of: Mexico, AUSTRALIA, UNITED ARAB EMIRATES (UAE), Indonesia, Singapore, Taiwan, Thailand, COLOMBIA, CHILE, Vietnam, India, China, Malaysia, BELGIUM, Japan, CANADA | Terminated | June 01, 2017 | 101 | 212354 | Not_Computer
Z-1468-2017 | 76615 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis fluoroscopic x-ray system |  |  | February 27, 2017 | 2017 | Siemens Medical Solutions USA, Inc | A gap in the housing of the wireless foot switch could potentially result in liquids penetrating the interior. Such liquids may include disinfectants, cleaning agents and/ or bodily fluids. In rare cases, the introduction of liquids in the wireless foot switch housing could result in the failure of the foot switch. | Process control | A Customer Safety Advisory Notice (CSAN) was mailed to affected customers on 02/27/2017 to inform them of the issue. The notice provided information on what actions to take should the foot switch fail to function, the risks involved and what corrective actions they are taking to fix the issue. |  | 8297 | 2,158 units distributed in US and 6,139 worldwide | Distributed Nationwide | Open, Classified |  |  | 8297 | Not_Computer
Z-1457-2017 | 76447 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | In some instances, the system will lock tabs within a study, even when a second user does not have the study open on a different workstation. | Software design | The recalling firm issued a letter dated March 28, 2016 via e-mail. If no email address was available, the letters were issued via certified mail. Non-responders were issued letters on July 11, 2016. The letter identified the affected product, problem and actions to be taken. For question send an email to recall@merge.com. |  | 361 | 361 sites potentially have the affected versions | US Nationwide Distribution. | Open, Classified |  |  | 361 | Not_Computer
Z-1460-2017 | 76429 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Arkon Anesthesia Delivery System, Model 99999. |  |  | February 08, 2017 | 2017 | Del Mar Reynolds Medical, Ltd. | Spacelabs has received reports of the Arkon Anesthesia Delivery System, Model 99999, shutting down without warning. The power supply has failed in a way that does not trigger the normal on battery alarms. | Software design | On 08 February 2017, a customer letter will be sent via priority service, return receipt requested, to all U.S. customers. Records of notification will be maintained to ensure complete notification.On 15 February 2017, a customer letter (translated as necessary) will be emailed to all international subsidiaries and distributors of record. Records of notification will be maintained to ensure complete notification.Upgrade is expected to be available in April 2017; all affected Arkon Anesthesia Delivery Systems will receive a software upgrade onsite. |  | 450 | 325 devices sold to 35 US consignees and 76 devices sold to 14 international consignees | Worldwide Distribution - US nationwide and internationally in the following countries: BOLIVIA, CANADA, CHINA, ECUADOR, FRANCE, IRELAND, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PANAMA, PERU, POLAND, UNITED KINGDOM. | Open, Classified |  |  | 450 | Computer
Z-1463-2017 | 76339 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The ABL800 Model#:All |  |  | January 26, 2017 | 2017 | Radiometer America Inc | Reports that the sample type "cord blood" has been changed into the two sample types "cord blood arterial" and "cord blood venous. | Device Design | Customer notification letters were sent 01/27/16. The letter provided the following instructions:The firm recommends that customers ensure that operators and clinicians using the "cord blood sample results from a HIS/LIS or a middleware system to determine treatment are made aware that arterial or venous is not transmitted.An upgraded version of the ABL90 and ABL800 series software will ensure that the full sample type cord blood arterial and cord blood venous will be transmitted to the HIS/LIS or middleware system. The new software version will be installed by local engineer when available.If not the end user customers are instruction to provide a copy of the recall notification letter to the end user.Any questions regarding this letter contact Radiometer America Technical Support at 1-800-736-0600 opt.1. |  | 17522 | 17,522 | U.S. distribution to the following; nationwide.Foreign distribution to the following; Netherlands, China, Czech Republic, Denmark, Germany, Spain, France, Hungary, India, Japan, Poland, South Africa, Singapore, Switzerland, Turkey, Norway, Austria, Italy, Sweden, Finland. | Open, Classified |  |  | 27210 | Computer
Z-1456-2017 | 76326 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VASOVIEW HEMOPRO VH-3500 |  |  | January 27, 2017 | 2017 | Maquet Cardiovascular, LLC | Maquet has received several complaints involving the VASOVIEW HEMOPRO VH-3500 which involve the device exhibiting the failure mode of intermittent or no power. | Under Investigation by firm | Maquet Inc. issued an Urgent Removal Correction Notice letter dated January 27, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check inventory for affected device. Remove and keep it in a secure location to preclude being used. Customers were asked to complete the Medical Device Removal Response Form and return to Maquet by e-mailing a scanned copy to HemoproPower@getinge.com or by faxing the form to 1-973-629-1518. |  | 4283 | 4283 units | US Nationwide Distribution | Open, Classified |  |  | 4283 | Not_Computer
Z-1398-2017 | 76175 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba Ultimax DREX-ULT80 X-ray generator |  |  | May 19, 2016 | 2016 | Toshiba American Medical Systems Inc | It was discovered that the generator of the system could possibly terminate the exposure prematurely during an examination | Radiation Control for Health and Safety Act | Toshiba America Medical Systems planned action to bring defect into compliance :1. A notification letter will be sent to customers with affected systems. This letter includes the statement that the repairs will be made at no cost to the customer.2. Modified firmware will be installed on the systems to prevent the defect from recurring.3. The consignees will be contacted by a Toshiba service representative for an appointment when the modified firmware becomes available4. The consignee will be informed that if this problem occurs before the new firmware is installed; they should reset the x-ray generator to complete the exam; and5. The CAP will be completed by May 1, 2017. For further questions please call (714) 730-5000 |  | 254 | 254 | USA ( nationwide ) Distribution to the states of : WA, OR, CA, OH, IL, CT, IA, KY, AL, LA, HI, PA, NJ, DE, MN, DC, NY, MI, SC, MD, TN, OK, WI, UT, ID, NV, TX, AL, IN, WV, KS, MO, CO, MT, MS, NC, GA, NH,UT, FL, AZ and ND. | Open, Classified |  |  | 254 | Not_Computer
Z-1465-2017 | 76433 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | BivonaÂż TTS (Tight to Shaft) Pediatric Tracheostomy Tube, Model No. 67SP035, Reorder No. 67SP035. |  |  | February 09, 2017 | 2017 | Smiths Medical ASD Inc. | Smiths Medical became aware that one lot of the BivonaÂż 3.5mm TTS Cuffed Pediatric tracheostomy tubes was incorrectly labeled for neonatal use instead of correctly labeled for pediatric use. | Labeling False and Misleading | Consignees were sent a Smiths Medical "Urgent Medical Device Field Safety Notice" dated February 08, 2017. The letter was addressed to "Clinician". The letter listed the Affected devices, Reason for Recall, Risk to Health, and Instructions to Customers. Advised customers to complete and return the "Urgent Medical Device Field Safety Notice (FSN) Response Form" to smithsmedical4390@stericycle.com. All affected devices must be returned to Stericycle for processing. If the product was further distributed, consignees were asked to immediately notify them of the Recall and retrieve all affected devices. For questions contact Stericycle via e-mail at the following address: smithsmedical4390@stericycle.com. |  | 142 | 142 | US: CA, ILM NC, GA, NE, MA, CO, CT, MO, TN, LA, MD, MN, FL, SC,M OK, MI, PA, TX, WA, WI. | Open, Classified |  |  | 142 | Not_Computer
Z-1712-2017 | 76695 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Enzyme 2 Calibrator, multi-analyte mixture |  |  | February 23, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed a positive bias of up to 41% in quality control and patient results for Aspartate Aminotransferase (AST) and Aspartate Aminotransferase concentrated (AST_c) assays following calibration with Advia Chemistry Enzyme 2 Calibrator lots 6JD079 and 6MD042. | Under Investigation by firm | Siemens mailed an Urgent Medical Device Correction Notice to affected customers on 02/21/2017. The notice informed calibrator customers of the reason for the correction, risk to health, and actions to be taken by customers who are in the process of converting from the fixed factor to the enzyme calibrated assay and information for customers who had already converted, until the investigation is complete. |  | 314 | 314 calibrators | Worldwide Distribution - United States Nationwide:DenmarkSwedenNorwayGreat BritainBelgiumNetherlandsSpainIsraelSwitzerlandFrancePortugalGreeceGermanySloveniaCzech RepublicHong Kong Colombia Korea Australia | Open, Classified |  |  | 314 | Not_Computer
Z-1477-2017 | 76358 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | C2 CryoBalloon Ablation System - Controller Cap. |  |  | January 06, 2017 | 2017 | C2 Therapeutics, Inc. | The Controller Cap of the C2 CyroBalloon Ablation System may crack as it is tightened onto the controller. That could compromise the integrity and result in damage to the Controller Cap. | Under Investigation by firm | Firm initially notified consignees who would be using the device in an upcoming procedure on January 6, 2017 to alert them of the potential issue and recommended that products with the affected cap not be used unless it can be checked by a C2 rep. Controller caps will be replaced within a few weeks; facilities will receive replacement caps by mail along with a shipping inventory label for return of those currently on hand. Formal recall letters were sent on January 19, 2017. Letters were sent by e-mail with receipt and read confirmation and follow up by phone call. |  | 755 | 755 controller caps | Worldwide Distribution - US Nationwide and the countries of Germany, Belgium, The Netherlands United Kingdom and China. | Open, Classified |  |  | 755 | Not_Computer
Z-1543-2017 | 76593 | No | MIH |  |  | Cardiovascular | PMA | Endurant / Endurant II 23mm and 25mm Bifurcated Stent Graft Systems |  |  | February 27, 2017 | 2017 | Medtronic Vascular, Inc. | Medtronic is initiating a voluntary Urgent Medical Device Recall for a subset of Endurant/ Endurant IIBifurcated Stent Graft Systems of specific models and serial numbers as this specific subset of stent grafts has greater susceptibility to fabric permeability variations that may be associated with endoleaks observed during the initial implant procedure. | Packaging change control | Medtronic sent an Urgent Medical Device Recall letter dated March 2017, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that your facility has received a potentially affected Endurant /Endurant II 23mm and / or 25mm Bifurcated Stent Graft System from the identified subset. As a result, Medtronic is asking that you take the following actions:Identify and quarantine unused Endurant /Endurant II 23mm and 25mm Bifurcated Stent Graft System from the identified subset of models and serial numbers that are in your inventory .Return unused products from the attached model and serial number list that are in your inventory to Medtronic by contacting Customer Service at 1-800-854-3570 (#4) and referencing this communication to initiate a return and credit of unused product. Your Medtronic sales representative can assist you in the return of this product as necessary.Complete the attached Customer Confirmation Certificate and email it to RS.CFQFCA@Medtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.For further questions please call Customer Service at 1 (800) 633-8766. |  | 3331 | 3331 devices | Worldwide Distribution - US Distribution including Puerto Rico and to the countries of :ArgentinaArmeniaAustraliaAustriaBelgiumBrazilBulgariaCanadaChinaColombiaCosta RicaCroatiaCyprusCzech RepublicDenmarkDominican RepublicEcuadorEgyptEl SalvadorFinlandFranceFrench PolynesiaGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIranIrelandIsraelItalyJapanJordanKazakhstanKoreaKuwaitLatviaLebanonLithuaniaLuxembourgMalaysiaMexicoMontenegroNetherlandsNew ZealandNorwayOmanPanamaParaguayPeruPhilippinesPolandPortugalReunionRomaniaRussian FederationSaudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSri LankaSwedenSwitzerlandTaiwanThailandTurkeyUnited Arab EmiratesUnited KingdomUruguayVenezuelaViet Nam | Open, Classified |  |  | 3331 | Not_Computer
Z-1710-2017 | 76688 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 assay Immunoassay System, Remanufactured, Part No. 386220 |  |  | March 07, 2017 | 2017 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Family of Immunoassay Systems because a circuit board that controls motor functions can affect device positioning on the pipettor and incubator belt while the instrument is running. The position errors cause the instrument to stop running and cancel tests that are in process. | Software Design Change | An Urgent Medical Device Recall letter was sent on 3/7/17 to inform customers that Beckman Coulter has determined through internal investigation that a circuit board that controls motor functions in the instrument systems listed above can affect device positioning on the pipettor and incubator belt while the instrument is running. Customers are informed of the impact, action to be taken, and the resolution. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Customer Technical SupportCenter:Âż From our website: http://www.beckmancoulter.comÂż By phone: call 1-800-854-3633 in the United States and Canada.Âż Outside the United States and Canada, contact your local Beckman Coulterrepresentative. |  | 851 | 851 units total for all products (291 units in US) | Affected geographies: Algeria, Austria, Australia, Bangladesh, Bulgaria, Brazil, Canada, China, Costa Rica, Cyprus, Colombia, Czech Republic, El Salvador, France, Germany, Ghana, Greece, Guatemala, India, Indonesia, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malawi, Mexico, Morocco, Myanmar, Nepal, Oman, Panama, Philippines, Pakistan, Paraguay, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, South Korea, Spain, Switzerland, Tunisia, Turkey, United Kingdom, United States, Vietnam, Zimbabwe | Open, Classified |  |  | 851 | Not_Computer
Z-1483-2017 | 76562 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | .045 X 6 ST Guide Wire, Part number 45-0029-S, Sterile 10 pack of WS-1106ST and has a 6 inch length and 0.045 inch diameter. |  |  | February 17, 2017 | 2017 | Acumed LLC | Guide wires sized as .045 X 6 were mis-packaged as .054 X 7 guide wires and those sized as .054 X 7 were mis- packaged as .045 X 6 guide wires. There is a possibility that using the incorrect guide wire size can cause health hazards, such as mal or nonunion, need for revision surgery, and histological response (if wire breaks). | Process control | Acumed sent an Urgent Medical Device Recall letter dated 02/16/2017, to all affected customers on 02/17/2017.Consignees were instructed to follow special instructions and send survey to RecallNotification@acumed.net. Domestic customers with questions or need to return products to Acumed, please contact Acumed Agent Inventory via email AgentInventory@acumed.net to obtain a Return Material Authorization (RMA) number or contact your agent directly. International customers with questions or need to return products to Acumed, please contact Acumed Business Service via email at BusinessServices@acumed.net or call to obtain a Return Material Authorization (RMA) number. Customers with questions about recall can call 877-627-9957, choose Option 4 (US and Canada), or call 001-503-627-9957, choose option 4 (outside of US and Canada). |  | 1.08 | 24 units of .045 X 6 ST Guide Wire and 25 units of .054 X 7 ST Guide Wire | Worldwide Distribution in Great Britain, Ireland, and Japan only. | Terminated | May 19, 2017 | 91 | 1.08 | Not_Computer
Z-1478-2017 | 76638 | No | NPK | Bone grafting material. | Dental | Dental | 510(k) | Inion GRT Tack, drill bit and tack appl., 4x, Sterile. |  |  | March 01, 2017 | 2017 | Inion Ltd. | Aluminum pouch seal was noticed to be defective (slightly open). | Other | Inion sent an Urgent: Medical Device Recall Notification letter dated March 1, 2017. The recall letter asked customers to inspect the aluminum pouch seals, remove any devices with defective seals from sales stock, and to return the product to Inion. The customers were also asked to inform Inion Oy of the outcome of their inspections using a feedback form. For further questions call (954) 659-9224. |  | 36 | 36 pieces | US Distribution to : FL and NC. | Open, Classified |  |  | 36 | Not_Computer
Z-1480-2017 | 76483 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | 882480: BrightView designed for single or dual detector nuclear imaging accommodating a range of ECT studies. In addition it can be used to perform planar static, dynamic, gated, total body, circular-orbit and noncircular orbit SPECT, gated SPECT (circular and noncircular)studies, computer-programmed protocol strings, and reference scans (dual detectors). |  |  | January 31, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | Four issues: 1. Motion controller problem stops scan and no data image produced. 2. Door interlock switch problem disables CT scan. 3. Detector contacts head holder when performing Patient Unload. 4. JETStream freezes during gated planar scan | Software design | Consignees received a customer information letter via certified mail, 1/31/2017. International consignees were notified by Philips Healthcare Global Market representatives. Field Service Engineers are contacting consignees to conduct appropriate corrections. Non-responders will be contacted multiple times. |  | 1218 | 1218 total | Worldwide distribution. US nationwide, ALGERIA, ARGENTINA, ARMENIA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CAYMEN ISLANDS, CHILE, CHINA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, FINLAND, FRANCE, GABON, GERMANY, GREAT BRITAIN, GREECE, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PALESTINE, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, &amp; VIETNAM. | Open, Classified |  |  | 1218 | Not_Computer
Z-1484-2017 | 76607 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-636, 60" Microbore extension set. |  |  | February 20, 2017 | 2017 | Vygon MFG, Inc., dba/ Churchill Medical Systems, Inc. | Incorrect priming volume printed on the device package. | Error in labeling | Recall letters were sent to the customer via email and fax on 2/20/2017. They were asked to take the following actions: The recall will be conducted at no cost to you. To effect return of the recalled product and replacementwith new inventory, we request your cooperation in the following steps:1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately; 2. Count and document your affected inventory on the "Recall Acknowledgement and Inventory Return Form; 3. Fax or email the form to Vygon including customer contact information; 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number; 5. Return the product as well as the original signed Recall Acknowledgement and InventoryReturn Form using the instructions from Customer Service involving product pick-up; 6. All inventory will be credited or replaced at no cost to you. If you have questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-4 73-5414, or by e-mail at customerservice@vygonus.com. |  | 30700 | 30,700 | US Nationwide Distribution in the states of NJ, TX, VA, FL, VA, MA, IL, MI, NY, OH, WA, PA, WV | Open, Classified |  |  | 30700 | Not_Computer
Z-1713-2017 | 76675 | general purpose reagent | PPM | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | NucliSENS Lysis Buffer, REF 200292 |  |  | March 07, 2017 | 2017 | BioMerieux SA | Problem with colored eluates for whole blood extractions | Under Investigation by firm | BioMerieux sent an Urgent Product Safety Correction Notice to all affected customers electronically on March 7, 2017. US customers will be sent a separate letter. The letter requested that they discard the product and notify all customers that use the product. The notification included a reply form to be returned. Customers with questions were instructed to contact their local bioMerieux Customer Service representative. For questions regarding this recall call 314-731-7301. |  | 1025 | 1025 | Worldwide Distribution - US including IL, OH, WA, and Internationally to Argentina, Australia, Brazil, Chile, Columbia, France, Hong Kong, Italy, Netherlands, Peru, Spain, Uganda, United Kingdom, and Zimbabwe. | Open, Classified |  |  | 1025 | Not_Computer
Z-1486-2017 | 76431 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Cardio study list does not show STAT studies without refreshing. | Software design | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. Nonresponders were issued letters dated 7/11/2016. |  | 217 | 217 sites potentially have the affected versions | Distribution was nationwide to medical facilities. There was also military distribution. There was no foreign/government distribution. | Open, Classified |  |  | 217 | Not_Computer
Z-1534-2017 | 76713 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | ARTIS Q, Interventional Fluroscopic x-ray system |  |  | February 22, 2017 | 2017 | Siemens Medical Solutions USA, Inc | In Artis systems with A100G generators, a component can be affected by aging which could result in the potential failure of a module in the high-voltage generator. | Other | A Customer Safety Advisory Notice (CSAN) was mailed on 02/22/2017 to customers affected by the issue on 02/22/2017 to inform them of the corrective action they plan to perform to prevent a possible hazard to persons and equipment. The notice explained the underlying issue, the impact on system operation and the potential risk to patients. It also informed users of the actions being taken and how they will be implemented. |  | 264 | 264 distributed worldwide (112 in the U.S.) | Distributed throughout the United States. | Open, Classified |  |  | 264 | Not_Computer
Z-1703-2017 | 76690 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 Intelligent Lab System and Dimension Vista 1500 Intelligent Lab System, analyzer, chemistry (photometric, discrete), for clinical use |  |  | March 03, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that in isolated cases when ECREA is processed immediately after the weekly automated Acid Clean routine during Probe Test, there is the remote potential for an elevation of greater than 15 percent in the ECREA result. | Under Investigation by firm | Siemens mailed an Urgent Medical Device Correction Notice letter dated March 3, 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Siemens Customer Care Center - Technical Solutions at (800-441-9250) or your local Siemens Technical Support Representative. |  | 2733 | 2733 systems | Worldwide Distribution - United States Nationwide and Puerto Rico, Canada, Qatar, Republic Korea, Saudi Arabia, Slovakia, Slovenia, Spain, Switzerland, United Kingdom, Australia, Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, and Portugal. | Open, Classified |  |  | 2733 | Not_Computer
Z-1499-2017 | 76047 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Arial Wireless Water-Resistant Call Pendant |  |  | June 22, 2016 | 2016 | Stanley Security Solutions Inc | Devices were incorrectly programmed during manufacturing therefore depressing the pendant button may result in an alarm not sounding as intended. | Software design (manufacturing process) | Stanley Healthcare sent an Important Recall Notice dated June 22, 2016, to all affected customers. The notice instructed customers to discontinue use of the device and to contact the recalling firm to arrange for return of the device. The recalling firm also asked the customer to complete a response form indicating they read, understood and complied with the recall notice. Customers with questions were instructed to contact Arial Customer Service at 888-716-0077. For questions regarding this recall call 781-216-9035. |  | 500 | 500 individual pendants | Nationwide Distribution to AR, AZ, CA, FL, GA, IA, ID, IN, KS, MD, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SD, UT, VA, and WI. | Open, Classified |  |  | 500 | Not_Computer
Z-1498-2017 | 76312 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Merge Eye Station Import Utility (ESIU). The firm name on the labeling is Merge Healthcare. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | System locks up which may result in potential patient injury or delay in diagnosis or treatment. | Software design | Merge Healthcare, Inc. sent an Urgent: Medical Device Recall letter dated March 28, 2016, via email on April 4, 2016. If no email address was available, the letters were issued via certified mail. Nonresponders were issued letters dated July 11, 2016, via email.The letter notifies the customer of the issue and informs them of the workaround and that a fix was not yet available, although they were working on one. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they understand the workaround. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. Three call attempts will made before closing the case.The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369. |  | 78 | 78 sites potentially have the affected accessory | Distribution was made USA (nationwide) to medical facilities. Government distribution was made. Foreign distribution was made to Canada. There was no military distribution. | Open, Classified |  |  | 78 | Not_Computer
Z-1500-2017 | 76405 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | EyeSuite i.8.2.1.0 Software for ophthalmic use including selection of Intra Ocular Lenses (IOLs). |  |  | January 05, 2017 | 2017 | Haag-Streit USA Inc | There is a possibility for data to be stored under the wrong patient on the DICOM Server following a non-standard workflow. | Software design | On January 5, 2017 the firm sent Field Safety Notice to their customers. |  | 163 | 163 units | US Nationwide distribution. | Open, Classified |  |  | 163 | Not_Computer
Z-1492-2017 | 76683 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance BigBore Oncology CT - 728243 |  |  | February 23, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | When the user performs a retrospective respiratory gated 4D CT scan with the Bellows pulmonary gating device, the CT images might be reconstructed at a single phase, while the annotations on the images incorrectly indicate that the reconstruction is from the requested specific phases (0% to 90%). | Software design | Philips sent an Urgent Field Safety Notice dated February 24, 2017, to all affected consignee's via certified mail with returned receipts for tracking purposes. Philips Key Market representatives will distribute the notices internationally. A Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. Implementation of the software update will be monitored. Consignees requiring further information or support concerning this issue were instructed to contact their local Philips representative. For North America and Canada, contact the Customer Care Solutions Center at 1-800-722-9377. |  | 11 | 11 | Worldwide Distribution - US to AR, CA, CO, CT, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OH, TX &amp; WI; Internationally to the following countries: Australia, Austria, Canada, China, Czech Republic, Germany, India, Israel, Jordan, Korea, Mexico, New Cldonia, Romania, Russian Federation, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 78 | Computer
Z-1496-2017 | 76717 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software with Issuer of Patient ID (IPID). The firm name on the label is Merge Healthcare, Hartland, WI. |  |  | September 22, 2015 | 2015 | Merge Healthcare, Inc. | For sites using the Issuer of Patient ID (IPID), the system will display the study list and images from different patients with different IPIDs as though they are for the same patient if they all have the first name, last name, and medical record number in common. | Process design | Merge sent an Urgent: Medical Device Recall letter dated March 28, 2016, via email the week of April 4, 2016, or via certified mail if they did not have an email address for the customer. On/about September 22, 2015, the recalling firm emailed "IMPORTANT PRODUCT INFORMATION" letters dated August 20, 2015, to their customers. The letter notifies the customer of the issue and informs them a fix for the issue is in the new software release. The notification instructs the customer to pass the information onto all other users, as well as any downstream customers if it was further distributed. The notification also informs the customer their response is required using the enclosed form and the return addressed envelope. The response form asks, in part, if they are interested in accepting the fix. If they decline the fix, they are asked to provide the reason why. The response was required no later than April 18, 2016.Nonresponders were issued a second letter dated July 11, 2016, asking for a response by August 2, 2016. The letter and response form contained the same information as the first letter. If the second letter results in additional unresponsive customers, a call campaign will be managed by Customer Support and documented via SalesForce. The recalling firm reported emails will be tracked for delivery confirmation and the letters will be tracked for delivery by USPS certified services.For further questions, please call (877) 741-5369.. |  | 16 | 16 sites were originally identified as potentially having the affected versions | US Distribution was made to medical facilities located in GA, IL, MD, MI, MO, OH, TN, and VT. There was no foreign/government/military distribution. | Terminated | April 18, 2017 | 574 | 16 | Computer
Z-1683-2017 | 76725 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hand Foot II, UVB-138. One device per box; device is surrounded by foam inserts inside a corrugated box. For UV phototherapy treatment. |  |  | January 27, 2017 | 2017 | National Biological Corp | The wiring may allow the lamps to turn on with the key switch rather than the timer. | Employee error | Individual patients were contacted via telephone. Replacement devices have been distributed. The distributor was contacted via email. The firm has requested that consignees return the recalled device to their firm to be held under quarantine. |  | 7 | 7 units | Worldwide distribution. US states: IA, IL, MI, NY, PA, and WI. Country: Jordan.There are no government accounts for this recall.There are no Canada accounts for this recall.There are no Mexico accounts for this recall. | Open, Classified |  |  | 7 | Not_Computer
Z-1509-2017 | 76283 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBed |  |  | November 18, 2016 | 2016 | Ohmeda Medical | Certain power cords could overheat. The affected power cords were manufactured by Electric-Cord and are referred to as Taller Bridge power cords. The overheating of the power cord could result in fire, charring, smoke, or sparking where the power cord connects to the wall outlet which could cause thermal injury. | Under Investigation by firm | According to the firm, this issue will be resolved by sending an Urgent Medical Device Correction customer letter to customers with affected Power Cords. This letter will inform the customer of the issue and provide safety instructions for immediate mitigation of the issue. The long-term solution is to re-place all affected Power Cords at no cost to the customer.According to the firm, the plan for public notification is to contact all customers with affected Power Cords by means of an Urgent Medical Device Correction letter. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to the following titles within the affected accounts; Nurse Managers, Labor &amp; Delivery/NICU, Bio-Medical Engineering Department Managers and Risk Management DirectorsWe will track all service requests for the field action to completion. These service requests will record the replacement of the affected Power Cords. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures. |  | 5993 | 5,993 (5,080 US; 913 OUS) | Worldwide Distribution - US (nationwide and Puerto Rico) Bahamas and Internationally to UK, Trinidad &amp; Tobago, Taiwan, Spain, Panama, Netherlands Antilles, Mexico, Lebanon, Japan, Ireland, Honduras, El Salvador, Ecuador, Costa Rica, Colombia, Canada, Brazil, and Venezuela, | Open, Classified |  |  | 5993 | Not_Computer
Z-2566-2017 | 76629 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Segmental malefemale taper, 200mm, Sterile, |  |  | February 28, 2017 | 2017 | Zimmer Biomet, Inc. | Zimmer Biomet Inc. initiated a field action on 02/28/2017 for sterile-packaged implants packaged in two different package configurations due to packaging design verification test failures. | Packaging | Zimmer Biomet initiated a field action for sterile-packaged implants packaged in two different package configurations due to packaging design verification test failures. Zimmer Biomet sent letters via Fed-Ex and e-mail dated 02/28/2017 to customers informing them of the recall, and instructing them to do the following: 1. Review this notification and ensure affected team members are aware of the contents.2. Complete the Certification of Acknowledgement portion of Attachment 1a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three(3) days.3. Assist the Zimmer Biomet sales representative quarantine all affected product.4. If after reviewing this notice customers have further questions or concerns, please call 574-371-3071 between 8:00 am and 5:00pm EST, Monday through Friday. Calls received outside of the call center operating hours will receive a prompt to record a voicemail or be transferred to an oncall representative in the case of an emergency. Alternatively, questions may be sent byemail to corporatequality.postmarket@zimmerbiomet.com.Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA:Âż MedWatch Reporting: Adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online, by mail, or by fax.Âż Online: www.fda.gov/medwatch/report.htmÂż Mail: Use postage paid, pre-addressed form FDA 3500, available at: www.fda.gov/MedWatch/getforms.htmÂż Fax: 1-800-FDA-0178Under 21 CFR 803, manufacturers are required to report any serious injuries where a product has contributed or may have contributed to the event. Please keep Zimmer Biomet informed of any adverse events associated with this product or any other Zimmer Biomet product by emailing zimmer.per@zimmerbiomet.com.For questions regarding this recall call 574-371-3071. |  | 9867 | 9867 | Worldwide Distribution - US (nationwide) Foreign: Argentina, Australia, Belgium, Brazil, Bridgend Wales, Canada, Chile, Colombia, Dominican Republic, England, Eschbach, France, India, Italy, Japan, Jordan, Korea, Malaysia, Mexico, Netherlands, Nicaragua, Panama, Peru, Russia, Saudi Arabia, Singapore, Spain, Switzerland, Taiwan, Thailand | Open, Classified |  |  | 9867 | Not_Computer
Z-1768-2017 | 76723 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Bedside monitor CSM-1901(Life Scope G9) with main unit CU-192RA. The Problem only affects CSM-1901(Life Scope G9) communicating with CNS-6201A (PU-621RA) and CNS-9701A (MU-971RA). |  |  | March 13, 2017 | 2017 | Nihon Kohden America Inc | The Pause function on central monitors will not automatically resume when connected to a Life Scope G9 patient monitor. | Nonconforming Material/Component | A technical bulletin was issued on March 7, 2017 to make Life Scope G9 users aware of following interaction: the CNS 6201A and CNS 9701A (using the noted software versions) will not Auto Resume when a patient is reconnected to a Life Scope G9 if the patient tile is Paused. Therefore, each time a patient is reconnected to the Life Scope G9, the Pause must be cancelled and monitoring manually resumed on the CNS 6201A or CNS 9701A.A customer letter dated March 13, 2017 was sent to all customers to inform them that Nihon Kohden America (NKA) is voluntarily notifying them that the Pause function on central monitors (models CNS 6201A and CNS 9701A; Software Versions 01-06, 02-10 and 02-40), will not automatically resume when connected to a Life Scope G9 patient monitor (model CSM 1901; versions 01-19, 01-20 and 01-21). Customers are informed of issue,risk to health, and instructions. Customers are instructed to complete the attached reply form and return it to the fax number or email address on the form. Customers with questions are instructed to contact NKA Technical Services Group at (800) 325-0283, Option 6 (M-F, 8am-5pm PDT) or by email at G9@nkusa.com. |  | 654 | 654 units total (230 units in US) | Worldwide - US Nationwide Distribution including the countries of Germany, Spain, Brazil, Paraguay, Bolivia, Chile, Colombia, Costa Rica, Mexico, Peru, Pakistan, Israel, Qatar, Saudi Arabia, South Africa, Tunisia, United Arab Emirates, Korea, Malaysia, Singapore, Thailand | Open, Classified |  |  | 654 | Computer
Z-1801-2017 | 76727 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VIDAS Estradiol II, Ref 30431-01, 60 tests |  |  | March 17, 2017 | 2017 | BioMerieux SA | Fulvestrant cross reacts with the VIDAS Estradiol II immunoassay and may cause falsely elevated estradiol results in patients treated with Fulvestrant. | Device Design | bioMerieux sent an Important Product Safety Notice A Product Safety Notice to all affected customers on March accounts on March 17, 2017, informing them not to use VIDAS Estradiol II test to test estradiol level in patients under Fulvestrant therapy. The letter did not request return of any of the product, and requested that they return an acknowledgement form acknowledging receipt of the notice and that they have followed the instructions and implemented the actions in the field notice. Customers with questions were instructed to contact their bioMerieux Customers Service representative. For questions regarding this recall call 314-731-8805. |  | 834 | 834 kits | Nationwide Distribution including AZ, CA, CO, FL, GA, IN, IA, LA, MI, MS, NY, NC, OH, OK, OR, PA, TX, WA | Open, Classified |  |  | 834 | Not_Computer
Z-1487-2017 | 76687 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | MotoCLIP(TM) Super Elastic Fusion System, Size 18mm x 14mm x 14mm, REF 1442-1814, Rx Only, STERILE R. |  |  | February 16, 2017 | 2017 | Crossroads Extremity Systems Llc | Seven 18mm x 18mm x 18 mm implants had been placed in kits, packed and labeled as 18mm x 14mm x 14mm. | Labeling mix-ups | The firm notified their distributors by email on 02/16-17/2017. The distributors were asked to return the unused lot #101866. The firm also contacted the implanting surgeons either by visit or telephone about the potential for an incorrect size of the implant and asked if there were any adverse events. |  | 20 | 20 units | Al, FL, IN, Il, NE, TX, NM, CO, AZ | Terminated | April 27, 2017 | 70 | 20 | Not_Computer
Z-1494-2017 | 76376 | No | LZS |  |  | Ophthalmic | PMA | Final Fit Software Version 1.11 and 1.12; |  |  | February 01, 2017 | 2017 | Nidek Inc | During treatment planning, the procedure was programmed with an unintended (wrong) correction. | Software change control | Nidek sent a Recall notification letter dated February 1, 2017 to all affected consignees by certified mail method beginning February 2, 2017. Letters advised of the issue received from a complaint and the steps the user took that created the issue. To prevent the issue, the letters provide steps that should be taken. Nidek's Customer Support team will contact customers by phone or email. Customers with questions were instructed to contact Customer Service at 1-800-223-9044 or email ff_support@nidek.com. |  | 29 | 29 | Nationwide Distribution to MN, NV, OH, NC, CA, CO, CA, GA, VA, MI, AZ, PA, TN, WA, TX, NY. | Open, Classified |  |  | 29 | Not_Computer
Z-1490-2017 | 76712 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR" Avant-garde |  |  | February 14, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software update | Software design | Siemens mailed an Customer Information Letter on 02/15/2017 to affected customers to inform them of a software update they plan to implement to reduce risk of collision when delivering automatically sequenced treatments with automatic movements of the gantry and/or the treatment table, using the SIMTEC-Auto Field Sequence Option. |  | 80 | 80 systems | Distributed throughout the United States | Open, Classified |  |  | 80 | Computer
Z-1814-2017 | 76769 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | AESCULAP MINOP InVent 30 Trocar System, non-sterile |  |  | March 07, 2017 | 2017 | Aesculap Implant Systems LLC | Aesculap Implant Systems LLC is recalling the Minop Trocar due to the possibility it may have sharp edges on the distal end which may lead to the abrasion of the insulation when removing the electrode. | Nonconforming Material/Component | On March 15, 2017, 16 facilities and 1 Sales Rep were sent an Urgent Medical Device Recall Notification letter. Letters were sent Fed-Ex overnight. Customers were asked to immediately discontinue use and quarantine the product. A Sales Representative will remove the affected product and return to Aesculap Inc. |  | 21 | 21 units distributed in U.S. | Product was distributed throughout the United States and Canada. | Open, Classified |  |  | 21 | Not_Computer
Z-1517-2017 | 76719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Cardio software. |  |  | September 22, 2015 | 2015 | Merge Healthcare, Inc. | Users can merge a device import file with an image study that already has a confirmed report, which may result in including information not present at the time of physician interpretation. | Software design | On 9/22/2015, the recalling firm issued letters dated 9/8/2015 and flagged "PRODUCT INFORMATION" via email to their customers. The recalling firm also issued letters dated 3/28/2016 via email the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A letter dated 7/11/2016 was issued to nonresponders. |  | 89 | 89 sites potentially have the affected software | US Nationwide Distribution. | Open, Classified |  |  | 89 | Not_Computer
Z-1520-2017 | 75353 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris System PC unit, model no. 8000 and 8015. The central programming, monitoring and power supply component for the Alaris infusion pump System. |  |  | November 01, 2016 | 2016 | CareFusion 303, Inc. | Reports where the Low Battery alarm and/or the Very Low Battery alarm are not being triggered before the battery is discharged and all infusion channels are stopped. | Use error | A customer letter dated 11/1/2016 will be sent to all customers to inform them that BD has identified an issue with the Alaris System PC unit. The letter informs the customers that BD has received reports for the Alaris PC units where the Low Battery alarm and/or the Very Low Battery alarm are not being triggered before the battery is discharged and all infusion channels are stopped. The letter informs the customers of the issue and recommended steps to take. Customers with recall related questions are instructed to contact BD Support Center at (888) 562-6018, 7-4pm PT, M-F, supportcenter@carefusion.com. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266, 24hrs/day, 7days/wk, customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229, 6am-5pm PT, M-F, DL-US-INF-TechSupport@carefusion.com. |  | 613800 | 613,800 total units (575,221 units in US) | Worldwide distribution. US nationwide, Europe, Australia, New Zealand, South Africa, Greater Asia, Middle East, and Canada. | Open, Classified |  |  | 613800 | Not_Computer
Z-1518-2017 | 76422 | No | LQE |  |  | Ophthalmic | PMA | Raindrop Near Vision Inlay, Model# RD1-1 |  |  | February 07, 2017 | 2017 | Revision Optics Inc | ReVision Optics has updated the instructions for use (IFU) for the Raindrop Near Vision Inlay to emphasize that only Balanced Salt Solution (BSS) may be used to irrigate under the flap and that topical medications or lubricants should not be administered until centration of the inlay and proper flap positioning has been confirmed at the slit lamp. | Use error | A Dear Doctor letter was sent to inform customers that ReVision Optics wantedto share a surgical update regarding the Raindrop Near Vision Inlay procedure.Customers are informed that no ophthalmic medications or lubricants should be used immediately prior to or during placement of the Raindrop Near Vision Inlay. Only Balanced Salt Solution (BSS) should be used to irrigate under the flap during the Raindrop procedure. Customers are informed that ophthalmic medications or lubricants should be administered only after verifying proper inlay centration and flap positioning as observed at the slit lamp. This ensures that no further inlay or flap manipulation is required. ReVision Optics has updated the instructions for use (IFU) for the Raindrop Near Vision Inlay to emphasize that only BSS may be used to irrigate under the flap and that topical medications or lubricants should not be administered until centration of the inlay and proper flap positioning has been confirmed at the slit lamp. Customers with questions are instructed to contact Luis Vargas regarding the surgical procedure update. |  | 1279 | 1,279 devices | US: AL, AZ, CA, FL, GA, HI, IL, IN, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, PA, SD, TN, TX, UT, WA | Terminated | June 21, 2017 | 134 | 1279 | Not_Computer
Z-1521-2017 | 76201 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | ETEST Ceftazidime TZ256 Foam packaging, Product Name: ETEST Ceftazidime TZ 0.016-256. |  |  | December 15, 2016 | 2016 | BioMerieux SA | Product Stability issues: The current shelf-life claims of the ETEST products are not supported by internal testing. | Storage | biomerieux sent an Urgent Field Safety Notice dated December 15, 2016, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to identify the affected lots and immediately order the replacement products. Laboratories should continue to follow their current QC procedures. When replacement product is received, consignees should discontinue using any discard affected product. Consignees should distribute this information to all appropriate personnel in their laboratory, retain a copy in their files, and forward the information to all parties that may use this product, including others to whom they may have transferred our product. Consignees should contact their bioMerieux representative for product compensation. Consignees should also complete and return the Acknowledgment Form in Attachment A by Fax to confirm receipt of the notice. Consignees with questions should contact their local bioMerieux Clinical Customer Service representative.For questions regarding this recall call 314-731-8516. |  |  | unknown | Worldwide Distribution - U.S.( Nationwide); Foreign: Utd.Arab Emir., Argentina, Austria, Australia, Brazil, Canada, Switzerland, Chile, Colombia, Czech Republic, Germany, Algeria, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Japan, South Korea, Mexico, Netherlands, Poland, Portugal, Russian, Fed. Sweden, Singapore, Thailand, Turkey, South Africa, Indonesia, Peru, Saudi Arabia, Gabon, Kazakhstan, Philippines, Ivory Coast, Venezuela, Kenya, Serbia, Tunisia, Macedonia, Lebanon, Romania, Benin, Bosnia-Herz., Oman, Uruguay,Lithuania, Burkina Faso, Slovenia, Qatar, Kuwait, Costa Rica, Paraguay, Dominican Republic, Croatia, Panama, Congo, Dem. Rep. Congo, Honduras, Estonia, Latvia, Cambodia, Bulgaria, Vietnam, Dutch Antilles, Nicaragua, Israel, Georgia, Palestine, Egypt, Pakistan (includ.Afghanistan), Malta, Jordan Bahrain, El Salvador, Nigeria, Dutch Antilles, Iran, Belarus, Zimbabwe, Morocco, Mauretania, Anguilla, Philippines, Malaysia, Turkmenistan, Guam, Brunei, Bolivia, Guatemala, Ecuador, Belarus, Mongolia | Open, Classified |  |  |  | Not_Computer
Z-1513-2017 | 76694 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Poly Component Trial, MPCS, packaged in the following sizes and configurations: |  |  | February 20, 2017 | 2017 | Medtronic Sofamor Danek USA Inc | The dovetails of poly trials, Beta 2.0, were observed as either cracked or broken. | Under Investigation by firm | Medtronic initiated their recall on 20 February 2017, by letter delivered by Fed Ex . The letter was addressed to the Risk/Materials Manager and stated that a Medtronic Clinical Specialist would coordinate the removal of the Responsive Orthopedics Total Knee Arthroplasty (TKA) System instrument sets from the premises. Customers with questions were instructed to contact the Medtronic Post-Market Quality Team at 901-399-3197. |  | 72 | 72 units | US distribution to AR, MA | Open, Classified |  |  | 183 | Not_Computer
Z-1519-2017 | 76263 | No | BTQ | Nasopharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | Teleflex Green Non-Latex Robertazzi Nasopharyngeal Airway, Catalog No. 123130 and 123132 |  |  | January 13, 2017 | 2017 | Teleflex Medical | Labeling error | Unknown/Undetermined by firm | Consignees were notified of the recall by letter on 1/13/2017. The letter requested that they immediately discontinue use and quarantine the recalled product. Distributors were requested to conduct a sub-recall. The letter included a reply form which was to be returned to Teleflex. |  | 7600 | 7,600 eaches | Worldwide Distribution - US Nationwide in the state of AZ, CA, CO, CT, FL, GA, IL, IN, KY, LA, MD, MA, MI, MN, MS, MO, NV, NJ, NM, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI, and the countries of Canada and Mexico. | Open, Classified |  |  | 7600 | Not_Computer
Z-1535-2017 | 76365 | No | ORZ | Eyelid thermal pulsation system. | Ophthalmic | Ophthalmic | 510(k) | LipiFlow Thermal Pulsation System/LipiFlow, System Activator (Disposable), Rx Only, Sterile |  |  | January 30, 2017 | 2017 | Tearscience, Inc | Labeling: The labeled storage temperature is not consistent with the temperature conditions tested for the entire four year shelf life duration. | Labeling Change Control | Tearscience sent an MEDICAL DEVICE LABELING CORRECTION/RECALL letter dated January 30, 2017..Actions to be Taken by Customer: Within 10 business days of receipt of this letter, please take the following actions: 1) Read the enclosed Summary of Important Labeling Changes for the LipiFlowÂż System. 2) Complete the Acknowledgement and Receipt form at www.TSVerify.pmcprograms.com or return the enclosed completed form by email or mail. 3) If you have stored Activators in an environment that is not temperature controlled (without heat and air conditioning), discontinue use of these Activators and call TearScience Labeling Change Hotline at (844) 872-5658.For any questions regarding this letter, please call (844) 872-5658 Monday through Friday, 8:00AM to 6:00PM, Eastern Time. We want to thank you for being a valuable TearScience partner, and as always would like to thank you for your business, your trust, and your commitment to treating MGD.For further questions please call (919) 459-4831. |  | 94095 | 94,095 units. | US (nationwide) | Open, Classified |  |  | 94095 | Not_Computer
Z-1542-2017 | 76373 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon MRI System |  |  | April 16, 2015 | 2015 | Hitachi Medical Systems America Inc | Image data transferred from the MRI system to a workstation showed errors on the slice position reference image. | Software in the Use Environment | Formal notification letters were sent to consignees on 2/13/17. Hitachi Service will update the software at each site during visits. |  | 68 | 68 units | US nationwide distribution. The products were distributed to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY.There are four (4) government accounts for this recall: VA Medical Center/Lyster Army Health Clinic, Building 301 - Andrews Avenue, Fort Rucker, AL 36362; Parma VA Medical Center, 8787 Brookpark Rd, Parma, OH 44129; Viera VA Outpatient Clinic, 2900 Veterans Way, Viera, FL 32940; Perry Point VA Medical Center, Avenue D, Suite 1A-120, Perry Point, MD 21902.There are zero (0) Canada accounts for this recall.There are two (2) Mexico accounts for this recall: Hospital General Tipo "B" de Alta Especialidad - SMH, "Fraccionamiento 1 De La Parcela 101, Del Ejido De Atapaneo", Morelia, MX 58300; Hospital General de Occidente, Zoquipan 1050 Avenue, Zapopan Jalisco, MX 45170The product was distributed to the following countries: Brazil. | Terminated | June 29, 2017 | 805 | 303 | Computer
Z-1781-2017 | 76002 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Imager II Angiographic Catheter, M001316581 (5-pack outer package UPN), M001316580, single unit, inner package UPN) |  |  | December 13, 2016 | 2016 | Boston Scientific Corporation | One lot of Imager II Angiographic Catheters is being recalled due to packaging mislabeling. The product may be labeled as Imager II Contra 2 curve catheters, when in fact the package contains Contra curve catheters. | Employee error | Customers were sent an Urgent Medical Device Removal letter, dated 12/5/2016, via overnight mail delivery beginning 12/15/2016. The letter identified affected product, described the issue, and asked for product to be returned. Removal instructions were provided and a Reply Verification Tracking Form was asked to be returned. Questions can be directed to Nicole Pshon, Field Action Manager at 763-494-1133 or BSCFieldActionCenter@bsci.com |  | 130 | 26 boxes of 5 (130 singles) | AZ, FL, MO, NJ, NY, OH, PA, and TX. | Open, Classified |  |  | 130 | Not_Computer
Z-1695-2017 | 76696 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 |  |  | March 09, 2017 | 2017 | Ion Beam Applications S.A. | The Proton Therapy System can allow gantry rotation even if the snout is not properly locked. | Component design/selection | A Field Safety Notice was mailed or handed to affected customers to inform them of the issue. The notice provides information on the issue, the risk for the patient and user and actions to be taken before starting gantry rotation. Corrective action developed by IBA are detailed. |  | 12 | 8 worldwide, 4 in U.S. | Distributed in MA, FL, VA, and TN. China, South Korea, France and Germany | Open, Classified |  |  | 12 | Not_Computer
Z-1764-2017 | 76628 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas 8100 uni-directional reformatter (BRF) module with Software Version 02-xx |  |  | March 01, 2017 | 2017 | Roche Diagnostics Corporation | Roche Diagnostics Corp. initiated a voluntary correction because a rack crash may occur on the Cobas 8100 bi-directional reformatter (BRF) and uni-directionalreformatter (BRF) modules with software version 02-xx, only when rack buffering is activated and the rack buffer is completely full. This issue can lead to sample spillage, posing a potential risk to operators / laboratory staff due to exposure to potentially infectious material, cross-contamination of samples in the affected racks and erroneous results due to sample carry-over. | Under Investigation by firm | Roche initiated a recall on 03/01/2017, notices were issued via UPS to customers and instructed customers to do the following:Roche is developing software version 03-01 to resolve this issue, and a Roche Representative will be contacting you in the near future to schedule an update. Until this software update is available, the rack buffering option for the BRF and URF modules must be disabled. If you are an authorized cobas 8100 automated workflow series operator, follow the steps outlined below to disable rack buffering. If your facility does not have an authorized cobas 8100 automated workflow series operator, please contact the Roche Support Network Customer Support Center at 1-800-428-2336 for further steps.If you have a major spill, set the affected module to offline and contact the Roche Support NetworkCustomer Support Center at 1-800-428-2336." If you are an authorized cobasÂż 8100 automated workflow series operator, follow the steps outlined inthis Urgent Medical Device Correction (UMDC) to disable rack buffering." If your facility does not have an authorized cobas 8100 automated workflow series operator, pleasecontact the Roche Support Network Customer Support Center at 1-800-428-2336 for further steps." Complete the attached faxback form and fax it to 1-888-345-0024." File this Urgent Medical Device Correction (UMDC) for future reference.Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 45 | 45 | Domestic: MA, IN, AL, OH, MO, NE, LA, CA, IA, NJ, AR, PA, SC, IL, TX, and MI.Foreign: None | Open, Classified |  |  | 45 | Computer
Z-1539-2017 | 76693 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Panta Nail, Rx only, Sterile, |  |  | March 03, 2017 | 2017 | Integra LifeSciences Corp. | Voluntary recall/Removal of Panta and Panta XI Nails because there may be the potential that the internal and external package seal does not provide the level of sterility intended. | Packaging change control | Integra sent "Urgent Medical Device Recall Letters and Return Response Forms" dated March 3, 2017 to their distributors and customers.All customers were asked to examine their inventory to determine if they had the recalled lots and if they did they are instructed to stop using them immediately. Distributors are asked if the recalled product was further distributed, those customers should be notified and instructed to follow the recalling procedures. The return response form should be completed, even if they do not have any recalled lots, annotating the amount of product and list a record of the lot numbers. The return response form should be sent back to Integra via email to: FCA2@integralife.com or via fax to: 1-609-750-4220. Upon receipt by Integra, if your company has identified recalled lots, Customer Service will contact the customer and provide an RMA number and directions to return the product, as well as input an order to replace the quantity indicated on the return response form. Should any customer have any questions regarding the recall instructions, they can contact Customer Service at 1-800-654-2873. To return units under field inventory, customer can contact 1-888-601-0203, option 2.For further questions, please call (717) 309-3966. |  | 4449 | 4,449 units | US Nationwide and Internationally | Open, Classified |  |  | 4449 | Not_Computer
Z-1769-2017 | 76518 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Custom Tubing Pack with Bioline Coating |  |  | February 16, 2017 | 2017 | Datascope Corporation | A complaint was received on January 1, 2017 stating the pigtail on line 2 of the oxygenator was assembled to the port where the yellow cap and the clear cap remained on the center port. | Device Design | Maquet Inc. was notified by their customer on January 1, 2017. This was a reverse notification. The one impacted customer said they will not return the product since they can easily correct the pigtail connection before use.For further questions regarding this recall, please call (973) 709-7442. |  | 21 | 21 kits | US Distribution to the state of : CA. | Open, Classified |  |  | 21 | Not_Computer
Z-1682-2017 | 76253 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum 360 Infusion Pump, Software Version 15.02. The infusion pump is capable of delivering fluids for a variety of therapies such as parenteral, enteral, or epidural infusions |  |  | December 30, 2016 | 2016 | Hospira Inc., A Pfizer Company | Depleted Battery alarm shows instead of Replace Battery. On battery power, ongoing therapy stops, alarms show and sound, pump shuts down after 3 minutes. On AC power, the pump reboots (takes about 25 seconds). Therapy can resume. AC power interruption of 3-7 seconds prompts an incorrect E323 alarm, ongoing therapy stop, and pump reboot. Longer power loss results in battery power operation. | Software in the Use Environment | On 12/30/2016, firm sent notification of the recall to Hospira direct consignees of record via traceable mailing. Delivery will be confirmed via traceable mailing. Consignees who further distributed the product were asked to notify their customers. Effectiveness of the communication will be confirmed by phone call with direct consignees who do not respond. The final disposition of the recalled product will be field correction (firm will contact consignees to upgrade the software in their Plum 360 infusion pumps). |  | 62847 | 62,847 units | Nationwide, Australia, New Zealand, Canada, Mexico, Chile, Philippines, Hong Kong, Greece, Ireland, Jordan, Kuwait, Oman, Poland, Saudi Arabia, Turkey, United Arab Emirates, Italy, France, United Kingdom, Spain | Open, Classified |  |  | 62847 | Computer
Z-1671-2017 | 76255 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Regard, Item Number: 800430004, Sterile, LD0499D - C Section Pk - Cabrini |  |  | January 09, 2017 | 2017 | Resource Optimization &amp; Innovation Llc | Custom procedure trays contain a component, a light glove, that has been recalled by the manufacturer because it can split during application to a light handle breaking the sterile field. | Nonconforming Material/Component | ROi sent an Urgent Recall Notice dated January 9, 2017, to all affected consignees. The recalling firm instructed consignees to quarantine affected product and complete the response form. The recalling firm stated that they will issue over lableing for the affected packs. Sterile replacement light covers will be sent directly to the end user if available. Consignees were instructed to inform their subaccounts of the recall. For questions regarding this recall call 417-820-2793. |  | 454 | 454 kits | US Distribution to LA and TX. | Open, Classified |  |  | 39233 | Not_Computer
Z-1497-2017 | 76596 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 7F PreludeÂż Short Sheath Introducer. Catalog Numbers: K15-00070, K15-00170, PSS-7F-4-035MT, PSS-7F-4MT. |  |  | February 16, 2017 | 2017 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for sheaths included with the 7F PreludeÂż Short Sheath Introducers because the sheath marker tips may detach during use. | Process control | Consignees notified on 2/16/17 to immediately stop using or distributing the affected lots and return them to Merit. For any questions, please contact your Merit Sales Representative or Merit Customer Service at 801.208.4381. |  | 1265 | 1,265 total | Nationwide and other foreign consignees. No Canadian or VA/govt/military consignees. | Terminated | August 08, 2017 | 173 | 1265 | Not_Computer
Z-1679-2017 | 76760 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Non-Sterile Double DIN to DIN "Y" Connector, |  |  | January 20, 2015 | 2015 | Molded Products Inc | A customer relayed a complaint of a leaking dialysis tubing connector. | Process control | On 1/20/2015, the recalling firm requested the firm's customer to return units from the affected lot. |  | 400 | 400 pieces | Product was distributed to Ontario, Canada. | Terminated | March 30, 2017 | 800 | 400 | Not_Computer
Z-1471-2017 | 76438 | No | JIH | Lithium test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products Li (Lithium) Slides, REF/Catalog Number/Product Code 1632660 (18 slide format, 90 slides per sales unit) and REF/Catalog Number/Product Code 8318925 (60 slide format / 300 slides per sales unit), Unique Identifier Nos. 10758750009466 and 10758750004720, RX ONLY, IVD --Ortho Clinical Diagnostics, Inc. Rochester, NY 14626 --- Background: |  |  | February 01, 2017 | 2017 | Ortho-Clinical Diagnostics | Intermittent imprecision in results generated using two specific lots of VITROS Li Slides. | Nonconforming Material/Component | On 2/01/ 2017, URGENT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2017-036, dated 2/01/2017) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received the affected VITROS Chemistry Products Li Slides to inform them of the issue and to request that they immediately discontinue use of this product. The correction is to the user level. Discuss any concerns regarding previously reported VITROS Li results obtained from the affected with their Laboratory's Medical Director to determine the appropriate course of action. Ortho will be shipping replacement product. ---Foreign affiliates were informed of the issue by e-mail on 2/01/2017 and instructed to notify their consignees of the affected product with this issue and the requirement to immediately discontinue use of the product. For questions, please contact our Ortho Care(TM) Technical Solutions Center at 1-800-421-3311. |  | 230 | US: 94 sales units; OUS: 136 sales units | US, Puerto Rico, Canada, EU, China, Brazil, Chile, and Dominican Republic. | Open, Classified |  |  | 230 | Not_Computer
Z-1417-2017 | 76545 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator, |  |  | February 24, 2017 | 2017 | Philips Electronics North America Corporation | HeartStart MRx Monitor/Defibrillator may may exhibit abnormal behavior when operating using asingle battery as the sole source of power may fail to deliver therapy or monitoring | Device Design | Philips Healthcare issued an Urgent Medical Device Correction notification/Field Safety Notice. on 2/24/17. The Urgent Medical Device Correction notification/ Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Action (1) Philips is voluntarily initiating this correction and is providing Instructions for Use (IFU) Addendum titled, HeartStart MRx Battery Connector Pins for all customers with affected units. Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:Philips is asking the customer when inspecting the device, to inspect the HeartStart MRx battery compartment connections. Include this step as part of a routine Operational Check to verify that the battery connector pins are clean, fully extended, not bent, and without residue.Action (2) Philips is voluntarily initiating this correction and is providing measures to prevent, or methods to allow the user to correct, the abnormal device behavior related to charging and pacing after disconnecting the device from AC mains. |  | 97794 | 97,794 total ( 47,362 US and 50,429 OUS) | Worldwide Distribution-US (Nationwide) including PUERTO RICO and countries of: AFGHANISTAN, ALBANIA, ANGOLA, ARGENTINA, ARUBA, AUSTRALIA,AUSTRIA,AZERBAIJAN,BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA,BOSNIA AND HERZEGOWINA, BRAZIL,BRUNEI DARUSSALAM, BULGARIA,CAMBODIA,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, CROATIA (local name: Hrvatska),CYPRUS,CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC,ECUADOR,EGYPT,EL SALVADOR, ESTONIA, ETHIOPIA,FINLAND,FRANCE,GABON,GERMANY,GHANA,GREECE, GUADELOUPE, HONG KONG,HUNGARY,ICELAND,INDIA, INDONESIA,IRAN (ISLAMIC REPUBLIC OF),IRAQ,IRELAND,ISRAEL,ITALY,JAPAN, JORDAN, KAZAKHSTAN, KENYA,KOREA, REPUBLIC OF,KUWAIT,LAOS,LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA,LITHUANIA,LUXEMBOURG, MACAU, MACEDONIA,MALAYSIA,MALTA,MAURITIUS,MEXICO,MONGOLIA,MOROCCO, MOZAMBIQUE,MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES,NEW CALEDONIA,NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA,PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL,QATAR,REUNION,ROMANIA,RUSSIAN FEDERATION,SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA,SOUTH AFRICA,SPAIN,SRI LANKA, SWEDEN,SWITZERLAND,TAIWAN, PROVINCE OF CHINA,TANZANIA, UNITED REPUBLIC OF,THAILAND,TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN,UGANDA,UKRAINE,UNITED ARAB EMIRATES,UNITED KINGDOM,UZBEKISTAN,VENEZUELA,VIET NAM,YEMEN. | Open, Classified |  |  | 97794 | Not_Computer
Z-1755-2017 | 76820 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | MEDTECH ROSA Spine 1.0.2 |  |  | February 27, 2017 | 2017 | Zimmer Biomet, Inc. | Unapproved change made by the supplier. | Component change control | Zimmer Biomet sent an Urgent Medical Device Recall Correction letter on February 27, 2017,to all affected customers informing them of the following:Description of the issue:A robot arm file parameter linked to arm calibration was modified by the robot arm supplier without notification to MEDTECH.As a result, once the robot arm is on trajectory, an error leading to a device reboot may decrease the devices applicative accuracy, one of the essential performance requirements, to a level above the specification. This may happen after a shutdown of the device, (voluntary or due to an error), without a return of the robot arm position to a Parking position.In the worst case scenario, this may cause an incorrect positioning of the instruments by the ROSA devices. If this is not detected by the user this could ultimately lead to ineffective treatment, serious injury or even death of the patient.MEDTECH did not record, in normal conditions of use, any customer reports linked to this issue.Required actions by users:Pending the implementation of a permanent correction, MEDTECH is recommending the following actions to all users of the ROSA device:1- After a shutdown of the device, (voluntary or due to an error), without a return of the robot arm in Parking position, ROSA device must be switched off with the I/O switch button on the device rear panel.2- Restart the device after 10 seconds, and after full extinction of emergency button light. If the Reset button at the device rear panel is red, restart the ROSA PC by pushing the Reset button.3- If robot position is NOT Âż Home Âż, load patient folder and main software interface will appear on the screen.a. Go to the Guidance menu, select trajectories and follow instructions.b. Click on Start button.c. Start recovery procedure and follow all instructions displayed on the screen until robot arm position is Home .d. Close the patient foldere. Click on Exit and follow all ins |  | 1 | 1 | Nationwide Distribution to CA, FL, IL, KY, MA, MN, NC, NY, OH, PA, and District of Columbia. | Open, Classified |  |  | 34 | Computer
Z-1731-2017 | 76730 | No | OXQ | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | OPHTHALMIC DRESSING TRAY |  |  | March 09, 2017 | 2017 | Centurion Medical Products Corporation | Centurion Medical Products is recalling Convenience Kits that contain Covidien Curity Oval Eye Pads due to supplier's recall because of potential breach in sterile packaging. | Nonconforming Material/Component | On March 9, 2017 Centurion Medical Products recalled specific lots of medical convenience kits due to a supplier initiated recall of Covidien Curity Eye Pads. All affected customers were notified, advised to discontinue use of associated kits, and requested to confirm the status of remaining stock to Centurion. |  | 210 | 210 kits | US Distribution: NY, NC, PA, CA, OH, WA, VA, NJ, IL | Terminated | August 07, 2017 | 151 | 7342 | Not_Computer
Z-1685-2017 | 76619 | No | LHL | Haemophilus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Legionella IFA (IF0950) intended for qualitatively detecting and semi-quantitating human serum antibodies to Legionella pneumophila. |  |  | February 15, 2017 | 2017 | Focus Diagnostics Inc | After a submission for CLIA database update from Focus Diagnostics to DiaSorin Molecular,the company found that their Legionella IFA device (IF0950) is not 510(k) exempt. | No Marketing Application | Focus Diagnostics/DiaSorin Molecular is issuing a medical device correction for the Legionella IFA (Immuno Fluoresence Assay). On February 1, 2017, the firm was notified that the Legionella IFA was not properly classified according to 21 CFR Part 807. The product is being withdrawn from the market. Use of the product should be immediately discontinued, and any remaining inventory should be immediately destroyed. The firm sent notice to its consignees, which should be reviewed with the facility's laboratory or mediclal director. Patient results obtained with all lots of the affected product should be evaluated by the clinician and if recommended, retested by another method. |  | 196 | 196 units | Product was distributed to three testing laboratories in California, New Jersey and Virginia | Terminated | April 13, 2017 | 57 | 1596 | Not_Computer
Z-1455-2017 | 76697 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | XTRAC Excimer Laser, Model No. AL8000, AL10000 |  |  | January 23, 2017 | 2017 | Strata Skin Sciences, Inc. | When the Liquid Guide ( LLG) is not used according to the operator manual, the switch control wires may become loose and shorted. Shorted wires may lead to unintentional laser firing. | Radiation Control for Health and Safety Act | STRATA Skin Sciences, Inc., planned action to bring defect into compliance: The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 1000 domestic units and 612 international units will consist of the following:1. Send Field Correction notification to all consignees by January 30, 2017,2. Verify all consignees received the notification by February 10, 2017,3. Ship all Rear Panel Boards to consignees by May 1, 2017, and4. All corrective actions will be performed at no cost to the purchasersFor further questions call Customer service phone no.: (877) 449-8722 |  | 1612 | US - 1000 , Foreign - 612 | USA (nationwide) Distribution | Terminated | June 05, 2017 | 133 | 1612 | Not_Computer
Z-1688-2017 | 76700 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.x, Picture archiving and communication system |  |  | March 06, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software changes now available to address several issues | Software design | Siemens sent a notice dated March 7, 2017, to affected customers on to inform them that software updates via the latest syngo.via software was now available to that would address several issues that were reported. Customers were asked to forward the notice to the new owner. For questions regarding this recall call 610-448-6481. |  | 104 | 104 systems | Nationwide Distribution. | Terminated | June 12, 2017 | 98 | 104 | Computer
Z-1687-2017 | 76377 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Philips V60 Ventilator with Version 2.20 Software, |  |  | February 01, 2017 | 2017 | Respironics California Inc | The V60 Ventilator with Version 2.20 software installed may falsely detect that the blower motor has stalled. If this condition occurs, the software will cause the ventilator to shut down (Vent Inop) and display Error Code 100E. Ventilatory support will cease. | Device Design | Philips sent an Urgent- Medical Device Letter dated February 1, 2017, to all affected customers. It will be sent the week of February 6, 2017, to inform customers that Respironics is voluntarily recalling all Philips V60 Non Invasive Ventilators (NIV) that have Version 2.20 Software installed. Customers are informed that the V60 Ventilator with Version 2.20 software installed may falsely detect that the blower motor has stalled. If this condition occurs, the software will cause the ventilator to shut down (Vent Inop) and display Error Code 100E. Ventilatory support will cease. An audible high-priority alarm will sound continuously for at least 2 minutes when the V60 shuts down for any Vent Inop condition and is operating on battery power. If the V60 is connected to AC power (mains supply), the alarm will continue to sound until an operator intervenes. If the V60 is connected to a remote alarm system, the alarm system will be activated until action is taken by the operator. Error Code 100E only exists in V60 Version 2.20 software. Therefore, V60s running Version 2.10 software are not subject to this particular Vent Inop condition. E-mail back the complete form attached to the rear of this recall notification to recall. response @ Philips. com or Fax to 1-877 499-7223. For further questions please call (760) 918-7300. |  | 2512 | 2512 units | Worldwide Distribution - US Distribution and to the countries of : Argentina, Austria, Belgium, Canada, Czech Republic, Finland, France, Germany, Hong Kong, Iceland, India, Indonesia, Italy, Japan, Kuwait, Malaysia, Mexico, Netherlands, Norway, Qatar, Saudi Arabia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 2512 | Computer
Z-1863-2017 | 76507 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | various polyethylene implants |  |  | February 17, 2017 | 2017 | Zimmer Biomet, Inc. | Endotoxin levels higher than process maximum limits were discovered in the identified polyethylene components. | Process design | On February 17, 2017 URGENT MEDICAL DEVICE RECALL (REMOVAL) - LOT SPECIFIC notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Distributors, Sales Representatives, and Distributor Operation Managers Your Responsibilities: 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet within 2 weeks. a. Request a Return Authorization Number via email to rgarequest@zimmerbiomet.com or through FAST/SMS. Specify RECALL as the RGA type. b. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com. c. Include a hardcopy of Attachment 1 with your shipment for immediate processing. d. Clearly mark RECALL on the outside of return boxes. 5. Hospitals that received direct shipments of the affected products from Zimmer Biomet will be sent a copy of the Risk Manager Recall Notice directly. Surgeons will also be notified. It is important that you review the list of hospitals and surgeons included with the email notification sent to your facility to identify additional accounts Zimmer Biomet has not notified. Using the Additional Accounts Form provided with the email notice, return contact information for any additional hospitals and surgeons that may have received or used the affected product. If there are no additional |  | 28253 | 28,253 in total | Worldwide - US Nationwide in the ARGENTINA AUSTRALIA CANADACHILE CHINA COSTA RICA ECUADOR GSCCINDIA ITALY JAPAN KOREA, THE REP OF (SOUTH)MALAYSIAMEXICO PANAMA SingaporeSWITZERLANDTHAILANDUK | Open, Classified |  |  | 28253 | Not_Computer
Z-1739-2017 | 76606 | No | HMP | Eye pad. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Covidien Curity Eye Pad |  |  | March 03, 2017 | 2017 | Covidien LLC | Product sterility is compromised due to breach of sterile barrier | Packaging process control | Medtronic/Covidien) initiated recall by letter on March 3, 2017 via Federal Express. The letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Contact Medtronic representative or Customer Service at (800) 882-5878For questions regarding this recall call 203-492-5000. |  | 7168000 | 7168000 | Worldwide Distribution. US (Nationwide), Internationally to Canada, Australia, Brunei, Japan, Mexico, Philippines, Singapore, Andorra, Belgium, Botswana, Czech Republic, Denmark, Egypt, France, Greece, Hungary, Ireland, Israel, Italy, Kenya, Libya, Luxembourg, Mauritius, Namibia, Netherlands, Norway, Poland, Portugal, Russia, Russian Federation, Spain, Saudi Arabia, Seychelles, Slovakia, South Africa, Switzerland, UAE, and United Kingdom. | Open, Classified |  |  | 70254852 | Not_Computer
Z-1691-2017 | 76610 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Altius M-INI OCT Posterior Spinal Fixation System, Model Number 1200-9216 |  |  | March 16, 2017 | 2017 | Zimmer Biomet Spine, Inc | Zimmer Biomet is recalling the Altius M-INI OCT Posterior Spinal Fixation System because the driver is unable to mate with the screw head tulip. | Device Design | Zimmer Biomet sent an Urgent Medical Device Recall letter on March 20, 2017, to all affected consignees. The letter identified the product, the problem and the action to be taken by the consignee. Consignee's Responsibilities1. Review this notification and ensure affected team members are aware of the contents.2. Prior to sending Altius Kits to central processing, perform a functional check of the driver withall screws in the kit to ensure they mate. If the drivers do not mate, please return inaccordance with Step 3b.3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory ReturnCertification Form  Phase I. b. If the driver is unable to mate with the screws in the set, immediately sterilize and return theaffected nonconforming product from your distributorship and affected hospitals within yourterritory along with a completed Attachment 1  Inventory Return Certification Form Phase I and Attachment 2  Certificate of Sterilization to Zimmer Biomet.i. For each return, send a copy of Attachment 1 tojeremy.hansmann@zimmerbiomet.comii. Include a hardcopy of Attachment 1 with your shipment for immediateprocessing.iii. Include a copy of Attachment 2  Certificate of Sterilization with returnedinstruments.iv. Mark the outside of the returns box(es) clearly with RECALL. 4. Distributors will notify their accounts to which the product was further distributed and to locateand return products from these accounts. Using the Additional Accounts Form provided withthe email notice sent to your facility, return contact information for any additional hospitals thathave the affected product. If there are no additional users to notify, please confirm this byannotating and returning the form provided indicating that no additional users have beenidentified.5. Retain a copy of your field action acknowledgement and product return forms for your records inthe event of a compliance audit of your facility.6. If |  | 141 | 141 | Worldwide Distribution - US (nationwide) and Internationally to the Netherlands | Open, Classified |  |  | 141 | Not_Computer
Z-1690-2017 | 75424 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | 8500M Handheld Pulse Oximeter |  |  | September 29, 2016 | 2016 | Nonin Medical, Inc | Nonin manufactures and sells two 8500 handheld pulse oximeters (Model 8500 with no memory feature and the 8500M which has a memory feature). Some of the devices distributed may not have the correctpulse oximetry board with the integrated memory feature. | Process control | Consignees were called and sent on 9-29-2016 a Nonin Urgent Medical Device Recal" letter dated 29 September, 2016. The letter described the problem and the product involved in the recall. Additionally, it described the Potential Risk and Required Action. Requested consignees to return the affected units and to forward the contact information if the unit is no longer in their inventory. For questions contact Kim.Aves@nonin.com or via direct telephone at 763 577 3196. |  | 15 | 15 (11 US; 4 OUS) | Distribution in US (MI, NC, OH, NC, VA), and CANADA | Open, Classified |  |  | 15 | Not_Computer
Z-1898-2017 | 76821 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Wingman 35 Crossing Catheter, 65cm, |  |  | March 22, 2017 | 2017 | ReFlow Medical | ReFlow Medical is recalling the Wingman 35 Crossing Catheter because it may be prone to tip detachment. | Under Investigation by firm | An Urgent Medical Device Recall letter was issued 3/22/17 for the Wingman 35 Crossing Catheter. The recall was issued due to a risk of catheter tip detachment which could cause injury to vasculature and thromboembolic events. Customers are asked to complete the Response Form and reply via email (quality@reflowmedical.com) or fax number (760.290.3216), even if they have no product to return.On 4/20/17, an Urgent Medical Device recall letter was sent to expand the recall to include all Wingman 35 Crossing Catheter lengths and lots manufactured since March, 2015.On 5/3/17, ReFlow Medical issue a press release to a news wire through PR.com to inform their customers that FDA is classifying this as a Class 1 recall. The press states that the Wingman35 Crossing Catheters have been found to exhibit tip splitting or separation, which has resulted in two adverse events. ReFlow has received 2 complaints of catheter-tip splitting and/or separation. A total of 2327 Wingman catheters are in distribution. Consumers with questions may contact ReFlow Medical Customer Relations at 1-949-481-0399, Monday through Friday, between 8:00 a.m. and 4:30 p.m. Pacific time or by email at info@reflowmedical.com.On 5/9/17, ReFlow Medical sent an updated press release to the newswire to correct the distribution dates from January 2015 and March 2016 to March 2015 and March 2017. |  | 2794 | 2,794 units | Worldwide Distribution - US Nationwide in the states of : NJ, FL, CA, GA, MI, OH and countries of: Germany, Turkey, and AU. | Open, Classified |  |  | 2794 | Not_Computer
Z-1700-2017 | 76768 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive software. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The versions allow images to be stored without pixel data, resulting in the system not being able to present all the prior studies, which could cause a delay in treatment in determining the progression of disease. | Software design | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A second letter dated 7/11/2016 was issued to nonresponders. |  | 371 | 371 sites potentially have the affected software | Distribution was nationwide to medical facilities. Foreign distribution was made to Canada, as well as other countries. Government and military distribution was also made. | Open, Classified |  |  | 371 | Computer
Z-1698-2017 | 76806 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA Digital Radiography X-Ray System DX-D100 |  |  | January 20, 2016 | 2016 | AGFA Healthcare Corp. | The DX-D100 User Manual already contained information about to move a DX-D100 mobile unit manually by using an Allen wrench, but did not indicate where the Allen wrench should be stored. | Other | AGFA Healthcare sent an On January 20,2016, an 'URGENT FIELD SAFETY NOTICE" dated January 20, 2016, to all affected customers by email or sent via FedEx to the US and Canadian consignees. The notice included an acknowledgment form to be sent back. The letter described background information and actions. Customers with questions were instructed to call 1-877-777-2432 and reference PR1405160001-VR0000288.For questions regarding this recall call 864-421-1984. |  | 308 | 308 | Worldwide Distribution - US including AL, AR, CA, CO, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NY, NC, OH, OK, OR, RI, SC, TN, TX, VA, WV, WI, WY, and Internationally to Canada. | Terminated | April 10, 2017 | 446 | 308 | Not_Computer
Z-1853-2017 | 76347 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | NAMIC Fluid Delivery Set - 72" (183cm) UPN H965913000051, Catalog 91300005 |  |  | January 13, 2017 | 2017 | Angiodynamics Inc. (Navilyst Medical Inc.) | The product has the potential to be non-sterile as it was not packaged in accordance to specification. | Under Investigation by firm | On January 13, 2017 Navilyst Medical initiated a recall action by contacting the 6 consignees by telephone and directing them to stop using the affected product and segregate until receipt of written Recall Notification. Recall Notifications dated January 19, 2017 were delivered by Federal Express to these consignees on January 20, 2017. Per the Recall Notification, Consignees are instructed to segregate and return all affected devices to Navilyst Medical, Inc. (an AngioDynamics Company) and complete and return the Reply Verification Tracking Form, provided in the recall notification. |  | 200 | 8 boxes/25 units | CA, KY, MD, OR, VA, TX | Open, Classified |  |  | 200 | Not_Computer
Z-1701-2017 | 76809 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i LTX DENTAL IMPLANTS |  |  | March 08, 2017 | 2017 | Biomet 3i, LLC | A single lot of product that potentially has a single part within the lot that does not match the product on the label due to a commingle at the firm supplier | Under Investigation by firm | On March 13 Zimmer Biomet contacted their international consignee and asked them to:1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form. a. Return a digital copy to postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship within your territory along with a completed Attachment 1  Inventory Return Certification Form to Biomet 3i.a. For each return, send a copy of Attachment 1 to postmarket@zimmerbiomet.com.b. Include a hardcopy of Attachment 1 with your shipment for immediate processing. c. Mark the outside of the returns box(es) clearly with RECALL. 5. Retain a copy of your recall acknowledgement and product return forms for your records in the event of a compliance audit of your facility.If after reviewing this notice you have further questions or concerns please contact us by email at postmarket@zimmerbiomet.com.1-800-342-5454 |  | 5 | 5 | Argentina | Open, Classified |  |  | 5 | Not_Computer
Z-1697-2017 | 76851 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive (ICEA) software. iConnect Enterprise Archive is intended for use as a vendor neutral archive for storage and communications of medical images and data |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | The study is archived but cannot be opened in iConnect Access and cannot send to PACS resulting in comparison studies not being available for physician review to determine the progression of disease. | Software design | MERGE sent an Urgent Medical Device Recall letter dated July 11, 2016, to all affected customers. The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A second letter dated 7/11/2016 was issued to nonresponders. |  | 298 | 298 sites potentially have the affected versions (290 of those in US) | Worldwide Distribution - US (nationwide) and Internationally to Canada, Belgium, Australia, Jordan, and the UK.. | Open, Classified |  |  | 298 | Not_Computer
Z-1694-2017 | 76591 | No | LKK |  |  | General Hospital | PMA | SynchroMed II implantable drug infusion pump, Model 8637-40, |  |  | February 09, 2017 | 2017 | Medtronic Neuromodulation | Medtronic received a complaint that there was an error codedisplayed on the programmer when the physician attemptedto interrogate an implanted SynchroMed II pump. The errorcode prevented the physician from updating the pump;however the pump was providing therapy. | Under Investigation by firm | A Medtronic representative visited the clinic on 2/9/2017 and corrected the SynchroMed II pump by resetting the pump memory using a SynchroMed Lab Programmer running on a Medtronic laptop with a validated protocol. Consignee was given a Medtronic "Urgent Medical Device Correction" letter dated February 8, 2017. The letter described the Issue and Actions to be taken by the consignee. For questions contact your Medtronic Representative Medtronic or Neuromodulation Technical Services at 1-800-707-0933. |  | 1 | 1 | IL | Completed |  |  | 1 | Not_Computer
Z-1696-2017 | 76763 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Prosound F75 or F75 |  |  | June 11, 2014 | 2014 | Hitachi Aloka Medical, Ltd. | Loosened monitor arm. | Process control | Hitachi Loki Medical sent a customer letter dated June 11, 2014, to affected consignees. The letter informed consignees of the issue and actions the firm planned to take. For questions regarding this recall call 888-782-5652. |  |  |  | Nationwide Distribution to PA, TX, FL, AZ, IN, OH | Open, Classified |  |  |  | Not_Computer
Z-1824-2017 | 76869 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee, Angiographic x-ray system |  |  | March 28, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software error - As a result of a software fault in Artis zee systems with software version VD11 and an A100 generator, the possibility exists that following the failure of a tube assembly focus, the Artis system functions properly until the next reactivation. Following reactivation, the system no longer initializes as intended. | Nonconforming Material/Component | Siemens mailed a Customer Safety Advisory Notice (CSAN) AX017/17/S to affected customers on 03/28/2017 to inform them of the corrective action that will be performed to prevent a possible hazard to persons and equipment. The notice also explained the underlying issue, the impact on system operation, potential risk to patients and actions they are taking to correct the issue. |  | 3 | 3 units distributed in U.S. | Distributed to: VA, CA, WI | Terminated | August 07, 2017 | 132 | 3 | Computer
Z-1767-2017 | 76749 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADDÂż Administration Set, 21-7339-24, Rx Only, Sterile EO. |  |  | March 17, 2017 | 2017 | Smiths Medical ASD Inc. | Smiths Medical has become aware that the drawing on the Instructions for Use (IFU) on certain yellow striped CADDÂż Administration Sets (part number 21-7339-24), incorrectly represent the actual filter size of the product. Although the text description on the IFU correctly identifies the product to include a 0.2 micron filter, the illustration on the IFU identifies the filter size as 2.0 microns. | Packaging process control | Smiths Medical sent an "Urgent Medical Device Field Safety Notice" dated March 13, 2017. The letter was addressed to "Risk Manager". It provided a list of the Affected devices, Risk to Health, Instructions to Distributors &amp; Customers. Advised consignees to notify their customers. There is no need to return the devices to Smiths Medical. Requested consignees to complete the Urgent Medical Device Field Safety Notice (FSN) Response Form and return to smithsmedical3947@stericycle.com. For questions contact Stericycle at 844-265-7407. For further questions regarding this recall , please call (844) 265-7407. |  | 38415 | 38,415 (38145 US; 2268 OUS) | US:AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WV, WY.OUS: Canada | Open, Classified |  |  | 38415 | Not_Computer
Z-1704-2017 | 76486 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter Chemistry Analyzer AU680/AU5800, Catalog No. MU907400 (Rack ID Labels) |  |  | February 13, 2017 | 2017 | Beckman Coulter Inc. | Beckman Coulter is recalling the Rack ID Labels because the Rack ID number displayed and the Rack ID barcode label do not match. | Process control | Beckman Coulter sent an Urgent Medical Device Recall letter dated February 13, 2017, was sent to customers to inform them that Beckman Coulter has identified a mismatch in Rack ID Labels (REF MU9074); the Rack ID number displayed and the Rack ID barcode label do not match. Rack ID labels numbered 0161 thru 0180 may be read as 0141 thru 0160 on the AU System. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions regarding the notice are instructed to contact their Customer Support Center: http://www.beckmancoulter.com/customersupport/support, call 1-800-854-3633 in the United States and Canada; Outside the United States and Canada, contact their Beckman Coulter representative. |  | 2 | 2 sheets (20 labels/sheet) | US Distribution : MA ., Internationally to Japan | Terminated | April 12, 2017 | 58 | 2 | Not_Computer
Z-1706-2017 | 64902 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile Diagnostic X-Ray System with Digital Panel |  |  | November 28, 2012 | 2012 | Philips Healthcare Inc. | When the operator moves an image to another view then the electronic marker might be placed at the wrong anatomical position in the resulting image. | Software design | Philips Medicals Systems distributed an "URGENT Field Safety Notice" alerting customers of the defect on November 28, 2012. Customers are requested to manually rotate images by using the rotation tool.Customers are also recommended to check whether the automatic marker is set correctly, especially after a manual image rotation is performed. The marker can be corrected manually with the annotation tool. |  | 220 | 220 systems | Nationwide (AR, AZ, CA, DE, FL, GA, IA, IL, IN, KS, MD, MI, MN, NC, NM, NY, OH, OR, PA, SC, TX, VT, WI) | Open, Classified |  |  | 220 | Not_Computer
Z-1829-2017 | 76872 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Arrow EZ-IO Needle Set 45 mm 15 ga, Ref 9079. |  |  | March 29, 2017 | 2017 | Teleflex Medical | Some of the individual unit packaging (pouches) may be missing the lot number an expiration date. | Incorrect or no expiration date | Consignees were notified of the recall by letter on March 29, 2017. The letter requested that they check their inventory and immediately discontinue use and quarantine any product and to inspect on hand for missing lot numbers and expiration date. The letter included a reply form to be returned to Teleflex. A separate letter was sent to distributors requesting a sub-recall. |  | 4760 | 4760 eaches | AL, AK, AZ AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY, Canada and Costa Rica | Open, Classified |  |  | 4760 | Not_Computer
Z-1708-2017 | 76660 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | eCare Coordinator |  |  | March 02, 2017 | 2017 | Philips Visicu | eCareCoordinator (eCC) is intended for use in data aggregation, patient interaction facilitation, storage and clinical information management with independent physiological devices and ancillary systems that are connected either directly or through networks. The software is intended to provide patient information from the patient location through networking technology to a remote care team.A software defect was discovered in the Philips eCareCoordinator (eCC) Clinical User Interface that can at times cause missing or redundant data to be saved without notification to the user. | Software Manufacturing/Software Deployment | Field Safety Notice (FSN) will be sent to all customers using eCareCoordinator all Versions. The FSN describes the problem and circumstances in which the design defects occurs and the action planned by Philips to correct the problem. The FSN will also detail the functionality change in detail.A software correction will be made and released.by the vendor. Customers will have this correction applied to their current version when made available by Vendor in order to fix this correction. |  | 26 | 26 | Worldwide Distribution-US Nationwide | Open, Classified |  |  | 26 | Computer
Z-1707-2017 | 76656 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue MX4O WLAN Patient Wearable Monitor |  |  | March 06, 2017 | 2017 | Philips Electronics North America Corporation | Philips lntelliVue MX4O WLAN Patient Wearable Monitor may not automatically switch toMonitor Mode with audible alarms when association with central monitoring system is unsuccessful [incomplete] | Software design | Philips issued Field Safety Notice on 3/6/17. The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer .Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice:When a No Data Tele INOP alert is displayed on a Philips IntelliVue Information Center iX ( PIIC iX ), check the MX40 as directed in the PIIC iX B.01 IFU, page 87. If the MX40 is in the Coverage Area, remove the battery, and reinsert the battery. If the MX40 is not in the Coverage Area, place the sector in the PIIC iX in standby as directed in the MX40 B.06 IFU, page 53, or return the patient and MX40 to the Coverage Area.Philips is releasing a software upgrade B.06.18 for affected Philips IntelliVue MX40 Patient Wearable Monitors. A Philips Healthcare representative will contact customers with affected devices to arrange for the installation of the software upgrade |  | 2648 | 2648 | Worldwide Distribution - US Nationwide including the countries of: AustraliaBelgiumCanadaDenmarkFinlandFranceGermanyGermanyIndiaItalyNetherlandsNorwayQatarSingaporeSouth AfricaSwedenSwitzerlandUnited Kingdom | Open, Classified |  |  | 2648 | Computer
Z-1866-2017 | 76895 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD Viper LT software, version 3.00H |  |  | December 14, 2015 | 2015 | Bd Diagnostic | BD has confirmed reports that lower than expected RFU values were generated on customer BD Viper LT Systems after the installation of software version 3.00H. As RFU values decrease, results approach the assay cutoff threshold which has the potential to impact test results. | Software design | Recall action was initiated on December 14, 2015 when software version 3.00H was removed from the two customer instruments. The 2 affected US customers were contacted via email and were also mailed a letter via UPS overnight express. |  | 2 | 2 | 2 US customers, 0 Ex-US regions, and 0 distributors 0. | Terminated | April 19, 2017 | 492 | 2 | Computer
Z-1864-2017 | 76900 | No | NSX |  |  | General Hospital | Contact ODE | Software version 4.44A utilized on the BD MAX System |  |  | July 29, 2015 | 2015 | BD LIFE SCIENCES | The BD MAX System software version 4.44A contains a software anomaly that has the potential to affect some Open System Reagent (OSR) customers utilizing User Defined Protocol (UDP) assays. Customers utilizing only IVD assays are not affected.This anomaly may cause the OSR customer system to incorrectly switch the columns of the truth table on the users display. If the user tries to fix their display data, the display will appear correct however, the database may become incorrect, potentially resulting in incorrect results reporting.BD is able to trace the issue to a limited number of instruments that have software version 4.44A installed. Note that the problem only affects customers that are either creating a new UDP with Result Logic or modifying the Result Logic of an existing UDP. | Software design | BD sent an Urgent Field Corrective Action Recall notification was initiated on July 29, 2015. The customer was initially contacted via a phone call and subsequently provided with a written communication. All US customers were contacted directly by phone and provided the written communication. Global holds were placed on July 21, 2015 and July 24, 2015. For further questions, please call (410) 316-4000. |  | 29 | There are 29 US customers and 5 ex-US regions with a BD MAX System running software version 4.44A. | Worldwide Distribution - US Distribution to the states of : CA, CO, FL, IL, MN, MO, MT, NE, NJ, NY, OH, OR, PA, TX, TN, WA, WI., and to the countries of : Canada, Hong Kong, Taiwan, Turkey and Europe.Federal Government Agency sales/distribution centers or foreign countries VA Fort Harrison and VAMC Austin. | Terminated | April 19, 2017 | 630 | 29 | Computer
Z-1921-2017 | 76836 | No | KIC | Flotation cushion. | Physical Medicine | Physical Medicine | 510(K) Exempt | ROHO HIGH PROFILE Single Compartment Cushion with Sensor Ready Technology and ROHO MID PROFILE Single Compartment Cushion with Sensor Ready Technology. |  |  | March 20, 2017 | 2017 | Roho Inc. | When the Smart Check tool is used to perform a Smart Setup with one of the affected cushions, it may adversely affect the setup process resulting in an improper cushion setup. Sitting on an over inflated or under inflated Sensor Ready Cushion may reduce or eliminate the benefits provided by the cushion, increasing the risk to skin or to other soft tissue. | Equipment maintenance | On 3/13/17, firm began notifying customers of recall via emailed or mailed letter. The notification sent to dealers and distributors includes a request to forward the communication to their customer and send an acknowledgement back once all of their customers have received a copy of the communication. Customers (product users) will not be required to send the product they are using back. Instead, the recall letter provided two options: 1) keep the cushion and follow additional set-up instructions provided within the letter that would mitigate the risk or 2) contact the supplier to have the cushion replaced. Product that had been in use and is replaced will be destroyed (evidence to be provided to ROHO, Inc. if destroyed by a customer). Unused product that is still in stock with the firm will be destroyed. |  | 5646 | 5646 units | Nationwide, Sweden, Germany, New Zealand, Finland, Australia, Singapore, Korea, Luxembourg, Czech Republic, South Africa, Canada, Switzerland, United Arab Emirates,Spain, Austria, Hong Kong, Iceland, Japan, Canada, Italy, Ireland, France, United Kingdom | Open, Classified |  |  | 5646 | Not_Computer
Z-1738-2017 | 76840 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-D600 - DXD Imaging Package |  |  | May 02, 2016 | 2016 | AGFA Healthcare Corp. | One of the dent markers from the DX-D600 dropped down from the overhead rail. | Labeling design | On May 2, 2016, an "URGENT FIELD SAFETY NOTICE" letter was either emailed or sent via FedEx to the US and Canadian consignees. The notice included an acknowledgment to be sent back. The letter described background information and actions to be taken. |  | 81 | 81 units | AL, CA, DC, FL, GA, HI, IL, IN, KS, LA, MA, MN, MO, NV, NJ, NM, NY, OH, OK, OR, SC, TN, TX, VA, and Canada | Terminated | April 10, 2017 | 343 | 81 | Not_Computer
Z-1724-2017 | 76754 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | STAPLER,IS4000; da Vinci Xi EndoWrist Stapler Release Kit (SRK); model number 470298. |  |  | March 15, 2017 | 2017 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware and further investigation has determined that misinterpretation of instructions provided on the IS4000 Stapler Release Kit instruction card may result damage of the tool. | Device Design | Intuitive sent an Urgent Medical Device Correction (ISIFA2017-02-C) letter dated March 16, 2017 were sent by FED EX to be received on March 17, 2017. Letters advised users of the reason for the notification, risk to health, affected products with part numbers and a copy of Appendix A - proper use of the Stapler Release Kit (SRK). Customers were instructed to distribute a copy of the letter to all users at their facility, follow all instructions using the Instrument Release Kit (IRK) tool to finish manually unclamping the stapler and return the SRK to Intuitive Surgical through the Returned Material Authorization (RMA) process. Customers were also instructed to log into the da Vinci Online Community Field Action resource (US only) to read or complete any request actions related to this issue. In the case where the da Vinci online resource cannot be used (including outside he US), complete the attached Acknowledgement Form and return it via fax to Intuitive Surgical as instructed on the form. Customers were asked to retain a copy of the letter with their Stapler User Manual and the acknowledgement form for their files.If you need further information or support concerning this notification , please contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (4 AM to 5 PM PST) ormail: customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8AM to 6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 03-5575-1362 (9 AM to 6 PM JST)For questions regarding this recall call 408-523-2100. |  | 2059 | 2059 shipped - all languages | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Canada, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Japan, Luxembourg, Monaco, Netherlands, Norway, Portugal, Puerto Rico, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom and Venezuela | Open, Classified |  |  | 2059 | Not_Computer
Z-1737-2017 | 76698 | No | PJC | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Bond Elut OMS (Dried Matrix Spotting)' Product; |  |  | June 17, 2015 | 2015 | Agilent Technologies, Inc. | The product was marketed without 510(k) clearance. | No Marketing Application | Agilent Technologies Letters via e-mail were sent to customers on June 17, 2015.Letters advised that the Bond Elut DMS Product was being withdrawn from the market. |  | 364 | 364 in total | US and Australia, Thailand, Great Britain, Germany, Japan, Netherlands, France, Taiwan, Canada, South Africa, China, Slovenia, Spain, Italy, Norway, India, Russia, Finland, Czech Republic, Singapore, Denmark, Korea, Ireland, United Arab Emirates, .Belgium, Switzerland and Sweden. | Terminated | April 05, 2017 | 658 | 364 | Not_Computer
Z-1762-2017 | 76910 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | iConnect Enterprise Archive (ICEA) software. The firm name on the labeling is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | An error message can occur resulting in the prior studies being unavailable for viewing. | Software design | The recalling firm issued a letter dated 3/28/2016 via e-mail the week of 4/4/2016. If no email address was available, the letters were issued via certified mail. A second letter dated 7/11/2016 was issued to nonresponders. |  | 333 | 333 sites potentially have the affected versions | Distribution was made nationwide to medical facilities. Foreign distribution was made to Canada, as well as other countries. Government/military distribution was also made. | Open, Classified |  |  | 333 | Not_Computer
Z-1748-2017 | 76766 | No | MOI | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens ADVIA Centaur BR assay |  |  | March 08, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc | Potential for the ADVIA Centaur Syphilis assay to interfere with the BR (CA 27.29) a assay on the ADVIA Centaur Systems | Under Investigation by firm | Siemens Healthcare issued An Urgent Medical Device Correction (UMDC) to customers in the United States on March 08, 2017 to be delivered to customers on March 09, 2017and (UFSN) was sent for distribution to customers outside the United States on March 08, 2017. These notices inform customers of the potential and mitigation action to be taken for the ADVIA CentaurÂż Syphilis assay to interfere with the BR (CA 27.29) assay on the ADVIA Centaur Systems. This issue affects all current and future lots of ADVIA CentaurÂż Systems BR reagents until a solution is implemented.Fax-back forms included with letter used to confirm that the customers received the Urgent Medical Device Correction to the Customer Care Center at 312-275-7795.For questions regarding this recall call 508-668-5000.. |  | 1900 | 1900 kits | Worldwide Distribution - US (Nationwide) Internationally to Australia, Brazil, Canada, Columbia, Cyprus, Japan, Spain | Open, Classified |  |  | 5609 | Not_Computer
Z-1751-2017 | 76716 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT model number 882482 |  |  | February 24, 2017 | 2017 | Philips Medical Systems (Cleveland) Inc | The brake hub was not engaging with the gear box shaft due to an error in the assembly of components within the brake hub mechanism after service, causing the detector to backdrive. | Nonconforming Material/Component | Philips Healthcare sent an Urgent Field Safety Notice letter informing consignees of the issue through: Field Safety Notice (FSN) CLE17-006_88200502_88200503. Customers will be notified by direct communication from Philips Field Service Engineers (FSE). A copy of the FSN will be provided to customers by the FSE during the onsite visit for the inspection. A Business Reply form is attached to the FSN for customers to acknowledge receipt and understanding of the FSN. Customers with questions were instructed to contact the Customer Care Solutions Center at 1-800-722-9377.For questions regarding this recall call 440-483-7600. |  | 90 | 90 systems | Worldwide distribution - US including AZ, CA, CO, CT, FL, GA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NY, OH, OR, PA, TN, TX, VA, WA, WI, and WV. Internationally to Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Egypt, France, Germany, India, Israel, Italy, Japan, Mexico, Republic of Korea, Malaysia Netherlands, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, South Africa, Spain, Switzerland, Taiwan, UAE, and UK.There are eleven (11) government accounts for this recall: Edward Hines Jr VA Hospital, 5th Ave &amp; Roosevelt Rd, Hines, IL 60141; VA Jessie Brown OPC, 2030 W Taylor St, Chicago, IL 60612; VA Medical Center, 4300 W 7th St, Little Rock, AR 72205; VA Medical Center, 2200 SW Gage Blvd, Topeka, KS 66622; VA Medical Center, 2501 W 22nd St, Sioux Falls, SD 57105; VA Medical Center, 3600 30th St, Des Moines, IA 50310; VA Medical Center, 2300 Ramsey St, Fayetteville, NC 28301; VA Medical Center Durham, 508 Fulton St, Durham, NC 27705; VA Medical Center of Orlando, 13800 Veterans Way, Orlando, FL 32827; Veterans Hospital, 1 Veterans Dr, Minneapolis, MN 55417; Womack Army Medical Center, Reilly Rd, Fort Bragg, NC 28307.There are thirteen (13) Canada accounts for this recall: CAMPHILL SITE, 1796 SUMMER ST, HALIFAX, NOVA SCOTIA B3H 3A7; COR CARE INC, 325 MILNER AVE STE 410, SCARBOROUGH, ONTARIO M1B 5N1; CSSS DE RIMOUSKI-NEIGETTE, ROULÂżAU 150, RIMOUSKI, QUEBEC G5L 5T1; FOOTHILLS MEDICAL CENTRE, 1403 29 ST NW, CALGARY, ALBERTA T2N 2T9; HEADWATERS HEALTH CENTRE, 100 ROLLING HILLS DR, ORANGEVILLE, ONTARIO L9W 4X9; HEALTH SCIENCES NORTH, 41 RAMSEY LAKE RD, SUDBURY, ONTARIO P3E 5J1; MEDICAL IMAGING CONSULTAN, 200 BOUDREAU RD STE 102ST., ALBERT, ALBERTA T8N 6B9; NORTH BAY REGIONAL HEALTH, 50 COLLEGE DR, NORTH BAY, ONTARIO P1B 0A4; PETERBOROUGH REGIONAL HEA, 1 HOSPITAL DR, PETERBOROUGH, ONTARIO K9J 7C6; ROCKY VIEW GENERAL HOSPIT, 7007 14 ST SW, CALGARY, ALBERTA T2V 1P9; ROYAL VICTORIA HEALTH CEN, 201 GEORGIAN DR, BARRIE, ONTARIO L4M 6M2; SOUTHERN ALBERTA INS. TEC, 1301 16 AVE NW, CALGARY, ALBERTA T2M 0L4; WOODSTOCK HOSPITAL, 310 JULIANA DR, WOODSTOCK, ONTARIO N4V 0A4.There is one (1) Mexico account for this recall: Arrendadora Medica Galena, AV Tulum#MZ 1LT 1#SM 12, Benito Juarez, Quintana Roo, 77505. | Open, Classified |  |  | 201 | Not_Computer
Z-1746-2017 | 76826 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Buffered Sodium Citrate: (9NC) Blood Collection Tubes, 4.5 mL, 13 x 75 mm, Mfd: Becton Dickinson and Company |  |  | March 16, 2017 | 2017 | Becton Dickinson &amp; Company | BD is initiating this product correction of multiple lots of BD Vacutainer Buffered Sodium Citrate: (9NC) Blood Collection Tube (13x75 mm x 4.5 mL). A limited number of tubes in these lots may exhibit stopper pullout from the tube as the user removes the needle from the stopper following specimen collection. | Nonconforming Material/Component | BD sent an "Urgent Product Correction Letters and Customer Product Correction Response Forms" dated March 16, 2017 via email or UPS to their customers and distributors. |  | 13064500 | 13,064,500 tubes (non-expired) | Worldwide - US Nationwide Distribution | Open, Classified |  |  | 13064500 | Not_Computer
Z-1831-2017 | 76843 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | NEODENT CM Intraoral Scanbody, Article Number: 108.139 |  |  | March 13, 2017 | 2017 | Instradent USA, Inc. | During internal testing, Neodent determined that the body diameter (3.8mm) of the Neodent 108.139 CM Exact Intraoral Scan body is not compatible with the platform diameter of CM Implants when the implant is intraosseoussly (subcrestally) positioned. | Component design/selection | Instradent sent an Urgent Field Safety Notice dated March 13, 2017, to all affected customers via UPS. Customers were asked to take the following actions: 1. Identify and segregate the affected units from your stock.2. If the product is still in your Inventory, return it for credit. 3. If you have used the product successfully, there is no need to take any action 4. For all cases complete and return the enclosed Customer Confirmation Form.For further questions, please call (855) 412-8883. |  | 170 | 170 pieces | Distributed to the states of: NJ, TX, NV, CA, PA, AZ, and IL. | Open, Classified |  |  | 170 | Not_Computer
Z-1702-2017 | 76595 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | The Medtronic StrataMR adjustable valves and shunts are designed to provide continuous cerebrospinal fluid (CSF) flow from the ventricles of the brain into the right atrium of the heart or the peritoneal cavity. The design enables the physician to non-invasive adjust valve pressure/performance level pre- and post-implantation by using magnetic adjustment tools without the need for radiographic confirmation. |  |  | February 22, 2017 | 2017 | Medtronic Neurosurgery | Incorrect valve rotor orientation due to misalignment and/or off-center positioning of the StrataMR Locator Tool: When the StrataMR valve is set in the incorrect orientation, higher valve opening pressures result leading to a corresponding underdrainage in CSF fluid performance. This condition may result in the following adverse health consequences: ventriculomegaly, headaches, nausea, vomiting and lethargy. If left untreated, underdrainage has the potential to lead to coma and death. | Device Design | The firm has initiated a CAPA and issued notification via letter to US Customers on February 22, 2017. As of the date of the report (2/22/17) the firm was still in process of notifying foreign customers. Recipients are instructed to: 1) Cease use of the listed product, 2) Account for number of product units implanted or discarded and return any unused units by completing an attached individualized customer product accountability form provided in the letter,3) If any of the indicated products have been implanted in patients, to adjust the valves according to instructions listed in an included attachment.Recipients are asked to fax a customer acknowledgment response form or scan/take a picture of the form and email to rs.mnsfca@medtronic.com with "StrataMR Recall" in the subject line. |  | 1796 | 2622 units (US = 1796, outside US =826) | Product was distributed throughout the US and to Australia, Belgium, Canada, Finland, Germany, Hong Kong, Israel, Italy, Japan, New Zealand, Norway, Poland, Portugal, the Russian Federation, Singapore, Spain, Sweden, Switzerland, Turkey, and the UK. | Open, Classified |  |  | 1796 | Not_Computer
Z-1756-2017 | 76482 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sterile Handpiece Sheaths, P/N 7-796-19, included in the Accessory Packages contained in the HYFRECATOR 2000 Electrosurgical Unit (ESU) cartons, 100 VOLTS AC, REF/Catalog Number 7-900-100, Rx Only. There is no problem with the Hyfrecator 2000 Electrosurgical Unit. -- ASSEMBLED IN MEXICO --- CONMED CORPORATION 525 FRENCH ROAD UTICA, NY 13502 |  |  | February 28, 2017 | 2017 | ConMed Corporation | For over a year, the Accessory Packages in the Hyfrecator 2000 ESU cartons were packaged in error by the manufacturer with non-sterile P/N 7-796-19 Handpiece Sheaths in pouches that were labeled as Sterile. This could cause an end user to use a non-sterile sheath as if it were sterile. | Mixed-up of materials/components | ConMed sent an URGENT: MEDICAL DEVICE RECALL Notices (dated 2/28/2017)to all affected US customers by USPS First Class Mail. International letters are in the process of translation and are scheduled to be sent by March 15, 2017. Customers were instructed to review their inventory for any of the affected product and were asked to complete the business reply form and return it to ConMed Corporation, 525 French Road, Utica, New York 13502.Customers with questions or requests should contact the Field Action Support Team at 1-800-448-6506, fax to 315-624-3225, or email hyfr2000@conmed.com.For questions regarding this recall call 315-624-3533. |  |  |  | Worldwide Distribution - US (nationwide) Internationally to Australia, Canada, Germany, Italy, UK, Great Britain, and Ireland. | Open, Classified |  |  |  | Not_Computer
Z-1765-2017 | 76665 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Immunodiagnostic Products NT-pro BNP Reagent Pack, REF/Catalog 680 2156, IVD --- Ortho-Clinical Diagnostics Pencoed, Bridgend CF35 5PZ, UK |  |  | March 07, 2017 | 2017 | ORTHO-CLINICAL DIAGNOSTICS | Increased frequency of calibration failures for VITROS Immunodiagnostic Products NT-proBNP Reagent Lots 1568, 1570, 1580 and 1590, due to background signals for these affected NT-proBNP lots that have been increasing with time since release testing, resulting in higher than expected VITROS Immunodiagnostic Products NT-proBNP level 1 calibrator signals produced, leading to calibration failures. | Nonconforming Material/Component | On 3/7/2017, IMPORTANT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2017-048) was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers and distributors who were shipped VITROS NT-proBNP Reagent Packs, Lots 1570, 1580 and/or 1590 to inform them of the issue and provided further instructions. |  | 1868 | US: 6324; Foreign: 1868 | Worldwide: United States, Puerto Rico, Bermuda, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Poland, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 1868 | Not_Computer
Z-1874-2017 | 76891 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport Medical HT70 and HT70- Plus Ventilators, Rx Only |  |  | March 30, 2017 | 2017 | Newport Medical Instruments Inc | Medtronic is issuing a voluntary field corrective action for all its Newport" HT70 and Newport" HT70 Plus ventilators because they may shutdown spontaneously during normal operation without an accompanying alarm. | Software in the Use Environment | Medtronic sent an Urgent Field Corrective Action Notice dated April 3, 2017, to all consignees via Federal Express or certified mail. The letter informed consignees of the remote potential for ventilator reset without accompanying alarm and the actions they should take. Customers were requested to acknowledge receipt and understanding of the Urgent Field Corrective Action Notice. Customers were also advised that a software service update will be available in May 2017 to address this issue. Medtronic issued a Press Release related to this voluntary field action on April 5, 2017.Consignees with questions were advised to contact the Technical Support Department immediately at 800-255-6774.For questions regarding this recall call 203-492-5415. |  | 12966 | 12,966 | Worldwide Distribution - US (Nationwide) Internationally to Canada, Australia, Brazil, Japan, Korea, Mexico, Bahrain, Egypt, GreeceHungary, Ireland, Israel, Jordan, Kazakhstan, Lithuania, MauritiusNamibia, Norway, Pakistan, Poland, Qatar, Romania, Russian FederationSouth Africa, Spain, Turkey, United Arab Emirates, United KingdomAlgeria, Armenia, Austria, Bahrain, Belarus, Belgium, BotswanaBulgaria, Cyprus, Czech Republic, Denmark, Egypt, Finland, FranceGermany, Greece, Hungary, Iran, Ireland, Israel, Italy, Kenya, LebanonLithuania, Macedonia, Malawi, Malta, Mauritius, Mozambique, NetherlandsNorway, Oman, Poland, Portugal, Qatar, Republic of Moldova, RomaniaSaudi Arabia, Serbia, Slovakia, Sudan, Sweden, Switzerland, TunisiaTurkey, Uganda, Ukraine, United Kingdom, Yemen, and Zimbabwe | Open, Classified |  |  | 12966 | Computer
Z-1775-2017 | 76885 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Refinity Rotational Intravascular Ultrasound (IVUS) Catheter; |  |  | March 29, 2017 | 2017 | Volcano Corporation | High complaint rates of kinking near the junction of the proximal shaft to telescope section of the catheter. Kinking could result in disruption of normal clinical workflow or loss of imaging. | Under Investigation by firm | Recall letters dated March 28, 2017 will be hand delivered by Volcano reps. Reps will also pick up any unused catheters for return and obtain a signed proof of delivery for both customer and Volcano Rep. |  | 718 | 718 total units | Distribution to US and Japan | Terminated | June 21, 2017 | 84 | 718 | Not_Computer
Z-1877-2017 | 76928 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ViperWire Advance with Flextip Peripheral Atherectomy Guide Wire, Part Number: 7-10026-01,Model Number: VPR-GW-FT14. |  |  | February 09, 2017 | 2017 | Cardiovascular Systems Inc | CSI is recalling three lots of VIPERWIRE ADVANCE WITH FLEXTIP PERIPHERAL ATHERECTOMY GUIDEWIRE . It was discovered that 0.018 in ViperWire Advance with Flextip product was packaged as a 0.014in ViperWire Advance with Flextip. | Process control | On 2/9/2017, consignees were sent a CSI "Urgent Medical Device Recall" letter dated 09 February 2017. The letter described the Affected Product, Recall Description, &amp; Instructions. Advised consignees to immediately remove, return the product &amp; complete the Customer Acknowledgement Form. For further information contact Customer Service, Cardiovascular Systems, Inc., 1225 Old Highway 8 NWSaint Paul, MN 55112 877-274-0901 - Tel, 612-677-3355 - Fax. |  | 595 | 595 | US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, NJ, NY, OH, OR, PA, SC, SD, TN, TX, and WA | Terminated | July 27, 2017 | 168 | 595 | Not_Computer
Z-1879-2017 | 76935 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Pressure Injectable Arrowg+ard Blue Plus Multi-Lumen Central Venous Catheterization Kit. |  |  | March 17, 2017 | 2017 | Arrow International Inc | Arrow International is recalling due to possible voids in the catheter hub, which could lead to an interlumen crossover. | Process control | Arrow International mailed an Urgent Medical Device Notification Letter dated April 4, 2017, to affected customers to inform them of the issue. Customers were asked to discontinue use immediately and return product by completing the enclosed Recall Acknowledgement Form then emailing it to recalls@teleflex.com or faxing it to 1-855-419-8507. A Customer Service Representative will provide a Return Goods Authorization Number for the products' return.For questions regarding this recall call 866-246-6990. |  | 310 | 310 units | US Distribution to NJ, IL, and PA | Open, Classified |  |  | 310 | Not_Computer
Z-1766-2017 | 76762 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | The PreludeSYNC Radial Compression Device is a sterile, single use disposable device used to assist in gaining and maintaining hemostasis of the radial and ulnar artery following catheterization procedures. It has a soft wristband with a secure hook and loop fastener. The band delivers adjustable compression of the puncture site with an inflatable bulb, and a check valve for easy inflation and deflation with a syringe. A clear curved backer plate provides optimal visualization of the puncture site and ease of placement. |  |  | March 03, 2017 | 2017 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. announces a voluntary field action for the PreludeSYNC Radial Compression Device because particulate generated during the manufacturing process may enter the valve, preventing a complete seal of the inflation valve. | Process control | Merit Medical sent an Urgent Product Recall Notice on March 3, 2017, via letter for consignees to immediately stop using or distributing the affected lots and return them to Merit. For any questions, please contact Merit Customer Service at (801) 208-4381. For further questions regarding this recall, please call (801) 253-1600. |  | 1940 | 1,940 units total | Nationwide, Canadian and other foreign consignees. No VA/govt/military consignees. | Terminated | June 21, 2017 | 110 | 1940 | Not_Computer
Z-1771-2017 | 76784 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ViperWire Advance Peripheral Guide Wire with Flex Tip, Model No. VPR-GW-FT18. Sterilized with Ethylene Oxide. |  |  | March 03, 2017 | 2017 | Cardiovascular Systems Inc | Cardiovascular Systems, Inc. is recalling one lot of ViperWire Advance with Flextip, model VPR-GW-FT18 because it may contain a 0.014" ViperWire Advance Flextip product instead of the 0.018" ViperWire Advance with Flextip. | Under Investigation by firm | Customers were sent an Urgent Medical Device Recall letter beginning March 03, 2017. The letter identified affected product, reason for recall, and asked for affected devices to be removed from use and return to CSI. Questions can be directed to CSI Customer Service at 877-274-0901, or your CSI Sales Representative. |  | 186 | 186 devices | Nationwide distribution in the states of AR, AZ, CO, FL, GA, HI, IN, KS, KY, LA, MS. NC, NM, RI, TX, VA | Open, Classified |  |  | 186 | Not_Computer
Z-1922-2017 | 76955 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | The JETSTREAM is a rotational atherectomy catheter system designed for use in debulking and treating vascular disease in the peripheral vasculature. The System includes multiple distal ports located at the catheter tip, which are designed to provide independent infusion and aspiration functions for the active removal of fluid, excised tissue and thrombus from the peripheral treatment site. |  |  | April 05, 2017 | 2017 | Boston Scientific Corporation | Boston Scientific is initiating a voluntary removal of two lots of JETSTREAM SC 1.85mm and one lot of JETSTREAM XC 2.1mm Atherectomy Catheters due to these units being mislabeled. Some 1.85mm JETSTREAM SC catheters may have been labeled as 2.1mm catheters, and some 2.1mm JETSTREAM XC catheters may have been labeled as 1.85mm catheters. | Process control | Consignees were sent a Boston Scientific "Urgent Medical Device Removal" letter dated April 5, 2017. The letter was addressed to Materials Manager / Field Action Contact. The letter described the issue and the prduct involved in the recall. Advised consignees to segregate and return the affected product to Boston Scintific. This recall is to be carried to the customer level. Requested consginees to complete and return the Account Reply Verification Tracking Form. Contact your local Sales Representative can answer any questions that you may have regarding this Removal. |  | 137 | 137 | US: AL, AZ, CT, FL, HI, IL IN, IA, KY, LA, ME, MI, MO, NJ,NY, OH, OR, PA, SC, TN, TX, VA, WA, WI.OUS; GERMANY, GREAT BRITAIN, POLAND, NETHERLANDS | Open, Classified |  |  | 137 | Not_Computer
Z-1774-2017 | 76862 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Kappa, Catalog Number: MS18853 in combination with Scio, Scio Four, Scio Four Oxi plus, Scio Four Oxi, Scio Four plus. Catalog Number(s): MS18597/MS18596/MS18853. |  |  | March 28, 2017 | 2017 | Draegar Medical Systems, Inc. | It was reported that a set low O2 alarm does not go off although the measured O2 level is below the alarm limit. | Under Investigation by firm | On March 28, 2017, US consignees were sent an Urgent Medical Recall letter and Customer Reply and Order card. Delta family monitors running software version 10.0 in facilities that also have at least one Scio module will be downgraded to software version VF9.1 as a temporary solution Free of Charge. Draeger Medical Systems, Inc. is developing a new software version to resolve the issue (VF10.1). Once available, all Delta family monitors that were running software version VF10.0 will be upgraded with VF10.1 Free of Charge.For all other countries, the responsible local Draeger Sales and Service will mail the letter and card. |  | 2156 | 2,156 (total for all devices) | Worldwide - US Nationwide distribution in the states of: AL, CA, GA, IL, LA, MA NY, OH, PA, TN, TX and VA and the countries of:ARATBEBGBNBOBYCACHCLCOCZDEEGESFIFRGBGRHRHUIDIEILINITLTMXMYNLNONPPAPEPKPLPTRORSRUSASESKTHTRTWUGUZVNZA | Open, Classified |  |  | 2156 | Computer
Z-1770-2017 | 76889 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Combi Dockable Neurosurgery table for the Siemens MAGNETOM Aera (1.5T) and Skyra (3T) MRI Systems |  |  | October 17, 2016 | 2016 | Siemens Medical Solutions USA, Inc | Due to a potential manufacturing error, a rare possibility exists for a sudden drop in table height of the Combi Dockable Table Neurosurgery. | Software design (manufacturing process) | Siemens mailed an Customer Safety Advisory Notice on October 17, 2016, to all affected customers to inform them of the corrective action they plan to take to correct the issue. The notice also informed users of the possible risks to patients and how the corrective action will be implemented. For further questions, please call (610) 448-6461. |  | 3 | 3 systems distributed in the U.S. | US Distribution to: DC and MD. | Completed |  |  | 3 | Not_Computer
Z-2733-2017 | 76884 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | ATAR Disposable Extension Cable, sold under brand names: ATAR Disposable Series, ATAR 53912 Disposable Series, and XI Disposable Series. |  |  | March 31, 2017 | 2017 | Oscor, Inc. | The Disposable Oscor ATAR extension cables could separate form the connector during use when the instructions for reuse are not being followed. | Other | Oscor sent an Important Information letter dated March 27, 2017, to all affected customers. The letter explained that they awee aware the Reusable Oscor ATAR extension cables could separate from the connector during use when instructions for use were not followed. They pointed out to consignees that if reused, they can only be reused a maximum of 2 times and must be resterilized by Oscor EtO.Customers were notified again, when the recall expanded on 06/01/2017, and were sent notifications on 06/05/2017. Instructions to customers included to examine inventory, quarantine any affected devices, notify customers if further distributed, and contact Oscor to arrange for replacement. Customers were notified a third time when the recall expanded on 06/13/2017, and notifications were sent between 06/23/2017 and 07/01/2017. Instructions included to check inventory, quarantine any affected devices, return the Acknowledgement Form, and return affected cables to Oscor.Customers with questions should call 727-937-2511. |  | 36242 | 36,242 units | Worldwide Distribution -US to AL, AK, CA, CO, FL, GA, ID, IL, IN, IA , KS, KY, ME, MD, MA, MI, MN, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI, DC. Foreign distribution to following countries: Macedonia, Pakistan, Austria, Canada, Guatemala, United Kingdom, Republic of Panama, Chile, Aruba, Saudi Arabia, Kuwait, Costa Rica, Albania, Estonia, New Zealand, Germany, Puerto Rico, Turkey, Virgin Islands, Iraq, Hong Kong, and Morocco, | Open, Classified |  |  | 110840 | Not_Computer
Z-1782-2017 | 64545 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS Web Diagnostic (WebDX) by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources within the system or across computer networks at distributed locations. |  |  | February 18, 2013 | 2013 | Ge Medical Systems Information T | It was reported from a complaint that study category labels of "Prior" and "Baseline" and hanging position appeared incorrectly swapped when the hanging protocol for three comparison studies is manually applied from the patient folder. | Software design | Firm sent notification letter to all impacted customer accounts. The letter provided the following instructions for customers to follow until the permanent fix for the issue was deployed: 1) The users should be instructed to look at the date on each exam displayed rather than just the labels assigned to the exam by the system. 2) If this problem is detected in your specific workflows, a system administrator can remove the study category labels (Current/Prior/Baseline)from all overlays on the images/exams. A software code change was deployed to customer sites to correct the issue. |  | 118 | 118 | Nationwide, Worldwide | Terminated | July 31, 2017 | 1624 | 118 | Computer
Z-1777-2017 | 76552 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn ProBP 2400 Digital Blood Pressure Device #2400, REF 901096, Rx ONLY, Made in China --- Distributed by Welch Allyn Inc., 4341 State Street Road, Skaneateles Falls, NY 13153 --- Microlife Corporation 9F, 431, RuiGuang Road, NeiHu, Taipei, 11492, Taiwan, R.O.C. --- Microlife AG Espenstrasse 139 9443 Widnau, Switzerland |  |  | February 27, 2017 | 2017 | Welch Allyn Inc | A defective electrical component in the battery charging circuit in two lots of the ProBP 2400 may potentially cause an over-voltage battery charging condition when the device is connected to the external power supply/battery charger. This potential defect resides with the device, not the battery. The over-voltage condition can result in high NiMH battery temperatures that, in some cases, can reach levels sufficient to cause melting of the plastic battery door and other adjacent plastic and foam device components. | Other | Welch Allyn sent an Urgent Medical Device Recall letter and Response Forms dated February 21, 2017. They were mailed to the first round of customers beginning on Monday February 27, 2017, via US Postal Service Priority Mail. Welch Allyn will follow up with second round of mailings to the additionally identified end users the week of March 20, 2017.All end users will be instructed to confirm that their unit is affected; and if affected, a replacement unit will be provided so that the affected unit can be returned and scrapped. --- Of the 1120 units, a total of 981 units were distributed from Welch Allyn warehouses and 139 units remained under Welch Allyn control and were not distributed. These units have since been destroyed.For further questions, please call (844) 360-8220. |  | 981 | 981 units | Worldwide Distribution - USA (nationwide) Distribution | Open, Classified |  |  | 981 | Not_Computer
Z-1783-2017 | 76774 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cataract Kit, part numbers CAT1467AM and CAT1538(B. |  |  | March 16, 2017 | 2017 | Windstone Medical Packaging, Inc. | AMS custom kits containing Medtronic Covidien Curity Eye Pad are recalled due to the potential for the sterile packaging of the eye pad to be compromised. | Nonconforming Material/Component | AMS sent an Urgent Medical Device Recall letter dated March 16, 2017, and provided yellow stickers "Recall Notice" to customers. They are advised to over-label the affected AMS kits at their facilities. Customers were instructed to place sticker in such a way that it would be immediately noticed but should not cover up or conceal any pertinent information such as Pack Number, Pack Name, Lot Number or Expiration Date on a kit. Customers should check inventory and complete the Recall Reply Form. Fax the completed form to Aligned Medical Solutions at 407-804-8460 or email the completed form to vdavis@alignedmedicalsolutions.com.Customers can call Aligned Medical Solutions at 321-527-7714. |  | 4065 | 4,065 kits | Nationwide Distribution including CA, CO, FL, GA, IA, KY, MA, MN, MO, ND, NJ, NV, OH, SD, TX, UT, WI, and WY. | Terminated | June 06, 2017 | 82 | 19512 | Not_Computer
Z-1780-2017 | 76805 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Vivid S60/S70/S60N/S70N, H45041SU, H45041SW, H45581MS, H45581PD. |  |  | March 10, 2017 | 2017 | GE Healthcare, LLC | GE Healthcare has become aware of an issue where a patient other than the intended is incorrectly selected by the operator in situations where DICOM Worklist search response time is slow. This issue is limited to certain Vivid Ultrasound systems and can result in the incorrect patient information showing on the screen during the exam. If not detected by the operator, images from the actual (intended) patient will be stored under the incorrect patient after the exam. This issue could lead to misdiagnosis. There have been no injuries reported because of this issue. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 76164,76168 dated March 10, 2017. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering, &amp; Managers of Radiology / Managers of Cardiology.The letter described the safety Issue, Safety Instructions, Affected Product Detail, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 396 | 396 (67 US; 329 OUS) | United Kingdom, United Arab Emirates, Turkey, Thailand, Switzerland, Sweden,Spain, South Africa, Slovenia, Saudi Arabia, Russia, Romania, Portugal, Poland,Philippines, Norway, New Zealand, New Caledonia, Netherlands, Myanmar, Kuwait, Korea (Republic of), Japan, Italy, ISRAEL, Ireland, Iran, India, Hungary, Hong Kong, Guatemala, Greece, Great Britain, Germany, France, Finland, Estonia, Egypt, Ecuador, Denmark, Czech Republic, Croatia, Colombia, Chile, Canada, Bulgaria, Brunei Darussalam, Brazil, Belgium, Austria, Australia, Argentina, Vietnam.Uruguay, Taiwan, Slovakia, Singapore, Reunion, Qatar, Panama, Pakistan, Jordan, Indonesia, Guadeloupe, Czech Republic | Open, Classified |  |  | 955 | Not_Computer
Z-1778-2017 | 76859 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo software. The firm name on the labeling is Merge Healthcare, Hartland, WI. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | Use of the software may show an incorrect value to the user when viewing the Fractional Flow Reserve (FFR) results during recording. | Software design | MERGE sent an Urgent Medical Device Recall letter dated April 4, 2016 to all affected customers via e-mail the week of April 4, 2016. If no email address was available, the letters were issued via certified mail. A second letter dated July 11, 2016 was issued to non-responders. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all users of the affected product be provided with the notice. Customers were also asked to reply using the enclosed form and the return addressed envelope. Customers with questions were instructed to send an email to recall@merge.com.For questions regarding this recall call 877-741-5369. |  | 24 | 24 sites potentially have the affected versions | Nationwide Distribution including AZ, CA, CO, FL, GA, KS, LA, MI, MO, OH, OK, TN, TX, VA, WA, WI, and WV. | Open, Classified |  |  | 24 | Computer
Z-1817-2017 | 75890 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Nordson Micromedics FibriJet Endoscopic Catheter Introducer, Catalog Number SA-3445. Sterilized using Ethylene Oxide |  |  | October 20, 2016 | 2016 | Micromedics, Inc. | The sterile packaging may contain small channels in the seal of the pouch, If the channels are present, the package integrity may be compromised which may affect sterility assurance. | Employee error | A Medical Device Recall Letter was sent to customers beginning 10/20/2016 via Federal Express in the US and via UPS ground in Israel. The letter identified the affected device and lot number, issue, asked for product to be discontinued from use, quarantine and notify Nordson Medical, via a response form. Response form can be faxed back at 888-504-0606, email to sharon.lynch@nordsonmedical.com or mailed to Nordson Medical, 1270 Eagan Industrial Road, Suite 120, St. Paul, MN 55121. |  | 273 | 273 units | AL, CA, FL, MS, NC, NJ, NY, TN, and TX. Outside the US: Israel | Completed |  |  | 273 | Not_Computer
Z-1812-2017 | 76787 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | Volt Resistance Heated Slipper (All sizes) |  |  | April 15, 2016 | 2016 | 2k Innovations Inc. | During an FDA inspection it was found that the firm claims unapproved therapy and medical claims on product information and website. | Labeling False and Misleading | 2k Innovations sent e-mail communication from 04/15/2016 thru 09/09/2016 informing their consignees that the firm is no longer linking their products to any medical and therapeutic claims and that new labels and boxes will be corrected and the claims will be taken out of the website. For questions customers should call 888-518-6871, ext 700. |  | 22757 | 22,757 items in total | Nationwide Distribution | Open, Classified |  |  | 22757 | Not_Computer
Z-1815-2017 | 76342 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System, Model 4.5, 4.7, 5.0 |  |  | January 18, 2017 | 2017 | RAYSEARCH LABORATORIES AB | An error in NVIDIA GPU (Graphics Processing Unit) card drivers can occur for certain software programs. In a non-standard Citrix environment where the GPU settings are configured with a virtual Citrix graphics board (graphics card) an erroneous calculation may be allowed. | Device Design | RaySearch Laboratories sent a Field Safety Notice dated January 18, 2017, to their customers via email. ACTIONS TO BE TAKEN BY THE USER include: 1) Verify the GPU settings on each of the Citrix XenApp servers according to the instructions in this section. A correct configuration of the RayStation GPU settings will show a graphics board that has been verified by RaySearch. The graphics boards that are verified for operation are listed in the section "Graphics card (GPU) for accelerated computations" in the System Environment Guidelines. 2) Check that displayed graphics board in the GPU Settings dialog is listed as verified for operation according to the section "Graphics card (GPU) for accelerated computations" in the SystemEnvironment Guidelines. The following figure shows an example of a correct installation. If the graphics board is not verified for operation according to the System Environment Guidelines, the installation is incorrect. In this case, disable GPU computation and contact RaySearch Support. 3) Check if "Citrix Systems Inc. Display Driver" is displayed instead of the graphics board name. Ifso, the Citrix XenApp installation is incorrect. If an incorrect installation of Citrix XenApp is detected: a. Set the Citrix XenApp server to maintenance mode or disable GPU computation on the erroneous server; b. Contact RaySearch Support for instructions on how to resolve the issue. Customers are advised to pass a copy of the safety notice and educate IT staff and all users that are members of the RayStation-Administration user group (or equivalent) about this (i.e., users with the permission to change GPU settings). Inspect your product and identify all installed units with the above software version number(s), then confirm you have read and understood this notice by completing and returning the last page via email or fax to: freddie.cardel@raysearchlabs.com, fax 888 501 7195. |  | 741 | 741 units | Nationwide Distribution to AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, LA, MN, MD, MA, MI, MS, MT, NJ, NY, NC, NV, OH, OK, OR, PA, RI, TN, TX, VA, VT, VI, WA, WI, and Washington D.C. | Open, Classified |  |  | 741 | Computer
Z-1816-2017 | 76506 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | TotalCare Bed Systems, Model Numbers P1900, PR1900, and U1900 with 3-Level Patient Position Monitoring and Scale Upgrade Kits, Model Numbers P1947A and P1948A). |  |  | February 23, 2017 | 2017 | Hill-Rom, Inc. | The bed exit alarm may not function properly if the alarm has been armed for a period longer than seventy-two (72) hours without the alarm being triggered by patient movement or removal of patient weight within that time period. | Software change control | On 2/23/2017, URGENT: MEDICAL DEVICE CORRECTION notifications were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers without a service contract received the following instructions:Action to be taken by the user: A software update for the TotalCareÂż beds GCI has been released to resolve this issue. Enclosed with this letter are a USB stick with the updated software, instructions on how to update the GCI software, and a response form to be returned within 30 days. 1) Identify any affected TotalCareÂż beds with 3-Level Patient Position Monitoring in your facility by using the criteria above. 2) Please use the enclosed instructions and update the software on each potentially affected TotalCareÂż bed immediately. 3) Record the serial numbers of the beds being updated on the response form and return to Hill-Rom within 30 days of receipt of this letter. Additional USB sticks with the updated software may be obtained by contacting Technical Support at 1-800-445-3720 (select Option 2, then Option 2 again). Until each of the affected TotalCareÂż beds are updated with the new software, we recommend you do one or more of the below to avoid any potential occurrences of bed exit not functioning. 1. Power cycle (unplug and plug in the bed mains power cord) once every 48 hours. 2. Remove the Patient and re-zero the scale once every 48 hours. 3. Turn off and do not use the Bed Exit Alarm functionality. Continued use of the bed exit alarm without the mitigations listed above may cause risk to patient. If you have an upgrade kit (P1957A) that has not been installed, please contact Hill-Rom for return and replacement. Customers with a service contract received the following instructions:Action to be taken by the user: A Hill-Rom representative per |  | 3037 | 3037 | US Nationwide. AUSTRALIABANGLADESHBRAZILCANADACHINAGERMANYINDIAITALYKUWAITPAKISTANSAUDI ARABIATHAILANDUNITED ARAB EMIRATESCYPRUSGREECENETHERLANDS | Open, Classified |  |  | 3037 | Computer
Z-1818-2017 | 76782 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Keystone Dental Inc. PrimaConnex Straight Implant TC, WD Internal Connection 5.0 x 11.5mm. |  |  | March 09, 2017 | 2017 | Keystone Dental Inc | Straight PrimaConnex Implants contained Tapered PrimaConnex Implants | Labeling mix-ups | Keystone Dental initiated a market removal communication by telephone to customers on 3/9/2017. Customer Relations Fax Back Forms were completed to document these calls. In addition, a letter was sent via Federal Express/or hand delivered to all distributors and customers on 3/13/2017. Product returns requested and if implanted to implement healing and prosthetic protocol. Toll Free: 866-902-9272 |  | 59 | 59 units | CA, IL, IN, FL, MA, MN, NJ, NV, PA, | Terminated | July 05, 2017 | 118 | 59 | Not_Computer
Z-1819-2017 | 76911 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi EndoWrist Suction Irrigator, 8 mm instrument; |  |  | March 31, 2017 | 2017 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware that in specific scenarios with system software P6 and the da Vinci XiSuction Irrigator (PN 480299-03), users can experience unexpected motion of a system arm. | Software design | Firm sent e-mails on 3-17-17 to affected sites instructing them to STOP USE of the XI Suction Irrigator and RMA. and return all unused Xi Suction/irrigation instruments for credit. A formal letters was sent to affected sites on March 31, 2017. Letters advised of the reason for recall with details on product name and product number. Letters requested that any unused devices be returned to Intuitive via the standard RMA process. All personnel who use the Da Vinci system should be made aware of the issue. Response form should be completed online or by returning copy enclosed with the letter. |  | 15 | 15 x 6 packs (90 units) | US Only - one location each in AL, CO, KS, NV, NY, and 2 in TX | Terminated | July 28, 2017 | 119 | 15 | Computer
Z-1823-2017 | 76746 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered SwissPlus Implant System, Model No. SPB10 |  |  | March 20, 2017 | 2017 | Zimmer Dental Inc | It was manufactured with the minor diameter of the internal thread that is below specification. | Nonconforming Material/Component | A customer letter was sent on 3/20/17 to inform customers that Zimmer Biomet is recalling a single lot of the Tapered SwissPlus 3.7mmD Implant, Lot No. 63498673. The letter informs customers that the recall is due to one (1) lot of part number SPB10 (lot 63498673) that was manufactured with the minor diameter of the internal thread below specification. This condition may cause mating issues with some screws. The condition occurs when the screw used has a minor thread diameter that is closer to the maximum allowed diameter (maximum material condition). However, screws closer to the least material condition will still fit into the implant. Customers with questions or concerns are instructed to call the customer call center at 1-800-854-7091 between 8:00 am and 5:00pm EST, Monday through Friday. Customers are also instructed to email to corporatequality.postmarket@zimmerbiomet.com. |  | 188 | 188 units | Worldwide: ZA, FR, IT, IR | Open, Classified |  |  | 188 | Not_Computer
Z-2036-2017 | 76934 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PERI-LOC VOLAR DISTAL RADIUS - RIGHT 3H STD HEAD LOCKING PLATE, 62 MM LENGTH, SST PLATE, REF 71823114, QTY: (1) NON-STERILE, Rx only, Smith &amp; nephew, Inc., Memphis TN 38116 |  |  | March 22, 2017 | 2017 | Smith &amp; Nephew, Inc. | This lot of left plates were laser etched and packaged as right plates. | Mixed-up of materials/components | Smith &amp; Nephew initiated their recall by email on March 22, 2017. Consignees are being requested to return all affected product in their possession. Distributors were instructed to notify their customers and ensure that the recall instructions were followed. For further questions, please call :1(800) 821-5700. |  | 12 | 12 units | Internationally to the country of : India | Open, Classified |  |  | 12 | Not_Computer
Z-2034-2017 | 76828 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Sysmex, Automated Hematology Analyzer XN Series, Sysmex Corporation, 1-5-1 Wakinohama-Kaigandori, Chuo-ku, Kobe 651-0073 Japan |  |  | March 23, 2017 | 2017 | Sysmex America Inc | Reported customer complaints of automated hematology analyzers catching fire due to misalignment of an internal protective sheet that shields electronic components from salty reagents. Build-up of salt deposits on unprotected electronic components may combust due to heat generated by electronics when instrument cycles. Completed CAPA identified the root cause as supplier providing defective component (sensor ready valve) and the misalignment of the electronic protective shield during manufacturing of the device. | Component design/selection | Sysmex sent an Field Correction letter dated April 2017. Product will not be returned, but corrected at the customer site. The correction requires a service visit that will be documented in the Sysmex service Customer Relationship Management (CRM) software. The service report is accepted by the customer at the time of service. Completion of the service records will be tracked by reports generated from the Service CRM. For further Technical Assistance or questions please call (1 888) 879-7639. |  | 3682 | 3,682 units | Nationwide, Canada | Open, Classified |  |  | 3682 | Computer
Z-1820-2017 | 76810 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper Release 8.2 (Laird chiller is always installed); Allura Xper FD10 722026; Allura Xper FD10/10 722027; Allura Xper FD20/10 biplane 722029; Allura Xper FD10 OR Table 722033; Allura Xper FD10/10 OR Table 722034; Allura Xper FD20/10 OR Table 722036; Allure Xper FD20/20 722038; Allura Xper FD20/20 biplane OR Table 722039 |  |  | March 22, 2017 | 2017 | Philips Electronics North America Corporation | Due to a leak in the detector cooling system, cooling liquid may leak outside the drip tray of the chiller. If this occurs the liquid may drip onto electrical components in the R cabinet located in the technical room, which could lead to damage to the system and potentially cause thermal events such as a burning odor, smoke or fire. | Under Investigation by firm | Consignees were notified of the recall via letter sent on 3/22/2017. The letter explains the issue, actions that should be taken by the customer, and actions to be taken by Philips. |  | 6992 | 6992 (total for all devices) | Nationwide, worldwide including Canada | Open, Classified |  |  | 6992 | Not_Computer
Z-1826-2017 | 76800 | No | MAF |  |  | Cardiovascular | PMA | Cordis S.M.A.R.T. Flex Vascular Stent System |  |  | February 16, 2017 | 2017 | Cordis Corporation | Deployment Difficulty. | Other | Email to the Cardinal Health and J&J; country entities sent February 16, 2017. Each country will perform the recall locally. (This is an OUS only Removal.) |  | 2700 | 2,700 | FOREIGN ONLY: ARGENTINA, ARMENIA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, FINLAND, FRANCE, GERMANY, IRAN, ISRAEL, ITALY, JORDAN, KOREA (South), KUWAIT, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, POLAND, PORTUGAL, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, and UNITED KINGDOM | Open, Classified |  |  | 2700 | Not_Computer
Z-1828-2017 | 76844 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Station and Merge Eye Care PACS. |  |  | April 04, 2016 | 2016 | Merge Healthcare, Inc. | During an antivirus program scan of the Eye Station or Eye Care PACS, the antivirus program may detect a newly captured image to be an unwanted file and delete the image permanently. | Software in the Use Environment | The recalling firm issued a letter dated 3/28/2016 via email the week of 4/4/2016 or via certified mail if they did not have an email address for the customer. A second letter dated 7/11/2016 was issued to nonresponders. |  | 1627 | 1,627 sites potentially have the affected versions | Distribution was nationwide to medical facilities. Foreign distribution was made to Canada, as well as to other countries. There was also government distribution. | Open, Classified |  |  | 1627 | Not_Computer
Z-1825-2017 | 76718 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Hemotherm CE, model 400CE, Heater-Cooling Device |  |  | March 09, 2017 | 2017 | Cincinnati Sub-Zero Products Inc | Improperly rated fuses were used in the manufacture of the device. | Component design/selection | Letters were sent to customers and distributors instructing them to immediately cease use and/or further distribution of affected product, remove and discard the affected fuses and replace with the new fuses provided by CSZ, and complete and return the Field Action Response Form. |  | 390 | 390 devices | Worldwide distribution. The recalled product was distributed to the following states: AL, AZ, CA, CO, FL, GA, IL, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, VA, and WI. The recalled product was distributed to the following countries: Mexico, Taiwan.There are no Canada accounts for this recall.There is one (1) Mexico account for this recall: Koolfer S.A. DE C.V., San Isidro 168, Col. Pedregal De Santa Ursula Coapa, Del. Coyoacan, CP. Mexico D. F. 04600.There is one (1) government account for this recall: WRNMMC -PAA, 8965 Brown Dr, Bethesda, MD 20889. | Open, Classified |  |  | 390 | Not_Computer
Z-1827-2017 | 76765 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Ferno PROFlexxÂż Model 28Z Chair Cot |  |  | March 08, 2017 | 2017 | Ferno-Washington Inc | The recall was initiated as a result of the detection of an incorrect nut used in manufacturing. | Mixed-up of materials/components | Consignees were contacted via telephone and email on 3/8/17. No product will be returned to the recalling firm. The firm and its service company will correct each product in the field at the customers location. |  | 34 | 34 units | Worldwide distribution. The recalled products were distributed to the following states: CA, FL, HI, IL, NY, OH. There are no government accounts for this recall. There are no Canada accounts for this recall. There are no Mexico accounts for this recall. The recalled product was distributed to the following countries: Taiwan, United Arab Emirates, Thailand, Singapore | Terminated | June 29, 2017 | 113 | 34 | Not_Computer
Z-1834-2017 | 76509 | No | GFJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Soft Face "No Bounce" Mallet |  |  | February 20, 2017 | 2017 | Signal Medical Corporation | Potential for poly tip of the mallet to come off during use. Potential for foreign material in the surgical field or delay in surgery. Reminder for customers that the poly tip must be removed for cleaning and sterilization to assure proper cleaning and sterilization of the entire mallet. | Labeling design | On 2/20/2016, URGENT: MEDICAL DEVICE RECALL notifications dated February 10, 2017 were sent to the affected consignees via e-mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are to retain the notice and replace the old Instructions For Use with the updated version. Also, they are to acknowledge receipt of the instructions by signing the notice and emailing to SignalMedical2017@gmail.com or faxing to 810-364-7072. The firm has attached the revised labeling to the letter. |  | 1640 | 1,640 | Nationwide Distribution. | Open, Classified |  |  | 1640 | Not_Computer
Z-1835-2017 | 76604 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | INFINITY Internal Hex Healing Abutment, 4.5mm Platform, 4.7mm CW-3mm CH Straight; Product Code: 20145473 |  |  | February 22, 2017 | 2017 | Ace Surgical Supply Co., Inc. | Incorrect device was packaged. | Labeling mix-ups | ACE sent a Customer Notice letter dated February 22, 2017 via certified mail to customers. The letter identified the affected product, problem, risk to health and actions to be taken. Customers were instructed to complete and return the enclosed Business Reply Form. For question call 1-800-441-3100 ext. 208. |  | 87 | 87 | Nationwide Distribution in the states of CA, CT, FL, NY, OH, PA, and UT. | Terminated | August 07, 2017 | 166 | 87 | Not_Computer
Z-2035-2017 | 76961 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | AQUIOS CL Flow Cytometer, Ref no. B30166, Software Versions 2.0 and 2.0.1 |  |  | April 03, 2017 | 2017 | Beckman Coulter Inc. | Beckman Coulter has confirmed that the Export feature located in the Results area of the AQUIOS System software is not correctly exporting Tetra Combo CRD files. | Other | An Urgent Medical Device Recall letter was sent on 4/3/17 to inform customers that area of the AQUIOS System software is not correctly exporting Tetra Combo CRD files. The letter informs the customers of the impact, action, and resolution. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Customer Support Center via website, http://www.beckmancoulter.com/customersupport/support, Via phone, call 1-800-369-0333 in the US. Customers outside the US are instructed to contact their local Beckman Coulter Representative. |  | 188 | 188 units total (24 units in US) | Worldwide - US Nationwide in the countries of: Angola, Australia, Belgium, Botswana, Canada, Chile, Columbia, Czech Republic, Denmark, Finland, France, Germany, Ghana, Israel, Italy, Malaysia, Nambia, Netherlands, Pakistan, Panama, Poland, Portugal, Romania, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, Uganda, United Kingdom and Zambia | Terminated | August 07, 2017 | 126 | 188 | Computer
Z-1895-2017 | 77011 | giemsa stain | HYF | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Shandon Rapid Chrome Kwik-Diff Kit Part Number: 9990700 |  |  | March 15, 2017 | 2017 | Richard-Allan Scientific Company | Methylene Blue, Solution #3 in the Kwik Diff Kit was not Methylene Blue but Crystal Violet instead. | Mixed-up of materials/components | On 3/15/2017, URGENT MEDICAL DEVICE RECALL letters were sent to the affected consignees. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. ACTIONS TO BE TAKEN BY THE CUSTOMER/USERPlease contact Richard Allan Scientific for directions on safe disposal or return if customers have affected lots. All other lots of the products, or similar products manufactured by Richard Allan Scientific, are unaffected. Upon identification of the affected lot, please notify Richard Allan Scientific of the quantity and the preferred disposal/return method. If product is being disposed of, evidence of disposal is required to be sent to Richard AllanScientific. If affected product has been used with no issues, please notify Richard Allan Scientific of the quantity that is unavailable for return or disposal.Please use the attached Recall Return Response Acknowledgement &amp; Receipt Form. |  | 390 | 390 | Worldwide Distribution -- US: CA, CO, CT, DE, GA, HI, IA, IL, KS, MA, MS, NC, NH, NM, NY, OH, PA, PR, RI, UT, VA, VT, WA, and WV; and, the countries of AE, BH, CA, CZ, FR, GB, HK, SA, and TR. | Terminated | July 27, 2017 | 134 | 390 | Not_Computer
Z-1830-2017 | 76804 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Proton Therapy System |  |  | March 21, 2017 | 2017 | Ion Beam Applications S.A. | In specific circumstances, the user is able to restart an aborted irradiation without having to select the correct partial irradiation data. | Software design | A Field Safety Notice was mailed to the customer affected by the issue on 03/21/2017. The Notice informed the user when resuming from a partial irradiation, to ensure that the correct delivered dose has been selected in the partial report dialog box before proceeding with irradiation. |  | 3 | 2 worldwide, 1 in U.S. | Distribution in US (to Louisiana), and France. | Open, Classified |  |  | 3 | Not_Computer
Z-1833-2017 | 76838 | No | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Access Ferritin Calibrators, Part No. 33025, Ref No. B50136C |  |  | March 10, 2017 | 2017 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access Ferritin Calibrators kit because it contains a calibration card with an incorrect S5 calibrator concentration. | Other | Beckman Coulter sent an Urgent Medical Device Recall letter dated March 10, 2017, to inform them that Beckman Coulter has determined through customer feedback that a portion of the Access Ferritin Calibrators kit lot listed above contains a calibration card with an incorrect S5 calibrator concentration. The S5 concentration value printed on the card and encoded in the barcode is 1.500 (ng/mL) rather than the correct concentration of 1500 (ng/mL). Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Customer Technical Support Center via http://www.beckmancoulter.com/customersupport/support, call 1-800-854-3633 the US and Canada and customers outside the US and Canada are instructed to contact their local Beckman Coulter Representative. |  | 889 | 889 kits total (6 kits in US) | Worldwide Distribution - US Distribution and to the countries of : Algeria, Bangladesh, Belgium, Botswana, Croatia, Czech Republic, Denmark, Djibouti, Egypt, France, Germany, Ghana, Greece, Hungary, Ireland, Italy, Kazakhstan, Lebanon, Lithuania, Morocco, Nepal, Netherlands, Oman, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, and Zimbabwe. | Terminated | June 27, 2017 | 109 | 889 | Not_Computer
Z-1836-2017 | 76384 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Valleylab Laparoscopic Handset; Material # E2750; Laproscopic Handset Non-Sterile Bulk; Material # E2750NSB |  |  | January 26, 2017 | 2017 | COVIDIEN MEDTRONIC | Customer reports of the device handset continuing to operate after release of the activation button. | Device Design | A recall was initiated on January 26, 2017. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 195 McDermott Road, North Haven, CT 06473, Attention: Field Returns Department. Disposition of product within Medtronic control and customer returned devices: Will be reworked. |  | 24703 | 24,703 Units | Worldwide - US Nationwide Distribution in the states of:AR AZ CA CO FL GA IA IL KS KY LA MA MI MN MO MS NC NE NM NY OH OR PA SC SD TN TX VA WA WV | Open, Classified |  |  | 24703 | Not_Computer
Z-1837-2017 | 76855 | No | PAX |  |  | Neurology | HDE - Humanitarian Device Exemption | Alcon Patient Care Kits |  |  | March 27, 2017 | 2017 | Alcon Research, Ltd. | Alcon is initiating a Voluntary Medical Device Removal because the product supplier has stated that there is potential for sterile packaging to be compromised. | Under Investigation by firm | Alcon sent an Urgent Voluntary Medical Device Recall letters dated March 23, 2017 to customers. The letter identified the affected product, problem and actions to be taken. The letter instructs the consignees to quarantine and return any unused products of the items/lots listed in the letter. For questions contact Customer Service at 1-800-862-5266. |  | 561702 | 561,702 | US Nationwide Distribution | Open, Classified |  |  | 561702 | Not_Computer
Z-1841-2017 | 76824 | No | GFG | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew ACUFEX TRUNAV Retrograde Drill 7MM |  |  | March 15, 2017 | 2017 | Smith &amp; Nephew, Inc. | Potential for the cutting blade to detach from the device and/or the distal drill head to fracture during retrograde reaming | Device Design | Smith &n; Nephew sent an Urgent Medical Device Recall Notice dated March 15, 2017, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return product and to follow instructions on the attached Response Form. Customers with questions were instructed to email FieldActions@smith-nephew.com.For questions regarding this recall call 508-261-3600. |  |  |  | Worldwide Distribution - US (Nationwide) and Internationally to AU, CL, DB, DE, ES, FR, GB, NL, SE | Open, Classified |  |  |  | Not_Computer
Z-1852-2017 | 76642 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | It is recommended for patients with weight/size between 4-22 lbs (1.8-10 kg). |  |  | March 03, 2017 | 2017 | Verathon Inc | Product labeled as GVL 2 Stat (blades), lot GC32825, was physically packaged with the incorrect size device (GVL 1 Stat (blades)). | Process design | Verathon provided a notification letter to their Canadian distributor on March 03, 2017, and distribution of the letter to the rest-of-the-world (ROW) foreign consignees was initiated on March 10, 2017.Verathon intends to send the Urgent Medical Device Recall Glidescope GVL2 Stat letter and Customer Response Form to US consignees by March 31, 2017.Consignees should quarantine and discard any affected product on site, and fill out the attached Recall Response Form and return it to Verathon by fax to (425) 883- 2896, or email to CSNotifications@verathon.com even if they do not have any Stats (blades) subject to the recall. Consignees who have any questions about this voluntary recall, please contact Verathon representative or Verathon Customer Care at (800) 331-2313 (U.S. and Canada). If consignees located outside of the U.S. or Canada, please call (425) 867-1348. Consignees can also send email to cservice@verathon.com. |  | 38 | Total 292 boxes (US= 254 boxes; International = 38 boxes) | Worldwide Distribution: USA ( nationwide) and to the countries of : Canada, Chile, Mexico, Panama, Philippines, Qatar ,South Africa, and The Netherlands. | Open, Classified |  |  | 38 | Not_Computer
Z-1854-2017 | 76729 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Peri-Loc(R) VLP 3.5 MM X 6 MM SELF-TAPPING CORTEX SCREW, QTY: (1), STERILE R, REF 71801306 |  |  | March 09, 2017 | 2017 | Smith &amp; Nephew, Inc. | The expiration dates on the labeling are incorrect, the actual expiration dates are before the date on the labeling. | Labeling Change Control | Smith &amp; Nephew sent an Urgent Medical First Recall Notice dated March 9, 2017, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested the quarantine and return of the devices, and for the distributors to notify their consignees and ensure the same actions. International consignees were notified by email. Consignees with questions were instructed to contact FieldAction@smith-nephew.com.For questions regarding this recall call 978-749-1440. |  | 1 | 1 unit | Worldwide Distribution - US (Nationwide) Internationally to Colombia, Denmark, Finland, France, Great Britain, Italy, and Singapore | Open, Classified |  |  | 317 | Not_Computer
Z-2166-2017 | 76770 | No | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SHBG IMMULITE 2000/IMMULITE 2000 XPi |  |  | January 31, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc. | Certain lots of Sex Hormone Binding Globulin (SHBG) do not meet the thirty (30) day open vial adjustor stability claim when reconstituted and stored at 2-8ÂżC, as published in the Instructions for Use (IFU). Adjustors stored at 2-8C after 14 days may show a bias in patient samples and quality control falling outside of published ranges. Lookback not recommended. | Nonconforming Material/Component | Siemens issued an Urgent Medical Device Correction (UMDC) to customers using the affected SHBG adjustor lots. Customers were advised to store the affected reconstituted adjustor lots at 2-8C or -20C for up to 14 days, after which they must be discarded. The IMMULITE/IMMULITE 1000 SHBG is not affected by this issue. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative. For questions regarding this recall call 914-631-8000. |  | 9100 | 9100 kits | Nationwide Distribution | Open, Classified |  |  | 9100 | Not_Computer
Z-2071-2017 | 77008 | No | PGW | Stereotaxic instrument. | Neurology | Ear Nose &amp; Throat | 510(k) | Stryker PROFESS Registration Kit. |  |  | March 10, 2017 | 2017 | Stryker Instruments Div. of Stryker Corporation | Product shipped proximate to or past the expiration date listed on the product label. | Storage | Stryker sent an " URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated March 10, 2017, to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Actions to be Taken by the Customer/User: 1. Immediately review this Recall Notification. 2. Immediately check all stock areas and/or operating room storage for affected equipment. Quarantine and discontinue use of any recalled PROFESS Registration Kits. 3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this Notification and identify how many affected items are currently in your inventory. Please complete and return the BRF even if you dont have any affected product on hand. Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in the Notification. 4. If you have further distributed this product, please forward this Notification and the attached BRF to all affected locations. Please indicate each location on the BRF. 5. Fax the completed BRF to Stryker Instruments Regulatory Department at 866-521-2762, or scan and email a copy to StrykerInstrumentsRecalls@stryker.com 6. If the BRF for your facility indicates that recalled product is currently on hand, a FedEx label will be emailed to you. This shipping label should be used to return recalled product. Upon receipt of the recalled product(s), replacement products will be shipped to your account. For further questions please call (269) 323-5330. |  | 65 | 65 | US (nationwide) Distribution to the states of : CA, CT, GA, IL, IN, MA, MD, NC, OH, PA, TN, and TX. | Terminated | July 21, 2017 | 133 | 65 | Not_Computer
Z-1870-2017 | 76927 | No | FCS | Fiberoptic light ureteral catheter. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Bush DL Ureteral Illuminating Catheter Set |  |  | March 07, 2017 | 2017 | Cook Inc. | Product may overheat, melt, and burn patient. | Device Design | On 3/7/2017, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Action To Be Taken:1. Examine inventory immediately to determine if you have affected products and quarantine affected products.2. Return the affected products to Cook Medical with a copy of the Acknowledgement and Receipt Form toreceive a product credit.3. Even if you do not have affected products on hand, you must still complete the Acknowledgement andReceipt Form and send it via fax (812.339.7316) or email to (fieldactionsna@cookmedical.com).4. Complete and return the Acknowledgement and Receipt Form attached to this letter within 30 days.5. Report adverse events to Cook Medical Customer Relations by phone at 800.457.4500 or 812.339.2235,Monday through Friday between 7:30 am and 5:00 pm (Eastern time), or by email viacustomerrelationsna@cookmedical.com.This notice must be shared with appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred.Should you have any questions or concerns, please contact Cook MedicalCustomer Relations at 800.457.4500 or 812.339.2235. |  | 19756 | 19756 | The devices have been distributed within the United States and the following countries: Bahamas, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Ireland, Israel, Italy, Jersey, Kuwait, Mexico, Morocco, New Zealand, Norway, Panama, Paraguay, Portugal, Romania, Saudi Arabia, Spain, Sweden, Switzerland, Trinidad and Tobago, United Arab Emirates and the United Kingdom. | Open, Classified |  |  | 19756 | Not_Computer
Z-1868-2017 | 76508 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | SmartTools Knee System |  |  | February 20, 2017 | 2017 | Orthosoft, Inc. dba Zimmer CAS | There has been an increase in the number of complaints regarding bent or broken Drive Pins of the Validation Tool manufactured with drawing Revision M to P | Other | Zimmer Biomet sent an Urgent Medical Device Recall - Lot Removal letter dated February 20, 2017, sent to the affected distributors and hospital Risk Managers and Surgeons via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Distributor responsibilities include: This letter is initiating Phase I of the lot specific field removal of the iASSIST Validation Tool field removal. You are receiving this letter because our records indicate that you have received an affected product that needs to be corrected. Zimmer Biomet is currently making preparations for replacement activities to follow. This document is provided to alert all users of the potential issue and to highlight proper usage of the instrument per the existing surgical techniques in order to minimize the chance of any failure pending a replacement. A separate field removal notification will be issued with detailed instructions in April 2017. You will be notified when a replacement is available. Do not return any product at this time as a part of this field action.Your Responsibilities 1. Review this notification and ensure affected team members are aware of the contents. 2. The affected products can continue to be used until replacements are available. To minimize the chances of bending or breakage during use, please follow the iASSIST Knee Surgical Technique (Ref. 97-9001-101-00 Rev 9) and/or iASSIST Knee Surgical Technique (2-Pod Version) (Ref. 97-9001-004-00 Rev 2), specifically the following warning on pages 36 and 37, respectively. 3. Inspect affected devices before and immediately after use to confirm that the Drive Pins are not bent or broken. In case of breakage, the Drive Pin head will disassemble, as shown below. In the unlikely case of a breakage, make sure that both parts are retrieved from the wound. 4. Please keep Zimmer Biomet |  | 84 | 84 | Nationwide Distribution | Open, Classified |  |  | 84 | Not_Computer
Z-1865-2017 | 76745 | No | OKV | Obstetric-gynecologic general manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | Vacutainer Tubes |  |  | March 15, 2017 | 2017 | Avid Medical, Inc. | Avid Medical Inc has received a complaint was received from an end-user indicating the vacutainer component 519700 (BD: 367863) expiring 12/16 was discovered inside of a kit labelled with an expiration date after 12/16. Upon receipt of complaint, AVID Medical promptly initiated an internal investigation and root cause analysis. Investigation found mixed component lots were contained in product 519700 component case. This occurrence resulted in the incorrect lot number and expiration being attributed to all work orders produced using the affected component. This issue was immediately contained and corrected. A full inventory analysis was conducted reviewing all vacutainer component inventories. The issue was found to affect a total of five vacutainer components. | Employee error | Avid Medical has identified all impacted products along with distribution centers and end-users as addressed in the tables supplied within this notification. An urgent medical device notice of correction will be issued to alert end-users about the inaccurate expiration date on convenience kits containing the vacutainer component(s) (See attachment). A correction response form will be used to effectively manage all distributer/end-user confirmation of required actions (See attachment). Avid Medical will make three attempts to confirm appropriate actions have been taken and acknowledged by each distributer/end-user for all affected product remaining in inventory. End-users will be provided a label template or actual labels to affix to kits in their inventory to alert users to confirm the vacutainer component expiration date prior to component use.Replacement vacutainer components will be offered for accompanied use with the affected kits. Execution of the correction will be conducted with direct oversight of the associated Avid Medical account representative. Avid Medical has determined that a notification /correction is appropriate because the convenience kits are critical and customized to hospital needs. Most of the impacted kits are sterile and are built in a configuration which is customized to hospital operating procedures according to each hospitals indicated requirements; therefore, new kit production and rework of affected kits would result in lack of critical supplies for the affected end-users. |  | 25308 | 25308 | All products are distributed through distribution centers confined to the US. All end-users for these products are also solely located in the US. | Open, Classified |  |  | 25308 | Not_Computer
Z-2085-2017 | 76992 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Systems Insulin Calibrator; Test Code: Calibrator IRI; Catalog Number: 04618899 |  |  | April 05, 2017 | 2017 | Siemens Healthcare Diagnostics, Inc | Current lots of ADVIA Centaur/XP/XPT/CP Insulin (IRI) assay recover approximately 40% higher than the World Health Organization (WHO) 1st IRP 66/304 standardization based on slope values. Other performance characteristics are being met per the assay Instructions for Use. Overall risk to health is negligible and Siemens is not recommending a review of previously generated results. | Nonconforming Material/Component | Siemens issued a UMDC to ADVIA Centaur/XP/XPT/CP IRI customers on April 5, 2017. They were informed of the issue &amp; that they may continue use of current product but should discontinue when corrected product is available (reagent lots ending in 201 with cal lots ending in 02). A shift down will occur with the corrected product. BioRad control ranges will be posted to QCnet.com for the corrected reagents. |  |  |  | USOUS: Afghanistan, Albania, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium,Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark,Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Georgia,Germany, Greece, Guadeloupe, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq,Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Latvia, Lithuania, Luxembourg, Macedonia,Malaysia, Malta, Martinique, Mexico, Myanmar, Nepal, Netherlands, New Caledonia, NewZealand, Nicaragua, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal,Qatar, Republic of Korea, Romania, Russian Federation, Saudi Arabia, Senegal, Serbia,Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan,Thailand, Tunisia, Turkey, Turkmenistan, United Arab Emirates, United Kingdom, Uruguay,Vatican, Vietnam, Austria, Belgium, Bulgaria, Switzerland, Cyprus, Czech Republic, Germany,Denmark, Estonia, Spain, Finland, France, Great Britain, Hungary, Italy, Luxembourg, Latvia,Netherlands, Norway, Greece, Croatia, Poland, Portugal, Romania, Sweden, Slovak Republic,Turkey | Open, Classified |  |  |  | Not_Computer
Z-1867-2017 | 76920 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | 216 Wilson Frame Pads, Model # 5319-37; |  |  | March 28, 2017 | 2017 | Mizuho OSI | Mizuho has discovered that the 5319-37 Wilson Frame Pads have a design deficiency in the construction of their mounting sleeves that may allow the patient to contact the underlying carbon fiber structure of the Wilson frame and could lead to a pressure injury. | Process control | Letters were sent to customers and distributors on March 28, 2017 by Federal Express or UPS. Information is also posted on firm's web site. Frames having 5319-37 pads should not be used and all of the affected pads retrieved and destroyed. |  | 216 | 216 sets | United States, Canada, Japan, Australia, New Zealand, Great Britain, Sweden, Spain, Germany, Switzerland, Kenya, Saudi Arabia, Hong Kong, South Korea. | Terminated | August 15, 2017 | 140 | 216 | Not_Computer
Z-1871-2017 | 76775 | No | EYZ | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SURESTEP Tray, Bardex I.C. Complete Care Drainage Bag, 2000 ml Drainage Bag, 16 Fr., Sterile. |  |  | March 09, 2017 | 2017 | C.R. Bard, Inc. | Misbranded: Product labeled, Not made with natural rubber latex, may in fact contain natural rubber latex. | Error in labeling | Consignees were notified by letter on/about 03/09/2017. |  | 5030 | 5030 units | CA, CO, FL, GA, IA, IL, IN, KY, LA, MI, MN, MS, NC, NE, NJ, NM, OH, OR, PA, RI, TN, TX, VA and WA. | Open, Classified |  |  | 5030 | Not_Computer
Z-1869-2017 | 76446 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Elcam Minimal Residual Volume Luer-activated Swabable-stockcock (MRVLS) 4w gxl:w1/r rbls |  |  | February 02, 2017 | 2017 | Elcam Medical, Inc. | The stopcock was found to have punctures. | Package design/selection | Elcam Medical sent their customers a recall letter/return response form dated February 2, 2017. Elcam is asking customers to examine their inventory for any affected products and to complete the return response form provided. For questions call 201-457-1120 ext. 102 or 201-457-1120 ext. 112. |  | 810 | 810 | U.S Nationwide Distribution | Open, Classified |  |  | 810 | Not_Computer
Z-1872-2017 | 76875 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PGCL [poly (glycolide-co-caprolactone)] Absorbable Suture, Manufacturer CP Medical, Portland OR, Sterile, Product Codes L923, L459, L493, L762, L463, L922, L495, L463, L497, L494, L493 |  |  | February 02, 2017 | 2017 | CP Medical Inc | The tensile strength minimum as directed by USP &lt;881&gt; was not met over the length of the product's shelf life. | Unknown/Undetermined by firm | The affected distributors/customers were contacted directly via telephone or by e-mail on February 6-7, 2017. Those end users not contacted by CP Medical directly will be contacted by their distributor |  | 6048 | 504 boxes/12 sutures - 6,048 individual sutures | AL, AZ, AR, CA, CO, CT, FL, ID, IL, IN, KY, LA, MT, NE, NJ, NM, NY, OK, OR, SC, TX, WA and South Africa | Open, Classified |  |  | 6048 | Not_Computer
Z-1873-2017 | 76744 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Laerdal Compact Suction Unit Âż 4 under the following labels: 1) 300 ml (3.3 lbs) canister version - CATALOG #(s): 88006101 - LCSU 4, 300 ml Complete unit (US), 88006103 - LCSU 4, 300 ml Complete unit (UK), 88006140 - LCSU 4, 300 ml Complete unit (EUR), 881151 - Replacement unit, 88006201 - LCSU 4, 300 ml, RTCA Complete unit (US), 88006203 - LCSU 4, 300 ml, RTCA Complete unit (UK), 88006240 - LCSU 4, 300 ml, RTCA Complete unit (EUR) and 881152 - Replacement Unit, RTCA; 2) 88005101 - LCSU 4, 800 ml Complete unit (US), 88005103 - LCSU 4, 800 ml Complete unit (UK), Âż 88005140 - LCSU 4, 800 ml Complete unit (EUR), 88005201 - LCSU 4, 800 ml, RTCA Complete unit (US), 88005203 - LCSU 4, 800 ml, RTCA Complete unit (UK), 88005240 - LCSU 4, 800 ml, RTCA Complete unit (EUR) |  |  | January 08, 2017 | 2017 | Laerdal Medical Corporation | The On/Off functionality in the LCSU 4 is controlled by electronic components, one of which might be damaged. | Nonconforming Material/Component | On January 8, 2017 Laerdal Medical distributed Urgent Device Recall Notification Letters to their customers thru Stericycle. The documents were dispatched via standard UPS ground delivery. Laerdal Medical has discovered that the LCSU4 may turn on by itself when connected to a power source and/or may not be able to be turned off with the on/off button without removing the battery or external power source. You can continue to use your suction units until your replacement arrives by following the steps: Exception: Some LCSU 4 units have a RW label on the bottom. These units have been reworked by Laerdal Medical and are not affected by this Field Safety Notice, even if they have manufacturing dates within the 20150501 to 20161028 range. **Check On/Off functionality**You can run a quick test to identify if the unit has failing On/Off functionality:Turn unit on;Turn unit off again.If the unit cannot be turned off using the On/Off button, the unit contains a damaged component. The device should be taken out of service.The safety and effectiveness of the device is not affected by the failure of the On/Off button as long as it has power.Note that even if the unit passes this test, the problem may occur later. The actions below must be taken to ensure continued safe use of the LCSU 4 until you receive a replacement unit.**Actions to be taken to ensure continued safe use of the LCSU 4 until you receive a replacement unit**You can continue to use your LCSU 4 unit until you receive a replacement unit, but it is recommended to implement the following extra safety measures:Store the LCSU 4 unit without the battery connected so that it cannot turn on by itself.Perform routine checks of the On/Off functionality by turning the unit on and off using the On/Off button.All units with failing On/Off functionality should be removed from service.** Actions to be taken if the On/Off functionality fails during clinical use**If the On/Off functionality fai |  | 4377 | 4,377 units | Nationwide | Open, Classified |  |  | 4377 | Not_Computer
Z-1875-2017 | 76987 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 18L6 HD transducer on the ACUSON S Family ultrasound systems with software versions VD10A or VD10C; |  |  | March 30, 2017 | 2017 | Siemens Medical Solutions USA, Inc. | When scanning with the 18L6 HD transducer on the ACUSON HELX" Evolution with Touch Control, the ultrasound system may display a triple image or an image with a dark band. For the triple image issue, the system repeats one-third of the aperture, but does not display the full field of view. | Software design | Siemens sent an Customer Safety Advisory Notification via certified mail on March 30, 2017. OUS consignees are notified through the implementation of the field correction with the distributors for all countries. Letters advised users of the issues with software version VD10A or VD10C and provided steps to avoid the potential risk. Once available, all systems will be updated with corrected software. For further questions, please call (650) 694-5398. |  | 2045 | 2,045 systems | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : United Arab Emirates, Argentina, Armenia, Austria, Australia, Azerbaijan, Bahrain,Bangladesh, Belgium, Bolivia, Brazil, Canada, Chile, Columbia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Ethiopia, Egypt, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovakia,Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, Venezuela, Vietnam and United Kingdom, | Open, Classified |  |  | 2045 | Computer
Z-1832-2017 | 76807 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) Pump Driveline Splice Kit |  |  | March 10, 2017 | 2017 | Heartware | In April 2015, HeartWare, now a part of Medtronic, notified users of a safety issue related to the HVAD Pump Driveline Splice Kit. Medtronic now intends to send a follow-up communication to identified clinicians informing them that a new, improved splice kit is available. | Device Design | HeartWare, now a part of Medtronic sent an Urgent Medical Device Recall letter on April 2015 and a follow-up letter on March 2017, to customers via overnight mail and/or email.. The letters identified the affected product, problem and actions to be taken. Customers were instructed to review the recall notice, forward the notice to individuals within their organization who need to be aware of the notice, complete, sign, and return the Acknowledgement Form . For questions contact your local HeartWare representative. |  | 9 | 9 units | Worldwide Distribution - US Nationwide in the states of FL, MO, TN, KY, NY and the country of Germany. | Open, Classified |  |  | 9 | Not_Computer
Z-1885-2017 | 77069 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 01.4mm x 457mm (18in) K-Wire, Threaded Blunt, Ref 9080B-18T, Qty: 1 per pack, Material: SS 316L, Full Temper, per ASTM F138, Amendia, 1755 West Oak Parkway, Marietta, GA 30062 |  |  | April 06, 2017 | 2017 | Amendia, Inc | Mislabeled guide wires - It was discovered that some Guide Wires labeled as stainless steel were actually manufactured from nitinol. | Process design | Amendia sent an Urgent Medical Device Recall Letter dated April 6, 2017, and April 25, 2017, to all affected customers. The letter requested that they check inventory, quarantine product for return to Amendia, and requested a sub-recall. The letter also included a reply form to be returned to Amendia. Customers with questions were instructed to call Amendia Customer Service at 770-575-5224.For questions regarding this recall call 404-610-7215. |  | 15 | 15 units | US distribution to GA, KS &amp; TX | Terminated | May 30, 2017 | 54 | 15 | Not_Computer
Z-2091-2017 | 77072 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | VORTRAN Automatic Resuscitator, VARPlus; |  |  | April 20, 2017 | 2017 | Vortran Medical Technology 1, Inc | Vortran received a customer complaint that the spin-nut-DSS connector was assembled in the reverse direction on the oxygen tubing; making the connector unable to be screwed on due to the wrong thread direction. | Process control | Medical Device Recall letters were sent on April 20, 2017.Letters explained the reason for recall and listed the specific lots sent to each distributor along with dates, PO numbers and quantity. Vortran Medical asks that each distributor provide a list of customers who received the affected product so that they can notify them directly of the recall. |  | 35 | 35 cases (10 each) | CA, MS, FL | Open, Classified |  |  | 49 | Not_Computer
Z-2056-2017 | 77116 | No | LOD | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Cobalt HV Bone Cement |  |  | April 13, 2017 | 2017 | Zimmer Biomet, Inc. | Six complaints were filed regarding loss of the Zimmer Biomet applied seal of a sterile Tyvek packaging. A breach in the packaging could lead to loss of sterility of the device. | Under Investigation by firm | On 4/13/2017 URGENT MEDICAL DEVICE RECALL  LOT SPECIFIC notifications were sent to the affected consignees via e-mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 4238 | 4238 | TX.JapanMexico | Open, Classified |  |  | 4238 | Not_Computer
Z-2059-2017 | 77113 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Bone Dowell Harvest Tubes |  |  | April 12, 2017 | 2017 | Zimmer Biomet, Inc. | mislabeled as 8mm tube but it isactually 9 mm in size. A 9mm subcomponent was erroneously substituted and etched as 8mm for the actual8mm subcomponent part. | Nonconforming Material/Component | On 4/12/2017 URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via e-mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. |  | 49 | 49 | No US distribution.JAPANNETHERLANDSSINGAPORESWITZERLAND | Open, Classified |  |  | 49 | Not_Computer
Z-2068-2017 | 77094 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Regenerex Patella |  |  | March 22, 2017 | 2017 | Zimmer Biomet, Inc. | pegs shearing post-operatively | Under Investigation by firm | On 3/22/2017 URGENT MEDICAL DEVICE RECALL REMOVAL notifications were sent to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Distributors, Sales Representatives, and Distributor Operation Managers Your Responsibilities 1. Review this notification and ensure affected team members are aware of the contents. 2. Immediately locate and quarantine affected product in your inventory. 3. Complete the Certification of Acknowledgement portion of Attachment 1  Inventory Return Certification Form. a. Return a digital copy to corporatequality.postmarket@zimmerbiomet.com within three (3) days. 4. Immediately return all affected product from your distributorship and affected hospitals within your territory along with a completed Attachment 1  Inventory Return Certification Form to Zimmer Biomet. a. Request a Return Authorization Number via email to rgarequest@zimmerbiomet.com or through FAST/SMS. Be sure to specify RECALL as the RGA type when requesting. b. For each return, send a copy of Attachment 1 to corporatequality.postmarket@zimmerbiomet.com. c. Include a hardcopy of Attachment 1 with your shipment for immediate processing. d. Mark the outside of the returns box(es) clearly with RECALL. 5. Note that any hospitals and surgeons that received direct shipments of this product from Zimmer Biomet or were consigned products, will be sent a copy of the Risk Manager and Surgeon Field Action Notice directly. It is important that you review the list of hospitals and surgeons included with the email notification sent to your facility to identify additional accounts Zimmer Biomet has not notified. Using the Additional Accounts Form provided with the email notice sent to your facility, return contact information for any additional hospitals and/or surgeons that may hav |  | 8154 | 8154 | US Nationwide.distribution and Canada. | Open, Classified |  |  | 8154 | Not_Computer
Z-1878-2017 | 76740 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | uNion Cervical Plate System |  |  | March 07, 2017 | 2017 | Ulrich Medical USA Inc | Surgical technique was revised to prevent static plate screws blocking mechanisms from disassociating. | Use error | On March 7, 2017, the recalling firm sent letters by mail informing customers of the field correction along with the updated surgical technique which contains the enhanced cautionary statement. Customers were instructed to destroy older versions of the surgical technique. The letters also contained a link to the recalling firm's website on which the recalling firm posted the current, updated copy of the surgical technique. |  | 4854 | 4,854 plates | US Nationwide in the states of: AL, AR, AZ, CA, CT, FL, GA, ID, IL, KS, LA, MO, MS, NY, OH, OK, OR, TN, TX, and WI | Open, Classified |  |  | 4854 | Not_Computer
Z-1876-2017 | 76930 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL Trypticase Soy Agar with 5% Sheep Blood (TSA II) and Chocolate II Agar (GC II Agar with Hemoglobin and IsoVitaleX)  I Plate |  |  | March 27, 2015 | 2015 | Becton Dickinson &amp; Co. | BD has confirmed that one lot of TSA II and Chocolate Agar I-Plates may fail to recover Haemophilus influenzae 10211 and Haemophilus parainfluenza 51505 on the Chocolate Agar side of the plate. All other QC organisms are correctly recovered. | Under Investigation by firm | Recall notification was initiated on March 27, 2015. US customers were contacted via letters sent by UPS. Customers were requested to return response forms indicating acknowledgement of the recall and quantity of remaining product that was at their facility. The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 100% of non-respondents to verify receipt of the customer letter. |  | 1600 | 1600 plates | US Nationwide Distribution in the states of AR, CA, CO, FL, IL, NJ, NV, OR, PA, TX, and WA | Terminated | April 26, 2017 | 761 | 1600 | Not_Computer
Z-1880-2017 | 77053 | No | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | Omega Incline Platform Lift |  |  | April 12, 2017 | 2017 | Savaria Concord Lifts, Inc. | Savaria Concord Lifts Inc has decided to initiate this recall of its Omega Incline Platform Lift because bolts can come loose causing the carriage to separate from the trolley and the unit to fail | Other | Savaria Concord Lifts Inc contacted the dealers to insure they received the parts and instructions and ask if they understand what is requested. If the dealers indicated they could no longer find the required parts new kits and instructions were sent on April 2015. Savaria Concord Lifts Inc contacted the dealers via e-mail and phone. For further questions, please call (800) 661-5112. |  | 74 | 74 | Worldwide Distribution - US Distribution and to North America. | Open, Classified |  |  | 74 | Not_Computer
Z-1881-2017 | 75614 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Navitrack System - OS Knee Universal, Orthopedic Stereotaxic Instrument |  |  | October 06, 2011 | 2011 | Orthosoft, Inc. dba Zimmer CAS | Zimmer CAS voluntarily conducted a retrospective recall of the Navitrack System - OS Knee Universal software ORTHOsoft-UniTkr-2.3.2.6, due to a calibration sequence crash. | Software design | Zimmer sent a Medical Device Field Notification letter dated December 6, 2011, to all affected customers. This is a retrospective report of a correction taken on 6 December 2011. The software issues were fixed in the new software version 2.3.3.0. Recall notices were distributed on December 6, 2011. Telephone: 18663367846 (US &amp; Canada) or 5143958883 (International),Email: cassupport@zimmercas.com Customers with questions should call 1-866-336-7846 (US &amp; Canada) 514-395-8883 (International).For questions regarding this recall call 574-372-4487. |  | 105 | 105 | Worldwide Distribution - US including AZ CA CO FL GA IL IN KS MD MI MN MS OH OR PA SC SD TX VA WA and WIInternationally to Canada, Australia Austria Colombia Denmark France Germany Korea New Zealand Russia South Africa and Thailand | Terminated | May 08, 2017 | 2041 | 105 | Computer
Z-1882-2017 | 77042 | No | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | The Savaria Stairfriend is a Residential, commercial, indoor, stair lift for curved stairs. |  |  | April 12, 2017 | 2017 | Savaria Concord Lifts, Inc. | Savaria Concord Lifts has decided to initiate a recall of the Stairfriend, because of product defect. The product defect relates to poor quality of welding holding the seat post . | Nonconforming Material/Component | The firm plans to notify its customers thru the phone and also via the dealers it sells it product too. |  | 226 | 226 | North America / South America, please see attached list. | Open, Classified |  |  | 226 | Not_Computer
Z-1886-2017 | 76739 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker SmartLife Large Aseptic Housing (7126-120-000) |  |  | February 28, 2017 | 2017 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments issued a notice of correction for an update to cleaning practices for the Stryker SmartLife Large Aseptic Housings following complaints of deterioration to the housing and the top and bottom housing separating at the weld. Using cleaning practices outside of the HIFU could lead to housing separating at the weld. | Use error | Stryker Instruments issued a written corrective notice via mail including a reminder to follow the IFU cleaning procedures for the Stryker SmartLife Large Aseptic Housings (7126-120-000). A small subset of customers have experienced product issues based on cleaning practices that are causing a deterioration of the housing, and speration of the top and bottom of the housing at the weld. Customers are instructed as follows:Several factors can adversely influence the structural integrity of the device. Below are some suggestions for you to review from the Heavy Duty Care Instructions/Instructions for Use to help reduce undue strain on the housing material:" Stryker recommends following sterilization/autoclaving instructions in the Care Instructions."Stryker has identified that some customers are applying a pre-treatment to the device, please note that the Care Instructions advise the following that pertain to the application of pretreatments (Reference Page 12 of the Heavy Duty Care Instructions): ALWAYS make sure the detergent solution is completely rinsed from the interior and exterior of the equipment before drying the equipment. ALWAYS prepare the detergent solution according to the manufacturers recommendations. Pay particular attention to the concentration used and the total dispersion. "Stryker recommends when using an auto-washer, assure the appropriate parameters (rinse time/settings) are used as instructed in the Care Instructions (Reference Page 21 of the Heavy Duty Care Instructions). Additionally: Unload the washer-disinfector and visually inspect the equipment for remaining soil or detergent solution. If soil or detergent solution remains, repeat the cleaning procedure using fresh detergent solution. Utilizing cleaning practices outside of the IFU could lead to the housing separating at the weld. The health risk associated with the separation of the housing at the weld, is a potential breach in sterility, which may requ |  | 302 | 302 | Domestic: NoneForeign: France, Germany | Open, Classified |  |  | 302 | Not_Computer
Z-1538-2017 | 76600 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare HVAD System Controller |  |  | February 03, 2017 | 2017 | Heartware | Product recall of all HVAD Controllers currently on the market and to exchange them for the next generation HVAD Controller (Controller2.0) | Software design | The firm will release a communication to their consignees stating the following: As a reminder, as with all HVAD Controllers, continue to reinforce the following with your patients and staff at all opportunities:" Patients should continue to have a backup HVAD Controller ready at all times in the event of a primary HVAD Controller failure." Staff only: The driveline extension cable is to be used during the pre-implant test only. It is not intended to be used after the pump is implanted in the patient.Hospital and Clinician Actions (to be executed in the following order):1) Review the enclosed notice and forms, and forward the notice to those individuals within your organization who need to be aware of its contents.2) Complete, sign, and return the Acknowledgement Form to HeartWare within thirty (30) days of receipt of this letter.4) Quarantine and replace affected HVAD Controllers, DC Adapters, Instructions for Use, Emergency Responder Guides and Patient Manuals in hospital inventory after training is complete.For every patient, quarantine and replace the following under clinical supervision in an environment where appropriate support equipment is readily available:- Primary and Backup HVAD Controller;- Affected DC Adapters; and- Patient Manual and Emergency Responder Guide.Clinicians are reminded not to perform an HVAD Controller exchange during an active electrical fault alarm as the HVAD Pump will be running a single stator. If an electrical fault is present, download patient log files and contact your HeartWare representative to resolve the electrical fault before executing the controller exchange.5) Return all quarantined HVAD Controllers and DC Adapters to HeartWare. Your HeartWare representative will assist you with this process.6) Completion Form. Once affected product in inventory has been identified and returned, complete and return the attached Completion Form to CON2.0@heartware.com or your HeartWare representative no late |  | 21488 | 8,343 US and 13,145 INTERNATIONAL | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Lithuania, , Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, UK, and Vietnam. | Open, Classified |  |  | 21488 | Not_Computer
Z-1887-2017 | 76551 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products VANC Reagent, REF/Product Code 680 1709, Unique Identifier No. 10758750006731, Rx ONLY, IVD --- Ortho Clinical Diagnostics - USA -- Ortho-Clinical Diagnostics Pencoed, CF35 5PZ, UK ---- The VITROS VANC reagent is for ORTHO Clinical Diagnostics manufactured by Siemens. The VANC product has a maximum on-analyzer stability of &lt; 7days. The on-analyzer stability of each VITROS VANC Reagent pack is determined by the system. Once the VANC reagent pack is loaded on the system, the on-analyzer stability time is continuously adjusted based upon both the amount of time the reagent pack is on the analyzer and the number of tests remaining in the pack. The method by which the on-analyzer stability of the VITROS VANC Reagent is determined is stated in the reagents Instructions for Use. |  |  | February 21, 2017 | 2017 | Ortho-Clinical Diagnostics | Potential for biased results generated using the specific lots of VITROS VANC Reagent. This can occur after the reagent packs are loaded onto the analyzer and stored within the on-analyzer stability date. | Under Investigation by firm | Ortho Clinical Diagnostics sent an IMPORTANT PRODUCT CORRECTION NOTIFICATION letter dated February 21, 2017, was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (for PO Boxes only) to all customers who received the affected VITROS Chemistry Products VANC Reagent lots to inform them of the potential for biased results and to request that they immediately discontinue use of this product. Foreign affiliates were informed by e-mail on February 21, 2017, of the issue and instructed to notify their consignees of the affected product with this issue and the requirement to immediately discontinue use of the product. The communication instructed customer to immediately discontinue use of the affected product. In order to minimize the disruption to the laboratory, the communication also provided the instructions for testing using the affected VANC lots until replacement product was received. If you have any further questions, please contact our OrthoCare Technical Solutions Ctr. at 1(800) 421-3311. For further questions regarding this recall, please call (908) 218-8776. |  | 616 | US: 2740; OUS: 616 | Worldwide Distribution - US Distribution including Puerto Rico and to the countries of : Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Poland, SpainSingapore and Venezuela. | Open, Classified |  |  | 616 | Not_Computer
Z-1883-2017 | 75013 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Add-On Bag Access Device, Model No. 10013365 |  |  | November 17, 2016 | 2016 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Add-On Bag Device due to reports of separation and/or leakages between the spike port and the drip chamber spike. A separationor leak can cause a delay of infusion, an interruption of infusion, exposure tomedication or hazardous infusates, or underinfusion. | Other | CareFusion sent an Urgent Medical Device Recall Notification letter dated November 2016, to inform their customers that CareFusion is recalling the SmartSite Add-On Bag Access Device Model 10013365 because of reports of leakages between the spike port and the drip chamber spike. The letter informs the customers of the lot numbers, potential risk and immediate actions to be taken. Customers with questions are instructed to contact CareFusion Support Center at (888) 562-6018 for recall related questions. Customers with adverse event report questions are instructed to contact Customer Advocacy at (888) 812-3266 or by email at customerfeedback@carefusion.com. Customers are instructed to complete and return the enclosed mandatory customer response card. |  | 384923 | 384,923 units | US Distribution | Open, Classified |  |  | 384923 | Not_Computer
Z-1888-2017 | 76846 | No | JSH | Differential culture medium. | Microbiology | Microbiology | 510(K) Exempt | Blood Agar (TSA w/ 5% Sheep Blood) Plate, For In Vitro Diagnostic Use, Remel, Lenexa, KS 66215 |  |  | March 24, 2017 | 2017 | Remel Inc | Confirmed complaint of surface contamination of Listeria monocytogenes. | Material/Component Contamination | On 3/24/2017, the recalling firm sent a recall notice to their customers explaining the reason for the recall and instructing customers to review their stock and discard any affected product. The notice also instructed customers to advise of the recall anyone they may have further distributed the recalled lot to. Attached to the recall notice is a response form which the recalling firm instructed their consignees to fill out and return. |  | 115900 | 1,159/100 plates/packs | US Nationwide distribution in the states of : AR, AZ, CO, IA, IL, KS, KY, LA, MI, MO, MS, MT, NC, ND, NE, NM, OH, OK, SD, and TX and Puerto Rico | Terminated | June 06, 2017 | 74 | 115900 | Not_Computer
Z-1884-2017 | 76842 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 5.5mm Diameter x 500mm CoCr Straight Rod, Model number X077-0420. The product is packaged in plastic bags with one piece per bag. |  |  | April 12, 2017 | 2017 | XTANT MEDICAL INC | The material type on the label may incorrectly state Ti6Al-4V ELI . The rods are composed of Cobalt Chromium. | Labeling mix-ups | XTANT sent an Urgent Medical Device Removal letter dated 4/12/2017 to customers with a Return Response Acknowledgement and Receipt form. The letter identified the affected product, problem and actions to be taken. For questions call 406-570-7028. |  | 97 | 97 units | Worldwide distribution - US Nationwide in the states of: AZ, CA, CT, GA, LA, NV, OH, TX and in the country of Mexico. | Terminated | August 16, 2017 | 126 | 97 | Not_Computer
Z-1889-2017 | 76938 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | 1 mL MedallionÂż Syringe. Catalog Numbers: 701989001, K01-05197P, MSS011-YP, K01-07946P, MSS011, MSS011-LB, MSS011-R, MSS011-Y, MSS011P, MSS011-LBP, MSS011-DG, K08-02926AP, K01-07742P, K10-05457P. |  |  | March 23, 2017 | 2017 | Merit Medical Systems, Inc. | Merit Medical Systems announces a voluntary field action for the 1mL MedallionÂż Syringes due to a defect with the syringe plunger tip. | Nonconforming Material/Component | Consignees began receiving notification via letter on March 23, 2017 to immediately stop using or distributing the affected lots and return them back to Merit. If there are any questions, contact Merit Customer Service at (801) 208-4381. |  | 23032 | 23,032 units | Nationwide and other foreign consignees. No VA/govt/military or VA consignees. | Open, Classified |  |  | 23032 | Not_Computer
Z-1894-2017 | 77075 | No | MMZ | Urological table and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Sonesta S2; Article numbers: 525-S2-110V and 525ÂżÂżS2-220V |  |  | March 08, 2017 | 2017 | Sonesta Medical Ab | The user / service manual for the Sonesta S2 Chair contain wrong information concerning the maximum patient weight. | Labeling mix-ups | Sonesta notified customers (distributors) on March 8, 2017, of the faulty labeling. Corrected labeling was sent to customers, they were requested to apply the new labels and confirm that the correction was implemented. For further information regarding this recall, please call (512) 327-9997. |  | 28 | 28 devices | Worldwide Distribution - US Distribution to the state of : MA., and to the countries of : Canada, Netherlands, Dubai, Taiwan, Lebanon, Italy | Open, Classified |  |  | 28 | Not_Computer
Z-1891-2017 | 77100 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | DOLPHIN INFLATION DEVICE CALIBER INFLATION DEVICE; Catalog Number: CL3030- 018SNA- 018SND 0185NA 0185ND; INTERNAL CODE NUMBER: 0185TR- 0185NA- 018SND 0185NA 0185ND |  |  | April 05, 2017 | 2017 | PEROUSE MEDICAL | Complaints regarding broken blister. | Packaging change control | Consignees were notified of the recall on April 5, 2017. They were informed of the issue and asked to take the following steps: 1. Inspect your stock to determine if you still have in your ownership devices among which the references and batch codes are listed in front page; 2. Control the potential damage of the blister according to the pictures above and the instructions mentioned on the labeling and the Instruction For Use: any damaged and/or cracked blister must be immediately discarded; 3. Use the compliant products without risk, after control. Indeed, the primary packaging will not deteriorate in time. If, to date, there is no crack in the blister, then the integrity of the product is not compromised; 4- Do not use the non-compliant products; 5. Complete and sign the attached reply form, then return it within 5 working days following the reception of this letter, and keep a copy; 6. Contact customer service for the modalities of return, replacement and associated support at +33 (0)4.72.39.74.13 or by fax at +33 (0}3.44.08.17.67 or by email at iblayon@vvgon.com (CC mhpourriere@vygon.com); 7. Communicate the safety information to any person concerned in your company and to all end customers to which the product were transferred. |  | 18431 | 18,431 | US: AZ (shipped through Belgium to USA), MN, TN | Open, Classified |  |  | 18431 | Not_Computer
Z-1890-2017 | 76436 | [6 empty] | MiniMed 620G and MiniMed 640G Insulin Pumps with Software Version 2.6 and below are affected. Catalog Numbers for these insulin pumps are MMT-1510, MMT-1710, MMT-1711, MMT-1712. |  |  | January 31, 2017 | 2017 | Medtronic Inc. | Medtronic has identified a software anomaly that can prevent the internal battery of the pump from charging. When the software anomaly occurs it leads to battery depletion and pump shutdown. | Software design | A urgent medical device recall letter was sent to customers on 1/31/17 to inform them that Medtronic has identified a software issue that could prevent the internal battery of the pump from charging. Customers are informed that should the software issue occur, an alarm is triggered and they will see the following message displayed on the pump screen: "Power error detected 00:00 25 Delivery stopped. Record your settings by uploading to CareLink or write your settings on paper. See user guide. OK." Customers are informed that If they experience the alarm and see the Power error detected message on their pump accompanied by the number 25 , then they are instructed to contact the Medtronic HelpLine team at XXXXX for assistance with troubleshooting the error. |  | 42215 | 42,215 units | Europe, Australia, Japan, Singapore | Terminated | May 04, 2017 | 93 | 42215 | Computer
Z-1897-2017 | 76294 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | BD FlowSmart Set MiniMed Pro-set Rx Only |  |  | December 23, 2016 | 2016 | Becton Dickinson &amp; Company | Patients are reporting high blood glucose (HBG) levels while using the MiniMed Pro-set and also noticed the catheter of the infusion set was bent/kinked. | Device Design | Becton Dickinson (BD) notified their sole customer Medtronic via e-mail on December 26, 2016. BD recommends that Medtronic notify their customers of the situation. For further questions, please call (201) 847-6800. |  | 57750 | 57,750 units | US Distribution to the state of : CA | Open, Classified |  |  | 57750 | Not_Computer
Z-1896-2017 | 76948 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | SurFlo IV Catheter |  |  | August 05, 2015 | 2015 | Terumo Medical Corporation | According to the Terumo (TMC) this device is used to access the human peripheral vascular system for short term use (less than 30 days) for the purpose of infusing fluids into patients. A small percentage of the catheters in the affected lots may have damage to the ETFE catheter which may allow the catheter to break during withdrawal. | Process control | Terumo sent an Urgent Product Bulletin letter dated August 3, 2015. The letter identified the product the problem and the action needed to be taken by the customer. 1. CUSTOMERS MUST IMMEDIATELY DISCONTINUE USE, AND QUARANTINE ANY AVAILABLE STOCK, OF THE LISTED PRODUCT CODES.2. DISTRIBUTORS MUST NOT DISTRIBUTE ANY OF THE LISTED PRODUCT CODES, QUARANTINE ANY AVAILABLE STOCK AND SEND THIS RECALL BULLETIN IMMEDIATELY TO EACH CUSTOMER THAT RECEIVED THESE PRODUCTS TO INFORM THEM OF THIS ISSUE.Terumo is taking this action to assure the safety of our customers and their patients even though it may cause a significant inconvenience. If this creates a critical shortage, please contact us immediately so we may try to provide assistance in procuring replacement or alternative product. For further questions, please call (410) 392-7226 |  | 2478200 | 2,478,200 | US Distribution.This product has been sold to both distribution companies with multiple locations and directly to end users/hospitals in US. | Terminated | May 01, 2017 | 635 | 2478200 | Not_Computer
Z-1899-2017 | 76870 | No | NDC | Predictive pulmonary-function value calculator. | Anesthesiology | General &amp; Plastic Surgery | 510(k) | Accu-Chek Connect Diabetes Management App versions 2.0.0, 2.0.1 and 2.1.0 for iOS and Android |  |  | March 14, 2017 | 2017 | Roche Diabetes Care, Inc. | A program error (bug) in the Bolus Advisor feature, which could result in incorrect bolus advice and a potential insulin over-delivery. | Software design | Roche Diabetes Care, Inc. discovered a software issue in the Accu-Chek Connect app versions 2.0.0, 2.0.1, and 2.1.0 that can result in incorrect bolus recommendations. This software bug may result in errors in the active insulin calculation, creating the potential for a false high insulin bolus recommendation. An insulin bolus delivered based on potentially incorrect bolus advice could lead to serious health consequences such as hypoglycemia. Customer notices were distributed via mail on 03/30/2017, and the firm's website was also updated with recall information on that day. Customers were instructed to do the following: discontinue use of the bolus advisor feature in app versions 2.0.0, 2.0.1, and 2.1.0 (iOS &amp; Android) until the corrected version is available.Âż Provide patients with a copy of the enclosed Urgent Medical Device Correction 17-001.Âż Complete the enclosed faxback form and fax it to Roche Diabetes Care, Inc., at 1-888-912-8456.Âż If you have questions about the information in this notification, please contact Accu-Chek Customer Care by visiting accu-chek.com or by calling 1-800-688-4578.The corrected version of the app is expected to be available on or about April 13, 2017. New patients will be able to begin using the bolus advisor feature once the corrected version is available. Please advise your patients to visit accu-chek.com for updates regarding the availability of the corrected version of the app. |  | 64324 | 64,324 | US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, DE, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV | Open, Classified |  |  | 64324 | Computer
Z-1900-2017 | 76914 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MR GHIATAS Beaded Breast Localization Wire, Product Code: 477201, 479201 |  |  | March 28, 2017 | 2017 | Bard Peripheral Vascular Inc | It may be at risk of containing a non-MR compatible GHIATAS Beaded Breast Localization Wire instead of a MR compatible GHIATAS Beaded Breast Localization Wire. | Other | Bard sent an Urgent Medical Device Recall notification letter dated March 28, 2017 to customers. The letter identified the affected product, problem and actions to be taken. For questions contact your Bard representative. |  | 570 | 570 units | Worldwide Distribution - US Nationwide in the states of TX, FL, MO, NJ, WI, AL, TX, CA, OK, AZ, NC, GA, ID, WA, NY, NM and the countries of: Brazil, Canada, China, Taiwan, Canada, and Belgium. | Terminated | June 29, 2017 | 93 | 570 | Not_Computer
Z-1902-2017 | 76994 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Dressing M/U XL, P/N 02344 |  |  | March 30, 2017 | 2017 | Breg Inc | They may not have been sealed prior to sterilization. Products with unsealed pouches will be non-sterile. | Employee error | Breg sent an Urgent Medical Device Correction Letter to customers on 4/13/17. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:Please take the following actions:1) Determine if any potentially affected products are in your stock by comparing each product's Lot and Date to the Lot Numbers and Expiration Dates in the table above. Refer to Picture 2 for the location of these fields on the product packaging.2) Once you have matched Lot and Date, inspect the product to identify any open (unsealed) pouches. Refer to Picture 1 below.3) Remove any open pouches from your inventory and discard.4) Complete the Response Form included as Attachment A, and email or fax to Breg as follows:a. E-mail: bregfieldcorrection@breg.com: Subject: Sterile Dressing Correctionb. Fax: 800-FAX-BREG (329-2734)- Subject: Sterile Dressing CorrectionCustomers with questions are instructed to contact Breg Customer Care at (800) 321-0607. |  | 2609 | 2,609 units total | US Nationwide Distribution | Terminated | June 23, 2017 | 85 | 2609 | Not_Computer
Z-1892-2017 | 76652 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The TM Ardis Interbody System implant is a convex, straight TLIF or PLIF device for interbody fusion of the anterior column of the spine. TM Ardis is designed for fusing the adjacent bony surfaces and may be used to replace a disc at one or two contiguous levels in L2-S1. The superior and inferior surfaces of the device are textured to provide increased stability and convex to conform to the vertebral endplates. The device also has two slots on the posterior end to mate with the insertion instrument. The TM Ardis device is wholly comprised of Trabecular Metal Porous Tantalum. Surgical instrumentation for use with the proposed system will be fabricated from surgical grade stainless steel and other applicable materials. |  |  | July 07, 2016 | 2016 | Zimmer Trabecular Metal Technology, Inc. | This field action is being initiated following the firm's discovery of a procedural error in the approval of literature resulting in the undo release of certain content.Information was disseminated to distributors and sales representatives during the 2016 National Sales Meeting held February 17-21, 2016. | Unknown/Undetermined by firm | Zimmer Biomet sent an Urgent Medical Device Correction Notice dated July 7, 2016, to all consignees on the list requiring them to destroy all literature and notify Zimmer Biomet of others who they have disseminated the material further. Retraining the sales force on the use of the TM Ardis. The details of the strategy are contained within the Draft Recall Strategy attached to the email transmitting this report. For further questions, please call (973) 576-0032. |  | 499 | 299 units; 200 brochures | Nationwide Distribution | Terminated | May 03, 2017 | 300 | 499 | Not_Computer
t | o | o |  | b | i | g | [18 empty] | Not_Computer
Z-2685-2017 | 77267 | No | JJN | Electrophoresis apparatus for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Multiphor II Electrophoresis System |  |  | May 15, 2017 | 2017 | Perkinelmer Life Sciences, Inc. | Potential for electric shock in the case of a failure to install the grommets, supplied with the system. | Device Design | Perkin Elmer issued URGENT: Field Correction letter on 5/15/17 and accompanied with the GE HealthCare Important Product Information letter The notifications will be delivered to the end customer by email facilitated through PerkinElmer Technical Support group located in Akron Ohio.Technical Support will ensure the customers receive the both letters and understand the recommended actions. Accounts requested to complete the enclosed Response Form.Questions regarding the actions to be taken, as outlined in the GE HealthCare letter, should be directed via email to orderFI@ge.com or via phone (in the US) 011 44 800 515313, (in the UK) 0800 515 313, and (rest of world) +44-800 515 313. |  | 241 | 241 instruments | US Nationwide Distribution | Open, Classified |  |  | 241 | Not_Computer
Z-2703-2017 | 77244 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Carestation 620 A1, Carestation 650 A1, and Carestation 650c A1 |  |  | May 02, 2017 | 2017 | GE Healthcare, LLC | Carestation 620, 650 and 650c has Potential for Elevated FiCO2 and Unexpected System Malfunction | Device Design | Consignees were sent on 5/2/2017 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 34082 dated May 2, 2017. The letter was addressed to Chief of Anesthesia, Director of Biomedical / Clinical Engineering&amp; Health Care Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. GE Healthcare will correct affected products. GE Healthcare representative will contact the firms to arrange for the correction. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 3448 | 3,448 (334 US; 3,114 OUS) | US Distribution including PR and to the states of : AL, AR, AZ, CA, CO, CT, FL GA, ID,IL, IN, KY, LA, , MA, MD, MI, MO, MS, MT, MN, NC, NH, NJ, NV, NY, OH, OK, PA, TN, TX, VA, WA, WI, and UT. OUS: AlgeriaArgentinaAustraliaAustriaBelgiumBoliviaBosnia and HerzegovinaBrazilBulgariaCambodiaCanadaChileColombiaCosta RicaCote d'IvoireCroatiaCzech RepublicEcuadorEgyptEl SalvadorEstoniaFinlandFranceGermanyGreeceHondurasHong KongHungaryIndiaIndonesiaIraqIrelandIsraelItalyJapanJordanKenyaKorea (Republic Of)KuwaitLatviaLebanonLithuaniaMalaysiaMaldivesMaltaMauritiusMexicoMoldovaMoroccoMyanmarNetherlandsNew ZealandNigeriaNorwayOmanPakistanParaguayPeruPhilippinesPolandPortugalRomaniaRussiaSaudi ArabiaSenegalSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanTanzaniaThailandTurkeyUnited Arab EmiratesUnited KingdomUruguayVietnamZimbabwe | Open, Classified |  |  | 3448 | Not_Computer
Z-2697-2017 | 77291 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Icon is a teletherapy device intended for stereotactic irradiation of head structures ranging from very small target sizes of a few millimeters to several centimeters e.g., metastatic tumors, recurrent glioblastomas, trigeminal neuralgia, medically refractory essential tremor, orbital tumors, ocular tumors, optic nerve tumors, benign diseases (such as meningiomas, vestibular schwannomas, post-surgical pituitary adenomas, craniopharyngioma, hemangioblastomas, schwannomas, arteriovenous malformations, cavernous malformations, chordomas, glomus tumors, hemangiomas), skull base tumors, head and neck tumors (such as unknown primary of the head and neck, oral cavity, hypopharynx, oropharynx, nasopharynx, sinonasal, salivary gland), and pediatric tumors (such as glioma, ependymoma, pituitary tumors, hemangioblastoma, craniopharyngioma, meningioma, metastasis, medulloblastoma, nasopharyngeal |  |  | May 23, 2017 | 2017 | Elekta, Inc. | If the gantry module is replaced after original installation the configuration settings may be missing. The identified risk for this issue is electrical safety for technicians doing maintenance on the X-ray generator. | Other | Elekta sent an Urgent Important Field Safety Notice dated June 2017, to all affected customers. This notice informs users of the specific product and version numbers affected by the issue. The affected devices are spare part Gantry Module (item number 1515698). Customers are also instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta representative. For further questions, please call (770) 670-2422. |  | 2 | 2 | US Distribution to the states of : NY and PA. | Open, Classified |  |  | 2 | Not_Computer
Z-2701-2017 | 76427 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | The Olympic Pasteurmatic 3000 and Olympic Pasteurmatic 3500 are intended for cleaning and high-level disinfection of respiratory therapy and anesthesia equipment by killing vegetative bacteria and viruses, in a hot water bath at ~160Âż - ~170Âż F for 30 minutes. |  |  | May 22, 2017 | 2017 | Natus Medical Incorporated | Olympic Pasteurmatic 3000/3500 Systems are recalled because they do not fully meet todays regulatory requirements. Given the age of the products design, component and technology obsolescence, and the time involved in fully complying with the latest design and development standards, the firm decided to recall of the Olympic Pasteurmatic 3000/3500 System devices by means of disabling the devices. | Device Design | Natus Medical sent an URGENT: RECALL NOTIFICATION MEDICAL DEVICE RECALL letter dated May 2017.Please refer to the accompanying Instructions to Disable the Pasteurmatic (DOC-022838), then complete and return the attached reply form. Should you require assistance in decommissioning your affected system please indicate option 2 on the reply form and you will be contacted to schedule an in- service. Natus may reimburse facilities for Olympic Pasteurmatic" 3000/3500 System purchases based on the age of the affected system as determined by the system serial number.There is no evidence to suggest there is a safety risk associated with the Olympic Pasteurmatic" 3000/3500 System device, nor have there been any safety incidences reported to Natus.Natus apologizes for any inconvenience. Please understand that this is part of our commitment to providing only the highest quality products and information to our customers. Please contact us at Natus_Quality_Programs@natus.com or (866) 940-7143 option #2 if you have any further questions. |  | 275 | total 275 units (97 units in US and 178 units international) | Worldwide Distribution - USA (nationwide) Distribution to the states of : AL, AR, CA, CO, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, and WV and in the following countries: Canada, Colombia, Costa Rica, Ecuador, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, Mexico, Panama, Peru, Qatar, Saudi Arabia, Taiwan, Thailand, and Viet Nam. | Open, Classified |  |  | 275 | Not_Computer
Z-2661-2017 | 77467 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Medacta GMK HINGE, Tibial Insert Fixed, SIZE 2, HEIGHT 23 mm, REF 02.09.0223H, STERILE EO |  |  | May 19, 2017 | 2017 | Medacta Usa Inc | One lot of product does not include a screw for attachment to the tibial baseplate. | Nonconforming Material/Component | On 5/19/17 all distributors with recalled product were immediately notified of the recall by email. On 5/23/17, all distributors were further notified of the recall via phone. On 5/30/17, a third notification was completed via phone. |  | 4 | 4 units | CA, AZ, NC, Israel, France | Completed |  |  | 4 | Not_Computer
Z-2662-2017 | 77001 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Logic Fit Tibial Tamp Head |  |  | May 25, 2017 | 2017 | Exactech, Inc. | The firm determined that the labeling (surgical technique) should be updated to clarify the proper technique and instrumentation to remove the tamp assembly from the bone, i.e., to include a caution statement about the potential for instrument breakage if the tamp handle/guide is misused by impacting in retro-grade. | Error in labeling | Consignees were sent a Field Advisory Notice dated 5/22/2017 by mail informing them updated surgical techniques are available on Exactech's website (www.exac.com). The notice requested consignees:Extend the information to accounts in possession of the productComplete &amp; return the attached Field Advisory Notice Response Form to Exactech |  | 573 | 573 | AL, AZ, CA, CA, CO, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, ME, MN, NC, NJ, NV, NY, OH, OK, OR, RI, SC, TN, TX, VA, WA, WI, Hawaii and Puerto RicoAustralia, France, Germany, India, Italy, Japan, Korea, Luxembourg, Singapore,Spain, Switzerland, The Netherlands and United Kingdom | Open, Classified |  |  | 573 | Not_Computer
Z-2671-2017 | 77077 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC Unit, Model 8015 |  |  | June 12, 2017 | 2017 | CareFusion 303, Inc. | BD initiated the recall of Alaris PC unit model 8015 after the firm identified five scenarios which can result in the occurrence of Systems Error Code 255-16-275 and can potentially result in interrupted infusions. | Software design | BD sent an Medical Device Safety Notification Letter dated June 12, 2017 , to customers to inform them that BD is issuing this Safety Notification letter to inform you of reports of customers experiencing a System Error (error code: 255-16-275) with Alaris System PC unit model 8015 that can result in interruption of infusions. Customers are informed of the issue, potential risk, and required actions to be taken by users and by BD. Customers with questions with the safety notification are instructed to call (888) 562-6018 or email supportcenter@carefusion.com. Customers with adverse event reports are instructed to call (888) 812-3266 or email customerfeedback@carefusion.com. Customers with technical questions are instructed to call (888) 812-3229 or email dl-us-inf-techsupport@carefusion.com. For further questions, please call (858) 617-5925. |  | 568283 | 568,283 units | Worldwide Distribution - USA (nationwide) and to the countries of : Europe, Australia, New Zealand,South Africa, Greater Asia, Middle East, and Canada. | Open, Classified |  |  | 568283 | Not_Computer
Z-2700-2017 | 77279 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module model 8100 |  |  | June 12, 2017 | 2017 | CareFusion 303, Inc. | There is a potential risk that could cause unintended flow in the older, centered sear door latch design in the Alaris Pump module model 8100. | Device Design | An Urgent Medical Device Recall letter will be sent on 6/12/17 to customers to inform them that BD is issuing this letter to inform you of a potential risk associated with the Alaris Pump module. Firm identified a specific scenario that could cause unintended flow in the older, centered sear door latch design in the Alaris Pump module model 8100. This scenario is reproducible under the following situations: a) the user does not close the roller clamp on the IV administration set before the pump door is opened as recommended, and b) opening the pump door using an atypical technique with the door latch with the centered sear design. The letter instructs customers If their device has a centered-sear that was manufactured between June 2002 to June 2004, clearly mark and sequester (e.g. Biomed department) the Alaris Pump module that exhibited the issue. Notify CareFusion Customer Advocacy at 888-812-3266 or customerfeedback@carefusion.com to coordinate an onsite remediation or send the devices to the BD Service Depot. |  | 35940 | 35,940 units | US and Canada | Open, Classified |  |  | 35940 | Not_Computer
Z-2704-2017 | 77514 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech Optetrak Logic CC 15mm Tibial Augment Screws |  |  | April 07, 2017 | 2017 | Exactech, Inc. | Missing Screw in carton. Nine (9) serial numbers were identified as being inadvertently packaged as one per carton when two per carton are required. | Process control | Consignees were initially contacted by phone on 04/07/2017. Consignees were sent a recall notice by mail, dated 04/13/2017. The notice requested that consignees immediately cease distribution/use of product, that the recall Extend the information to accounts in possession of the productIdentify/quarantine any subject devices in your inventoryComplete &amp; return the attached Recall Inventory Response Form to Exactech within 5 business days at 2411 NW 66th Ct, Gainesville, FL 32653 or faxed to 352-337-3915. Questions from customers are to be directed to Kaya Davis at kaya@exac.com or 800-392-2832. |  | 9 | 9 devices | CA and NY. | Open, Classified |  |  | 9 | Not_Computer
Z-2740-2017 | 77442 | No | LKK |  |  | General Hospital | PMA | PROMETRA II Programmable Infusion Pump, Implantable Infusion Pump, Catalog Number/REF 13827, Part Number PL-31790-02. Manufactured by: Flowonix Medical Inc., --- The correction involves the labeling of the device plus the following Prometra II pump labeling is also being revised: Intrathecal Catheter IFU, REF 11823, PL-22790-02; Catheter Revision Kit IFU, REF 11 830, PL-21798-02; Refill Kit IFU, REF 11825, PL-21794-00; Prometra II Patient Implant Card (Permanent) PL-32300-01; Prometra II Patient Implant Card (Temporary): PL-32375-01; Prometra II Patient Guide: PL-31912-01. |  |  | May 22, 2017 | 2017 | Flowonix Medical, Inc. | Flowonix Medical received a report of a patient implanted with the Prometra II Programmable Pump who may have received a fatal drug overdose during an MRI procedure. The Prometra II pump has an FDA-approved design feature intended to permit safe exposure to an MRI without removing drug from the reservoir. | Under Investigation by firm | Flowonix sent an URGENT MEDICAL DEVICE CORRECTION Letter and Response Forms dated May 22, 2017, with updated instructions were sent via UPS 2nd Day Air on May 25, 2017 to healthcare professionals: physicians and MRI facilities and technicians. The documents will also be available at www.flowonix.com. A Correction Letter will also be sent to patients to inform them of the labeling changes. For further questions, please call (973) 426-9229. |  | 2204 | US: 2204 units | Nationwide Distribution | Open, Classified |  |  | 2204 | Not_Computer
Z-2705-2017 | 77528 | timer, general laboratory | JBS | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | O-Two Adult Ventilation Timer |  |  | June 23, 2015 | 2015 | O-Two Medical Technologies, Inc. | Malfunction of an O-Two Adult Ventilation Timer, where the flash rate of the malfunctioning device was reported to be twice that of the rate noted on the label. | Mixed-up of materials/components | O-Two Medical Technologies sent a Recall Notification on June 23, 2015, to all affected customers. The letter notified customers of the malfunctioning device, description, brand name, model/catalog number, target population, quantity, Lot Number, Shelf Life, hazard potential, and an image of labeling. The letter also provided instructions to customer for action and response. Customers were to return the affected product and O-Two will ship back the unaffected medical devices. Customers were instructed to acknowledge receipt of this recall notice by fax to 1-905-677-2035. For questions regarding this recall call 905-677-9410 or 1-800-387-3405. |  | 4382 | 4382 units (4082units distributed USA) | Worldwide Distribution - US (nationwide) Canada, Europe Denmark and Finland | Completed |  |  | 4382 | Not_Computer
Z-2702-2017 | 77571 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Penumbra 3D Revascularization Device |  |  | June 09, 2017 | 2017 | Penumbra Inc. | Penumbra has identified an issue in these four lots involving a raw material component of the delivery wire. This issue could result in breakage of the delivery wire, which could potentially lead to serious patient injury or death. | Component change control | Penumbra sent an Urgent Voluntary Field Removal Notice dated June 9, 2017. Customers were instructed to inspect current inventory and remove any affected units for return. Customers were also instructed to return the response form. Penumbra has identified an issue in these four lots involving a raw material component of the delivery wire. This issue could result in breakage of the delivery wire, which could potentially lead to serious patient injury or death.Penumbra personnel will contact customers directly to arrange return of affected units and replacement of returned product at no charge to them.Customers with question should call 1-510-748-3288. |  | 155 | 155 units | Nationwide Distribution to the following states: AZ, CA, CO, FL, IN, MA, MI, MN, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WV | Open, Classified |  |  | 155 | Not_Computer
Z-2731-2017 | 77534 | No | NAE | Hydrogel wound dressing and burn dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Burn ReliefÂż |  |  | September 30, 2015 | 2015 | O-Two Medical Technologies, Inc. | For lack of evidence to validate effectiveness of sterilization dose, O-Two Medical Technologies Inc. hereby announces a For lack of evidence to validate effectiveness of sterilization dose, O-Two Medical Technologies Inc. hereby announces a voluntary medical device recall of the suspected batches of Burn ReliefÂż products to prevent potential contamination. . | Process control | O-Two sent a Voluntary Recall Notification dated September 30, 2015, to all affected customers. Customers were instructed to acknowledge receipt of the recall notice by fax to 1-905-677-2035; or email to cs@otwo.com. Customers with questions should call 1-905-677-9410.For questions regarding this recall call 905-677-9410 . |  | 624 | 624 units | Worldwide Distribution - US (nationwide) Canada, Belgium, Norway, and Malaysia | Completed |  |  | 624 | Not_Computer
Z-2736-2017 | 77566 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CS 100i Intra-Aortic Balloon Pump |  |  | June 16, 2017 | 2017 | Maquet Datascope Corp - Cardiac Assist Division | The device failed to pump due to an electrical test failure code #58 (power up vent test fail), maintenance code #3, and an autofill failure which has been associated to a patient death due to the failure of the device to initiate therapy. | Component design/selection | Datascope Corp/MAQUET issued a press release on June 16, 2017. Datascope Corp/MAQUET sent an Urgent Product Recall Letter on June 19, 2017, to their customers via FedEx next day delivery service. Pursuant to the WARNINGS section of our CS100i, CS100 or CS300 IABP Operating/User Instructions, clinicians are instructed not to leave the patient unattended during IABP therapy. Until the service is performed, we recommend the IABP be powered on prior to initiation of therapy to allow the IABP to successfully complete its self-test. This action will take less than 60 seconds to perform. In the event the IABP fails to successfully complete the self-test and exhibits electrical test failure code 58, please remove theIABP from service and contact your local Maquet/Getinge Sales &amp; Service Office. Customers will be instructed to complete and return the attached Medical Device Field Correction Response Form via fax to1-973-807-9217 or send via email to IABP2017@getinge.com.ABP2017@getinge.com.If you are a distributor who has shipped any affected products to customers, please forward this document to their attention for appropriate action.For U.S. customers with technical questions, please contact ourTechnical Support Department (at 1-888-627-8383 and press 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST |  | 12319 | 12,319 units total | Worldwide Distribution - (in over 100 countries) US (Nationwide) | Open, Classified |  |  | 12319 | Not_Computer
Z-2739-2017 | 77361 | No | FKH | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Conductivity Standard Solution, 14.0 mS/cm @ 25 degrees Celsius, specification +/- 0.02 mS, Discard 30 days after opening, Sodium Chloride (CAS# 7647-14-5) and Deionized Water (CAS# 7732-18-5), 32 oz (950 mL) |  |  | May 16, 2017 | 2017 | Mesa Laboratories, Inc. | When a meter is calibrated with a solution that is out of tolerance, the dialysis machine can read "functioning properly" when is not. | Nonconforming Material/Component | MesaLabs sent an Device Correction Notice dated May 15, 2017, to all affected customers. Actions Required:Mesa Laboratories, Inc. will replace all solution with this Lot Number. The defective product should be discarded and not returned to Mesa Laboratories, Inc.We recommend that you obtain replacement solution for any bottles that you have on hand of the affected lot numbers. To obtain replacement solution, please fill out the attached form and FAX or email it to Mesas Customer Service department. For further questions, please call (303) 987-8000 ext 120. |  | 9900 | 9900 bottles | USA (nationwide) Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VI, WA, WI, WV, WY, Hawaii and Puerto RicoCanada. | Open, Classified |  |  | 9900 | Not_Computer
Z-2718-2017 | 77560 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 28 cm (11") PUR Bifuse Ext Set w/6-Port NanoClaveÂż Manifold, Check Valve, 2 NanoClaveÂż, Rotating Luer. |  |  | May 30, 2017 | 2017 | ICU Medical, Inc. | The connection between the Female Luer and the NanoClaveÂż manifold has the potential for developing a leak. | Nonconforming Material/Component | Customers were notified via letter on approximately 05/30/2017 and were instructed to examine their inventory, quarantine affected products, complete and return the response form, and return to ICU Medical. If there are any questions or assistance needed, customers can call ICU Customer Service Monday through Friday between 8:30 AM and 4:00 PM Pacific time: 866-829-9025. Or they can email: Productreturns@icumed.com. |  | 50 | 50 units | Distribution to Italy and France, and the following states in the US: AZ, OH, TX, WA. | Open, Classified |  |  | 2750 | Not_Computer
Z-2714-2017 | 77551 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | ArjoHuntleigh Flowtron ACS900, medical pump, software V1.099 |  |  | June 14, 2017 | 2017 | Arjo, Inc. dba ArjoHuntleigh | Reports of the device failure involving delivering continuous pressure to the calf garment and not deflating as intended, without triggering an alarm. | Under Investigation by firm | ArjoHuntleigh sent a Field Safety Notice and Customer Response form to affected customers describing the reason for the recall and the actions to be taken. The instruction to customers is to review the Field Safety Notice and complete the Customer Response Form to acknowledge receipt and understanding of the Field Safety Notice and the required actions. These actions taken are as follows: 1. Ensure that all caregivers and users of the ArjoHuntleigh Flowtron ACS900 pump are made aware of this Field Safety Notice (FSN) and all listed devices at facility are available to be upgraded to software version V2.000 during the service technician visit. 2. Complete and sign the enclosed Customer Response Form and return this form to the local ArjoHuntleigh office. Note: ArjoHuntleigh will contact the designated facility representative listed on the Customer Response Form to schedule service to upgrade the software on all listed devices at your facility, free of charge.The device may stay in use until the upgrade will be performed, on the condition that the patients limbs are monitored frequently and the garments are checked to ensure that they are correctly fitted to the patient and that deflation of the garments is occurring on a regular basis. It is also recommended that the pump LCD display is regularly checked to ensure that it shows that the garments are inflating and deflating correctly (in the limited number of complaints, the LCD display was blank). This Field Safety Notice must be distributed to those individuals who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. Additionally they should maintain awareness of the notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. The letter states that if the customer has additional questions or they require assistance completing the Customer Response Form, they should contact ArjoHu |  | 40675 | 40,675 units worldwide | Worldwide Distribution - US (Nationwide) and internationally to the following countries: United Kingdom, Australia, New Zealand, Japan, Italy, United Arab Emirates, France, Canada, Brazil, Austria, Saudi Arabia, Qatar, India, Switzerland, Bahrain, Kuwait, South Africa, Lebanon, Netherlands, Poland, Spain, Ecuador, Hong-Kong, Colombia, Malaysia, Paraguay, Mauritius, Cyprus, Greece, Oman, Norway, Ireland, Pakistan, Czech Republic, Mexico, Sweden, Tanzania, Hungary, Singapore, Jordan, Israel, Denmark, Chile, and Germany. | Open, Classified |  |  | 40675 | Computer
Z-2746-2017 | 77354 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Install CD, Spine tools, Plus and S7, Version 25. Model Number 9731958. |  |  | May 10, 2017 | 2017 | Medtronic Navigation, Inc. | Software issue related to the StealthStation S7 system and the Synergy Spine application Version 2.1 configured with Spine Tool Install CD version 25. Issue may result in user being unable to navigate the Navigated Elevate Inserter with the StealthAiR Spine Frame during spine surgical procedures. | Software in the Use Environment | Medtronic sent an Urgent - Product Advisory Notice dated May 10, 2017, to all affected customers. Instructions indicated that a Medtronic representative will be in contact to update the system with version 26 and will remove any copies of the Spine Tools Install CD, version 25. Please quarantine any copies of version 25 until they can be removed.If you have further questions regarding this communication, please contact your local Medtronic Navigation Technical Services Representative, your local Medtronic Spine Sales Representative, or call (800 595-9709. |  | 108 | 108 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Brazil, Germany, Switzerland and Great Britain. | Open, Classified |  |  | 108 | Computer
Z-2770-2017 | 77373 | No | PTJ | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative test kits containing colistin (cs01n) |  |  | May 16, 2017 | 2017 | Biomerieux Inc | The investigation demonstrated a high rate of very major errors (resistant isolates calling susceptible) with the VITEKÂż 2 AST-GN colistin (cs01n) compared to agar dilution (the reference method used for cs01n development) and compared to broth-microdilution (recommended method by EUCAST/CLSI). | Other | Customers were notified on or about 05/16/2017 and then notified again with an updated letter. Instructions to customers included requesting that the attached acknowledgement form be completed, distribute the recall notification to customers, and provide the notice to future customers. |  | 1520000 | 1,520,000 cards | Domestic: NC. Foreign: Argentina, Austria, Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, France, Germany, Guatemala, Hong Kong, India, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, Netherlands Antilles, Nicaragua, Panama, Paraguay, Peru, Philippines, Poland, Singapore, South Africa, South Korea, Spain, Thailand, Turkey, United Kingdom, Uruguay, Venezuela, &amp; Vietnam. | Open, Classified |  |  | 1520000 | Not_Computer
Z-2699-2017 | 77604 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arrow Glide Thru Peel-Away Sheath/Dilator Introducer |  |  | June 12, 2017 | 2017 | Arrow International Inc | Arrow is recalling additional lots that were identified as part of an active recall. Arrow is recalling the affected product due to the possibility that the catheter peel-away component hub tabs may prematurely detach when the practitioner begins to peel apart the sheath body from the catheter. | Device Design | Teleflex/Arrow International mailed an Urgent Medical Device Recall Notification Letter to affected customers on 06/12/2017 to inform them of the issue. Arrow requested that customers examine their inventory immediately for the affected lots and discontinue use and quarantine any products with the associated product codes identified in the notice and complete the Recall Acknowledgement Form and fax back to the number included in the notice. |  | 13542 | 9,037 units in the U.S. and 4,505 Internationally | Distributed to SC, AL, NJ, IN, MA, GA, CA, PA, AZ, VA , WA and Bangkok | Open, Classified |  |  | 13542 | Not_Computer
Z-2339-2017 | 77292 | No | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VIDAS Testosterone, Ref 30418 |  |  | May 15, 2017 | 2017 | bioMerieux, Inc. | Complaints were received from customers observing falsely overestimate results or external quality control higher results than expected when performing tests with VIDAS Testosterone. | Under Investigation by firm | An Urgent Product Safety Notice was sent on 5/15/2017 to consignees. The Notice requested that they discontinue use and discard any remaining inventory of the product, and to identify any possible falsely over estimated results and determine appropriate actions if relevant. The Notice also included a response form to be returned to Biomerieux. |  | 742 | 742 | AR, CA, CO, GA, IL, IA, MI, MN, MT, NY, OK, TN, TX, WY, Columbia, Costa Rica, El Salvador, France, Hong Kong, Jordan, Nicaragua, Russian Fed., Sri Lanka, Sweden, Ukraine | Open, Classified |  |  | 742 | Not_Computer
Z-2670-2017 | 77402 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Jelco, Hypodermic needle ProÂż needles 19G X1 1/2" 800/CA Model No. EL 1915 &amp; 19G X 2" 600/CA Model No. SB 5021, Rx Only, Sterile EO. |  |  | April 04, 2017 | 2017 | Smiths Medical ASD Inc. | Smiths Medical became aware on January 11, 2017, that syringe assemblies provided by Nipro Medical Corporation contained pin holes in the side wall of the hub. The defect was not detected during Smiths Medical incoming inspection. The syringe assemblies were used to manufacture seven (7) finished good lots of Hypodermic Needle ProÂż needles, Smiths Medical part numbers EL 1915 and SB5021. | Process control | Smiths Medical sent an "Urgent Medical Device Recall Notice" dated April 04, 2017. The letter was addressed to Quality Assurance Manager &amp; Associate Sourcing Consultant. The letter listed the Affected devices described the Reason for Recall, &amp; Risk to Health. Requested consignee to complete Urgent Medical Device Recall Response Form within 10 days ofreceipt and send it to: fieldactions@smithsmedical.com.For further questions, please call (603) 352-3812 |  | 736000 | 736,000 | US Distribution to the state of : IN | Open, Classified |  |  | 736000 | Not_Computer
Z-2660-2017 | 77603 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Nussloch GmbH ASP6025 Tissue Processor |  |  | June 19, 2017 | 2017 | Leica Microsystems, Inc. | Incorrect labeling for specified voltage for the Alarm Connectors on the rear side of the unit. | Labeling mix-ups | Customers were notified via email on 06/19/2017. Instructions to customers included to clean surface around the Alarm Connectors and place the newly provided label at the Alarm Connectors (over the engraved specification). They were also instructed to update their current Instructions For Use and update the online help function in the software, as well as return the response form. |  | 812 | 812 units | Domestic: AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WA, WI, WV, &amp; DC. Foreign: Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, England, France, Georgia, Germany, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lithuania, Malaysia, Maldives, Mexico, Morocco, Myanmar, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, &amp; United Arab Emirates. | Open, Classified |  |  | 812 | Computer
Z-2672-2017 | 77620 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 and Proteus ONE proton therapy systems |  |  | May 18, 2016 | 2016 | Ion Beam Applications S.A. | Recall is being initiated to address on-site software installations not being performed consistently with IBA standard operating procedures. | Employee error | An Urgent Field Safety Notice MID-54097 was mailed to affected customers to inform that IBA plans to deploy a new software by March 2017 to correct the issue. Customers were reminded to follow standard protocol until the software updates are deployed. |  | 28 | 18 worldwide and 10, in the U.S. | Products distributed to FL, VA, PA, OK, IL, NJ, WA, TN, LA, TX, South Korea, Germany, France, Czech Republic, Italy, Poland, and Sweden | Open, Classified |  |  | 28 | Computer
Z-2669-2017 | 77190 | [6 empty] | Electrophysiology catheter cables are designed |  |  | May 08, 2017 | 2017 | Sterilmed Inc | Sterilmed is recalling all lots of specific product codes of unexpired Sterilmed Reprocessed EP Ablation Cables because Sterilmed determined that the cables are Class III devices and thus require, but do not currently have, premarket review and approval. | No Marketing Application | Sterilmed sent an Urgent Notice Medical Device Recall letter dated April 25, 2017, to all affected customers. The letter identified affected product, stated reason for recall. Affected devices were asked to be quarantined, until returned to Stericycle, by August 31, 2017 to qualify for credit. A response form was asked to be returned. Questions can be directed to Sterilmed's Customer Care Team at 1-888-541-0078. For further questions, please call ( 877 ) 787-0369. |  | 963 | 963 devices | US Distribution to the states of : AL, AR, CO, FL, IL, MI, MO, MS, MT, NC, OH, PA, TN, TX, VA WI and WV. | Open, Classified |  |  | 963 | Not_Computer
Z-2664-2017 | 77004 | No | HTX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | EXACTECH Equinoxe Glenoid Reamer - Catalog #: 321-05-38, Catalog #: 321-05-42, and Catalog #: 321-05-46 |  |  | May 25, 2017 | 2017 | Exactech, Inc. | Instructions for use (IFU) is updated to clarify the proper technique and instrumentation to ream the Glenoid, i.e., to include a caution statement that the surgeon should ream on-axis and that reaming off-axis may result in instrument breakage. | Error in labeling | On May 22, 2017 Exactech sent a Notification to all their Direct Accounts an Distributors telling them that a new updated technical techniques LIT # 718-01-30 Rev J is now available in their website (www.exac.com) or can be requested by contacting the Customer Service Representative at 800-392-2832, fax 352-378-2617 or via e-mail at Kaya@exac.com. This updated the technical techniques for implantation for Equinoxe Platform Shoulder System to clarify the proper technique for use of the Pilot Tip Reamers, i.e., to include caution statement that the surgeon should avoid applying a bending force to the pilot tip reamer and should avoid using the reamer to retract the humeral head, as this may result in instrument breakage. The notification also request to forward and informed this notice to all the accounts that have received this product and to complete and return the attached Field Advisory Notice Respond Form within 5 days of receipt. For more information please contact: Kaya Davis, Inventory Representative at kaya@exac.com or by phone at (352) 377-1140 |  | 813 | 813 | AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, MA, MD, ME, MN, NC, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, Hawaii and Puerto RicoAustralia, Austria, Brazil, Canada, China, France, Germany, Greece, India, Italy, Korea, Luxembourg, Singapore, South Africa, Spain, Sweden, Switzerland, The Netherlands and United Kingdom | Open, Classified |  |  | 4390 | Not_Computer
Z-2755-2017 | 77623 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Visualase(R) Cooled Laser Applicator System (VCLAS) 15MM TIP, REF 9735561, (1) Laser Diffusing Fiber - 600um Core, 15mm Tip, (1) Cooling Catheter with 3M(TM) SteriStrip(TM), (1) Pump Tubing Set, (1) Drainage Bag. Medtronic Navigation, Inc. Louisville, CO 80027 |  |  | June 07, 2017 | 2017 | Medtronic Navigation, Inc. | Medtronic has become aware that the VCLAS 15MM TIP cap, also referred to as Tuohy Borst Adapter (TBA), may not be capable of creating a seal around the optical fiber to ensure adequate saline flow through the cooling system. When the TBA is unable to seal around the optical fiber, low retention force to the optical fiber may occur, causing saline to leak out of the TBA. | Component design/selection | Medtronic sent an Urgent Medical Device Safety Alert dated June 1, 2017, to all affected customers.Actions to be taken:1. Please examine your inventory and if any of the affected products listed above are found, immediately quarantine them for return to Medtronic.2. Complete and follow instruction on the attached consignee response form. Sign and date the bottom of the form and then return the form to Medtronic at RS.NavFCAÂżmedtronic.com or fax to Medtronic Technical Services at 651-367-7075.If you have affected products, contact Medtronic Technical Services at 1-800 595-9709 to receive a return material authorization (RMA) and arrange for their return and no charge replacements. Once an RMA number is obtained, ship the affected product to Medtronic. For further questions please call (720) 890-3241. |  | 214 | 214 units | US Distribution to the states of : CA, GA, MN, FL, NY, IL, TX | Open, Classified |  |  | 214 | Not_Computer
Z-2684-2017 | 77621 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235, Proton Therapy System |  |  | December 13, 2016 | 2016 | Ion Beam Applications S.A. | Software issue | Software design | An Urgent Field Safety Notice was mailed to affected customers to inform them of the issue. The notice indicated possible software failure, the associated risks and actions users should implement to correct the issue. IBA plans to follow-up with a software update which will be deployed on impacted sites to correct this issue. |  | 11 | 6 worldwide and 5 in the U.S. | Distributed to FL, VA, IL, NJ, WA and South Korea | Open, Classified |  |  | 11 | Computer
Z-2688-2017 | 77324 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Sensis Vibe Systems with Software Version VD10B, Model Numbers 10765502, 10910620, 11007641, 6648161 --- Programmable diagnostic computer, |  |  | May 16, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Software error. In Sensis Vibe systems with software version VD10B, a software error can result in: problems generating a report and/ or- information from different examinations of the same patient being combined into one report. --- The error causes information from two examinations to be combined into one report. This does not result in a hazardous situation; however, if the erroneous report has been used for treatment or diagnosis, this can potentially result in an incorrect treatment or diagnosis. | Software design | Siemens sent an Customer Safety Advisory Notice dated May 16, 2017, was mailed to all users on May 16, 2017. The letter informs customers of the software issue and potential impact. Siemens Medical Solutions will contact users to make arrangements to update the system with updated software to correct the issue. For further questions, please call (610) 448-6461. |  | 13 | US: 13 units | US Distribution to the states of : IN, MI, IA | Open, Classified |  |  | 13 | Computer
Z-2687-2017 | 77482 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Tenaxis Medical ArterX(TM) Surgical Sealant, (a) 4mL, REF 9002, and (b) 5mL REF 9006, STERILE R, Store at 2oC - 8oC |  |  | October 17, 2014 | 2014 | Mallinckrodt Manufacturing LLC | Product was improperly labeled leading to improper storage. | Other | The direct consignee was notified by email on October 17, 2014.The action requested to be taken on receipt of the notice:" Remove all ArterX delivered by CJ Medical from inventory and quarantine for return to CJ Medical." Return the product to CJ Medical for a credit or product replacement" If no ArterX is found in inventory, please comment to that effect on the attached form and return the form to CJ Medical.The consignee was also requested to circulate the Field Safety Notice among all users affected by this product. For further questions, please call (314) 654-2000. |  | 246 | 246 units | Worldwide Distribution to the country : UK | Terminated | July 15, 2017 | 1002 | 246 | Not_Computer
Z-2690-2017 | 77362 | No | PER | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2, Part Number 378600, |  |  | May 22, 2017 | 2017 | Beckman Coulter Inc. | When samples are prepared using the PrepPlus or PrepPlus 2 sample preparation systems and then run on a Flow cytometer, there is an imprecision in the absolute counts on only the first tube in a run. | Software design | The recalling firm initiated their recall by letter disseminated via email on 05/22/2017. The letter requested the following actions be taken:1. Insert an empty daughter tube in place of a patient sample as the first tube in the carousel that is programmed for the addition of Flow-Count Fluorospheres.2. Refer to the COULTER PrepPlus Operators Guide / PrepPlus 2 Instructions for Use (IFU) for instructions on adding Flow-Count Fluorospheres to daughter tubes.3. Review all results following the flow cytometer data review instructionsreferenced in the associated IFU.Beckman Coulter also requested the consignee share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If the consignee forwarded any of the affected product(s) listed above to another laboratory, They were asked to provide them a copy of the letter. |  | 473 | 473 units | Worldwide Distribution | Open, Classified |  |  | 516 | Not_Computer
Z-2696-2017 | 77509 | No | MHY |  |  | Neurology | PMA | Infinity 7, Implantable Pulse Generator, REF 6663, Rx only, STERILE EO |  |  | June 02, 2017 | 2017 | St. Jude Medical, Inc. | The firm discovered a risk of loss of Neuromodulation therapy due to exposure of the IPG to monopolar electrosurgical devices (ESU) during medical procedures. | Device Design | The notification material is in the form of a Physician Letter and Patient letter which will be delivered to physicians and patients in the US. The physician mailing will be initiated June 2, 2017 for delivery on June 7-9, 2017. The physicians were provided with patient management recommendations, Step for the patient controller updates, and Clinician's Manual Revised Warning. The Patient mailings will begin the week of June 12, 2017 to allow for clinician notification and are expected to be delivered within 3-5 business days. The patient letter informs of them of the potential for the loss of therapy and the steps to receive the software upgrade. |  | 40 | 40 units | US Nationwide Distribution | Open, Classified |  |  | 19219 | Computer
Z-2757-2017 | 77686 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Converters(R) Under Buttocks Drape with Fluid Control Pouch II, REF 8482, STERILE R, Qty. 1, Cardinal Health, Waukegan, IL |  |  | June 27, 2017 | 2017 | Cardinal Health 200, LLC | Outer label of some units of product number/lot number 8482, Lot 16LFS235 and product number/lot number 9414, Lot 16LUS317 do not match the drape packaged inside. | Mixed-up of materials/components | Cardinal Health sent an Urgent Product Recall letter dated June 27, 2017, to all affected consignees. In the letter, the firm requested the following actions be taken:1. Inspect your inventory for the affected product code and lot number (listed on the package)2. segregate and quarantine all on-hand product3. Return the enclosed acknowledgment form via fax to 847-689-9101 ir614-652-9648, whether or not you have affected product, indicating the product code, lot and quantity of product you've quarantined or discarded.4. Notify any customers to whom may have distributed or forwarded product affected by this product.5. Contact the appropriate Customer Service group to arrange for return and credit of any affected product:Hospital - 800-964-5227Federal Government - 800-444-1166Distributor - 600-635-6021All other customers - 888-444-5401For questions regarding this recall call 800-327-7503. |  | 1591 | 1591 each | Worldwide Distribution - US (Nationwide including PR) and Canada | Open, Classified |  |  | 2119 | Not_Computer
Z-2675-2017 | 64542 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ Infusion Systems are a multipurpose family of devices that provides precise delivery of multiple therapies across the general spectrum of clinical care. The Plum A+ family infusion pumps are self contained, microprocessor based infusion devices. The pumps are available in single channel and in triple channel pumping mechanism types. |  |  | February 01, 2013 | 2013 | ICU | A possible interruption of therapy when an E321 error code occurs while using the infuser, which is caused by battery issues. | Other | Hospira has contracted with Stericycle Inc., in Indianapolis, IN to send letters and a reply form to Hospira customers. Receipt of the letters will be confirmed by UPS tracking. Customers will be requested to return a reply form to Stericycle indicating they have received the letter, and notified their customers. The customers will also be asked to supply the number of Plum A+ infusers at their facility. |  | 180 | 180 units | Nationwide, Worldwide | Open, Classified |  |  | 596267 | Not_Computer
Z-2761-2017 | 77342 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Atrium iCAST Balloon Expandable Covered Stent; 7mm x 22mm x 120cm; Model/Part Number 85455 (Marketed in U.S. only) |  |  | May 25, 2017 | 2017 | Atrium Medical Corporation | Firm received a complaint that a customer who ordered an ADVANTA V12 7mm x 22mm x 120cm received an ADVANTA v12 7 mm x 38 mm x 80 cm, due to the manifold (catheter hub) was printed as 7mm x 38 mm x 80 cm. The complaint was the result of incorrect dimensional information (7mm x 38mm x 80cm) printed on the manifiold (catheter hub). The complaint device physical dimensions matched the dimensional information displayed on the outer packaging and the device inner pouch. This was the only complaint received for this non-conformance. | Under Investigation by firm | The firm, Getinge/Maquet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated May 25, 2017 via FedEx overnight to affected US consignees via FedEx Next Day Air. Getinge sent a second "URGENT MEDICAL DEVICE RECALL LETTER" dated June 2, 2017 via FedEx overnight to the affected customers in order to correct a product code/part number due to a typo in the initial letter and to assist in identifying the labeling problem. The letters described the product, problem and actions to be taken. Customers are instructed to examine your inventory; remove affected products; quarantine and place in a secure location; if you have the affected product, contact Getinge Customer Service at 1-603-880-1433, Monday through Friday between 9:00 am to 5:00 pm for RA and shipping instructions to return affected product to the Manufacturer, and complete and Return the enclosed Medical Device - Recall Response Form to email: catheterhub.us@getinge.com or fax to 1-973-807-9207.If you have any questions, please contact your Getinge/Maquet representative or call Maquet Customer Service at 1-603-880-1433, Monday through Friday between the hours of 9:00 am to 5:00 pm Eastern Time Zone. |  | 246 | 246 (US: 183 units; Foreign: 63 units) | Worldwide distribution-US (nationwide) and countries of: Germany, Great Britain, Austria, Ireland, Netherlands, France, Italy, Sweden, and San Marino. | Open, Classified |  |  | 246 | Not_Computer
Z-2754-2017 | 77627 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Merge Eye Care PACS Viewer 5.2 |  |  | June 13, 2017 | 2017 | Merge Healthcare, Inc. | When the user has not set up any user preference on the sorting order to render the study images (OS/OD/etc.) and reports, the ECP may get an incorrect study index. If the user attempts to delete an image, ECP may then read that image as OS when it is, in fact, OD. | Software design | The firm initiated the recall by certified letter the week of 06/26/2017. The notice provided the work-around and requested a response. |  | 18 | 18 units | FL, IL, CO, MI, CA, NY, ND, OH, CT | Open, Classified |  |  | 18 | Not_Computer
Z-2706-2017 | 76899 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultrasonic Endoscope |  |  | February 17, 2017 | 2017 | Fujifilm Medical Systems U.S.A., Inc. | Fujifilm is recalling multiple endoscopes after a retrospective review. | Other | Fujifilm sent an Urgent - Medical Device Removal letter dated February 17, 2017. Customers are provided information regarding the recall and asked to complete the attached form and return to Fujifilm. Questions can be directed to sales representative. For further questions, please call (973) 709-2202. |  | 4 | 4 units | US Distribution to : AL, HI and MO | Open, Classified |  |  | 4 | Not_Computer
Z-2710-2017 | 77599 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | APLIO 300 TUS-A300 |  |  | May 01, 2017 | 2017 | Toshiba American Medical Systems Inc | When Continuous Trace method is used as the tracing method on Spectral Doppler in velocity trace measurement, an incorrect value may be displayed in the ultrasound systems and in the ultrasound workstations listed below. | Software Manufacturing/Software Deployment | Toshiba Medical sent an Urgent Medical Device Correction letter on May 1, 2017. The consignee is instructed not to use the Continuous Trace method for cardiac velocity trace measurements to quantify the MPG values.Modified software is being provided. A Toshiba service representative will contact the consignee to make an appointment to install the modified software to fix the problem.Customers were asked to complete the form and fax it to the toll-free number shown at the top of the form. The form can also be sent by email to RAffairs@tams.com.Customers with questions were instructed to call 800-521-1968. |  | 90 | 90 units | Nationwide Distribution | Open, Classified |  |  | 433 | Computer
Z-2707-2017 | 77086 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | The Schiller PB 1000, also known as Merge Hemo Patient Data Module (PDM), a major component of the Merge Hemo software. |  |  | March 10, 2016 | 2016 | Merge Healthcare, Inc. | If the PDM is utilized to output a pressure for computing fractional flow reserve (FFR), the analog output may not reflect the true representation of the selected pressure waveform being provided by the PDM to the Link Assembly which provides the analog output. | Software design | The recalling firm issued recall letters dated 3/10/2016, 5/3/2016, and 8/5/2016 via email or via certified mail if the email address is not available. |  | 555 | 555 PDM's | Distribution US nationwide. | Open, Classified |  |  | 555 | Not_Computer
Z-2713-2017 | 77308 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Video Duodenoscope |  |  | January 17, 2017 | 2017 | Pentax of America Inc | Pentax initiated a field correction/safety alert for two (2) models of the Video Duodenoscope to determine how soiling may have occurred on the surface of the suction cylinder and under the distal cap during testing. | Device Design | The firm, Pentax Medical, sent a "FIELD CORRECTION' letter dated January 17, 2017 to their affected customers via USPS Certified Mail. The letter described the product, problem and actions to be taken. The customers were instructed to immediately remove any affected product from use; follow product labeling; ensure all reprocessing personnel are knowledgeable and thoroughly trained on the instructions of Use for manual reprocessing of the devices; clean elevator recesses and follow all reprocessing instructions; and complete and return the FIELD CORRECTION RESPONSE FORM via a member of the PENTAX Service Department as part of an on-site-visit or Fax to: QA/RA Department at 201-799-4063 (alternate 201-391-4189) or a pdf copy to customeradvisiories@pentaxmedical.com. The firm will replace any affected product with its current generation model.If you have questions or request further information or assistance contact PENTAX Medical Customer Support at 1-800-431-5880 (8:30AM - 5:00PM, Monday-Friday, EST) or email: customeradvisories@pentaxmedical.com. |  | 2015 | 2,015 (US - 519 (ED-3490TK) and 5 (ED-3270K); OUS - 1,491) | Worldwide Distribution-US (Nationwide) including states of: AL, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, PA, PR, TN, TX, UT, VA, WA, WI, WV, WY;. and country of: CANADA. | Open, Classified |  |  | 2015 | Not_Computer
Z-2715-2017 | 77565 | Yes | LWQ |  |  | Cardiovascular | PMA | On-X Prosthetic Heart Valve |  |  | June 12, 2015 | 2015 | CryoLife, Inc. | Incorrect product label. The type of Sewing Ring was mislabeled as being a Standard Sewing Ring instead of an Anatomic Sewing Ring. | Labeling mix-ups | On-X LTI Management Team had assessed this event to not be reportable. Correction were done by the On-X LTI in 2015. |  | 52 | 52 | Worldwide Distribution - US Nationwide in the states of AL, CA, FL, IA, IN, LA, MA, MN, MO, MS, NC, OR, PA, and TN.and the countries of Finland, United Kingdom and Israel | Open, Classified |  |  | 52 | Not_Computer
Z-2732-2017 | 77532 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova Biomedical Stat Profile Prime Analyzer as follows: |  |  | June 19, 2017 | 2017 | Nova Biomedical Corporation | Software- scheduled A-QC analysis could initiate priorto a patient result being released by the operator and result in the pending patient result being overwrittenwith the A-QC test results. | Software design | Nova Biomedical issued Urgent Field Correction Notice on June 19, 2017, to all affected customers. The letter advises users of the problem and Nova Biomedical or an authorized Nova Biomedical representative will schedule a no charge update of your Stat Profile Prime Analyzer software to version VI1.0.883.0. A representative will contact customers to schedule this update. Questions: contact Nova Biomedical Technical Support at 1-800-545-6682, in Canada contact us at 1-800-263-5999, or outside of N. America contact your authorized Nova Biomedical subsidiary or distributor.For questions regarding this recall call 781-647-3700. |  | 179 | 179 units (US) and 904 (OUS) Total 1083 | Worldwide Distribution - US (Nationwide)Foreign:AlbaniaAlgeriaArgentinaAustraliaAustriaBangladeshBeninBoliviaBrazilCanadaChileChinaColombiaCosta RicaCzech RepublicECUADOREgyptEl SalvadorFinlandFranceGermanyGhanaGreeceGuatemalaGuyanaHondurasHong KongHungaryICELANDIndiaIranIsraelItalyIvory CoastJapanJordanKenyaKoreaLatviaMacedoniaMalaysiaMexicoNepalNetherlandsNigeriaPakistanPakistanPalestinePeruPhilippinesPolandPuerto RicoRepublic of PanamaRepublica DominicanaRussiaSaudi ArabiaSlovak RepublicSloveniaSouth AfricaSri LankaSwitzerlandTaiwanThailandThe NetherlandsTrinidad,W.I.TunisiaTurkeyUKUruguayVietnam | Open, Classified |  |  | 179 | Computer
Z-2725-2017 | 77481 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Medica ISE Module Calibrant A; 500 mL - 140 mmol/L Na+, 125 mmol/L Cl-, buffers and preservatives; For use with ISE Module; |  |  | February 29, 2016 | 2016 | AMS Diagnostics, LLC | Lot was not functioning properly and not exhibiting the correct physical properties. | Labeling mix-ups | AMS Diagnostics contacted consignees by telephone on 02/29/2017. Representatives advised consignees to discard any recalled product on hand, to be replaced with working product. Customers should not return recalled product to the firm. |  | 9 | 9 units | US Nationwide Distribution in the states of AZ, CA, FL, MI, NC, NY, and SC. | Open, Classified |  |  | 9 | Not_Computer
Z-2730-2017 | 77364 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter Synchron AQUA CAL Level 1, Part No. 471288 |  |  | May 15, 2017 | 2017 | Beckman Coulter Inc | The firm identified a negative shift of 5mmol/L for sodium recovery in Synchron Control Level 3 when calibrated with AQUA Cal 1. | Nonconforming Material/Component | Recall notices will be sent via mail or email to customers the week of 05/15/2017. Customers are instructed to discard the affected calibrator and request a replacement. Call Client Services at (800) 526-3821 (US only) |  | 4445 | 4444 AQUA CAL Level 1 kits | US Nationwide Distribution in the states of AL, AR, AK, CA, CO, DE, FL, GA, ID, IL, IN, IO, KS, KY, LA, MA, MD, MS, MI, ME, MO, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WV, WI, and Puerto RicoBrunei, Canada, China, Czech Republic, France, Guatemala, Hong Kong, India, Ireland, Italy, Japan, Lebanon, Mexico, Netherlands, Panama, Paraguay, Philippines, Qatar, South Africa, Taiwan, Thailand, Turkey, United Arab Emirates, &amp; United Kingdom. | Open, Classified |  |  | 4445 | Not_Computer
Z-2727-2017 | 77569 | No | JAF | Nonfetal ultrasonic monitor. | Radiology | Radiology | 510(k) | NicoletÂż EliteÂż 2 MHz OB Water Proof Probe, Model #: X1L008, Catalog Code: NW20. |  |  | June 09, 2017 | 2017 | Natus Neurology Inc | Some of the Elite 2 and 3 Obstetric probes and 5 MHz Vascular Probes may exhibit excessive background noise that may mask the users ability to clearly hear heart rate and/or blood flow. | Radiation Control for Health and Safety Act | Customer notification via an Urgent Recall Notification/Field Notice began 06/09/2017. The letter identified the affected product and the reason for the recall. Customers were asked to complete and return the attached verification form. The firm will replace functioning probes that returned free of charge. Customers are to follow the instructions provided for returning any affected devices. |  | 2335 | 2,335 units total | Worldwide Distribution -- US, to the following states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH , OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV and, to the countries of Australia, Canada, Chile, Columbia, Germany, India, Japan, Jordan, Kuwait, Netherlands, Papua Guinea, Philippines, Qatar, Saudi Arabia, Singapore, Sweden, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 2335 | Not_Computer
Z-2735-2017 | 77632 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa Healthcare NX 3.0.8950 Imaging Processing Software |  |  | December 21, 2016 | 2016 | AGFA Healthcare Corp. | A customer reported that when using an NX workstation with software version NX 3.0.8950 software and selecting the affected patient/exam from closed exams, initially the wrong image was linked to the exam and appeared. After a short time the wrong image was replaced by the correct image, however the wrong image was used for transmitting to PACS. | Device Design | On December 22, 2016, an URGENT FIELD SAFETY NOTICE letter was sent either by mail or email. The letter included an acknowledgement form to be returned. The letter informed the customer of the problem and actions to be taken. |  | 221 | 221 units | AL, AZ, AR, CA, DC, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WV and Canada | Open, Classified |  |  | 221 | Computer
Z-2734-2017 | 77505 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Medical Systems Infinity Acute Care System (M540) |  |  | June 06, 2017 | 2017 | Draegar Medical Systems, Inc. | Cockpits with revision index 06 or higher that contain 4GB RAM modules may not annunciate audio or visual alarms on the Cockpit and Central Station. | Software design | Draeger issued on June 6, 2017 Urgent Medical Recall letter with Addendum US serial numbers via Certified Mail addressed to Risk Manager and Director Clinical Engineering The letter instructs the consignees, if the serial number of the cockpit is on the serial number list, to remove the Cockpit from use until it can be determined whether the 4GB RAM Module is among those affected. The IACS can be used without the Cockpit and the M540 will function as the primary monitoring device once the Cockpit is disconnected. A DrÂżgerService Representative will schedule a time to inspect the Cockpit and, if necessary, replace the 4GB RAM Module. The inspection and 4GB RAM replacement (if necessary) will be performed free of charge. For questions regarding the Operation and/or servicing the DrÂżger InfinityÂż Acute Care System Cockpit contact DrÂżgerService Technical Support at 800-437- 2437 (press 4 at the prompt and then press 4 again).For OUS the responsible local Draeger Sales and Service will mail the letter with Addendum serial number list. |  | 719 | 37 US and 682 OUS | AZ, NH, SD, TN, TXForeign:AustraliaAustriaBelgiumCanadaChinaCosta RicaFranceGermanyGreat BritianHong KongIndiaIrelandItalyJapanMalaysiaNamibiaNetherlandsPolandPortugalRussian Fed.South AfricaSouth KoreaSpainSwedenSwitzerlandThailandTurkeyUnited Kingdom | Open, Classified |  |  | 719 | Not_Computer
Z-2863-2017 | 77474 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Turbo-Elite Laser Atherectomy Catheter Peripheral Rapid Exchange Catheters |  |  | June 06, 2017 | 2017 | Spectranetics Corporation | Potential for flouroscopially visible marker band at the tip of the Rapid Exchange Turbo Elite Atherectomy Catheters to become loose or disengaged rom the catheter when the device is in use in an off-label manner. | Device Design | Spectranetics sent a Field Safety Notice letter via certified mail on June 7, 2017, to all customers who have ordered the affected model numbers since 23May2016 reiterating the risks associated with off-label use of the device (i.e., not activating the laser within a stent and not until all contrast media is flushed from the treatment area). The product continues to meet specifications and does not need to be returned.Customers with questions were encouraged to call their local Spectranetics Sales Representative or Spectranetics Customer Service Department. Customer Service can be reached by phone at 800-231-0978 (Option 2), by fax at 877-447-2022, or by email at customer.service@spnc.com. Customer Service is available Monday thru Friday, From 7AM-5PM MST. |  |  |  | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, Hawaii and Puerto RicoForeign distribution to Bahrain, Brazil, Canada, Italy, Japan, and Netherlands | Open, Classified |  |  |  | Not_Computer
Z-2686-2017 | 77201 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee,x-ray system, angiographic |  |  | February 03, 2017 | 2017 | Siemens Medical Solutions USA, Inc | Failures can result due to a component defect on Artis zee systems with an A100 Plus generator of a certain delivery lot and 2-focus Megalix Cat Plus tube unit. | Component design/selection | CDRH approved the CAP subject to the following conditions:1. You completed this recall as of May 6, 2017. Report any failures regarding this recall immediately to the District Recall Coordinator.2. If this recall is expanded to include any additional consignees, models, or end users, any future customer notification letters to be sent must comply with the requirements of 21 CFR 1003.21. This includes, but is not limited to, the prescribed packaging and the statement that you will remedy the defect or bring the product into compliance at no charge. |  | 109 | 97 distributed worldwide, 12 distributed in United States | Worldwide Distribution - US to AR, KS, MI, IA, IL, WI, MN, WV, OH, and NY | Open, Classified |  |  | 109 | Not_Computer
Z-2614-2017 | 77581 | No | RCR |  |  |  |  | Standex International Bakers Pride Microwave Ovens Model E300 |  |  | January 01, 2017 | 2017 | Standex Intl Corp | It was discovered that all five microwave ovens fail to have concealed or inaccessible safety interlocks. | Device Design | The Standex Cooking Solutions Group plans to replace all 149 Bakers Pride Model E300 Speed Ovens remaining in the field as follows :1.Product design changes include (a) an updated interlock switch mounting that reduces the effects of shipping and handling on switch positioning; (b) a revised mechanism for the secondary interlock switch to prevent operation by any part of the human body or any object with a straight insertable length of 3.9 inches. A 200,000-cycle leakage measurement is conducted and will be finished by July 31, 2017. 2.The updated replacement oven will be shipped with the improved, stronger bracing materials and added cover over the handle with markings and changes in the door mechanism.3.New compliant ovens will be delivered by an installer to distributors and owners.4.All shipping and installation costs for delivery of the new units and return of the old units will be paid by Standex International Corporation. 5.The updated and corrected pages of the Service Manual and Owners Manual will be included with the delivery of the replacement oven, and the corrected warning label and Certification label will be permanently affixed to the replacement ovens.6.Standex International Corporation expects to deliver the replacement units after testing, which will be completed by July 31, 2017. Distributors and customers will then be given a period of 4 weeks to return the old unit. 7.All returned ovens will be updated to revised construction and subjected to the full Unit Test Procedure required for new products. Ovens that cannot be updated will be destroyed. 8.Progress reports will be provided to FDA on the 15th of every month until all non-compliant ovens have been replaced with new compliant ovens, returned to the factory and modified to the revised and compliant construction, as verified by conducting the Unit Test Procedure.CDRH approves the CAP subject to the following conditions:1.Notification of all dealers and |  | 149 | 149 | US Distribution | Open, Classified |  |  | 149 | Not_Computer
Z-2741-2017 | 77553 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | MAKO RIO Standard System, internal cPCI Card CageP/N: 201337. |  |  | June 07, 2017 | 2017 | Mako Surgical Corporation | An intermittent electrical problem that could lead to a loss of system power due to a loose screw connection. | Employee error | Customers were sent notices titled "Urgent Medical Device Correction Notification" on 06/12/2017 via UPS Next Day Mail. The firm notifies consignees of the reason for recall and instructs that users of the recalled product be provided the recall notice. Users may continue to use their MAKO system; and, Stryker/MAKO Service personnel will contact users o schedule time to inspect and replace the card cage, if necessary. Also, users are instructed to return the completed Business Reply Form within 5 business days. The completed form can be sent to strykerortho7391@stericycle.com or faxed to 855-544-4807. Any questions can be directed to Clayton Odor at 954-628-0502. |  | 20 | 20 | Distributed to the states of AZ, CA, CO, CT, FL, GA, ID, LA, MA, MO, MS, NC, OH, PA, TN, TX, &amp; WA. | Open, Classified |  |  | 20 | Not_Computer
Z-2744-2017 | 77137 | No | GCW | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | 9400 |  |  | April 13, 2017 | 2017 | Pentax of America Inc | Pentax Medical did not always provide transformers with 9175 isolation transformers are used with 7245C, 7245C/E, 7245D, 9200Cs, 9200Ds, 9310HDs, and 9400s computer systems. | Device Design | Pentax notified their customers on 4/13/2017 via USPS. |  | 241 | 241 units in total | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 241 | Not_Computer
Z-2752-2017 | 77513 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10mL Syringe Luer-Lok Tip Catalog number 301997 |  |  | September 16, 2016 | 2016 | Becton Dickinson &amp; Company | BD identified that several lots did not receive the specified dose of irradiation necessary to meet the Sterility Assurance Level (SAL) required. | Process control | BD sent recall letter/return response form on 9/16/2016. |  | 547600 | 547,600 units | Nationwide Distribution | Open, Classified |  |  | 3173100 | Not_Computer
Z-3074-2017 | 77729 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | POLYURETHANE ENFit(R) AMERITUS NG FEEDING TUBE, WITH SIDE PORT, REF Numbers: |  |  | June 23, 2017 | 2017 | Kentec Medical Inc | Inadequate detectability or radiopaque stripes | Process control | The firm notified their consignees by telephone and email on 06/23/2017, and followed with a letter sent certified mail, on 07/07/2017. The letter instructed the consignee to cease use of the product, isolate it, and contact Kentec. |  | 100526 | K100526 | TX, VA, IL, CA, KS | Open, Classified |  |  | 202132 | Not_Computer
Z-2759-2017 | 77791 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smiths Medical CADD Blue-Striped IV Administration Set, REF 21-7036-01, Sterilized using ethylene oxide. |  |  | December 09, 2016 | 2016 | Smiths Medical ASD Inc. | The label of the CADD Blue-Striped Administrative Set is missing the expiration date, which is considered a regulatory compliance issue for Canada. | Unknown/Undetermined by firm | Customers were sent an Urgent Medical Device Safety Notice, dated 12/9/2016. The letter identified affected product and stated the reason for recall. The letter also provided a list of devices that each consignee ordered, along with the date of manufacture and associated expiration date. Customers were asked to locate devices and compare to the list and determine if affected devices are within expiration date. If the device is expired customers are to dispose of them per their procedures for dealing with expired product. If devices are not expired, they may continue to use them through the expiration date. A response form was asked to be returned within 10 day s of receipt of the letter. Questions can be directed to smithsmedical6185@stericycle.com |  | 641070 | 641,070 units | Canada only | Completed |  |  | 641070 | Not_Computer
Z-2756-2017 | 77668 | No | PJC | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | PerkinElmer 226 Sample Collection Device |  |  | May 31, 2017 | 2017 | PerkinElmer Health Sciences, Inc. | PerkinElmer 226 Sample Collection Devices Lot Number 317114 had serial numbers printed on the pack labels that did not match the serial number range of the devices contained within the pack. | Packaging | PerkinElmer sent an Urgent Medical Device Correction letter dated May 30, 2017, to the sole consignee. PerkinElmer requested the consignee implement the following actions:1. Discontinue use of the pack label to track the device serial numbers distributed to your sample submitters.2. Verify submitter of devices received without provider information prior to demographic entry.3. Verify the serial number of the pack label matches the device serial numbers within the pack of your undistributed inventory prior to distributing to your sample submitter or,4. Return undistributed inventory to PerkinElmer for inspection , and rework if required.For further questions please call (864) 299-8787. |  | 99968 | 99,968 units | US Distribution to the state of : AL | Terminated | July 27, 2017 | 57 | 99968 | Not_Computer
Z-2772-2017 | 77717 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | STOCKERT Heater-Cooler System 3T, 240V/60 Hz |  |  | June 30, 2017 | 2017 | Sorin Group USA, Inc. | Plastic heat exchanger fibers separating the blood compartment from the water compartment in certain oxygenators may allow diffusion of hydrogen peroxide in a quantity that exceeds allowable limits. | Device Design | "Consignees were sent a recall notification, dated 06/30/2017, via mail. Instructions included a recommendation from LivaNova to use its oxygenators, or other oxygenators that are equipped with compatible heat exchangers, to ensure that the permeability rate is within the prescribed allowable limits (ALs). Depending on whether or not the oxygenator was from LivaNova, there were specific instructions to follow for each type and source. Consignees were also instructed to complete and return the Confirmation Form.If product was further distributed, the firm asks that further customers be notified of the recall and that customers contact the LivaNova Quality Assurance Team at usfsn@livanova.com. Questions regarding the recall can be directed to 800-986-4702 or 3T.US@livanova.com. Any adverse reactions are to be reported to FDA's MedWatch Adverse Event Reporting program." |  | 7546 | 7,546 units | "Domestic: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY. Foreign: AL, AE, AR, AT, AU, AZ, BD, BE, BG, BH, BR, BY, CA, CH, CL, CN, CO, CR, CY, CZ, DE, DK, DZ, EC, EE, EG, ES, ET, FI, FR, GB, GE, GR, GY, HK, HR, HU, IE, ID, IL, IN, IQ, IR, IT, JO, JP, KR, KW, KZ, LB, LK, LT, LU, LV, LY, MA, MC, MK, MN, MT, MU, MX, MY, NG, NL, NO, NP, NZ, OM, PA, PE, PH, PK, PL, PR, PS, PT, QA, RE, RO, RS, RU, SA, SE, SG, SK, SV, SY, TH, TN, TR, TT, TW, UA, VN, &amp; ZA." | Open, Classified |  |  | 7546 | Not_Computer
Z-2762-2017 | 77340 | No | KQJ | Fibrinogen determination system. | Hematology | Hematology | 510(k) | Pacific Hemostasis FDP (Fibrinogen Degradation Products) Assay Kit (30 Determinations), REF/Model 100650, UDI 00845275000863, IVD; --- Fisher Diagnostics, a division of Fisher Scientific Company LLC ,a part of Thermo Fisher Scientific Inc., Middletown, VA 22645-1905 USA --- For determination of Fibrinogen Degradation products in serum and urine. |  |  | May 19, 2017 | 2017 | Fisher Diagnostics | An issue was discovered with the Pacific Hemostasis Fibrinogen Degradation Products (FDP) kit, list number 100650, lot 948546 and FDP Latex Reagent, list 100651, lot 890199. While the kit performs accurately at FDP concentrations greater than 20 Âżg/ml, there is an observed reduction in sensitivity when FDP levels are less than 20 Âżg/ml. This reduction in sensitivity has been tied to FDP Latex Reagent lot 890199 and may result in false negatives. | Under Investigation by firm | ThermoFisher Scientific sent an MEDICAL DEVICE RECALL Letter dated May 16, 2017, to all affected customers and response forms were sent to customers on May 19, 2017. Customers are informed of the recall and potential health risks and instructed to review their test results. Review your inventory and any product on hand is to be discarded. --- The firm ceased distribution of the lots and began its investigation. The held product will ultimately be destroyed. Level A effectiveness checks will be conducted to the consignees and repeated attempts will be made to non-responding customers. The firm will track all mailed, emailed and faxed return forms for the duration of the recall.For further questions, please call (540) 869-8188. |  | 331 | 331 units | USA (nationwide) including Puerto Rico, and Internationally to Colombia | Open, Classified |  |  | 358 | Not_Computer
Z-2769-2017 | 77712 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Diagnostics Ultrasound System |  |  | April 24, 2017 | 2017 | Konica Minolta, Inc. | The Doppler velocity measurement on the Sonimage HS1 ultrasound system is not accurate. | Software design | Konica Minolta notified their customers on April 24, 2017. |  | 36 | 36 units | CA, CO, KY, FL, NJ, NY, IA TX, &amp; WA | Open, Classified |  |  | 36 | Not_Computer
Z-2779-2017 | 77508 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Myelotec Video Guided Catheters, REF 2010 3.0 mm OD x 30 Cm Working Length, Two 1.3mm Working Channels |  |  | March 02, 2017 | 2017 | Myelotec, Inc. | Obstructed/blocked port from defective VGC Access Port Body component. | Device Design | Myelotec sent an Urgent Product Recall Notification letter dated April 10, 2017, to all affected customers. Customers are instructed to notify customers of any product further distributed, identify any recalled product on hand, and to complete the attached recall response, to be sent back to the distributor. For further questions, please call (770) 664-4656 Ext. 140. |  | 2980 | 2980 units | Worldwide Distribution - US Distribution to the states of : CA, FL, KY, MN, and NJ ., and to the countries of : Brazil, Iran, Kuwait, Malaysia, Singapore, South Korea, Taiwan, and Turkey | Open, Classified |  |  | 2980 | Not_Computer
Z-2778-2017 | 77607 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | INVUITY(R) PhotonBlade(TM), REF PB1 |  |  | June 12, 2017 | 2017 | Invuity, Inc. | A small number of customer observations were recently received which related to possible unintended energy discharge with the potential to cause tissue damage to patients. | Device Design | Official notices were sent to customers via FedEx overnight on 6/14/17. Email notices were sent to the sales personnel 06/12/2017 regarding the return of their trunk stock inventory and asking them to also contact customers who had been sent letters. |  | 1328 | 1328 units | nationwide | Open, Classified |  |  | 1328 | Not_Computer
Z-2780-2017 | 77490 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Oxford/Oxbridge AFN End Cap |  |  | April 19, 2017 | 2017 | Ortho Solutions Inc | There is a material mislabeling of the AFN End Cap. | Error in labeling | Orthosol notified their customers on 4/19/2017 via USPS. |  | 277 | 277 units | US distribution to MD, NY &amp; PA | Open, Classified |  |  | 277 | Not_Computer
Z-2782-2017 | 76514 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath, Model D140010 (small curve) and D140011 (large curve). |  |  | February 15, 2017 | 2017 | Greatbatch Medical | One side of the MobiCath Bi-Directional Guiding Sheath pouch may not be sealed. | Process control | Consignee was sent a Greatbatch "Urgent Medical Device Recall" letter dated February 15, 2017 via UPS and email . The letter was addressed to Biosense Webster, Inc. The letter described the product involved in the recall, Reason for Recall, Risk to Health,&amp; Instructions to Biosense Webster. Advised consignee to quarantine and return the devices, complete and return the field Recall Response Form to Stericycle, and to communicate the recall to their customers and end users. For questions contact your local Greatbatch Medical representative, Monday through Friday via telephone at 1-763-951-8184 from 8:00 AM (CST) to 5:00 PM (CST), or e-mail us at FieldActionCenter@Greatbatch.com |  | 11360 | 11,360 | CA | Open, Classified |  |  | 11360 | Not_Computer
Z-2785-2017 | 77810 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | SORIN GROUP, SMART PERFUSION PACK, SMA PRIMO2X OPTI PACK, REF 088512900, STERILE EO, Rx Only, 1 EA |  |  | June 29, 2017 | 2017 | Sorin Group USA, Inc. | Medtronic informed LivaNova that some of their DLP Pressure Disposable Display Sets are packaged in pouches which may have small pinholes (approximately 1mm or less in diameter). The result is a potential sterile barrier breach in the display sets. These display sets areincluded with LivaNova Perfusion Tube Sets (PTS) which were then distributed to customers. | Other | The firm initiated their recall verbally on 06/29/2017 and following with letters mailed on 07/13/2017.The actions to be taken by the customer:1. Notify staff to not use the Medtronic labeled device contained within the LivaNova Perfusion Tube Set packs.2. If the set is a standalone component within the pack, remove the device upon opening the pack at the time of the procedure.3. If the device is connected to the circuit, remove the device and attach a new unit upon opening the pack.4. Identify the Medtronic device with one of the barcode pack labels attached to the lidstock of the LivaNova perfusion pack.5. Collect the devices for return to LivaNova until all affected perfusion sets have been consumed. In order to assist in controlling inventory, the LivaNova sales representative will have labels available to identify packs containing the affected devices.6. Contact LivaNova Customer Support at (800) 650-2623 (M-F 8AM to 5PM MT) to arrange for the return of the affected product, and to order equivalent replacement product. |  | 524 | 524 units | Nationwide | Open, Classified |  |  | 2197 | Not_Computer
Z-2781-2017 | 76701 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | ViaPeel PTFE Peelable Introducer 5FR, Part Number 10890-006 |  |  | November 11, 2016 | 2016 | Greatbatch Medical | Greatbatch received complaints with the ViaPeel Introducer Sheath 5FR, of handles detaching during use. | Process control | Consignee was sent via email on 11/11/2016 a Greatbatch "Urgent Medical Device Recall" letter dated November 2016. The letter described the product involved in the recall, Reason for Field Corrective Action, Risk to Health, Replacement Devices &amp; Instructions to Customer. Advised consignees to quarantine the product, contact end users &amp; complete and return the Field Recall Response Form to Stericycle at Greatbatch7893@stericycle.com or fax at 888-473-8015. For questions contact your local Greatbatch Medical representative, contact us at 1-763-951-8235, or e-mail us at FieldActionCenter@Greatbatch.com |  | 5000 | 5,000 devices | NY | Open, Classified |  |  | 5000 | Not_Computer
Z-2799-2017 | 77671 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Stroke Fast Pack(TM) Trevo XP ProVue Retriever 3 x 20, Trevo Pro 14 Microcatheter, AXS Catalyst 060x132CM - US, UPN M0033PK32021001 |  |  | June 26, 2017 | 2017 | Stryker Neurovascular | Stryker Neurovascular has become aware that some 3-Pack Stroke Fast Packs weremanufactured using a carton sleeve where the pre-printed contents on the back of the sleeve did not match the physical contents of the pack. | Packaging process control | Stryker notified their consignees by letter on 06/21/2017. The letter stated the following:This potentially impacts all 3-pack Stroke Fastpacks manufactured in the US. The product quality of the individual products in the Stroke Fastpack is not impacted. All units were manufactured to specification.We request that you read this notice carefully and complete the following actions:1. Immediately check your internal inventory for impacted Catalog numbers.2. Remove and discard the Stroke Fastpack carton sleeve.3. Circulate this Field Safety Notice internally to all interested/affected parties.4. Maintain awareness of this notice internally until all required actions have been completed within yourfacility.5. Inform Stryker if any of the subject devices have been distributed to other organizations.a) Please provide contact details so that Stryker can inform the recipients appropriately.6. Please inform Stryker of any adverse events concerning the use of the subject devices.7. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. Therefore please complete the form even ifyou no longer have any of the subject devices in your physical inventory.8. Return the completed form to your nominated Stryker Representative or to NVFieldActions@stryker.com |  | 43 | 43 packs | nationwide | Open, Classified |  |  | 293 | Not_Computer
Z-2777-2017 | 77756 | No | RCE |  |  |  |  | GE Inspection Technologies Blade¦line Industrial Cabinet X-ray system |  |  | June 28, 2017 | 2017 | GE Inspection Technologies, LP | GE Inspection Technologies, LP determined that the Blade¦line cabinet x-ray system fails to comply with the cabinet x-ray performance standards emission limit of 0.5 milliRoentgen in one hour (21 CFR Âż 1020.40(c)(1)(i)) underneath the bottom center of the system. The crawl space affected by the excess emission is approximately 22 inches high, normally covered by decorative panels, and is not intended to be accessed. | Radiation Control for Health and Safety Act | GE Inspection Technologies will bring the defect into compliance: 1. Four panels of sheet metal/lead/sheet metal construction (6 mm lead thickness) will be permanently mounted using an industrial epoxy to the inside floor of the cabinet to prevent emissions from exceeding the limit in the performance standard. A radiation emission measurement will be made following the repairs of each system to assure the issue was corrected.2. GE will notify affected customers within 5 days of GEs receipt of this approval letter.3. The notification will recommend that customers allow GE to visit the site and install the new shielding within 120 days following GEs receipt of this letter.4.The text of the notification to customers will include the required statement that confirms that GE will, without charge, bring the affected cabinet x-ray systems into compliance with the performance standard. For further questions, please call (717) 447-1278. |  | 13 | 13 units distributed in the US | US Distribution to the states of : AL, SC, OH, MI, PA, and Internationally to Germany and Singapore | Open, Classified |  |  | 13 | Not_Computer
Z-2803-2017 | 77640 | No | PFQ | Osmolality test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | InflammaDry(R) AN RPS DIAGNOSTIC SOLUTION, Sample Collectors |  |  | June 22, 2017 | 2017 | Quidel Corporation | In-house testing revealed the presence of plastic spurs at the tips of some Sample Collectors. | Process design | Quidel sent an Urgent Medical Device Recall letter dated June 22, 2017, to all affected customers. The notice instructed consignees to check inventory for affected lots (11705359 and 1705405), to remove and destroy any affected sample collectors from affected kits, and complete and return the attached Inventory Assessment/Certificate of Destruction form to Quidel by fax at 858-552-7905 or email to technicalsupport@quidel.com. Any questions by consignees should be directed to Quidel's Technical Support Team at 800-874-1517 or 858-552-1100, or by email at technicalsupport@quidel.com. It is requested that consignees keep recall notification and response in their records. |  | 22638 | 1078 kits, 21,560 sample collectors | Nationwide Distribution including AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NY, OH, OK, PA, SC, SD, TX, VA, WA, &amp; WV. | Open, Classified |  |  | 22638 | Not_Computer
Z-2867-2017 | 77405 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MULTIGENT Creatinine (Enzymatic), Catalog Numbers 8L24-31 and 8L24-41; IVD. Intended for the quantitative determination of creatinine in human serum, plasma, and urine using the ARCHITECT c Systems |  |  | May 22, 2017 | 2017 | Sentinel CH SPA | Reagent 1 component may result in Quality Control values shifting out of acceptable range, which is not resolved by recalibration. | Nonconforming Material/Component | Sentinel Diagnostics notified the US Distributor on 5/22/17 to notify US customers by letter. Actions to be Taken1. Discontinue the use of the impacted lots immediately and destroy any remaining inventory of impacted lots according to your laboratory procedures.2. If you have an alternate reagent lot available in inventory, switch to this alternate lot, or use a new lot provided.3. Please review the content of this communication with your Medical Director and retain this letter for any future reference.Questions regarding this information, U.S. Customers should contact Customer Service at 1-877-4 ABBOTT.Customers to complete Field Safety Notice Receipt. |  | 2562 | 2356 kits Worldwide; US 206 kits | Nationwide Distribution | Open, Classified |  |  | 2562 | Not_Computer
Z-2865-2017 | 77669 | No | NSX |  |  | General Hospital | Contact ODE | SoftLab Software |  |  | May 31, 2017 | 2017 | Soft Computer Consultants, Inc. | Display of lab results based on incorrect LOINC code/test descriptions for tests that were performed at a reference lab, saved incorrectly, and sent to systems that display the EMR. | Use error | The firm, SCC Soft Computer, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 05/31/2017to its consignees. The letter described the product, problem and actions to be taken. The consignees were informed that the issue will be corrected with a mandatory hotfix. The Consignees were instructed to acknowledge receipt of the recall notice by signing the recall notice and calling Technical Support at 800-763-8522 to grant permission to load the hotfix.For questions or concerns, please call 800-763-8522 Technical Support, available 24/7, and for any adverse reactions/quality problems experienced with the use of this product are to be reported to FDA's MedWatch Adverse Event Reporting Program either online, by regular mail, or by fax. |  | 18 | 18 units | US Distribution to states of: CA, IL, LA, MA, MD, MI, MN, NJ, NY, PA, and TX. | Open, Classified |  |  | 18 | Not_Computer
Z-3024-2017 | 77488 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR450w 1.5T |  |  | May 11, 2017 | 2017 | GE Medical Systems, LLC | Possible incorrect wiring of the Uninterruptable Power Supply, which could contribute to a potential shock hazard. | Mixed-up of materials/components | Consignees received a written recall notice dated 05/17/2017 notifying them of a correction to all units using an Uninterruptable Power Supply. Consignees were instructed not to unplug the unit and to continue normal use. Units with the Eaton 9130 Uninterruptable Power Supply connected to the MR operator console would receive a correction at no charge. This would involve possibly replacing the power cable. GE Representatives will be contacting consignees to arrange corrections. Any questions can be directed to GE Healthcare Service at 1-800-437-1171 or your local service representative. |  | 68 | 68 units in total | US Nationwide in the states of : AK, CA, CO, CT, DE, FL, IA, IL, IN, KS, KY, LA, MD, MO, ND, NE, NJ, NY, OH, PA, TX, WA, and WI | Open, Classified |  |  | 68 | Not_Computer
Z-2866-2017 | 77653 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | OptiLITE(TM) Products for Laser Surgery, Catalog Numbers: HTM1020F, HTM1040F, HTM1060F, HTM1010F; Manufactured by CONVERGENT LASTER TECHNOLOGIES |  |  | June 23, 2017 | 2017 | Xintec Corp | It has been determined that the reprocessing instructions may not provide sufficiently detailed information for the cleaning, disinfection and sterilization of the products listed below. The preliminary investigation indicates that validation data related to the reprocessing of these devices may not meet newly enacted standards. | Labeling False and Misleading | The firm, Convergent Laser Technologies, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated 06/23/2017 to its Consignees. The letter described the product, problem and actions to be taken.The consignees were instructed as follows: "Please immediately review your inventory and quarantine any affected product. Please return any affected products within 30 days according to the instructions provided in the attached Acknowledgement and Receipt Form. Upon receipt of the product and response form, we will arrange for a credit to your account. Also, whether or not you are returning any affected product, please send a copy of the Acknowledgement and Receipt Form to our Regulatory Department via fax at 510-832-1600 or via e-mail to regulatory@convergentlaser.com."If you have any further questions or concerns, please contact our Customer Service Department at 510-832-2130 or 800-848-8200. |  | 4343 | 4343 units | Worldwide distribution- US (Nationwide) and countries of: Canada and Sri Lanka. | Open, Classified |  |  | 4343 | Not_Computer
Z-2760-2017 | 77759 | No | RBY |  |  |  |  | Olympus Scientific Solutions Americas Corporation ( OSSA ) Model: VANTA Analytical X-Ray System, X-Ray Fluorescence |  |  | July 11, 2017 | 2017 | Olympus Scientific Solutions Americas | Instruments may, under certain specific circumstances listed below, fail to stop when a test is initiated and no sample is present (proximity shut down fails to end test). | Radiation Control for Health and Safety Act | Olympus Scientific Solutions Americas Corp will bring defect into compliance:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Tracking of customer/owner self-installation(s) of the software update should be done to the maximum possible extent via serial-number-tracked downloads, follow-up email reporting, return- mail card or by other similarly effective means. For further questions, please call (781) 419-3900. |  | 146 | 146 units | Worldwide Distribution - USA (nationwide) Distribution | Open, Classified |  |  | 146 | Computer
Z-2868-2017 | 77648 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | ROSA Brain, 3.0.0 |  |  | October 11, 2016 | 2016 | Zimmer Biomet, Inc. | Complaint of the head holder connector locking up mechanically when tightened. | Process design | Product was/will be corrected by Field Technician staff in the form of a visit. |  | 20 | 20 units (16 US, 4 OUS) | Worldwide Distribution - US including CA, FL, MA, MN, NC, NY, OH, PA, DC; Foreign - Australia, and France | Open, Classified |  |  | 20 | Not_Computer
Z-2869-2017 | 77781 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Non-Locking screws in the following lengths: |  |  | July 13, 2017 | 2017 | In2Bones, SAS | The firm received several customer complaints reporting some perioperative breakages of NE0411 Non-Lockingscrews when inserting the screw into the radius diaphysis | Device Design | In2Bones sent a Field Safety Corrective Action letter dated July 13, 2017, to all affected customers. The consignee was directed to remove the products from distribution and return it to the recalling firm. Customers with questions were instructed to contact the Quality and Regulatory Affairs team at: +33 4 72 29 26 26/ +337 61 88 44 30 or by email: qualite@in2bones.com. |  | 287 | 287 screws | Worldwide Diribution - US including Tennessee; Foreign - Italy, Portugal, Spain, Switzerland, UK, and France | Open, Classified |  |  | 287 | Not_Computer
Z-2855-2017 | 77658 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Piston Syringe, 10 CC L/L, 100 per box, 12 boxes/case, 1,200/case |  |  | June 23, 2017 | 2017 | International Medsurg Connection, Inc. | Insufficient documentation to support product sterility claims. | Under Investigation by firm | The firm, International Medsurg Connection Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 23, 3017 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine inventory for affected product, quarantine affected product, return the product, and notify customer to either return unprocessed product in shipping cases or destroy items in opened cartons and to Acknowledge receipt of this notification by one of the following: email: quality@intlmedsurg.com; phone at 847-619-9926 or fax to 847-619-9927.If you have any questions, contact Quality Manager at 847-619-9926 or email: quality@intlmedsurg.com. |  | 54204100 | 54,204,100 units total | Nationwide Distribution including OH, MA, GA, FL, NC, MS, MO, PA, IN, NJ, TX, KS, LA, NE, WA, AL, UT, AK, CA, IL, MN, MD, NY, and MI. | Open, Classified |  |  | 54204100 | Not_Computer
Z-2872-2017 | 77733 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARD DISPOZ-A-BAG Leg Bag with Flip-Flo Valve with Drain Extension, Extension Tube and Latex Straps, 19 oz. Packaged as single unit in thermoformed trays. There are 50 units per case. |  |  | June 02, 2017 | 2017 | C.R. Bard, Inc. | Customer complaint of a cut in the side of the leg bag, indicating a potential breach of the product's sterile barrier. | Process control | Bard Medical a Medical Device letter dated June 2, 2017 to affected customers via FedEx. The letter identified the affected product, problem and actions to be taken. The notification requests that consignees examine inventory and identify any product on hand. The firm indicates that product return is not required, but customers are encouraged to call 1-800-526-4455 if product doesn't meet their needs or if they have any questions regarding the recall. If product was further distributed, the firm requests that the provided recall notification be forwarded to the respective organization. |  | 2350 | 2,350 units | US Nationwide in the states of CA, CO, FL, IL, IN, LA, MN, MO, NC, NJ, NY, OH, OK, PA, TX, WY, including PR. | Open, Classified |  |  | 2350 | Not_Computer
Z-3017-2017 | 77590 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE Software 2.2 |  |  | June 19, 2017 | 2017 | Elekta, Inc. | Incorrect PEC values. | Unknown/Undetermined by firm | Elekta sent an Urgent Important Field Safety Notice 618-01-303-026 letter dated June 2017 to all affected customers. The letter identified the affected product, problem, the potential clinical impact and actions to be taken.. It also provides a workaround to avoid the issue. |  | 110 | 110 units | Worldwide Distribution in the states of AK, CT, GA, KY, LA, MA, MD, MN, NC, NM, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA and countries of HawaiiAustralia, Austria, Belgium, Denmark, France, Germany, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Netherlands, New Zealand, Poland, Slovenia, Spain, Sweden, Switzerland, United Kingdom, &amp; Vietnam. | Open, Classified |  |  | 110 | Not_Computer
Z-2873-2017 | 77500 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PALLAS M/MAXIMIS 5.5 mm Instrument Tap. Part of MAXIMIS Pedicle Screw Spinal Fixation System. Valorem Surgical, Chicago, IL. Product labeled as Pallas M Spinal Fixation System through 2015. |  |  | June 21, 2017 | 2017 | Valorem Surgical LLC | Devices are not consistent with Quality System Requirements. | Device Design | Valorem Surgical sent an Urgent Medical Device Recall letter dated June 21, 2017, informing consignees of the product removal and instructing consignees to return any product to Valorem Surgical LLC to minimize potential risk to patients. Consignees are to fill out the attached Medical Device Recall Return Response form and send it back to the firm by email at global.safety@ValoremSurgical.com or by fax at 312-275-7117, or by mail. For questions pertaining to this recall, please submit an email inquiry to global.safety@ValoremSurgical.com |  | 2225 | 2,225 units total | Nationwide Distribution - US including CA and VA. | Open, Classified |  |  | 2225 | Not_Computer
Z-2871-2017 | 77699 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sorin Group Aortic Arch Cannula, 7 mm x 10 in x 3/8 in, Rx Only, Sterile |  |  | June 09, 2017 | 2017 | Sorin Group USA, Inc. | Identification of excess plastic on the tip of the cannula. | Process control | LivaNova sent an Urgent Medical Device letter dated June 16, 2017 to affected customers via certified mail or e-mail. The letter identified the affected product, problem and actions to be taken. The notice instructs customers to remove all recalled product from inventory and contact LivaNova Customer Support at 800-650-2623 to arrange for product return and replacement. |  | 105770 | 105,770 units | Worldwide Distribution - US Nationwide in the states of : AR, AZ, CA, GA, IL, IN, KS, MI, MN, MO, NC, NE, NY, OK, PA, SD, TX, &amp; VA. and foreign countries of: Canada, Iran, Mexico, &amp; New Zealand. | Open, Classified |  |  | 105770 | Not_Computer
Z-2099-2017 | 77099 | No | MIH |  |  | Cardiovascular | PMA | Zenith Alpha Thoracic Endovascular Graft |  |  | March 22, 2017 | 2017 | Cook Medical Incorporated | There have been five reports of graft thrombosis/occlusion duringglobal commercial use, each following treatment for BTAI.One case resulted in patient death, and three cases resulted inreintervention. | Device Design | The firm, Cook Medical, sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 22, 2017 to the affected consignees via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. On 6/22/2017, updated "URGENT: MEDICAL DEVICE CORRECTION AND REMOVAL" notifications were sent to the affected consignees via courier. The update included removal of the indication for use in BTAI, as follows: The Zenith Alpha Thoracic Endovascular Graft is indicated for the endovascular treatment of patients with aneurysms or ulcers of the descending thoracic aorta having vascular morphology suitable for endovascular repair including:" Iliac/femoral anatomy that is suitable for access with the required introduction systems " Nonaneurysmal aortic segments (fixation sites) proximal and distal to the aneurysm or ulcer: o with a length of at least 20 mm, and o with a diameter measured outer-wall-to-outer-wall of no greater than 42 mm andno less than 20 mm.Additionally, the following warning has been added to describe the thrombus risk that has been observed when the device is used to treat BTAI:" Risk of in-graft thrombus has been observed when the Zenith Alpha ThoracicEndovascular Graft has been used to treat BTAI.Because of the IFU correction to remove BTAI from the indication, it is necessary to remove specific sizes of this device (grafts with a proximal or distal diameter 18  22 mm) that would likely be used only for BTAI.This notice must be shared with appropriate personnel, including down to the user level, within your organization or to any organization where the potentially affected devices have been transferred.Should you have any questions or concerns, please contact Cook MedicalCustomer Relations at 800.457.4500 or 812.339.2235. |  | 4916 | 4916 | Nationwide Distribution. | Open, Classified |  |  | 4916 | Not_Computer
Z-3018-2017 | 77592 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Truliant Tibial Trial Handle. |  |  | June 23, 2017 | 2017 | Exactech, Inc. | The Truliant Tibial Trial Handle's pin may disassociate from the main body. | Under Investigation by firm | Exactech sent an "Important Notification" letter dated June 23, 2017, and Customers were notified via email on June 23, 2017. Instructions included to cease distribution of the product, notify customers if distributed, quarantine any affected devices, return affected devices, and complete and return the response form. For further questions, please call (352) 377-1140. |  | 68 | 68 devices | US Distribution to the states of : AZ, CA, CO, FL, IL, OH, OK, NV, NY, SC and TN. | Open, Classified |  |  | 68 | Not_Computer
Z-3030-2017 | 77696 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | VersaSafe Plastic Cannula - 11 Gauge, Model 9391-0200 |  |  | June 30, 2017 | 2017 | CareFusion 303, Inc. | Punctures in blister packaging that was detected during the packaging process. | Packaging process control | BD sent an Urgent Medical Device Recall letter dated June 2017 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use and return the affected lot to CareFusion for replacement. For recall related questions contact BD Support Center 1-888-562-6018. |  | 1546800 | 1,546,800 units | Worldwide Distribution in the states of : CO, IL, IN, MD, MI, NY, OH, &amp; TX. and foreign countries of: Australia, New Zealand, &amp; Japan. | Open, Classified |  |  | 1546800 | Not_Computer
Z-3020-2017 | 77698 | No | LYJ |  |  | Neurology | PMA | VNS TherapyÂż AspireHCÂż Generator, VNS TherapyÂż AspireSRÂż Generator. Model No. 106. |  |  | June 13, 2017 | 2017 | Cyberonics, Inc | Manufacturing process used to assemble the circuit board may result in some devices experiencing a faster than expected reduction in device longevity. | Manufacturing material removal | Affected hospitals were notified beginning on 06/13/2017 and prescribing physicians were notified beginning on 06/28/2017. Instructions for affected hospitals included examine inventory, quarantine any affected devices, return the devices, complete and return the response form. Instructions for prescribing physicians included to monitory patients frequently and perform diagnostic testing at each visit, ensure patients are using the device regularly to verify stimulation is felt, ensure patients notify their physician if stimulation changes, and complete and return the response form. |  | 1909 | 1,909 units | Worldwide Distribution - US Nationwide and Puerto Rico and the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, Puerto Rico.Distribution internationally to Argentina, Austria, Belgium, Canada, Croatia, Czech Republic, Finland, Germany, Ireland, Israel, Italy, Japan, Lebanon, Netherlands, Norway, Poland, Qatar, Saudi Arabia, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 2305 | Not_Computer
Z-1498-2016 | 73612 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | iVue and iVue 500 (now called iScan) OCT systems with software version 2016.0.0.127 (USA only) and 2016.1.0.127 (out of US) |  |  | March 16, 2016 | 2016 | Optovue, Inc. | A software anomaly was detected that affects Optovue iVue and iVue 500 with Normative Database that may display images of the same eye rather than right and left if user scrolls on the map. | Under Investigation by firm | Recall Letters were sent to customers on March 22, 2016 to inform them of the recall. |  | 40 | 40 | PA, CA, IL, HI. MO, CO, GA, ME, IN, AZ, WV, VA, FL, AR, UT, NE, IN, NV, LA, TN, KS, OK. Foreign distribution forthcoming | Terminated | April 04, 2017 | 384 | 40 | Computer
Z-3031-2017 | 77624 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum 360 Infusion System |  |  | June 15, 2017 | 2017 | ICU Medical Inc | Potential for the Connectivity Engine Module of the affected unit to disengage from the main chassis, causing the unit to shut down without an alarm notification. This would delay initiation of therapy or interrupt an infusion. | Under Investigation by firm | ICU Medical sent an Urgent - Medical Device Correction letter dated June 15, 2017, to all affected customers. The firm contracted with Stericycle for recall notification, response tracking, and effectiveness checks. The customer notification was also posted on icumed.com. The notification requested that consignees verify their unit is working and testing the CE Module. If the unit's CE Module is loose or a blank display occurs, customers are asked to record the S/N and contact ICU Medical Technical Service Support Center at 800-241-4002, option 4. If the unit is functional, return to service. If the unit powers down without an alarm notification, use another unit and contact ICU Medical Global Complaint Management at 800-441-4100. The firm will contact reporting consignees of correction scheduling when parts become available. Adverse reactions or quality problems related to this problem should be reported to FDA's MedWatch Adverse Event Reporting Program by mail, fax, or online. Consignees are advised to inform potential users within their organization of the recall and to complete the recall response form attached to recall notification. The completed response form can be returned by email to ICUMedical7087@stericycle.com or faxed to 855-544-4814. If product was distributed, ask customers to request them to contact Stericycle for a copy of recall notification at 855-544-4813. |  | 68559 | 68,559 units | Worldwide Distribution - USA (nationwide) Distribution to the states of including Puerto Rico : AL, AR, AZ, CA, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NE, NH, NJ, NM, NY, OH, PA, SD, TN, TX, VA, WA, WI, WY and District of Columbia., and to the countries of : Australia, Canada, Chile, Colombia, France, Greece, Hong Kong, Ireland, Italy, Jordan, Kuwait, Mexico, New Zealand, Oman, Peru, Philippines, Poland, Saudi Arabia, Spain, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 68559 | Not_Computer
Z-3032-2017 | 77637 | unit, drying | JRJ | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ThermaSure(TM), a CENORIN(TM) product, SERIES 130 MEDICAL DEVICE DRYER: (a) Model 135; (b) Model 136 |  |  | June 07, 2017 | 2017 | Cenorin, LLC | Dryer unit not heating. Manufacturing issue causes the terminal block to overheat, char and/or burn. | Process control | Cenorin initially contacted all domestic customers at the facility and distributor level via phone on June 7, 2017, which was followed by an Urgent Medical Device Recall letter dated July 20, 2017, to all affected customers. Cenorin sent initial notification to international distributors via email. Customers were asked to immediately examine their inventory for the affected product. If the product exhibits any symptoms of failure, immediately cease using it and contact Cenorin immediately at recall@cenorin.com or call Technical Support at 800-426-1042, ext 224. |  | 34 | 34 units | Worldwide Distribution - US (Nationwide) Internationally to Japan, Hong Kong, Brunei, and Israel | Open, Classified |  |  | 95 | Not_Computer
Z-3035-2017 | 77695 | No | DOF | Lead test system. | Toxicology | Toxicology | 510(k) | LeadCare II Blood Lead Test Kit, containing LeadCare II Lead Controls |  |  | January 20, 2016 | 2016 | Magellan Diagnostics, Inc. | LeadCare II Test Kits controls out of range | Device Design | Magellan notified customers initially by Email, fax and/or verbal communication with Magellan Product Support representatives beginning on January 20, 2016 to advise of the problem and provide revised Assignment Control Value Matrix.. |  | 12931 | 12,931 kits | Worldwide - US Nationwide distribution and the countries of:COLOMBIA, CANADA, THAILAND, INDIA, UAE, AUSTRALIA, GERMANY, CANADA, SOUTH AFRICA, SPAIN, NEW ZEALAND, KOREA | Completed |  |  | 12931 | Not_Computer
Z-3042-2017 | 77794 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Arctic Sun 5000 Temperature Management System, 100-120 V North America and 100-120 V US |  |  | March 31, 2017 | 2017 | C.R. Bard, Inc. | Potential failure of the Arctic Sun 5000 Temperature Management System may occur when the continued use is in excess of 2000 system hours without preventive maintenance. | Unknown/Undetermined by firm | Bard sent an Urgent Customer Notification letter dated March 31, 2017, to each of the U.S. consignees. The letter contained recommended actions and included an Acknowledgement Form to be returned. For further questions, please call (770) 784-6119. |  | 3450 | 3,450 units | Worldwide Distribution - USA (nationwide) to the states of :AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WY., and to the countries of : Australia, Korea, Japan, Canada, Hong Kong, China, Singapore, &amp; EMEA | Open, Classified |  |  | 3450 | Not_Computer
Z-3041-2017 | 77711 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice Pro Automated PD system Model numbers: 5C8310, 5C8310R |  |  | July 05, 2017 | 2017 | Baxter Healthcare Corporation | Baxter Healthcare Corporation has been made aware that users may not be following the instructions in the Operators Manual and incorrectly (using sharp objects) opening disposable set packaging while setting up their Peritoneal Dialysis (PD) therapy, damaging the cassettes for the HomeChoice or HomeChoice PRO cyclers. If damaged cassettes are used, the cyclers may not consistently detect very small holes/cuts in the sheeting of the cassette in the patient valve region and the cycler may deliver air into the patient. | Use error | The firm notified their consignees by letter on 07/05/2017 of the potential for damaging the disposable set if tools are used to open the packaging and directing them to only open the packaging by hand by letter. |  | 23732 | 23732 units | Worldwide Distribution | Open, Classified |  |  | 38876 | Not_Computer
Z-3038-2017 | 77803 | Yes | LWS |  |  | Cardiovascular | PMA | EMBLEM MRI S-ICD Model 219, Subcutaneous-Implantable Cardioverter Defibrillator, Rx. |  |  | June 29, 2017 | 2017 | Boston Scientific Corporation | The device can deliver an atypical amount of energy due to memory corruption inside the device. | Software design | The firm issued notifications dated June, 2017, beginning 6/29/2017 to physicians (implanting and patient follow-up physicians) via hand delivery by their sales representatives. The firm estimated approximately 30% of the U.S. physicians would be receiving the notification via hand delivery. Hand delivery by affiliates in foreign countries to customers began approximately 6/30/2017. Overnight mail letters were issued to the U.S. and foreign physicians starting on/about 7/7/2017 who did not receive hand- delivered notifications. |  | 9200 | Approximately 9,200 devices | Worldwide - US Nationwide distribution, including Puerto Rico, U.S. Virgin Island, and Guam, There was also worldwide foreign distribution, including Canada. | Open, Classified |  |  | 38400 | Not_Computer
Z-3043-2017 | 77820 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Encore(TM) 26 Inflation Device, REF Catalog Nos.: (a) H74904526011, (b) M001151050, (c) M00566670, (d) M0067101140 |  |  | June 15, 2017 | 2017 | Boston Scientific Corporation | Boston Scientific (BSC) has received complaints for cracked trays on the Encore device. | Under Investigation by firm | The firm initiated their recall on 06/15/2017 by letter delivered by Fed Ex in the US, on 06/15-29/2017 in the Americas delivered by Courier, email and hand delivery, on 06/16/2017 in Europe by DHL, on 06/16-19/2017 in the Arab Emirates and East Asia by email and hand delivery and on 06/136/2017 in Japan by hand delivery. In the notice, the firm described the problem and requested the following action: 1. Immediately check the inventory and identify whether there are products with UPNs and lot numbers affected by the recall. 2.Inspect the device tray in the area highlighted above. If any units with a cracked tray are identified in the supply, immediately discontinue use and segregate affected units in a secure location.3.Arrange for the damaged units to be returned to Boston Scientific by completing the attached Reply Verification Tracking Form.4.If no damage is identified, the device may be used as normal. Please complete and return the attached Reply Verification Tracking Form |  | 305779 | 305,779 units | Worldwide | Open, Classified |  |  | 305779 | Not_Computer
Z-3050-2017 | 77823 | No | NCF | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | ORA System with VerifEye, Catalog Number 8065998300 For use during intraocular lens surgery |  |  | June 30, 2017 | 2017 | Alcon Research, Ltd. | Some ORA Carts have the potential to return an incorrect IOL power measurement during cataract surgery. This issue appears to have been caused by a software coding error that results in the lens coefficients for an IOL model being downloaded from the Alcon server in an incorrect order. | Software design | The Market Action was initiated with Initial telephone contact to all affected customers starting June 30, 2017. These customers were informed of the issue, and advised not to use their ORA Carts for calculations with the lens model affected on their specific system. A confirmatory letter (Attachment 4) was also sent via overnight mail to these customers July 12, 2017. As of July 12, 2017, Alcon has identified 8 ORA Carts that are affected by the software coding error. Alcon has reset the IOL databases on all 8 of the identified affected ORA Carts. |  | 367 | 367 units | Worldwide distribution. US Nationwide and Argentina, Australia, Belgium, Brazil, Canada, China, Colombia, France, India, Japan, Mexico, Netherlands, Panama, Portugal, Romania, Spain, Thailand, UAE, and United Kingdom | Open, Classified |  |  | 796 | Computer
Z-3046-2017 | 77917 | sudan iv | KKK | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Convertors Robotic Drape with Attached Leggings; Product is packaged as Single Sterile Drape/8 per Case |  |  | July 28, 2017 | 2017 | Cardinal Health 200, LLC | Product 29474 LAVH drape was packaged as 29574 Robotic Drape with Leggings. | Packaging process control | Cardinal Health sent an Urgent: Product Recall letter dated July 28, 2017.Distributors are requested to notify their customers of the action. An acknowledgement form is included as part of the notification. Responses are tracked. All returned product will be quarantined. All inventory recovered from the field will be destroyed. For further questions, please call (800) 292-9332. |  | 200 | 200 units | US Distribution to the states of : NC, UT, OR, NY,TX, GA, NJ and FL. | Open, Classified |  |  | 200 | Not_Computer
Z-3048-2017 | 77655 | [6 empty] | VASER PRO Amplifier, VASER 2.2 VASERlipo System stem, Rx Only, 115/230V - 60/50Hz 300VA |  |  | June 01, 2017 | 2017 | Solta Medical Inc | Potential inaccurate calibration by service depot repair, which could contribute to patients being burnt by the device. This only impacts units serviced between the dates of 09/08/2016 and 05/23/2017. | Equipment maintenance | Consignees were telephoned on June 1, 2017, as initial recall notification. At that time, all consignees agreed to discontinue use of the affected product until it was returned or repaired. An Urgent Medical Device Recall letter dated June 14, 2017, requested that consignees that have not been contacted regarding this recall and have a device repaired between the dates of 09/08/2016 and 05/23/2017 contact Solta Medical immediately at 877-782-2286. Any questions regarding the recall can be directed to Solta Medical Product Support Team at 877-782-2286 or 510-259-5299, option 2. |  | 7 | 7 units | US Distribution to the states of : AL, CA, CO, FL, NY, PA, &amp; TX. | Open, Classified |  |  | 11 | Not_Computer
Z-3056-2017 | 77739 | No | PGW | Stereotaxic instrument. | Neurology | Ear Nose &amp; Throat | 510(k) | Medtronic Instrument Kit 9733908 Fusion ENT, reusable, RX. This kit includes Kit 9733452 (containing 9733449, Straight Suction, 9733450, 70 deg Curved Suction, and 9733451, 90 deg Curved Suction). The firm name on the label is Medtronic Navigation, Inc., Louisville, CO. |  |  | June 29, 2017 | 2017 | Medtronic Navigation, Inc. | The instruments are not able to pass instrument verification which is performed prior to the use of each instrument. | Device Design | Medtronic sent an Urgent Medical Device Recall letter dated June 25, 2017, via overnight mail on June 29, 2017, to all affected customers. Letters to foreign customers are being handled by geography-specific Regulatory Specialists. Customers were instructed examine their inventory for the affected product and immediately quarantine them for return for Medtronic. Customers with affected product were instructed to contact Medtronic Technical Services at 1-800-595-9709 to receive a return material authorization (RMA) and arrange for their return and no charge replacements. Customers were asked to complete and follow instructions on the attached Consignee Response Form. Sign and date the bottom of the form and then return the form to Medtronic at RS NavFCA@medtronic.com or fax it to Medtronic Technical Services at 651-367-7075. Customers with questions were instructed to call 1-800-595-9709. |  | 15 | 15 kits | Worldwide Distribution - US (nationwide) and Internationally to Canada. | Open, Classified |  |  | 603 | Not_Computer
Z-3066-2017 | 77518 | [6 empty] | Tissu-Trans SFILL 360, 3-TT-SFILL 360. |  |  | June 01, 2017 | 2017 | Shippert Medical Technologies | Seal integrity of the Tissu-Trans product sterile packages cannot be assured. | Packaging | Customers were notified of the recall between 06/01/2017 and 06/14/2017 via FedEx letter. Instructions included remove and quarantine all affected devices, complete and return the response form, arrange for the return of affected devices, and notify any customers if further distributed. |  | 5188 | 5,188 units total | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, MD, MI, MN, MO, NC, NJ, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, WV. and the countries of: Italy, Austria, Turkey, United Arab Emirates, Germany, Spain, Saudi Arabia, Switzerland, Japan Kuwait, Denmark, Sweden, United Kingdom, Belgium, France, Australia, Netherlands, Greece, South Korea. | Open, Classified |  |  | 5188 | Not_Computer
Z-3059-2017 | 77962 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Various custom surgical kits manufactured for CMS, CUSTOM MEDICAL SPECIALTIES, INC.: |  |  | June 15, 2017 | 2017 | Custom Assemblies Inc | The firm manufactured and distributed custom surgical kits which contained BD SafetyGlide" 22GA 1Âż inch Needles. BD subsequently recalled these needles because, " Loose polypropylene particulate in the fluid path could be expelled out of the needle during injection. | Material/Component Contamination | BD sent an Urgent Medical Device Recall letter dated June 12, 2017, to all affected consignees. Consignees were instructed to examine their inventory and quarantine any affected product. Consignees with questions should call 201-847-4267. For questions regarding this recall call 919-202-8462. |  | 2155 | 2155 kits | Nationwide Distribution to NJ, TN, NY, KS, OH, CA, | Completed |  |  | 2155 | Not_Computer
Z-3069-2017 | 77839 | No | MHY |  |  | Neurology | PMA | DBS Lead Depth Stop contained in Medtronic Deep Brain Stimulation (DBS) lead kits, models 3387/3387S, 3389/3389S, and 3391/3391S, The lead kit contains one DBS lead in addition to multiple accessories used in DBS lead implant procedures. Among these accessories is the DBS lead holder, also referred to as the lead depth stop. |  |  | March 17, 2017 | 2017 | Medtronic Neuromodulation | Medtronic received reports that the DBS depth stop did not adequately secure to the lead, which can result in DBS lead placement beyond the intended target. | Process control | Medtronic sent an Urgent Medical Device Correction letter dated March 2017, to all affected consignees with important safety information on the use of the device. No product retrieval is involved at this time. The letter was addressed to Neurosurgeon. The letter described the product involved in the recall, provided information on the "Background", "Potential Risk to Patients", "Actions", &amp; "Additional Information". For questions, contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7 am - pm CDT. |  | 86024 | 86,024 | Worldwide Distribution - US including AL, AR, AZ, CA,CO, CT, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN,M MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV.Foreign Distribution: India, ANZ, Taiwan, Hong Kong, Canada, China, Japan, Korea, Malaysia, Philippines, Singapore, Thailand, Indonesia, Vietnam, Albania, Algeria, Austria, Belarus, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hungary, Iceland, India, Iran, Iraq, Israel, Italy, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Morocco, Netherlands, Norway, Pakistan, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, Argentina, Brazil, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Mexico, Panama, Peru, Puerto Rico, Uruguay, and Venezuela. | Open, Classified |  |  | 86024 | Not_Computer
Z-3067-2017 | 77380 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ultrasound Video Gastroscope, Model EG-3670URK, |  |  | May 23, 2017 | 2017 | Pentax of America Inc | Error in reprocessing instructions in Revisions 13 and 14 of IFU Z845, used with EG-3670URK and EG-3870UTK ultrasound gastroscopes. In the disinfection step, the IFU has a warning indicating the cleaning detergent solution should remain in contact with the all internal channels and external endoscope surfaces. The disinfecting solution, not the cleaning detergent, should be used for this step. | Error in labeling | Recall Notifications were sent to 104 consignees via USPS Certified Mail on 5/23/2017. Included with the Field Safety letter was the customer response form, and the addendum to IFU Z845. These were also posted on Pentax of America web page until the recall is closed. |  | 118 | US: 118 units | Nationwide | Open, Classified |  |  | 234 | Not_Computer
Z-3036-2017 | 77384 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Samsung GU60A system- |  |  | May 24, 2017 | 2017 | NeuroLogica Corporation | When performing the Stitching function on GU60A system, unintended area is exposed to radiation | Device Design | Neurologica will bring defect into compliance : 1.All facilities were notified via a customer notification letter.2.Samsung will release improved software to avoid unintended exposure. When setting the stitching area, if there is an over exposed area, the height of the swivel arm will be lowered over the exposed area and the first region (first image) will be collimated by the bottom blade.3.A software patch was developed: SW3.01 was released on 05/08/2017 and SW1.00 was released on 07/28/17.4.The software patch will be installed on the affected devices.5.Samsung has provided the statement and assurance that all corrections will be performed free of charge.6.All units will be provided with updated software by November 11, 2017.For further questions, please call (978) 564-8500. |  | 87 | 87 units | Nationwide Distribution | Open, Classified |  |  | 87 | Computer
Z-3071-2017 | 77555 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD(R) Medication Cassette Reservoir with clamp and female Luer. Non vented stopper included, REF 21-7002-24, Rx only, STERILE EO, Manufacturer Smiths Medical ASD, Inc. |  |  | June 05, 2017 | 2017 | Smiths Medical ASD Inc. | The medication cassette reservoir, part number21-7002-24, with lot number 16X659, may have been manufactured with the incorrect pressure plate and the tubing used on the cassette may have been routedincorrectly. | Process change control | Consignees (may include distributors and hospitals) in Japan were notified the week of June 5, 2017. Recall letters were delivered by Smiths Medical, Japan sales representatives during the week of 06/19/2017.The consignees were directed to return any affected devices in their possession and receive replacements. Their Smiths Medical sales representative will collect the affected product and replacement product will be sent to them. |  | 17280 | 17280 units | Japan, China, US | Open, Classified |  |  | 17280 | Not_Computer
Z-3072-2017 | 77706 | No | GBT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | mcompass 2-channel Balloon Catheters |  |  | June 14, 2017 | 2017 | Medspira, Llc | Potential failure in the balloon bond in the inner catheter stem, leading to the separation of the balloon from the inner stem, leaving it in the rectal cavity. | Employee error | The firm Medspira plans to contact all foreign and domestic consignees in regards to the recall. International consignees are to destroy recalled product by cutting the catheter shaft and balloon with scissors. It is also requested that their own customers be notified of the recall. Medspira plans to provide product replacement. Domestic customers are asked to return affected catheters to the recalling firm. Medspira indicates that they will cove shipping costs and replacements. All consignees are asked to complete and return the attached recall response form. For further questions please call, (612) 789-0013. |  | 355 | 355 catheters (255 US - 100 foreign. ) | Worldwide Distribution - US Distribution to the states of :CA, FL, IA, ID, and TX., and to the countries of : Austria, Italy, Malaysia, Philippines, Sweden, and United Kingdom | Open, Classified |  |  | 355 | Not_Computer
Z-3076-2017 | 77799 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEKÂż 2 Compact 30, REF numbers: 27530 and 27530R (clinical), and 27630 and 17630R (industry) software and reagent cards designed for the identification (ID) and antimicrobial susceptibility testing (AST) of bacteria and yeast. |  |  | July 12, 2017 | 2017 | Biomerieux Inc | Customers have reported that some VITEKÂż 2 cards are staying in preliminary status, not finalizing after ejection from the instrument, and not allowing cards in subsequent carousel slots to be processed. The issue was reported to occur on VITEKÂż 2 Compact 15 and Compact 30 systems following a system software update to version 8.01. | Other | Consignees (Subsidiaries) of bioMÂżrieux, impacted by the identified issue, were notified of the FSCA and corrective action (FSCA notification) electronically on 07/12/2017. The direct consignees (subsidiaries and distributors) have until 08/12/2017 to notify their respective impacted users.The directions to the user are as follows:Please implement the following actions at this time:" Confirm this letter has been distributed to, and reviewed by, all appropriate personnel within your organization." Refrain from performing the VITEKÂż 2 Systems Software 8.01 update.o A solution has been identified, and your local bioMÂżrieux representative will contact you with further instruction in the coming weeks." If you have updated VITEKÂż 2 Systems Software to version 8.01 within the last seven (7) days, please contact your local bioMÂżrieux representative." Please store this letter with your bioMÂżrieux instrument documentation." Complete the Acknowledgement Form and return it to your local bioMÂżrieux representative." |  | 9234 | 9234 units | Worldwide | Open, Classified |  |  | 11100 | Computer
Z-3077-2017 | 77802 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Fogarty Adherent Clot Catheter Model Numbers 1408010 and 140806 |  |  | July 06, 2017 | 2017 | Edwards Lifesciences, LLC | The diameter of the Fogarty Catheter balloon, measures 6mm instead of 10mm. | Packaging | Recall notification letters were sent to affected customers on 7/6/2017. The recall notice included a response form and instructed consignee to return any unused devices.. |  | 200 | 200 units | Worldwide Distribution - US Nationwide GB, NL, DE, IT, FR, BE, ES, NO, CZ, BG, SE, AT, CH, PL, RU, FI | Open, Classified |  |  | 200 | Not_Computer
Z-3078-2017 | 77593 | No | OJJ | Hydrogel wound dressing and burn dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Burn Sheet, sterile - 60" x 96". Included in First Aid Kits: 36M - Delmarva Power/ACE - w/Logo, Certified Safety Manufacturing logo, and sold blank. |  |  | June 26, 2017 | 2017 | Certified Safety Mfg Inc | Burn Sheet has possibility of being shredded. | Nonconforming Material/Component | Customers were notified via email on approximately 06/26/2017. Instructions included removing any affected burn sheets from first aid kits in stock, notify customers to do the same, return the response form and make arrangements for replacement items if necessary. |  | 214 | 214 kits total (1 burn pad per kit) | US Nationwide Distribution in the states of KY, MD, NY, PA, VA | Open, Classified |  |  | 214 | Not_Computer
Z-3080-2017 | 77845 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | AirLife(TM) Reduced Heparin ARTIRIAL BLOOD SAMPLER, REF 9025RHTR, 1ML SYRINE 25G 5/8", LUER SLIP PROTECTIVE SHEATH, 3.2 IU, STERILE R, Rx ONLY, CAREFUSION |  |  | July 18, 2017 | 2017 | Vyaire Medical | Lot #0001047501 of the 9025RHTR ABG sampling kits have been manufactured containing the incorrect sub-assembly which includes a larger needle (23 Gage, 1" needle), larger syringe (3mL), and larger amount of heparin (113IU), than what is intended for the kit. | Process control | The firm, Vyaire Medical, sent an "URGENT RECALL NOTIFICATION" letter dated 07/18/2017 to its consignees. The letter described the product, problem and actions to be taken. Thee consignees were directed to inspect their current inventory and remove the recalled lot. Distributors and direct consignees were instructed to complete the Customer Response Form and return to GMB-GLB-VSFieldActions@Carefusion.com and destroy the product. If you wish to obtain credit/replacement contact Vyaire Customer Support at (800) 323.9008 (option #1) Monday-Friday 8:00am CST - 5:oopm CST. Distributors were instructed to notify their customers and have them respond directly to the distributor.For any additional questions and support concerning this voluntary Recall, please contact Clinical Risk Coordinator, at (224)706-6830 or Lindy.Schenning@Vyaire.com. |  | 5400 | 5,400 units | Worldwide Distribution- Nationwide distribution to states of: FL, OK, VA, MI, TH, NM, MA, MO, WA, AND country of: Australia. | Open, Classified |  |  | 5400 | Not_Computer
Z-3045-2017 | 77608 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Neusoft NeuViz 128 Multi-slice |  |  | May 19, 2017 | 2017 | Neusoft Medical Systems Co., Ltd. | Software defect | Software design | Neusoft Medical Systems will bring defect into compliance: The defect will be remedied by updating software. Field Change Order I related to software updated will be released to the affected systems in ' Aug, 2017. The affected systems will be upgraded with updated software in the field free of char e in 6 months after the FCO release. |  | 154 | 154 | US Nationwide distribution | Open, Classified |  |  | 388 | Computer
Z-3081-2017 | 77480 | No | HRM | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LX3 Floor Stand for LuxOR Ophthalmic Microscope; Endure LuxOR Q-VUE Ophthalmic Microscope. |  |  | May 25, 2017 | 2017 | Alcon Research, Ltd. | Multiple shipments or forceful impact to a caster on a LuxOr/LX3 microscope may stress the microscopebase, which could contribute to a caster breaking or detaching from the base. | Packaging process control | Alcon sent an Urgent Medical Device Correction letter dated May 25, 2017, to all affected customers. Demonstration units in inventory will be scheduled for retrofit of base and/or upgraded caster configuration through onsite service calls. Customers with questions or concerns were instructed to call 1-800-832-7827.For questions regarding this recall call 817-293-0450. |  | 1074 | 1,074 units | Worldwide Distribution - US (nationwide and Puerto Rico): Units were also distributed globally to the following countries: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Herzegovina, Bulgaria, Canada, Chile, Columbia, Costa Rica, Croatia, Curacao, Czech Republic, Denmark, Dominican Republic, Egypt, El Salvador, Ecuador, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hungary, India, Italy, Jamaica, Japan, Jordan, Kazakhstan, South Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Mongolia, Morocco, New Zealand, Nicaragua, Nigeria, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, Ukraine, United Kingdom, and Uzbekistan | Open, Classified |  |  | 1074 | Not_Computer
Z-3079-2017 | 77633 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MEDLINE SIRUS Surgical Gown, Strong Protection, Poly-Reinforced, Breathable Film Sleeve, XXL-XLONG Sterile Single-Use Gown; Sold in kit with Sterile Gown/Hand Towel/CSR Wrap and boxes of 18 Gowns. |  |  | June 01, 2017 | 2017 | Master &amp; Frank (Pinghu) Co., Ltd. | The sleeve seam of the breathable material was found to have inconsistent sealing, which has the potential to allow fluid to flow pass through the protective seam. | Other | Master &amp; Frank Enterprise Co., LTD is working directly with Medline Industries, Inc., on this recall. A written notification wasemailed to Medline Industries, Inc. on June 1, 2017 notifying them of the recall and further instructions. Consignee was instructed to quarantine and destroy any recalled product and to respond to the recall communication. The attached recall response form requested that the consignee notify all of their own customers of the recall. |  | 2160 | 2,160 units | Domestic: IL. | Open, Classified |  |  | 2160 | Not_Computer