Z-0917-2011 | 57068 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva with Digital Image (DI) Rel.2 system. |  |  | March 06, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Images from two different patients may be combined in one patient folder in the Digital Image (DI) or may be combined as one (printed) image. It is possible that an image of a previous patient registered in the DI will be overwritten with an image of a later patient. | Other | On March 6, 2009, Philips sent an "URGENT-Device Correction Easy Diagnost Eleva, MultiDiagnost Eleva, and OmniDiagnost Eleva with DI Rel. 2 systems notification" to their consignees/customers.The consignees/customers are advised that before diagnosing images, always check the stored (Digital Image, Viewform or PACS) or printed images to ensure that these represent the data and images captured from the examination of the respective patient. Do not use images if they do not represent the patient that was examined. Consignees will be contacted by Philips to schedule a software upgrade of their units.The consignees can contact Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FCO 70800070" with any questions about this recall. |  | 43 | 43 units | Nationwide distribution: USA including states of: AZ, CA, CO, CT, DE, FL, GA, HI, ID, LA, MA, ME, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OR, PA, SC, TN, TX, VA, WA, WI, and WY. | Terminated | June 16, 2011 | 832 | 95 | Software | Display/Image | Software update | Software Update
Z-0923-2011 | 50074 | No | MGA | Alpha-globulin immunological test system. | Immunology | Immunology | 510(K) Exempt | K-ASSAY Alpha-1 Microglobulin Reagent Kit. Kit contains: 1 vial Reagent 1, Buffer 50 mL, and 1 vial reagent 2, Antibody Reagent 50 mL. The product is packaged in a kit box along with with the package insert. |  |  | September 17, 2008 | 2008 | Kamiya Biomedical Company, LLC | Standardization of the reagent performance data did not match the standardization used to assign the calibrator values. | Employee error | Kamiya Biomedical Company, LLC sent an urgent medical Device Recall letter dated September 17, 2008, and the corrected package insert via e-mail. . The letter identified the product, the problem, and the action the customer should take. Customers were instructed to replace their previous package inserts with the corrected version and asked to sign and FAX the letter back once that had been completed. As one customer is also a distributor, the recalling firm told the this distributor's representative to advise their single customer of this correction and replace the package insert, and return the signed letter back to the firm. On September 22, 2008, both customers were contacted by phone and reminded to fax the signed letters once completed.For questions regarding this recall call 1-800-222-0342. |  | 5 | 5 kits | Kits were distributed to one laboratory in MN and to one distributor in PA who further distributed product to one customer. | Terminated | January 20, 2011 | 855 | 5 | Not_Computer
Z-0922-2011 | 56356 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker brand Cart with Isolation Transformer, |  |  | April 16, 2010 | 2010 | Stryker Endoscopy | The electrical cord power plug design may allow the prongs to develop a fracture inside or outside the plug, if the plug is misused. | Device Design | Stryker Endoscopy sent an URGENT: Device Recall letter dated April 16, 2010, to Stryker Sales reps and all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Locate the affected units.Visually inspect each cord for damage.If the cord has bent or cracked prongs the unit should be taken out of service immediately. The customer was instructed to call a Regulatory Affairs Recall Representative at (408) 754-2738 for further instruction. If the cord did not show signs of damage, it could remain in service.An acknowledgement form was included to document confirmation of receipt of the letter. Upon completion of the power cord inspection, customers were to sign the acknowledgement form indicating that all of their units were inspected. Customers were asked to mail the form to Stryker Endoscopy or fax to (408) 754-8378 or scan the card and email to PowerCord@stryker.com.For questions regarding this recall call the Stryker RA/QA Department at (408) 754-2738 or (800) 435-0220, Reference RA 2010-034. |  | 7371 | 7371 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MA, ME, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, MY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Canada, Chile, China, India, Korea, Mexico, Singapore, Taiwan and Latin America. | Terminated | January 19, 2011 | 278 | 7371 | Battery
Z-0926-2011 | 57573 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DPOC-20, Direct Peroral Cholangioscopy Balloon, G54951, 20 mm/17 mL, Rx Only, Cook Medical, Winston-Salem, NC 27105 |  |  | December 23, 2010 | 2010 | Cook Endoscopy | There is a risk of gastrointestinal perforation or air embolism due to balloon deflation difficulty, resulting in poor patient outcomes. | Other | The firm, Cook Medical, sent an "URGENT: PRODUCT RECALL" letter dated January 7, 2011, to all customers. The letter described the product, problem and the actions to be taken by the customers. The customers were instructed to quarantine any affected product and return to Cook Medical. If they are distributors, they were instructed to contact their customers, direct them to quarantine any affected product and coordinate the return of the affected product to Cook Medical. The customers were also instructed to complete and return the bottom of the letter and send to Cook Medical as well as the PRODUCT RECALL EFFECTIVENESS SURVEY via fax at 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. Note: All non-responding customers will receive a total of three (3) recall notifications ( two (2) follow up requests for a response in addition to the initial recall notification). Should you have any questions, please contact the Customer Relations Department at Cook Medical for more information. Please use CustomerRelationsNA@cookmedical.com or call toll free at 1-800-457-4500, press 4 then enter extension 2146. |  | 1 | see line 1. | Worldwide distribution: USA including states of: AL, CA, FL, IL, IN, LA, MA, NC, NY, OH, PA, SC, TN, TX, WA, and WI; and countries of: Belgium, Canada, Germany, Hong Kong, and the Netherlands.. | Terminated | August 04, 2011 | 224 | 1 | Not_Computer
Z-0928-2011 | 56931 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio Fixed Femoral A/P Sizing Guide; non-sterile instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; A subsidiary of Stryker Corporation |  |  | September 27, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that there is a potential for the shoulder screw on the Scorpio Sizing Guide (the "Sizing Guide") to disassociate from the main body of the instrument. | Device Design | Stryker Orthopaedics sent an 'Urgent Update Product Recall' letter dated December 30, 2010, along with Product Acknowledgement forms to all customers (US and international). The letter described the product, problem and action to be taken by the customers. The customers were asked to complete and return the PRODUCT RECALL ACKNOWLEDGMENT FORM within five (5) days via fax at 201-831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product; and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team, at (201) 831-5028. |  | 6 | six units | Worldwide distribution: USA and countries including: Australia, Brazil, China, Columbia, France, Greece, India, Italy, Korea, Poland, Singapore, South Africa, and United Kingdom. | Terminated | July 31, 2013 | 1038 | 243 | Not_Computer
Z-0924-2011 | 56859 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho ProFx 6850 Operating Room Table, with orthopedic accessories, AC Powered, manufactured by Mizuho OSI, Union City, CA. |  |  | March 30, 2010 | 2010 | Mizuho Orthopedic Systems Inc | Device component spar gear box locks up and may result in failure of push pull tube, which may result in release of breaking and impact. | Device Design | All customers will be contacted to arrange for removal and shipping arrangements. Verification will be done by the Repair and Verification instructions for the leg spars, via registered letters.*******UPDATED INFO---The firm sent another notification letter---dated 12/22/10---which states in part: "****The upgrade involves the correction of your 6850 PROfxÂż spars (6850-2100) at your facility by a Qualified Technician. The spars shall be upgraded to remove the 6850-2123, Drive Bracket, Spar Push Rod and replace it with the 6850-2170, Drive Bracket, Spar Push Rod. Upon completion of the upgrade, the 6850-2123, Drive Bracket, Spar Push rod shall be returned to Mizuho OSI for processing. ******* |  | 17 | 17 units | Worldwide Distribution -- USA &amp; Germany. | Terminated | May 17, 2011 | 413 | 17 | Not_Computer
Z-0929-2011 | 56403 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CG8+ Cartridge; |  |  | August 03, 2010 | 2010 | Abbott Point of Care Inc. | Two cartridge lots were identified that may generate falsely elevated PCO2 results. | Other | Abbott Point of Care Inc. sent an Urgent Recall Notice letter dated July 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the cartridges from the identified lots and follow the instructions on the enclosed reply card to receive replacement cartridges. If cartridges were distributed from the lot numbers to another institution, a copy of the recall letter should be provided to them.For any questions regarding this recall contact Abbott-Point of Care Technical Support at 800-366-8020, Option 1. |  | 55350 | 55350 total cartridges | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, and WY and the countries of Oman and Turkey. | Terminated | December 05, 2011 | 489 | 55350 | Not_Computer
Z-0933-2011 | 57334 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter FloGard 6201 Volumetric Infusion Pump. Baxter Healthcare Corporation, Medication Delivery Division. Product code 2M8063 |  |  | November 17, 2010 | 2010 | Baxter Healthcare Corp. | The FloGard Infusion Pump was released from service without having the Current Leakage test performed. This could lead to the patient/user to receive an electrical shock. | Process control | The firm, Baxter, telephoned the one customer on November 17, 2010, informing them that electrical safety testing was not performed on their FloGard pump serial number 11091330FA during the last service. Since the test was not conducted on the pump, the device may not function as designed and could lead to a hazard of electrical shock. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter advised the customer that a Baxter Field Service Engineer would be sent to their facility to perform electrical safety testing or alternatively, the customer may choose to send the pump to Baxter for testing.If you have any questions regarding this information, please contact Vice President, Quality, GIS at 847-270-5507. |  | 1 | 1 pump | Nationwide distribution: WA only | Terminated | March 16, 2011 | 119 | 1 | Hardware | Physical Safety Hazards | Replace/Return | Remove or Replace
Z-0936-2011 | 57572 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Super Arrow-Flex(R) PSI Set |  |  | December 20, 2010 | 2010 | Arrow International Inc | Packaging may have been compromised, and sterility of product cannot be guaranteed. | Other | The firm, Teleflex Medical thru its subsidiary Arrow International, Inc., sent an "Urgent Field Safety Notice" letter dated December 20, 2010, with Field Acknowledgement form to all customers. The letter described the product, problem and action to be taken by the customers. The letter instructed the customers to check their stock for affected product, cease use and distribution of product and quarantine all products from affected lot immediately; contact their local Teleflex Medical Customer Service/Sales Rep (or Distributor) for a return authorization number prior to filling out the FIELD SAFETY CORRECTIVE ACTION ACKNOWLEDGEMENT AND STOCK STATUS FORM; and complete and return the form immediately by email. If you do not have the product from the affected lot, indicate it on the form and fax it to your local Teleflex Medical Customer Service/Sales Rep (or Distributor). The Teleflex Medical Customer Service/Sales Rep (or Distributor) will coordinate the product return with the customer.If you have any questions, please contact the Regulatory Affairs Project Manager at 610-378-0131 ext 3318 or email: Jody.Cadd@teleflexmedical.com. |  | 320 | 320 units | International distribution only: countries include: Austria, Belgium, Czech Republic, Germany, Greece, Italy, and Luxembourg. | Terminated | September 25, 2014 | 1375 | 320 | Not_Computer
Z-0931-2011 | 57515 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Single Channel Volumetric Infusion Pump. Baxter Healthcare Corporation, Medication Delivery Division. Product Code: 2M8161 |  |  | December 01, 2010 | 2010 | Baxter Healthcare Corp. | While the infusion pump was being serviced at the Baxter Service Center, it was inadvertently routed to the packaging area before the final inspection. | Employee error | Baxter telephoned the one customer on December 1, 2010, informing them that that an outgoing Quality Control Evaluation was not performed on their Colleague pump serial number 15030799CC during the last service. Since this was not conducted on this pump, the device may not function as designed and could lead to an interruption of therapy. Baxter requested that they locate the pump and remove it from use as soon as possible. Baxter will then send a shipping container for retrieval, reinspection and return of the device back to the customer at no cost.For questions regarding this recall call (800) 437-5176. |  | 1 | 1 pump | Texas | Terminated | May 13, 2011 | 163 | 1 | Not_Computer
Z-0932-2011 | 57527 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC); |  |  | December 14, 2010 | 2010 | Varian Medical Systems, Inc. Oncology Systems | Multiple isocenter plans will load on the 4DITC and are available for imaging and treatment, but the 4DITC software does not recognize that mulitplex isocenter is present. | Software design | Urgent Medical Device Correction/Field Safety notification letters were distributed to all affected users, and to Varian Sales, Marketing and Service Organizations on December 14, 2010 by certified mail. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters described the issue and recommended user actions along with Varian's actions. Users are to use caution when acquiring actual couch positions for any plan. In addition, each plan should be reviewed for quality assurance purposes in accordance to set departmental or organizational policies. The firm updated the procedure and enclosed a copy with the letter. It was also posted to the customer support site. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 4188 | 4188 | Worldwide Distribution. | Terminated | July 22, 2011 | 220 | 4188 | Software | Treatment/Delivery/Therapy | Procedure update | Other
Z-0934-2011 | 57612 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Lubricating Jelly, Sterile, Non-staining, Made in the USA, packaged into the following size containers 5 gram tube, 2 ounce and 4 ounce tubes, 4 ounce bottle, 3 and 5 gram packets |  |  | December 23, 2010 | 2010 | H &amp; P Industries, Inc. dba Triad Group | Triad Group is recalling all lots of Sterile Lubricating Jelly manufactured by Triad Group. This recall has been initiated due to concerns expressed by the Food and Drug Administration regarding the validation of the gamma radiation sterilization cycles for these products. | Process control | The firm, Triad Group, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 22, 2010, to all customers via US Postal Service certified. The letter described the issue, identified affected product as that which contained a lot number beginning with the digits 7, 8, 9, or 0, and the action to be taken by the customers. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall; notify their customers at once if the recalled product was further distributed, and complete and return the enclosed Recall Acknowledgement form via fax at 262-538-2947 or mail along with the recalled product to: Recall Coordinator, Triad Group, 700 West North Shore Drive, Hartland, WI 53029.If you have any questions, please call Triad Group Customer Service Monday thru Friday, between the hours of 8:30 a.m. and 4:00 p.m. Central Time at 262-538-2900 ext 2761. |  | 453787 | 453,787 cases | Worldwide distribution: USA including states of: CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MI, MO, MS, MY, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI and Puerto Rico; and countries including: Australia, Belgium, Canada, China, Lebanon, New Zealand, Philippines, Northern Marianna Islands, Saudi Arabia, Singapore, Taiwan, and United Arab Emirates. | Terminated | June 11, 2012 | 536 | 453787 | Not_Computer
Z-0935-2011 | 57516 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX and CXE Infusion Pumps, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8161 and 2M9161. |  |  | December 01, 2010 | 2010 | Baxter Healthcare Corp. | The Air In Line (AIL) test was not properly performed by the service technician at the Service Center. | Employee error | Baxter telephoned the two customers on December 1, 2010, informing them that that a post repair test was not performed correctly on their Colleague pumps (serial numbers 12070599CC and 17086323CP) during the last service. As a result, the device Air Sensor may not function as designed and could lead to undetected air in the IV line. Baxter wants to ensure that the devices are properly tested, safe and effective. Baxter requested that they locate the pumps and remove them from use as soon as possible. Baxter will then send a shipping container for retrieval, reinspection and return of the device back to the customer at no cost.*** Baxter expanded the recall to include twelve additional pumps on June 22, 2011 by telephone calls to the affected accounts. The recall strategy remains the same as the December 1, 2010 instructions. **** |  | 14 | 14 pumps | Nationwide Distribution -- Arizona, Indiana, *** California, Iowa, New Jersey, Ohio, Virginia and West Virginia***. | Terminated | January 20, 2012 | 415 | 14 | Not_Computer
Z-0938-2011 | 50791 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Crosser 14S Catheter, Catalog/REF CRU14S, Lot number 5053 manufactured by Flowcardia Inc, Sunnyvale, CA |  |  | January 09, 2009 | 2009 | Flowcardia Inc. | Mislabeled expiration date: Some units were mislabeled with a three year shelf life, but the product is validated only for a two year shelf life. | Employee error | The firm, FlowCardia, Inc., sent an "URGENT! Medical Device Recall" letter dated January 9, 2009 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to return the product to FlowCardia, at the firms expense, for replacement and Confirm that this product has been used by completing and returning the attached fax document via fax to 408-616-0053. The customers were ask to please make an attempt to contact the customer service representatives, Operations specialist at 888-456-7286 or Vice President, Clinical, Quality &amp; Regulatory Affairs at 408-617-0352 x302, within the next 30 days. Note: FlowCardia Sales, Clinical and Customer support will send notification via certified mail or traceable courier and track response.If you have any questions or need any additional information email (dmichaels@flowcardia.com) or telephone at 408-617-0352 x302. Alternatively you may contact customer service at 888-456-7286. |  | 35 | 35 units | Nationwide distribution: CA, NJ, LA, TX, GA, MI and CO. | Terminated | January 25, 2011 | 746 | 35 | Not_Computer
Z-0937-2011 | 56804 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn PIC 50 Portable Intensive Care System. Model 971081 series. |  |  | September 17, 2010 | 2010 | Welch Allyn Protocol, Inc | Potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail. The loss of the ability to view the patient's Electrocardiogram (ECG) on the screen could cause the user to delay or fail to administer defibrillating shocks needed to resuscitate a patient. | Other | On 9/17/10, Welch Allyn began sending out the Important Product Alert letter to all their direct consignees. On 9/24/2010, the firm sent out the letters to their distributors.The Product Alert letter informs the consignees of a low potential likelihood that the Portable Intensive Care (PIC) 50 Defibrillator's screen display may fail.The consignees are advised to follow the instruction in the PIC 50 user's manual requiring routine shift equipment checks daily or at each shift change to minimize the likelihood of the malfunction from occurring. The checks ensure that the PIC 50 and its component (including the screen display) are fully operation and undamaged.The letter further states "our analysis indicates that routinely doing such checks will likely lead to discovery of this malfunction as it occurs, and thus timely maintenance and repair".The distributors are advised to communicate this Product Alert to their customers if the devices were further distributed. For customers in US and Canada, distributors are advised to provide the electronic Excel file with customer's name, address, serial number and part number to WA2221@Stericycle.com. For customers in other countries, they are advised to provide information to PIC50Product Alert@welchallyn.comConsignees can call Welch Allyn at 800-289-2500 for any questions during normal business hours. |  | 3606 | 2002 units to US consignees and 1604 units to foreign consignees | Worldwide Distribution -- US, including states: AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and countries ofAustralia, Austria, Belgium, Canada, China, Columbia, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Jordan, Kuwait, Kyrgyzstan, Latvia, Lebanon, Liechtenstein, Lithuania, Malaysia, Mexico, Oman, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and Yemen. | Terminated | July 25, 2011 | 311 | 3606 | Other | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0949-2011 | 57470 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | The Stryker InterPulse system is a single patient use, sterile, disposable, pulsed lavage device. The system consists of an irrigation handpiece, irrigation tubing, irrigation tips, suction tips, suction tubing, and splash shields. |  |  | October 04, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | There is potential for a breach in sterility in the blister packaging. This can lead to the use of non-sterile product which, if contaminated, could lead to a patient infection. | Packaging process control | Stryker sent a Medical Device Recall Notification letter dated October 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and dispose of all affected product and to complete the business reply form and return to Stryker to receive replacement interpulse products at no charge.For questions regarding this recall call (269) 323-7700. |  | 954 | 954 | Worldwide Distribution - USA (including all 50 states) and the countries of Brazil, Canada, France, Germany, Japan, Netherlands, South Africa, South Korea, Spain, Sweden, Switzerland and the UK. | Terminated | June 06, 2011 | 245 | 3187 | Not_Computer
Z-0955-2011 | 57380 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Newdeal (Newdeal Compression Plate) Model #UNI-CP 339900 Sets, Lot Numbers: E71H, E7V1, E86R, EELE, EGTC |  |  | November 18, 2010 | 2010 | Integra LifeSciences Corp. | Integra LifeSciences Corporation has decided to initiate a voluntary recall of the Uni-Comp Kits due to inaccurate information relative to the recommended sterilization parameters described in the products Instructions for Use (IFU). The sterilization parameters noted in the Uni-Comp IFU were incorrect. It is noted that the sterilization recommendations in the UNI-Comp IFU in which the Exposure | Other | The firm, Integra, contacted the Sales Representatives on November 8, 2010 via telephone conference calls. Integra also sent a "URGENT: PRODUCT RECALL NOTIFICATION" letter dated November 9, 2010 all customers via FedEx overnight delivery. The letter described the product, problem and the action to be taken by the customers and the firm. The customers were instructed to immediately examine their inventory, QUARANTINE THE PRODUCT, and complete and return the RECALL ACKNOWLEDGMENT AND RETURN FORM via mail: Integra LifeSciences Corporation, Attn: Sandy Yates, 311 Enterprise Drive, Plainsboro, NJ 08536, fax to (609)-275-5363 or Scan and email to sandra.yates@integralife.com. The firm will have a Integra LifeSciences Corporation Sales Representative contact the customer to assist them in retrieving and returning their Uni-CP Compression Plate System to Integra. Integra will also ship replacement Uni-CP Compression Plate (Uni-CPII and Large Uni-Clip) Systems back to the customers as soon as possible at no charge.Should you have any additional questions, contact the Director of Marketing, Hindfoot-Extremity Reconstruction or your regional managers at 513-533-7929, 609-275-0500 or toll free at 800-654-2873. |  | 67 | Sixty seven (67) Kits | Nationwide distribution: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA and WI. | Terminated | November 18, 2011 | 365 | 67 | Not_Computer
Z-0966-2011 | 57708 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological, Inc. Koyle Diaper Stent 6.0 Fr, Sterile, GPN G17667, REF KDSCF-060050. |  |  | December 13, 2010 | 2010 | Vance Products Inc. dba Cook Urological Inc. &amp; Cook OB/GYN | The firm received a complaint that 15 units of the Koyle Diaper Stent were received in packages that had not been sealed. Each stent should be sealed in a poly-Tyvek pouch. | Packaging process control | The firm sent an URGENT RECALL NOTICE to one consignee on 12/13/2010 (the other consignee returned all their product as a complaint). The letter informed the consignee of the recall and asked the customer to return the recalled product. |  | 25 | 25 stents | CA, CT | Terminated | January 31, 2011 | 49 | 25 | Not_Computer
Z-0954-2011 | 57498 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien, LigaSure, Blunt Tip Laparoscopic Sealer/Divider, 5 mm-37 cm, REF LF1537, Sterile EO, Covidien, Valleylab, a division of Tyco Healthcare Group LP, 5920 Longbow Drive, Boulder, CO 80301 |  |  | December 08, 2010 | 2010 | Covidien LP | During surgical procedures, the device may energize when clamping without pushing the activation button, which could result in unintended energy delivery to tissue. | Device Design | The firm, Covidien, sent a "URGENT: PRODUCT RECALL" letter dated December 6, 2010, to the customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to remove product from inventory, immediately quarantine the product and return any remaining stock of the lots listed to: Covidien Energy-based devices, Attention: R82, Building 4/Receiving, 5920 Longbow Drive, Boulder, CO 80301. The customers were also instructed to complete and return the Customer Response Form immediately via mail (in the envelope provided) or fax to (303) 516-8307.To return any identified product for credit or replacement, contact Customer Service at 1-800-722-8772 Option 1.Should you have any medical questions concerning this recall, please contact your local sales representative or the Covidien Clinical information Hotline at (800) 255-8522 (option 1). |  | 25009 | 25,009 units | Worldwide distribution: USA and countries including: Australia, Canada, France, Germany, Japan, Puerto Rico, Singapore, and Uruguay. | Terminated | March 27, 2012 | 475 | 25009 | Not_Computer
Z-0939-2011 | 57630 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Linemaster Wireless Footswitch, VASER System, REF Vaser 2.0, Sound Surgical Technologies LLC, 357 South McCaslin Blvd., Suite 100, Louisville, CO 80027 |  |  | December 21, 2010 | 2010 | Sound Surgical Technologies, LLC | Surgical equipment, activated by a wireless foot switch, may remain powered on when switch is no longer depressed. | Process control | Customers were notified by e-mail with an Advisory Notice on 12/21/2010. The Notice identified the affected product and described the issue. The Notice also explained what the potential hazard involved along with a mechanism to eliminate the potential hazard. Sound Surgical would arrange a service call to implement the correction at the firm's expense. If customers experience this potential failure, they are to contact Sound Surgical Technologies at 1-888-677-8765 or 1-888-471-4777. |  | 27 | 27 units | Nationwide Distribution -- AL, AR, AZ, CA, CO, DC, FL, GA, IA, IL, IN, MA, MI, MO, NJ, NY, PA, and TX . | Terminated | February 08, 2011 | 49 | 27 | I/O
Z-0967-2011 | 57188 | [6 empty] | StaXxÂż XD Expandable Cartridge Convex Expandable Implant 25mm (L) x 9mm (W) x 7mm (H) |  |  | October 22, 2010 | 2010 | Spine Wave Inc | Package Insert Incorrect | Packaging process control | Spine Wave notified each affected customer on October 22,2010, via telephone of the product, the issue and action they should take. Customers were asked to segregate the product from other products and immediately return to Spine Wave.For questions regarding this recall call 203-944-9494. |  | 12 | 12 units | Nationwide Distribution including Delaware, Massachusetts, Maryland, and New York | Terminated | March 15, 2011 | 144 | 12 | Not_Computer
Z-0963-2011 | 57609 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | SRD S5 Mast Roller Pump System 85, S5 Component, Part No.: 50-80-70, Made in Germany, Distributed by: Sorin Group USA, Inc., Arvada, CO 80004. |  |  | November 12, 2010 | 2010 | Sorin Group USA, Inc. | Touch screen may become unresponsive, inhibiting user input. | Device Design | All affected US customers and distributors were notified by certified mail on Nov 12, 2010 via a Field Correction letter. They were told that they would be contacted by Sorin to arrange an appointment to inspect and replace affected touch screens on site. They were told that they could safely continue using their S5 systems in accordance with the Operator's Manual and previously supplied instructions until the inspection and necessary replacement had been completed. A customer response form was also included for completion and return to the firm. Customers can contact the firm if they have any questions. |  | 16 | 16 units | Worldwide Distribution. | Terminated | July 13, 2012 | 609 | 6671 | Not_Computer
Z-0968-2011 | 57231 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sternal Saw II Sterile Blades for the Sarns Sternal Saw System, Catalog number 5589, Terumo Cardiovascular System Corp, Ann Arbor, MI. |  |  | October 20, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The firm discovered that some outer box labels contain incorrect information regarding the method of sterilization. The label incorrectly indicates the product was ethylene oxide sterlized but the product actually was sterilzed with gamma radiation. | Labeling design | Terumo Cardiovascular Systems Corporation firm sent out an 'URGENT MEDICAL DEVICE CORRECTION: SAFETY ADVISORY" letter dated October 20, 2010, to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Safety Advisory, assure that all users were aware of the notice, and confirm receipt of the Safety Advisory by faxing the attached Customer Response Form to the fax number on the form. For questions regarding this recall call 1-800-521-2818. |  | 181902250 | 42,650 units, 4265 boxes of 10 each | US, Indonesia, UAE, Hong Kong, Japan, Singapore, Malaysia, Taiwan, Belgium, Uraguay, Venezuela, Columbia, Mexico, Guam, Vietnam and Thailand. | Terminated | December 14, 2011 | 420 | 181902250 | Not_Computer
Z-0986-2011 | 54350 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LT -D Laryngeal Tube, Size 4, KLTD204, King Systems Corporation 15011 Herriman Boulevard, Noblesville, IN 46060. |  |  | December 15, 2009 | 2009 | King Systems Corp. | King Systems was issued a warning letter from the Center for Devices and Radiological Health, Office of Compliance on October 28, 2009. The warning letter asserts that Kings Systems KLT(S)-D airway exceeds cleared indications in that it is promoted for "airway management in patients over 4 feet in height, for controlled or spontaneous ventilation, difficult and emergent airway cases, and ambulat | Labeling False and Misleading | On October 29, 2009, the firm, Kings System, sent a "Urgent - Product Correction" letter to the customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the promotional material from the King System website which includes literature, and Instructions for Use. The firm instructed all Sales force and distributors to cease dissemination of promotional materials, literature and communications.Note: The firm revised indications, instructions for use, product labels and labeling so it is aligned with the 510Ks. The customers were also instructed to replace current instructions for use with revised instructions for use and destroy the current IFU; return Kings LT(S)-D oropharyngeal airway product inventory to King Systems; and complete and return the attached Instructions for Use Destruction Signature Form via fax to King Systems at 317-776-5197.For further information, please call 317-776-6823 ext 257 or email: LTSD@kingsystem.com. |  |  |  | Worldwide distribution: USA and countries including: Canada, Costa Rica, and Panama. | Terminated | August 16, 2012 | 975 |  | Not_Computer
Z-0973-2011 | 53198 | No | LPB |  |  | Cardiovascular | PMA | RF Ablation System Foot Switch, used with the EPT-1000 RF Generator, Foot Switch Model Number UPN M0048400, and in kits with UPNs M0044810 and M0044811. |  |  | July 14, 2009 | 2009 | Boston Scientific Corp | Product malfunction causes inadvertent delivery of energy. | Nonconforming Material/Component | The firm, Boston Scientific, sent an "URGENT: Medical Device Recall" letter dated January 14, 2009, with attached recall instructions to all customers. The letter described the product, problem, and the action to be taken by the customers. The customers were instructed to complete and return the Reply Verification Tracking Form and Obtain an RGA via fax to 408-957-6204; and return the recalled product using BSC Federal Express Number 1-809-4713-2 via second day delivery to Boston Scientific Corporation, Distribution Center, 500 Commander Shea Blvd., Quincy, MA 02171. Boston Scientific Corporation sales representatives will contact each customer.If you have any questions, please contact 408-935-4952 or email: wareb@bsci.com. |  | 1044 | 1044 units | Worldwide distribution | Terminated | January 27, 2011 | 562 | 1044 | Not_Computer
Z-0987-2011 | 57501 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Physiological Monitoring, Patient Monitor |  |  | December 17, 2010 | 2010 | Philips Healthcare Inc. | Some PAS-201 active speaker assemblies in use with the Philips IntelliVue Information Center (PIIC) have experienced intermittent audio or loss of audio. Speaker failures may delay recognition of an alarm condition, if operators do not recognize that there is no alarm sound from the monitor and the monitor is not connected to another audible alarm annunciation system. | Device Design | Philips notified all affected customers with an Urgent Medical Device Correction notification letter dated December 17, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to observe their system for any of the described symptoms. If they experience intermittent audio or there is no sound from their PAS-210 speaker while it is not connected to another audible alarm annunciation system, customers were instructed to remove the speaker from use and contact their local Philips service representative. Customers were advised not to rely exclusively on the audible alarm system for patient monitoring, the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.ACTIONS PLANNED BY PHILIPSPhilips will replace affected speaker assemblies free of charge.A Philips Healthcare representative would contact the customer when the replacement was available. Field Service Engineers will remove and replace defective speakers to correct the problem. The defective speakers will be returned to vendor (Beach Wire) to be discarded. This correction will be done free of charge for customers. |  | 1071 | 1071 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, MT, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA WI, WV, WY, and Canada | Terminated | April 18, 2017 | 2314 | 1071 | I/O | Alarm/Message | Replace components | Remove or Replace
Z-0970-2011 | 57703 | [6 empty] | RLSense, Restless Legs Sensor, Proto6, Rev-J Firmware |  |  | September 15, 2010 | 2010 | Sleep Modus | Device had no 510(k) clearance prior to patient use. | Other | Sleep Modus, Inc. notified consignees by e-mail on 9/15/2010 and 10/7/2010. The e-mail identified the product, the problem, and the action to be taken. Customers were instructed to let them know if they experienced any of the issues identified in the e-mail. For questions call 214-405-4055. |  | 110 | 110 devices (2 per kit) | Nationwide Distribution including Texas | Terminated | January 31, 2011 | 138 | 110 | Not_Computer
Z-0972-2011 | 57051 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Bio Med Devices Disposable Pediatric Breathing Circuit, One Set |  |  | October 19, 2010 | 2010 | Biomed Devices | Breathing circuit may disconnect between the adapter adjacent tubing cuff | Component change control | The firm,Bio Med Devices, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 18, 2010 via fax to the customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to immediately examine their stock for the affected lots, if found, discontinue distributing and using the product and promptly return via UPS #025106 to Bio-Med Devices 61 Soundview Rd, Guilford, CT 06437, Attention: 4040LC Recall Return. The customers were also instructed to immediately contact their accounts, if the product was further distributed , advise them of the recall situation and have their outstanding product returned, and to complete and return the RETURN FAX-IMMEDIATE RESPONSE REQUIRED form via fax to 203-458-0440. Note: Bio-Med informed the customers that they will credit the return product; however, they only plan to replace the product in an emergency situation, because they have very few of the 3/8" x 15mm adapters left and are looking to discontinue the 4040LC due to the discontinuance of this product by the manufacturer.If you have any questions, please contact (203) 458-0202 x 217. |  | 15 | 15 cases (20/cs) | Nationwide distribution: OH, WI. | Terminated | October 23, 2014 | 1465 | 16 | Hardware |  | Return | Remove or Replace
Z-0991-2011 | 57522 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | EndoPlege Sinus Catheter Kit, REF EP, 9 Fr x 18.9", Contains Phthalates, Sterile EO, Sterile R, Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | December 03, 2010 | 2010 | Edwards Lifesciences Llc | Defective product may result in the silicone valve becoming dislodged from the introducer which can result in injury. | Device Design | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated December 3, 2010 to customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to check their inventory and return the unused devices; complete and return the INTROCSC Introducer Recall Response form to the attention of Customer Service, Edwards LifeSciences, One Edwards Way, Irvine, CA at 949-250-3489. If you have any questions, please contact Edwards Customer Service at 800-424-3278 from 8:00am - 4:30pm Pacific Time. |  | 3427 | 3,427 kits | Nationwide Distribution | Terminated | October 27, 2011 | 328 | 3517 | Not_Computer
Z-0941-2011 | 57540 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Prelude Short Sheath Introducer, 7F-SMT, 4 cm, REF No: PSS-7F-4-038MT, Sterile EO, Merit Medical, South Jordan, Utah. |  |  | December 17, 2010 | 2010 | Merit Medical Systems, Inc. | Introducer tips may detach and embolize during use, or the tip material may elicit a thrombogenic response. | Component change control | Merit Medical Systems, Inc., issued an "Urgent Product Recall Notice" dated December 17, 2010, and contacted all customers (sales reps and authorized distributors) via email, phone and/or fax. The notice described the product, problem and action to be taken by the customers (Sales reps/Distributors). The customers were instructed to contact all their customers to identify, quarantine and return all unused inventory and extend the recall to any of their customers who may have received the affected products. The customers were instructed to complete and return the Manufacturer RECALL NOTIFICATION and Product Retrieval Form via fax to 1-801-208-3378 and send original to: the Manager, Customer Service, Merit Medical Systems, Inc., 1600 West Merit Parkway, South Jordan, UT 84095-2416.Questions should be directed to the Senior Product Manager at 801-208-4344 or the Manager, Customer Service at 801-208-4365. |  | 378 | 378 units | Nationwide Distribution | Terminated | July 11, 2011 | 206 | 378 | Not_Computer
Z-1001-2011 | 56354 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C Series ClinicÂżs (Includes Trilogy, and Novalis Tx), Model Numbers: H14, H26, H27, H29, Distributed by and/or Manufactured by: Varian Medical Systems Inc., Palo Alto, CA |  |  | December 01, 2009 | 2009 | Varian Medical Systems Oncology Systems | A circuit design anomaly whereby a faulty balance potentiometer can lead to an undetected electron beam asymmetry. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 1, 2009 via Federal Express to consignees/customers. The letter described the product, problem and actions to be taken by the customers and the firm. The customers were instructed to perform a verification of field symmetry on electron beam energies to determine the threshold and report any field symmetry issues to their local Varian service office for attention. The firm, Varian, is developing new functionality to detect the symptoms of this component failure condition under all conditions of electron beam operations and asserting a beam interlock if it is detected. Note: Varian indicated until the solution can be developed, they will conduct an audit of installed Clinacs to identify those the suspect "blue" electron beam balance potentiometers installed, and will execute a program of prophylactic replacement of these potentiometers.If you require further clarification, please contact your local Varian Customer Support District or Regional Manager at phone: USA and Canada-1-888-VARIANS (888-827-4265); Europe- +41 41 749-8844 or email support-emea@varian.com; North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdes@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems - www.myvarian.com; Varian Medical Systems -www.varian.com |  | 4411 | 4411 | USA and countries including: Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bolivia, Brazil, Canada, Chile,, China, Columbia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France,French Polynesia, Germany, Guadeloupe, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Latvia, Lebanan, Luxembourg, Maca, Macedonia, Malaysia, Martinique, Mexico, Moldova, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan,, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea,, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay.. | Terminated | March 14, 2012 | 834 | 4411 | Hardware | Device Operation | Instructions | Safety Notice/Insructions
Z-0988-2011 | 56963 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Southmedic/Personna Safety Scalpel with Personna Plus Blades, Reorder No. 73-0610, Size #10, General Surgical Use, Sterile, unless package is damaged or opened., Southmedic Inc., 50 Alliance Blvd., Barrie, ON., Canada L4M 5K3 www.southmedic.com, Made in Canada. |  |  | September 03, 2010 | 2010 | Southmedic, Inc. | Scalpels labeled as sterile have the potential to be unsealed and therefore not sterile. | Packaging process control | On 9/3/2010, the firm notified consignees about the recall by facsimile letter. The letter instructed consignees to return scalpels for credit or replacement. The firm enlarged the scope of the recall on 11/15/2010. An additional 927 cases were recalled. This includes 11 additional lots. The firm sent out URGENT MEDICAL DEVICE RECALL letters dated 11/15/2010 to its direct accounts. The letter identified that affected product and the reason for the recall. The consignees were instructed to check their inventory, complete and return the attached Voluntary Recall Notification Confirmation Form, return the affected product to the manufacturer. Product return question should be directed towards Marc Pauze at 800-463-7146 ext 299 and regulatory questions should be directed towards Tish Anger at 800-463-7146 ext 342. |  | 1330 | 1330 all products | Worldwide Distribution -- USA, including states of TX, MO, NJ, LA, IL, NY, KY, PA, VA, WA, CA, MA, MD, IN, NV, TN, NC, AZ, IA, WI, UT, NE, GA, FL, MN, ME, OH, SD, and KS and countries of Australia and Canada. | Terminated | May 18, 2011 | 257 | 1330 | Not_Computer
Z-1005-2011 | 53780 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 10032746 Acuson Antares PE Product Version 5.0, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | October 02, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | ECG waveform does not accurately align with 2D, M-Mode or Spectral Images. | Nonconforming Material/Component | The firm's Customer Safety Advisory notice was mailed to all affected customers. The letter describes when the malfunction occurs and the potential risk. It also recommended that users collect corresponding evidence with spectral waveforms that would confirm or deny findings based on ECG, in order to avoid potential risks for misdiagnosis. The firm planned to exchange the faulty hardware component in the system. Customers would be contacted to schedule for the hardware exchange. Questions regarding this issue should be directed to the local service support person. |  | 67 | 67 units total | Worldwide Distribution | Terminated | March 29, 2011 | 543 | 67 | Hardware | Output/Calculation | Replace component | Remove or Replace
Z-1008-2011 | 54535 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System, with MultiPlan Treatment Planning System Software version 3.5 |  |  | December 29, 2009 | 2009 | Accuray Inc | If electron density values are left empty, the calculation of radiation dose in a patient will be modeled as air-like density material rather than correct density. A plan may be created and saved, thus creating the risk of mistreatment. | Employee error | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated December 29, 2009, to all customers. The letter described the product, problem, and action to be taken by the customers. The customers were instructed to verify correctness of the CT number to relative electron density and Mass Density Calibration curves. Note: A patch is being developed to address this issue. Accuray Customer Support will contact the customers regarding further resolution.If you have any questions or concerns regarding this issue, please contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non USA) or by email: customersupport@accuray.com. |  | 71 | 71 units are subject to this action | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Italy, France, the Netherlands, the UK, India, Spain, Greece, Saudi Arabia, Switzerland, Germany and China. | Terminated | February 11, 2011 | 409 | 71 | Software | Output/Calculation | Software update | Software Update
Z-1002-2011 | 56357 | No | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stryker Endoscopy, Infravision Esophageal Kit, |  |  | May 11, 2010 | 2010 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter upon removal of the e-kit from the body. | Device Design | On 5/11/10, Stryker Endoscopy initiated the recall notification letters explaining the reason for the recall and requesting the product be returned. |  | 4125 | 4125 | Nationwide and 3 foreign consignees (Canada, Singapore and Switzerland). | Terminated | February 01, 2011 | 266 | 4125 | Not_Computer
Z-1003-2011 | 56349 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand C-Series Clinacs and Acuity, Exact Couch Top, Model Numbers: |  |  | April 19, 2010 | 2010 | Varian Medical Systems Oncology Systems | The product has been shipped since 2005 with the wrong label regarding max weight to be used on accessory rails. | Labeling mix-ups | Varian issued via customer technical bulletin on April 19, 2010, to Varian Sales, Marketing and Service organization. Notification. The bulletin identified the product, the problem, and the action to be taken by the customer. Customers were asked to fill out the attached form and send it to their local Varial Medical Systems Field Service Office.For questions regarding this recall call 650-483-3153. |  | 13 | 13 | Worldwide Distribution - Japan only | Terminated | December 23, 2011 | 613 | 13 | Not_Computer
Z-0940-2011 | 56868 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Maquet Heartstring II Proximal Seal System (HS-1045) manufactured by Maquet Cardiovascular, Wayne, NJ |  |  | May 05, 2010 | 2010 | Maquet Cardiovascular, LLC | Deployment tube may detach during use due to insufficient adhesive, which may result in product failure to operate. | Process design | The firm, MAQUET Getinge Group, sent an "URGENT DEVICE REMOVAL IMMEDIATE ACTION REQUIRED" letter dated May 5, 2010, to all customers via Federal Express. The letter described the product, problem and actions to be taken by the customer. The customers were instructed to immediately examine their stocks to determine if they have any of the affected product, if found, discontinue dispensing (distributing) the lot and complete the Field Action Response Form; immediately contact accounts if they acted as a distributor, and acknowledge receipt of this notification whether or not they have any affected products by completing the Field Action Response Form via fax to (U.S.): 1-888-899-2874.If you have any additional questions, please contact your local MAQUET Cardiovascular sales representative or our Customer Service at 1-888-880-2874. |  | 1630 | 1630 | Worldwide distribution: USA and countries including: Japan, Mexico, Germany, Switzerland, Israel | Terminated | June 21, 2011 | 412 | 1630 | Not_Computer
Z-1004-2011 | 56759 | No | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General &amp; Plastic Surgery | 510(k) | Optovue iVue optical coherence tomography system, computer controlled opthalmic imaging system, manufactured by Optovue, Fremont, CA |  |  | July 20, 2010 | 2010 | Optovue Inc. | Use of the CAM attachment may result in a positioning anomaly and mistreatment in the eyes of the patient. | Software design (manufacturing process) | Optovue Inc. sent a Field Modification/Correction Notice dated July 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that they could continue to use the iVue until the correction has been made. Customers were instructed to review any existing iVue patient data to ensure that this situation and an improper designation of eye assignment has not occurred. Upon completion of the updates, customers were to sign the attached form indicating they acknowledge that their iVue has been updated and they understand that they must review existing data. For any questions regarding this recall call 1-866-344-8948 or by mail at 45531 Northport Loop W, Fremont, CA 94538. |  | 22 | 22 devices in the US. | Nationwide | Terminated | February 08, 2011 | 203 | 22 | Not_Computer
Z-1009-2011 | 57576 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AutoSuture SURGIPRO Polyproylene Mesh Clear 3" x 5" Sterile |  |  | December 15, 2010 | 2010 | Covidien LP | Sterility compromised due to breach in sterile barrier | Packaging process control | Covidien notified customers via letter dated November 15, 2010. The letter identified the product, the problem, and action to be taken by the customer. Customers were instructed to examine inventory and return the affected lot to Field Returns Department, 195 McDermott Road, North Haven, CT 06473. For questions regarding this recall call 203-492-8091. |  | 132 | 132 units | Worldwide Distribution - USA including CA, CT, FL, HI, IL, IN, KS, MA, MD, MI, MN, NC, NJ, NY, OH, PA, SC, TX, and WI and the countries of Canada and Italy | Terminated | January 10, 2013 | 757 | 132 | Not_Computer
Z-1010-2011 | 57595 | No | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Fraxel re:storeÂż Dual Laser System----a laser surgical instrument for use in general and plastic surgery and in dermatology. |  |  | December 15, 2010 | 2010 | Solta Medical, Inc. | The product has the potential to cause inadvertent laser firing which may lead to patient/operator injury if the problem recurs. | Software design | The firm, Solta Medical, issued three "DEVICE CORRECTION" letters, one dated December 15, 2010, the second one date December 17, 2010 and the third one dated December 20, 2010, to all affected customers. The letters described the product, problem and action to be taken by the customers. The customers were instructed to follow the instructions carefully, sign and date the bottom of the instruction form upon completion of the update and promptly return the signed form and USB stick in the enclosed, return-addressed envelope (Product Support, Solta Medical, 25881 Industrial Blvd., Hayward, CA 94545) Note: Solta Medical sales personnel will assist with the software update process so, if your local representative has already updated your system, you do not need to do anything further.If you have any questions, please call Solta Product Support at (510) 259-7291. |  | 352 | 352 units | Worldwide Distribution: USA and countries including: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Egypt, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Kuwait, Mexico, Netherlands, New Zealand, Pakistan, Philippines, Poland, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, United Kingdom, and Vietnam | Terminated | November 14, 2011 | 334 | 352 | Software | Physical Safety Hazards | Software update | Software Update
Z-1017-2011 | 57661 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Port Dressing Tray, Reorder DT12885, Sterile, Centurion Medical Products Corp. Williamston, MI. |  |  | December 06, 2010 | 2010 | Centurion Medical Products | The firm is conducting a recall of kits that contain the MONOJECT PREFILL ADVANCED flush syringes containing heparin. The syringes are under recall by their manufacturer Coviden because of testing that showed "that the crude heparin used to produce a specific batch of purified heparin sodium USP contained low level of Over-Sulfated Chondroitin Sulfate (OSCS)". | Other | The firm, Covidien, sent an "URGENT PRODUCT RECALL" letter dated November 12, 2010, to affected customers via Certified Mail, Return Receipt Requested. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory and in-use stock to determine if they have any units of the affected product, if found, immediately remove from inventory and quarantine; if they are a distributor they were ask to notify their customers of this letter; and complete and return the PRODUCT RECALL FORM whether or not they have the affected product via fax to: 203-822-6009 or scan and email to: sdfeedback@covidien.com. Note: The returned product will be collected and counted prior to destruction or return to the vendor.Should you have any medical questions concerning this recall, please call 508-261-6393.For Customer Service questions regarding the recall, please use one of the following methods: phone 1-800-962-9888 option 1 (8am to 6:30pm Eastern); email: sdfeedback@covidien.com or fax: 302-822-6009. |  | 2250 | 2250 kits. | Nationwide Distribution: USA including states of: OH, WI, IL, MO and TX. | Terminated | February 16, 2011 | 72 | 8580 | Not_Computer
Z-1012-2011 | 57259 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow WH-02220 - Epidural Catheterization Kit with FlexTip Plus(R) Catheter |  |  | November 11, 2010 | 2010 | Arrow International Inc | Sterility may be compromised. | Employee error | The firm, Arrow International, issued an "URGENT MEDICAL DEVICE RECALL" letter dated November 11, 2010 to their consignees/customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to check their inventory for the affected products, cease use and distribution, quarantine all affected product immediately, return any affected product freight collect along with the completed form via mail to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, Attn: Asheboro Quality Assurance Manager, and complete and return the FIELD SAFETY CORRECTIVE ACTION ACKNOWLEDGEMENT AND STOCK STATUS FORM via fax to: 800-343-2935.If you require additional information or clarification regarding this matter, please contact your sales representative, or Arrow's Customer Service Department at 866-396-2111. |  | 2320 | 2320 | Nationwide Distribution: USA states including IL and LA. | Terminated | July 05, 2012 | 602 | 2320 | Not_Computer
Z-1011-2011 | 57275 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Blood Analyzer Part Number: 393-090. |  |  | November 01, 2010 | 2010 | Radiometer America Inc | Due to a software error, a patient mix up can occur on the ABL90. If a sample is pre-registered and the sample "fails" the next result will inherit the patient data from the previous result. | Software design | On 11/16/10 the firm sent Medical Device Field Correction letters to their customers. The letter states that all systems will be upgraded and until each system is upgraded, they instructed their customer to "always query for patient ID prior to introducing the sample." They also instruct each customer to complete the "Fax Return Form" acknowledging receipt of the letter. If customers have questions regarding the letter or the upgrade, they should contact Technical Support at 1-800-736-0600 Opt 4. |  | 432 | 432 units | Worldwide Distribution -- USA, including states of MI, NJ, &amp; PA and countries Australia, Austria, Belgium, Denmark, Finland, France, Germany, Indonesia, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, Poland, Russia, Spain, Sweden, Switzerland, United Arab Emirates &amp; United Kingdom. | Terminated | May 09, 2012 | 555 | 432 | Software | Output/Calculation | Software update | Software Update
Z-1027-2011 | 52903 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER component of MOSAIQ versions 1.30 (all builds), 1.40 (all builds), 1.50 (all builds), 1.60 (all builds), and 2.0 (all builds). Manufactured by IMPAC Medical Systems, Inc. Sunnyvale, CA. |  |  | July 17, 2009 | 2009 | Impac Medical Systems Inc | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion), and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Software design | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to:European customers: +44 1293 654401 American &amp; rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or callEuropean customers: 44 1293 654320American &amp; rest of world customers: 800-488-4672. |  |  |  | Worldwide Distribution - USA including AL, CA, CT, FL, GA, IA, IL, KY, MI, MN, MS, ND, NJ, OH, OR, PA,TN, TX, VA, WA, WI, and WV and the country of Canada. | Terminated | November 25, 2011 | 861 |  | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-1367-2011 | 53012 | Yes | NIQ |  |  | Cardiovascular | PMA | Endeavor Resolute Zotarolimus-Eluting Coronary Stent System with Rapid Exchange Delivery System, UPN 00613994400505, model #IDERES25024X. Manufacturer: Medtronic, Inc. Minneapolis, MN. Manufactured in: Medtronic Ireland, Galway, Ireland. Product is packaged inside a sterile, inner pouch, labeled in part "Use immediately". Inner pouch is supposed to be within a foil pouch containing desiccant, and moisture absorber. Foil pouch label reads in part "RX Endeavor Resolute Zotarolimus-Eluting Coronary Stent System...2.5 mm X 24 mm...REF IDERES25024X...Contents: One (1) Endeavor Resolute zotarolimus-eluting stent mounted on a rapid exchange stent delivery system...Guide Catheter/Minimum Inner Diameter &gt;=5F/0.056" (1.42 mm)...Maximum Guidewire Diameter 0.014" (0.36 mm)...Maximum Stent I.D. 3.00 mm (0.12"...Nominal Pressure 9 atm (912 kPa)...Rated Burst Pressure 15 atm (1520 kPa)...MR Conditional...Sterilized using ethylene oxide Consult instructions for use...Manufacturer: Medtronic, Inc. 710 Medtronic Pkwy NE Minneapolis, MN 55432 USA...Manufactured In: Medtronic Ireland Parkmore Business Park West, Galway, Ireland Tel:+353-91-708000 Fax:+353-91-757524 Investigational Devices / Returned Goods: Tel: (800) 556-4247 / Fax: (877) 518-2469...BioLinx Medtronic Polymer System". Foil pouch contents and Instructions for Use are placed within an outer box, which is labeled identically to the foil pouch. |  |  | August 06, 2009 | 2009 | Medtronic Vascular Galway Limited | The product was removed from investigational sites due to a packaging issue, in that the device units were packed without the foil pouch, oxygen absorber and Instructions For Use. | Packaging | The firm, Medtronic, sent their initial "Urgent Medical Device Recall Notice" to customers with affected product via e-mail on August 7, 2009. An additional "Urgent Medical Device Recall Notice" was sent to all customers on August 12-13, 2009 via FedEX and UPS. The notice described the product, problem and action to be taken by the customers. The customers were instructed to remove the affected units from their inventory and isolate the units in a secure location to prevent additional use. A Medtronic CardioVascular Clinical Customer Service (CCS) representative will contact their site to arrange for return of the units. This will be followed by a shipment of replacement units to their sites. If there were any questions regarding the return process, customers were to contact CCS at 800-556-4247.If you have any further questions or concerns regarding this notice, please contact the Sr. Clinical Research Manager at 707-543-2284 or ann.depaoli@medtronic.com. |  | 49 | 49 units | Nationwide distribution: AL, AR, CA, FL, IN, MD, MI, NY, OH, PA, RI, TX, VA, and VT. | Terminated | February 17, 2011 | 560 | 49 | Not_Computer
Z-1022-2011 | 53014 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PEAK brand PlasmaBlade EXT, Tissue Dissection Device, |  |  | May 19, 2009 | 2009 | Peak Surgical | A customer reported that a patient suffered a burn while using the device. | Other | Peak Surgical notified afftected customers by phone on May 19, 2009. They were notified of the product, the problem, and the action to be taken. Peak Surgical sales representatives removed PlasmaBlade EXT devices from consignees affected by the recall. |  | 493 | 493 units | Worldwide Distribution - United States including CA, CO, FL, IL, MD, MN, NV, NY, OH, OK, TN, TX, UT, and WAand the countries of Turkey and Ireland | Terminated | April 21, 2011 | 702 | 493 | Not_Computer
Z-1021-2011 | 56923 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital Connect 4.0 &amp; 4.1, radiological image processing software |  |  | August 23, 2010 | 2010 | Vital Images, Inc. | There is a potential alignment error of fused PET/CT images in the Vitrea Core component of Vitrea Enterprise Suite 1.2 and 1.3. The error occurs during panning of the fused PET/CT images. Because the PET data has a larger voxel size than the CT data, it is possible to pan the CT image in more frequent, smaller steps than the PET image, so that the fused PET image does not move until the CT data | Software design | The firm, Vital Images, Inc,., sent a "Urgent Software Alert" letter dated August 23, 2010 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to not use the PET/CT fusion functionality until the software updates are provided and to re-examine any past results. If you have any questions or comments, contact Customer Support at support@vitalimages.com or at 1-800-208-3005. |  | 178 | 178 | Worldwide Distribution: USA including states of: AK, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, PA, SC, TN, TX VA, WV, WI, and WY; and countries of: AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, and ITALY.AUSTRALIA, HONG KONG, UNITED KINGDOM, NETHERLANDS, FRANCE, BELGIUM, CANADA, CHINA, AUSTRALIA, SWITZERLAND, ITALY. | Terminated | November 06, 2012 | 806 | 178 | Software | Display/Image | Software update | Software Update
Z-1023-2011 | 54658 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation therapy device. Medical charged particle radiation therapy system with software version. Model Numbers: Multi Plan Treatment Planning System, Software version 3.0, 3.1, 3.5.0, and 3.5.1. |  |  | January 15, 2010 | 2010 | Accuray Inc | When a user modified the align center of an existing plan during the treatment planning process the effective depths of each beam in the plan were not updated by the system. These depths were incorrect and led to an incorrect dose calculation. Software patch /update to address this issue. Implementation at affected sites will initiate, once the patch is available. | Software design | The firm, Accuray, sent an "URGENT DEVICE CORRECTION" letter dated January 18, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to load the plan and recalculate the dose to reflect the current state of beam, compare the reloaded plan and the recalculated plan to verify maximum dose and prescription values, and re-prescribe and save to overwrite to existing plan if necessary. The firm is developing a patch/software update to address the issue and a Accuray Customer Support person will contact the customer regarding further resolutions.If you have any questions or concerns regarding this issue, contact Accuray Customer Support at +1-877-668-8667(USA) or +1-408-716-4700 (non USA) or email: customersupport@accuray.com. |  | 99 | 99 units | Worldwide Distribution: USA and countries including: Canada, Turkey, India, Russia, Italy, France, Belgium, Netherlands the UK, Spain, Athens, Saudi Arabia, Switzerland, the Ukraine, Germany, Korea and Hong Kong. | Terminated | February 11, 2011 | 392 | 99 | Software | Output/Calculation | Software update | Software Update
Z-1025-2011 | 57555 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb, Renu Fresh Multi-purpose Solution, 2 fl oz (60 mL), Travel Kit, Manufactured by: Bausch &amp; Lomb Place, Rochester, NY 14609 |  |  | December 22, 2010 | 2010 | Bausch &amp; Lomb Inc | There is a potential for some packages to not meet sterility requirements. | Other | On December 8, 2010, Target was visited by FDA/PHI-DO and was informed on the violation. The Target customers that received lot GH0078 were requested to return all affected product directly to the Bausch &amp; Lomb Greenville, SC Distribution Center. Bausch and Lomb sent an Urgent Voluntary Medical Device Recall to all first consignees who were shipped the affected lot GH0066 of Renu Fresh Multipurpose Solution on/about December 22, 2010. Since Target customers had already been notified previously regarding lot GH0078, this lot number was not included in the subsequent recall notification. All consignees were requested to review their inventory, and return any of the affected product to Stericycle Pharmaceutical Services, 2670 Executive Drive, Indianapolis, IN 46241. Stericycle is conducting this recall under the control of Bausch &amp; Lomb.For questions abut the recall process contact Stericycle, Inc. at 1-888-345-8316.For questions regarding this recall call 1-800-828-6974. |  | 292992 | 146,496/ 2 oz Bottles | Worldwide Distribution - USA including AL, CA, CT, FL, GA, HI, IA, IN, LA, MA, MI, NE, NJ, NY, OH, PA, RI, SC, TN, TX, UT, VA, and WI and the countries Taiwan and Brazil | Terminated | September 04, 2012 | 622 | 360288 | Not_Computer
Z-1024-2011 | 57464 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Laryngoscope Handle, Reusable, Rx, Product Code 4558GSP or 4559LED, Vital Signs Colorado, Inc., 11039 E. Lansing Circle, Englewood, CO 80112. |  |  | December 20, 2010 | 2010 | Vital Signs Colorado Inc. | Battery which powers the laryngeal scope light may dislodge resulting in loss of illumination. | Device Design | Healthcare Systems sent an URGENT MEDICAL DEVICE CORRECTION letter dated December 20, 2010, to all affected consignees advising them of the possible problem and giving them interim safety instructions until new battery cartridges were available. Customers were instructed to complete and return the Confirmation Form which was included with the letter.For information regarding this recall call 973-956-5300. |  | 2202 | 2202 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, MA, MN, MO, NH, NJ, NY, OH, OK, SC, TN, TX, WA, and WV and the countries of Australia, Belgium, Canada, Dominican Republic, Ecuador, Germany, Hong Kong, Latvia, New Zealand, Taiwan, Turkey, and the UK | Terminated | July 01, 2011 | 193 | 2202 | Not_Computer
Z-1020-2011 | 56392 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | i-STAT Level 2 Control Value Assignment Sheets |  |  | August 03, 2010 | 2010 | Abbott Point of Care Inc. | Three i-STAT Level 2 Control Value Assignment Sheets and the i-STAT CLEW A20 Software System Manual Update CD contain the incorrect range for pCO2, BUN and Glucose for the i-STAT EC8+, 6+, EC4+, E3 and G cartridge types. | Other | Abbott Point of Care Inc., sent an URGENT RECALL NOTICE to all affected customers on 8/3/2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to replace all copies of the value sheets with the ones provided in the letter. ASLO, if they have a disc with the Value sheets on it, it is to be replaced with the updated disc and the original discarded. If the control materials, discs or value sheets have been forwarded to another facility, a copy of the letter was to be sent to them. The Business Reply card is to be completed and returned to Abbott. Questions or comments are to be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1. |  | 8293 | 8293 Value sheets worldwide | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of Argentina, Austria, Belgium, Bolivia, Brazil, Chile, China, Columbia, El Salvador, France, Georgia Republic, Germany, Hong Kong, Israel, Italy, Kuwait, Mexico, Netherlands, New Zealand, Panama, Paraquay, Peru, Poland, South Africa, Sweden, Switzerland, Taiwan, Turkey, UAE, UK, Uruquay, and Yemen. | Terminated | September 28, 2011 | 421 | 8293 | Other |  | Manual update | Other
Z-1030-2011 | 57570 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table |  |  | March 10, 2008 | 2008 | Elekta, Inc. | Possible error in the table longitudinal position. | Other | Elekta sent an updated IMPORTANT NOTICE letter dated December 24, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Elekta recommended that all users, and Eleckta field service engineers do regular QA checks to compare the movement of the treatment table in a given axis with the digital readout. Elekta Service representatives would attach the screws with Loctite and calibrate the axis on these sites. For questions regarding this recall call 770-670-2548. |  | 597 | 597 units | Nationwide Distribution and Puerto Rico | Terminated | August 09, 2012 | 1613 | 597 | Not_Computer
Z-1031-2011 | 57571 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Presice Cryosurgical System, model FPRCH2047, and the following system is not approved in the US: CryoHit-P Cryoblation System, model FPRCH2058. The system is a software controlled cryoblation system designed to destroy targeted tissue using a minimally invasive percutaneous approach. Freezing and thawing of the tissue is achieved through the use of the Joule-Thompson effect, using a high-pressure gas distributed through closed-end needles placed into the patient;s target organ. |  |  | December 20, 2010 | 2010 | Galil Medical, Inc. | This voluntary correction has been initiated due to the discovery of a damaged optocoupler on one of our Presice Systems during periodic preventive maintenance. The optocoupler sends a signal across the locking mechanism in the Presice and Seednet MRI Systems, enabling the software to recognize if the channel on the system manifold is locked or unlocked. If an optocoupler is damaged or malfunction | Device Design | Galil Medical sent an "Urgent Medical Device Correction" letter dated December, 30, 2010, to all affected consignees. The letter described the product, the problem involved, and the action to be taken by the consignee. Consignees were instructed to make sure the locking mechanism on the Precise and Seednet MRI Systems were in the "locked" position when a needle is inserted into the manifold. Gently pull on the black plug to ensure needle is secure in locking mechanism. Apply the enclosed alert sticker on the bottom of the distribution panel and the bottom of the system monitor of the Precise System and on the side of the mobile unit itnerface panel and the top of the system monitor of the Seednet MRI System. Consignees were asked to take a picture of the alert sticker after it has been applied to the appropriate location on their Presice or Seednet MRI System(s) and e-mail to yoni.kadosh@galilmedical.com. Consignees were instructed to complete and return the enclosed form as soon as possible.For questions regarding this recall call Customer Service at (877) 639-2796.problem involved. |  | 63 | 63 | Worldwide Distribution - USA including CA, CT, FL, GA, MA, MD, MN, MO, NC, NY, OH, PA, SC, TN, TX, VA, and WA and the countries of CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, and JAPAN.CANADA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, ISRAEL, ITALY, KOREA, NETHERLANDS, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, CHINA, JAPAN. | Terminated | March 27, 2012 | 463 | 85 | Computer
Z-1033-2011 | 57650 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, CADDÂż-Solis Ambulatory Infusion Pump, Model Number: 2100 and 2110, Reorder Number 21-2101-00, 21-2101-0100-50, 21-2101-0100-51, 21-2101-0200-02, 21-2101-0200-03, 21-2111-0100-51, 21-2111-0200-02, and 21-2102-51. |  |  | January 13, 2011 | 2011 | Smiths Medical ASD, Inc. | Reports of the Cassette Lock mechanism on certain CADDÂż-Solis Pumps becoming stuck in the locked position, making it nearly impossible to remove the medication cassette reservoirs (or administration set). The purpose of the Cassette Lock is to secure the medication cassette reservoir (or administration set) to the pump. | Nonconforming Material/Component | The firm, Smiths Medical, sent an "URGENT MEDICAL DEVICE CORRECTION NOTICE" dated January 13, 2011 to all consignees/customers. The notice described the product, problem, and action to be taken by the customers. The customers were instructed to have their Pumps undergo a correction which consist of removal and reattachment of the Cassette Latch handle using a new adhesive process. The customers were also instructed to remove the Pump from service if unable to remove cassette (or administration set); to inspect their inventory and identify Pumps; prepare a plan for correction of these Pumps; complete and return the attached CONFIRMATION FORM via fax to 1-800-237-8033 or by email to CADDDSolis_CassetteLock@smiths-medical.com, and if they have distributed the affected product to other persons or facilities, promptly forward this notice to the recipients. If you should have any questions regarding this information, please contact Smith Medical at 1-866-216-8806. |  | 2903 | 2903 (2539 US; 364 OUS) | Worldwide Distribution: USA including states of: AK, AZ, C, CO, DC, FL, GA, ID IL, MA, MI, MN, MO, MT, NE, NH, NM, NY, NC, OK, OR, PA, RI, TN, TX, UT, VT, VA, WA, WI, and WY; and countries including: Australia, Austria, Belgium, Canada, France, Germany, Great Britain, Ireland, Italy, Luxembourg, New Zealand, Qatar, Saudi Arabia, Switzerland and The Netherlands. | Terminated | April 17, 2012 | 460 | 2903 | Not_Computer
Z-1038-2011 | 57721 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Irrigation Kit, Large, Vinyl, REF: 6114, Apogee Medical, LLC, 90 Weathers Street, Bournesville, NC 27596. |  |  | January 18, 2011 | 2011 | Apogee Medical | This recall has been initiated due to concerns expressed by the Food and Drug Administration regarding the validation of the gamma radiation sterilization cycles for these products of Sterile Lubricating Jelly manufactured by Triad Group. | Other | The firm, TRIAD Group, sent a "Urgent Medical Device Recall" letter dated January 18, 2010 to all consignees/customers. The letter described the product, problem and the action to be taken by the customers. The customers were instructed to review there inventory, cease distribution of the affected lots and to complete and return the enclosed Product Recall Response Form via fax to 1-919-570-9611 or e-mail to Apogee Medical. If they distributed the product, they were instructed to notify their customers of the recall; to include a copy of the recall notification, and to instruct them to return the recalled product to the wholesaler who would then return the product to Apogee Medical. Non-responders to the recall notice will be contacted on a weekly basis.Should you have any questions regarding this notification, please call the Regulatory Affairs/Quality Assurance Manager at (919) 435-5409. |  | 1 | see line 1. | Nationwide Distribution: USA including states of: AL, AZ, CA, FL, GA, IL, IN, KS, LA, MA, MO, NC, NJ, NY, OH, OK, PA, TN, TX, VA and WA. | Terminated | February 28, 2012 | 406 | 4454545 | Not_Computer
Z-1028-2011 | 57504 | No | BZQ | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | ThermiSense model 5700-A5 |  |  | December 21, 2010 | 2010 | Salter Laboratories, Division of Regulatory Affairs | The incorrect length and connector were utilized in lot 110810 of the ThermiSense Alice 5, making the product unusable as it will not interface with the sleep labs equipment. | Employee error | Salter Laboratories, Inc. sent an URGENT PRODUCT RECALL LETTER dated December 20, 2010, to all affected customers. The letter identfied the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for this product and lot number. Contact Salter to arrange for the return of the product. Complete the attached reply form, confirming receipt of the letter, and fax it back to Salter Labs at 1-800-628-4690.For questions regarding this recall call 661-854-3166. |  | 41 | 41 units | Nationwide Distribution including CA, CO, DE, ID, IN, KY, MD, MO, and OH, | Terminated | June 06, 2011 | 167 | 41 | Not_Computer
Z-0993-2011 | 57507 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arstasls One Access System; Models: AAD100, AAD101, Part Numbers: FG-02279, FG-03010; The system consists of the following components: (1) Device Sheath/Anchor, (2) Shaft |  |  | October 19, 2010 | 2010 | Arstasis, Inc | Device has the potential to fracture and sheath separation during use. | Device Design | The firm, Arstasis, initiated a market withdrawal and sent notification via letters dated October 19, 2010, to customers informing the customers that Arstasis is implementing a product improvement and will replace all existing inventory that the customer has on their shelf with the improved product. The customers were instructed to work closely with their local Arstasis territory manager to ensure that their product replacement is handled satisfactorily, and complete and return the PRODUCT REPLACEMENT and PRODUCT RETURN forms to Arstasis Customer Service via fax at 1-877-688-8417 or email to customerservice@arstasis.com.If you have any questions or concerns, contact Customer Service toll free at 1-877-594-4545. |  | 330 | 330 units | Worldwide Distribution | Terminated | February 16, 2011 | 120 | 330 | Not_Computer
Z-1029-2011 | 55087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER for Mosaiq interfaced to Electa SL75-5 machines. |  |  | March 10, 2010 | 2010 | Impac Medical Systems Inc | Computer interface is defected. Computer based verification system for recording patient treatment on a radiotherapy treatment machine fails to record treated field. | Software design | A User Notice was sent by the firm to its consignees. It identified the affected product and the reason for the recall. It also discussed the clinical impact and interim mitigations. The firm states in the letter that a fix for the problem and been developed and is available. Customers are to contact IMPAC Support for the patch to the SL75-5 interface. If customers have questions, concerns, or requests, they can email support@impac.com or one of the numbers listed in the letter for the applicable country. The letter should be distributed to all users affected by the issue at the customer's organization. |  | 40 | 40 | Worldwide Distribution -- US, Canada, Germany, and France. | Terminated | April 06, 2011 | 392 | 40 | Software | Device Operation | Software update | Software Update
Z-1042-2011 | 57549 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Vertex Max Drill Tube, Catalog Number 9733140, a component of the Navigated Vertex Max Drill Guide Kit, Catalog Number 9733370. Medtronic, Inc., Louisville, CO. |  |  | December 15, 2010 | 2010 | Medtronic Navigation, Inc | Incorrectly manufactured drill tube could cause incorrect screw placement during spinal surgery. | Component design/selection | Consignees were notified by a Product Correction Notification letter on 12/15/2010 and given the affected product and a description of the problem. The letter stated that a Medtronic representative would contact customers to visit their facility and replace the affected guide with one that has been correctly manufactured. Additional information is available at 800-595-9709. |  | 15 | 15 devices | Worldwide Distribution -- USA, including states of GA, IL, IN, MN, MO, OH, OR, PA, and TX and countries of Denmark, Finland, Singapore, Spain, and UK. | Terminated | May 04, 2011 | 140 | 15 | Not_Computer
Z-1019-2011 | 57558 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva. The marketing brochure is labeled in parts: "Philips Medical Systems is part of Royal Philips Electronics... www.medical.philips.com... Philips Medical Systems Nederland B.V". |  |  | December 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | There is a risk of selecting the wrong patient, when patient is selected in the patient list and the examination tab pressed with background query are running at the same time. | Software design | On 1/30/2009, Philips began sending out URGENT-Device Correction PCR Eleva notifications to their consignees.The consignees were informed of the hazards as when a patient is selected in patient list and examination tab pressed with background query running at the same time, there would be a risk that the wrong patient is selected.The firm issued the Field Corrective Order (FCO) 73200029 to their representative to correct the problem. The corrective action would be a software upgrade that resolve the issue.All consignees can contact Philips Call Center at 800-722-9377, #5, #2 and reference "FCO 73200029" with any questions. |  | 686 | 686 units | Nationwide Distribution -- USA, including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WI, WV, and WY and to LANDSTUHL ARMY BASE, APO in United Arab Emirates. | Terminated | August 25, 2011 | 975 | 686 | Software | Output/Calculation | Software update | Software Update
Z-1046-2011 | 56448 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnifit M/S PSL Porous Coated Acetabular Shell 40MM; |  |  | August 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics determined that specific lots of the Omnifit Acetabular Shells may have a screw hole wall thickness that falls above the design specification. | Device Design | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated August 2, 2010, via Federal Express to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory to identify the affected product; retrieve all affected product and return it their branch or agency warehouse for reconciliation; reconcile all product(s) on the attached Product Recall Acknowledgement Form and fax to (201) 831-6069, and return the affected product via Inter-Org Transfer and ship to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  |  |  | Worldwide distribution: USA including states of: AL, AR, CA, FL, IN, ME, MI, MS, NC, NJ, OK, OR, WA , and WI; and countries including: China, Hong Kong and Korea. | Terminated | November 09, 2012 | 830 |  | Not_Computer
Z-0943-2011 | 57193 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Discovery ST (16 Slice) System; Discovery ST (16 Slice) Systems with technical Reference Manual numbers 2351785-xxx and 2390935-xxx distributed from November 2001 to August 2010. |  |  | October 27, 2010 | 2010 | GE Healthcare, LLC | Although the system x-ray output is accurate for the technique factors given, and the dose display is also accurate, certain CTDI adjustment factors used for the manual calculation from the Technical Reference Manual (TRM) are incorrect. CTDI100 Aperture Adjustment Factors for 4x0.625 mode contain incorrect values. This issue may lead to inaccurate CTDI manual calculation. The results of the diffe | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be provided together with the TRM addendum informing the customers of the issue and related hazards. The letter also identified the affected product and gave instructions, which include referring to the TRM addendum, along with the product correction. This correction will be executed via a Field Modification Instruction IFMI 254281 released in November, 2010 with an estimated completion date of May, 2011. Questions or concerns should be directed to the call numbers listed or a local GE Healthcare Service Representative for countries not listed. |  | 1450 | 1450 units total in US | Worldwide Distribution | Terminated | July 13, 2012 | 625 | 1450 | Not_Computer
Z-1064-2011 | 53054 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Linear Medical Accelerator ARTISTE MV systems, when equipped with a syngo RT Therapist, Therapist Express Basic, syngo RT Therapist Assist or Therapist Express Assist workspace with software version NET VA10A for RTT 4.1.122 or NET V10B for RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA. |  |  | July 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Potential for Mistreatment-- when treatment plans intended for non-ARTISTE machinese are loaded, if the plans contain unspecified beam parameters, attributes or unknown block codes, the therapy system may remove those parameters. | Software design | The firm,Siemens, sent two Customer Advisory letters (TH004/09/S and TH008/09/S) contained in the Update Instructions to all consignees/customers with active devices. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow the updated instructions. Siemens service representative will sign a document noted as CORRECTED upon completion of the software installation.If you have any questions, call 1-925-602-8157. |  | 151 | 151 units | Worldwide distribution: USA and countries including: Germany, France, Poland, Norway, Australia, Malaysia, Spain, New Zealand, the UK, India, Republic of Korea, Canada, Ireland, Italy, Egypt, Belgium, Philippines, Saudi Arabia, Netherlands, Hungary, Norway and South Africa. | Terminated | April 12, 2011 | 621 | 151 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1063-2011 | 57531 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Inrad sample master aspiration biopsy needle with hiliter sharp bevel needle, 22 ga x 2 3/8'' (6 cm), model 41222, Inrad, Inc. Kentwood, MI. |  |  | December 20, 2010 | 2010 | Inrad, Inc | This lot of product has the wrong expiration date on the carton label. The pouches contained the correct expiration date. | Employee error | The firm, Inrad Inc., sent an "URGENT FIELD CORRECTION" notice dated December 20, 2010 to its customers. The notice described the product, problem and actions to be taken by the customers. The customers were instructed to immediately examine their inventory for the affected product. If the recalled product was in stock, the customers were told to removed the affected product from the carton and dispose of carton; to disregard carton label expiration date and reference the expiration date appearing on the pouch label only, and to complete, sign and return the attached "Carton Expiration Date" INFORMATION FORM via fax to 616-301-779 or mail to Inrad Inc., ATTN: Director RA/QA, 4375 Donker Court SE, Kentwood, MI 49512.If you have any additional questions about this action, please call 800-558-4647 extension 102. |  | 250 | 250 | Nationwide Distribution: USA including states of: FL, GA, KY, NH, NJ, NM, PA , SC, WA, and WI. | Terminated | February 09, 2011 | 51 | 250 | Not_Computer
Z-1045-2011 | 57610 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian High Energy Linear Accelerator: |  |  | December 10, 2010 | 2010 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was identified whereby the bolts used to fasten the counterweight to the gantry may not, in some cases, be fully tightened to the required torque specification. | Process control | An Urgent Medical Device Correction letter was sent to all affected users with a description of the problem and user corrective action steps. Letters were sent on 12/10/2010. Users are to immediately investigate any unusual noises coming from the affected product and report them to whoever is responsible for servicing the device. For customers that do have a service contract with Varian, Varian service will perform an inspection of the counterweight bolts as part of the regular Periodic Maintenance Inspection procedure. Customers are to advice the appropriate personnel of the contents of the letter. For further clarification, customers can contact their local Varian Customer Support Services Manager. |  | 1882 | 1882 | Worldwide Distribution. | Terminated | August 02, 2011 | 235 | 1882 | Not_Computer
Z-0998-2011 | 57524 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Diabetes Care, Precision Point of Care Blood Glucose Test Strips; |  |  | December 22, 2010 | 2010 | Abbott Diabetes Care, Inc. | Certain lots of Precision Xtra, Precision Xceed Pro, MediSense Optium, Optium, OptiumEZ and ReliOn Ultima Blood Glucose Test Strips may generate erroneously low blood glucose results. | Process control | On 12/22/10, Abbott Diabetes Care notified customers via Press Release and letters explaining the reason for the recall and requesting the products be returned. |  | 359000000 | 359 total lots, 1,000,000 strips per lot. | Products were released for distribution in US and Puerto Rico. | Terminated | October 25, 2011 | 307 | 359000000 | Not_Computer
Z-1065-2011 | 57333 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Bond Polymer Refine Detection Kit; an in vitro diagnostic immunohistochemistry reagent kit; Catalog No. DS9800; |  |  | November 16, 2010 | 2010 | Leica Microsystems, Inc. | Users of the Bond Polymer Refine Detection System, DS9800, may experience instances of "weak staining" due to the DAB Part B component, which may give a false negative result. | Other | Leica Microsystems sent Medical Device Field Correction letters dated 11/16/10 to the affected users of the Bond Polymer Refine Detection System, DS9800 via first class mail on the same date, informing them that they may experience instances of "weak staining" due to the DAB Part B component in the affected lots, which may give a false negative result. The users were requested to test their DS9800 System prior to use to rule out that the DAB Part B components have been affected, using the enclosed testing protocol. If their system is affected, the users were instructed not to use it and contact Leica Technical Assistance Center for a replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 653 | 653 units | Worldwide Distribution -- USA and Canada. | Terminated | October 17, 2011 | 335 | 653 | Not_Computer
Z-1067-2011 | 57676 | [6 empty] | Fujinon Endoscope Series; |  |  | January 07, 2011 | 2011 | Fujinon Inc | The FICE feature installed in video endoscopic processors is being disabled since this feature should have been reviewed under the premarket notification/510k process. | PMA | Fujinon sent an Urgent Medical Device Corrective Action Letter, dated 1/7/2011, via overnight mail to all consignees. The letter advises users that the FICE feature should no longer be used and that a Fujinon Field Service Engineer will be coming to physically disable the FICE feature on the processor. Instructions labeled as Attachment B are provided to assist in ceasing use of the FICE function. Attachment A is provided for frequently asked questions. Customers are to complete the Response/Acknowledgement Form and either fax it or email back to the firm. If there are additional questions, customers can call 800-385-4666 or email products.fice.@fujinon.com. |  | 783 | 783 units | Nationwide Distribution | Terminated | June 04, 2012 | 514 | 783 | Not_Computer
Z-1066-2011 | 57321 | No | CKA | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abaxis brand Piccolo Comprehensive Metabolic Panel; Used with the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer; Catalog/Part Number: 400-0028; |  |  | November 19, 2010 | 2010 | Abaxis, Incorporated | Abaxis has determined that during the manufacture of product lot 0396BC5 may report inaccurate values of Alanine Aminotransferase (ALT) when used in the Piccolo Blood Chemistry Analyzer or Piccolo xpress Chemistry Analyzer. | Other | The firm, Abaxis Inc., sent an "URGENT PRODUCT RECALL NOTICE" dated November 19, 2010, to all customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to check their inventory of the lot for boxes #0665 or higher; if found, to discontinue use; contact Abaxis for immediate replacement; complete and return the enclosed PRODUCT RECALL RESPONSE form to Abaxis Technical Support via fax at 1-877-349-2087, and ensure that their laboratory staff is informed of this recall notice.If you have any questions regarding this Product Recall Notice, please contact Abaxis Medical Technical Support at 800-822-2947, or at medtechsupport@abaxis.com. |  | 797 | 797 units | Worldwide distribution: USA and countries of: Kowloon and Germany. | Terminated | August 23, 2011 | 277 | 797 | Not_Computer
Z-1080-2011 | 56391 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Centurion Medical Products Insertion/PPE Kit CV12590, Sterile. |  |  | June 02, 2010 | 2010 | Centurion Medical Products | The kits contain heparin lock flush solution USP code 504505 that is under recall by APP Pharmaceuticals. APP Pharmaceuticals is recalling the heparin because of incomplete documentation associated with test results. | Nonconforming Material/Component | Centurion Medical Products sent out Urgent Product Recall Notices on 6/3/2010. A second recall notice was sent to nonresponding consignees on 6/21/2010. The letter instructed the consignees to quarantine the recalled product and contact Centurion so that the recalled product can be retrieved. |  | 20 | 20 | Nationwide distribution. | Terminated | February 09, 2011 | 252 | 4080 | Not_Computer
Z-1094-2011 | 57202 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald, CELL-DYN Emerald CPU Board Hematology System; |  |  | October 15, 2010 | 2010 | Abbott Laboratories | There is the potential for the instrument to not power up due to the flash memory issue. This will result in an error message displayed is "No Memory Available" and has the potential for delay in generating patient results. | Other | The firm, Abbott Laboratories, sent a "Product Correction" letter dated October 15, 2010 to all customers. The letter described the product, problem and action to be taken by the customer. The customers were instructed to install a new printer driver in the software location using the provided instructions and to complete and return the Customer Reply Form via fax 1-800-777-0051 or email to QAGCO@abbott.com even if they no longer have the instrument.If you have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT; customers outside the U.S., please contact your local hematology customer support representative. |  |  |  | Worldwide distribution: USA including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, O, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WI; and countries including: Argentina, Australia, Brazil, Canada, Dominican Republic, Germany, Singapore, South Korea, St. Eustalius &amp; St. Martin, Trinidad and Tobago. | Terminated | June 07, 2011 | 235 |  | Hardware | Device Operation | Install a new printer driver | Other
Z-1061-2011 | 57651 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMedimplantable drug pump |  |  | January 14, 2011 | 2011 | Medtronic, Inc. - Neuromodulation | This letter provides important reminders concerning the potential for a pocket fill during a SynchroMedÂż II or SynchroMed EL implantable drug pump refill procedure, and important patient management recommendations that will be added to our product labeling. A pocket fill is the inadvertent injection of all or some of the prescribed drug into the patients subcutaneous tissue, which includes the pum | Other | The firm, Medtronic, sent an "Urgent Medical Device Correction Cover " letter dated January 2011 to all customers. Enclosed with each letter, an "Urgent Medical Device Correction Important Clinical Information about Pocket Fills" letter dated January 2011 as well as a Clinician Refill Reference Card, Refill Appointment Cards, Physician Reply Card and a Self Addressed Stamped Envelope. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to always check the needle placement within the pump septum throughout the procedure to help perform a successful refill; be aware that the pump refill difficulty may increase due to a variety of factors (listed); monitor the patient closely for signs and symptoms of overdose according to the provided instructions. Note: Medtronic is in the process of supplementing the labeling related to improper injection with the information included in this communication.Patients with questions are encouraged to talk with their physician or contact Medtronic Patient Services at 1-800-510-6735, Monday - Friday, 8 a.m. to 5 p.m. CDT. Physicians with medical questions related to this issue or Medtronic therapies should contact Medtronic at 1-800-328-0810, Monday - Friday, 8 a.m. to 5 p.m. CDT. UPDATE - June 2013 - Medtronic Field Representative began notifying Healthcare Professionals June 3,2013,via visit to deliver and discuss the Pump Refill Procedure Safety Update letter. The letter informs clinicians that SynchroMed Implantable Infusion System labeling and the Clinician Refill Reference Card have been updated. The labeling and reference card were updated to provide important reminders concerning the potential for a pocket fill during a SynchroMed II or SynchroMed EL implantable drug pump refill procedure. The Medtronic Rep will complete a Physician Notification Confirmation form documenting the delivery of the information. A copy of the recall notification letter a |  | 91369 | 91,369 total for all devices. | Worldwide distribution: USA including DC and PR and countries including: AUSTRALIA, NEW ZEALAND, HONG KONG, TAIWAN, CHINA, SINGAPORE, MALAYSIA, INDONESIA, SOUTH KOREA, THAILAND, ARUBA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, FAROE ISLANDS, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HUNGARY, ICELAND, INDIA, IRELAND, ISRAEL, ITALY, JAMAICA, JORDAN, KUWAIT, LEBANON, LITHUANIA, LUXEMBOURG, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NETHERLANDS ANTI LES, NORWAY, PANAMA, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, SAN MARINO, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VATICAN CITY. | Terminated | February 04, 2015 | 1482 | 91369 | Not_Computer
Z-1222-2011 | 56860 | No | FZT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet HEARTSTRING Aortic Cutter 4.3 mm, AC-3043, Lot numbers 9081071 and 9082071, Manufactured by Maquet Cardiovascular, Wayne, New Jersey. |  |  | June 04, 2010 | 2010 | Maquet Cardiovascular | Pouch integrity may be compromised, resulting in loss of sterility. | Process control | Urgent Device Removal letters were sent out on June 10, 2020 by Federal Express. The letter identified the affected product, described the reason for the recall, and asked for all affected product to be returned. The firm asked customers to examine their inventory and to discontinue distributing the affected lots. In addition, customers are to complete and return the Field Action Response form. If the affected product was further distributed, those customers also need to be notified of the recall and return their affected product to their distributor. All affected returned products may be exchanged with other unaffected HEARTSTRING Aortic Cutter 4.3 MM devices. Customers are to contact their local MAQUET Cardiovascular sales representative or Customer Service at 1-888-880-2874 to arrange for an exchange and if they have any questions. |  | 395 | 395 units | Worldwide Distribution -- USA, including states of IA, KY, IL, WI, TX and NC and countries of Japan, Germany and Hong Kong. | Terminated | June 21, 2011 | 382 | 395 | Not_Computer
Z-1221-2011 | 57733 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bariatric Bed, P1840, Hill-Rom, Batesville, IN. |  |  | February 03, 2011 | 2011 | Hill-Rom, Inc. | The firm received one report that a patient was weighed on the bed at 53 Kg when the patient actually weighed 73 Kg. The patient expired the next day and it is inconclusive as to whether the scale was a contributing factor. The firm investigated and found that the software for the scale system did not clear the last patient weight from the history of the scale system. | Software design | The firm, Hill-Rom, sent a "URGENT MEDICAL DEVICE CORRECTION" letter dated January 19, 2011, along with a software upgrade (USB stick/flash drive) with instructions, to all affected customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to requests that facility service personnel immediately upgrade their software; complete and return the Customer Response Form via fax to 812-931-2410 or mail to: Hill-Rom Company, 1069 State Route 46 E, Batesville, IN 47006, Attn: MOD Coordinator J12 and forward a copy of this notice to any other facility personnel. Note: The customer can request that Hill-Rom technicians complete the software upgrade.If you have any questions concerning this Urgent Medical Device Correction, please contact Hill-Rom Technical Support at 800-445-3720. |  | 40 | 40 beds. | Worldwide distribution: USA and countries including: AUSTRALIA, BRAZIL, CANADA, CHINA, EGYPT, GERMANY, GREECE, ITALY, JAPAN, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, OMAN, POLAND, QATAR, SAUDI ARABIA, SPAIN, SWEDEN, THAILAND, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 26, 2012 | 601 | 945 | Computer
Z-1224-2011 | 57597 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pump in Style Advanced (PNSA)Breast Pump; a daily use, double pump, electric powered breast pump; Manufactured exclusively by and for Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; the affected breast pumps were sold in the following configurations: |  |  | January 10, 2011 | 2011 | Medela Inc | The performance of the pumps may be compromised due to the potentiometer used on the circuit board. Affected pumps may not turn on, may turn off on its own, may reduce or lose suction during the pumping process, or may make a clicking noise. | Vendor change control | This firm, Medela, sent an "Important Device Recall" letter dated January 11, 2011 to all of direct accounts via UPS or US Mail. The accounts were informed that a component on the control board, after a period of usage, may cause the motor to slow down or stop, resulting one or more of the following symptoms: the pump will not turn on; the pump turns off on its own; the pump reduces or loses suction during the pumping process or the pump is making a clicking noise. If the users have a pump that exhibits any of these symptoms, they were requested to visit Medela's website at www.medelareturns.com to expedite shipment of a replacement unit. Questions about the recall process were directed to Medela Customer Service at 1-800-435-8316 or e-mail at medelareturns@medela.com.Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelareturns.com on January 11, 2011. |  | 11154 | 11,154 units | Nationwide, including Puerto Rico | Terminated | May 21, 2012 | 497 | 11154 | Hardware | Device Operation | Replace device | Remove or Replace
Z-1093-2011 | 57405 | No | KQG | Coagulation instrument. | Hematology | Hematology | 510(k) | Hemochron Whole Blood Coagulation Systems Activated Clotting Time (ACT); Package label states -International Technidyne Corporation Response Celite ACT Tube HRFTCA510; Lot number and expiration date are listed. Each box contains 95 ACT tubes tests. |  |  | November 30, 2010 | 2010 | International Technidyne Corp. | Some Activated Clotting Time (ACT) Tubes may contain a component that does not meet its specification, making it possible for a failure to detect a tube or the presence of a clot. | Process control | ITC notified their distributors and end user customers with an Urgent Medical Device Recall letter dated November 29, 2010 through UPS overnight mail. The letter identified the affected product and explained the reason for the recall. Distributors and customers should check their inventory for any affected product. Distributors should immediately discontinue shipment of the affected product and place the product on hold. Then regardless of whether they have any remaining inventory, they should complete the enclosed Distributor Account Tracking Form. The firm will contact them to arrange for any affected product to be returned and will credit them for any unused, returned product. Customers are to stop using the affected product and remove it from their inventory. They should also complete the attached Customer Account Tracking Form and return it to the firm. The firm will contact them to arrange for the return of any affected inventory and will credit the customers for any unused, returned product. If customers believe that they have experienced a problem with the product, they should take any additional actions regarding their records or patients. Questions should be directed to ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International) or e-mail techsupport@itcmed.com. |  | 265 | 265 boxes (95 ACT Tubes per box) | Worldwide Distribution | Terminated | August 22, 2012 | 631 | 265 | Not_Computer
Z-1095-2011 | 57660 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Femoral Distal Fixation Peg; Catalog Number 5575-X-00; Manufactured by Stryker Howmedica Osteonics Corp., Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 USA A Subsidiary of Stryker Corp. |  |  | December 22, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the labels on the Triathlon Distal Femoral Fixation Pegs contain an incorrect expiration date of January 2015 (2015-1) when the actual expiration date should read June 2014 (2014-6). | Incorrect or no expiration date | The firm, Stryker, sent a "UPDATED (Corrected Lot Code) URGENT PRODUCT RECALL" letter dated January 6, 2011, via Fed Ex ,with return receipt to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to fax the attached Product Recall Acknowledgement form to (201) 831-6069; contact the hospitals in their territory that have the affected product to arrange return of the product, and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive Mahwah, NJ 07430.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100 or the Associate Product Manager, Knee Marketing Team at (201) 831-5028. |  | 15 | 15 units | Nationwide distribution. | Terminated | October 03, 2012 | 651 | 15 | Not_Computer
Z-1219-2011 | 57730 | No | MIE |  |  | Cardiovascular | Contact ODE | Pulmonary Valve &amp; Conduit SG |  |  | January 07, 2011 | 2011 | CryoLife, Inc. | A pre-processing culture performed on a companion allograft (Aortic Valve &amp; Conduct SG) tested positive for Staphylococcus aureus. | Other | The firm, Cryolife, sent an "Urgent-Human Tissue Recall" letter dated January 7, 2011 to the consignee/customer. The letter describe the product, problem and action to be taken by the customer. The customer was instructed to determine if any additional patient surveillance or intervention was necessary, and to complete and return the enclosed postcard affirming that the notification was received.If you have any questions, please contact Cryolife's Medical Director at 1-678-290-4403. |  | 1 | One | Nationwide distribution: USA state of: MI | Terminated | November 01, 2011 | 298 | 1 | Not_Computer
Z-1236-2011 | 57505 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker 5190 TPS Series Burs. Powered simple cranial drills, burrs, trephines, and their accessories. |  |  | November 30, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | Product was packaged with an incorrect packaging material and therefore has a potential to be non-sterile which, if contaminated, could lead to a patient infection. | Packaging | The firm, Stryker, sent a "MEDICAL DEVICE RECALL NOTIFICATION" letter dated December 2, 2010 to all customers. The letter described the product, problem and the action to be taken by the customers. The customer were asked to complete and fax in the business reply form indicating they have received the notification and to indicate whether or not they have affected product on hand, and send back any affected product to Regulatory Affairs at Stryker Instruments. Stryker will then send a replacement product to the customer.Upon receipt of the completed business reply form, the customer will be contacted to arrange return of any affected product they have in their inventory. Once the affected product is received, a replacement 5190 TPS Series Bur will be sent to the customer.For questions regarding this recall, please contact Stryker Instruments at 1-800-800-4236 ext. 3584 or kathleen.kelly.smith@stryker.com. |  |  |  | Worldwide distribution: USA including states of: CA, FL, LA, MI, MT, NE, NH, OH, PA, SD, TN, TX, VT, and WA; and countries including: Sweden and Australia. | Terminated | April 13, 2011 | 134 |  | Not_Computer
Z-1231-2011 | 57550 | No | GYK | Shunt system implantation instrument. | Neurology | Neurology | 510(K) Exempt | Integra Peritoneal Introducer Sheaths, 61 cm;Integra NeuroSciences/REF 901-124; Manufacturer Integra NeuroSciences Implants S.A., 2905 Route Des Dolines, 06921 SOPHIA ANTIPOLIS Cedex, France; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, NJ 08536, USA |  |  | December 22, 2010 | 2010 | Integra LifeSciences Corp. | Peritoneal Introducer Sheaths were incorrectly labeled as 61 cm instead of 46 cm. | Labeling mix-ups | Integra Life Sciences sent an Urgent Product Recall letter dated December 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine any of the affected product. Contact Integra Customer Service at (800) 654-2873, select option 2 to arrange for product returns. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating the current status of quantities of affected product in their inventory. For questions call (609) 936-2237. |  | 95 | 95 units | Nationwide Distribution including CO, FL, GA, MO, NC, NJ, PA, TX, UT, and WI. | Terminated | May 07, 2012 | 502 | 95 | Not_Computer
Z-1228-2011 | 57679 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Thomas Medical Products, Acuity Bleed back Control Valve, Acuity Rotating Hemostasis Valve, Acuity Universal Cutter, and Contamination Sleeve for use during vascular interventional procedures. Contamination Sleeve part no. CON-006-02; Acuity Bleedback Control Valve part no. CSP-001-00; Acuity Rotating Hemostasis Valve part no. CSP-002-01; Acuity Universal Cutter part no. CSP-003-00; Acuity Universal Cutter part no. CSP-003-01. |  |  | December 16, 2010 | 2010 | Thomas Medical Products Inc | Compromised sterility/weak or open seals. | Packaging process control | GE Healthcare Thomas medical products Recall notification dated 16 December 2010 (GEHC Ref # 35003) was sent by letter via UPS Next Day Air and/or USPS certified mail. |  | 28679 | 28,679 for all products | Distribution Nationwide USA and Germany. | Terminated | August 08, 2012 | 601 | 28679 | Not_Computer
Z-1230-2011 | 57735 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Bipolar cord for use with the VirtuoSaph Endoscopic Vein Harvesting System, MCBICORD1, Terumo Cardiovascular System, Ann Arbor, MI. |  |  | January 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The Bipolar cord product specification indicates the product has a service life of 50 uses. However the IFU does not specify an end of service life. The firm has recieved 46 complaints on the Bipolar cord. The firm investigated and found that the units involved in the complaints were often over 12 months old. The firm believes the units involved in the complaints were beyond their product life | Other | Terumo Cardiovascular Systems Corporation sent an URGENT MEDICAL DEVICE CORRECTION letter dated January 10, 2011, to Operating room managers. The letter informed the customers of the Bipolar cord product life of 50 uses. The letter states that the IFU for the bipolar cord is being updated to include the information regarding the limit of 50 uses and sterlization cycles. The letter recommends that user keep a sterile back up bi-polar cord in a convienent place during endoscopic vein harvesting procedures. The firm provided a tracking sheet in each package to help users track the number of sterlization cycles for the cord. Customers were instructed to review the Medical Device Correction notice, assure that all users were aware of the notice, and confirm receipt of the communicaiton by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call Terumo CVS Customer Service at 1-800-521-2818. |  | 2782 | 2782 cords | Worldwide Distribution - United States including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, , NY, MO, MS, MT, NC, NE, NH, NJ, NM, OH, OK, OR, PA, SC, TN, TX, VA, WI, AND WV and the countries of Chile, Hong Kong, Malaysia, United Arab Emirates, Belgium, Japan, Korea, The Phillipines, Singapore, Taiwan, Thailand, Germany, and United Kingdom. | Terminated | July 13, 2011 | 184 | 2782 | Not_Computer
Z-1223-2011 | 57607 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan, Model 5.34. |  |  | October 01, 2005 | 2005 | Elekta, Inc. | Investigation found that if a user accidentally selects the wrong image/tube position images can be displayed flipped. | Other | Product bulletin Leksell GammaPlan - "Leksell GammaPlan version 5.34" (dated October 2005) was sent to affected Customers as notification of the problem. Field Change Order FCO-DCO-05-0006 (dated November 22, 2005) and Technical Notes Leksell GammaPlan Leksell SurgiPlan Important Notice - "Potential issue using flipped angiographic images" (dated October 2005) was sent to Elekta Service Engineers as notification of the problem. The letters described the issue and the FCO asked the Service Engineers to confirm their receipt of the documentation. The Technical Notes also discussed image visualization requirements and listed recommendations. |  | 272 | 272 units | Worldwide Distribution | Terminated | March 02, 2011 | 1978 | 272 | Other |  |  | Other
Z-1229-2011 | 56058 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS brand GreenLight PV Laser System, Model-P/N: 0010-9230; |  |  | February 08, 2010 | 2010 | AMS Innovative Center - San Jose | A failure mode whereby the fiber attachment to the GreenLight PV laser system could inadvertently be broken. If fiber were to bend sharply at point of fiber connection, it may break and cause a burn. | Other | AMS Innovative Center - San Jose issued Important Product Update Information letters via express mail on February 8, 2010, and September 15, 2010 to all its direct consignees. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to attach the Fiber Guard to their GreenLight PV Laser System, and to complete the Response Form in the enclosed, self-addressed, postage paid envelope to AMS. For any questions regarding this recall call the Customer Care Department at 1-800-356-7600 or send an e-mail to CustomerCare@ammd.com. |  | 242000 | ~242,000 units | Worldwide Distribution including US | Terminated | February 14, 2011 | 371 | 242000 | Not_Computer
Z-1258-2011 | 51399 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Blood Glucose Monitoring System Test Strips and Freestyle Lite Blood Glucose Monitoring System Test Strips, for use with FreeStyle and FreeStyle Lite Test Strips respectively, 100 count, manufactured by Abbott Diabetes Care, Alameda, CA. |  |  | March 13, 2009 | 2009 | Abbott Diabetes Care, Inc. | Product dispositioned for destruction was potentially diverted to end users. | Other | The firm, Abbott Diabetes Care (ADC), will be posting notification on its Abbott Diabetes Care website. Also, the firm issued an "IMPORTANT: MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2009, to customers ( registered users of Free Style, Free Style Flash, FreeStyle Freedom or Free Style Navigator). For foreign consignees/customers, the firm has issued communication worldwide. The letter described the product, problem and action to be taken by the customers. The customers were instructed to please check the lot number on the test strip carton or vial for the affected product number; if found, discontinue use and call the ADC Customer Care line at: 1-800-310-1500 and the ADC representative will provide information on how to return the test strips and arrange to have replacement test strips sent.If you have any questions, please call ADC Customer Care line at 1-800-310-1500. |  | 1860 | 1860 units | The firm, ADC, quarantined the eight lots and disposed them to scrap. These eight lots were not released from the firm ADC's control; therefore, the firm was unable to determine the number of consignees. | Terminated | February 15, 2011 | 704 | 1860 | Not_Computer
Z-1259-2011 | 54746 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II Turbo Link Hemoglobin Testing System, including Hemoglobin A, Program Reorder Pack 1600 Tests, Program Buffer A, Buffer B, Cartridge Set, Analytical Cartridge, Guard Cartridge, Calibrator, Accessory Box, Instruction Manual and Quick Guide |  |  | February 19, 2010 | 2010 | Bio-Rad Laboratories Inc | The results from the blood device had the occurrence of some ramping baselines on their chromatograms. The ramping baseline can affect quantitation of the blood results. | Nonconforming Material/Component | The firm, Bio-Rad Laboratories, Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated February 5, 2010, to all customers and letter dated February 16, 2010 to subsidiaries via Fed-Ex or by fax. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review there chromatograms form the lots listed and compare them to what documented; follow the noted instructions; destroy reorder packs from the affected lots if a ramping baseline issue is observed, and complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to (510) 741-3954, attention: Bio-Rad CSD Regulatory Affairs Department. If you have any questions, please contact your regional Bio-Rad Office or (510)724-7000. |  | 27 | 27 | Worldwide distribution: USA including states of: OH, MN, SC and VA; and country including: Japan. | Terminated | March 03, 2011 | 377 | 27 | Not_Computer
Z-1249-2011 | 56182 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Radiotherapy delivery system (C-Series Clinac, Trilogy and Novalis Tx), model numbers H14, H17, H18, H23, H26, H29, HCX |  |  | June 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Position errors may result in misdelivery of radiation. | Device Design | Varian sent Urgent Medical Device Correction letters on June 23, 2010 by Federal Express to all customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Field Safety notification including customer distribution lists, was e-mailed to Varian international Customer Support Managers. Managers will translate to local language if necessary, notify sites by appropriate methods. |  | 5741 | 5741 units | Worldwide distribution. | Terminated | May 03, 2011 | 314 | 5741 | Not_Computer
Z-1248-2011 | 57485 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex Dental Systems, Gendex eXpert DC Intraoral X-Ray System; a dental x-ray system; Model 110-0205G1, catalog number EXPERTDC75NA |  |  | December 17, 2010 | 2010 | Gendex Dental Systems | The intraoral x-ray systems were installed improperly with incorrect screws which allow the unit to pull away from the wall when the arm is fully extended. This condition may harm the patient. | Component design/selection | Gendex sent a Urgent Medical Device Field Correction letter dated December 17, 2010 to the consignee, informing them that the x-ray units were installed with improper sized lag bolts and requesting that they contact Gendex Customer Service at 1-800-323-8029 to complete the field correction. |  | 5 | 5 units | California | Terminated | February 15, 2011 | 60 | 5 | Not_Computer
Z-1246-2011 | 57587 | No | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | ClassOne Orthodontics Buccal SS Tubes Label ID Reference Number 351-327 (NF DB UR 6 Convertible .018-14T 14DO HK-Each); |  |  | November 01, 2010 | 2010 | Ortho Organizers Inc | Products were mis-labeled as nickel free. | Pending | An Urgent: Product Recall Notice letter, dated November 1, 2010 was sent to consignees at the distribution level who received recalled products. Consignees were instructed to do the following:1) Carry out a physical count of current inventory and record this data on the business reply card provided with this notice. Return the enclosed Reply Card with the amount remaining inventory to Ortho Organizers, Inc., Carlsbad, CA 92008 within 10 days of receipt of this notification.2) Immediately cease distribution/use of this product and remove from your shelves.3) Notify all of your retail customers with the Urgent: Product Recall Notice (Retail Level) letter who purchased the affected item numbers from you that the affected items are stainless steel and contain nickel.4) Contact your sales representative to return any unused items. Return any unused product within 30 days. Please clearly mark the return carton Recall Material Enclosed. Ortho Organizers shall accept any product covered by this recall and credit your account for all product and shipping costs. |  | 1300 | 1300 | Worldwide Distribution -- USA, including states of CA, CO, IL, LA, MI, MN, PA &amp; TX and countries of CA (Canada), BE (Belgium), NO (Norway), and CO (Columbia). | Terminated | April 27, 2011 | 177 | 1300 | Not_Computer
Z-1255-2011 | 57626 | No | BYL | Nonpowered oxygen tent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Face Tent |  |  | January 07, 2011 | 2011 | Convatec Inc. | Face tents and Oxygen masks may have a sticky glue-like substance on the inside or outside of the mask consisting of dioctyl phthalate, epoxidized soybean oil and glycerol monoleate. | Nonconforming Material/Component | Unomedical, Inc issued recall letters with attached return response forms dated January 7, 2011 via UPS. The letters request the customer stop use and quarantine product. The recall response form is to be completed and returned to Unomedical.The letters/return response form were also sent out in Spanish.Another recall letter was sent dated February 16, 2011 to capture additional lots that werent included in the initial letter. |  | 1579100 | 1579100 pcs. | Worldwide distribution, including USA, Puerto Rico, Canada, Mexico, Colombia, El Salvador, Paraguay, UK, Austria, Belgium, Cyprus, Czech Republic, Denmark, Estonia, Faroe Islands, Finland, France, Germany, Hungary, Iceland, Ireland, Italy, Netherlands, Norway, Poland, Spain, Sweden and Switzerland, Israel, Jordan, Lebanon, Libya, Morocco, Saudi Arabia, UnitedArab Emirates, Russia, Australia, New Zealand,Japan and Vietnam. | Terminated | February 27, 2013 | 782 | 2692300 | Not_Computer
Z-1253-2011 | 57592 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | FlexicairÂż II (Model number 13000) Low Airloss Therapy Units |  |  | December 06, 2010 | 2010 | Hill-Rom Manufacturing, Inc. | Due to product age and parts availability the units are being withdrawn to reduce a potential health risk to patients when bed rails are not used. | Other | Hill-Rom Manufacturing, Inc. sent an Urgent Medical Device Withdrawal letter dated December 15, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The FLEXICAR MC3 AND FLEXICAIR 13000 were no longer available for rental as of November 15, 2010. Customers in Michigan were required to switch to another product by December 31, 2010. All other customers were required to switch to another product by March 4, 2011. Customers advised to continue the use of the product until arrangements could be made to replace the product. were notified by letter on/about 12/06/2010. The letter indicated the units will no longer be available for rental as of 11/15/2010. Patients in Michigan currently using the products will be required to switch to another product by 12/31/2010. All other patients using the products will be required to switch to another product by 03/04/2010. They were advised that a Hill-Rom representative will be contacted them to discuss details of the product withdrawal and possible alternative surfaces for their patients.For questions regarding this recall call 800-445-3720. |  | 1 | 1 Flexicair II unit | Nationwide Distribution including CA, FL, GA, IL, MI , MS, NJ, NY, OH and PA | Terminated | November 01, 2011 | 330 | 46 | Not_Computer
Z-1069-2011 | 57580 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | FlexTray Procedure Delivery System, Product Codes: KBW71SXL- Bariatric Surgery Trays, TBW13SXL- Bariatric Surgery Trays, XBB14B - Bariatric Surgery Trays, and XBB65S - Bariatric Surgery Trays. |  |  | March 27, 2009 | 2009 | Ethicon Endo-Surgery Inc | The firm initiated this second subsequent recall because the affected devices can remain closed and clamped on tissue during surgical procedures due to the fact that the mechanism which connects the articulation joint to the device shaft is insufficient. | Other | Ethicon Endo-Surgery issued "Urgent Expanded Device Recall Event 1981" notification letters dated April 7, 2009, to their customers. They provided a list of all Affected Product Codes of the ETS FLEX45 Articulating Endoscopic Linear Cutters (including the Product Code: ATS45 from the 1st recall) and Affected Flex Tray Codes subject to recall. The notification document reminds the Operating Room and Materials Management customers of the problem which necessitated the previous recall (Z-0361-2011). It goes on to instruct their customers to immediately discontinue use of the affected products. They are to complete and fax back the Business Reply Form, Affected product is to be returned with a copy of the completed Business Reply Form to Stericycle, Replacement product should be received within 30-60 days after return, Customers can contact Stericycle concerning this recall at 1-888-365-4680, reference Event 1981. |  | 135 | 135 kits | Nationwide Distribution -- USA including the following states, AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MO, MS, MT, NC, NE, NH, NJ, MN, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, and Puerto Rico. | Terminated | February 15, 2013 | 1421 | 29538 | Not_Computer
Z-1257-2011 | 57499 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-301A AQUILION ONE (MPDCT0297EA) - Whole Body X-Ray CT Scanner |  |  | September 07, 2010 | 2010 | Toshiba American Medical Systems Inc | Incorrect Dose Calculation. The software that performs the pre-scan dose calculation uses an incorrect parameter for an exam Plan using a tube voltage of 80kV or 100kV. This results in a pre-scan expected dose that is approximately 10% lower than what the displayed value should be when the scanner executes the exam Plan. | Software design | Toshiba America Medical Systems (TAMS) issued notification to consignees 9/7/2010, via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. |  | 396 | 44 units distributed between 09/2007 - 07/2010 | Nationwide distribution. | Terminated | April 20, 2012 | 591 | 396 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1251-2011 | 57604 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Elekta Leksell Stereotactic System, used for localization (spatial reference) for cranial surgery using X-ray or CT and MRI Image data. |  |  | October 02, 2009 | 2009 | Elekta, Inc. | New protocols for MR sequences may result in higher RF energies deposited during MR scanning, generating heat in uninsulated fixation posts. | Labeling design | Elekta sent Important Notice Affected Product: Fixation Posts, A331, dated June 22, 2009. The Important Notice will be added to the Customer User manual upon receipt. |  | 379 | 379 units | Nationwide | Terminated | February 17, 2011 | 503 | 379 | Not_Computer
Z-1247-2011 | 57665 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Mantis Tap Sleeve,Mfg by: Stryker Spine, 2 Pearl Court, Allendale, NJ 07401-1677. packaged in the following products: |  |  | January 07, 2011 | 2011 | Stryker Spine | Stryker Spine has received reports that after sterilization but prior to surgery, cracks were noticed at the tantalum marker / PEI tube interface. This can likely be attributed to differences in the rate of thermal expansion between the markers and the sleeve during the sterilization cycle. | Other | Stryker Spine sent an Urgent Product Recall letter dated January 07, 2011, to Stryker Spine sales branches and agencies where affected product was distributed. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital locations to identify the product.Reconcile any Mantis Tap Sleeves with the catalog and lot numbers referenced in the letter. Complete the Customer Response Form with distribution sheet provided to assist with product reconcilation. Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance at 201-760-8370.Customers were to retrieve and return the affected product using the enclosed pre-paid mailing label to Regulatory Compliance, Stryker Spine 6 Pearl Court, Allendale, New Jersey 07401. For questions call 201-760-8180. |  | 25 | 25 units | Nationwide Distribution including MA, MN, and PA. | Terminated | March 08, 2015 | 1521 | 25 | Not_Computer
Z-1252-2011 | 57731 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | January 05, 2011 | 2011 | Elekta, Inc. | Flexmap calibration did not cover the full 360 degrees of gantry rotation (0 to 355 degrees) and that this caused an isocenter displacement with 2D PlanarView images required at 180 to -180. | Other | The firm, Elekta Limited, sent an "Important Notice A344" dated January 5, 2011 to all consignees/customers. The notice described the product, problem and action to be taken by the customers. The customers were instructed to follow any instructions or recommendations covered in the notice. The notice provides suggestions on how to check and prevent this situation from occurring; it instructs the User to file the document in the Important Notice section of the appropriate User Manual; to complete a FlexMap Check each time they acquire a new FlexMap and amend their local procedures to include the FlexMap Check and complete and return the Important Notice Confirmation of Receipt via Fax to: Field Support Administration, Elekta Limited, Linac House, Fleming Way Crawley, West Sussex, RH10 9RR UK +44 (0)1293 654401.If you have any questions, contact Elekta Limited at +44 (0)1293 654200 or www.elekta.com. |  | 223 | 223 units | Nationwide distribution. | Terminated | April 07, 2011 | 92 | 223 | Other |  |  | Other
Z-1250-2011 | 57159 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Xeleris Processing and Review Workstation 1.0, 1.1, 2.0, 3.0 |  |  | January 12, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of safety issues relevant for the Xeleris Workstation 1.0, 1.1, 2.0, 2.1, 3.0:1. Measurements taken on the CT Images on the CT system are different from the Xeleris Workstation2. Inaccurate PET SUV values may be calculated when data is retrieved from PACS after being processed in the Xeleris.3. Wrong marking of Early and Late result series on display in Bra | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 12, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to please ensure that all potential users in their facility are made aware of this safety notification and the recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare call center in US or Canada at 800-437-1171 and other countries at 262-896-2890. |  | 6046 | 6046 | Worldwide distribution: USA including DC, PR, and Guam; and countries including: AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA, AZERBAIJAN, BELGIUM, BOSNIA &amp; HERZEGOVIA, BULGARIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, LUXEMBOURG, LEBANON, MALY ASIA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, PORTUGAL POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SPAIN, SOUTH AFRICA, SOUTH KOREA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, and VIET NAM. | Terminated | October 16, 2012 | 643 | 6046 | Software | Output/Calculation | Software update | Software Update
Z-1245-2011 | 57379 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare IT, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | December 03, 2010 | 2010 | Ge Healthcare It | The GE Centricity Laboratory System Software has safety issues with the "Telephone List" function, the "Anatomic Pathology" application, and the Instrument Interface IO module that may impact patient safety. | Software design | GE Healthcare, sent "Urgent Medical Device Correction" letters dated December 12, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the product, the problem, and the action to be taken by the customer. The letters provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 26 | 26 units | Worldwide Distribution -- USA, including AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, NC, NH, NJ, NY, OR, TN, VA, and WI, , and internationally to the countries of Australia, England, Hong Kong, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 224 | 26 | Software | Physical Safety Hazards | Software update | Software Update
Z-1334-2011 | 50094 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | S T-Bone Plate; |  |  | May 26, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A package labeled as a S T-Bone Plate, Catalog Number 57-05260, Lot Code G4M00F1576, may actually contain a M T-Compression Plate, Catalog Number 57-10250, Lot Code G4N00F1584 and vice versa. | Process design | Stryker Orthopedics sent an Important Market Withdrawal letter dated May 26, 2006, by Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory, including their hospital consignment location, to identify the product.If the affected product remained in its original packaging, customers were instructed to retrieve all of the affected product lots and return it to their branch or agency warehouse for reconcillation.Reconcile all product on the attached Product Accountability Form and fax a copy of the form to 201-831-6069 within 5 days of receipt of the notice.Customers were asked to return the affected product via Inter-Org Transfer and ship to Product Surveillance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, USA 07430-2104. Customers were also instructed to contact their Stryker Orthopaedics Customer Service representataive to re-order the product that they were returning.For questions regarding this recall call (201) 831-5825. |  | 21 | 21 | Nationwide Distribution | Terminated | December 08, 2014 | 3118 | 21 | Not_Computer
Z-1320-2011 | 57295 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | WalkMed Infusion Triton 90" 20 drop IV Administration Set with Needlefree Y-site |  |  | November 16, 2010 | 2010 | Churchill Medical Systems, Inc. | Routine testing conducted to challenge the success of sterilization was not fully completed for sterile products. | Vendor change control | The firm, Churchill Medical/Vygon, sent an "URGENT MEDICAL DEVICE RECALL" letter dated November 16, 2010, via overnight mail to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to check their stock for affected product; immediately cease use and distribution and quaratine the affected product; contact Vygon's Customer Service Department at 1-800-473-5414 for a return Merchandise Authorization number; complete and return the Recall Acknowledgement and Inventory Return Form via fax to 215-390-2019 and mail return product with form to: Vygon, 103A Park Drive, Montgomeryville, PA 18936, and if the customer is a distributor, they were to provide their customers with a copy of this notification and attached form. Note: An additional follow up communication will be provided to the customers after reconciliation of the returned inventory and all inventory will be credited or replaced.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email at customerservice@vygonus.com. |  | 12150 | 12150 units | Nationwide Distribution: USA including states of: CO, IA, OH, MA, MD, MO, NY, FL, SC, FL, and VA. | Terminated | May 01, 2012 | 532 | 80232 | Not_Computer
Z-1341-2011 | 57454 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Celite ACT Cartridge |  |  | December 03, 2010 | 2010 | Abbott Point Of Care Inc. | Celite ACT cartridges have increased numbers of values outside the accepted allowable error. | Device Design | On behalf of Abbott, Genco Pharmaceutical Services sent an "URGENT RECALL NOTICE" dated November 2010, on December 3, 2010 via Fedex to all consignees/customers. Abbott sent out notification packages to foreign consignees/customers on December 3, 2010. The notice described the product, problem and action to be taken by customers. The firm recommended that customers use cartridges within 12 hours of removal from refrigerated storage based on the instructions provided with the product. The customers were instructed to complete and return the Business Reply cards to acknowledge receipt of the communication and provide the letter to any customers that have recieved the product. If you have any additional questions, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care Representative.On February 27, 2012 Abbott sent an Important Product Information letter via Fed Ex to inform customers that i-Stat Celite ACT cartridges manufactured from the beginning of 2012 can now be stored for two weeks at room temperature as indicated on the cartridge box. Foreign consignees were notified on 2/28/2012. Users are instructed to complete business reply cards to acknowledge receipt of the communication and provide the letter to any customers that have received the product. Any additional questions can be answered by contacting 8003668020 option 1. |  | 1164275 | 1164275 cartridges | Worldwide distribution. | Terminated | July 16, 2013 | 956 | 1164275 | Not_Computer
Z-1261-2011 | 57318 | No | KTW | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 02.122.550, 50mm Spiral Blade, 14mm, Right - Synthes Spiral Blade Component of the Synthes Modular Blade Plate System |  |  | October 27, 2010 | 2010 | Synthes USA (HQ), Inc. | Potential for out of specification star grind teeth. | Other | Synthes initiated their recall on October 18, 2010, notifying affected hospital accounts via consultant communication. Affected Synthes Trauma Sales Consultants were notified via email (receipt requested). A copy of the Urgent Medical Device Recall letter dated October 20, 2010, was mailed via UPS second day air on October 20, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected devices and call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. A replacement order would be taken at that time.Customers were also instructed to complete the Verification Section at the end of the letter by checking the appropriate box indicating the affected product had been located or indicating in the appropriate box that no affected product was located. Also, customers were to indicate the number of devices found and note the Return Authorization Number. Return the Verification Form with the recalled product to:Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.For questions regarding this recall call 1-800-620-7025, x 6883. |  | 19 | 19 | Nationwide Distribution including AL, CA, GA, IA, IL, IN, MA, MI, MS, NC, NJ, NV, OH, PA, SC, SD, TN, and WV. | Terminated | April 16, 2015 | 1632 | 561 | Not_Computer
Z-1340-2011 | 57792 | No | HMR | Ophthalmic surgical marker. | Ophthalmic | Ophthalmic | 510(K) Exempt | Walcott Rx Products, Surgical Division, Corneal Marking Pad, REF RX40490, Sterile. For use in ophthalmic surgery. |  |  | December 23, 2010 | 2010 | Aspen Surgical Products, Inc. | Fragments from the pad may disengage and stick to the instrument. The condition could cause fragments to come into contact with the intact cornea prior to surgery. | Component design/selection | Aspen Surgical sent URGENT MEDICAL DEVICE RECALL NOTICES dated 12/23/2010 by Certified Mail. The letters instructed the users to examine their inventory and place the recalled products in quarantine. Consignees were instructed to return the product and complete the response form. Distributors were asked to notify their customers or requesting that Aspen notify the customers.The firm does not intend to continue manufacturing or distributing the product. |  | 77850 | 77850 all products | Worldwide distribution including USA Canada, England, Sweden, Saudi Arabia, England, Hong Kong, Belgium and Israel. | Terminated | October 07, 2011 | 288 | 77850 | Not_Computer
Z-1362-2011 | 57544 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConMed System 2450 Electrosurgical Generator, REF 60-2450-100, 100 V, ConMed Electrosurgery, Centennial, CO. |  |  | December 27, 2010 | 2010 | ConMed Electrosurgery | Electrosurgical generator may allow possibility of electrical shock and/or burns to user or patient. | Device Design | ConMed Electrosurgery sent an Urgent Device Recall letter dted December 27, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify anyone who may have received the affected product.. They were to review their inventory and return a Business Reply Form via fax to 303-699-9854 and to contact ConMed Electrosurgery Technical Services at 800-552-0138 to schedule a time to have their equipment upgraded. |  | 119 | 119 units | Worldwide Distribution, including USA (nationwide) and the countries of Australia, Belgium, Bolivia, Canada, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Honduras, Hong Kong, Hungary, India, Iraq, Israel, Italy, Jamaica, Japan, Jordan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, Vietnam. | Terminated | August 01, 2011 | 217 | 2323 | Not_Computer
Z-1346-2011 | 57287 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ARTERIAL LINE KIT/Product SKU (Code) AK-04510 |  |  | October 29, 2010 | 2010 | Arrow International Inc | Sterility may be compromised. | Other | Arrow International, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated October 29, 2010, to all affected customers. The letter identified the products, the problem, and the action to be taken by the customer. Customers were instructed to quarantine the product.If they had no affected product in stock, check the appropriate box on the Recall Acknowledgement &amp; Stock Status Form and fax the form to Arrow's Customer Service fax number 800-343-2935.If they had affected stock, immediately quarantine the product and contact Arrow Customer Service at 800-523-8446 for a Return Authorization Number. If customers wished to return product for replacement, they were to consult their sales representative or Arrow Customer Service for information on product availability, substitutability, and lead times. Complete the enclosed Recall Acknowledgement &amp; Stock Status Form and immediately fax the compoleted form to Arrow, Fax number 800-343-2935, Attn: Customer Service.Return any affected product freight collect, along with the original completed Recall Acknowledgement &amp; Stock Status Form to:Arrow International, Inc.312 Commerce PlaceAheboro, NC 27203Attn: Asheboro Quality Assurance Manager.For any questions regarding this recall call 800-523-8446. |  | 361640 | 361,640 | Worldwide Distribution -- United States - all 50 States, and the countries of Canada, South America, China, South Africa, Australia, Puerto Rico, Philippines, Singapore, South Korea, New Zealand, Mexico, Japan, India, Indonesia, Bahamas, Taiwan, Thailand, Finland, France, Cyprus, Czech Republic, Belgium, Germany, Ireland, Italy, Latvia, Netherlands, Poland, Slovakia, Spain, Bahamas, Costa Rico, Dominican Republic, Ecuador, Guatamala and the United Kingdom. | Terminated | September 29, 2012 | 701 | 963380 | Not_Computer
Z-1342-2011 | 56387 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, Model Number/Catalog No.: 270-2417, which contains the update kit CD-ROM, Model/Catalog No.: 270-2423, Distributed and Manufactured by: Bio-Rad Laboratories, Inc., Hercules, CA. |  |  | March 24, 2010 | 2010 | Bio-Rad Laboratories Inc | The firm received customer complaints associated with the Update Kit CD-ROM not being able to generate a valid calibration after being uploaded, resulting in patient results not being generated. | Other | Medical Device Correction letter dated March 24, 2010 was sent to customers. The letter identified the affected product and explained the reason for recall. Customers were asked to destroy Lot No. AA00190 and replace it with the Lot No. AA00838 CD-ROM that is enclosed with the letter. Customers are to complete and fax back the Medical Device Correction Response Form at the number provided. |  | 48 | 48 units | Worldwide Distribution -- USA and Italy. | Terminated | March 04, 2011 | 345 | 48 | Not_Computer
Z-1370-2011 | 57850 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | MitraXs" Self Adjusting Atrial Retractor, Sterile EO, Models (REF): A0030/1; A0031/1; A0032/1; A0033/1, Cardio Life Research s.a., Green Hill, Hoge Wei, 16, B-1930 Zaventem, Belgium. |  |  | January 26, 2011 | 2011 | St. Jude Medical Cardiovascular Division | St. Jude Medical Cardiovascular Division is conducting a voluntary recall of all batches and model numbers of the MitraXs Self Adjusting Atrial Retractor. St. Jude Medical has been distributing this product since their acquisition of CLR in late 2008.They have determined that some distributed batches have very small burrs on the polycarbonate material at the distal edge. As these burrs may | Process control | The firm, St. Jude Medical, sent an "Urgent Medical Device Recall Notice" dated January 26, 201,1 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to discontinue use of the product (MitraXs Self Adjusting Atrial Retractor batches) and remove from inventory. St. Jude began production of a new batch of units that will bear the labeling of St. Jude Medical. Shipment of these units started on January 10, 2011. Note: A St. Jude Medical Sales Representative will contact the customer to facilitate removal and reconciliation of the affected product.If you have any questions or concerns, please contact your St. Jude Medical Sales Representative or 651-490-4470 and/or 612-310-7936. |  | 450 | 450 | Worldwide distribution: USA including states of: NJ, OH, MI, and MN; and countries including: BELGIUM, FRANCE, GERMANY, SWITZERLAND, and UNITED KINGDOM. | Terminated | May 31, 2012 | 491 | 450 | Not_Computer
Z-1374-2011 | 56394 | No | LZS |  |  | Ophthalmic | PMA | AMO brand WaveScan WaveFront System, OPM software versions 3.67 and 3.67, |  |  | July 13, 2010 | 2010 | AMO Manufacturing USA, LLC | The failure to comply with Minimum System Requirements, increased system errors could result, and a miscreated treatment table (incorrect LASIK refractive treatment) could result. | Software design | Abbott Medical Optics, Inc. sent an Urgent Medical Device Field Correction letter to all affected customers starting July 15, 2010. The letter identified the product, the problem, and the action to be taken by the customer.Customers were asked to take the following actions:Surgeons were instructed to complete the facsimile attachment to acknowledge receipt of the notification letter and fax to 714-247-4530.For those using the affected product a 3.68 OPM disc must be requested by e-mailing a request to OPM@pacful.com.For any questions please contact AMO Technical Service Representatiave at 1-877-266-4543. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britian, Greece, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Malaysia, Mexico, Netherland, New Zealand, Peru, Philippines, Portugal, Russia, Russia Federation, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Taiwan, ThailandChina, India, South Korea, Singapore, Taiwan, Thailand, Turkey, United Arab Emirates, Venezuela, and Vietnam. | Terminated | June 22, 2012 | 710 |  | Other | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1373-2011 | 57502 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; part number 110-0234G1 (domestic) and 110-0234G3 (international) |  |  | December 21, 2010 | 2010 | Gendex Dental Systems | The Automatic Exposure Control (AEC) feature located on the control panel of the Orthoralix 8500 Panoramic Dental X-Ray System was functional, but was subsequently disabled in units manufactured starting 5/17/10. North American units manufactured between 12/4/09 and 5/16/10 had functioning AEC buttons, but their manuals stated they were disabled. Units intended for International distribution bet | Labeling Change Control | Gendex sent a recall letter to their accounts on December 21, 2010, stating that the AEC button on the control panel was a non-functional feature that becomes illuminated upon activation and that Gendex would send them a kit to correct this issue and inactivate the button. A second letter was sent to the accounts on January 31, 2011, explaining that a corrective kit was no longer necessary, letting them know if the AEC feature was active or inactive for their system, and providing them with the appropriate user manual addendum for their system.Customers were asked to complete and return the required Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com.For questions regarding this recall call 800-323-8029. |  | 574 | 574 systems | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WA, and WI and the countries of Australia, Canada, Germany, New Zealand, Taiwan, and Vietnam. | Terminated | November 09, 2011 | 323 | 574 | Not_Computer
Z-1260-2011 | 57567 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device Product Code STRAP25. |  |  | December 29, 2010 | 2010 | Ethicon, Inc. | The Ethicon SecureStrap 5 mm Absorbable Strap Fixation Device may have the potential for the sterile barrier to be compromised due to a defect in the packaging. | Pending | Ethicon sent Urgent Voluntary Product Recall letters dated, December 28, 2010, to their customers on December 29, 2010 via UPS. The letter identified the affected product and described the reason for recall. Customers are to discontinue use and sampling immediately. They are to return the product using the enclosed prepaid shipping label along with the enclosed business reply card regardless of whether they have any affected product or not. Customers should contact 1-877-551-7147 or speak to their manager if they have questions regarding this recall. |  | 4110 | 4,110 | Worldwide Distribution -- USA and Canada. | Terminated | November 07, 2012 | 679 | 4110 | Not_Computer
Z-1372-2011 | 57736 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Zoll Medical (Part # 16381, HeartSync, Ann Arbor, MI. |  |  | December 17, 2010 | 2010 | Heart Sync, Inc | The seal on the electrode pouch may be open, exposing the electrode to the outside environment. The package does have a caution statement that states, " Do not open package until immediately prior to using electrodes. and Electrodes may dry out when removed from packaging and exposed to air. | Packaging | Heart Sync, Inc sent an URGENT MEDICAL DEVICE RECALL letter dated December 23, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Examine their inventory and quarantine product that was subject to the recall that was not sealed properly. Contact any customers if they further distributed the product and notify them at once of this product recall. Contact the firm if they had any unused product. The firm would issue a UPS Call Tag to have the product returned to them and the firm would send replacement product as soon as they were notified.Complete and return the Recall Return Response Form to the firm in the attached prepaid envelope.For any questions regarding this recall call 734-213-5530. |  | 300 | 300 | Worldwide Distribution - USA including AK, AZ, CA, FL, IL, KS, LA, MA, ND, NM, OH, TX, WA, WI and WV; and Canada | Terminated | July 13, 2011 | 208 | 300 | Not_Computer
Z-1378-2011 | 57666 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | InfusO.R. Infusion Pump, an Rx syringe infusion pump; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A.; product codes 2L3100, 2L3100R, 2L3100U, 6461500 and 6461500R. These product codes includes refurbished units and InfusO.R. pumps previously sold under the Bard label. |  |  | January 10, 2011 | 2011 | Baxter Healthcare Corp. | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches), an electrical circuit will not be completed and the InfusOR will not operate. The reason being that the InfusOR (2L3100) Battery Spring has a flat contact that measures .53 inches Âż.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | Other | Safety Alert letters dated January 11, 2011 were mailed via first class mail to the InfusOR direct accounts, to the attention of the Director of Biomedical Engineering, the Director of Anesthesia and the Director of Materials Management, on the same date, informing them that over-the-counter Energizer C size alkaline batteries should not be used with the InfusO.R. Pumps (Product codes 2L3100, 2L3100R, 2L3100R, 2L3100U, 2L3100N, 6461500 and 6461500R). That particular battery incorporates a feature that may inhibit electrical contact and affect the ability of the pump to power-up. The accounts were requested to check their inventory of InfusO.R. Pumps for the presence of over-the-counter Energizer C size alkaline batteries. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST.*** Baxter expanded the scope of their action to include the Mini-Infuser System and sent Urgent Device Correction letters dated 3/4/11 via First Class mail to all InfusOR and Mini-Infuser customers who had purchased either of the affected product lines. The letters informed the accounts that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The accounts were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future refer |  | 39725 | 39,725 pumps | Nationwide Distribution | Terminated | August 27, 2012 | 595 | 39725 | Battery | Device Operation | Instructions | Safety Notice/Insructions
Z-1377-2011 | 57270 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Optima MR450w 1.5T MR systems |  |  | November 18, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware that during certain service procedures, there is a potential for electric shock associated with the gradient cooling manifold hoses of your MR750, MR450, and MR450w MR systems that may impact service or maintenance personnel's safety.The gradient cooling manifolds used on the affected systems may contain electrically conductive hoses, which become live when the gra | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 16, 2011, to all affected customers. The letter was addressed to Hospital Administrator / Risk Manger, Radiology Department Manager and Maintenance / Biomedical Engineering Department. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product correction and Contact Information.Customers were instructed to follow appropriate lockout/tagout (LOTO) procedures.Customers were instructed to contact GE Healthcare engineering at 843-687-7213 or 843-697-3009 for instructions on how to proceed. For questions concerning the notiifcaton customers were instructed to contact their local Support Engineer. |  | 68 | 68 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, LA, MD, MA, MI, MO, NV, NJ, NY, NC, OR, PA, SC, TN, TX, WA, WI. and the countries of AUSTRALIA, BRAZIL, CANADA, EHYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, ITALY, JAPAN, KUWAIT, QATAR, REPUBLIC OF KOREA, SWITZERLAND, TAIWAN, and the UNITED KINGDOM.UNITED KINGDOM, TAIWAN, SWITZERLAND, REPUBLIC OF KOREA, QATAR, KUWAIT, JAPAN, ITALY, INDIA, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, CANADA, BRAZIL, AUSTRALIA. | Terminated | February 04, 2012 | 443 | 104 | Not_Computer
Z-1390-2011 | 56670 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The GMD UNIVERSAL URINARY INCONTINENCE SLING (Code 1010) is supplied sterile and is intended for single patient use only. |  |  | February 14, 2011 | 2011 | Generic Medical Devices | Protective sheath sticks to itself and sleeve breakage has resulted in complaints and MDRs. | Component design/selection | The firm, GMD, issued an "URGENT MEDICAL DEVICE RECALL" letter dated February 14, 2011 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed not to use any GMD Model 1010 Universal Slings remaining in their inventory; return any Model 1010 slings in their possession to the Distributors or GMD; notify their customers to whom they shipped Model 1010s of this voluntary recall and request them to return any Model 1010 slings in their possession to them; and return all Model 1010 inventories to GMD using the attached Recalled Model 1010 Return Form. Note: Full credit in the amount of the returned products purchase price will be issued for all sterile, unexpired product returned to GMD.Contact the GMD Product Performance Senior Manager, at phone: (253) 853 3512 or email: 1010recall@gmd-us.com, if you have any questions/issues related to this recall. |  | 1711 | approximately 1711 | Worldwide distribution: USA and countries including: Australia, Argentina, Brazil, Canada, South Africa, and Scandinavia. | Terminated | September 24, 2012 | 588 | 1711 | Not_Computer
Z-1392-2011 | 53691 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | V5M TEE Transducer, when used with an ACUSON S2000 ultrasound system that has cardiac options installed and is at software version VA10, VA10A, VA10B, VA10C, VA15, VA15A, VA15B, VA16 or VA16A, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | August 28, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | If the transducer is not disconnected from the system after an automatic shutdown due to thermal limits being exceeded, then the next time it is used, the thermal safety guards do not take effect. Normally, the thermal limit before automatic image freeze (and therefore power shut off) occurs at 43 degrees Centigrade. | Software design | Siemens sent a Customer Safety Advisory Notification letter on August 28, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their local Customer Service Engineer would install a software update as soon as it becomes available. Customers were instructed to pass the Customer Safety Advisory Notification letter to all who need to be aware of these issues until the corrective action is completed. For questions regarding this recall call 650-969-9112 or 800-422-8766. |  | 97 | 97 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, China, Czech Republic, Denmark, Germany, Great Britian, Hong Kong, Indonesia, Italy, Japan, Norway, Poland, Portugal, Russian Federation, Singapore, South Korea, Taiwan, and Turkmenistan. | Terminated | March 29, 2011 | 578 | 97 | Software | Physical Safety Hazards | Software update | Software Update
Z-1395-2011 | 57715 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | SP Brand Strep A Cassette; Part Number: B1077-26; Lot #: STA0010037 |  |  | January 12, 2011 | 2011 | Biosite Inc Dba Innovacon Inc. | The recall was initiated after the Investigation of customer-reported complaints, testing of market retention and returned product, It was determined that this lot of product hasdemonstrated an elevated Incidence of non-specific binding, potentially causing false positive results. A falsepositive test result could lead to the administration of antibiotics that are unnecessary. This exposes the | Pending | The firm, Alere San Diego, sent a "URGENT MEDICAL DEVICE RECALL" letter dated January 17, 2011. The letter described the product, problem and action to be taken by customers. The customers were instructed to: (1) Immediately discontinue use of the affected lot (if applicable) and inventory remaining full or partial kits at their site (if applicable);(2) Destroy any remaining inventory of this lot and document destruction per their site requirements (if applicable);(3) Complete and FAX the enclosed Verification Form within 10 days to confirm their receipt of this notice and to receive replacement of their remaining inventory of this lot (if applicable). The customers were ask to please return the form even if they have no remaining inventory or do not have any involved product in order to confirm receipt of this notification; and (4) to please ensure that all users of the device have received a copy of this Urgent Medical Device Recall.Customers with questions about the information contained in the notice were instructed to contact Technical Services at 888-246.7483, option 2 or by e-mail attechservices@alere.com. |  | 19877880 | 814 kits/24,420 devices | Nationwide distribution | Terminated | June 26, 2012 | 531 | 19877880 | Not_Computer
Z-1383-2011 | 57790 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AUTOMIX High Speed Compounder System, product 2M8077; automated nutrition compounders; Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield, IL 60015 |  |  | January 27, 2011 | 2011 | Baxter Healthcare Corp. | The AUTOMIX compounder may not respond correctly to the key pressed by the operator using the keypad. The incorrect keypad response issue can occur both in manual data entry mode (STD Mode) and when entering order numbers to download from host software (MAN ID Mode). It is possible that an incorrect volume and/or an incorrect solution could be admixed into the final bag. | Other | Baxter Healthcare Corp. sent Urgent Device Correction letters dated January 27, 2011, via first class mail to all affected customers. informing personnel that use Automix compounders to exercise care when entering data, and verify the accuracy of the keypad response. The accounts were requested to convey this information to all users of the Automix compounders, emphasizing the need to verify data accuracy each time the keypad is used. The accounts were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667.After determining the root cause of the incorrect keypad responses, Baxter issued a follow-up correction letter on 4/27/11 outlining the recommended work practices to reduce fluid penetration into the Automix compounder keypad/control module. |  | 270 | 270 units | Worldwide Distribution - USA - Nationwide including Puerto Rico, and internationally to Brazil, Canada, Colombia, Costa Rica, Czech Republic, El Salvador, Egypt, Germany, Guatemala, Honduras, Luxembourg, Mexico, Panama, Saudi Arabia, Singapore, Switzerland, Taiwan, Trinidad, Turkey, United Arab Emirates and United Kingdom. | Terminated | February 16, 2012 | 385 | 5160 | Computer
Z-1393-2011 | 56367 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Tyvek Self-Seal Pouch with STERRAD Chemical Indicator 12332. |  |  | June 23, 2010 | 2010 | Advanced Sterilization Products | The recall was initiated after Advanced Sterilization Products (ASP) discovered a non-conforming lot. The misaligned perforations may have caused the seal to be incomplete on some of the pouches, creating a potential breach for pouches being used to store sterilized equipment on shelves. | Process control | Advanced Sterilization Products (ASP) released Urgent: Product Recall notifications to its consignees identifying the affected product and explaining the reason for recall. The letter also discussed the impact of using this affected product and return instructions. The firm asked customers to examine their inventory, discontinue use of Tyvek Self-Seal Pouch with STERRAD Chemical Indicator 12332 and to return all unused affected products to Stericycle. Customers are to complete the Business Reply Card and return it to Stericycle. |  | 470 | 470 | Worldwide Distribution -- USA, Argentina, Australia, Canada, Chile, Czech Republic, Hong Kong, India, Japan, Panama, Peru and Puerto Rico. | Terminated | March 14, 2012 | 630 | 470 | Not_Computer
Z-1394-2011 | 57082 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet Cardiopulmonary AG, MAQUET RPM 20: Roller Pump for Heart Lung Machine HL20. |  |  | September 28, 2010 | 2010 | Maquet Cardiovascular Us Sales, Llc | A wrong ball bearing was installed in the pump heads of product. This may make the pump head wear out and cease function. | Component change control | Maquet hand delivered a letter to the consignee on 10/4/2010. |  | 4 | 4 units | One hospital in New York City | Terminated | August 11, 2011 | 317 | 4 | Not_Computer
Z-1388-2011 | 57245 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-ROM Distal Femoral Cutting block Guide SZ X- SM/SM, REF 216311001, non-sterile, Depuy International Leeds, England. |  |  | October 19, 2010 | 2010 | Depuy Orthopaedics, Inc. | The firm was notified that the Hinge Block Femoral Box Guide cut being off alignment when using the S-ROM Distal Femoral Cutting Block Guide SZ SML/X-SML pin holes to locate. The firm then found that the same issue was found int the MED block size. | Device Design | The firm, DePuy, sent an "URGENT INFORMATION - DEVICE CORRECTION" email and letter dated October 19, 2010 to all customers. The notices described the product, problem and actions to be taken by the customers. The notices stated that: Depuy is issuing a device correction; however, the product may still be used; the devices do not need to be returned; the distributors were to provide a letter to clinicians that were using the device; the surgical technique will be revised in October 2010; the product will undergo a design change, and the recalled product will be traded out in December 2010 and January 2011.If you have any additional questions, please contact DePuy's Scientific Information office at 1-888-554-2482.On 3/4/2011, The firm sent out a revised URGENT DEVICE RECALL asking that all affected cutting guides be taken out of service immediately. Devices should be returned to DePuy and new cutting guides will be sent out within 2 weeks. |  | 229 | 229 | Worldwide distribution: USA including states of: AZ, AR, CA, CO, FL, GA, IA, ID,IL, IN, KY, KS, LA, MA, MD, MI, MN, NC, NV, PA, TN, TX, VA, WA, and WI; and countries including: Canada and Ireland. Australia, Austria, Canada, Denmark, Ireland, Israel Italy, New Zealand, Russia, Sweden, Switzerland, Spain, and UK. | Terminated | December 03, 2012 | 776 | 448 | Not_Computer
Z-1387-2011 | 57700 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Masimo Hemoglobin |  |  | December 23, 2010 | 2010 | Masimo Corporation | Masimo Corporation is initiating a voluntary recall on the Pronto-7 rainbow 4D Resuable Sensor because when performing an SpHb measurement, the sensor may incorrectly read actual finger temperature. As a result, Pronto-7 Check Pulse CO-Oximeter may provide an incorrect SpHb measurement. | Process design | Masimo sent an Urgent Voluntary Medical Device Recall letter dated December 21, 2010, to all affected customers via FexEx on December 23, 2010. The Urgent Voluntary Recall Notice provides instructions on the actions that the customers and field personnel must take to correct the issue.Customers were instructed to choose any of the three options below. Option 1 would ensure accurate results are achieved when using the Pronto-7 rainbow 4D Reusable Sensor with the Pronto-7 Spot Check Pulse CO-Oximeter. (1) Option 1: Prior to performing an SpHb measurement, wait at least seven (7) minutes with the patient's finger inserted inside the Pronto-7 rainbow 4D Reusable Sensor. The seven minute wait will help ensure that the actual finger temperature and the sensor temperature are close to equalized.Masimo is working to quickly develop a revision to the Pronto-7 rainbow 4D Reusable Sensor that will address this issue and plan to have it available in early 2011. We will contact you when the new revision of the sensor becomes available to replace the current sensor;(2) Option 2: Return only the Pronto-7 rainbow 4D Reusable Sensor and request a loaner Pronto device with rainbow DCI sensor (part numbers 9167 and 2646) to be utilized until the new version of the sensor is available for use with the Pronto-7. Please let us know if you wish to try the Pronto or see the product brochure for Pronto. We will contact you when the new revision of the sensor becomes available;(3) Option 3: In addition, the Pronto-7 rainbow 4D Reusable Sensor and the Pronto-7 Spot Check Pulse CO-Oximeter may be returned for a refund.Customers were also instructed to complete the Tracking/Verification Form provided with the recall notice and fax it to Masimo at: 1-949-297-7499. |  | 497 | 497 | Worldwide Distribution - Nationwide in the US and worldwide to Argentina, Austria, Bahrain, Belgium, Benin, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebaron, Libyan Arab, Jamahiriya, Malaysia, Malta, Mauritius, Morocco, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Reunion, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United kingdom | Terminated | March 26, 2013 | 824 | 497 | Not_Computer
Z-1391-2011 | 57805 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT Cardiovascular - autotransfusion System |  |  | January 25, 2011 | 2011 | Haemonetics Corporation | Update Instructions for Use for the CardioPAT&reg; Manual for Leak Detection. User may mis-interpret the air-leak rate displayed by the cardioPAT device when making a determination about chest-drain removal, which could result in the patient experiencing a pneumothorax. | Labeling design | Haemonetics issued a Field Correction Notice on 1/25/11 . with the Air Leak Detection addendum and index be inserted into the appropriate document, applied to all cardioPAT@ devices on site. The serial number for each cardioPAT device on site should be recorded on the attached acknowledgementand the completed acknowledgement should be returned to your Haemonetics Sales Consultant. A Haemonetics Sales Consultant will train on the content of these items and insert them into the Operator Manuals |  | 309 | 309 units | Worldwide distribution, including USA, Australia, Belgium, Canada, Switzerland, China, Germany, Finland, France, United Kingdom, Hong Kong, Israel, Italy, Japan, Netherlands, Russian Federation, and Sweden. | Terminated | January 14, 2013 | 720 | 309 | Not_Computer
Z-1401-2011 | 57781 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Verio |  |  | January 10, 2011 | 2011 | Siemens Medical Solutions USA, Inc | On some systems the gradient cable connections did not meet firm's specifications after installation | Other | Siemens Medical Solutions USA, Inc sent customers a letter on January 10, 2011 notifying them of an Update Instruction and informing them that recalling firm would preventatively check the gradient cable connections of all systems. For questions regarding this recall call 610-219-6300. |  | 92 | 92 | Nationwide Distribution including AK, AZ, CA, CT, FL, HI, IL, IN, KY, MA, MD, MI, MN, MO, MT, NC, NJ, NM, NV, OH, OK, OR, PA, RI, TN, TX, VA, WI, and WV | Terminated | September 25, 2012 | 624 | 92 | Not_Computer
Z-1421-2011 | 57775 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM SIRESKOP SD, SIRESKOP SX, AXIOM LUMINOS |  |  | January 14, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm became aware of a potential malfunction and possible hazard to patients in the event the tilt buttons on the OPTI Grip become stuck under the housing | Other | Siemens Medical Solutions USA, Inc issued a Customer Safety Advisory Notice via Update Instruction XP005/10/S to affected customers. This letter informs customers of the potential malfunction and of possible danger for patients, users or other persons. A future Update Instruction will be released for the replacement of the tilt buttons on all OPTI Grip handles. Customers were instructed to forward the safety notice to any new owners of this device and to provide Siemens Medical with names of the new customers. For any questions regarding this recall call 610-448-3237. |  | 1323 | 1323 | Nationwide Distribution including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | July 16, 2013 | 914 | 1323 | Not_Computer
Z-1164-2011 | 57663 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Spectrum Central Venous Tray, Minocycline/Rifampin Impregnated Double Lumen Polyurethane, 5.0 Fr/ 15cm, C-UDLMY-501J-LSC-ABRM-HC-IHI-FST, G48067, Sterile, Cook Medical, Cook Incorporated, Bloomington, IN. |  |  | October 04, 2010 | 2010 | Cook, Inc. | Cook, Inc. is recalling Single, Double, Triple, and Five-Lumen Central Venous Catheter Trays AND Single and Double Lumen PICC Peripherally Inserted Central Venous Catheter Trays. These kits and trays contain 0.9% Sodium Chloride Injection USP 5 mL fill in 6 mL flush syringes that have been recalled by Excelsior Medical. During a routine test for syringe sterility, a possible integrity failure wa | Nonconforming Material/Component | Cook, Inc. sent URGENT DEVICE RECALLS letters to its customers dated 10/4/2010. The letters asked customers to examine their inventory and notify Cook for product return. The letter included a response form. |  | 51 | 51, 473 total trays | Worldwide distribution: USA, Chile, Colombia, Panama, Barbados, Canada, Guatemala, and Saudi Arabia. | Terminated | June 21, 2011 | 260 | 51 | Not_Computer
Z-1400-2011 | 57818 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Integra Radionics HRAIM Intubation Head Ring Assembly |  |  | January 12, 2011 | 2011 | Integra LifeSciences Corp. | Overall length of the intubation hoop in the HRAIM Intubation Head Ring Assembly is too long and will not allow a device to to attach | Manufacturing material removal | Integra sent a Medical Device Field Corrective Action letter dated January 12, 2011, to all affected customers via Federal Express. Integra Sales Specialists were provided with the replacement part and written instructions on how to replace the affected intubation hoops from customer inventory identified in the traceability report. Upon completion of the field correction each sales specialist was instructed to return the affected hoop along with a completed copy the Recall Acknowledgement and Return Form. If additional information is required contact me at (609) 936-2485 or james.tillman@integraJife.com. |  | 8 | 8 units | Worldwide Distribution - USA, Canada, and Australia | Terminated | April 04, 2011 | 82 | 8 | Not_Computer
Z-1396-2011 | 57861 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Comfort Sling Plus Model 300, Part number 35300305, Size Medium, LIKO, Sweden. Liko a Hill-Rom Company. |  |  | February 09, 2011 | 2011 | Hill-Rom, Inc. | Slings were improperly assembled and the loops that are used to support the sling when hung from a lift were too long, potentially putting the patient being lifted at risk of falling out. | Pending | The firm sent an URGENT MEDICAL DEVICE RECALL notice dated 02/09/2011 to Facility Risk Manager/Facility Administrators. The letter Identified the affected product, explained the recall situation, and asked firms to check their facility for the affected serial numbers. Consignees were asked to return the recalled products, and a replacement sling would be sent to them. Customers can contact Hill-Rom Technical Support at 800-445-3720, if they have any questions concerning this matter. |  | 10 | 10 units | Worldwide Distribution -- US, Switzerland, Germany, Denmark, Finland, France, UK, Italy, Netherlands, Sweden, and Canada. | Terminated | October 07, 2011 | 240 | 10 | Not_Computer
Z-1399-2011 | 57431 | No | MKT |  |  | Microbiology | PMA | Abbott RealTime HBV Assay, List 2N40, in vitro diagnostic; Abbott Molecular Inc., Des Plaines, IL 60018 USA |  |  | December 20, 2010 | 2010 | Abbott Molecular | Internal testing identified the potential of under filled and over filled enzyme reagent vials which may result in obtaining less than the total number of tests per kit as described on the package insert. | Other | Abbott Molecular sent a Field Advisory Notice letter to all RealTime HIV-1, CT/NG and HBV assay customers on December 20, 2010, informing them that internal testing identified the potential of under filled and over filled enzyme vials which may result in obtaining less than the total number of tests per kit as described in the package insert. The customers were told that specimens that meet acceptance criteria, per the product labeling are not affected by the situation and are considered valid results, and that the issue has no impact on results for those runs that generate valid results. Any questions were directed to Abbott Technical Support at 1-800-553-7042, Option 2. |  |  |  | Nationwide Distribution | Terminated | November 19, 2012 | 700 |  | Not_Computer
Z-1432-2011 | 57084 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | 9) Signa Excite 1.5 T (Excite II) |  |  | November 09, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the gradient cables being inadvertently swapped while servicing the gradients on some GE Healthcare MR Systems that may impact patient safety.If the gradient cables are inadvertently swapped while servicing the gradients, and the Service Engineerdoes not perform the geometry check as required in the service procedure, images m | Nonconforming Material/Component | On November 9, 2010, the firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated November 5, 2010 to all Consignees/Customers. The letter described the product, problem and action to be taken by the firm. GE instructed the customers to contact their local GEHC support Engineer, if they have any questions about the image orientation, or the procedures for geometry checks on thier system. Additionally, a GE Healthcare representative will install a device, which will reduce the potiential for the gradient cables to be swapped.Contact your local GEHC Support Engineer if you have any questions concerning this notification or phone: 262-548-2731. |  | 1175 | 1175 | Worldwide distribution: USA including DC and PR; and countries including: VIET NAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKEY, TUNISIA, FORMER YUGOSLAV REPUBLIC OF MACEDONIA, THAILAND, TAIWAN, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, REPUBLIC OF KOREA, PORTUGAL POLAND, PHILIPPINES, PERU,PARAGUAY, PANAMA, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MAURITIUS, MARTINIQUE,MALTA, MALAYSIA, LUXEMBOURG,LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, KENYA, KAZAKHSTAN, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, ISLAMIC REPUBLIC OF IRAN, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUYANA, GUATEMALA, GUADELOUPE, GREECE, GERMANY, GEORGIA, FRENCH POLYNESIA, FRANCE, FINLAND, EL SALVADOR, EGYPT, ECUADOR, DOMINICAN REPUBLIC, DENMARK, CECH REPUBLIC, CYPRUS, CROATIA, COLOMBIA, CHINA, CHILE, CANDA, BULGARIA, BRAZIL, BOSNIA AND HERZEGOVINA, BOLIVIA, BELGIUM, BELARUS, AZERBAIJAN, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Terminated | July 03, 2012 | 602 | 8937 | Not_Computer
Z-1434-2011 | 57821 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cardiovascular Review Station |  |  | January 27, 2011 | 2011 | AGFA Corp. | Echocardiography volume measurements performed within CRS were not reflected correctly in associated Results Management (RM) reports. | Other | The firm, AGFA Healthcare, sent a "Safety Notice" dated January 27, 2011 to all customers. The notice described the product, problem and action to be taken by firm and customer. AGFA informed the customers that a software upgrade was made between December 2008 and February 2010 and instructed the customers to distribute this information to all individuals within their facility and to complete and return the attached URGENT SAFETY NOTICE FEEDBACK FORM via fax to 864-421-1664 or email to debbie.norris@agfa.com. Note: Service Bulletin DD+DIS100.09E, was issued March 2009 and an upgrade from Software Versions 2.12.05 through 2.14.03.SU1 to Software Version CRS 2.14.03.SU2 was required for the impacted sites. Most identified consignees/customers required the software upgrade, but some consignees/customers had already been upgraded prior to the event and did not require additional upgrades.If you have any questions about this matter, please contact the Post Market/QARA North America via email at debbie.norris@agfa.com or call at 864-421-1754. |  | 107 | 107 | Worldwide distribution: USA and country of: Canada | Terminated | March 29, 2011 | 61 | 107 | Software | Output/Calculation | Software update | Software Update
Z-1422-2011 | 57704 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE MV with Artiste Digital Linear Accelerators in combination with the Flat Panel Positioner; Part No. 8139789; |  |  | December 03, 2010 | 2010 | Siemens Medical Solutions USA, Inc., Concord, CA | The firm become aware of a potential issue with ARTISTE system that may result in patient injury should the positioner belt break, and the cam disc (secondary safety mechanism) is not installed and if the flat panel is over the patient. | Other | The firm, Siemens, issued a "Customer Information" letter on December 3, 2010 via mail to all its direct consignees/customers. The letter described the product, problem and actions taken by the firm to the customer. Siemens arranged for an inspection of their customers systems to make sure the safety mechanism is fully functional. Siemens instructed the customers to include the Customer Information letter in their ARTISTE System Owner manual chapter 'Safety Advisory Letters' .If you any questions, please call 925-246-8407. |  | 20 | 20 units | Worldwide distribution. | Terminated | June 30, 2011 | 209 | 20 | Not_Computer
Z-1423-2011 | 57629 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | HomeFill Oxygen Compressor IOH200 Homefill, The serial numbers are as follows: 10JF034081, 10JF034082, 10JF034083, 10JF034084, 10JF034085, 10JF034086, 10JF034087, 10JF034088, 10JF034089, 10JF034090, 10JF034091, 10JF034092, 10JF034093, 10JF034094, 10JF034095, 10JF034096, 10JF034097, 10JF034098, 10JF034099, 10JF034100, 10JF034101, 10JF034102, 10JF034103, 10JF034104, 10JF034105, 10JF034106, 10JF034107, 10JF034108, 10JF034109, 10JF034111, 10JF034112, 10JF034113, 10JF034114, 10JF034115, 10JF034116, 10JF034117, 10JF034119, 10JF034120, 10JF034121, 10JF034122, 10JF034123, 10JF034124, 10JF034125, 10JF034126, 10JF034127, 10JF034128, 10JF034129, 10JF034130, 10JF034131, 10JF034132, 10JF034133, 10JF034134, 10JF034135, 10JF034136, 10JF034177, 10JF034178, 10JF034179, 10JF034180, 10JF034181, 10JF034182, 10JF034183, &amp; 10JF034184. |  |  | January 11, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that inadequate clearance between a motor ground clip and the motor windings may result in an increased risk for shock to service personnel. It also may result in contact with the motor winding, resulting in a shorted winding; this condition has the potential to result in motor failure/shutdown or to cause the motor | Other | The firm, INVACARE Corp., sent an "URGENT PRODUCT RECALL: IOH 200 Homefill Recall" letter dated January 10, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to "immediately examine their inventory and quarantine product subject to recall. In addition, if they may have further distributed this product, please identify their customers and notify them at once of this product recall." The customers were also instructed to complete and return an enclosed response form. Note: All product is to be return to the recalling firm.If you have any questions, contact the Manager, Regulatory Affairs at (440) 329-6356. |  | 62 | 62 Compressors | Nationwide distribution: USA including states of: FL, GA, KY, LA, MO, NC, NE, NJ and SC. | Terminated | November 21, 2011 | 314 | 62 | Not_Computer
Z-1409-2011 | 57487 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Navigation System - iNtellect Cranial Planning Software - Enable, Part number 6000-652-300. It is part of the product series of the Stryker Navigation System. The system comprises software for pre-operative surgical planning and supports computer assisted surgery based on a wireless optical tracking localization device for the use in navigated cranial surgery. It includes automatic multi-modality correlation. The system provides an easy-to-use work flow concept for image importation, FESS, Craniotomy, Biopsy, etc. Additional, automatic segmentation of brain surface, tumor, vessels, etc. are supported. |  |  | November 10, 2010 | 2010 | Stryker Instruments Division of Stryker Corporation | When users imports more than one image set and changes the viewing direction for both images, the orientation of every other image series might be stored incorrectly, as a mirror image. This may lead to insufficient or inaccurate correlation results or incorrectly shown orientation of every second image series during navigation. | Software design | Urgent: Medical Device Recall Notification letters were sent on 12/3/2010, and software upgrades were to be performed in the field by the end of 2010. Specific Customer Instructions: 1) Locate the units listed in the notification.2) Distribute this notification to all departments affected.3) Return the enclosed business reply form to confirm receipt of this notification and identify how many affected units are currently in your inventory. 4) Once the new software is available, Stryker will dispatch a service tech in order to update the software. Note: Software upgrades for all applications will be available by the end of the year. The upgrades will be scheduled as we receive your business reply form.Note: If you currently have a Time License or Pay Per Use card, a new one will be issued to you at the time of the update.5) If you have loaned or sold any of the units listed in this notification, please, forward a copy of this notice to the new users and advise Stryker of the new users name and their new location.6) If you have disposed of any of the affected units and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.Questions regarding this recall should be directed to Stryker Field Action Team at 800-800-4236, ext 3808, 3584, or 4354 between the hours of 8 a.m. - 5 p.m. Monday - Friday. |  | 1155 | 1,155 total | Worldwide Distribution -- USA, Canada, Mexico, Australia, Argentina, Brazil, United Kingdom, Japan, Poland, S. Korea, Hong Kong, Taiwan, India, and New Zealand. | Terminated | June 01, 2011 | 203 | 1155 | Computer
Z-2809-2011 | 57829 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ***EXPANDED RECALL*** Curlin Infusion Administration Set, REF 340-4165, Sterile EO, Zevex, Inc., 4314 Zevex Park Lane, Salt Lake City, UT 84123. |  |  | January 24, 2011 | 2011 | Zevex International, Inc. | Defect in tubing may trigger false Air-In-Line warning. | Process control | The firm, MOOG, sent an "Expansion of Moog/Curlin Ambulatory Pump Administration Set Recall" letter dated May 2, 2011 to its customers. This letter informed the customers that this recall has been expanded to included specific lot numbers of Curlin model ambulatory infusion administration sets. The customers were instructed to follow the outlined process in the letter: remove suspect product code from inventory, contact MOOG Customer Service at 1-800-970-2337, prompt #7, contact your distributor to arrange replacement product, request a shipping return label as per the distributor's process and ship back to their location and follow the temporary solution outlined in the letter. Moog Customer Service will provide a call tag for convenient product return.If you have any questions contact MOOG Customer Service at 1-800-970-2337. The firm, MOOG, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 24, 2011 to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to remove suspect product code from inventory; contact their distributor to arrange replacement product, if purchased from distributor and request shipping return label per distributors process and ship back to their location or contact MOOG Customer Service at 800-970-2337, prompt #7, if purchased directly from MOOG and MOOG Customer Service will provide a call tag for convenient product return; follow the temporary solution provided until replacements are administered; and complete and return the MEDICAL DEVICE CORRECTION REPLY FORM in the self addressed envelope provided.If you have any questions, contact MOOG Customer Service at 800-970-2337. |  | 8080 | 8,080 units | Worldwide Distribution: USA including states of: AR, CA, FL, IA, KS, MS, NE, NY, OH, OK, TN, TX, and UT; and countries including: Australia, Croatia, France, Germany, Kuwait, New Zealand, Qatar, Saudi Arabia, South Africa, United Arab Emirates, and UK. | Terminated | September 27, 2011 | 246 | 3024740 | Not_Computer
Z-1435-2011 | 57820 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Results Manager/Results Manager Administration Tool |  |  | April 15, 2009 | 2009 | AGFA Corp. | User could modify a measurement's units and that change would not be reflected in echocardiography formulas used within the Adult echocardiography RM. | Other | The firm, Agfa Healthcare, sent individual letters dating from April 15, 2009 to August 26, 2009, to the customers. The letters described the product, problem and actions to be taken. Agfa provided the affected customer with a complete list of their impacted patients and the measurements. Agfa also scheduled unit corrections with their customers and received acknowledgement of the letters at the time of scheduling the unit corrections. Service Bulletin DIS026.09E was issued January 2009 to inform users that Agfa will no longer allow customers to request unit changes for measurements that have formulas. Note: Technical Support ran diagnostic scripts for each customer and documented in GSC cases for each consignee/customer.If you have any questions, contact Agfa Healthcare at (949) 830-4959. |  | 6 | 6 devices | Nationwide distribution: USA including states of: IL, MT, NY, OR and SC. | Terminated | October 13, 2011 | 911 | 6 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0893-2011 | 54768 | No | RFZ |  |  |  |  | Hemochron Signature Elite Whole Blood Microcoagulation (IVD) System. |  |  | March 31, 2010 | 2010 | International Technidyne Corp. | There is a mismatch between the class of the laser barcode scanner and the labeling of the Hemochron Signature Elite: some of the devices were inadvertently equipped with class II laser barcode scanner instead of the intended class I barcode scanner. | Labeling mix-ups | Urgent - Medical Device Recall - Hemochron Signature Elite letters, dated February 24, 2010, by Federal Express to all direct customers. The letter informed customers of the issue and stated that the problem does not negatively affect the performance of the product and to continue using the product as intended. In addition, customers are to notify the appropriate personnel about the notice, complete and return the attached form, and an ITC representative will contact customers to discuss arrangements to re-label their devices. Customers can contact ITC Technical Support at 800-579-2255 (US) or 732-548-5700 ext 4700 (International) if they have any questions regarding this issue and notice. Customers can also email sigelitescan@itcmed.com with questions. |  | 1640 | 1640 units (1303 US, 337 foreign) | Worldwide Distribution -- United Arab Emirates, Belgium, Canada, Switzerland, Germany, Denmark, France, United Kingdom, Ireland, Iceland, Italy, Kuwait, Liechtenstein, Netherlands, Oman, Qatar, Saudi Arabia, Sweden, and Singapore. | Terminated | September 25, 2012 | 909 | 1640 | Not_Computer
Z-1457-2011 | 57402 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | QuickVue One-Step hCG Urine, Quidel Catalog Number: 20109, Lot Number: 706842 |  |  | November 23, 2010 | 2010 | Quidel Corporation | The recall was initiated after Quidel Corporation determined that a small number of QuickVue One-Step hCG Urine kits have been incorrectly packaged with QuickVue H. pylori gll cassettes inside the kit boxes. These foil pouches and testcassettes are clearly labeled as H. pylori. Nevertheless, there is a risk that customers may overlook this labeling. | Pending | Quidel Corporation initiated an "Urgent Medical Device Recall" letter on November 23, 2010, with an attached "Distributor Recall Notification Form" to the affected distributor consignees notifying them of this issue and requesting return of all kits from the affected lot. Quidel is further requesting that distributor consignees notify end-users of the issue and request return of kit boxes containing the H. pylori pouched cassettes. The recall notices sent to the distributors were sent via US Postal, certified, with return receipt.The consignees were instructed to:(1) Locate kits with the affected lot number and immediately remove the kits from your existinginventory and place them into a quarantine location, (2) Contact Quidel Corporation for instructions on how to return the kits and obtain replacement product.Quidel Technical Support(800) 874-1517, Monday-Friday, 7am - 5pm PST,(3) Fax the enclosed Inventory Form to Quidel at (858) 431-3522. Indicate the total number of kits currently on hand with lot number 706842 or if you do not have any kits of this lot numberremaining in inventory. |  | 90600 | 3624 kits = 90,600 tests total | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, HI, IL, KS, KY, MA, MD, MN, NC, NJ, NM, NV, OK, OR, PA, TN, TX, UT, VA, and WA and Canada. | Terminated | June 26, 2012 | 581 | 90600 | Not_Computer
Z-1458-2011 | 57561 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The Panorama HFO system has the Traceable Item Identification (12NC): |  |  | February 05, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | In the ExamCard MobiFlex the Patient Position is "Prone" which must be "Supine". The result is mislabeling of left and right side of the patient. If this is not detected by the operator there is a risk of wrong treatment or the need of additional X-RAY during treatment. | Other | Philips began sendng out the URGENT-Device Correction Notice letter dated February 5, 2010, ro all affected consignees on February 10, 2010. The consignees were advised to change the patient position from "Prone" immediately to "Supine" and store the corrected ExamCard in the hospital database if they use the ExamCard, MobiFlex dual Injection.Consignees were instructed to call Philips Healthcare Call Center at 800-722-9377 #5 (then follow prompts) and also reference "FCO 78100318" with any communication. |  | 82 | 82 units | Nationwide Distribution including the following states: AL, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NE, NH, NJ, NY, OH, OR, SC, TN, TX, VA, VT, WI, and WV.AL, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NE, NH, NJ, NY, OH, OR, SC, TN, TX, VA, VT, WI, and WV. There was no foreign consignees. | Terminated | August 10, 2011 | 551 | 82 | Not_Computer
Z-1472-2011 | 57514 | No | NRS | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Veridex, Cell Search Circulating Tumor Cell Control Kit. Product code 7900003 |  |  | November 09, 2010 | 2010 | Veridex, LLC | There is a discrepancy between the assigned range of values on the high control label and the range listed on the Cell control Assigned Value Label for the Cell Search Circulating Tumor Cell Control Kit. | Process control | Veridex notified all affected consignees on November 11, 2010 by Federal Express overnight letter and prodivded with the informaiton regarding the label discrepancy. Foreign consignees were also notified on November 11, 2010 of the issue and instructed to notify their affected consignees. Consignees are asked to return the confirmation of receipt form by 11/29/2010. |  | 150 | 122 US, 28 International | Worldwide distribution, including: USA, Canada, France, UK, Germany, Italy, China, Spain &amp; Portugal. | Terminated | March 28, 2014 | 1235 | 150 | Not_Computer
Z-1459-2011 | 57560 | No | MKT |  |  | Microbiology | PMA | COBAS AmpliPrep/COBAS TaqMan HBV v2.0 US-IVD; |  |  | December 28, 2010 | 2010 | Roche Molecular Systems, Inc. | There is an issue regarding liquid level detection of the CS1 Magnetic Glass Particles (MGP) Reagent Cassette of the COBAS AmpliPrep /COBAS TaqMan HBV v2.0 Test on the COBAS AmpliPrep instrument. | Employee error | Roche sent Urgent Medical Device Correction Letters to their consignees on December 28, 2010 via Fed Ex. The letters advise the customers to verify the correct configuration of the CSI reagent cassette, remind customers to follow the package insert instructions for including all required controls for each batch and advise that review of previously generated results may be warranted. A UMDC Faxback form was attached for the consignee to complete.For questions regarding this recall call 908-253-7569. |  | 22212 | 93 kits to US; 22,119 kits to foreign | Worldwide Distribution - USA including CA, CT, IN, MI, MN, NC, NE, NY, OR, TX, WA, and WI and the countries of Australia, Austria, Azerbaijan, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, India, Israel, Italy, Japan, Kuwait, Luxembourg, Malaysia, Mauritania, Mongolia, Netherlands, New Zealand, Norway, Poland, Romania, Saudi Arabia, Singapore, Slovakia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, and the United Kingdom | Terminated | August 27, 2012 | 608 | 22212 | Not_Computer
Z-0286-2011 | 56922 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel/Edge/Elite Select Computed Tomography X-Ray System. |  |  | May 01, 2010 | 2010 | GE Healthcare, LLC | Failure to properly document the CTDI in the Technical Reference or User Manual. | Radiation Control for Health and Safety Act | Planned action(s) to repair defect or to bring product into compliance: GE Healthcare will bring the Technical Reference Manual into compliance by correcting the noncompliance in forward production and the installed base units. This will be accomplished by delivering an "Important Electronic Product Radiation Warning" letter and a TRM addendum which contains the corrected information to customers. This will be carried out via a Field Modification Instruction IFMI 229251 planned to be released in October, 2010 with an estimated completion date of April, 2011. The letter identifies the affected product and describes the defect and related hazards. It also asks customers to refer to the TRM addendum for the scan time accuracy, the helical scan exposure time, and the CTDI adjustment factors for 4 and 8 slices. Contact information is also provided. |  | 103 | 103 units | Nationwide Distribution | Terminated | December 05, 2013 | 1314 | 103 | Not_Computer
Z-1495-2011 | 53690 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFAversion 2.1A software, model number 426-0246, a component of the Bio-Rad PhD System. Manufactured by Bio-Rad Laboratories, France, Distributed by Bio-Rad Laboratories, Hercules, CA. |  |  | September 29, 2009 | 2009 | Bio-Rad Laboratories Inc | Incorrect patient results are being generated from out-of-range samples tested withc ompetitive enzyme immunoassays on the PhD system with software version 2.1A. | Software change control | BIO-RAD sent an URGENT: MEDICAL DEVICE CORRECTION letter dated September 29, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their customer records to identify all PhD System customers that are using version 2.1A of the EIA/IFA Software and Methods CD-ROM.Customers who are not running competitive immunoassays on the PhD System may continue using the 2.1A software.Customers who run competitive immunoassays must revert to software version 2.0B.Customers were instructed to translate if necessary and add local contact information to the customer letter and distribute to all PhD System customers that are using 2.1A of the EIA/IFA Software and Methods CD-ROM. A Customer Medical Device Correction Response Form must be collected from each customer to ensure that they have received this important communication. Once forms have been collected from all applicable customers, they should be faxed to Bio-Rad CSD Regulatory Affairs Department along with the completed Subsidiary Medical Device Correction Response Form.For questions regarding this recall call 510-741-3954. |  | 262 | 262 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, FL, IL, MA, MD, MN, MO, MS, MT, NC, NY, OH, PA, SD, TN, TX, UT, and VT and the countries of Australia, New Zealand, Singapore, China, France, Sweden, and Canada | Terminated | March 11, 2011 | 528 | 262 | Software | Output/Calculation | Revert to a previous version | Software Update
Z-1479-2011 | 56385 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio Rad brand Autoimmune EIA ANA Screening Kit, 96 Tests, |  |  | March 01, 2010 | 2010 | Bio-Rad Laboratories Inc | The firm received customer complaints of increased positivity rates associated with the ANA Screening kit due to a decline in stability. | Nonconforming Material/Component | An Urgent: Medical Device Recall Notification, dated March 1, 2010, was provided to the firm's consignee, identifying the affected product and stating the reason for the recall. The letter also asked the customer to examine their inventory and quarantine the affected kits. If the product had been further distributed, those customers must be identified, notified, and asked to discontinue use of the affected product. Customers were to discard all affected product and complete and return the enclosed Medical Device Recall Response Form. The letter also stated that replacement kits would be available in mid-March. |  | 2109 | 2,109 bulk packs | Nationwide Distribution -- New York, US. | Terminated | March 11, 2011 | 375 | 2109 | Not_Computer
Z-1477-2011 | 57427 | No | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA 125 II Assay packaged in 100 and 500 test kits as follows: |  |  | December 14, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc | Positive bias (40-55%) with patient results greater than the upper limit of normal and less than the upper limit of normal | Process control | Siemens Healthcare Diagnostics on 12 /14/10 issued Urgent Device Recall notice via Federal Express to accounts. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Device Recall notice directs customers to immediately discontinue the use of ADVIA Centaur CA 125 II assay kits that contain CA 125 II reagent lot 144. It also recommends that patient results obtained with this reagent lot be reviewed. A confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory action. Technical questions, please contact the Technical Solutions Center at 1-877-229-3711 |  | 2960 | 2960 units | Worldwide distribution, including USA, Canada, Argentina, Austria, Belgium, China, Denmark, France, Germany, Great Britain, Greece, India, Italy, Japan, Korea, Lavitia, Lithuania, Malyasia, Mexico, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Romania, Saudi Arabia, Switzerland, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Czech Republic, Taiwan, Turkey, Uganda, and Venezuela | Terminated | April 12, 2013 | 850 | 2960 | Not_Computer
Z-1478-2011 | 55293 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | DiaSorin ETI-ANA Screen Kit, 96 Test, Model Number N0093, manufactured by Bio-Rad Laboratories. |  |  | March 01, 2010 | 2010 | Bio-Rad Laboratories Inc | Inaccurate results-- Consumer complaints of increased positivity rates with the ANA screening kit. | Nonconforming Material/Component | Recall was initiated on 3/1/2010. DiaSorin S.R.I. (the sole consignee) was alerted via email using the firm's Medical Device Recall Notification. Consignee informed of the recall and instructed to dispose of the affected product according to local laws. |  | 999 | 999 kits | Foreign Only:One distributor, DiaSorin S.R.I. in Italy.No wholesalers.No retailers.No domestic distribution, No Federal Government sales/distribution centers. | Terminated | March 11, 2011 | 375 | 999 | Not_Computer
Z-1473-2011 | 57083 | [6 empty] | GE ADVANTAGE SIM MD versions 5.x, 6.x, 7.x up to and including version 7.7.0 |  |  | November 24, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware that inconsistent DRR (Digitally Reconstructed Radiographic) visualization may occur when the blended DRR feature or the Isocenter Move mode feature is used in combination with specified other functions of the AdvantageSim application on Advantage Workstation. These inconsistencies may affect patient safety. Other AdvantageSim functions are not affected. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 23, 2010. The letter was addressed to Hospital Administrators / Risk mangers, Oncology Department Managers / Radiation Oncologists, &amp; Radiation Physicists / Dosimetrists / Therapists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For any questions or concerns regarding this recall, please call one of thefollowing phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 1531 | 1531 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MI, MN, MS,MO, NT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, PAKISTAN, NOWAY, NIGERIA, NETHERLANDS, MEXICO, MALAYSIA, LITHUANIA, LATVIA, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, ECUADOR, DOMINICA REPUBLIC, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BRAZIL BELGIUM, AUSTRIA, AUSTRALIA. | Terminated | August 03, 2012 | 618 | 1531 | Other |  |  | Other
Z-1476-2011 | 57692 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN Treatment Planning System |  |  | January 28, 2008 | 2008 | Elekta, Inc. | PrecisePLAN generates digitally Reconstructed Radiographs (DRR;s) with a shift in the superior direction. | Other | Elekta sent an Important Notice A308 letter dated January 28, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Elekta representative directly if they had any questions. . |  | 277 | 277 units | Worldwide Distribution - USA and the countries of Bulgaria, Switzerland, Czech Republic, Germany, Spain, France, Hungary, Ireland, Italy, Liechtenstein, Netherlands, Poland, Portugal, Sweden, Slovak Republic, Australia, China, Hong Kong, and Japan, | Terminated | April 21, 2011 | 1179 | 277 | Not_Computer
Z-1474-2011 | 57419 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Aesculap MicroSpeed Uni Control Box, Part #GD670 |  |  | November 22, 2010 | 2010 | Aesculap Inc | Potential safety issue in that the control box can misinterpret the information sent from the low speed motor if the control magnet inside the motor changes its posthion. | Software design | AESCULAP sent an IMPORTANT CORRECTION NOTICE letter dated November 22, 2010, to all affected consignees. The letter identified the product, the problem, and the action the consignee should take. Consignees were instructed to not use the motors for thread cutting and removal of screws.If any unusual reactions of any cutting tools were observed consignees should check to ensure that the motor is not running in the opposite direction.If any error message is displayed, consigness were instructed to not use the motor.Consignees were also instructed to check both directions of running before use and do not use a motor which produces an error message in one direction. Aesculap would be providing all customers with a software update to the MicroSpeed Uni control box either by the unit being returned to Aesculap for update or by firm representative doing the upgrade on site at their facility.For questions regarding this recall call 800-258-1946 x5254 or 800-258-1946 x5267. |  | 91 | 91 units | Worldwide Distribution - USA including CA, CO, DE, FL, GA, IN, LA, MO, NC, NJ, OK, PA, TN, TX, and VA and the country of Canada. | Terminated | July 05, 2012 | 591 | 91 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1475-2011 | 57100 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Vu a-POD Intervertebral Body Fusion Device. Packaged as surgical kits containing all implants and instruments required for a procedure. Also packaged as individual instruments and implants in poly bags distributed as replenishment for surgical kits. Theken Spine, LLC, Akron, OH 44306. |  |  | October 19, 2010 | 2010 | Integra LifeSciences Corp. | The IFU provided with the device contains incorrect information in the Indications for Use Section, stating the following: The Theken Spine Vu a"POD Intervertebral Body Fusion Device is intended for use as a stand alone implant when used with the SpinPlate provided in the set. | Employee error | On 10/19/2010 the firm sent "URGENT : FIELD CORRECTION ACTION Letters to their customers. The letters identified the affected product and explained the reason for the recall. The letter also stated that the issue is not believed to have a health impact on patients. Physicians are asked to continue to exercise their clinical judgment with any periodic patient follow up. Attached to the letters were IFUs containing the correct information. Customers are to complete and return the enclosed Recall Acknowledgement Form per the instructions provided. Questions can be directed to Suzanne B. Wojcik, Sr. Dir Regulatory Affairs and Quality Assurance, at 330-835-8528. |  | 2163 | 2,163 pieces | Nationwide Distribution -- AL, AZ, CA, CT, ID, KS, MO, NH, NV, NY, OH, OK, TX &amp; UT. | Terminated | September 19, 2012 | 701 | 2163 | Not_Computer
Z-1497-2011 | 57884 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | High Viscosity Dermabond ProPen Topical Skin Adhesive 2-Octo Cyanoacrylate, Sterile, Rx only, Ethicon, Inc. |  |  | February 02, 2011 | 2011 | Ethicon, Inc. | Some units of Dermabond Topical Skin Adhesive and Dermabond ProPen Topical Skin Adhesive contain discolored product. | Nonconforming Material/Component | Ethicon sent their recall letters with business reply cards attached on February 2, 2011 via UPS overnight mail. Customers are informed of the issue and asked to return the attached business reply card. |  | 725722 | 725,722 each (total) | Worldwide distribution, including USA, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Czech Repub, Chile, China, Denmark, Dominican Repub, Egypt, Finland, France, French Guiana, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, ireland, Israel, italy, Japan, Korea, Lithuania, Luxembourg, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Palestine, Puerto Rico, Reunion, Russia, Senegal, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Ukraine, and Venezuela. | Terminated | November 07, 2012 | 644 | 1451494 | Not_Computer
Z-1494-2011 | 54620 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Male Driver 4.0mm (catalog # MDr140L) and Male Driver 3.2mm (catalog # MDr132L), an instrument used for the insertion of the Fassier-Duval Telescopic IM System intramedullary nail, for bone fixation. |  |  | January 29, 2010 | 2010 | Pega Medical Inc. | Possible breakage. Structural weakness has been detected on the Male Driver 3.2mm and 4.0mm sizes. In certain cases, this weakness could lead to a premature failure of the instrument if it is submitted to large torsional loads higher than the normal forces required to introduce the male component. | Process control | On 01/29/2010, Pega Medical began contacting their consignees by e-mail and telephone.Pega Medical followed up with their consignees by sending out the replacement components along with the explanation letter. The subject on the letter is Notice for 3.2 mm and 4.0 mm Male Drivers of the Fassier-Duval IM Telescopic System.The consignees were informed that there could be a structural weakness in some of the Male Driver instruments. This problem has been detected on the 3.2mm and 4.0mm sizes. In certain cases, this weakness could lead to a premature failure of the instrument if it is submitted to large torsional loads higher than the normal forces required to introduce the male component. Consignees can contact Pega Medical at 450-688-5144 x 242. |  | 12 | 4 units of MDr132L (3.2 mm) and 3 units of MDr140 (4.0 mm) distributed in US. | Components were distributed to four hospitals in the US in the state of GA, OK, TN, and UT. | Terminated | March 26, 2011 | 421 | 12 | Not_Computer
Z-1493-2011 | 57827 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Neotrode Wire Attached, Catalog Number: 1731-003; Neotrode II Wire Attached, Catalog Number: 1741-003; Softrace Neonatal Wire Attached, Catalog Number: 2321-003; Softrace Pediatric Wire Attached, Catalog Number: 2331-003; Softrace Limb Band, Catalog Number: 2310-003; TRU-LINK ECG Electrode (Neonatal Pre-Wired), Catalog Number: 685-0037-00; for use with electrocardiographic monitoring equipment to detect neonatal/pediatric heart action voltages. |  |  | September 23, 2010 | 2010 | ConMed Corporation | Poor signal or loss of signal. Certain recalled devices have a loose wire attachment that may lead to intermittent or poor signal quality. | Nonconforming Material/Component | ConMed Corporation distributed an Urgent Device Recall notification letter on Tru-Link ECG Electrode (Neonatal Pre-wired) dated September 23, 2010. ConMed Corp. has advised to cease the use of the devices immediately and conducted a recall to the user level. Direct consignees were asked to contact sub-accounts if recalled devices were further distributed and notified them of such recall. Recalled devices were to return to ConMed Corp. at 525 French Road, Utica, NY 13502 with completed Attachment II to indicate the amount returned. Customers with no recalled devices were asked to fill out Attachment II and faxed to 315-624-3225. Questions can be directed to Ms. Patricia Cotter, ConMed Recall Coordinator at 315-624-3237 or fax to 315-624-3225 or email to neorecall@conmed.com. |  | 9009000 | 9,009,000 Pediatric/Neonatal ECG Electrodes with Attached Wire | Worldwide distribution, including USA, Dubai, Ecuador, Egypt, Germany, Holland, Indonesia, Israel, Italy, Jamaica, Kenya, Korea, Lithuania, Malaysia, Mexico, New Zealand, Norway, Qatar, Saudi Arabia, Singapore, South Africa, Sweden, Switzerland, US Virgin Islands, Canada and Belgium. | Terminated | October 28, 2014 | 1496 | 9009000 | Not_Computer
Z-1379-2011 | 57566 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | XenMatrix Surgical Graft Sterile, Porcine Dermal Matrix |  |  | January 07, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Davol cannot guarantee that all Xenomatic product released from July to October 2010 meets FDA guidelines for endotoxin limits; and therefore, the grafts are potentially contaminated with bacteria. | Process control | Davol issued recall notification to accounts by letter dated 1/06/11 requesting recalled units be returned. contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. A fax sheet attached to the letter to be completed and fax to 1-401-825-8753. |  | 1456 | 1456 units | Nationwide | Terminated | October 31, 2013 | 1028 | 5227 | Not_Computer
Z-1518-2011 | 57552 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | LICOX CMP Tissue Oxygen Monitor; Catalog numbers REF AC31. Manufacturer GMS Gesellschaft fur medizinische, Sondentechnik mbH, Dorfstrasse 2, 24247 Mielkendorf, Germany; Distributed by Integra NeuroSciences, 311 Enterprise Drive, Plainsboro, New Jersey 08536 USA. |  |  | December 17, 2010 | 2010 | Integra LifeSciences Corp. | Certain LICOX CMP Tissue Oxygen Monitors could allow for underestimation of temperature values. | Nonconforming Material/Component | Integra Life Sciences notified their Integra Neurospecialist via email and their customers via Fed Ex overnight mail on December 17, 2010 with an Urgent: Product Recall letter. The letter identified the affected product and the reason for the recall. The letter states that an Integra Neurospecialist will contact customers to arrange for the return of their Licox monitor to the Service Center for recalibration. A loaner will be provided until the recalibration is complete. The attached "Recall Acknowledgement and Return Form" must be completed and returned to Integra. If there are any additional questions, customers should contact Gary Rabinovich at 609-936-6864. |  | 26 | 26 units | Worldwide Distribution | Terminated | November 05, 2012 | 689 | 26 | Not_Computer
Z-1522-2011 | 56660 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, ApexPro Telemetry System |  |  | January 28, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential for loss of patient monitoring when ApexPro Telemetry System is used in conjunction with a mirrored view CARESCAPE CIC Pro in combination monitoring mode (combo mode). The safety issue includes the potential for a delay in annunciation of audible alarms when used in combo mode, and may result in thedelay of treatment of a cardiac event. | Software design | GE sent an "Urgent Medical Device Correction" letter dated December 29, 2010. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Product Correction, Safety Instruction and Contact Information.Customers were instructed to contact Technical Support or their local Service Representative with any questions regarding this medical device.For any questions regarding this recall call 262-573-4122. |  | 2346 | 2346 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT,NE, NJ, NM, NY, NC, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY. and the UNITED KINGDOM, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH KOREA, RUSSIA, NORWAY, NEW ZEALAND, NETHERLANDS, MALAYSIA, JAPAN, ITALY, ISRAEL, IRELAND, GERMANY, FRANCE, FINLAND, DENMARK, CUBA, CHINA, CANADA, AUSTRIA, AUSTRALIA, POLAND. | Terminated | May 29, 2012 | 487 | 2346 | Other | Alarm/Message |  | Other
Z-1521-2011 | 57847 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vari Lase WireFiber Laser Fiber, RED 7137, Rx ONLY, Sterile EO, Vascular Solutions Inc., 6464 Sycamore Court, Minneapolis, MN 55369 |  |  | January 10, 2011 | 2011 | Vascular Solutions, Inc. | Investigations of a recent Device Experience Report has made us aware of a potential problem with our Vari-Lase WireFiber, Model 7137, manufactured with the following lot numbers: 5500898 and 550896.Recently it was reported that the endovenous laser fiber cage and guidewire tip of the WireFiber became separated. It is possible that separation could occur during patient procedure and has the p | Nonconforming Material/Component | Vascular Solutions sent an "Urgent Medical Device Recall" letter dated January 10, 2011. The letter described the issue and the product affected. Advised consignees to immediately check their inventory , place affected product in a secure location, and to complete and return the Customer Inventory Form.For questions regarding this recall call 763-656-4210. |  | 41 | 41 | Nationwide Distribution - USA including CT, MA, MN, NY, VA | Terminated | May 03, 2012 | 479 | 41 | Not_Computer
Z-1498-2011 | 57303 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida; System, Diagnostic Ultrasound |  |  | September 28, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the SSH-880CV Aplio Artida Diagnostic Ultrasound System due to software issues. | Software design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter to consignees via USPS return receipt mail. The letter contained a return reply form that is to be faxed to TAMS at 877-349-3054 as proof of notification.Customers were instructed to share the information with all users and reviewing radiologists as well as clinical engineering or biomedical group at their facility.For questions regarding this recall call 800-421-1968 or 800-521-1968. |  | 40 | 40 units | Nationwide Distribution to CA, GA, IA, IL KY, LA, MA, MN, MT, OH, OR, PA, TX, WI, and Puerto Rico. | Terminated | May 02, 2012 | 582 | 40 | Software |  | Instructions | Safety Notice/Insructions
Z-1520-2011 | 57841 | [6 empty] | VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices. The VCare IFU is packaged one per carton of eight VCare Pouched Devices. The Instructions for Use is an 8.5" x 12" 12-page booklet and is packed loose in the carton. Only the carton will have to be opened to replace the old IFU; there is no need to open the sterile packaging of the individual VCare pouches. |  |  | February 02, 2011 | 2011 | ConMed Corporation | Product manufactured after 4/11/2008, but before 8/01/2009, contains the earlier Instructions For Use, IFU (PN 17416, Rev. D) which does not contain the revisions on the addition of Warning Notes 10 and 11 and the expansion of Warning Note 6 to the VCareÂż Instructions for Use. The reason for the field correction is to provide a copy of the revised IFU to all customers who may have unexpired VcareÂż | Device Design | ConMed Corporation sent a communication to provide its direct consignees with important supplemental information concerning the Instructions for Use (IFU) for the VCAREÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator uterine manipulators. The direct consignees were instructed to place a copy of the enclosed Instructions for Use, Part No. 14071 (Rev. A), with any inventory in your possession with an expiration date prior to August 2011 (2011-08) and replace any other Instructions for Use in thier possession. A copy of the enclosed Instructions for Use is already packaged with all devices bearing later expiration dates. Any inventory in their consignees possession with an expiration date prior to January 2011 (2011-01) is expired and should not be used. If the devices were further distributed to other facilities or distributors, or their consignees are a distributor, they were instructed to forward this information as appropriate to any known users of these devices.For questions regarding this recall call 315-624-3225. --- The IFU Warning changes are as follows: A new warning statement # 6 asks the user to verify the presence of all the components of the device upon removal of the VCAREÂż device from the patient. Although the prior IFU contained warnings about verifying that all components of the device were properly removed following use during surgery, the new statement expands the earlier, briefer statement. 6. Upon removing VCARE, the surgeon should visually inspect the VCARE device, and the patient, to make sure that the entire VCARE device was properly removed and that no components or fragments of these components were retained in the patient. There are 5 parts/components to the VCARE Cervical Elevator Retractor. These are: 1) the balloon; 2) the forward "cervical" cup; 3) the back or vaginal cup; 4) the locking assembly with thumb screw; 5) the metal shaft and handle with balloon inflation valve. --- A new warning statement # 10 further addresses |  | 31028 | 31,028 VCareÂż Vaginal-Cervical Ahluwalia's Retractor-Elevator devices | Worldwide Distribution - USA including NY, TX, CA, IA, AZ, PA, IN, and KS and the countries of Austria, Jordan, Turkey, China, Portugal, Chile, Belgium, Ecuador, Brazil, Romania, Japan, Malaysia, India Switzerland, Colombia, Netherlands, Spain, Saudi Arabia, France, and Canada. | Terminated | August 30, 2012 | 575 | 31028 | Not_Computer
Z-1501-2011 | 57702 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's LAP CHOLE KIT 5 GK; Model Number: GK105 |  |  | September 10, 2010 | 2010 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of specific lot numbers of the Direct Drive Laparoscopic Clip Applier, model number CA090, due to a potential defect in the jaws, possibly resulting in jaw breakage. | Process design | Applied Medical sent an URGENT: MEDICAL DEVICE RECALL letter dated September 10, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory and complete the attached Voluntary Recall Notification Confirmation Form and fax it to 949-713-8832. Customers were asked to return all affected product.Secondary consignees were notified by Applied Medical on September 16, 2010. For any questions regarding this recall call 949-713-8041. |  | 2843 | 2,843 units for all products | Worldwide Distribution - USA including AL, AR, AZ, CA, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VT, WA, and WI and the countries of Australia, Europe, Japan, and Kuwait. | Terminated | March 09, 2011 | 180 | 2843 | Not_Computer
Z-1562-2011 | 54551 | No | JNB | Phenylalanine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Quantase Neonatal Elution Buffer, a component of Quantase Neonatal Accessory Kits (#1, 2 and 5), Model number 532-5035, included in 532-5036, 532-5037 and 532-5040, Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | January 26, 2010 | 2010 | Bio-Rad Laboratories Inc | Pressure contained within the Elution Buffer vial may expel the rubber stopper, causing the potential for injury. | Other | Direct accounts were notified by phone, followed by fax (domestic). Urgent Device Recall notification letters were sent out on January 26, 2010 to subsidiary along with a Bio-Rad customer form to be send to each customer. The letters identified the affected product and explained the reason for the recall. Customers were asked to examine their inventory and quarantine the affected lots listed. In addition, they are to complete and return the enclosed response form according to the instructions provided. Customers are contact their regional Bio-Rad office for Elution Buffer replacement. |  | 714 | 714 | Worldwide Distribution -- USA, including states of MI and MN and Bio-Rad facilities in Italy, France, Czech Republic, Canada, Argentina, Panama, United Kingdom, Hong Kong, China, Canada, Singapore, India, South Korea, Brazil, and Sweden. | Terminated | March 17, 2011 | 415 | 714 | Not_Computer
Z-1523-2011 | 56344 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE System Digital Linear Accelerator with the syngo RT Therapist version 4.1 Connected to MOSAIQ, Part No. 8168754, Distributed by Siemens Medical Solutions Oncology Care System, Concord CA |  |  | June 29, 2010 | 2010 | Siemens Medical Solutions USA, Inc | A software problem that has the potential for data loss during the transfer of treatment records to the OIS, which may not be recorded in MOSAIQ, and subsequently may lead to mistreatment. | Other | Siemens sent Urgent Medical Device Correction, Safety Advisory Notices to all affected customers on June 29, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their System Owner Manual, Chapter, "Safety Advisory Letter.For questions customers were instructed to contact their Siemens Service Engineer or Application Specialist.For questions regarding this recall call 925-602-8157. |  |  |  | Worldwide Distribution - USA including NC,OH, and the countries of Belgium, Sweden, Germany, Lebanon, Japan, Australia, and Iran. | Terminated | March 09, 2012 | 619 |  | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-1559-2011 | 57500 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Wall Mounting Plate, an accessory to the Gendex eXpert DC (REF 110-0205G1) and 765DC (REF 110-0155G1)Intraoral X-Ray Systems; Gendex Dental Systems, Des Plaines, IL USA |  |  | December 20, 2010 | 2010 | Gendex Dental Systems | The X-Ray unit may break loose from the mounting plate and fall off the wall due to defects in the mounting plate. | Nonconforming Material/Component | The firm, Gendex, sent an "URGENT MEDICAL DEVICE FIELD CORRECTION" letter dated December 20, 2010 to their customers. A revised notification letter dated January 20, 2011, was sent to the customers providing further clarification on instruction to correct the problem, as well as a template for end user notification. The letters described the product, problem and actions to be taken. The customers were instructed to contact their customers, inform them of the issue and replace the wall-mounting plate. The customers were also instructed to complete and return the attached acknowledgement form via fax to 215-997-5665 Attn: Regulatory Affairs or email to: regulatory.affairs@gendex.com and contact Gendex Customer Service at 1-800-323-8029.If you have any questions or concerns, contact the Regulatory Affairs Associate at 847-364-3290 or regulatory.affairs@gendex.com. |  | 34 | 34 wall plates | Worldwide distribution: USA and countries including: Australia, Canada and Germany. | Terminated | August 14, 2012 | 603 | 34 | Not_Computer
Z-1539-2011 | 57605 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cylindrical Bullet Nose Reamer 14.0 mm x 250 mm, REF 475816, Biomet Orthopedics 66 East Bell Drive Warsaw IN. |  |  | December 23, 2010 | 2010 | Biomet, Inc. | The firm was notified by their supplier that during manufacture of this instrument, the 12 degree cutting edge was not created. Testing found that if the cutting edge is missing, the reamer will not perform as intended. | Other | The firm, BIOMET, sent an" URGENT MEDICAL DEVICE RECALL NOTICE" dated December 23, 2010 to its customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to immediately locate, discontinue use of product and return the product to Biomet. Biomet distributors were told to notify hospital personnel of the recall if they had distributed the product to hospital customers. The Customers/Distributors were instructed to confirm receipt of the notice by calling 800-348-9500 extension 3755 and 3983 and to complete and return the Response Form included with the recall notice via fax to 574-372-1683.Questions related to this notice should be directed to 574-371-3755 or 574-372-3983, Monday through Friday, 8am to 5pm. |  | 51 | 51 all products | Worldwide distribution: USA including states of: FL, OH, PA, IN, CA, and NJ; and countries of Argentina, Canada, Mexico, Netherlands and Spain. | Terminated | October 01, 2012 | 648 | 51 | Not_Computer
Z-1548-2011 | 57563 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK12 Defibrillator/Monitor. |  |  | October 01, 2010 | 2010 | Physio Control, Inc. | L1FEPAK 12 Operating Instructions that describethe process for disconnecting the device from AC (wall) power are not always being followed. When theinstructions are not followed and the Power ON button is pushed within 2 seconds after disconnectingthe device from AC power, the service indicator will illuminate and the device may fail to charge the defibrillator and/or initiate pacing. | Device Design | Physio Control, Inc. began mailing the URGENT-MEDICAL DEVICE INFORMATION letter to their consignees on October 1, 2010. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were advised to follow the current LIFEPAK 12 Operating Instructions-Section 7.5-AC Power Adapter Operation and wait at least 2 seconds to push the Power ON button after disconnecting the device from AC power. If the instructions are not followed and the Power ON button is pushed within 2 seconds after disconnecting the device from AC power and the service indicator illuminates, it's important to be aware that turning the device off, then back on again, will reset the device and restore full functionality.The complete copy of the current LIFEPAK 12 Operating Instructions can be downloaded from the company's website at www.physio-control.com/products/ defibrillators, Select LIFEPAK 12, Select Documents.The consignees can call the firm's Technical Support Team at 1-800-442-1142, option 5, between 6:00 A.M. to 4:00 P.M. (Pacific), Monday - Friday. |  | 39 | 39, 180 units (17168 units internationally and 22012 units in US) | Worldwide Distribution - USA and the countries of ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE.Devices were distributed to the following countries:ARGENTINA, AUSTRALIA, BAHRAIN, BERMUDA, BOLIVIA, BRAZIL, BRUNEI, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, EGYPT, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PUERTO RICO, QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SRI LANKA, TAIWAN, THAILAND, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UAE, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN, and ZIMBABWE. | Terminated | April 05, 2012 | 552 | 39 | Battery | Device Operation | Advice | Safety Notice/Insructions
Z-1556-2011 | 57693 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN Treatment Planning System |  |  | February 05, 2008 | 2008 | Elekta, Inc. | When using the 'Beam Edit" function to modify a bean in an unlocked Fraction Group, there are two conditions that may cause the user to unitentionally modify a locked fraction group. | Other | The firm, Elekta, sent an "Important Notice A312" titled "Unintentional Modification of a locked fraction group in PrecisePLAN" dated February 5, 2008 to all affected customers. The notice described the product, problem and action to be taken. The customers were instructed to use care when using the "Make Opposed" function and when changing the Fraction Group of a beam. The customers were also advised to review the final plan and any locked fraction groups to make sure the planned beam arrangement is what was intended. Note: A Field Change Order 540 073 "All PrecisePLAN 2.15 users" was issued on September 30, 2008 to all customers.If you have any questions, call 770-670-2548 or +44(0)1293 654400. |  | 282 | 282 units | Worldwide distribution: USA and country of: Mexico. | Terminated | April 07, 2011 | 1157 | 282 | Not_Computer
Z-1549-2011 | 57589 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Precedence Imaging System; 16 Slice 5/8, 16 slice 3/8; |  |  | December 28, 2010 | 2010 | Philips Medical Systems | Philips Healthcare Nuclear Medicine determined that the detector arm assembly could be compromised. | Other | Urgent - Medical Device Correction notifications were sent to all domestic consignees on December 28, 2010 via Federal Express. The letter identified the affected products, the reason for the recall, the hazard involved, and actions to be taken by the customer/user. Customers have been instructed to immediately stop using the system and contact a Philips service representative for an immediate inspection if one of the listed reasons apply. Customers were also instructed to follow the instructions in the system's Instructions For Use to avoid a collision or force on the detector. The firm's service will contact customers to arrange a time for Philips to conduct an inspection and to conduct the appropriate field safety correction. The letter should be placed in customers' Instructions For use. Notices to international consignees will be sent as soon as translation is complete. If customers need further information or support concerning this issue, they should contact their local Philips representative. |  | 540 | 540 units, all varieties | Worldwide Distribution | Terminated | April 02, 2012 | 461 | 540 | Not_Computer
Z-1555-2011 | 57688 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop Pro 4.2, 5.0, 5.1, 6.0, 6.2 |  |  | August 23, 2007 | 2007 | Elekta, Inc. | Under certain conditions it has been found that movement of the collimators can be induced during multi segment beams such as IMRT and Omniwedge which will not be detected by the Desktop Pro if high values are sent within the iCom tolerance table. | Other | The firm, Elekta, sent an "Important Notice A302" dated August 23, 2007 titled "Update to iCom Tables for Desktop Pro when used with Third Party R&V; Systems",to all affected customers. On September 28, 2007, Elekta issued a second Important Notice A304 which reiterated the same information as was in Important Notice A302. The notices described the product, problem and actions to be taken. This notice provides a workaround on how to avoid the issue. In order to avoid this issues documented, Elekta recommended that the customers set the tolerance for collimators in the iCom tolerance table to be 3.0 degrees-no less than 3.0 and to use the Linac Record, if configured, as a means of checking the final collimator position. Elekta noted that the issue will be resolved in a future release of Desktop Pro.If you any questions related to this issue, please contact your local Elekta, third party R&V; system representative or Elekta Limited at +44 (0)1293 654200. |  | 229 | 229 units | Worldwide distribution: USA including Puerto Rico and countries including: Canada and Mexico. | Terminated | April 07, 2011 | 1323 | 229 | Not_Computer
Z-1525-2011 | 57854 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TI ANCHOR 2.5MM W / 2-0 SUTURE, G-FORCE (R) SUTURE ANCHOR, Tissue Anchor, REF 86TA-2520, Rx ONLY, Sterile EO, Wright, Arlington, TN USA |  |  | January 11, 2011 | 2011 | Wright Medical Technology Inc | Driver may break during use. | Other | Wright Medical Technology Inc initiated their recall of the product by telephone to their distributors on January 11, 2011. The firm followed with letters sent on February 1, 2011, to distributors, hospital administrators, and surgeons. Distributors were asked to immediately complete the attached response form, confirming receipt of the notice, and return a copy to Wright Medical by fax to 901-867-4788. For questions they were instructed to call 800-874-5630.For questions regarding this recall call 901-867-4788. |  | 45 | 45 units | Nationwide Distribution including FL, NC, OH, IL, IN | Terminated | July 11, 2011 | 181 | 55 | Not_Computer
Z-1526-2011 | 57851 | [6 empty] | Fresenius 2008T Hemodialysis Machine |  |  | January 19, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Cursor on the dialysis screen may disappear and not respond to user input | Device Design | Fresenius Medical Care contacted customers by mailing a Customer Notification letter "Urgent correction FMNCA Model 2008T Hemodialysis Machine" via Certified-Mail, Return Receipt requested on January 19, 2011. The Technical Services department began contacting consignees by telephone on January 13, 2011, to schedule an appointment for installing the electronic board on the 2008T machines at the customers' facilities. For questions call Technical services 1-800-227-2572. |  | 1268 | 1268 units | Worldwide Distribution - USA and Canada | Terminated | July 08, 2013 | 901 | 1268 | Other |  |  | Other
Z-1557-2011 | 58008 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | AquaRefresh, Multi-Purpose Solution, 2 fl.oz/60mL, Made in the Netherlands, Distributed by Marietta Vision, 397 N. Sessions Street, Marietta, GA 30060. |  |  | February 21, 2011 | 2011 | Marietta Vision | Product is misbranded; Product labeled as, Made in the Netherlands, was actually manufactured in Mexico. | Other | John Patterson, President of Marietta Vision called each optical retailer on February 21, 2011, and told them of the misbranding issue in regards to country of origin. In addition, an Urgent Product Recall letter was sent to each account on February 23, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their AquaRefresh inventory for lot number KD005, and discontinue distribution of the samples immediately and to return any outstanding stock to Marietta Vision for exchange. Marietta Vision would return the misbranded product back to the Mexican factory at their expense to have the product packaging corrected. If they elect not to do so, Marietta Vision would have the product disposed of.Customers were instructed to fax the enclosed form back to Marietta Vision at 678-797-1099.For questions regarding this recall call 866-300-6257. |  | 1000 | 500/2 fl.oz. Bottles | Nationwide Distribution including AR, AZ, CA and TX. | Terminated | November 03, 2011 | 255 | 1000 | Not_Computer
Z-1561-2011 | 57974 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Pulmonary Hemi-Artery SG |  |  | February 17, 2011 | 2011 | CryoLife, Inc. | Staphylococcus aureus (not methicillin resistant) identified in companion tissue. | Other | CryoLife Inc. issued a letter to the physician on February 17, 2011, advising him of the results of the testing performed on a companion cardiac allograft. The letter did not recommend explant of the allograft or special antimicrobial treatment of the patient unless medically necessary.The physician was asked to complete and return the enclosed, stamped, self-addressed postcard affirming the receipt of the notification.For any questions the physician was advised to call 770-419-3355. |  | 1 | One | Nationwide Distribution including AZ &amp; NV | Terminated | April 24, 2012 | 432 | 1 | Not_Computer
Z-1554-2011 | 57988 | No | OIG | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Medline Industries, Inc. "Cedar" brand Vinyl Economy Exam Glove, Powder-Free, Large Size, Catalog Number: MSV403P, Lot Number: JCU04958569 |  |  | February 03, 2011 | 2011 | Medline Industries Inc | The recall was initiated because the affected lot, lot number JCU004958569, of gloves did not meet the requirements for the water-leak test. | Release of Material/Component prior to receiving test results | MEDLINE INDUSTRIES, INC. sent a RECALL NOTICE IMMEDIATE ACTION REQUIRED letter by US Mail on February 10, 2011, to all affected cusomters. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to:1. Immediately quarantine ALL inventory of the affected lot for return to Medline. 2. Complete the attached recall verification form. Please complete this form EVEN IF you have none of the affected lot in stock.3. Fax or email the completed form to Medline. The email address is Recalls@Medline.com or faxed the form to 866-767-1290. Customers with questions regarding the recall were instructed to contact Medline at 1-866-359-1704 or via e-mail at Recalls@Medline.com. |  | 696 | 696 cases (696,000 units) | Nationwide Distribution including CA, HI, and NV. | Terminated | March 19, 2012 | 410 | 2796 | Not_Computer
Z-1558-2011 | 57696 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Polaris Spectra Position Sensor Unit, a component of the StealthStation S7 System, Medtronic Navigation, Louisville, CO 80027. |  |  | December 22, 2010 | 2010 | Medtronic Navigation, Inc | Position Sensor Unit may generate tracking errors or may stop tracking immediately. | Component design/selection | The firm, Medtronic, sent a "Product Correction Notification" letter dated December 22, 2010, to all US Consignees/Customers on December 23, 2010 and foreign Consignees/Customers on December 27, 2010. The letter described the product, problem and actions to be taken. The customer were instructed to discontinue use of the impacted navigation system(s) until a replacement camera (PSU) is installed. A Medtronic Navigation Representative will visit each facility to remove the navigation camera and install a replacement. If you have any questions, contact your Medtronic Navigation Representative or Medtronic Navigation Technical Support services at 1-800-595-9709 or +1-720-890-3200. |  | 11 | 11 (US), 8 (Intnl) | Worldwide distribution: USA including states of: CA, CO, GA, SC, TN, and WI; and countries including: Australia, Canada, Chile, Russia, Saudi Arabia, and South Africa. | Terminated | March 11, 2011 | 79 | 11 | I/O
Z-0905-2011 | 58077 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | BuckyDiagnost general purpose x-ray system with Optimus Generator. |  |  | January 17, 2011 | 2011 | Philips Healthcare Inc. | Affected systems were determined not to prevent run-on of a selected technique factor if the technique selector and the exposure control were activated simultaneously. | Radiation Control for Health and Safety Act | A Customer Information BuckyDiagnost Diagnostic X-Ray System letter, dated January 6, 2011, was sent to customers. The letter identifies the affected product, the reason for the recall, actions that are to be taken by customers, and actions that are planned by the firm. Customers are advised to make sure that the desired exposure values are displayed and that they have released the technique button BEFORE they start the exposure by pressing the exposure switch whenever manually changing the exposure parameters and in order to avoid the problem. Philips will issue a Field Change Order (FCO) to resolve the issue by updating the system software free of charge. Philips will contact customers to implement the FCO. Further information or support concerning this issue can be obtained by contacting your local Philips representative. The Field Service Engineers began software implementation the week of March 21, 2011. |  | 1810 | 1810 units in the US | Worldwide Distribution | Completed |  |  | 1810 | Software | Device Operation | Software update | Software Update
Z-1568-2011 | 56120 | No | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Polycarbonate prescription single vision lenses for use in sport goggles. |  |  | June 30, 2010 | 2010 | Luxottica Optical Manufacturing | The polycarbonate prescription lenses may not meet the frame manufacturer's minimum lens thickness specification of 2.0 millimeters. LNAD made the decision to remove these sport goggles from the market because of the potential risk that the lenses could break or disengage from the frame in the event of an impact. | Other | Luxottica Optical Manufacturing sent notification letters on June 30, 2010, to their customers informing them of the problem. Customers were instructed to complete the enclosed Sport Goggle Return Form, and to indicate whether they would like a replacement pair of goggles or a refund.Place the completed Sport Goggle Return Form and their sport goggles (without the sport goggle case) in the envelope provided and return it to the firm via the United States Postal Service. Customers were instructed not to return to the retail store where purchased. If customers were unable to return their sport goggles because they were lost or discarded, customers should complete the appropriate information on the Sport Goggle Return Form, being sure to indicate whether they would like a replacement pair of goggles or a refund, and then sign the form. Place the completed Sport Goggle Return Form in the envelope provided and return it to the firm via the United States Postal Service. For any questions, call 1-877-488-2407. They instructed their customers to return the product for replacement or refund. |  | 2916 | 2,916 individual lenses (1,458 pair of sport goggles) | Worldwide Distribution - USA including the following states: AK, AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV; and the country of Canada. | Terminated | September 19, 2012 | 812 | 2916 | Not_Computer
Z-1592-2011 | 56994 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, dental implant planning system. |  |  | September 30, 2010 | 2010 | Straumann Manufacturing Inc. | Software: Surgical Error may occur when the implant length is changed after positioning the surgical sleeve. | Software design | Straumann notified accounts of this correction by letter datedSeptember 30, 2010 and delivered via Federal Express Delivery Service. The upgrade is available on-line by logging into the coDiagnostiX system and enabling the automatic upgrade to occur. |  | 20 | 20 systems | Nationwide | Terminated | September 19, 2011 | 354 | 20 | Software | Device Operation | Software update | Software Update
Z-1594-2011 | 57915 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ARCOS Modular Revision System Torque Limiting T-Handle 55 inch- pounds, REF 31-301850, Biomet Orthopedics Warsaw, IN. |  |  | January 28, 2011 | 2011 | Biomet, Inc. | The firm received two complaints of the tabs on the Torque Limiting T-Handle fracturing during surgery. The firm performed internal testing and found that the tabs may fracture when used in the "OFF" or "locked" setting for hand-reaming the femoral canal. | Device Design | Biomet, Inc. made telephone calls to distributor accounts on January 28, 2011, with specific instructions that the Arcos Torque Limiting T-Handle was not to be used in the "OFF" position. Distributors were required to use the Arcos Torque Limiting T-Handle on the "55" setting for insertion of locking screws and troch bolt. Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE letter dated February 21, 2011, to their distributors. The letter stated that the firm was currently manufacturing replacement handles which would be distributed to the field shortly. Consignees were advised to immediately discontinue use of the T-Handle in the "OFF" position.Send back the current T-Handle as soon as they receive a replacement.Carefully follow the instructions on the enclosed "FAX Back Response Form."Fax a copy of the Response Form to 574-372-1683 prior to return of product.Use priority carrier for shipment.If the product has been further distributed, consignees must notify hospital personnel of this action, via the enclosed "Dear Bioment Customer" notice. This letter must be give to hospital personnel responsible for receiving recall notices. Consignees are charged with the location and return of these products. Consignees were instructed to confirm receipt of this notice by calling, 800-348-9500, ext. 3755 or 3983. For any questions call 574-371-3755 or 574-372-3983. |  | 40 | 40 units | Worldwide Distribution - USA including the states of CA, NY, FL, VA, UT, IL, MI, KY, OH, MA, PA, SD, MO, TX, NV, Puerto Rico, and the countries of Australia, New Zealand, Canada. | Terminated | July 27, 2011 | 180 | 40 | Not_Computer
Z-1572-2011 | 57732 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | January 11, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The potential for injury to a patient should the treatment data be incorrectly uploaded to patient records. | Software design (manufacturing process) | Sent an Urgent Medical Device Correction Safety Advisory Notice to customers on 1/11/11. Letters will be delivered by Siemens Customer Service Group or sent by certified mail. Letter identified the affected product and stated there is a potential risk due to a problem identified by the manufacturer. Customers are asked to refer to the attached Safety Notice from Elekta IMPAC Software. A Siemens Service Engineer will contact customers site to arrange for a patient matching criteria audit to determine if you have a problem. Customers should include the Safety Advisory Notice including the Safety Notice from Elekta IMPAC Software in their product related documentation until further notice. If customers have any questions, they should contact their responsible SIEMENS Service Engineer or Application Specialist. |  | 10 | 10 units | International Only Distribution -- Germany and Japan | Terminated | April 19, 2012 | 464 | 10 | Other | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-1593-2011 | 57519 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Hemashield Finesse Ultra-Thin Knitted Cardiovascular Patch |  |  | December 07, 2010 | 2010 | Maquet Cardiovascular, LLC | One unit of Hemashield Finesse Ultra-thin Knitted Cardiovascular Patch which did not meet the internal release criteria for water permeability was distributed. | Other | The firm, Maquet, contacted their sole consignee/customer via telephone informing the customer of the recall and instructing the customer to put the unit to the side. On December 8, 2010 a sales representative visited the site to ensure the unit had been put to the side and to ensure the unit was shipped back to Maquet. If you have any questions, call the Manager, Regulator Affairs at 973-709-7487. |  | 1 | one unit | Nationwide distribution: SC only. | Terminated | September 12, 2012 | 645 | 1 | Not_Computer
Z-1591-2011 | 57432 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Anchor Tissue Retrieval System; ETO sterilized, single use disposable device; port size 15 mm, 1550 mL capacity; 3 units per box; Manufacturer: Anchor Products Co., 52 Official Rd, Addison, IL 60101 U.S.A.; Ref TRS175SB |  |  | December 09, 2010 | 2010 | Anchor Products Company, Inc. | The metal component within the unit can break and cause a malfunction. Use of the device may represent a potential health hazard if the metal component is not removed after breakage. | Device Design | The firm, Anchor, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 8, 2010 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any remaining units to their distributor for replacement; and complete and return the attached ANCHOR RECALL RETURN RESPONSE FORM via fax to (314) 961-4535 or mail to: Progressive Medical, Inc., Attn: Marketing Manager, 11085 Gravois Industrial Court, St. Louis, MO 63128, even if they have no inventory. A representative of Anchor Products then visited each of the direct accounts to help them implement their subrecall of the devices from their customers.For additional information please contact: Vice President, Anchor Products Company at 1-800-543-9124 or 1-630-543-9124; and the distributor at 1-800-969-6331 or 1-314-961-4535. |  | 2256 | 2,256 units | Worldwide distribution: USA including state of: Missouri; and countries including: Australia, Belgium, Canada, France, Greece, Ireland, Japan, Spain and the United Kingdom. | Terminated | January 20, 2012 | 407 | 91041 | Not_Computer
Z-1580-2011 | 57396 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PolyAx Locking Plate System, POLYAX FEMORAL PLATE - RIGHT, 15 hole, Sterile, REF 8141-30-115, Depuy Orthopaedics, Inc., Warsaw, IN. |  |  | November 01, 2010 | 2010 | Depuy Orthopaedics, Inc. | The firm determined that the 8.0 mm hole threads in the PolyAx Distal Femoral Locking Plate were incorrectly manufactured, which prevents the mating screw from locking in place. The screw should lock with the plate, instead the screw passes through the plate. | Other | Urgent Information - Recall Notice letters, dated November 1, 2010 were sent to customers. The letters described the reason for recall and identified the product affected along with the affected units. The letters also provided clinical guidance and stated that the problem had been corrected and that DePuy was in the process of replenishing its inventory. Hospitals were to determine if any of the recalled components they received were still on hand, and return them immediately to their DePuy Representative or directly to DePuy for credit. |  | 246 | 246 all products | Nationwide Distribution -- ME, MA, FL, VA, CO, TX, WY, AZ, MO, WI, GA, PA, AL, NY, OH, NJ, NC, and CA. | Terminated | June 27, 2012 | 604 | 246 | Not_Computer
Z-1569-2011 | 57314 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | MONOJECT Blood Collection Tubes containing heparin |  |  | November 15, 2010 | 2010 | Covidien LP | The heparin sodium USP used to manufacture blood collection tubes was produced with crude heparin containing Over-Sulfated Chondroitin Sulfate. | Material/Component Contamination | Covidien sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using the product, and return it to the firm. Distributors were instructed to notify their customers. All customers were required to respond to the notice regardless of their remaining inventory. Responses will be tracked for follow-up and effectiveness checks. For Customers Service questions regarding this recall call 1-800-962-9888 option 1. |  | 18100 | 18,100 | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, KS, KY, MA, MI, MN, MO, MS, NC, ND, NE, NJ, NM, MV, MY, OH, OK, PA, SC, SD, TN, TX, WA, and WI and the country of Canada | Terminated | December 04, 2015 | 1845 | 176100 | Not_Computer
Z-1589-2011 | 57965 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Biomet INNERVUE DISPOSABLE SCOPE KIT 1.2MM WITH CANNULA SET, PRODUCT NUMBER 922141, STERILE |  |  | February 10, 2011 | 2011 | Biomet, Inc. | The cannula could not be connected to the scope and scopes disengage during use. The outer hub of Innervue scopes may become detached from the inner hub of the scopes during assembly of the scope to the Innervue hand piece. This condition may prevent the Innervue scope from being securely attached to the Innervue hand piece. | Process design | On 2/10/2011, Biomet issued an "URGENT MEDICAL DEVICE NOTICE" to their Direct Accounts instructing customers to identify and discontinue use of the recalled products and return them to Biomet. Response forms were provided. |  | 264 | 264 | Distribution USA nationwide and to the Netherlands | Terminated | August 27, 2013 | 929 | 607 | Not_Computer
Z-1566-2011 | 57969 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 100/120V System 1 Base, REF 801763, Terumo Cardiovascular Systems Corp., Ann Arbor, MI. |  |  | February 14, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) has received 19 reports of a circuit breaker for the Terumo Advanced Perfusion System 1 tripping and causing the system to switch to battery back-up power. In two cases, the incident was caused when the circuit breaker came into incidental contact with the base of an IV pole; in others there was no known cause. Terumo CVS' investigation has determined that | Device Design | Terumo sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 14, 2011, addressed to Chief of Perfusion, Department of Cardiovascular Surgery or Director of Operationd Room Services. The letter explains the problem with the circuit breakers and hazards of the problem. The letter states that a service technician will visit the facility to install a guard. The service will be performed at the next scheduled preventative maintenance visit for those consignees with a service contract. For customers without a service contract, Terumo CVS will contact the consignees to schedule an appointment for the correction.Customers were instructed to review the Medical Device Correction notice.Assure that all users are aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions or concerns customers were instructed to contact Terumo CVS Customer Service at 1-800-521-2818. |  | 728 | 728 devices. | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, ME, MI, MO, MT, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV and the countries ofCANADA, COLUMBIA, JAPAN, RUSSIA, EGYPT, TURKEY, PAKISTAN, KUWAIT, COSTA RICA, TAIWAN, SAUDI ARABIA, INDIA, SOUTH KOREA, JORDAN, THE PHILLIPINES, INDONESIA, HONDURAS, DOMINICAN REPUBLIC, REPUBLIC OF GEORGIA, HONG KONG, UNITED ARAB EMIRATES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, GERMANY, BELGIUM, MEXICO, THAILAND, VEITNAM, SOUTH AFRICA | Terminated | March 20, 2012 | 400 | 1113 | Hardware | Device Operation | Install a guard | Other
Z-1577-2011 | 57885 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needle, Soft Tissue Biopsy Needle, NAC-1825M, 18ga (1.3mm) x 25cm (10), Remington Medical, Inc. |  |  | February 07, 2011 | 2011 | Remington Medical Inc. | There is a potential for the needle hubs to break during use. | Device Design | Remington Medical consignees were notified via facsimile, email and certified mail starting 02/07/2011. Distributors are requested to communicate this recall to their customers. If a customer does not respond to the recall notice within one week of the recall notification, the customer will be contacted by telephone. If this is unsuccessful, the customer will be sent a second Recall Notice letter and Customer Inventory Reconcile Fax by certified mail by 28 February 2011. If there is no response at that point, RMI will consider the customer's inventory depleted. All return product for this recall will follow RMIs Returns Goods Authorization procedure (SOP7-200). All product returns and customer responses will be documented. Upon completion of the recall, the status will be updated and submitted for the FDA's recall closure process. |  | 8161 | 8161 units | U.S. Nationwide | Terminated | August 12, 2011 | 186 | 46878 | Not_Computer
Z-1565-2011 | 57944 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | VersaCare P500 Microclimate Management Surface Upgrade Kits P3218A and P3219A Hill-rom Batesville, IN. |  |  | February 21, 2011 | 2011 | Hill-Rom, Inc. | Hill-Rom has become aware of a potential issue with some TotalCareÂż and VersaCareÂż beds equipped with the integrated P500 Microclimate Management(MCM) therapy surface, and with the VersaCareÂż P500 MCM surface upgrade. Because of an unauthorized vendor part modification, it is possible that the MCM blower may not perform as originally designed. Given this modification, the MCMblower may experi | Other | Hill-Rom sent an "IMPORTANT MEDICAL DEVICE CORRECTION" notice dated February 21, 2011, to Facility Risk Manager/Facility Administrators. The letter states that because of an unauthorized vendor part modification, it is possible that the MCM blower may not perform as originally designed. Given this modification, the MCM blower may experience the following: delayed start, slow running, or failure to operate.This issue does not pose a hazard to patients as long as normal patient monitoring is in effect in relation to managing skin temperature and moisture. Customers were instructed to continue to use the product as they normally would. The mattress system would provide an alarm in the event the MCM blower does not operate properly. A Hill-Rom representative would be contacting their facility to make arrangements for correcting the affected MCM blowers. If customers had any questions concerning this request or the procedure as outlined, they were instructed to contact Hill-Rom Technical Support at 800-445-3720.For questions regarding this recall call 812-934-7777. |  | 740 | 740 units | Worldwide Distribution - USA including AL, CA, CO, CT, DC, DE, FL, HI, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TN, TX, VA, WA, and WI, and the countries of ARUBA, AUSTRIA, CANADA, FRANCE, GERMANY, LEBANON, PORTUGAL and SAUDI ARABIA | Terminated | August 06, 2012 | 532 | 2465 | Not_Computer
Z-1605-2011 | 57611 | No | IXL | Spot-film device. | Radiology | Radiology | 510(K) Exempt | Philips BV Pulsera, Rel. 2.3, 718095. |  |  | December 14, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | The fixing strap of the belt holding the CBX Laser to the Image Intensifier may be detached during a procedure when used with the BV Libra and the BV Pulsera/Endura Systems. | Device Design | On 01/06/11, Philips sent an URGENT Device Correction Notice, dated 01/05/2011, to their consignees. The consignees were informed that the fixing strap of the Laser Aiming Device Image Intensifier may inadvertently detach when used with BV Libra, BV Endura Rel. 2, and BV Pulsera Rel 2.3 Mobile X-Ray systems. Specifically the draw latch of the fixing strap holding the CBX Laser to the Image Intensified can be accidentally released during a procedure when used with the above Mobile X-Ray systems. While waiting for Philips to implement their corrective action, the consignees are advised to fix the closed latch body of the draw latch to the metal band using a Tie Wrap. Without securing the latch with a Tie Wrap, the device should not be used.The consignees can contact Philips Call Center at 1-800-722-9377, select 5, and reference FCO 71800035 with any questions about this recall. |  | 214 | 214 units | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and Washington, DC. | Terminated | February 28, 2012 | 441 | 227 | Not_Computer
Z-1606-2011 | 56427 | No | OAW |  |  | Immunology | HDE - Humanitarian Device Exemption | MESOMARK Humanitarian Device. Authorized by Federal Law for use as an aid in monitoring of patients diagnosed with mesothelioma. Each kit contains 96 tests. |  |  | July 29, 2010 | 2010 | Fujirebio Diagnostics, Inc. | presence of material attached to the bottom of some wells in places visible to the naked eye, which caused an overestimation of the sample. | Other | Notification dated 7/7/10 was sent to the customer via e-mail and FedEx Priority Overnight.FUJIREBIO Diagnostics, Inc. sent an URGENT PRODUCT CORRECTION NOTIIFICATION letter dated July 7, 2010, to all affected customers via e-mail and FedEx Priority Overnight. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to immediately discontinue using MESOMARK kit. Lot 57K04510 and discard any remaining inventory.Complete and return the attached Confirmation of Receipt form upon receipt of this notification, no later than July 30, 2010.Notify your ordering physician or health care provider as appropriate.For questions regarding this recall call 877-861-7246 or 610-240-3800. |  | 42 | 42 kits | Nationwide Distribution including UT | Terminated | July 05, 2012 | 707 | 42 | Not_Computer
Z-1596-2011 | 57892 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Power cord use on Symbiq One-Channel infuser and |  |  | October 27, 2010 | 2010 | Hospira Inc | Complaints of broken, bent or missing prongs, charring, sparks, visible smoke, burnt smell have been reported on Symbiq AC Power cords. | Component design/selection | Hospira Recall notification letters were sent by Federal Express on November 1, 2010. |  | 32376 | 32,376 power cords | Nationwide distribution and to Australia, Canada, and Malaysia. | Terminated | November 07, 2013 | 1107 | 32376 | Not_Computer
Z-1598-2011 | 57697 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Plum XL Micro/Macro Infusion Pump, Single Channel, with a 0.1-999 mL/hr flow rate range; list number 11846 |  |  | February 14, 2011 | 2011 | Hospira Inc. | Hospira has received reports of the Plum XL, XLM and XLD Infusion Pumps with no audible alarm at the low audio level setting. | Labeling design | Hospira sent Urgent Device Recall letters dated February 14, 2011 to their customers on the same date, informing them that Hospira had received complaints of the Plum XL, XLM and XLD infusion pumps in which the audible alarm failed. If the audible alarm fails and the user does not notice the visual alert, the user may not be aware of the change in pump status such as air-in-line or occlusion. This may result in a delay or interruption of therapy which may result in serious injury and/or death. Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. They were also instructed to ensure that all personnel in their facility are following Hospira's Technical Service Manual for routine decontamination of the pumps and cleaning of the buzzers during preventative maintenance, including inspection of the buzzers for leg lifting. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance.The accounts were also requested to complete the attached reply form and return it to Hospira via fax at 1-866-382-4228. |  | 23385 | 23,385 pumps | Worldwide distribution, including United States, the U.S. Virgin Islands, Puerto Rico, Argentina, Australia, Bahamas, Bahrain, Barbados, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, EI Salvador, Finland, France, Greece, Guatemala, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Luxembourg, Mexico, New Zealand, Oman, Panama, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, San Marino, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Venezuela. | Open, Classified |  |  | 205689 | Not_Computer
Z-1607-2011 | 57759 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex Free NeoGuard Transducer Cover, Ref 676-117, Sterile (2.6 x 30cm) cover, packaged 24/box. The firm name on the label is CIVCO, Kalona, IA. |  |  | January 21, 2011 | 2011 | Civco Medical Instruments Inc | Non-sterile transducer covers were incorrectly labeled as sterile and contained incorrect temperature storage information. | Labeling design | The recalling firm notified their customers via phone and e-mail letter entitled Urgent - Medical Device Correction, dated 1/21/11, explaining the reason for recall. The letter stated that the firm will replace the affected product and requested that customers dispose of the product from the affected lots. A response form was enclosed to report final disposition of the product and should be returned to the firm. If there are any questions, customers can call their customer service representative at 800-445-6741 between 8 am and 5 pm CST. |  | 31 | 31/24-cover boxes | Nationwide Distribution -- WA, CA, and MO. | Terminated | July 09, 2013 | 900 | 31 | Not_Computer
Z-1471-2011 | 57767 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Dressing Change Tray, DT5880A, Sterile, Centurion Medical Products Corp., Williamston, MI. |  |  | January 14, 2011 | 2011 | Centurion Medical Products Corporation | The firm is conducting a subrecall of kits containing Triad Groups alcohol prep pads. Triad initiated a recall of prep pads due to a concern about potential contamination of the products with an objectional orgainism, Bacillus cereus. | Nonconforming Material/Component | The firm, Centurion Medical Products, sent an "URGENT PRODUCT RECALL NOTICE" dated January 12, 2011 to affected customers via Certified Mail, Return Receipt Requested. The notice described the product, problem and action to be taken. The customers were instructed to identify and remove all inventory of the recalled kits in their possession; forward a copy of this notice to any customers that this product was further distributed to, and complete and return the attached Accountability Record form via fax to 517-546-3356 or email to: lcarpenter@centurionmp.com. The notice states that a Centurion representative will contact them regarding retrieval of the product from the facility.Should your require additional information, please contact the Director of Quality Assurance &amp; Regulatory Affairs at (517) 546-5400 ext. 1135. |  | 68056 | 68056 all kits | Nationwide distribution: USA including state of: CT, ID, NM, MA, MO, MT, VA and SC. | Terminated | June 21, 2011 | 158 | 68056 | Not_Computer
Z-1595-2011 | 57819 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | AMPLATZER Sizing Balloon II, REF 9-SB-034, Sterile EO, AGA Medical Corporation, 5050 Nathan Lane North, Plymouth, MN 55442. |  |  | December 17, 2010 | 2010 | AGA Medical Corporation | AGA Medical is initiating a voluntary recall of a specific lot of mislabeled AMPLATZER Sizing Balloon II. AGA Medical received 2 reports from customers who received sizing balloons that were labeled as 34 mm; however, the package contained 24 mm devices. The proper sizing (24 mm) was printed on the device hub visible through the product package. | Employee error | AGA Medical Corporation sent an Urgent Medical Device Voluntary Recall letter, dated December 17, 2010, via e-mail to all affected customers. A hard copy was sent via Fed Ex to customers on December 20, 2010. The letter identified the product, described the issue, and the action that should be taken by the customer. Customers were requested to cease distribution and use of the devices, check inventory for unused devices, and complete and return a response form to facilitate return of the affected devices to AGA.For questions regarding this recall call 1-763-531-3175. |  | 54 | 54 | Worldwide Distribution - USA including AZ, CA, KY, PA, SC, TX and OUS to include: Argentina, Germany, Greece, Netherlands, Peru, Poland, Portugal, Slovakia, United Kingdom | Terminated | March 29, 2012 | 468 | 54 | Not_Computer
Z-1609-2011 | 57846 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Defibrillator Monitor (M3535A/M3536A) with software (SW) versions F.01.00 or F.01.01 and with all of the following hardware options EtCO2, IP and Temp. |  |  | February 07, 2011 | 2011 | Philips Healthcare Inc. | Software: EtCO2 and Ventilation Rate values may be labeled incorrectly. | Software design | Philips Healthcare representatives issued an Urgent Medical Device letter on 2/7/11. Customers are asked to follow the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notification. In this section, customers are given the following instructions: Philips software (SW) version F.01.03 (for M3535A and M3536A) and R.01.00 (for M3536A, English only) corrects this problem.If you have affected product a Philips Service Engineer will contact you to install software version F.01.03 or R.01.00. Until your software is updated your HeartStart MRx monitor/defibrillator can remain in service. If you are using the device in one of the use modes listed under problem description, make sure that operators are aware that even though the EtCO2 and ventilation rate values can display interchangeably (unexpectedly or randomly switch) under either the ventilation rate label or the EtCO2 label, the values themselves can be identified by their color (for example if the ventilation rate label is white, the ventilation value is displayed as white). Customers are to contact their local Philips representative at 1-800-722-9377, if they need further information or support regarding this issue. |  | 2382 | 2382 units | Worldwide -- USA, Canada, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA,FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KENYA, LATVIA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, REUNION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | March 16, 2017 | 2229 | 2382 | Software | Output/Calculation | Software update | Software Update
Z-1610-2011 | 57936 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative/Surgical System (57mm), 2 1/4 inch |  |  | February 01, 2011 | 2011 | Convatec Inc. | Natura Durahesive ConvaTec Moldable Technology Post Operative/Surgical System (57mm) units may contain pouches with 45 mm Natura flanges, therefore the pouch, low pressure adapter and barrier would not fit together. | Employee error | Convatec sent a Medical Device Recall letter dated February 1, 2011, via UPS with return response forms attached to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop distributing and quarantine the recalled lot (0100814).Perform a count of recalled product currently in inventory. Complete the enclosed Recall Response Form and return it to the address on the response form. Return the attached Recall Response Form even if no recalled product is in inventory. Ensure that their account number is correctly identified on the attached Recall Response Form. To return the recalled product call the toll free number 1-800-582-6514 to receive an RGA number and instructions on how and where to return the recalled product. If the product has been distributed to wholesalers then customers should forward the letter to them and ask that they follow the Required Actions and return the Recall Response Form to the address listed on the form. Product must be labeled for shipment to: UPS Supply Chain Solutions1920 Outer LoopDock Door #123Louisville, KY 40219For questions regarding this recall call 908-904-2149. |  | 77 | 77 market units | Nationwide Distribution including AZ, IL, KS, MA, MI, NC, OK, OR, PA, SC, and TX. | Terminated | January 11, 2012 | 344 | 77 | Not_Computer
Z-1611-2011 | 58045 | [6 empty] | MMP 200 Articulating Booms, Catalog Number 0682-000-604SX, Product Number 0682-001-060C. Used for operating room equipment. |  |  | January 28, 2011 | 2011 | Stryker Communications Corp | Utilization of incorrect screws resulted in inadequate thread engagement reducing the load bearing capability of the product. | Process control | Stryker notified consignees verbally on January 28, 2011 and by letter dated February 11, 2011. |  | 2 | 2 | MA, CO | Terminated | March 30, 2011 | 61 | 2 | Not_Computer
Z-1608-2011 | 57211 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ORTHOFIX Firebird Spinal Fixation System, Pedicle Screw Spinal System, Reduction Case, Part #68-0111, Lot #16518-LC18. |  |  | October 22, 2010 | 2010 | Orthofix Inc | There is a possibility that certain Anti-Splay Caps may not mate properly during assembly of the construct to the Screw Body with Multi-Axial Reduction Screw Assemblies. | Other | The firm, Orthofix, sent an "URGENT: PRODUCT RECALL" letter dated October 22, 2010 to all customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to: 1) identify and remove the affected product 2) complete the enclosed Tracking and Verification Form via fax to: Orthofix Inc. at 214-937-2764, and 3) contact their Account Services Representative at 1-888-298-5700 and obtain a return authorization number to return product to Orthofix, Inc.The letter also states that Orthofix Inc., will provide replacement Anti-Splay Caps via priority delivery.If you have any questions regarding the removal and return of this product to Orthofix Inc., please contact your Account Service Representative at 1-888-298-5700. Should you have any other questions or concerns, please contact the Director of Regulatory Affairs at 214-937-2061 or by email: DarlaChew@Orthofix.com. |  | 166 | 166 Units | Nationwide distribution: USA including states of: CA, CO, FL, LA , MD, MI, MO, NJ, NY, OH, TN, VA, and WA. | Terminated | April 25, 2011 | 185 | 166 | Not_Computer
Z-1619-2011 | 56746 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GLU Reagent (GLU cartridge assay) used on Synchron LX and UniCel DxC Synchron Clinical Systems, Part Number 442640; |  |  | August 10, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Cartridge Glucose (GLU) on UniCel DxC and Synchron LX instruments may give falsely high results for moderately hemolysed samples.The Impact:(1) Neonatal samples drawn by lancets are more likely to be hemolysed to a degree sufficient tocause interference with glucose estimation and are a special at-risk population,(2) Other hemolysed samples may be similar | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 10, 2010 with an enclosed fax-back Customer Response Form to the customers. The letter provided the customers with an explanation of the problem identified, the impact and an action to be taken. The customers were instructed to run samples using another method/technology and to complete and return the enclosed FAX BACK RESPONSE FORM via fax to (714) 961-4232 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd., PO Box 8000, Brea, CA 92822.Firms with questions regarding this Product Corrective Action letter, were instructed to call our Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada http://www.beckmancoulter.com/customersupport/support/default.asp. |  | 81871 | 81,871 units | Worldwide distribution: USA and country: Canada. | Terminated | October 17, 2013 | 1164 | 81871 | Not_Computer
Z-1613-2011 | 57831 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | STA-Flex SUBTALAR SCREW, 8 mm x 14 mm, REF 982362, Sterile, Biomet Sports Medicine, 56 East Bell Drive, Warsaw, IN. |  |  | January 21, 2011 | 2011 | Biomet, Inc. | The firm has initiated this recall following internal fatigue testing that shows possible initiation of fatigue cracking in implants under load. | Pending | The firm sent URGENT MEDICAL DEVICE RECALL NOTICES dated 1/21/2011 to distributors, implanting physicians, and medical facilities. The letter to the distributors and medical facilities states the recalled product should be immediately located and returned to Biomet. In addition, the "FAX Back Response Form" should be completed and returned to the number provided. If product has been further distributed, distributors are to contact hospital personnel via the enclosed "Dear Biomet Customer" notice. The letter to the implanting physicians informs the physician of the recall and lists the patients that were implanted with the device. The physician letter quotes from the surgical technique "the implant should also be removed if any postoperative complications occur". The letter to the implanting physicians does not recommend any action be taken with the devices that are implanted. The notice just makes the surgeons aware of the recall. Distributors and Physicians are to confirm receipt of the letter by contacting 800-348-9500, ext 3755 or 3756. Also, questions related to this matter should be directed to 574-371-3755 or 574-371-3756. |  | 254 | 254 all products | Nationwide Distribution. | Terminated | November 10, 2011 | 293 | 254 | Not_Computer
Z-1621-2011 | 56637 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT G3+ Cartridge: PCO2 |  |  | September 25, 2006 | 2006 | Abbott Point Of Care Inc. | Certain lots of G3+ cartridges may be difficult to fill. | Other | The firm, Abbott Point of Care, sent a "Important notice" letter dated September 25, 2006 to all customers. The letter described the product, problem and actions to be taken. The customers were to destroy and discard the i-STAT G3+ Cartridges and complete and return the customer reply form. Abbott sent replacements cartridges to their customers.If you or any of the healthcare providers you serve have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1. |  | 12800 | 12,675 in US, 125 ROW | Worldwide distribution: USA and country of : Canada. | Terminated | February 27, 2012 | 1981 | 12800 | Not_Computer
Z-1492-2011 | 58015 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK FlexLink Plus, subcutaneous infusion set. Part Number 05511089001. |  |  | February 24, 2011 | 2011 | Roche Insulin Delivery Systems Inc. | There is the potential to kink or bend the cannula when inserting the ACCU-CHEK FlexLink Plus infusion set. This can lead to under delivery and elevation of blood glucose levels. | Device Design | Roche Insulin Delivery Systems sent urgent Product Recall letters, dated 2/21/2011. The letter identified the affected product and described the issue regarding kinked or bent cannulas. The firm requests that customers discontinue use of the ACCU-CHEK FlexLink Plus infusion set and return the unused stock. Customers are to contact their health care provider to determine if changes to their pump therapy are needed and how to temporarily continue insulin pump therapy without the ACCU-CHEK FlexLink Plus infusion set. The attached Customer Reply Card should be completed and return via mail or fax per the instructions on the card. The letter also lists some additional information pertaining to signs and symptoms. Questions regarding this recall should be directed to ACCU-CHEK Customer Care at 1-800-688-4578, 24 hours a day, 7 days a week. |  | 1454918000000 | 1,454,918 million worldwide | Worldwide Distribution -- US and Canada. | Terminated | July 27, 2011 | 153 | 1454918000000 | Not_Computer
Z-1438-2011 | 55432 | No | LQF | Mycobacterium tuberculosis immunofluorescent reagents. | Microbiology | Microbiology | 510(k) | Mycobacterium Tuberculosis Complex Culture Identification Test, 2860. |  |  | December 30, 2008 | 2008 | Gen-Probe Inc | This recall has been initiated following the identification of certain assay tubes that do not contain probe reagent. The affected kits may contain a combination of tubes with correct probe amount, partial probe amount and some probe tubes that are empty. Possible false negative results. | Process control | Gen-Probe issued a Recall Notice to customers recommending that they discontinue the use of the identified batch (lot) numbers listed. For customers that have used these batches and obtained negative results, they recommend that the referring physician be notified about the possibility of a false negative result.With the Recall Notice, consignees received a Customer Response Form to complete. Gen-Probe requested that they segregate the affected batches listed on the Customer Response Form, then reconcile, and destroy all partial and/or full kits left in inventory, being sure to document the amount destroyed on the Customer Response Form. Gen-Probe should be updated via phone, fax, and email information. Customers should then sign and return the completed form to Gen-Probe via fax to 858-410-8250, or scanned and emailed to technicalsupport@genprobe.com.The Technical Support team should be contacted for replacement kits. Technical Support was also available for additional information at 888-484-4747 or technicalsupport@gen-probe.com. |  | 1293 | 1293 kits | Worldwide Distribution -- USA, including Puerto Rico and countries of Australia, Guatemala, Malaysia, Bahamas, Hong Kong, Mali, Singapore, Bermuda, India, Mexico, South Africa, Canada, Israel, Netherlands, South Korea, Finland, Kuwait, Panama, Thailand, France, Latvia, and Philippines. | Terminated | June 08, 2011 | 890 | 2659 | Not_Computer
Z-1625-2011 | 54384 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Medtronic brand KyphX HV-R Bone Cement with Kyphon mixer, a polymethyl bone cement kit consisting of a quantity of powder (20 g) and a volume of liquid (10 ml/9 g) mixed in a Kyphon mixer for up to two minutes. Model Number: CO1B; Manufactured by Elmdown LTD, London, England, for Medtronic Spine LLC, Sunnyvale, CA. |  |  | November 10, 2009 | 2009 | Medtronic Spine LLC, formerly Kyphon Inc | Product may not reach proper viscosity (dough state) and inconsistency or lack of homogeneity during cement preparation and during delivery period. | Other | The firm, Medtronic Spine LLC, sent an "URGENT DEVICE RECALL NOTICE" dated November 9, 2009 to the customers. The notice described the product, problem and actions to be taken. The customers were instructed to return the affected lots to Medtronic Spine; to not use any of the affected product that they may have received (Medtronic will replace any affected product in their inventory), the customer should contact their Medtronic Spine sales representative or customer service at [866-959-7466 or for geography] for assistance with replacement product, and complete and return the URGENT DEVICE RECALL NOTICE: Action Required form to Medtronic Spine LLC via fax at (408) 543-6192 or mail to: Medtronic Spine LLC, Attn: Recall Coordinator, 1221 Crossman Avenue, Sunnyvale, CA 94089. If you have any questions about this notification, please do not hesitate to contact Sr. Director, Regulatory Affairs via email: pamela.n.segale@medtronic.com or phone at (408) 548-5235. |  |  |  | USA including states of: AL, AR, AZ, CA, CO, DC, FL, GA, IA, IN, KY, MA, MD, MO, MT, NC, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries including: Argentina, Austria, Belgium, Brazil, Canada, Chile, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Ivory Coast, Jordan, Kuwait, Lebanon, Luxembourg, Lybia, Malaysia, Malta, Mexico, Morocco, Netherlands, New Caledonia, Norway, Panama, Philippines, Poland, Portugal, Puerto Rico, Reunion, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Terminated | March 11, 2011 | 486 |  | Not_Computer
Z-1627-2011 | 57506 | No | DHX | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | CEA ELISA Kit. |  |  | December 22, 2010 | 2010 | BioCheck Inc | Product labels indicate that the products are for IVD use and they do not have an approved 510(k). | PMA | The firm, BioCheck, sent an "URGENT: DEVICE RECALL" letter dated December 22, 2010 to all customers. The letter described the product, problem and actions to be taken by customers. The customer were instructed to examine their inventory for the affected lots, if found, discontinue distributing the lots and promptly return via Fedex to BioCheck at 323 Vintage Park Drive, Foster City, CA 94404, Attn: RETURNED GOODS; immediately contact their accounts, if they further distributed the product, and inform them of the Recall, and complete and return the response form via email at: info@biocheck.com or fax to (650-573-1969) immediately.If you have any questions, please call (650) -573-1968 anytime during office hours (8:30am to 5:30pm Pacific time). |  | 25000 | estimate 25,000 kits, all varieties | Nationwide distribution. | Terminated | March 25, 2011 | 93 | 25000 | Not_Computer
Z-1623-2011 | 54468 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 4000 Vent Head Spring Assembly, a component of the CELL-DYN 4000 Automated Hematology Analyzer, List Number 07H95-01, manufactured by Abbott Diagnostics Division, Santa Clara, CA |  |  | December 29, 2009 | 2009 | Abbott Laboratories | A component used in assembly is out of specification, resulting in failure to retract, eject specimens, exposure below the safety zone, piercing of specimen tubes, or error message to display. | Other | The firm, Abbott Laboratories, sent an "Urgent Field Safety Notice Product Recall" letter dated December 28, 2009 to the Customers via Federal Express. The letter described the product, problem and action to be taken. The customers were instructed to destroy any CELL-DYN 4000 Vent Head Spring Assemblies that they have on hand and have not installed; (No action is necessary for assemblies that are currently installed on the CELL-DYN 4000 Analyzer); complete and return the Customer Reply form (even if they do not have the product) via fax to 1-800-777-0051 or email: QAGCO@abbott.comProduce Recall; indicate the number of Vent Head Spring Assemblies destroyed. the number identified will be replaced with Vent Head Spring Assemblies that meet specifications, and to please keep this communication with their CELL-DYN 4000 System Operator's Manual. If you have any questions regarding this communication, U.S. customers should call Customer Support at 1-877-4ABBOTT (1-877-422-2688). Customers outside the U.S., please contact your local hematology customer support representative. |  | 157 | 157 assemblies | Worldwide distribution: USA including states of: CA, CO, FL, GA, KY, NC, NJ, NY, OR, VA, and WV; and countries including: Argentina, Australia, Canada, Germany and Japan. | Terminated | March 11, 2011 | 437 | 157 | Not_Computer
Z-1647-2011 | 56044 | No | LRM | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio Rad brand Autoimmune EIA Anti-dsDNA Test Kit, 96 Tests, |  |  | March 02, 2010 | 2010 | Bio-Rad Laboratories Inc | The firm received customer complaints of increased positivity rates associated with the Anti-dsDNA test kits. | Nonconforming Material/Component | The firm, Bio-Rad, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated March 2, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory and quarantine the affected kit lots; discard all affected product in accordance with all local, state and federal disposal procedures; complete and return the enclosed MEDICAL DEVICE RECALL RESPONSE FORM via fax to: (510) 741-3954, attn: Bio-Rad CSD Regulatory Affairs Department, and to please re-evaluate all positive patient samples and retest if necessary. If you have any questions, contact Bio-Rad CSD Regulatory Affairs Department at 510-724-7000. |  |  |  | Worldwide distribution: USA and countries including: Canada, Carribean, Chile, France, Germany, Hong Kong, India, Italy, Japan, New Zealand, Panama, Singapore, South Korea and United Kingdom. | Terminated | March 15, 2011 | 378 |  | Not_Computer
Z-1651-2011 | 57534 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure LC 2.0 Instrument, catalog number 05197686001, Roche Diagnostics, Indianapolis, IN |  |  | October 26, 2010 | 2010 | Roche Diagnostics Operations, Inc. | If the bar code is printed from the Batch Results Screen, a wrong batch ID may be printed, while the correct batch ID is still assigned to the batch in the data repository of the MPLC 2.0 software. This may result in sample mis identification. If printed from the ordering screens, the bar code is printed with the correct batch ID. | Software design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 04, 2010 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to print batch ID barcode labels from the ordering or stage set up screens only; print Cooling Block barcode labels from the Post Elution Edit screen only; complete and return the attached UMDC Faxback Form for MagNA Pure LC 2.0 Instrument [UMDC 10-228] via fax to 1-877-551-7152 and file this correction notice for future reference. The firm plans to issue a software update in first quarter 2011. If you have questions about the information contained in this UMDC, for Clinical customers, please contact Roche Molecular Diagnostics Technical Service at 1-800-526-1247; for Research customers, please contact Roche Applied Sciences Technical Service at 1-800-262-4911. |  | 119 | 119 devices | Worldwide distribution: AL, AZ, CA, CO, DC, FL, GA, HI, IL, KY, LA, MA, MD, ME, MN, MO, MT, NC, ND NE, NH, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, VT, WI and Puerto Rico; and country of Europe. | Terminated | August 30, 2011 | 308 | 119 | Software
Z-1644-2011 | 57724 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; the pumps were sold under the following configurations: |  |  | February 14, 2011 | 2011 | Hospira Inc. | Hospira has received customer reports of the Plum A+ infusion pumps with no audible alarm conditions, which have been associated with failure of the piezoelectric assembly due to component quality issues. | Nonconforming Material/Component | The firm, Hospira, sent "URGENT DEVICE RECALL" letters dated February 14, 2011 to their customers on the same date. The letter described the product, problem and action to be taken by the customers. The Customer/Clinicians were instructed to weigh the risk associated with continued use of the device versus removal from service. If they elect to continue the devices, they were provided with instructions to perform alarm tests prior to each clinical use (during their pump cleaning process). If the alarm is not audible, they were instructed to discontinue use of the pump and contact Hospira Global Product Safety and Complaints at 1-800-441-4100. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-888-216-7330. |  | 144477 | 144,477 pumps | Worldwide distribution: United States including the U.S. Virgin Islands and Puerto Rico, and countries including: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, United Arab Emirates and Vietnam. | Open, Classified |  |  | 167595 | Hardware | Alarm/Message | Remove | Remove or Replace
Z-1649-2011 | 57914 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-ART System, Model # H-0000-0003 |  |  | January 26, 2011 | 2011 | TomoTherapy Incorporated | It was determined that the Treatment Planning Station (TPS) can potentially under dose. | Software design | Consignees were sent a TomoTherapy "Field Safety Notice Medical Device Correction" letter dated January 26, 2011. The letter was addressed to "TomoTherapy Customer". The letter described the Planning Issue, Related DQA Consideration, Graphical Presentation of Results, Product Affected, Recommended Action and Resolution. The letter suggests to use the information presented in order to avoid creating plans with an increased potential for discrepancies in planned vs. delivered dose. Specifically include measurements for small, off-axis targets that could be susceptible to the effects of this issue. DQA dose measurements should be performed at each unique full fraction dose that will be used for patient treatment, without using DQA scaling.If Customers have any questions they are to contact TomoTherapy Interaction Center by email or phone. |  | 326 | 326 | Worldwide Distribution -- US, including States of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NJ, NM, NY ,NC, ND,OH, OK OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, and WY and countries of SPAIN, GERMANY, FRANCE, INDIA, KOREA, SWEDEN, ITALY, MALAYSIA, SWITZERLAND, CANADA, UNITED KINGDOM, SINGAPORE, TURKEY, TAIWAN, AUSTRALIA, JAPAN, BELGIUM, CHINA, MEXICO, NETHERLANDS, POLAND, SAUDI ARABIA, and UNITED ARAB EMIRATES. | Terminated | March 18, 2014 | 1147 | 326 | Software
Z-1658-2011 | 54136 | No | LKV |  |  | Clinical Chemistry | PMA | Rapid fFN for the TLi IQ system and specimen collection kit in the following configurations: 0797Specimen Collection Kit (8 pack); manufactured by Hologic, Sunnyvale, CA. |  |  | December 15, 2009 | 2009 | Cytyc Prenatal Products/Hologic Inc. | Product name change requires conversion to a different type of PMA and implies that the product could be used to assess whether a patient's pregnancy is full term, or to predict imminent delivery in women who are at full term gestation. | PMA | The firm, Hologic, sent a "Important-Product Recall/Correction" letter dated December 18, 2009 to its customers. The letter described the product, problem and the actions to be taken. The customers were instructed to 1) replace the directional inserts for the specimen collection kit and health care providers' brochure with provided instructions, 2) over-label specimen collection kit boxes with provided stickers and individual kits, 3) permanently destroy any marketing and educational materials with the "full term" name and 4) complete and return the CUSTOMER RECALL RESPONSE FORM via fax at 408-744-1905 or scan and email to recallfullterm@hologic.com.If you have any questions about this notification, please contact the Hotline (800-442-9892, Option 7# or 508-263-8510). |  |  |  | Worldwide distribution. | Terminated | March 16, 2011 | 456 |  | Not Computer
Z-1681-2011 | 56332 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SCORPIO CR TIBIAL TRIAL, non sterile, made in the USA |  |  | July 13, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is the potential for the Scorpio T72 tibial Insert Trials to be oversized by a maximum of .024" (0.6mm). | Other | Stryker Howmedica Osteonics Corp.sent Urgent Product Correction letters Correction letters dated July 23, 2010, via Federal Express to all affected customers. were sent to all Stryker Branches on July 23, 2010 via Federal Express. Asia Pacific, Canada, Europe, Middle East, Africa, Japan and Latin America sites were sent letters on July 26, 2010 by Federal Express. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the attached Product Correction Acknowledgement Form by fax as soon as possible to 201-831-6069.For any questions customers were to call 201-972-2100 or 201-831-5028. |  | 159095 | 159,095 units, all types. | Worldwide Distribution - USA including OH, TX, MS, and CO and the countries of Australia, Hong Kong, Taiwan, India, Korea, Japan, Chile, Brazil, Columbia, Venezuela, , Panama, Santo Domingo, Mexico, Argentina, Costa Rica, Guatemala, Peru,Sweden, Germany, Netherlands, France, Spain, Italy, Romania, South Africa, United Kingdom. and Canada. | Open, Classified |  |  | 159095 | Not_Computer
Z-1665-2011 | 56407 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE System |  |  | July 22, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has found that communication may be cut off in the Aquilion ONE/ Premium Systems, owing to the failure of a part used in the console. | Software design | The firm, Toshiba America Medical Systems (TAMS), sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 22, 2010, to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that a new relay will be installed to prevent the occurrence of the problem within the next 90 days and that their Toshiba Service Representative will contact them to install the part to their Aquilion ONE/Premium System. The customers were instructed to contact their Toshiba Service Representative, if the problem occurs before the corrective action has been taken; to share the information with all users and review radiologists as well as clinical engineering or Biomedical group at their facility, and to complete and return the attached CUSTOMER REPLY FORM via fax to 877-349-3054 or email to raffairs@tams.com.If you have any questions regarding this letter, please feel free to contact the Director, Regulator Affairs at (800) 421-1968 or your local Toshiba Representative at (800) 521-1968. |  | 43 | 43 units | Nationwide distribution: USA including states of: AL, AR, AZ, CA, CO, FL, GA, IA, KY, LA, MA, MD, MI, MN, MT, ND, NJ, NV, NY, OH, PA, PR, TX, VA, WI, and WY. | Terminated | April 05, 2011 | 257 | 43 | Software
Z-1675-2011 | 56415 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Toshiba DUA-450F Kalare Radiographic Tilting Table. |  |  | July 15, 2010 | 2010 | Toshiba American Medical Systems Inc | Top table lateral movement on their DUA-450F Kalare Table cannot be performed, or the table top lateral movement is performed at an extremely low speed. As a result, positioning of the patient cannot be performed by tabletop movement. | Software design | Toshiba America Medical Systems (TAMS) issued Urgent Medical Device Correction notification dated July 8, 2010, to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. |  | 238 | 238 systems | US Nationwide | Terminated | April 20, 2012 | 645 | 238 | Software
Z-1674-2011 | 57199 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | BLACKSTONE Medical Inc. Firebird Spinal Fixation System 6.5mm x 30mm Bone Screw, Self Tapping, Cannulated. BLACKSTONE Medical Inc. Firebird Spinal Fixation System 6.5mm x 30mm Multi-Axial Screw, Self Tapping, Cannulated. |  |  | October 22, 2010 | 2010 | Orthofix Inc | There is a possibility that certain Multi-Axial Screws have spherical heads measuring less than the allowable tolerance. | Other | Orthofix sent an URGENT: PRODUCT RECALL letter dated October 22, 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their accounts immediately and advise them of the recall situation and have them return any affected inventory.For any questions customers were instructed to contact their Account Services Representative at 1-888-298-5700.For questions regarding this recall call 214-937-2061. |  | 27 | 27 (19 of 77-7630) (8 of 77-8630) | Nationwide Distribution including AZ, NY, TX, MS, LA, FL, NC, and AR. | Terminated | May 31, 2011 | 221 | 27 | Not_Computer
Z-1679-2011 | 57513 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Acetaminophen Test System for use on the Roche/Hitachi, part number 03255379190, Roche Diagnostics Corporation, Indianapolis, IN. |  |  | December 07, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The current product labeling for Acetaminophen does not provide specific interference information for bilirubin, hemoglobin or lipemia. Additional testing has been performed to quantify the impact of these interferences at different concentrations and different levels of acetaminophen. Low serum concentrations of acetaminophen were included in the study. | Other | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 07, 2010 to its customers. The letter described the product, problem, and action to be taken. The customers were instructed to be aware that endogenous interferents (i.e., bilirubin, hemolysis, and/or lipemia) may produce a falsely elevated value for samples containing concentrations less than 50 ug/ml of acetaminophen; complete and return the attached UMDC Faxback Form for Acetaminophen Assay via fax to 1-817-868-5343, and file this letter for future reference. The firm ask the customers to please consult with the physician or pathologist at their facility to determine specific clinical implications for their patients. The firm is also updating the product labeling to include information regarding the interference of icteric, hemolytic or lipemic samples.Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have any questions about the information contained in this Urgent Medical Device Correction (UMDC). |  | 8668 | 8,668 units | Nationwide distribution. | Terminated | March 13, 2012 | 462 | 8668 | Not_Computer
Z-1669-2011 | 57690 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs of the Computed Tomography X-Ray Systems under the following Brands and Model #s: Brilliance 6, M/Ns 3001-3708, 30001-30131; Brilliance 10, M/Ns 4001-4067, 40001-40035 ; Brilliance 16, M/Ns 5001-6180, 50002-50139 ; Brilliance 16P, M/Ns 6007-6163, 60001-60029 ; Brilliance 40, M/Ns 9009-9272, 29001-29153, 90028-92034; Brilliance 64, M/Ns 9501-10300, 95003-95697; Brilliance Big Bore, M/Ns 7001-7520; Brilliance iCT, M/Ns 100003-100169 &amp; Brilliance iCT SP, M/Ns 200004-200034. |  |  | December 10, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems has decided to recall the patient support vertical brake hub in all its Brilliance systems due to the collapse in the vertical direction. The short key connecting the vertical brake hub to the motor shaft can disengage due to the lack of thread locker on set screws. | Other | Philips Medical Systems (Cleveland) Inc sent Field Safety Notice to all affected customers on December 10, 2010. The Safety Notices identified the product, the problem, and the action to be taken by the customer. Customers were instructed to monitor their units and if the vertical brake hub malfunctioned to immediately contact the recalling firm. The letter also stated that a Field Service Engineer would visit to perform a brake hub rework on the units. For further information or support customers were to contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center 1-800-722-9377, option 5: Enter Site ID or follow the prompt(s).For questions regarding this recall call 440-483-7672. |  | 5700 | 5700 units/vertical brake hubs | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO,CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT,NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY, and the countries of Afghanistan; Algeria; Angola; Argentina; Australia; Austria; Azerbaijan; Bahamas; Bahrain; Bangladesh; Belgium; Bolivia; Bosnia &amp; Herzegovina; Brazil; Brunei Darussalam; Bulgaria; Cambodia; Canada; Chile; China; Columbia; Costa Rica; Cote d'Ivoire; Croatia; Cyprus; Czech Republic; Denmark; Dominican Republic; Ecuador; Egypt; El Salvador; England; Equatorial Guinea; Finland; France; French Polynesia; Germany; Greece; Hong Kong; Hungary; Iceland; India; Indonesia; Iran; Iraq; Ireland; Israel; Italy; Jamaica; Japan; Jordan; Kazakhstan; Kenya; Kuwait; Latvia; Lebanon; Libyan Arab Jamahiriya; Lithuania; Luxembourg; Madagascar; Malaysia; Malta; Martinique; Mexico; Monaco; Mongolia; Morocco; Mozambique; Namibia; Nepal; Netherlands; Netherlands; Antilles; New Zealand; Nicaragua; Nigeria; Norway; Oman; Pakistan; Panama; Paraguay; Peru; Philippines; Poland; Portugal; Romania; Russia; Saudi Arabia; Serbia &amp; Montenegro; Singapore; Slovakia; Slovenia; South Africa; South Korea; Spain; Sweden; Switzerland; Syrian Arab Republic; Taiwan; Tanzania; Thailand; Tunisia; Turkey; Ukraine; Unidentified KM; United Arab Emirates; United Kingdom; Uzbekistan; Venezuela; Viet am &amp; Yemen. | Terminated | August 21, 2013 | 985 | 5700 | Not_Computer
Z-1667-2011 | 57880 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Innovate Esmark Bandages, a component of various Presource Custom Sterile Packs/Presource PBDS Modules; Packaged by Cardinal Health Medical Products and services, McGaw Park, IL 60085 U.S.A.; There is no label on the Esmark bandage in the Presource pack/module, Some are placed in Tyvek pouches and some are just loose in the pack, but neither have any labeling on them. |  |  | January 28, 2011 | 2011 | Cardinal Health | Esmark bandages placed in some Presource kits/modules may not have been adequately sterilized. | Nonconforming Material/Component | The firm, Cardinal Health, sent an "URGENT: PRODUCT RECALL" letter dated January 28, 2011 to all affected consignees/customers via UPS next day air. The letter described the product, problem and action to be taken. The customers were instructed to quarantine and return the packs/modules to Cardinal Health for destruction; contact the appropriate Customer Service group listed in the letter for return instructions, and complete and return the enclosed Acknowledgment Form via fax to 847-689-9101, regardless of whether or not they have inventory. Should you have any questions, relating to this recall, please feel free to contact Cardinal Health Quality Systems, 800-292-9332. |  | 4.395 | 4.395 modules | Nationwide distribution. | Terminated | September 04, 2012 | 585 | 4.395 | Not_Computer
Z-1677-2011 | 57926 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OSTEOMATRIX bone graft substitute. OMG10. |  |  | January 31, 2011 | 2011 | Biostructures, LLC | The OsteoMatrix Synthetic Bone Graft Composite Instructions for Use document, P4G10A/Rev. 01 had language that was not consistent with the cleared indications. | Other | An email was sent on 06/15/2010 to the firm's distributors that had products with lot #0809D110 or lot #0307I208 in their possesion. The email was in regards to the distributors' OMG10 granule inventory. The distributors need to return all the recalled granule inventory to be swapped out for new inventory. |  | 248 | 248 units | Products were distributed to 6 consignees that were located in Michigan, Ohio, Maryland, California, Texas, and Utah. | Terminated | March 17, 2011 | 45 | 248 | Not_Computer
Z-1654-2011 | 57777 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura Xper Systems with Velara Generator, x-ray system. FCO72200184. |  |  | October 07, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | A wire in the connector of the Generator Grid Switch Supervisor box might short-circuit between 12V power and ground, and it will cause a System breakdown. | Nonconforming Material/Component | On 10/15/2010, Philips began sending out the URGENT-Device Correction Notice Failure Causing the Generator to Stop Working notification letter to their consignees.The consignees were informed that a short circuit may occur due to a faulty wire connection in the Supervisor Box Connector in the Generator, and it will cause the generator stops working. When the system fails during a critical interventional case, this may cause risk to patients (no Fluoro available).The systems involved in this recall are Allura Xper with the Velara Generator, Allura Integris with OMCP Generator, MultiDiagnost Eleva, and OmniDiagnost Eleva Imaging system.Consignees can contact Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 70800114, 70800115, 72200184, and 72200185. |  | 1556 | 1556 units distributed in US (1452-3107=1556) | Devices were distributed through out the US. | Terminated | February 23, 2012 | 504 | 2739 | Hardware | Device Operation | Notification | Safety Notice/Insructions
Z-1676-2011 | 57932 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Remote Cardiology Review Station, a cardiovascular information system. Model Number CV7.8 |  |  | February 10, 2011 | 2011 | AGFA Corp. | A discrepancy in the validation testing which resulted in the product not performing as intended. | Device Design | AGFA sent an "URGENT SAFETY NOTICE" letter dated 2/10/11 via FED-EX to the customer. The letter included an acknowledgment which was to be FAX-Back indicating that the information was received and understood. Agfa has discussed with the customer how the de-installation of their CRS Remote product - Software Version 7.8_HL would be performed. |  | 1 | One | TX | Terminated | September 15, 2011 | 217 | 1 | Other | Device Operation | Software update | Software Update
Z-1664-2011 | 57967 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | LUCAS 2 Chest Compression System. |  |  | February 08, 2011 | 2011 | Jolife AB | LUCAS 2 Chest Compression System may have malfunctions resulting from the failure of a hood probe on the device that connects to the battery. The device may stop in fail safe mode and the LEDs will start flickering or the device will switch off. | Other | Medical Device Correction letter, dated February 2011, was sent to customers and states that Jolife AB with the servicing support assistant of Physio-Control is conducting a voluntary upgrade of the LUCAS 2 Chest Compression System with the Serial number range of 30090016 through 30101932.Consignees were advised to keep LUCAS 2 in service and continue to use it in accordance with the Instruction for Use, Section 3.8-Safety Precautions. WARNING-MALFUNCTION. If there are interruptions, or the compressions are not sufficient or something unusual occurs during operation: Push ON/OFF for 1 second to stop LUCAS and remove the device. Start Manual chest compression.Physio-Control service representative will contact the consignees to schedule a service call to replace the affected battery connector board in the hood of the device. The new connector board will include a new probe design in which current cannot affect the spring.Consignees can contact Physio-Control Technical Support at 1-800-442-1142, #5 between 6 am-4 pm (Pacific Time), Monday-Friday or visit the website www.physio-control-notices.com/LUCAS2. If customers no longer have the affected device or the device location has changed they should contact Physio-Control Technical Support. |  | 1843 | 1843 units | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, China, Croatia, Czech Republic, Denmark, France, Germany, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirates, and United Kingdom. | Terminated | April 24, 2012 | 441 | 1843 | Other
Z-1673-2011 | 57747 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ovation Hip Stem, REF 201-0104, Offset Stem Inserter Neutral, Non-Sterile, Ortho Development, Draper, Utah 84020. |  |  | January 11, 2011 | 2011 | Ortho Development Corporation | Component of total hip arthroplasty may break during surgical procedure. | Process design | Ortho Development Corporation notified Field representatives of the issue and conditions that could lead to failure by e-mail on December 14, 2010. Redesigned product is being produced for replacement.For questions regarding this recall call 801-619-3450. |  | 33 | 33 units | Nationwide Distribution including AZ, CA, CO, FL, MI, NV, OR, and UT. | Terminated | November 04, 2011 | 297 | 80 | Not_Computer
Z-1666-2011 | 58053 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Labeling on box reads: MAQUET BEQ-HG 0284 3/8" X 3/8" Connectors, Lot 70041098 but the product packaged inside was MAQUET BEQ-HG 1113 3/16" x 3/16" Connectors, Lot 70041096. The HG 0284 and HG 1113 connectors are packaged sterile as single components in a container holding 20 units. |  |  | February 23, 2011 | 2011 | Maquet Cardiovascular, LLC | Tubing connectors incorrectly packaged and labeled. | Other | The firm, Maquet, sent a "PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION" letter dated February 22, 2011 via Federal Express to its customer. The letter described the product, problem and the action to be taken. The customer was instructed to examine their inventory immediately to identify and remove the affected connectors and complete and return the attached RECALL- Device Corrective Action Response Form via fax to Director of Regulatory Affairs at (973) 709-7016 or email: whitney.torning@maquet.com. If you have any additional questions, please contact your local MAQUET Cardiovascular Sales Representative or customer service at 1-800-777-4222. |  | 11360 | 40 BEQ-HG 0284 3/8" x 3/8" Lot 70041098 Connectors | Nationwide distribution: USA state of CA. | Terminated | September 13, 2012 | 568 | 11360 | Not_Computer
Z-1701-2011 | 51955 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott CELL-DYN 3700 Hematology Analyzer CS (Closed Sampler), , and CELL-DYN 3700 SL (Sample Loader), Manufactured by Abbott Diagnostics Division, Santa Clara, Ca |  |  | March 27, 2009 | 2009 | Abbott Laboratories | FIRE HAZARD-- Some power supply modules may have incorrect fuses installed, which may result in fires and smoking. | Component design/selection | Abbott Laboratories sent Product Correction letters, dated March 22, 2010, to its customers. The letter listed the additional products that are now included in the recall and gave an explanation of the reason for recall. The letter also discussed the safety impact and provided a list of necessary actions for customers. Customers are to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3000 Analyzer. If the fuse is correct, then the steps provided should be followed. If the fuse is incorrect, then customers are to replace the fuse with the correct fuse and follow the instructions provided. Customers are to contact their local Customer Support Representative in the event that they do not have their replacement fuse that was provided with the Accessory Kit. Customers can refer to the Product Information Letter dated April 20, 2010 for additional labeling to assist with the correct fuse replacement. The Customer Reply form should be completed and the letter should be kept with the Operator's Manual. Questions should be directed to Customer Support at 1-877-4ABBOTT.The firm, Abbott Laboratories, sent a "Product Correction" letter dated March 27, 2009 to its customers. The letter included a Customer Reply form and CELL-DYN accessory kits to inform new customers of the necessity of replacing fuses correctly. The letter described the product, problem and actions to be taken. The customers were instructed to check that the ampere rating of the fuse matches the voltage of operation that is currently being utilized with their CELL-DYN 3700 System and follow the instructions provided, and complete and return the Customer Reply form via fax to 1-800-777-0051 or e-mail QAGCO@abbott.com, even if they no longer have the instrument. The firm intends to service items in the field. In April 2010, additional labeling was sent to customers providing additional information when replacing fuses. |  | 2961 | 2961 units | Worldwide Distribution -- USA, including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of Argentina, Australia, Bahamas, Barbados, Brazil, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Curacao, Dominica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Haiti, Honduras, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand Trinidad and Tobago, Turks &amp; Caicos, Uruguay, and Venezuela. | Terminated | December 21, 2011 | 999 | 6881494 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-1698-2011 | 54665 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | The Waste Line Assembly and the Waste Outlet Tubing and the Waste Bottle Cable are components that are used on one or more CELL-DYN Systems. |  |  | November 30, 2009 | 2009 | Abbott Laboratories | The use of the device assemblies has been extended to a 6 month useful life. | Other | The firm, Abbott Laboratories, sent a "Product Correction" letter dated November 30, 2009 to all consignees/customers. The letter described the product, problem and actions to be taken. The letter stated that Abbott recommended that the customers change their Waste Line Assembly and Waste Outlet Tubing every six months and included a tag in the letter to record installation and replacement dates. Additionally, the customers were instructed to keep this communication with their CELL-DYN System Operator's Manual and complete and return the Customer Reply form via fax to 1-800-777-0051 or email QAGCO@abbott.com even if they no longer have the instrument.If you have any questions regarding this communication, U.S. customers should call Customer Support at 1-877-4ABBOTT (1-877-422-2688). Customers outside the U.S. please contact your local hematology customer support representative. |  | 874 | 874 | Worldwide distribution: AL, AR, AZ, CA. CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO MS, MT , NC, NE, NJ, NM, NY, OH, OK, PA, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries including: Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Dominican Republic, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Panama, Peru, Puerto Rico, Singapore, South Korea, Thailand, Trinidad and Tobago, and Uruguay. | Terminated | March 18, 2011 | 473 | 874 | Not_Computer
Z-1668-2011 | 56968 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Diluent Reagent, 10 mL, List Number: 09H48-02, |  |  | September 30, 2010 | 2010 | Abbott Laboratories | The product bar code label, when scanned, indicates that the volume is 10 mL rather than the correct volume of 10,000 mL. Analytical results are not impacted, although, there is a potential for delay in generating results. | Error in labeling | Product Recall notification letters were sent on 9/30/2010. The letter identified the affected product, an explanation of the problem, patient impact, and necessary actions. Consignees are to contact their local hematology customer support representative to arrange for replacement of the CELL-DYN Emerald Diluent Lot. Once the replacement lot has been received, customers should destroy the affected lot according to their laboratory procedures. A copy of the letter should be retained for their records and the Customer Reply Form should be completed and returned via fax or e-mail. US customers should contact Customer Support at 1-877-4ABBOTT, if they have questions. |  | 322 | 322 units | Worldwide Distribution -- USA, including states of PA, NY, CA, NC, IL, TX, KS, UT, VA, FL, AR, TN, MN, NE, MI, and DE and the country of Germany. | Terminated | April 07, 2011 | 189 | 322 | Not_Computer
Z-1691-2011 | 57734 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Sabina EE Lift, Liko, Sweeden. |  |  | January 27, 2011 | 2011 | Hill-Rom, Inc. | The firm has received reports of injuries related to the Sabina Sit-in-Stand lifts. Of six injury reports, two were injuries in patints during use and four injuries were in non-patients when they tripped over the lift device and fell on the unit. | Device Design | The firm, Liko, issued two "URGENT FIELD SAFETY NOTICES" one dated January 27, 2011 and a follow-up notice dated February 9, 2011 to its consignees/customers. The notices described the product, problem and actions to be taken. The customers were instructed to keep the device in a separate unoccupied area when not in use; use extra care and support when lifting patients with the device; not allow unauthorized persons, especially children, to play around or operate the lift; examine the sling hooks daily to make sure they are secure, and to forward a copy of this notice to any other facility personnel they deem appropriate. The notices state the firm is investigating a design change to mitigate the risk of injuries with the device.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 1009 | 1,009 all products | Worldwide distribution: USA including: AL, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK ,OR, PA, RI, SC, SD, TN ,TX, UT, VA, WA, WI, and WV; and country of: Canada. | Terminated | September 19, 2014 | 1331 | 1009 | Not_Computer
Z-1699-2011 | 58075 | No | EPP | External nasal splint. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | An external nasal splint is a rigid or partially rigid device intended for the use externally for immobilization of part of the nose. |  |  | February 21, 2011 | 2011 | Medtronic Xomed, Inc. | Medtronic Xomed, Inc. Jacksonville, FL issued a recall on Feb 21, 2011 for all lots of the alcohol pad packaged with the Medtronic's External Nasal Splints REF 1528116 (sm), 1528126 (med), 1528136 (lg) and the Thermasplints REF 1529100 (sm), 1529110 (med), and 1529120 (lg) distributed from Jan 2008 through Aug 2010. The alcohol supplier, Cardinal Health, issued a recall on the alcohol pads on Jan | Material/Component Contamination | Medtronic Xomed, Inc. sent an URGENT PRODUCT SAFETY ADVISORY (PSA) letter dated February 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately discontinue use and dispose of the alcohol prep pad co-packaged in their External Nasal Splint or Thermasplint kits. If there has been further distribution of this item, the Urgent Product Safety Advisory letter should be forwarded to the appropriate recipient. For any questions regarding this recall call 904-279-7532. |  | 26305 | 26,305 units (5,261 boxes) | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Singapore, Canada, Belgium, Croatia, Cyprus, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Latvia, Martinique, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Switzerland, United Arab Emirates, United Kingdom, Hong Kong, Taiway, Korea, Puerto Rico, and Costa Rica. | Terminated | April 03, 2012 | 407 | 34690 | Not_Computer
Z-1044-2011 | 57763 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE DS FULL FIELD MAMMOGRAPHY SYSTEM |  |  | July 20, 2010 | 2010 | GE Healthcare, LLC | It was noticed that 3 units received in Japan were missing the rating plate on the X-ray tube cover. There is a potential that other X-ray tube covers are missing the rating plate. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter was issued to customers identifying the affected product and describing the defect along with the related hazards. The letter states that there are no actions that are required to be taken on the part of the customer. A GE Healthcare Service Representative will apply the missing rating as necessary and free of charge. Questions or concerns regarding the letter should be directed to 1-800-437-1171 in the US. |  | 112 | 112 total | Worldwide Distribution | Terminated | May 16, 2012 | 666 | 112 | Not_Computer
Z-1709-2011 | 57970 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LTS-D, SIZE 4, QUANTITY 5 EACH, |  |  | January 14, 2011 | 2011 | King Systems Corp. | King Systems has been notified by Triad Group, our supplier of sterile lubricating jelly, that a recall has been initiated for this product because the product may be "inadequately sterilized" and use of this product may result in "patient infection". Because this particular sterile lubricating jelly is packaged in the KING LT-D Kit and KING LTS-D Kit, King Systems is sending a Notice of the Reca | Process control | The firm, King Systems, sent an "URGENT PRODUCT RECALL" letter dated January 14, 2011 to its customers. King Systems sent a follow-up "URGENT PRODUCT RECALL" letter dated March 3, 2011 which "SUPERSEDES" the January 14, 2011 recall letter with attached instructions to its customers. The letters described the product, problem and actions to be taken by the customers. The customers were instructed to notify all their purchasers of this recall; immediately cease sales and use of the affected products; identify all inventory on hand or returned to them by their customers; REMOVE Triad Group lubricant from King LT-D and LTS-D kits according to King instructions; quarantine the Triad lubricant, re-seal the LT-D and LTS-D kits and destroy Triad lubricant or return to Triad according to King instructions, and complete, sign and return Recall Acknowledgement Form via fax to: 317-776-6827 or email to: triadluberecall@kingsystems.com. King Systems will provide replacement lubricant packs to all customers. Customers with questions regarding product return, replacement, hospital notification or concerning the information requested above, please contact Customer Service at 1-800-642-5464. If you have any regulatory questions about the recall, please contact Director of Quality Assurance and Regulatory Affairs at 1-800-642-5464. |  | 577426 | 577,426 total units | Worldwide distribution: USA including: CA,CO, FL, ID, IN, KS, NC, NH, NY, OH, PA, PR, SC, TN, and TX; and Country of: Canada. | Terminated | February 20, 2013 | 768 | 577426 | Not_Computer
Z-1690-2011 | 57900 | No | JDK | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DePuy, a Johnson &amp; Johnson company, CEMENT RESTRICTOR, Size 5, Sterile, DePuy Orthopaedics. This device is intended to restrict bone cement migration into the distal medullary canal during implantation of a femoral hip prosthesis |  |  | January 27, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm became aware of a manufacturing error resulting in a potentially incomplete seal of the outer pouch. Although the sterility of the product is not compromised, the sterility of the inner pouch containing the product cannot be assured. Lack of sterility of the inner pouch could potentially cause contamination of the sterile field, including the implant. Non-sterile product and/or contaminat | Employee error | DePuy Orthopaedics sent an URGENT INFORMATION  RECALL NOTICE dated January 27, 2011, to DePuy Area Directors, FMDs, Distributors, and Office Managers. The notice identified the product, the problem, and the action to be taken. All territories were required to ensure that all recalled products that remain in customer accounts and field inventory are returned immediately to Warsaw to the attention of Returns. Hospitals that purchased and potentially implanted the now recalled product should be provided appropriate notification of the recall and should complete and return a signed Reconciliation Form(s) within seven business days. If the hospitals have any inventory on hand that is subject to the recall, distributors should remove it or help the hospitals return it for credit. All Reconciliation Forms should be faxed to 574-372-7567.For questions regarding this recall call 574-372-7333. |  | 20 | 20 | Nationwide distribution including the states of MA, PA, FL, LA, MN, WA, WI, CO, GA, NY, NC and SC and the countries of Canada, Colombia, Ireland and Argentina. | Terminated | April 23, 2013 | 817 | 140 | Not_Computer
Z-1704-2011 | 55617 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Stiletto Electrosurgical Probe (E-Probe); Model number: 0270-070-550, 0270-070-551, 0270-070-552, 0270-070-553, 0270-070-554, 0270-070-555; Manufactured by Stryker Endoscopy Puerto Rico. |  |  | February 25, 2010 | 2010 | Stryker Endoscopy | The heat shrink, PTFE E-Probe Tip were out of specification (OOS) in the longitudinal change (+5% -20%). | Device Design | All Sales reps were notified of the recall on February 25, 2010. Urgent Device Removal advisory notices were also sent to all consignees.Device Removal letters identified the specific part numbers and lot numbers involved. Letters also list the potential risk to patients as a result of the issue. Customers are asked to inspect their inventory for part and lot numbers and quarantine any affected product. If recalled product is found, customers are to contact Stryker Customer Service at 800-624-4422 and request an RMS to send back recalled units. New units will be provided at no charge. The acknowledgement form is to be completed and returned. Questions or concerns should be directed to 408-754-2000 or e-mail at eprobe@stryker.com. |  | 31014 | 31, 014 units | Worldwide Distribution -- US, Australia, Canada, Germany, Greece, Iberia, Korea, Switzerland, and Newbury. | Terminated | March 18, 2011 | 386 | 31014 | Not_Computer
Z-1685-2011 | 56399 | No | CGF | Estrone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Estrone Radioimmunoassay (RIA) Kit, DSL-8700. |  |  | April 22, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter has confirmed that eight Estrone RIA kit (Ref DSL 8700) lots are inaccurately calibrated to the firm's internal reference calibrator. Lots 992637, 100115A, 100205A, 100205RA, and 100226A have now expired and should not be in active use any more. | Other | Beckman Coulter released an Urgent: Product Corrective Action letter dated July 9, 2010 identifying the affected product and informing consignees of the reason for recall, asking them to take the following actions:*For lots 992637, 100115A, 100205A, 100205RA, and 100226A, review historical external quality control results relevant to the patients' results obtained with these lots to ensure the assay performance was within laboratory specifications.*For lots 100319A, 100409A, 100409RA, review patents' results obtained with the results of external quality control and interpret them in the light of additional applicable tests, patient's clinical status and clinical history of patient.Beckman Coulter has finalized the process of re-adjusting Estrone calibrators' values assignment to the new internal reference calibrator starting with lot 100403A.If consignees need products for additional tests they were told to follow replacement instructions below:*In the United States, please contact Customer Service at 1-800-526-3821*Outside the United States, contact your local Beckman Coulter Representative.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. They were to also complete and return the enclosed response form within 10 days so the firm can be assured that consignees received the notification.Consignees were instructed to forward any other questions regarding the notification to the DSL Technical Support Center at 1-800-231-7970 in the United States, or contact their local Beckman Coulter Representative. |  | 917 | 917 units | Worldwide Distribution -- US, including states of CA, FL, GA, IL, IN, KS, MI, NC, NY, OH, OR, TX, VA, WA, and WI and country of Canada. | Terminated | May 04, 2012 | 743 | 917 | Not_Computer
Z-1733-2011 | 55146 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott brand CELL-DYN Emerald Hematology Analyzer System; List Numbers: 09H39-01; The product is manufactured in France for Abbott Diagnostics Division of Abbott Laboratories, Abbott Park, IL |  |  | March 23, 2010 | 2010 | Abbott Laboratories | Abbott has determined that gain settings for the non-standard specimen types may not match the gain settings for the standard specimen type. There is the potential for impact to product performance. | Software design (manufacturing process) | The firm, Abbott, sent two "Product Correction" letters one dated March 23,2010 and a follow-up letter dated July 1, 2010, via express mail to all its consignees/customers. The letters described the product, problem and actions to be taken. The customers were instructed to perform the following steps, noted in the letter, to print and verify their gain setting. Abbott informed the customers that they would receive a mandatory software update and were instructed to install the CELL-DYN Emerald according to the attached instructions, and to complete and return the Customer Reply form via fax to 1-800-777-0051 or email to QAGCO@abbott.com.Should you have any questions or would like to discuss in more detail, contact Director Quality Assurance and Regulatory Affairs Abbott Hematology at 408-567-3403. |  | 891 | 891 units | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, and WV; and countries including: Australia, Brazil, Canada, Columbia, Dominican Republic, Germany, Puerto Rico, Singapore, South Korea, St. Eustalius &amp; St. Martin, and Turks &amp; Caicos. | Terminated | March 18, 2011 | 360 | 891 | Other |  | Software update | Software Update
Z-1727-2011 | 56962 | No | OJZ |  |  | Orthopedic | HDE - Humanitarian Device Exemption | Stryker Biotech OP-1 Putty (OP-1, BMP-7, bone morphogenetic protein) |  |  | October 25, 2010 | 2010 | Stryker Biotech | A change to the Preparation for Use sectionwas identified where 2.5cc was replaced with 2cc-3cc volume and was not approved | Labeling Change Control | The firm, Stryker Biotech, sent a "Medical Device Correction" letter on October 15, 2010 via return receipt letter to all customers. The letter contains labeling with the approved Preparation for Use Section of 2.5 cc volume and included the revised copy of the Package Insert. Additionally, the letter described the product, problem and actions to be taken. The customers were instructed to follow the new instructions provided and to keep the new instructions in mind when using the OP-1 Putty.If you have any questions concerning this communication, please contact 508-416-5326 or Stryker Biotech's Pharmacovigilance department at PVG@stryker.com. |  | 2861 | 2861 units | Nationwide distribution. | Terminated | March 20, 2012 | 512 | 2861 | Not_Computer
Z-1722-2011 | 55875 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 20 defibrillator/monitor and LIFEPAK 20e. Product part numbers: 3202487-xxx, 3202488-000, U3202487-000, 3202488-015, and U3202488-000. |  |  | May 27, 2010 | 2010 | Physio Control, Inc. | A component failure on the AC power supply assembly results in "NO AC POWER" operation.A failure of AC power can result in a delay of defibrillation therapy if backup battery power is depleted. | Device Design | PhysioControl, Inc. began sending an Urgent Medical Device Correction letter via certified receipt to all affected customers describing the potential issue. Customers were advised to keep the defibrillators in service and to follow the daily Operator's Checklist. All power supplies will be updated.For questions customers were instructed to contact Technical Support at 1-800-442-1142.For questions regarding this recall call 425-867-4000. |  | 53550 | 53550 units | Worldwide distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MI, MD, MI, NC, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY and the coutnries of ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Terminated | September 29, 2014 | 1586 | 53550 | Battery | Treatment/Delivery/Therapy | Instructions on cheklists/Power supply update | Safety Notice/Insructions
Z-1732-2011 | 57615 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Acute Care Troponin I TestPak, Troponin I Assay |  |  | December 23, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Low frequency of non-repeatable falsely elevated CTNI results without an associated error message. | Other | The firm, SIEMENS, sent an "Urgent Field Safety Notice" letter dated December 2010 to all Stratus(R) CS Instrument customers. The letter described the product, problem and actions to be taken. The customers were instructed to repeat testing of samples with CTNI results above 0.07 ng/mL; discuss content of this letter with their laboratory director regarding the need to review test results reported from CTNI TestPak lots 230200002 or greater; complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to Technical Solutions Center at 302-631-8467, and forward this notification to anyone to whom they may have distributed this product.If you have any questions, please contact the Siemens Technical Solutions Center at 800-405-6473. |  | 11635 | 11635 | Worldwide distribution: AL, AK, AZ, AR, CA, CO, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY.Product was distributed to foreign accounts in Algeria, Armenia, Australia, Austria, Bahrain, Bangladesh,Belgium, Bosnia &amp; Herzegovina, Canada, Canary Islands, Czech Republic, Cyprus, Croatia, Denmark, Finland, France, Germany, Great Britain, Greece, Guadeloupe, Indonesia, India, Israel, Ireland, Italy, Japan, Libya, Malaysia, Morocco, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Arab Emirates. | Terminated | March 19, 2013 | 817 | 11635 | Not_Computer
Z-1731-2011 | 57757 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter MAXM Hematology Analyzer and MAXM Hematology Analyzer with Autoloader. |  |  | December 14, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the Coulter MAXM, MAXM AL, HmX and HmX AL Analyzers omit the tilde (~) character when thetilde is used as part of the Sample 10 within a barcode label scanned by the primary mode barcode reader. There is a potential for Specimen or Patient misidentification to occur. | Other | The recall communication was initiated on 12/14/2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTER HmX Hematology Analyzer, COULTER HmX Hematology Analyzer with Autoloader, COULTER MAXM Analyzer, COULTER MAXM Hematology Analyzer with Autoloader. The letter provides the customers with an explanation of the problem identified and an action to be taken.Customers were instructed not to use the tilde (~) character in Specimen or Patient identifiers. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 4288 | 4288 units total | Worldwide Distribution -- US, Algeria, Andorra, Antigua and Barbuda, Australia, Austria, Bahrain, Bangladesh, Bermuda, Bosnia and Herzegovina, Botswana, Brunei Darussalam, Bulgaria, Burkina Faso, Burundi, Canada, Cayman Islands, China, Congo, Cote d'Ivoire, Croatia, Czech Republic, Djibouti, Egypt, Finland, France, French Guiana, French Polynesia, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Italy, Japan, Jordan, Republic of Korea, Kuwait, Latvia, Lebanon, Libyan Arab, Jamahiriya,Lithuania, Macao, Malawi, Mexico, Mongolia, Morocco, Namibia, Netherlands, New Caledonia, New Zealand, Nigeria, Oman, Pakistan, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia,Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sudan, Swaziland, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, U.S. Virgin Islands, Yemen, and Zambia. | Terminated | September 02, 2014 | 1358 | 4288 | Not_Computer
Z-1724-2011 | 57674 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | W Vital Health Pulsating Clean Rechargeable Toothbrush; Distributed by: Walgreen Co., 200 Wilmot Rd., Deerfield, IL 60015-4616, Made in China; |  |  | January 14, 2011 | 2011 | Brushpoint Innovations, Inc. | The power toothbrush may make a loud noise and the battery cap becomes separated from the main body of the toothbrush. The battery cap can become a projectile that could cause potential injury if it comes in contact with the body. | Device Design | The firm, BrushPoint Innovations, sent a email dated January 7, 2011 notifying their customers of the recall. The email described the product, problem and actions to be taken. The customers were instructed to pull and quarantine the toothbrushes due to potential product quality concerns, and hold their stocks until January 14, 2011, call in the number of toothbrushes quarantined, and then return them to the Walgreen distribution centers. A stop sale was placed on the affected item numbers in their cash register scanners.If you have any questions, call 905-833-5122 or 905-944-1000. |  | 8100 | 8,100 units | Nationwide distribution. | Terminated | December 16, 2012 | 702 | 8100 | Not_Computer
Z-1041-2011 | 57784 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE ESSENTIAL FULL FIELD MAMMOGRAPHY SYSTEM (SENO ESSENTIAL) |  |  | January 19, 2011 | 2011 | GE Healthcare, LLC | During installation it was discovered that the duplicate of the 21 CFR 1010 Certification label for the X-ray Generator Cabinet was missing from the installation kits that allow installation of the Senographe Essential in a van (mobile unit). | Radiation Control for Health and Safety Act | The firm plans to repair the defect or to bring product into compliance. The field action is being implemented to inspect the affected units and if found to be missing the certification label, a certification label will be applied. This will be carried out via a Field Modification Instruction planned to be released in March, 2011 with an estimated completion date of July, 2011. |  | 14 | 14 | Nationwide Distribution | Terminated | May 16, 2012 | 483 | 14 | Not_Computer
Z-1726-2011 | 57778 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PhD System EIA/IFA Software; Model/Catalog Number: 426-0246; |  |  | January 20, 2011 | 2011 | Bio-Rad Laboratories, Inc. | Firm's Correction involves advising customers who have implemented a Laboratory Information System (LIS) that they need to ensure quantitative results have passed all validation criteria prior to reporting patient results. When a patient result fails a Validation Rule, the qualitative result will indicate " INVALID". | Other | The firm, Bio-Rad, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 20, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the PhD Worklist Results to verify that all validations have passed and that the quantitative results are valid; to translate (if necessary) the customer letter, add local contacts information, distribute to all active PhD System customers in their region and provide tracking documentation to CSD-RA for all customers who received the customer letter.If you have any questions, contact the Bio-Rad CSD Regulatory Affairs Department at (510) 741-4618. |  | 650 | 650 units | Worldwide distribution: USA and countries including:Canada, China, Australia, New Zealand, Czech Republic, France, Hong Kong, India, Italy, Korea, Singapore, Sweden, Taiwan and United Kingdom. | Terminated | August 17, 2011 | 209 | 650 | Not_Computer
Z-1737-2011 | 56639 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus PLAC Test ELISA Kit, Enzyme Immunoassay for Quantitative Determination of Lp-PLA2 in human plasma or serum, part number 90123, manufactured by diaDexus Inc, South San Francisco. |  |  | August 20, 2010 | 2010 | DiaDexus, Inc | Product fails stability tests prior to expiration date. | Nonconforming Material/Component | A phone script has been prepared to contact customers. Customers were informed of the affected product and its reason for the recall. If customers didn't use the product after May 11, 2010, the action is considered complete. If the product was used after May 11, 2010, a follow up letter will be sent to the customer and they will be asked to re-test any samples beyond May 11, 2010 with an alternate lot of reagent. |  | 482 | 482 units | Worldwide Distribution -- US, including states of UT, OH, CA, NJ, NC, FL, OK, MA, NY, TX, NJ, WA, and MI and countries of Germany and Israel. | Terminated | March 21, 2011 | 213 | 482 | Not_Computer
Z-1736-2011 | 57443 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Total BhCG (BhCG) Assay Part Number 33500 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter has confirmed that non-reproducible falsely elevated test results may occur when using the Access Total BhCG assay. These elevated results are often attributed to pre-analytical factors and are particularly noted at the low end of the analytical measuring range.The role of preanalytical factors in laboratory testing has been described in a variety of published literature. | Pending | Beckman Coulter sent an Urgent Product Corrective Action letter dated November 11, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:." Evaluate current laboratory protocols to ensure the necessary procedures are in place for identifying and mitigating falsely elevated test results." Follow the sample handling instructions provided in the Access Total ÂżhCG Instruction for Use (IFU)." Review the enclosed pre-analytical sample handling information bulletins:1. The Role of Preanalytical Factors in Chemistry and Immunoassay Testing (P/N A34880A).2. Sample Handling Makes a Difference (P/N A28226A)." At the discretion of the laboratory director, any total ÂżhCG result that is questioned due to lack of clinical correlation should be repeated.NOTE: Due to the clinical utility of the Access Total ÂżhCG assay as a biomarker of pregnancy, Beckman Coulter feels it is unnecessary to require a review of all previously obtained Access Total ÂżhCG test results. The Access Total ÂżhCG assay is intended as a marker for pregnancy at the time of testing. A discrepant or questionable result would have been brought to the attention of the laboratory. The pregnancy status of a patient at this point, post testing, would be known to the patient and the physician." Please share this information with your laboratory staff and your phlebotomy director and/or phlebotomy staff and retain this notification as part of your laboratory Quality System documentation." Complete and return the enclosed response form within ten days so they can be assured that this important notification has been received.If you have any questions regarding this notification, please contact the Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call1-800-854-3633 in the United States and Canada. Outside the United States and Canada |  | 158919 | 158,919 units | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Antigua and Barbuda, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Congo, Croatia, Czech Republic, Djibouti, Egypt, Ethiopia, Finland, France, French Polynesia, Georgia, Germany, Ghana, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Moldova, Monaco, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Zambia, Zimbabwe. | Terminated | September 02, 2014 | 1413 | 158919 | Not_Computer
Z-1738-2011 | 57809 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thrombelastograph Hemostasis Analyzer System; a multipurpose system for in vitro coagulation studies; Haemonetics, Haemoscope Division, Niles, IL 60714; list number 07-022 |  |  | January 28, 2011 | 2011 | Haemoscope Division of Haemonetics Corp | Haemonetics has received a report of smoke from the unit with a power supply failure. | Other | Haemoscope sent letters to all of their U.S. accounts on January 24, 2011, via certified mail. The letters listed guidelines regarding proper use of the TEG 5000 System Power Supply to eliminate any potential risks such as electrical shorting or power supply damage. The information, which was printed in the letter, will also be available in an updated TEG User Manual soon to be released. A Caution label with installation instructions was included with the letter, to be applied to the back of the TEG unit. The accounts were requested to complete the enclosed acknowledgement form and fax it back to Haemoscope at 847-588-0455. Any questions were directed to 1-800-438-2834. |  | 1155 | 1,155 units | Nationwide Distribution including CA, FL, MD, PA, OH, SC, TN, TX, and VA | Terminated | March 12, 2012 | 409 | 1155 | Battery | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-1734-2011 | 57232 | microtome, cryostat | IDP | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific HM 550 Cold Disinfection Cryostat, Microm International GmbH, Waldorf, Germany |  |  | September 23, 2010 | 2010 | Thermo Fisher Scientific | Through internal quality review, firm determined the need to conduct a more comprehensive system validation to support the claim of complete bacterial, fungal and viral disinfection efficacy. | Device Design | ThermoFisher sent Advisory Notice letters beginning on September 28, 2010, to all affected customers. The letters stated that the firm needed to conduct more comprehensive system validation data on the Cold-Disinfection sytem option available on the HM550-D cryostat to support the claim of complete bacterial, fungal and viral disinfection efficacy. The letter recommended that the Cold Disinfection option on the HM550-D not be used as the sole disinfection method on the HM550-D until the validation work is completed. The firm provided addtional safety protocols pertaining to the cryostat disinfection. Customers were instructed to complete and fax the Confirmation of Receipt Form to Technical Support, COLD-D Advisory Notice at (269) 372-2436. |  | 100 | 100 units | Worldwide Distribution - USA including the states of LA, OH, FL, MD, PA, TX, MO, WI, KS, AL, CA, VA, MS, IL, NE, MA, and the countries of Australia, Belgium, China, Czech Republic, Denmark, Estonia, Finnland, France, Germany, Hong Kong, India, Italy, Japan, Netherlands, Norway, Poland, Portugal, Russia, Spain, Sweden, Taiwan, and the UK. | Terminated | April 02, 2012 | 557 | 100 | Not_Computer
Z-1739-2011 | 57366 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Gentamicin (GENT) Flex(R) reagent cartridge, K4012 |  |  | November 30, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | False elevation of Dimension Vista(R) Gentamicin results for patients and QC when prior reagent aspirated by the reagent probe is any of the following plasma protein methods: ASL, B2MIC, B2MU, CRP,CYSC, HCYS, hcCRP, IGE, IGG3, IGG4, RF, and STFR resulting in inappropriate patient treatment. | Other | Siemens Healthcare Diagnostics, Inc. sent an Urgent Field Safety Notice letter dated November, 2010, to all affected customers. Customers were requested to run GENT on a Dimension Vista(R) system that does not run the listed plasma protein methods. Instructions were provided for disabling methods on the instrument to prevent inadvertently testing Gent along with plasma protein methods.Customers were instructed to :Forward the notification to anyone to whom they may have distributed the product. Complete and return the Field Correction Effectiveness Check Form to Siemens Healthcare Diagnostics at (302) 631-8467.For further information contact Siemens Technical Solutions Center at 1-800-441-9250 or their local Siemens representative. |  |  |  | Nationwide Distribution including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | March 19, 2013 | 840 |  | Not_Computer
Z-1743-2011 | 56397 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON System(s) Chemistry RF Rheumatoid Factor Part Number 475136 |  |  | June 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter testing indicated that Rheumatoid Factor (RF) reagent lot M907325 will not remain stable to the labeled expiration date. | Other | The firm Beckman Coulter, sent a "URGENT: PRODUCT CORRECTIVE ACTION" letter dated June 25, 2010 to the consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions:*Discontinue use and discard appropriately Rheumatoid Factor (RF) reagent lot M907325*Contact their Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada to discuss reagent replacement options.*Complete and return the FAX BACK RESPONSE FORM via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd., Brea, California 92821, Mail Code E2.SE.08Consignees/customers were asked to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If consignees/customers forwarded the affected products listed above to another laboratory, then they need to also provide a copy of the Product Corrective Action letter to the laboratory.If they have any questions regarding this Product Corrective Action letter, consignees were told to call Beckman Coulter Customer Support Center (Hotline) at (800) 854-633 in the United States and Canada. |  | 58 | 58 units | Worldwide distribution: USA including states of: AZ, CA, CO, IA, PA, NH, NV, NY, and WA; and countries of: Belgium, China, Czech Republic, France, Germany, Hong Kong, India, Italy, Japan, Kuwait, Lebanon, Lithuania, Mexico, Netherlands, Russian Federation, South Africa, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | July 02, 2012 | 742 | 58 | Not_Computer
Z-1368-2011 | 57106 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura XPER FD 10 and Field Extension for XPER Cardiovascular System and Allura XPER FD 20 and Allura CV20. |  |  | November 04, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Adjustment of the micro-switches in the footswitch used on the Philips cardiovascular X-ray systems was not implemented. If it is programmed for exposure, it is possible that after pressing the pedal on the footswitch it does not release itself. The patient may be exposed to X-ray longer than intended. | Pending | On 11/5/10, Philips began sending the Electronic Product Radiation Warning notification letter, dated November 4, 2010, to their consignees.The consignees were informed that the middle pedal of the foot switch of the Allura FD10, FD20, CV20 systems, and the Allura Field Extension series does not not always become inactive (sticks) when pedal is released. The function of the middle pedal on the foot switch is freely programmable by the customer. If it is programmed for exposure and if the pedal does not release by itself, the patient will be exposed to X-ray longer than intended.The consignees were advised to check if their footswitch is possibly affected. Footswitch identification 452270000141 is noted. If the pedal does not release by itself press the red emergency stop button on the system.Consignees can call Philips Healthcare Call Center at 800-722-9377, #5 and Reference FCO 72200189. |  | 4 | 4 units | Worldwide Distribution -- US, including states of CA, IL, MA, NC, and OC. | Terminated | July 21, 2011 | 259 | 11 | Not_Computer
Z-1622-2011 | 57947 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F250NRe Hemodialyzer, Single Use, Catalog Number: 0500325E, Advanced Fresenius Polysulfone. |  |  | February 16, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | There is an increased risk of an internal blood leak. | Process control | FMCNA notified consignees on 2/16/11 by telephone and follow-up with an Urgent FMCNA Optiflux F250NRe Recall letter via Certified Mail, return Receipt requested. The firm states in the letter that there is an increased risk of an internal blood leak affecting a small percentage of the two lots of dialyzers. They are asking customers to examine their inventory, to discontinue use, and segregate the affected product in a secure place. Customers are to contact their FMCNA Customer Service Team at 1-800-323-5188 for instructions on how to return product. Used dialyzers should not be returned. If there are any questions, Customer Service should be contacted. |  | 38364 | 38,364 units | Worldwide Distribution -- USA and Canada. | Terminated | February 12, 2013 | 727 | 38364 | Not_Computer
Z-1643-2011 | 57916 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | AFFIXUS Lag Screw Drill, REF 2112-01-303, Non-Sterile, DePuy Orthopaedics, Inc., Warsaw, IN; DePuy International Leeds, England. |  |  | January 12, 2011 | 2011 | Depuy Orthopaedics, Inc. | The firm recieved a complaint that the cannula for the drill is not the same size at both ends. The drill is grabbing the guidewire and advancing and not allowing to finish reaming, resulting in the possibility of the guide wire progressing through and penetrating the acetabulum. | Other | On 1/12/2011 DePuy contacted their distributors by telephone and directed the distributors to cancel all pending surgeries and discontinue use of the sets until the investigation is completed. On 1/14/2011 DePuy again contacted all distributors by telephone and directed all distributors to return the instruments to DePuy for inspection as part of the investigation. |  | 296 | 296 units | Worldwide Distribution -- US, including the states of AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, NC, NV, NY, OH, PA, TN, TX, VA, WA, WI; and country of Switzerland. | Terminated | June 27, 2012 | 532 | 296 | Not_Computer
Z-1744-2011 | 57937 | No | LOM |  |  | Microbiology | PMA | ETI-AB-AUK PLUS kits,REF P001931, DiaSorin S.p.A., 13040 Saluggia (VC), Italy, Distributed by: DiaSorin Inc., 1951 Northwestern Avenue, P.O. Box 285, Stillwater, MN 55082-0285 |  |  | August 19, 2010 | 2010 | Diasorin Inc. | Through internal testing, DiaSorin has become aware that the kit lots identified (ETI-AB-AUK PLUS kits, catalog number P001931, lot numbers 0650270A and/or 0650270A/1) may not consistently exhibit stability to defined expiration dating, resulting in decreased binding and lower absorbance. Internal testing has confirmed that when defined assay validation criteria pass, analytical performance is co | Other | The firm, DiaSorin, sent a "Customer Notification Letter" dated 19 August, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to destroy all remaining inventory and complete and return the receipt of notification and destruction of product form, found in the bottom portion of the letter, via fax to: (651) 351-5669, ATTN: TECHNICAL SERVICES or mail to: ATTN: TECHNICAL SERVICES, 1951 Northwestern Ave., Stillwater, Minnesota 55082-0285.If you have any questions or concerns, please contact DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. |  | 710 | 710 | Worldwide distribution: USA including states of: AL, AR, CA, CT, DE, DC, FL, GA, IL, IA, KY, LA, MD, MA, MI, MN, MO, NH, NJ, NM, NY, NC, ND, OH, OK OR, PA, TX, VT, VA, WA,and WI; and country of: SOUTH KOREA. | Terminated | August 16, 2012 | 728 | 710 | Not_Computer
Z-1723-2011 | 57960 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R4.5 |  |  | October 25, 2010 | 2010 | Elekta, Inc. | If another patient is selected on the imaging system (XVI) release 4.5) during the transmission of end of treatment data from Desktop Pro R7.01 to R&V; system, the XVI information is prioritized and the end of treatment data is never received by the R&V; system. | Other | Important Notice A341 titled "Mandatory Upgrade Desktop Pro R7.01 Service Pack and Cancelled Notices" dated January 26, 2011, is in distribution to all affected customers. Field Change Order 20000501011, "Service Pack 2 upgrade to Desktop Pro R7.01 Service Pack 1", dated October 25, 2010 will be utilized by Elekta Service personnel to upgrade the Desktop Pro to include Service Pack 2. The Important Notice letter instructed customers to file the letter in the Important Notice section of the appropriate User Manual. Before doing so, users are to ensure that Desktop Pro R7.01 SP2 is installed on their treatment control system. If it is not installed, customers are to contact their local Elekta representative. Once the installation of Service Pack 2 has taken place, the notices listed in Table 1 of the letter will no longer be applicable and should be removed from the User Manual. Customers should direct their questions to their local Elekta representative. |  | 13 | 13 units | Class 2 Recall -- Nationwide Distribution -- Including CA, MI, MO, ND, OH, OR, RI, VT, and WI. | Terminated | May 24, 2011 | 211 | 13 | Software | Device Operation | Software update | Software Update
Z-1745-2011 | 52904 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Siemens Medical Solutions USA, Inc., ACUSON AcuNav ultrasound catheters when using SwiftLink catheter connectors; |  |  | March 20, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Screen Image Quality. An assembly issue on SwiftLink connectors may cause reduced/degraded image quality when using ACUSON AcuNav ultrasound catheter transducers. | Process design | Siemens Field correction was initiated on 3/20/2009. A device correction letter was distributed to all affected users, with a description of the problem and its corrective actions. Siemens service personnel will make a site visit. |  | 21 | 21 units | Distribution in US and Internationally. | Terminated | April 01, 2011 | 742 | 21 | I/O
Z-1746-2011 | 57824 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf Interactive burn and wound dressing, supplied as a living, bi-layered skin substitute |  |  | January 30, 2011 | 2011 | Organogenesis, Inc. | Product sterility compromised | Material/Component Contamination | Organogenesis notified accounts on Sunday, January 30, 2011, via fax and phone. The Company will collect written confirmation from the affected customers that they have been notified of the recall. Sales representatives for the affected customers have been directed to visit the customers as soon as possible Monday morning to assist in the recall and prevent additional applications of the product. Apligraf Medical Affairs personnel will provide clinical support to treating physicians, and update them as new information on microbial contamination and antibiotic susceptibility becomes available. Units not applied have been requested to be returned to Organogenesis |  | 50 | 50 units | nationwide. | Terminated | July 28, 2011 | 179 | 50 | Not_Computer
Z-1748-2011 | 58137 | No | NEZ | Ingestible telemetric gastrointestinal capsule imaging system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Given Imaging Limited, PillCam Express Delivery Device, For use with PillCam SB, endoscopic imaging device. Part numbers FGS-0367 (3-Pak) or FGS-0366 (1-Pak). |  |  | March 08, 2011 | 2011 | Given Imaging Ltd. | The capsule holder can break away from the catheter during endoscopic procedures, necessitating retrieval. | Device Design | Given Imaging, Inc. consignees were notified via FedEx letter on/about 03/08/2011. They instructed to return all affected product. A Customer Response form with return instructions was included to be completed and returned to Given Imaging. |  | 389 | 389 | U.S. Nationwide, Canada, Germany Israel and Australia. | Terminated | December 02, 2011 | 269 | 389 | Not_Computer
Z-1747-2011 | 58138 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf - Interactive burn and wound dressing, supplied as a living, bilayered skin substitute |  |  | February 21, 2011 | 2011 | Organogenesis, Inc. | Product sterility may be compromised | Packaging process control | Organogenesis contacted the accounts by telephone and fax on 2/24/11 to advise of the contamination and identification fo the organism for appropriate patient follow-up. |  | 7 | 7 units | Nationwide. | Terminated | July 28, 2011 | 157 | 7 | Not_Computer
Z-1796-2011 | 57177 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | ASR 300 Size 46, DEPUY ASR 300 ACETABULAR CUP SYSTEM. |  |  | August 23, 2010 | 2010 | Depuy Orthopaedics, Inc. | As part of our ongoing post-market surveillance of all products, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports.DePuy Orthopaedics issued a Field Safety Notice in March 2010 after receiving new data from the | Device Design | DePuy Orthopaedics issued a Field Safety Notice in March 2010 to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the Field Safety Notice with their organization and any organization where the ASR products may have been transferred. Customer were instructed not to implant the ASR devices. Patients who received the ASR System should be informed of this recall and instructed to return for a follow up visit.Patients with radiographic changes indicative of product failure should be addressed according to normal procedures.All other patients should be followed according to the April 22, 2010, and May 25, 2010 UK Medicines and Healthcare products Regulatory Agency (MHRA) Device Alerts. Follow up for all patients implanted with ASR acetabular cups at least annually for five years postoperatively. Beyond five years, follow up in accordance with locally agreed protocols.For patients who are symptomatic or implanted with a cup angle greater than 45 degrees, particulary where a small component has been implanted.:" Consider measuring cobalt and chromium ion levels in whle blood and/or performing cross sectional imaging including MRI or ultrasound scans." If metal ion levels in whole blood are elevated above 7 parts per billion (ppb) for either metal ion, a second test should be performed three months after the first in order to identify patients who require closer surveillance, which may include cross sectional imaging" If MRI or ultrasound scan reveals soft tissue reactions, fluid collections or tissue masses, then revision surgery should be considered.DePuy representatives were to assist with returns of any remaining inventory.For questions regarding this recall call 574-372-7333. |  | 270 | 270 | Worldwide | Terminated | August 29, 2013 | 1102 | 120984 | Not_Computer
Z-1818-2011 | 57946 | No | JNL | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | Access Thyroglobulin Antibody II, Part Number: A32898 |  |  | January 26, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed customer reports of reagent pack to pack variability for the Access Thyroglobulin Antibody II (TgAb) reagent lot numbers identified in the recall. Erroneous results for patient and quality control samples can be generated. The ThyroglobulinAntibody II (TgAbII) reagent lots may produce:(1) Erroneously high TgAb results for both quality co | Other | Beckman Coulter sent a Product Corrective Action (PCA) letter on January 26, 2011 via US Postal Service to the affected customers. The letter explained the problem identified and the action to be taken.Customers were instructed to:(1) Discontinue use of Access Thyroglobulin Antibody II reagent lot numbers identified in the tableabove;(2) If customers obtained questionable thyroglobulin antibody results, or received feedback from clinicians on reported thyroglobulin antibody results using the reagent lot numbers identified above, rerun or request a redraw of the patient sample;(3) Share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded any of the affected product listed above to another laboratory, they should provide a copy of the letter to them;(4) Complete and return the enclosed response form within 10 days so the firm can be assured that you have received this important notification and are taking appropriate action.For questions regarding the recall communication, customers were told to contact Beckman Coulter Customer SupportCenter:(1) Via our website, http://www.beckmancoulter.com/customersupport/support(2) Via phone, call 1-800-854-3633 in the United States and Canada(3) Outside the United States and Canada please contact your local Beckman Coulter representative. |  | 3658 | 3658 units (1800 in US) | Worldwide Distribution - USA including AR, CA, FL, GA, HI, IA, IL, KY, MA, MI, MN, MO, MS, NC, NJ, NM, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and the countries of Argentina, Armenia, Australia, Belgium, Brazil, Bulgaria, Canada, China, Czech Republic, France, French Polynesia, Georgia, Germany, Greece, Hungary, Israel, Italy, Kazakhstan, Kuwait, Lebanon, Morocco, Netherlands, New Zealand, Oman, Poland, Puerto Rico, Romania, Russian Federation, Serbia, South Africa, Spain, Switzerland, Taiwan, Tunisia, Turkey, Ukraine, United Arab Emirates, Yemen | Terminated | May 04, 2012 | 464 | 3658 | Not_Computer
Z-1819-2011 | 57962 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation (EBW) a component of the Computed Tomography X-Ray System Model Numbers EBW (CT) 728260 &amp; EBW (Portal) 728269. |  |  | October 03, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | During an FDA inspection on 10/1/2010 the firm was notified by FDA that their field action they initiated in Oct 2009, constituted a recall. The firm determined that the functional analysis feature in the Comprehensive Cardiac Analysis (CCA) was displaying incorrect values on the polar fields. The calculated results for the myocardial wall thickening was incorrect in that the numbers displayed f | Other | Philips sent an URGENT - Field Safety Notice letter dated September 29, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware that the numbers displayed for "Wall Thickening" actually represent a reduction in wall thickness from ES to ED.Philips will be providing a sotware update to resolve the issue. For further information or support customers were instructed to contact their local Philips representative:Customer Care Service Center in the U.S. at 1-800-722-9377, option 5:Diagnostic Imaging option 1: CT or their local Philips representative. |  | 2068 | 2,068 units | Worldwide Distribution - USA including the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX , UT, VA, VT, WA, WI, WV &amp; WY and the following countries: ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BANGLADESH, BELGIUM, BRAZIL, BRUSSEL, CANADA, CHILE, CHINA, COLUMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EDUADOR, EGYPT, FRANCE, GERMANY, GREECE, GUADELOUPE, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, LATVIA, LEBANON, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, MIRATES, UNITED KINGDOM, VENEZUELA &amp; YEMEN. | Terminated | September 21, 2012 | 1084 | 2068 | Software | Output/Calculation | Software update | Software Update
Z-1820-2011 | 57564 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 S/N 9089, 9551, 9548. |  |  | December 20, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | During the version 2.6 software upgrade, the system was not recalibrated with the infant phantom for the infant protocol. | Other | Philips sent an URGENT - Medical Device Correction letter dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers may consider reviewing any infant images that may have been affected during the Field Test to confirm reading accuracy. For further information or support concerning this issue, customers could contact their local Philips representative or local Philips Healthcare office for North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts) |  | 3 | 3 | Worldwide Distribution - USA including CO, IN, and TN and the countries of Norway, Israel, and Germany | Terminated | February 22, 2013 | 795 | 3 | Software | Device Operation | Instructions to review images accuracy | Safety Notice/Insructions
Z-1725-2011 | 58123 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Millar Instruments, Inc. Millar Sensors Systems Solutions Mikro-Tip Angiographic Catheters Sterile and Nonpyrogenic Model SPC-454D and SPC454F. |  |  | March 08, 2011 | 2011 | Millar Instruments, Inc | Possibility of epoxy debris particles inside the catheter lumen. | Process design | The firm, Millar, notified their consignees/customers by email and sent hardcopy of email on March 8, 2011. The notification described the product, problem and actions to be taken. The customers were instructed to return the catheter(s) to Millar or report back to Millar if the catheter(s) has been consumed.If you have any questions, please do not hesitate to contact Director, RA/QA at 832-667-7000. |  | 1080 | 1080 units | MN, NY, Washington DC, NC, MA, MN, PA, CA, MI, Australia, Germany, Canada, Japan, United Kingdom, and Norway. | Terminated | June 27, 2011 | 111 | 1080 | Not_Computer
Z-1824-2011 | 57455 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Reagent, Calcium Test System for cobas c 111, Part number 04718933190, Roche Diagnostics Corporation, Indianapolis, IN. |  |  | October 11, 2010 | 2010 | Roche Diagnostics Operations, Inc. | Internal investigations have confirmed imprecision and erroneous patient results (up to 35% discrepant results) for specific lots of the Calcium assay.Internal investigations revealed the presence of precipitate in the R2 of the affected calcium reagent lots. The precipitate observed was classified as o-cresolphthalein complexone (o-CPC), a reactive component of the calcium R2.The root cau | Other | The firm, Roche, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated October 11, 2010 to all consignees/customers via UPS(signature required) . The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the recalled product; discard from inventory according to their site's local regulations; complete and return the attached Faxback form via fax to 1-877-598-5705 even if they do not have the affected product, and file this letter for future reference. The firm added lot 63020401 to the recall in 3/2011. The firm is mailing an URGENT MEDICAL DEVICE REMOVAL notice to customers that purchased that lot of product.Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have questions about the information contained in this letter. |  | 14528 | 14,528 units | Nationwide distribution. | Terminated | September 13, 2011 | 337 | 27933 | Not_Computer
Z-1821-2011 | 58017 | No | JCY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF 24k100***QTY 1***24k Arthroscopy Outflow/Suction Tubing Set***MADE IN USA***Rx Only***CONMED LINVATEC, CONMED LINVATEC 11311 CONCEPT BLVD, LARGO, GL 33773-4900*** |  |  | September 10, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | On 09/13/2010 ConMed Linvatec, Largo, FL initiated a recall on 24k100 Arthroscopy Outflow/Suction Tube Set used with the ConMed Linvatec 24k Pump. The 24k100 Arthroscopy Outflow/Suction Tubing Set Lots 22747-000, 22748-000, and 2749-000 may have been incorrectly assembled and could potentially allow waste to enter the sterile field. | Other | Linvatec Corp. dba ConMed Linvatec sent an Urgent Medical Device Recall Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their facility's inventory for the affected product. If they had any of the product they were instructed not to use and segregate the product and return for credit.Return the affected products and the completed REPLY FORM to their local ConMed Linvatec office.If the product was transferred to another facility, customers were instructed to complete the REPLY FORM documenting the transfer of the product and mail or fax the REPLY FORM to their local ConMed Linvatec office at 727-319-5701.If the customers do not have the affected product they should complete the REPLY FORM and mail in the enclosed envelope or fax to 727-319-5701. |  | 4598 | 4,598 | Worldwide Distribution - USA (nationwide) and internationally to Saudi Arabia, Italy, France, Canada, Portugal, South Africa, Switzerland, Denmark, Germany, New Zealand, Australia, and Spain. | Terminated | January 31, 2012 | 508 | 4598 | Not_Computer
Z-1822-2011 | 58002 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ossur Total Knee, Model Number: 2100 |  |  | February 11, 2011 | 2011 | Ossur Engineering, Inc | The recall was initiated based on a finding by Ossur that some units of the Total Knee 2100 produced from October 19th, 2010 until January 21st 2011, contained back links that may compromise knee function prematurely. | Process design | Ossur initiated a Safety Notice (Voluntary Product Recall) letter on February 11, 2011, via US Postal Service Registered Mail. The recall letter informed the affected customers of the reason for the recall and the product affected.Ossur requested their distributor and providers take the following action: (1) Please identify the users of each of the Total Knee 2100 model devices listed above; (2) Also, please check your inventory and isolate any recalled devices you have on hand but have not fitted on users. Customers were also informed that a representative from Ossur customer service will be contacting them to coordinate replacement of the affected devices.The letters were followed by phone calls placed by a trained Ossur Customer Service Representative to each customer after US distributors provided Ossur with customer contact information.For questions regarding this recall call 517-629-8890. |  | 142 | 142 units | Worldwide Distribution - USA including CA, CO, FL, LA, MD, MI, MO, NJ, NY, OH, TN, VA,, and WA and the countries of Australia, Canada, Czech Republic, Germany, France, United Kingdom, Ireland, Italy, Japan, The Netherlands, Russian Federation, Sweden and Turkey. | Terminated | February 29, 2012 | 383 | 142 | Not_Computer
Z-1834-2011 | 54706 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Coherence Dosimetrist, version 2.2. |  |  | November 06, 2009 | 2009 | Siemens Medical Solutions USA, Inc | A problem in the software has the potential to be associated with the dose; when a template is created the user has the possibility to predefine organs. For the Target the user defines the max. dose and the min dose, for the OAR (Organs at Risk) the user only defines the max. dose. The firms engineering team determined that the software defect was repeatable. | Software design | Affected consignees will be notified by the Customer Service Engineer to schedule a software update beginning November 2009. |  | 60 | 60 units | Worldwide distribution: USA including: CA, KY, MI, NJ and TN; and countries of: Canada and Mexico. | Terminated | April 06, 2011 | 516 | 60 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1832-2011 | 54842 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson S2000 Ultrasound System with software and imaging. |  |  | March 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | "Any 2D measurement performed in Exam Review or offline, on a zoomed Retrospective clip capture may be erroneous. This applies to 2D, 2D/PW, 2D/CW and 2D/M-mode." | Software design | Siemens Medical Solutions USA, Inc. sent a Consumer Safety Advisory Notification letter on February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the Customer Safety Notice to all those within their organization who need to be aware of the issue until the corrective action was completed. For questions regarding this recall call 650-694-5580. |  | 1903 | 1903 | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Australia, Belgium, Bolivia, Brazil, Bahamas, Canada, Switzerland, Chile, China, Colombia, Czech Republic, Germany, Denmark, Ecuador, Eguypt, Spain, Finald, France, Great Britian, Germany, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Ireland, Italy, Jordan, Japan, Kenya, South Korea, Kuwait, Lithuania, Mexico, Malaysia, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, Slovak Republic, Thailand, Turkmenistan, Tunisia, Turkey, Taiwan, Ukraine, Venezeula, Yemen, and South Africa.. | Terminated | June 18, 2012 | 840 | 1903 | Software
Z-1837-2011 | 56840 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound systems (all versions) with cardiovascular options installed; |  |  | August 17, 2010 | 2010 | Siemens Medical Solutions USA, Inc., Mountain View, CA | There is an error in the estimation of mean pressure gradient in the Cardiac measurement package. The currently implemented calculation will result in an underestimation of the mean pressure gradient (PGmean) when estimated from a trace of the Doppler waveform of the stenotic flow. | Software change control | Siemens Medical Solutions USA, Inc., sent a Customer Safety Advisory Notification letter on August 17, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the notice on to all those within their organization who need to aware of the issue until the corrective action was completed. For questions customers were to contact their local service support person. For questions regarding this recall call 650-694-5993. |  |  |  | Worldwide Distribution - USA (nationwide)and the countries of Antares distribution includes: United Arab Emirates-AE, Austria-AT, Australia-AU, Azerbaijan-AZ, Brazil-BR, Belarus-BY, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Egypt-EG, Spain-ES, France-FR, Great Britain-GB, Guatemala-GT, Croatia-HR, Hungary-HU, Indonesia-ID, Ireland-IE, India-IN, Italy-IT, Jordan-JO, Japan-JP, South Korea-KR, Moldova-MD, Mexico-MX, Malaysia-MY, Netherlands NL, Norway-NO, New Zealand-NZ, Poland-PL, Portugal-PT, Romania-RO, Serbia-RS, Russian Federation-RU, Saudi Arabia-SA, Sweden-SE, Singapore-SG, Slovak Republic-SK, Suriname-SR, Turkey-TR, Taiwan-TW, Ukraine-UA, United States-US, Uruguay-UY, Uzbekistan-UZ, Venezuela-VE, Vietnam-VN, South Africa-ZA S2000 distribution includes: Austria-AT, Australia-AU, Brazil-BR, Canada-CA, Switzerland-CH, Chile-CL, China-CN, Czech Republic-CZ, Germany-DE, Denmark-DK, Spain-ES, Finland-FI, France-FR, Great Britain-GB, Hong Kong-HK, Indonesia-ID, India-IN, Italy-IT, Japan-JP, South Korea-KR, Kazakhstan-KZ, Lebanon-LB, Mexico-MX, Norway-NO, Philippines-PH, Poland-PL, Portugal-PT, Romania-RO, Russian Federation-RU, Sweden-SE, Slovak Republic-SK, Turkmenistan-TM, Turkey-TR, Taiwan-TW, United States-US X300 distribution includes: AE (United Arab Emirates) , AL (Albania) , AN (Netherlands Antilles), AO (Angola), AR (Argentina), AT (Austria), AU (Australia), AZ (Azerbaijan), BA (Bosnia-Herzegovina), BD (Bangladesh), BE (Belgium), BG (Bulgaria), BN (Brunei), BO (Bolivia), BR (Brazil), BY (Belarus), CA (Canada), CH (Switzerland), CL (Chile), CN (China), CO (Colombia), CV (Cape Verde), CZ (Czech Republic), DE (Germany), DK (Denmark), DZ (Algeria), EC (Ecuador), EG (Egypt), ES (Spain), ET (Ethiopia), FJ (Fiji), FR (France), GB (United Kingdom), GE (Georgia), GR (Greece), HK (Hong Kong), HN (Honduras), HR (Croatia), HU (Hungary), ID (Indonesia), IE (Ireland), IN (India), IR (Iran), IT (Italy), JO (Jordan), JP (Japan), KE (Kenya), KR (Korea, South), KW (Kuwait), KZ (Kazakhstan), LT (Lithuania), LV (Latvia), LY (Libya), MD (Moldova), ME (Montenegro), MK (Macedonia), MM (Myanmar), MU (Mauritius), MX (Mexico), MY (Malaysia), NG (Nigeria), NI (Nicaragua), NL (Netherlands), NO (Norway), NP (Nepal), NZ (New Zealand), PE (PerÂż), PH (Philippines), PK (Pakistan), PL (Poland), PS (Palestine), PT (Portugal), QA (Qatar), RS (Serbia), RO (Romania), RU (Russia), RW (Rwanda), SA (Saudi Arabia), SC (Seychelles), SE (Sweden), SG (Singapore), SI (Slovenia), SK (Slovakia), SV (El Salvador), SY (Syria), TH (Thailand), TM (Turkmenistan), TN (Tunisia), TR (Turkey), TT (Trinidad/Tobago), TW (Taiwan, ROC), UA (Ukraine), US (United States of America), UY (Uruguay), UZ (Uzbekistan), VE (Venezuela), VN (Viet Nam), YE (Yemen), ZA and (South Africa) | Terminated | June 18, 2012 | 671 |  | Software
Z-1831-2011 | 57887 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The enGenTM Laboratory Automation System consists of 3 main components including Instrument Management Subsystem (Instrument Manager), Automated Track Subsystem (TCAutomation) and Configuration File developed by Ortho-Clinical Diagnostics (OCD). The Instrument Manager is a Windows-based software product that operates on a Pentium-class (or higher) PC processor. The instrument management subsystem is capable of connectivity with clinical analyzers, a Laboratory Information Systems (LIS), barcode label printers and other third party devices. The TCAutomation is a pre-analytical and post-analytical sample management system. The configuration file enables operational support of the site specific layout and instrument selection. |  |  | January 27, 2011 | 2011 | Ortho-Clinical Diagnostics | A software anomaly that can occur with analyzers connected to the Bypass modules listed below on an enGen" Laboratory Automation System using TCAutomation" Software Version 3.1.1 and below. When this anomaly occurs, it is possible for a sample other than the intended sample to be aspirated, potentially leading to test results being assigned to the wrong patient without error notification. | Software design | The firm, Ortho Clinical Diagnostics (OCD), sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated January 27, 2011, by Federal Express overnight mail to all customers. Foreign affiliate consignees/customers were notified by email informing them of the issue on the same day. The letter described the problem, product and actions to be taken. The firm also included temporary procedures to help prevent the occurrence of the software anomaly. The customers were instructed to follow the enclosed Procedure to STOP the Automation System and/or Bypass Module; complete and return the attached Confirmation of Receipt - Important Response Required form by February 4, 2011, via fax to: Ortho Clinical Diagnostics, 1-888-557-3759 or 1-585-453-4110, and display this communication near the enGen Laboratory Automation System to ensure compliance by all operators until updated TCAutomation Laboratory Automation System software is installed. A Technical Bulletin will be sent to customers in the near future to be stored with user documentation. Ortho will implement a TCAutomation Laboratory Automation System software change in the near future to resolve this anomaly.If you have any questions regarding this notification, please call Customer Technical Services at 1-800-421-3311 (Options 1, 2). |  | 63 | 63 units | Worldwide distribution: USA and countries including: Canada, Chile, England, France, Germany, Italy, Mexico, Singapore and Spain. | Terminated | May 09, 2017 | 2294 | 63 | Software | Output/Calculation | Software update | Software Update
Z-1839-2011 | 58104 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | 10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue, Manufacturing Part Number 01287-09, Catalog Number D128709. |  |  | February 15, 2011 | 2011 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster has confirmed that there is an inaccessibility of electronic version of the instruction for use (e-IFU) for the "10 ft. Interface Cable, 10 Pin-Shielded Tip Pins to 12 Pin-Blue", catalog number D128709, through the J&J; Gateway (ie, e-IFU website). | Other | Biosense Webster, Inc. sent an Urgent Field Safety Notice letter dated March 16, 2011, to all affected customers (via express mail). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the enclosed Acknowledgement Form to confirm that all customers who received the product have been notified accordingly of this concern. If customers have any questions or require additional information, they are to contact their BWI representative and return the Acknowledgement form to:Biosense Webster, Inc.l5715 Arrow HighwayIrwindale, CA 91706Attn: Recall Coordinator - Virgil EllsworthPhone: (909) 839-8500Fax Number: (909) 839-7207.For questions regarding this recall call 909-839-7207. |  | 606 | 606 units total | Nationwide Distribution - including GA, KY, NV, WA, PA, and UT | Terminated | May 07, 2012 | 447 | 606 | Not Computer
Z-1833-2011 | 58025 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | REF 32-485050 VANGUARD KNEE INSTRUMENTATION MICROPLASTY SLIDEX ANTERIOR/POSTERIOR, QTY: 1 |  |  | February 10, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action due to the scale on the stylus of the Vanguard MP Slldex AP Sizer being incorrect for this assembly resulting in the scale being out of position. | Process change control | The firm, Biomet, sent a "FIELD SAFETY NOTICE" dated February 11, 2011 to all customers (heads of Orthopaedic Departments/Operating Departments /Sterile Services Departments/Procurement /Supplies/ Risk Management). The notice described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of the product; return to Biomet or their local Biomet distributor; ensure that the operating staff are made aware of this issue, and complete and return the FAX BACK RESPONSE FORM to Biomet UK Ltd or their local Biomet distributor via fax at +44 (0) 1656 645454. If you have any questions regarding this communication, please contact the Regulatory &amp; Compliance Manager at +44 (0) 1793 645317 or email: edward.spearpoint@biometeurope.com. |  | 17 | 17 | Foreign only; UK, Malta, The Netherlands, Italy, Turkey, Switzerland. | Terminated | May 16, 2011 | 95 | 17 | Not_Computer
Z-1835-2011 | 57857 | [6 empty] | Triathlon PKR Peg Drill, Catalog Number: 5650-4-512 (Size 1-2 Peg Drill); Catalog Number: 5650-4-536 (Size 3-6 Peg Drill); |  |  | February 09, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has identified a trend of complaints involving reports of the peg drill cold welding to the drill guide, causing the Peg Drill to cease functioning. If cold welding of these components occurs, the following potential patient and/or surgeon or surgical staff harms may results. | Component change control | Stryker sent Urgent Product Recall letters on February 9, 2011, and February 10, 2011, via Fed Ex. to all affected customers. The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069.For questions regarding this recall call 201-972-2100. |  | 13990 | 5650-4-512- 754 units; 5650-4-536- 880 units | Worldwide Distribution - USA , Sweden, Germany; France, Portugal, UK, Canada, Australia, and New Zealand, | Terminated | November 18, 2013 | 1013 | 13990 | Not_Computer
Z-1830-2011 | 57899 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Interventional Pain 6" 15 G (1.85 mm) Percutaneous Discectomy Probe, REF 407-266, Sterile, Stryker Instruments, Kalamazoo, MI. |  |  | February 03, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | The firm received five complaints that the stylet did not fit down the cannula. The investigation revealed that the incorrect probe assembly was packaged for two DeKompressor part numbers with the lot number of 10214012. | Other | Stryker Instruments Division of Stryker Corporation sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their inventory for any recalled product and quarantine any affected product. If the product was further distributed, customers were asked to forward the letter and the completed Reply Form to all affected locations. Customers were to be contacted to coordinate return of the affected product. Upon receipt of the product, a replacement would be issued to their account. For questions regarding this recall call 1-800-253-3210. |  | 44 | 44 units both products | Worldwide Distribution - USA including the states of PA, NY, KS, NJ, TX, MI, OK, and FL, and the countries of Italy and Switzerland. | Terminated | June 01, 2011 | 118 | 44 | Not_Computer
Z-1846-2011 | 55905 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | The VARIANT Sickle Cell Short Program ROM Card; Distributed by and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA |  |  | April 22, 2010 | 2010 | Bio-Rad Laboratories Inc | The presence of parameters for Version 40400-S on the labeled VARIANT Sickle Cell Short ROM Card instead of Version 80300-S. | Other | The firm, Bio-Rad, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated April 22, 2010, to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and quarantine the product; complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE FORM via fax to: (510) 741-3954, Attn: Bio-Rad CSD Regulatory Affairs, and contact there regional Bio-Rad office for ROM Card replacement. If you have any questions, contact Bio-Rad CSD Regulatory Affairs Department at 510-724-7000. |  | 10 | 10 units | Worldwide distribution: USA including: MD, NJ and UT; and countries of: Italy and Singapore. | Terminated | April 07, 2011 | 350 | 10 | Computer
Z-1852-2011 | 58028 | container, specimen, non-sterile | NNI | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | Large Port Specimen Sock, used to convert the pour/accessory port of a Cardinal Suction canister into a specimen retainer. The sock is used as an accessory in a suction collection system. Its placed in a suction canister to remove solids from the liquid aspirated during surgery. The solids may consist of patient tissue that are then sent to pathology or retained for other purposes. There is no direct contact between this device and the patient. |  |  | February 18, 2011 | 2011 | Cardinal Health Inc | In one lot of product (lot# 1091745), the label for the small version of the sock (catalog number 65652-123) was improperly used for the larger version of the product (65652-124) | Labeling mix-ups | On February 18, 2011 Cardinal Health sent a consignee letter with the following instruction: 1) Quarantine either Catalog Number with the inticated Lot Number.2) Contact the appropiate Customer Service group for return instructions:Hospital - 800-964-5227Federal Government - 800-444-1166Distributirs - 800-635-6021All other customers - 888-444-54403) Return the enclosed acknowledge form via fax to 847-689-9101, even if you do not have affected product, as Cardinal Health is required to confirm receipt of this notification from you, and to prevent further notices.4) Notify any customer to whom you may have distributed product affected by this recall. |  | 3360 | 3,360 each | Worldwide Distribution - USA including AL, AR, AZ. CA. CO,, CT, DC, DE, FL, GA, HI, IA, ID,IL,IM,KS, KY, LA, MA, MD, ME. MI MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, &amp; Puerto Rico - and the country of Canada. | Terminated | October 04, 2011 | 228 | 3360 | Not Computer
Z-1850-2011 | 58105 | No | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | Pelton &amp; Crane Spirit 3000 dental chair (Model #SP30) |  |  | March 04, 2011 | 2011 | Pelton &amp; Crane Company | The Link Arm Assembly that connects some rear mounted options to the Spirit 3000 or Spirit 1800 dental chairs may potentially break and fall off the dental chair. | Process control | Pelton &amp; Crane sent an "Urgent Medical Device Recall" letter dated 3/04/2011. They were instructed to: Locate the affected devices this recall is in regards to. Schedule service with the end user for Pelton &amp; Crane representative to replace the affected components this recall is in regards to. International accounts were instructed to order link arm assemblies from Pelton &amp; Crane and replace the affected components. All consignees were to complete the Recall Acknowledgement/Return Form and return the form to Pelton &amp; Crane. |  | 9 | 9 units | Distribution to USA, Canada, and the Netherlands. | Terminated | August 16, 2011 | 165 | 15 | Not_Computer
Z-1847-2011 | 58066 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix" Panel NMIC/ID-123, Catalog # 448723, packaged 25 panels/carton, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA (800) 638-8663*** |  |  | February 28, 2011 | 2011 | Becton Dickinson &amp; Co. | In vitro diagnostic test kit is defective and may cause false positive or false negative results. | Mixed-up of materials/components | BD Diagnostics Systems sent an Urgent Product Recall letter dated February 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinsue use of the affected product and discard any remaining packages. BD would issue replacements for the discarded material. Customers were instructed to complete the attached form whether or not they had any inventory remaining so that the firm may acknowledge receipt of this notification. Complete and fax the form to BD Regulatory Compliance to 410-316-4258. For further assistance contact BD Customer Service at 1-800-675-0908. For other inquiries, contact BD Technical Services at 1-800-638-8663. |  | 377 | 377 cartons | Worldwide Distribution - USA including AZ, CA, FL, GA, MA, MD, NC, NY, PA, TX, WA, WI, and WY and the countries of The product was distributed to medical facilities nationwide and to foreign consignees in Argentina, Brazil, Chile, and Mexico. | Terminated | January 30, 2013 | 702 | 377 | Not_Computer
Z-1817-2011 | 57004 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | 115V Blanketrol III Model 233 Hyper/Hypothermia System, Part number: 86107; and 115V CoolBlue Model 25-01 Hyper/Hypothermia System, Part number: 86000. |  |  | October 18, 2010 | 2010 | Cincinnati Sub-Zero Products Inc | The firm initiated this recall due to changes that have been made to the Blanketrol III Model 233 device's Operation and Operation/Technical Service Manuals, in order to stay in compliance with labeling regulations. | Other | On 10/18/2010 &amp; 11/5/2010 the firm sent URGENT Medial Device Field Action letters, dated 10/18/2010, to their customers. The letters identified the affected product and stated the reason for recall. Customers were instructed to immediately collect and discard all affected product Operation and Operation/Technical Service manuals currently in their possession. They were to replace all the discarded manuals with the revised versions that were provided with the letter. The enclosed response form should be completed and returned by fax to 513-772-9119, scanned and e-mailed to B3manual@cszinc.com. If the product had been further distributed, then those customers should also be identified and notified of the field action via a copy of the letter. Customers should direct questions to the field action coordinator, David Niehaus at 513-773-8810 ext. 3234. |  | 1385 | 763 - 115V Blanketrol III Model 233; &amp; 133 - 115V CoolBlue Model 25-01 | Worldwide Distribution -- US, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, &amp; WI and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, CHILE, CROATIA, ECUADOR, EL SALVADOR, GERMANY, HONG KONG, INDONESIA, ISRAEL, ITALY, KOREA, MALAYSIA, MEXICO D.F., NEW ZEALAND, PEOPLE'S REPUBLIC OF CHINA, PERU, REPUBLIC OF SIGNAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY &amp; VENEZUELA. | Terminated | March 08, 2013 | 872 | 1385 | Not_Computer
Z-1845-2011 | 57613 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CellQuest Pro SW, version 4.0.2. and later; |  |  | September 07, 2007 | 2007 | BD Biosciences, Systems &amp; Reagents | Results and gates could be mismatched if you utilize the undo function to undo a gate that was deleted. | Software design | Important Product Information letters, dated September 2007, were sent the week of 9/14/2001 via Federal Express to all affected customers. The letters identified the affected product along with the reason for recall. The letter states that "as long as customers do not use the Undo function to undo a gate deletion, the software can be used with full functionality." The letter also states that BD will change the issue in the next software version. If customers believe that they may have analyzed sample files by utilizing the Undo function, then it is recommended that the sample files be re-analyzed. If customers need more information they should contact the Customer Support Center at 877-232-8995. |  | 4552 | 4552 | Worldwide Distribution | Terminated | March 31, 2011 | 1301 | 4552 | Software | Device Operation | Software update | Software Update
Z-1849-2011 | 57602 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HEMOCUE AB Hb 201 DM ANALYZER, in vitro diagnostic device for quantitative hemoglobin determination. |  |  | October 07, 2010 | 2010 | Hemo Cue, Inc. | Incorrect configurations settings in software. As a result, certain error codes do not trigger appropriately. These error codes are: E01, E02 and E05 which are all related to the optronic performance of the analyzer. Possible incorrect results. | Process control | Hemocue AB sent Customer Notification "Notice of Medical Device Correction on HemoCue Glucose 201 DM and HemoCue Hb 201 DM Analyzers" dated October 7, 2010, to customers notifying them of incorrect settings on error codes. The corrective action consists of an update of the configuration settings on the the affected HemoCue Glucose 201 DM and Hb 201 DM analyzers that are in the field. The affected units in the field will be corrected in accordance with a rework procedure that has been established by HemoCue AB. |  | 79 | 79 | Worldwide distribution, including USA, and Glucose 201 DM analyzers to Norway and the Netherlands; Hb 201 DM analzyers to the Netherlands, France, Ireland, Great Britain, Poland and Germany. | Terminated | June 22, 2011 | 258 | 102 | Software | Alarm/Message | Configuration update | Other
Z-1853-2011 | 54844 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound Imaging System with software versions 4.0 and 5.0. |  |  | March 01, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | Image captured from the device may be erroneous. | Software design | Acuson Customer Safety Advisory notification letters were sent March 1, 2010 by certified mail. |  | 3539 | 3539 | Worldwide distribution to USA, Andorra, United Arab Emirates, Netherlands Antilles, Angola, Austria, Australia, Azerbaijan, Bosnia, Bangladesh, Belgium, Bahrain, Brunei Darussalam, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, Faroe Islands, France, Great Britain, Georgia, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iran, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Sri Lanka, Lithuania, Latvia, Morocco, Moldova, Mexico, Nicaragua, Netherlands, Norway, New Zealand, Oman, Peru, Philippines, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Suriname, Thailand, Turkey, Trinidad and Tobago, Taiwan, Ukraine, UA, Uruguay, Venezuela, Vietnam, Yemen, and South Africa. | Terminated | June 18, 2012 | 840 | 3539 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-1854-2011 | 57773 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HGBPROKIT which contains one meter and one container of 25 test strips. HGB TEST 100; HGB PRO TEST STRIPS -100, includes four containers of 25 test strips per container; International Technidyne Corporation |  |  | January 21, 2011 | 2011 | ITC-Nexus Dx, Inc. | ITC has determined that results obtained using certain lots of Hgb Pro Professional Hemoglobin Test Strips may differ from results obtained using a reference hemoglobin testing system. | Other | ITC-Nexus Dx, Inc. sent an Urgent Medical Device Recall letter with return response forms to end users and distributors dated January 21, 2011 via UPS overnight. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to determine if they had an of the affected lots.If they have full boxes or loose containers of the affected lots customers were instructed to stop using the product and remove it from inventory.Customers were requested to fill out the return response form and send it back to ITC via fax, email or mail.For questions customers were instructed to call ITC Technical Support at 800-631-5945 (US) or 732-548-5700 (International)), Extension 4707 or e-mail techsupport@itcmed.com. |  | 50563 | 2500 containers + 50,563 addtnl containers =53063 containers | Worldwide Distribution - USA (nationwide) and the countries of India, Oman, South Africa, and Singapore | Terminated | November 02, 2012 | 651 | 50563 | Not_Computer
Z-1859-2011 | 57987 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Stanley Bed-Check 90-Day LT Bed Sensormat, With 5.5 Ft. Cord, Model 74001, 5 mats per inner box, 5 inner boxes per overpack box (total 25 units). The firm name on the box is Stanley, 1-800-523-7956. The mat contains only the phone number. |  |  | November 25, 2009 | 2009 | Stanley Security Solutions, Inc. | The cord could come loose from the mat potentially causing a short to the internal wiring. | Process change control | The firm, Stanley Healthcare Solutions, issued an "Important Technical Bulletin" dated November 25, 2009 to all customers via e-mail. In addition, the firm issued a "Compliance Notice" dated August 30, 2010 via certified mail because the firm had not collected all of the units with the defect. The bulletin described the product, problem and actions to be taken. The customers were instructed to check the cords on the mats to make sure they weren't pulled out, if the cords were pulled out, they were to return the mats for replacement; contact Customer Service at 1-800-824-2996 ext: 3440 to obtain a return authorization number and take immediate action to return the units to the firms location in Lincoln, Nebraska. The Compliance Notice informed the customers who had not sent in their units to promptly read the notice and take the necessary steps to respond appropriately for their facility. Additionally, the customers were instructed to complete and return the Bed Check Sensormat Return Form and page 2 of the "Compliance Notice" via at 402-475-4281 or mail to: Compliance Engineer, Stanley Security Solutions, Inc., 1550 N. 20th Circle, Lincoln, NE 68503. Follow-up phone calls to nonresponders were made using a phone script. A "Second Notice Compliance Notice" dated October 20, 2010 was issued via certified mail to customers asking for immediate response to the previous communication. In addition, a third notice dated October 28, 2010 which was also flagged as "Second Notice" was again used asking for a response from the customers.If you have any questions regarding the notice or any action that may need to be taken by you, please contact Quality Manager at 1-800-824-2996 extension 3419 or Compliance Officers at 402-742-9414 or 402-742-9425. |  | 3500 | 3,500 mats | Worldwide distribution: USA and the country of Canada. | Terminated | March 31, 2011 | 491 | 209183 | Not_Computer
Z-1862-2011 | 57852 | No | HTJ | Depth gauge for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Xcelerate Specialty Patella Milling System Per File K2399; Catalog Number I-K2399PR00, I-K2399PR10; Stryker Howmedica Osteonics, Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430 |  |  | January 27, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that the 10 mm on the bone remaining depth gage is incorrect thus allowing for the reamer to leave as little as 8.9mm (in a nominal condition) patella bone. | Device Design | Stryker sent an Urgent Product Recall letter dated January 27, 2011, via FedEX for January 28, 2011, delivery with return receipt. The letter identified the product, the problem, and the action to be taken by the customer. Customers were insructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069. Customers were also instructed to contact the hospitals in their territory that have the affected product to arrange return of the product. Product should be returned to the attention of: Regulatory ComplianceStryker Orthopaedics325 Corporate DriveMahwah, New Jersey, 07430. The orange Product Remediation sticker should be attached to the return, indicating the Product Remediation # RA 2009-296. Mark the outer box with words "Product Recall."For questions regarding this recall call (201) 972-2100. |  | 52 | 52 units | Nationwide Distribution | Terminated | October 15, 2012 | 627 | 52 | Not_Computer
Z-1866-2011 | 58136 | No | MGP |  |  | General &amp; Plastic Surgery | 510(k) | HydroHydrofilm is for use as a post-operative and primary dressing to protect against secondary infection and to cover non-exudating wounds and for use as a secondary dressing for fixation of catheters or cannulae.film, 10cm x 12.5 cm, Paul Hartmann AG, 89522 Heidenhiem, Germany |  |  | February 11, 2011 | 2011 | Hartmann USA, Inc | Misbranded; the expiry date on the primary packing incorrectly indicates 2010-10, instead of 2015-10. Consignees were notified by letter on/about 02/11/2011. They were advised to examine their inventory and not to use any of the affected product and to return it to Paul Hartmann. A response form was included to be completed and returned to Paul Hartmann. They were further instructed to pass the no | Other | Hartmann USA, Inc. sent letter to the Consignees on 2/11/11. They were advised to examine their inventory and not to use any of the affected product and to return it to Hartmann USA, Inc. A response form was included to be completed and returned to Hartmann USA, Inc. They were further instructed to pass the notice on to their consignees if distributed. For questions, call (803) 325-7600. |  | 1250 | 100 boxes (10 x 12.5 cm per box) | Nationwide -- Including CA, GA, IL, MO, MS, TX, and VA. | Terminated | August 12, 2011 | 182 | 1250 | Not_Computer
Z-1875-2011 | 58051 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | CVS Pharmacy Plastic Sesame Street Adhesive Bandages; a sterile, latex free, assorted sizes with multiple designs adhesive bandages; 20 bandages per box, 24 boxes per case; Distributed by: CVS Pharmacy, Inc., One CVS Drive, Woonsocket, RI 02895, Made in Egypt; REF CVS 826159, UPC 0 50428 95118 7 |  |  | February 21, 2011 | 2011 | Pharmaplast S.A.E. | The bandages are labeled as sterile, but failed sterility testing. | Process control | Pharmaplast S.A.E. called the direct account on 2/17/11 and sent a Urgent Medical device recall letter dated 2/21/11 to the direct account via overnight mail, requesting that they examine their inventory for the non-sterile adhesive bandages and quarantine them for return to Pharmaplast via sea freight. The account was instructed to contact their customers concerning the recall and return their stocks of the affected lot, which was done by telephone calls on 2/17/11.For questions regarding this Recall please call 00203-4500264 |  | 5336 | 5,336 cases | Nationwide Distribution - Including Illinois. | Terminated | August 25, 2011 | 185 | 5336 | Not_Computer
Z-1865-2011 | 57890 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Axxent Flexishield Mini |  |  | February 03, 2011 | 2011 | Xoft Inc. | Xoft Axxent Flexishield Mini product may may shed particles identified as tungsten, and may look like suspicious calcifications on follow up scans if the product was used during Intraoperative radiation therapy. | Device Design | Xoft, Inc. sent an Urgent Field Removal Action letter dated February 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and return all units and lots of Flexishield Mini Catalog Number: F5300 in their inventory. Customers were instructed to fill out the attached form and fax to 866-222-3404, then call Customers Service toll free directly at 877-963-8327 to make arrangements to return the product and receive a credit if applicable. Customers were asked to complete and return the form even if they did not have any units to return.For questions regarding this recall call 408-419-2300. |  | 165 | 165 units | Nationwide Distribution including AL, CA, CO, FL, IL, MI, MO, MS, NJ, NY, OK, PA, TN, and WI. | Terminated | June 08, 2011 | 125 | 165 | Not_Computer
Z-1863-2011 | 57742 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cardiac Marker Control Level 1, List Number 136601-06F12-02; Cardiac Marker Control Level 2, List number 136602/06F13-02; and Cardiac Marker Control Level 3, List number 136603/06F14-02. |  |  | December 17, 2010 | 2010 | Abbott Point Of Care Inc. | Presence of CO2 in the head space of the vial can result in the generation of high results in Cardiac Marker Control Level 1, 2, and 3 which is used along with the Cardiac Marker Control Calibration Verification Set. | Device Design | Urgent Recall Notice letters with business reply cards were sent Fed ex on December 17, 2010 to all accounts. International notifications were sent on December 21, 2010. The letter identified the affected product and lot numbers along with a background explanation of the issue. The firm recommends that customers remove the cap from the vial; squeeze the vial 10 times, prior to thawing; replace cap on the vial; and then follow the remaining steps for use as outlines in section 14-12 of the i-STAT 1 System Manual. Customers are to complete and return the Business Reply card to acknowledge receipt of the letter. If the affected product had been further distributed then a copy of the letter should be provided to those customers as well. If additional information is needed, customers should contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or their Abbott Point of Care representative. |  | 2360 | 2360 units total | Worldwide Distribution | Terminated | October 16, 2012 | 669 | 2360 | Not_Computer
Z-1620-2011 | 51208 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD FACS Sample Prep Assistant (SPA and SPA II) Instrument, |  |  | January 03, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | To prevent a needle stick while replacing the probe, the firm has initiated a Field Correction, and issued an improved Users Guide procedure (User Bulletin, Part Number 645050). | Nonconforming Material/Component | The firm, BD Biosciences, sent an information letter dated December 2008 with an enclosed New Probe Replacement Service/User Bulletin, to the customers. The letter described the product, problem and action to be taken. The customers were instructed to review the attached User Bulletin (Part Number 645050); ensure that their laboratory follows the updated procedure for removing and replacing probe; place the User Bulletin in their BD FACS SPA or SPA II User's Guide; always follow the instructions in User Bulletin to properly remove and replace probe, and complete and return the Effectivity Form via fax to 408-577-1409. Should you have any comments or questions, please contact BD Biosciences at 1-877-232-8995 in the US. Outside the US, contact your regional BD office or field service engineer for assistance. |  | 70 | 70 | Worldwide distribution: USA and countries including: Hong Kong, Australia, Belgium, Canada, El Salvador, Hong Kong, Japan, and Mexico. | Terminated | March 10, 2011 | 796 | 70 | Not_Computer
Z-1650-2011 | 57404 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BIO-RAD BioPlex 2200 Syphilis IgG Control Set. Lot # 46636 |  |  | November 23, 2010 | 2010 | Bio-Rad Laboratories | An upward shift in the control values can result in a control value that falls outside of the current manufacturing ranges. | Other | The firm, Bio-Rad, sent a "Voluntary Field Correction: QC Ranges for Syphilis IgG Control Set Lot#46636" letter to Domestic customers via FedEx on November 24, 2010 and to the International customers via e-mail on November 23, 2010. The letter described the product, problem and actions to be taken by customers. The Customers were advised to temporarily establish their own means and acceptable ranges for Control Lot 46636 for the r15 and r47 markers. The firm will provide adjusted recommended means and acceptable ranges. Additionally, the customers were to complete and return the Customer Response Form for USA and/or the Subsidiary Response Form (Outside US) via fax to (510) 741-4846, Attn: Bio-Rad BPD-Product Support.If you have any questions about this notification, contact your local Bio-Rad Technical Support. |  | 202 | 202 | Worldwide distribution: USA and countries including: Canada, Czech Republic, France, Germany, and Italy. | Terminated | August 04, 2011 | 254 | 202 | Not_Computer
Z-1519-2011 | 57107 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pacific Hemostasis Âż Thromboplastin-DS, catalog #100354, packaged in 4 mL lypholyzed vials, 10 bottles/carton. Vials are labeled, in part ***Fisher Diagnostics a division of Fisher Scientific Company, LLC Middletown, VA 22645-1905 USA*** |  |  | October 06, 2010 | 2010 | Fisher Diagnostics, A Company of Fisher Scientific LLC | An in vitro diagnostic reagent used in coagulation testing may cause inaccurate clotting times in quality control and patient plasma samples. | Nonconforming Material/Component | The firm, ThermoFisher Scientific, sent two "Urgent Medical Device Removal" letters one via Federal Express on October 20, 2010 and certified mail for domestic customers on October 21, 2010. The letters described the product, problem and action to be taken by the customers. The customers were instructed to indicate the quantities of all affected product currently in stock at their facility; complete and return the attached Response Form/table to Fisher Diagnostics via fax to (540) 869-8129 or by mail to the attention of Technical Support; dispose of any remaining inventory of the affected lot numbers; provide a copy of this letter to their customers, and file this letter for future reference.If you have any questions regarding this notification, please contact our Technical Support staff at 800-528-0494, Option 2, or email techsupport.diagnostics.mtn@themofisher.com. |  | 1190 | 1190 kits | Worldwide distribution: USA including states of: Alaska, Florida, Kentucky, Ohio, New Jersey, New York, and Texas; and countries including: Argentina, Dubai, Italy, Mexico, Philippines, Romania, and Turkey. | Terminated | November 02, 2011 | 392 | 1190 | Not_Computer
Z-1740-2011 | 57705 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | UC Tibial Insert, Size 4, 13mm, REF 163-3413, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | January 07, 2011 | 2011 | Ortho Development Corporation | Components of a replacement knee system could be mislabeled. | Error in labeling | The firm, Ortho Development, sent a "Field Notification Memo" date January 7, 2011 to its customers. The memo described the product, problem and action to be taken. The customers were instructed to promptly return any remaining parts of the listed product; locate all the parts in their inventory; identify any of the listed parts that may have been used in surgery; verify the quantity they have on hand and the quantity used in surgery and reconcile them with the quantity sent; if discrepancy exists, notify customer service resolution; prepare items for return and contact customer service to obtain RMA number. Once reconciled, arrangements would be made for sending new inventory and the return of unused parts.If you have any questions regarding this issue, please notify Regulatory Affairs Manager at (801) 619-3450. |  | 25 | 25 units | Worldwide distribution: USA including states of: AZ, CA, FL, MI, TX, and UT; and country of: Japan. | Terminated | December 27, 2011 | 354 | 72 | Not_Computer
Z-1728-2011 | 57814 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hc2 System Software Suite 4.0 version 3.0, available as component in Qiagen user guide, catalog #5050-00162, and as a component within a Qiagen PC, available as catalog # 5050-11140A. |  |  | January 12, 2011 | 2011 | QIAGEN Gaithersburg, Inc. | The software component of a diagnostic medical device may be defective and could cause incorrect patient values. | Software design | The firm, Qiagen, sent a "Notice of Product Field Action" letter dated January 12, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue the process of running multiple assay protocols on a single microplate to eliminate the potential for inaccurate patient result reporting, and to complete and return the acknowledgment section via fax to QIAGEN 661-702-3854 or scan the document and email to techservice-na@qiagen.com. The firm informed users that they will provide further detailed instructions for a workaround and software fix. A QIAGEN representative will contact the customers who indicated that they run multiple assays on a plate and are currently using the HC2 System Software Version 3.0 (Suite 4.0), to assist in the identification of those patients that may have been affected by the software defect.If you have any questions or concerns, please contact Qiagen Technical Services at 800-344-3631, Option 2. |  | 163 | 163 | The recalled product was sold to medical laboratories nationwide and in Canada. | Terminated | November 14, 2011 | 306 | 163 | Computer
Z-1729-2011 | 57631 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) Creatinine Flex(R) Reagent Cartridges |  |  | December 27, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed the potential for falsely elevated creatinine results with specified CREA Flex(R) reagent cartridge lots when processing tests from a well that has been punctured. This issue may affect creatinine results for quality control (QC) and patient samples. | Material/Component Contamination | The firm, SIEMENS, sent an "Urgent Field Safety Notice" dated December 2010 to all customers. The letter described the product ( Dimension(R) CREA Flex(R) Reagent Cartridges DF33A lots BB1286, CA1260 and FB1251), problem and actions to be taken. The customers were instructed to please discontinue use and discard any remaining inventory of the product; complete and return the fax form to the Customer Service Support for replacement of the product; if they must continue use of the product until they receive a replacement, open a new well set every 8 hours to minimize any impact to CREA test results, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the Technical Solution Center at 302-631-8467.If you have any questions, contact the Technical Solution Center at 800-441-9250. |  | 20177 | 20,177 | Worldwide distribution: United States and the following foreign countries: Austria, Belgium, Bosnia and Herzegovina, Canary Islands, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Great Britain, Greece, Hungary, Italy, Latvia, Lithuania, Kuwait, Macedonia, Morocco, Norway, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey and the United Arab Emirates. | Terminated | December 21, 2013 | 1090 | 20177 | Not_Computer
Z-1877-2011 | 57445 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5, Part Number: 6607073. |  |  | September 16, 2010 | 2010 | Beckman Coulter Inc. | A labeling error resulted in the specimen age stability and the prepared sample stability being stated incorrectly in the instructions for use. There is no impact to the product safety or functionality. | Pending | An Urgent: Product Corrective Action letter dated, September 27, 2010, was sent to customers. The letter identified the affected product and stated the reasons for the recall. Customers were also asked to take the following actions: ACTION/RESOLUTIONFor issue 1: The specimen stability claims should be as follows:- 48 hours for both % positives and absolute counts for the CD45/4/8/3 tube- 24 hours for both % positives and absolute counts for the CD45/56/19/3 tubeFor Issue 2: Results produced are not impacted when samples are prepared and run within the post collection time frames provided in the product instructions for use:- Preparation within 72 hours of collection for the CD45/4/8/3 tube- Preparation within 24 hours of collection for CD45/56/19/3 tubeBeckman Coulter requested consignees complete and return the enclosed response form within 10 days so they may maintain their records. The firm also asked that they share the information with laboratory staff, and retain this notification as part of their Quality System documentation.Questions regarding the Product Corrective Action were directed to (800) 526-7694 in the United States or (800) 463-7828 in Canada. If outside the US or Canada, consignees were instructed to contact their local Beckman Coulter Representative. |  | 26822 | 26822 units | Worldwide Distribution -- United States, Albania, Australia, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Bulgaria, Burundi, Cameroon, Chile, China, Colombia, Croatia, Czech Republic, Denmark, El Salvador, Finland, France, French Guiana, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Republic of Kuwait, Libyan Arab Jamahiriya, Macao, Malaysia, Mexico, Monaco, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan Panama, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | June 14, 2012 | 637 | 105810 | Not_Computer
Z-1883-2011 | 55554 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems enhanced AccuTnl Reagent Kit, Part Number: A78803 |  |  | December 10, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that under certain circumstances, the following assays may produce elevated values at the low end when run in conjunction with the enhanced Access Accu Tnl assay (P/N A 78033):(1) Access Total BhCG (P/N 33500);(2) Access HYPERsensitive hTSH/Fast hTSH (P/N 33820);(3) Access hFSH (P/N 33520).The issue can occur only when the assays are | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated December 16, 2009 to all customers. The letter described the product, problem and actions to be taken. Beckman Coulter instructed the customers to do the following mitigation to reduce the risk of encountering an erroneous result:Systems with multiple reagent pipettors (UniCel Dxl 800 and 600; UniCel DxC 880i, 860i, 680i, and 660i):1. Designate enhanced Access AccuTnl to a different reagent pipettor (s) than Access Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH.2. Discard any open Access Total BhCG, hTSH/Fast hTSH and hFSH reagent packs3. Calibrate Access Total BhCG, hTSH and hFSH using fresh reagent packs4. Run QC following acceptable calibration curves. Proceed with operation following acceptable QC performanceSystems without multiple reagent pipettors (Access, Access 2; SYNCHRON LXi 725, UniCel DxC 600i):1. Discard any open Access Total BhCG, hTSH/Fast hTSH and hFSH reagent packs.2. Calibrate Access Total BhCG, hTSH/Fast hTSH and hFSH using fresh reagent packs3. Run QC following acceptable calibration curves.Ensure that controls and patient samples for Access Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH are not run immediately following enhanced Access AccuTnl controls or patient samples.For example, during routine QC runs, ensure that these assays are requested prior to enhanced Access AccuTnl.4. Proceed with operation following acceptable QC performance.NOTE: An enhanced AccuTnl sample loaded with a system priority of STAT will interrupt routine Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH sample aspiration. Avoid loading enhanced AccuTnl sample aspiration. Avoid loading enhanced AccuTnl samples with a system priority of STAT during aspiration of Total BhCG/Dil-hCG2, hTSH/Fast hTSH and hFSH.Diagnostic test results should be interpreted in light of the clinical presentation of the patient, including: symptoms, clinical history, data from other |  | 22032 | 22,032 kits | Worldwide distribution: USA and country of Canada. | Terminated | May 07, 2012 | 879 | 22032 | Not_Computer
Z-1884-2011 | 52820 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | TEE Transducer V5M, when used with the Acuson S2000 Ultrasound System, software versions: VA15, VA15A or VA15B; Model Numbers: 10041461 (Acuson S2000) and 8263703 (V5M Transducer); |  |  | April 07, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Device function failure--When performing a TEE exam with the transducer, CW (Continuous Wave) mode is not functional. Due to work flow, this may not be apparent until the transducer is inserted into the patient's esophagus. | Software design | Siemens sent a Customer Safety Advisory Notification letter on April 7, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to pass the notice on to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-969-9112. |  | 61 | 61 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Great Britain, Hong Kong, Indonesia, India, Italy, South Korea, Norway, Turkmenistan, and Taiwan | Terminated | April 05, 2011 | 728 | 61 | Other |  |  | Other
Z-1879-2011 | 57810 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King System , King LTS-D Kit Size 5, Sterile, REF KLTSD405, Manufacturer: King Systems Noblesville, IN. |  |  | January 24, 2011 | 2011 | King Systems Corp. | The firms contract manufacturer reported that the products were below the specification for hardness. This could contribute to a higher incidence of problems during the insertion and final positioning of the airway in the patient. In rare instances the softness of the tube could cause the tube to fold over. | Other | King Systems Corp. sent an URGENT DEVICE RECALL NOTICE letter dated January 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. The letters instructed users to cease all sales of the recalled product, to notify the consignees for any product that was further distributed, and to return all recalled product to King Systems. The notice did include a response form.For questions call 317-776-6823, ext 257. |  | 1430 | 143 cases of 10 each case | Worldwide Distribution - USA and Canada | Terminated | June 29, 2012 | 522 | 2120 | Not_Computer
Z-1882-2011 | 58044 | No | LLC |  |  | General Hospital | 510(k) | iSite PACS with iSite ViewForum Applications, R7.4V6 L1 L2, L3, |  |  | February 23, 2011 | 2011 | Philips Medical Systems | A software anomaly that occurs when clinicians use Volume Vision to view images in 3D and uses the measurement tool to measure regions of interest. This could lead to an incorrect interpretation of the measurements which could cause an incorrect diagnosis and incorrect treatment plan. | Software design | Philips sent an Urgent Field Safety Notice letter dated February 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to assure that all colleagues were aware of the problem.When using the Volume Vision software, check the correctness of measurements (graphics and numerical values) prior to saving a captured image containing measurements.Save captured images containing measurements immediately after creating them and before altering the measurements in the original image view.For any subsequent measurements, create a new capture and save it immediately.For further information or support concerning this issues, customers should contact their local Philips representative at +1-877-328-2808. |  |  |  | Worldwide Distribution. | Terminated | August 18, 2011 | 176 |  | Software
Z-1880-2011 | 58013 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | POUCH/CARTON LABEL: |  |  | November 30, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | On Nov 30, 2010 ConMed Linvatec, Largo, FL recalled their Micro Sagittal Blade, Fine Lot number 180209, Product Number 5023-138. Some units of Micro Sagittal Blade, Fine from Lot 180209 may have compromised seals in the sterile packaging. | Process control | Linvatec Corp. dba ConMed sent an Urgent Medical Device Recall Notification letter dated November 30, 2010, by FedEx, email, Fax, etc) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the reply form within 30 days of receipt. Immediately check their facility's inventory for the affected product and lot number. If they had any of the product with lot number 180209. Do Not Use. Segregate and return for credit. If the product was transferred to another facility, customers were to complete the REPLY FORM documenting the transfer of the product and fax the form to 727-319-5701.For questions regarding this recall call 727-319-5701 or 800-535-8536. |  | 415 | 415 | Worldwide Distribution - USA including CT, FL, GA, HI, IN, MO, MS, MT, NJ, NY, OH, PA, SC, and TX and the countries of Canada, and South Africa. | Terminated | November 09, 2011 | 344 | 415 | Not_Computer
Z-1896-2011 | 58011 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Rusch Irrigation Tray w/Piston Catalogue Number 67890 |  |  | February 24, 2011 | 2011 | Teleflex Medical | The irrigation trays contain alcohol prep pads which are included in the scope of a recall initiated by Triad Group, due to a potential bacterial contamination. | Other | the firm,Teleflex Medical, sent two "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letters one dated February 24, 2011 and the other dated March 30, 2011 to all consignees and customers. The letters described the product, problem and actions to be taken. The consignees were instructed to communicate the recall to any of their customers who have received product included within the scope of the field corrective action via the customer letter. The customers were informed in the letter of concerns about potential bacterial contamination of the alcohol prep pads in the affected trays. If contaminated alcohol prep pads are used, there is a potential for infection to occur. The customers were instructed "that upon opening the affected trays for use, they should discard the prep pads as medical waste". The customers were also instructed to immediately discontinue use and quarantine any products with the product numbers and lot numbers listed; return the product, and complete and return the Recall Acknowledgement Form, even if they no affected product, via fax to 1-866-804-988, Attn: Customer Service (a customer service representative contact the customer with a Return Goods Authorization number and instructions for the return of the product to Teleflex Medical). If you have any questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 95647 | 95,647 eaches (total for both Catalogue Numbers) | Worldwide distribution: USA including states of: AL, AZ, CA, FL, GA, IL, MS, MI, MO, MS, NC, NJ, NH, NY, OH, PA, TX, UT, VA, Puerto Rica and the country of: Canada. | Terminated | April 17, 2012 | 418 | 95648 | Not_Computer
Z-1893-2011 | 58009 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF VP6528***Qty 1***Articulating Arm, Secondary Flat Panel***RX Only***See Instructions for Use***CONMED LINVATEC, CONMED LINVATEC, 11311 CONCEPT BLVD, LARGO, FL 33773-4908 USA*** |  |  | September 17, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | If the VHM Arm is left in the locked position, the gas spring can lose pressure and releasing the lock can result in a sudden downward movement. | Device Design | An Urgent Service Bulletin Notification was mailed to each ConMed Linvatec customer who received the affected product. The letter identified the affected product and stated the reason for notification. The letter asks customers to review the GCX Corporation Service Bulletins and updated Installation/User Manual. ConMed Linvatec recommends that customers periodically check the VHM Arm to see that it is functioning properly. In addition "Sudden Movement" stickers, provided by GCX, were also included to all consignees as part of the mailing. If customers in the US have questions or need technical support, they can contact ConMed Linvatec Endoscopy Customer Service at 888-292-0100 or e-mail Customerserv1@linvatec.com. |  | 373 | 373 | Worldwide Distribution -- Australia, Canada, Morocco, Saudi Arabia, Taiwan, Province of China, and throughout the U.S. | Terminated | May 17, 2011 | 242 | 373 | Not_Computer
Z-1894-2011 | 58109 | [6 empty] | BD Beaver XStar Safety 2.5mm Crescent Knife; |  |  | February 28, 2011 | 2011 | Beaver-Visitec International Inc. | Loose blades in the handle of knives due to uncured epoxy. | Nonconforming Material/Component | Beaver Visitec International, Inc.sent an Urgent Recall Notification letter dated February 25, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all unused product and to complete a Customer Response Form (to be returned via fax to 866-906-4304.For questions regarding this recall call 781-906-7917. |  | 1680 | 1680 units (1130 domestic, 550 foreign) | Worldwide Distribution - USA including AZ, MA, MI, MO, NC, NJ, NY, OH, OR, PA, RI, TN, TX, VA and the country of China | Terminated | June 14, 2012 | 472 | 1680 | Not_Computer
Z-1886-2011 | 58107 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Harmony LA/LED Monitor Arm Adapter. Model Number: LB61 |  |  | February 18, 2011 | 2011 | Steris Corporation | The arm adapter does not contain labelling specifying the applicable weight and moment limitations of the arm adapter | Other | Steris Corporation sent an Urgent Field Correction Notice letter dated February 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Steris representative would contact customers to arrange for replacement of the product. For questions regarding the firms visit to their facility, customers were instructed to contact Steris Field Service Dispatch at 1-800-333-8828.For any questions regarding this matter call 440-392-7705. |  | 14 | 14 Units | Worldwide Distribution - USA including CO, IL, IN, MO, NC, NV, NY and OH and the countrries of Canada, China, and Australia. | Terminated | March 13, 2013 | 754 | 14 | Not_Computer
Z-1868-2011 | 57939 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Curlin 6000 IOD Ambulatory Infusion Pump |  |  | February 28, 2011 | 2011 | Zevex International, Inc. | Infusion pumps may exhibit false error codes which render the device inoperable until a recovery process is performed. | Software design | The firm, MOOG, sent an "URGENT DEVICE RECALL NOTIFICATION" letter dated February 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The Customers were instructed to do the following:1) Identify the affected Pumps:a) If they have an affected pump, fill out and return the enclosed Business Reply Card within five (5) working days.b) If their pump is not affected by this recall, the pump can be used for its intended purposes.2) Removal of affected Product:a) remove affected products from service and use non-affected pumpb) If you do not have an alternate pump available for use, MOOG recommended that they weigh the risks and benefits for delivering medications using the affected pump versus the risk of not using any pump for each patientWARNING: Affected pumps should not be used to deliver critical medications where a delay in therapy could result in serious adverse health effects, injury or death; and affected pumps should not be used with the neonatal population. 3)Should they Have an Affected pump which exhibits Error Code 45 or is used in therapeutic situations specified in the warning, please contact MOOG Customer Service at (800) 970-2337, Monday thru Friday, 7:00am to 5:00pm MST to arrange for a loaner pump and the return of your affected pump.If you have any additional questions regarding this recall, please contact MOOG's Customers Service Department at (800) 970-2337. |  | 3407 | 3,407 units | Worldwide distribution: USA and countries including: Australia, Canada, Costa Rica, Croatia, Dubai, France, Germany, Ireland, Italy, Japan, Jordan, Kuwait, Malaysia, New Zealand, Qatar, Saudi Arabia, Singapore, South Africa, Taiwan, Turkey, UAE, United Kingdom. | Terminated | August 06, 2012 | 525 | 18306 | Software
Z-1900-2011 | 51497 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Siemens brand ACUSON S2000 Ultrasound System with software versions: |  |  | January 14, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | The product has a software problem in which previous patient measurement data gets associated with another patient's image. | Software change control | The firm, SIEMENS, issued a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that the first image in Exam Review is from the current patient; ensure that the Heart Cycle measurement tool is not selected on the Cale menu when entering Exam Review; be aware of this issue as described in the letter and perform four (4) or less measurements of the same measurement label. Should you have any questions, please call (650) 694-5398. |  |  |  | Worldwide distribution. | Terminated | April 06, 2011 | 812 |  | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1828-2011 | 52193 | No | RCE |  |  |  |  | Phoenix x/ray nanotom series. |  |  | March 16, 2009 | 2009 | GE Inspection Technologies, LP | Some of the systems do not have a safety interlock on the service access panel. Although the systems have two independent means for indicating when x-rays are being generated, some systems may have one indicator that is not discernible. Also, some systems lack proper placement of a caution label and others do not have this label affixed. | Radiation Control for Health and Safety Act | The firm has completed all of the following necessary actions and no further actions have been indicated:A Field Service Memo was issued to customers. It described the affected product and described the issue. GE Sensing and Inspection provided and installed retrofit kits, and the systems in question were tested by qualified Field Service Engineers at no charge to customers. GE asked customers to post the notification letter on the affected equipment until retrofit was completed. Questions were to be directed to the GE Service Center at 800-766-6062. |  | 118 | 118 units total | Nationwide Distribution | Terminated | April 17, 2017 | 2954 | 118 | Not_Computer
Z-1898-2011 | 54684 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Device brochure, labeled in part : BIO -RAD***L70242709 VARIANT TM II TURBO***Hemoglobin A 1C***Cartridge Set, Cartridge Set, Analytical Cartridge, Guard Cartridge Insert***UNITED STATES, Bio-Rad Laboratories, Inc. Hercules, CA 94547***" |  |  | December 08, 2009 | 2009 | Bio-Rad Laboratories Inc | Wrong lot number was found on the Hemoglobin Cartridge Insert . The insert did not match the lot number to the device cartridge or the device cartridge label or on the Update Kit CD. | Other | The firm, Bio-Rad, sent a "MEDICAL DEVICE CORRECTION" letter dated November 29, 2009 to all customers by Federal Express. The letter described the product, problem and action to be taken. The customers were instructed to do the following:1) If their Lot No. LV90170B VARIANT II TURBO Hemoglobin A(10) Cartridge Set includes an insert referencing Lot No. 80080, please destroy the insert and use the new insert enclosed with the letter.2) If their Lot No. LV90170B VARIANT II TURBO Hemoglobin A(10) Cartridge Set includes an insert referencing Lot No. 90170, they may continue using the insert.3) Complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to: Bio-Rad CSD Regulatory Affairs Department at (510) 741-3954.If you have any questions, please call 510-724-7000 or contact your regional Bio-Rad office. |  | 51 | 51 kits (Lot # 70292242; 45 kits( Lot # 70292318) | Worldwide distribution: USA including states of: CA, CT, GA, IL, IN, MA, ME, MS, NH, NV, OH, PA, TX, VA and WV; and countries including: France, Hong Kong, Italy, Korea and New Zealand. | Terminated | June 29, 2011 | 568 | 51 | Not_Computer
Z-1891-2011 | 57972 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal F2 Anesthesia Breathing Circuit, REF 5D440-61X-17, Manufacturer: King Systems, Nobleville, IN. |  |  | January 24, 2011 | 2011 | King Systems Corp. | The firm is conducting a subrecall of facetent masks because the masks could have a sticky substance on the inside or outside of the mask that may manifest itself 6-8 months after the manufacture date of the product. The sticky substance could result in inconvenience or in some remote cases cause patient harm. | Other | The firm issued an URGENT PRODUCT RECALL letter dated 1/24/2011. The letter identified the affected product and stated the reason for recall. The letter instructed consignees to cease all sales, to immediately examine their inventory, and to quarantine the product. The letter included a response form that is to be completed and returned. Distributors were instructed to inform their customers of the recall. King Systems Customer Service Department should be contacted at 1-800-642-5464 in order to receive a Return Authorization Number. The letter stated that consignees should return the recalled products to King Systems and the firm will replace the products or provide an acceptable substitute. Customers should contact Customer Service if they have any questions regarding the recall and the letter. |  | 12480 | 416 cases of 30 per case 12480 units | Nationwide Distribution -- NM, GA, TX, FL, IL, VA and LA. | Terminated | August 14, 2012 | 568 | 82905 | Not_Computer
Z-1899-2011 | 57855 | No | KSD | Heat-sealing device. | Hematology | Hematology | 510(K) Exempt | 2380 MINI Hand-Held Tube Sealer Power Source |  |  | January 31, 2011 | 2011 | Haemonetics Corporation | Battery pack may have exposed wire becoming hot | Equipment maintenance | The firm, Haemonetics, issued a "Field Notification: 2380 Sealer battery pack" letter dated January 28, 2011 to its customers. An international English version of this letter (CL101020-IE) has been released to foreign Distribution countries. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Record the manufacturing date of battery, equipment model, serial number, and photo of the affected battery.2) Report the defect through the Haemonetics PIR system by emailing: sebra2380@haemonetics.com with all of the information listed including the appropriate contact information for their company.3) Recycle or dispose of the nickel-cadmium (NiCad) battery pack in accordance with the local, state, federal, or country specific regulations. Under no circumstances should the battery pack be incinerated as this can cause explosion and personal injury. The operation manual will be updated to provide additional information about the proper installation and removal of the battery. This updated manual will be distributed with new devices upon release.Should you have any questions or concerns regarding this matter, please contact the firm at sebra2380@haemonetics.com. |  | 50031 | 50,031 units | Worldwide distribution: USA and countries including: Australia, Austria, AEU, Belgium, Bangladesh, Canada, China, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Korea, Malaysia, Mexico, New Zealand, Pakistan, Philippines, Singapore, Sri Lanka, Taiwan and Turkey. | Terminated | March 06, 2012 | 400 | 50031 | Not_Computer
Z-1903-2011 | 58241 | No | GAB | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | REF 906764, CURVTEK EYE NEEDLE, 12 MM LARGE, STAINLESS STEEL GAMMA STERILIZED, PKG/3, BIOMET SPORTS MEDICINE, 56 EAST BELL DRIVE, P.O. BOX 587, WARSAW, IN 46581 USA, STERILE |  |  | March 10, 2011 | 2011 | Biomet, Inc. | The firm initiated a recall due to the discovery that the devices were not sterilized, as required, prior to distribution. | Process control | The firm, BIOMET, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated March 10, 2011 to all customers. The letter described the product, problem and actions to be taken. The letter noted that -Use of a non-sterile Curvtek Eye Needle may lead to patient infection, possibly requiring a revision procedure. The customers were instructed to immediately locate, discontinue use of the product , remove the product from circulation and return the product to Biomet; carefully follow the instructions on the enclosed FAX BACK RESPONSE FORM; complete and return the response form via fax to 574-372-1683 prior to return of product; use priority carrier for shipment, and please confirm receipt of this notice by calling 800-348-9500, extension 3755 or 3756. The firm also informed physicians that had previously utilized the device to monitor patients for infection and provide treatment as appropriate. Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8am to 5pm. |  | 35 | 35 | Worldwide distribution: USA including states of: IN and KY; and countries of: Canada and The Netherlands. | Terminated | July 27, 2011 | 139 | 35 | Not_Computer
Z-1901-2011 | 58208 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prismaflex System, Catalog Number 107493, |  |  | March 14, 2011 | 2011 | Gambro Renal Products, Incorporated | Prismaflex touch screen may become unresponsive during Continuous Renal Replacement Therapy procedure. | Software change control | The firm, GAMBRO Renal Products, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated March 14, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were provided with response procedures to use if the frozen screen problem occurs while using the product. The customers were instructed to closely monitor the return line for potential air and/or clots in the circuit; review the instructions provided in the attached copy of Chapter 5: Troubleshooting, "Softkeys won't work." of the Prismaflex Operators Manual in conjunction with the attached Troubleshooting Checklist; be aware if performing Manual Termination with Blood Return, air detection is not provided, and complete and return the Customer Reply Form via fax to: Regulatory Affairs at 1-866-447-4471 or 303-222-6916, or by email to: GRPRegulatory.Affairs@us.gambro.com. If you have any questions or concerns, please contact for Clinical questions: ICON 800-554-1312 (available 24/7) or your local Gambro Representative; for Technical questions: Gambro Customer Support 800-525-2623, option 2 (Available 6:00am to 5:30pm MT, Monday thru Friday). |  | 1300 | 1300 units | Nationwide distribution. | Terminated | September 01, 2011 | 171 | 1300 | Other |  | Instructions | Safety Notice/Insructions
Z-1904-2011 | 58099 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | Zoll Auto Pulse Resuscitation System; Zoll Circulation Inc., Sunnyvale, CA 94085 |  |  | January 28, 2008 | 2008 | Zoll Circulation, Inc. | Complaints of batteries discharging prematurely and that during use the Auto Pulse Units would cease operating (providing compressions). | Other | The firm, Zoll Circulation Inc., initiated a Corrective and Preventive Action Report on January 28, 2009. According to the firm, the customers were contacted by telephone calls and were provided with a Battery Maintenance program. For further questions you may call (408) 541-2140 ext.341. |  |  | not provided | Nationwide Distribution -- including, VA, NY, OH, IN, MO, UT, FL, NC, ME, WA, SD, NM, PA, MS, ID, TX, GA, CA, CO, IL, AR, WV, TN, OR, and KS. | Terminated | April 28, 2011 | 1186 |  | Battery | Device Operation | Repair | Repair
Z-1905-2011 | 58112 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA versions 8.0 and 8.1 |  |  | February 18, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was identified in the ARIA version 8.0 and 8.1, when an image sequence template is added to a field in RT chart and a CIAO (Completely Irradiated Aperture Opening) is calculated, the resulting CIAO display at 4DITC may be incorrect. | Other | Varian Medical Systems sent an Urgent Medical Device Correction letter dated February 18, 2011, to all affected customers via Federal Express.. The letter identified the product, the problem, and the action to be taken by the customer. Varian has made design changes to address the issue in ARIA Radiation Oncology versions 8.2 and later. A corrective action will be undertaken to upgrade all affected items.For questions regarding this recall call US and Canada 1-888-VARIAN5 (888-827-4265.Europe +41 41 749 8844 |  | 645 | 645 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Azerbaijan, Bahrain, Bangladesh, Barius, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guadeloupe, HOng Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Lithania, Malaysia, Maurtama, Morocco, Netherlands, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Yemen. | Terminated | June 19, 2012 | 487 | 645 | Computer
Z-1915-2011 | 56977 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Penner Manufacturing Transfer Electric Bariatric with Scale, Patient Transfer/Lift System, Cascade Transfer, Models 383500-1 (right entry with scale). The firm name on the label is Penner Manufacturing Inc., Aurora, NE. |  |  | January 19, 2009 | 2009 | Penner Mfg Inc | Incomplete weld on the pillar mounting bracket | Employee error | The firm, Penner, sent an undated letter flagged "Important Safety Inspection: Action Required" via certified mail on January 19, 2009 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect the Penner Patient Transfer for any missing Welds of the Pillar Mounting Bracket; inspect the areas indicated in the enclosed photos to confirm the bracket is welded on all edges to the main frame of the base; complete the inspection, and complete and return the attached WELD INSPECTION RESULTS OF PENNER PATIENT TRANSFER form via fax to 402 694-5844 or mail to Penner Manufacturing Inc, 102 Grant Street, Aurora, NE 68818, Attn: Quality Assurance Manager. Penner stated in the letter, if welds are not complete (no weld), discontinue use and contact the Quality Assurance Manager at Penner Manufacturing immediately at 402-694-5003.If you have any questions, contact the Quality Assurance Manager at 402-694-5003. |  | 20 | 20 lifts | Worldwide distribution: USA and countries including: Canada and Mexico. | Terminated | April 08, 2011 | 809 | 2290475 | Not_Computer
Z-1929-2011 | 57288 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE |  |  | March 08, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential safety issues associated with the CARESCAPE" Monitor B850 when Patient Data Module (PDM), Tram Acquisition Module (Tram), Patient Side Module (E-PSM or E-PSMP), E- MASIMO or E-(P)RE(S)TN modules or Bed-to-bed alarming feature is used.1. When CARESCAPE Monitor B850 is used with the above listed modules, some SpO2 alarms may be delayed or missed.2. Whe | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the recall product the problem, Safety Issues, Affected Product details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to follow the safety instructions provided in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available.If you have any questions regarding this medical device correction or the identification or affected items please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 3817 | 3817 | Worldwide distribution: USA including states of:AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA,WA, WI, and WV; and countries including: JERUSALEM, UNITED KINGDOM, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLAND, MALTA, MALAYSIA, LEBANON, KUWAIT, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CZECH REPUBLIC, CHILE , CANADA, BELGIUM, BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | June 05, 2012 | 455 | 3817 | Software | Alarm/Message | Software update | Software Update
Z-1925-2011 | 57866 | No | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Inhibin A ELISA kit, Part Number: DSL-10-28100-4 |  |  | December 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the lots of Inhibin A identified above may contain microbial contamination in the conjugate diluent bottlesincluded in the kit. | Pending | Beckman Coulter sent an Urgent Product Corrective Action letter dated December 21, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product and discard all remaining inventory. In addition customers were requested to share the information with their laboratory staff and retain the notification as part of their laboratoryQuality System documentation. Customers were instructed to complete and return the enclosed response form within 10 days soBeckman could be assured that customers received this important notification.For product replacement:1. In the United States, please contact Customer Service at 1-800-526-3821, option 1.2. In Canada, please contact Customer Service at 1-800-463-7828.3. Outside of the United States and Canada, contact your local Beckman Coulter representative.Customers who needed assistance or had any questions regarding this notification, were instructed to contact Technical Support at 1-800-854-3633 in the United States and Canada. For those customers outside the United States and Canada, instructions were given to contact their local Beckman Coulter representative. |  | 740 | 740 total units for both products | Worldwide Distribution - USA including NY, IN, PA, NM, OH, TX, AZ, NC, TN, MI and CT and the countries Denmark, Germany, Spain, Turkey, United Arab Emirates, United Kingdom. | Terminated | August 31, 2012 | 619 | 740 | Not_Computer
Z-1907-2011 | 58006 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | 8V3c transducer used with the ACUSON SC2000 ultrasound system. 8V3c transducer model number 10433823; ACUSON SC2000 model number 10433816. |  |  | February 04, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | When using the 8V3c transducer in color mode, the derated Ispta 3 (Spatial Peak-Temporal Average Intensity) may exceed the limit stated in the instructions for use. | Software design | A Customer Safety Advisory Notification was initiated on 2/4/2011 and sent via certified mail. The letter identified the affected product and described when the problem occurs along with the associated potential risk. The firm recommends that users not use the 8V3c until the software issue is fixed. The firm is working to release a software update and customers' local Customer Service Engineer will install new software on their system when it is available. Customers should direct their questions to their local service support person. The information in the letter should be shared with all personnel within the organization who need to be aware. |  | 386 | 386 units | Worldwide Distribution. | Terminated | October 04, 2011 | 242 | 386 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1909-2011 | 58093 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | SPIDER Surgical Platform |  |  | February 25, 2011 | 2011 | Transenterix, Inc | Plastic shards may break off and fall into the abdominal space which may require retrieval during surgery. | Other | TransEnterix sent an Urgent Device Recall letter dated February 25, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the lot numbers affected, complete the attached inventory reconciliation form and return the form, along with any affected product toTransEnterix using the enclosed shipping label. Customers were asked to return the inventory form even if they had none of the lot numbers affected in their inventory.Their sales representative and customer service will work with them to replace the product in full at not cost to them.For questions regarding this recall contact TransEnterix Customer Care at 1-888-879-4111. |  | 305 | 305 units | Nationwide Distribution including AZ, FL, NC, NY, OH, OR, SC, TX, VA and WA | Terminated | November 07, 2011 | 255 | 397 | Not_Computer
Z-1928-2011 | 58014 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilator with software version V2.00x and Neonatal Option or V2.01x and Neonatal software. Hamilton-G5: catalog number 159001; |  |  | February 16, 2011 | 2011 | Hamilton Medical Inc | The adaptive volume/pressure regulation can be automatically and permanently reduced to PEEP + 3 mbar when a closed suctioning maneuver is performed on the patient and may be undetected by the operator. | Software design | A Medical device Safety Alert and Corrective Action along with a cover letter, was sent to all locations on February 16, 2011. The letter identified the affected product along with the reason for recall. The letter states that operators deactivate the "high tidal volume" alarm as described in appendix K.4 of the operator's manual until the ventilator is upgraded with a corrected software. Distributors will upgrade the affected devices with the corrected software as soon as possible. Questions can be directed to 800-426-6331, ext 215. |  | 97 | 97 units | Nationwide Distribution -- MA, FL, GA, NC, MI, NC, MO, TX, and WV. | Terminated | March 21, 2012 | 399 | 97 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1930-2011 | 58275 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650 |  |  | March 23, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue associated with the CARESCAPE Monitor B650 &amp; B850 Cardiac Output measurement using Patient Data Module (PDM), and activation of ECG Leads Off alarm in specific cases.1. When the CARESCAPE Monitor is used to measure Cardiac Output values with the Patient Data Module (PDM) or TRAM Module, Cardiac Output values may be calculated incorrectl | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated March 23, 2011 to its consignees/customers. The letters were addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letters described the Safety Issues, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. The customers were instructed to follow the safety instructions documented in the letter. GE Healthcare is working on a software correction that will be provided to their customers once it is available and will contact the customer to schedule the software update.If you have any questions regarding this medical device correction or the identification of affected items, please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 1443 | 1443 | AL, AZ, AR, CA, CO, CT, FL,GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA,PR, SC, TN,TX, UT, VA, WA, WV, WI.UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SWITZERLAND, SWEDEN SLOVENIA, SLOVAKIA, SPAIN, SINGAPORE, SAUDI ARABIA, QATAR, PORTUGAL POLAND, NORWAY, NORTH KOREA, NEW ZEALAND, NETHERLANDS, MALTA, MALAYSIA, LITHUANIA, LEBANON, KUWAIT, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, DENMARK, ESTONIA, CZECH REPUBLIC, BELGIUM BAHRAIN, COLOMBIA, CHILE CANADA, BRAZIL, AUSTRIA, AUSTRALIA. | Terminated | February 04, 2012 | 318 | 8345 | Computer
Z-1927-2011 | 58143 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Inflow/Outflow Cannula; |  |  | March 07, 2011 | 2011 | Stryker Endoscopy | Insufficient cleaning instruction were included for the Inflow/outflow Cannula in the user manual. | Other | The firm, Stryker, issued an "URGENT: DEVICE CORRECTION" memo dated March 8, 2011 to all customers. The memo described the product, problem and actions to be taken. The customers were instructed to discard any old Instructions for Use (1000-400-816 rev. A or B) for their Inflow/Outflow Cannual; download the new IFU (1000-400-816 rev. C) from the URL link: http://stryker.com/InflowOutflowCannual; ensure they follow the new cleaning instructions provided in the IFU and complete and return the enclosed "acknowledgement of receipt" form and fax to (408) 754-8378 or scan in and email to cannula@stryker.com. Stryker stated in the memo that "It is important to send the self addressed confirmation forms back to Stryker RAQA". If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at cannula@stryker.com. |  | 2618 | 2618 devices | Worldwide distribution: USA including states of: AL, ALASKA, AK, AR, CA, CO, CT, DC, DE, FL, GA, ID IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of: Australia, Canada, China, France, Germany, India, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Switzerland and United Kingdom. | Terminated | October 04, 2011 | 211 | 2618 | Not_Computer
Z-1924-2011 | 57141 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Quest MPS and MPS 2 Consoles, Quest Medical, Inc., An Atrion company. Single software controlled device that incorporates a pump, a heater/cooler, temperature monitoring, pressure monitoring, a heat exchanger, an arrest agent pump and an additive pump. The MPS console is used with the MPS sterile disposables with a built in heat exchanger. The device is used with on-pump or off-pump cardiac surgery. |  |  | July 29, 2010 | 2010 | Quest Medical, Inc | User update regarding installation of external heater-Cooler units with MPS and MPS2 consoles. | Other | Quest Medical Incorporated sent a "User Update" letter in July 2010 to consignees providing information concerning use of external heater-cooler untis with their device - Myocardial Protection System ( MPS) consoles. If you are installing an external heater-water cooler unit or have questions about this user update, please contact the Field Service department at Quest Medical, Inc. The number is 800.627.0026 or 1 972.390.9800. |  | 771 | 771 consoles | Nationwide, Canada, Saudi Arabia, Kuwait, Japan, and Italy | Terminated | April 26, 2011 | 271 | 771 | Not_Computer
Z-1932-2011 | 54483 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand VARIANT II TURBO Hemoglobin Alc Program, 1600 and 2000 tests Cartridge Sets. Catalog Numbers: 270-2417 and 270-2415; Distributed and Manufactured by: Bio-Rad Laboratories, Inc. Hercules, CA |  |  | December 01, 2009 | 2009 | Bio-Rad Laboratories Inc | Some customers have reported the occurrence of some ramping baselines on the chromatograms. The ramping baseline can affect hemoglobin test results. | Software change control | Bio Rad Laboratories, Inc. sent an Urgent Medical Device Correction letter dated December 1, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediatey examine their records to identify all customers who received one of the affected lots.Collect a Customer Medical Device Correction Response Form from every customer.Once forms were received from all applicable customers, complete the following Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or pdf to CSD Recall Coordinator (excluding USSD).For questions regarding this recall call 510-724-7000. |  | 1099 | 1099 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Taiwan, Singapore, India, China, S. Korea, Hong Kong, France, Sweden, United Kingdom, Italy, Hungary, Germany, Brazil, Mexico and Colombia. | Terminated | August 30, 2011 | 637 | 1099 | Not_Computer
Z-1943-2011 | 57189 | No | DTY | Cardiopulmonary bypass in-line blood gas sensor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems Corporations, CDI 101 Hematocrit/Oxygen Saturation Monitoring System. In-line blood-gas sensor used in cardiopulmonary bypass. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Inaccurate result. Possibility for displayed values for potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other measured blood parameters to be inaccurate. | Device Design | On 11/1/2010, Terumo sent "URGENT MEDICAL DEVICE CORRECTION" notices to Chief Perfusionist, Department of Cardiovascular Surgery or Director of Operating Room Services. The letter informed the consignees about the incidents of the CDI 100 and 101 displaying inaccurate values. The letter states that when a displayed value is significantly different from expectations based on the clinical situation, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operators manual. |  | 468 | 468 monitors | US, India, Colombia, Taiwan, Venezuela, Vietnam, Puerto Rico, Malaysia, Sri Lanka, Brazil, Indonesia, Pakistan, Chile, Germany, Hong Kong, Japan, AE, Singapore, Austria, Belgium and Thailand. | Terminated | April 17, 2012 | 533 | 1500 | Not_Computer
Z-1941-2011 | 58098 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Transmitter Sensor Cables, Part Numbers: 1002008, 1004587, 1007914-NAV, GE Healthcare Surgery, Salt Lake City, UT 84116. |  |  | February 24, 2011 | 2011 | GE OEC Medical Systems, Inc | Sensor cables may suffer material degradation when exposed to certain sterilization procedures. | Component design/selection | GE OEC Medical Systems, Inc sent an Urgent Recall Notice dated March 1, 2011, by certified letter on February 24, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if their Transmitter Sensor Cables or Receiver Sensor Cables have been exposed to any STERRAD sterilizer other than STERRAD 1005, discontinue use of these sterilizers for this product. The Transmitter Sensor Cables should be inspected for signs of material degradation.Further information is available at 800-874-7378. |  | 708 | 708 cables | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NY, OH, OR, PA, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI and WV and the countries of Austria, Canada, China, Denmark, England, France, Germany, Hong Kong, India, Ireland, Italy, Japan, San Salvador, Slovakia, Spain, and Switzerland. | Terminated | September 09, 2011 | 197 | 708 | Not_Computer
Z-1933-2011 | 57953 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Adjustable (retractable) Strap and Handle, type TGN-R100GR; a component of a Lifting Pole, an accessory to long term care hospital beds such as ArjoHuntleigh's Minuet 2 Community bed and Enterprise series hospital beds; Distributed by ArjoHuntleigh, a Branch of Arjo Limited Med AB, Wednesbury, United Kingdom |  |  | February 24, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | Part of the housing of the strap retracting mechanism built into the handle may break while the patient is using the handle and strap. | Other | ArjoHuntleigh sent an Urgent Field Safety Notice to all affected end users via UPS Ground on February 24, 2011. The notice informed the accounts of the potential for the strap retracting mechanism to break during use. The accounts were requested to check their inventory for the presence of adjustable straps and handles procured from ArjoHuntleigh, note if they have a red button on the retractor housing and the affected production dates as explained in the photographs in Figures 1 and 2 in the letter. Any straps and handles found as identified in Figures 1 and 2 should be removed from use and returned to ArjoHuntleigh. Any straps and handles with a yellow button design as shown in Figure 3 are not affected and can continue being used. The accounts were requested to complete and fax or mail the enclosed Customer Response Form to Arjo at 630-576-5012, indicating the number of strap and handles checked and how many were removed from use. Questions were directed to the ArjoHuntleigh Regulatory Department at 1-630-785-4885. |  | 90 | 90 units | Nationwide Distribution including California, Georgia, Hawaii, Michigan, South Carolina and Texas | Terminated | June 14, 2011 | 110 | 90 | Not_Computer
Z-1940-2011 | 57952 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 660i, 680i, 860i, 880i |  |  | January 26, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated due to Beckman Coulter receiving a report of results that were associated with incorrect patientidentifications when two sample racks advanced into the sample presentation area in one push. The firstrack of the two was not scanned. This is referred to as a "pushed pair." If a pushed pair event occurs, results may be associated with the incorrect patient identification. | Pending | Beckman Coulter sent an Urgent: Field Safety Notice Product Correction letter dated January 26, 2011, via US Postal Service to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:1. To prevent a pushed pair, be sure your sample racks are free of debris prior to use. Do not applyadditional labels to the sample racks. Extra labels can obscure the edges of a sample rack, orobstruct the gap between two racks. Be sure sample labels do not extend beyond the edges of asample rack;2. If a pushed pair occurs on the UniCel DxI system, the system generates a yellow Event Log message- Clear the presentation tray and press ROUTINE button. This Event Log message is associatedwith a variety of common Sample Presentation Unit (SPU) errors. For that reason, you must reviewthe Event Log details to confirm the pushed pair event has not occurred;3. Detach the Appendix included with this notification and store it near your UniCel DxI system. Use theAppendix as a reference if you suspect a pushed pair event has occurred on your system.Beckman Coulter also requested their customers share the recall information with their laboratory staff, and retain pages 1 and 2 of this notification as part of your laboratory Quality System documentation.For questions regarding ththe recall notification, customers were instructed to contact Beckman Coulter Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support, or call1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact theirlocal Beckman Coulter Representative. |  | 2639 | 2639 units (892 in US) | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Liechtenstein, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | June 04, 2012 | 495 | 2639 | Not_Computer
Z-1945-2011 | 58163 | No | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Free T4 Calibrator, Part Number: 33885, Lot Numbers: 014769 |  |  | November 11, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the stability claim for Access Free T4 calibrator lot 014769 is not met. The affected T4 calibrator lot can produce quality control and patient results that are erroneously low (by greater than 2 standard deviations). An erroneously low patient result may lead to misdiagnosis and/or delayed treatment. | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter dated November 29, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: ACTION/RESOLUTION:" Discontinue use of Access Free T4 calibrator lot 014769." Review historical quality control performance and determine whether any reported patient results were obtained while using a faulty calibration. If results were affected, you will see a shift downward in control recovery that produces out-of-range low QC results." If patient results were obtained with Access Free T4 calibrator lot number 014769, it is possible they are erroneously low. If any patient results were reported in the period between the time the erroneous calibration with the affected lot was performed and when a different calibrator lot was used, the laboratory should consider advising physicians to whom results were reported so they can evaluate whether diagnosis or treatment decisions should be revisited." No further orders of Access Free T4 Calibrator lot number 014769 will be shipped. Alternate lots will be provided for current and future replenishment orders or requests.Consignees/customers were asked to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation; and complete and return the enclosed RESPONSE FORM via fax to Beckman Coulter at 786-639-4000. For product replacement:" In the United States, please contact Client Services at 1-800-526-3821, option 1." In Canada, please contact Customer Service at 1-800-463-7828." Outside of the United States and Canada, contact local Beckman Coulter representative.Questions and answers regarding this issue were enclosed for consignee/customer reference. Additional questions, were directed to the Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customer |  | 459 | 459 units | Worldwide distribution: USA and countries including: Argentina, Belgium, Bosnia and Herzegovina, Canada, China, Colombia, Czech Republic, French Polynesia, Hong Kong, Ireland, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Oman, Panama, Philippines, Russian Federation, South Africa, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Venezuela, and Zambia. | Terminated | July 11, 2012 | 608 | 459 | Not_Computer
Z-1961-2011 | 52928 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Auxiliary 2 MHz Continuous Wave (CW) transducer, model number 8251046, when used with ACUSON Sequoia ultrasound systems at SW revision 12.X. Manufactured by Siemens Medical Solutions USA, Inc., Ultrasound Business Unit, Mountain View, CA. |  |  | March 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Overheating. A software bug may cause the transducer to overheat when used with certain transducer models. The transducer can become noticeably hot within a few minutes, which may cause patient discomfort or burns. | Software design | The firm, SIEMENS, sent a "Customer Safety Advisory Notification" to its customers on 3/21/2009. The notification described the product, problem and actions to be taken. The customers were instructed to disconnect any transducers that may be in an active transducer port and chose another imaging transducer prior to use of the Aux CW transducer. If they do not have another transducer to select other than those listed in the notification, they should close the study in progress and start a new study before using the Aux CW transducer-this procedure should be followed for each new patient and /or exam type. In addition, the customers were instructed to pass this notice on to all those within their organization who need to be aware of this issue until the corrective action is completed. SIEMENS will release a software update that will correct the issue. The customers local service engineer will install the updates on their system as soon as it becomes available.Should you have any questions regarding this notification, you may contact your local service support person and/or call (650) 969-9112. |  | 2591 | 2591 units | Worldwide distribution. | Terminated | March 09, 2012 | 1084 | 2591 | Software | Physical Safety Hazards | Instructions to disconnect transducer/Software update | Software Update
Z-1841-2011 | 58076 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Ultra Q LQP3106-U ophthalmic laser |  |  | February 15, 2011 | 2011 | Ellex Medical Pty Ltd | Unintended laser emission while the user was moving the joystick for targeting. | Process control | Ellex notified all distributors that have received the affected devices per Global Customer Service Bulletin GSB-11 -01a, dated 25 February 2011. |  | 233 | 233 units | worldwide distribution | Terminated | October 02, 2012 | 595 | 2335 | Not_Computer
Z-1958-2011 | 58122 | No | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | Dental chairs distributed under several brand names such as DCI Equipment; Marus MaxStar; Marus NuStar; and Pelton &amp; Crane. |  |  | March 01, 2011 | 2011 | Dental Equipment LLC | Link arm assembly that mounts the rear mounted accessories to the dental chair may potentially break and fall off the dental chair. | Process control | On 03/08/11, Dental Equipment LLC. started sending out the Medical Device Recall notification letter via USPS 1st class mail to US consignees and Federal Express mail to foreign consignees.The firm also sent a recall notification letter to consignees who only purchased dental accessories such as Rear Mounted Cuspidor with Coverset, Arm Articulating Rear Mounted Cuspidor - Marus Gray, Maxstar Orbit Cuspidor utility W/ Porcelain bowl, and zinc plated bolts.The letter identified the affected dental chair as Marus brand with the following model numbers: DC1490, DC1690, DC1700 and DC1702.The firm's corrective action including the replacement of the affected link arm assembly with a new link arm assembly. The firm's field representative will contact each consignee for an appointment to replace the affected link arm assembly.Consignees can contact the Dental Equipment LLC at 503-537-3617 for any questions. |  | 66 | 55 chairs and 11 replacement bolts distributed to US and foreign consignees outside the US. | Worldwide distribution to USA, Australia, Canada, Ireland, South Africa, and Vietnam. | Terminated | September 21, 2011 | 204 | 66 | Not_Computer
Z-1959-2011 | 58110 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos FX and FX Plus System |  |  | February 17, 2011 | 2011 | Siemens Medical Solutions USA, Inc | During product monitoring, firm discovered a possible hazard to patients when using the detector alpha rotation of all Axiom Aristos FX and FX Plus systems with serial numbers up to and including 1216. When using the detector alpha rotation, the detector may descend unexpectedly during movement or in vertical or table mode. If this occurs in vertical modes, the possibility exists that the detect | Other | The firm, SIEMENS, sent a "Safety Advisory Notice" dated February 15, 2011 to the affected customers via Update Instructions XP016/10/S on February 17, 2011. The notice described the product, problem and actions to be taken. The customers were instructed to immediately instruct their personnel accordingly; ensure that the safety advisory is placed in the systems's instructions for use; if they have sold this device/equipment and it is no longer in their possession, forward this safety notice to the new owner of the device/equipment; inform SIEMENS about the new owner of the device/equipment, and complete and return the Acknowledgement of receipt form. SIEMENS is preparing a modification of AXIOM Aristos FX and FX Plus that will resolve this potential malfunction. The field modification will be available from February 2011 and is expected to be carried out May 31, 2011. If you have any questions, call 09131-18-9908 or +1-888-826-9702. |  | 101 | 101 | Nationwide Distribution: USA including states of: AZ, AK, CA, CO, FL, GA. IL, IN, IA, KS, KY, ME, MI, MN, MS, MO, NE, NV, NJ, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, and WI. | Terminated | July 16, 2013 | 880 | 101 | Not_Computer
Z-1957-2011 | 58072 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Varis Vision - RT Chart version 7.3.10. For radiation therapy. |  |  | February 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue with the Varis Vision software where treatment can occur without the dynamic Multileaf Collimator component and may deliver a higher than expected dose of radiation. | Software design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected customers on February 7, 2011. |  | 421 | 421 | Worldwide Distribution. | Terminated | December 23, 2011 | 319 | 421 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1956-2011 | 57990 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | Coulter GEN's |  |  | January 31, 2011 | 2011 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter identified a problem with some date fields that are transmitted to the Laboratory Information System (LIS). Starting Jan 1, 2011, the date fields will transmit as a blank field to the LIS if the year is greaterthan 2010. The following dates are affected: Run date, Control expiration date, Date of birth(For individuals born in 2011 or later), a | Software design | The firm, Beckman Coulter, sent an "URGENT: PRODUCTION CORRECTION" letter dated January 31, 2011 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Coulter LH 750 Analyzer, Coulter LH 780 Analyzer, Coulter LH 500 Series System, and Coulter GEN S System. The letter describes the product, problem and actions to be taken. The customers were instructed to work with their US vendor to assess how this affects their laboratory. The actions they choose are dependent on the configuration of their LIS.ACTION FOR ISSUE 1:Use Sample Drawn Date and Date of Birth information available at your LIS in lieu of the analyzer transmitted fields BIRTH and UDATE.Use the current date when the results are received at the LIS in lieu of the transmitted Run Date [DATE]. Print results from the Workstation to keep a copy of the analysis date on the analyzer, if necessary.Use the control expiration date from the analyzer display or printout.ACTION FOR ISSUE 2:If you use the LIS 1G1 transmission format, use the General Information group field names in place of the Retic Parameter group field names to identify each reticulocyte run (see Table in Issue 2 section).Check with your LIS vendor to determine if it is possible to switch to the LH 750 Workstation transmission format. Format requirements are LIS driver specific.For patients with Dates of Birth in 2011 for both Issue 1 and Issue 2, results should be reviewed to determine if decision rules or age-related limits sets were triggered inappropriately.RESOLUTIONThese issues will be corrected in a future software release for the LH Series products. The GENÂżÂżS product has been discontinued; no software revisions will be forthcoming for the GENÂżÂżS product line. Customers were also instructed to complete and return the PRODUCT CORRECTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 310B06, 11 |  | 9419 | 9419 (total units distributed in U.S. &amp; foreign countries) | Worldwide distribution: USA (nationwide) and countries including: Algeria, Andorra, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bermuda, Bulgaria, Burundi, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, France, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Monaco, Morrocco, Myanmar, Namibia, Netherlands, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Virgin Islands (US). | Terminated | September 25, 2012 | 603 | 9419 | Computer
Z-1934-2011 | 57450 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Waste Bottle |  |  | October 05, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter confirmed Access waste bottles that ruptured and expelled liquid waste, exposing the operator to the system liquid waste. Over time, waste bottles can degrade, become brittle and potentially crack. | Pending | The recall communication was initiated on October 5, 2010 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access immunoassay Systems, UniCel DxC 600i Access, and Immunoassay Systems, Synchron LxI 725 Access Clinical Systems. The letter explains the problem identified, the impact, and actions required to be taken by customers. The letter instructs customers to inspect the waste bottle for signs of aging and deterioration. New bottles are opaque white in color.Signs of deterioration include:o Any color change to the bottle (yellow, orange or pink)o Crackingo Crazing (spider web cracking)Any defect that is not removed by routine cleaning is a sign of deterioration. It is important to routinely inspect the area nearest the handle and along all seams of the bottle. If upon inspection consignees detect any issues with their waste bottles, they were asked to contact Beckman Coulter Customer Service using the enclosed Replacement Order Form to obtain a new waste bottle, and keep this information should they require a replacement bottle in the future.Consignees were also asked to share this information with laboratory staff and retain the notification as part of laboratory Quality System documentation. They should also complete and return the enclosed Product Corrective Action Response Form within 10 days so Beckman Coulter can be certain they received this important notification.Beckman Coulter offers a service kit which may be used to replace the waste bottle and direct liquid waste to a floor drain. Beckman Coulter requires a consultation with consignees technical support representative to determine if their system can be fitted with a floor drain kit. If assistance is needed or consignees have any questions regarding this notification, they were asked to contact Technical Support at 1-800-854-3633 i |  | 7700 | 7700 units total (3178 in US) | Worldwide Distribution -- United States, Afghanistan, Algeria, Angola, Antigua and Barbuda, Australia, Austria, Azerbaijan, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Congo, Costa Rica, Croatia, Czech Republic, Djibouti, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, French Polynesia, Georgia, Germany, Ghana, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestinian Territory, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, Yemen, Zambia, and Zimbabwe. | Terminated | July 25, 2012 | 659 | 7700 | Not_Computer
Z-1952-2011 | 57830 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ArthroCare Opus SpeedLock Knotless Fixation Device; Catalog Number: OM-7500. |  |  | September 30, 2010 | 2010 | ArthroCare Corporation | During surgical procedure the product has the potential for separation of the plug guide. Separation of the plug guide may compromise the ability of the inserter to fully drive the implant completely into the bone, which may also compromise the ability to tension the suture properly. | Process control | On 9/30/2010, Arthrocare notified customers via letters explaining the reason for the recall and requesting the product be returned. |  | 170 | 170 units | Product was released for distribution in US. | Terminated | August 01, 2011 | 305 | 170 | Not_Computer
Z-1950-2011 | 57815 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medstorm Adult Radiolucent Leads-Out Electrode - Physio-Control, Part 16242. |  |  | December 05, 2008 | 2008 | Heart Sync, Inc | The "leads out" connector had bent pins which would not allow connection to the Medtronic Physio Control Defibrillator cable during a product demonstration. | Equipment maintenance | The firm sent a recall notice to their distributors on 12/3/2008 asking them to quarantine product and return it to HeartSync. The letter stated that customers' current inventory would be exchanged for new product. Customers were to call 1-800-828-4681 to provide the necessary information. |  | 21000 | 2100 (210 cases of 10 each) | Nationwide Distribution | Terminated | April 13, 2011 | 859 | 21500 | Not_Computer
Z-1947-2011 | 57385 | No | HBX | Cranioplasty material forming instrument. | Neurology | Neurology | 510(K) Exempt | DePuy Mitek FMS Irrigation Intermediary Tube Without One-way Valve Arthroscopy tubing |  |  | December 03, 2010 | 2010 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product sterility may be compromised due to breach in sterile packaging | Packaging process control | Depuy Mitek sent an Urgent Voluntary Product Recall letter dated December 2, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to isolate all inventory of the affected product and call 1-800-382-4682, PROMPT 3, for return good authorization and replacement product. Customers were instructed to return all identified affected product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324Record their assigned return material authorization (RMA) number on the outer shipping box.Use the Prepaid Federal Express shipping label for the return of the product.Include a photo copy of the completed business reply form.For questions or concerns call 1-800-382-4682, PROMPT 3. |  | 7344 | 7344 units (372 boxes x 24 pieces) | Worldwide Distribution - USA (nationwide) and the coutnries of United Arab Emirates, Australia, Belgium, Switzerland, Czech, Denmark, France, UK, Italy, Netherlands, Poland, and Slovakia. | Terminated | April 02, 2013 | 851 | 149720 | Not_Computer
Z-1967-2011 | 58304 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ExactaMix EVA Container, 3000mL, REF: 741, Sterile R, Baxa Corporation, Englewood, CO 80112. |  |  | March 08, 2011 | 2011 | Baxa Corporation | TPN bags may leak during filling operation, due to bag manufacturing error. | Nonconforming Material/Component | The firm, BAXA, sent a "ExactaMix Bag Product Notice" dated March 8, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to do the following:If they do not intend to use the manual addition port, verify that the cap is firmly seated onto the port by gently tugging on the cap. If they do intend to use the manual addition port, they should snap off the additive port cap prior to compounding the bag.If they encounter a faulty additive port, contact Baxa Technical Support to receive credit and/or replacement on defective product only.To return affected bags, follow the instructions provided in the notice if the bags were purchased directly from Baxa or purchased from wholesaler, distributor or other 3rd party vendor. Additional copies of the return form can be retrieved by going to www.baxa.com/ebug OR by calling Baxa Technical Support at 800-678-2292.If you have any questions or concerns, contact the Sr. Vice President, Quality and Regulatory at (303) 400-2838 or Baxa Technical Support at 800-678-2292. |  | 336650 | 336,650 bags | Worldwide distribution: USA (nationwide) and countries including: Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, Hong Kong, Ireland, Japan, Malaysia, Mexico, New Zealand, Panama, Puerto Rico, Singapore, South Korea, Taiwan, UK, Uruguay, Venezuela. | Terminated | March 23, 2012 | 381 | 6561200 | Not_Computer
Z-1960-2011 | 58290 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | STERIS SYSTEM 1E Liquid Chemical Sterilant Processing System |  |  | March 14, 2011 | 2011 | Steris Corporation Hopkins Facility | The printout, display and Operator Manual for the STERIS System 1E Liquid Chemical Sterilant Processing System are subject to recall/correction by the firm because in the event of an incomplete or aborted cycle, the language of the warning messages used in the printout, display, and Operator Manual omits the qualifying term: *liquid chemically* (not **sterilized), when alerting the user to inco | Other | The firm, STERIS, Issued an 'Urgent Field Correction Notice' dated March 14, 2011 addressed to their Hospital Administrator, or Manager, and Risk Manager customers on March 18, 2011. The notification letter describes the product subject to field correction and the problem that exists with the language omission present in the printout display and Operator Manual messages which warn of an incomplete or aborted processing cycle. The notification goes on to state that there is no health hazard involved in this error, as the users are still made aware that the processing cycle has not been successfully completed. The notification goes on to remind the user/customers that in case of any incomplete, cancelled or aborted cycle, the users are instructed to process devices in the cycle following the directions contained in the Operator Manual. Lastly, the notification letter informs the user/customers that the company will be installing new software on the recalled systems in order to revise the language on the cycle printout and the display screen warning of a disrupted cycle. It also states that a STERIS Service Representative will be contacting each facility in order to schedule installation of the new software. The notification also provides the customers with a 1-800 telephone number for STERIS Field Service Dispatch, to be used if the customers have questions or want to request additional information. It also provides a local telephone no. 1-440-392-7455, for obtaining additional information from the firm's Director, Div. of Low Temperature Sterilization at: 1-440-392-7455. |  | 106 | 106 units | Nationwide distribution: USA including states of :CA, DE, FL, GA, IA, ID, IL, KS, MA, MO, MT, NE, NJ, NM, NY, OH, OR, SD, TX, VA, WA, WI, and WV. | Terminated | March 13, 2013 | 730 | 106 | Software | Alarm/Message | Software update | Software Update
Z-1949-2011 | 58169 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Total T4 Reagent Kit, Part Number: 33800, |  |  | March 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the concentration of total T4 in samples from pregnant women is erroneously low (20%)when measured using the Access Total T4 assay. Erroneously low results which are actually high (outside of the normal reference range)could mask hyperthyroidism. Failure to diagnose or treat hyperthyroid disease associatedwith pregnancy could result in serious, permanent harm | Other | Beckman Coulter sent an Urgent Product Correction letter dated March 20, 2011, to all affected consignees. The letter identified the product, the problem and the action to be taken by the consignee. Consignees were issued the following instructions:(1) If you suspect pregnant patient results were evaluated at your laboratory using only Total T4 to report thyroid status, consider reviewing Total T4 results reported in the last 9 months. Consult with your Laboratory Medical Director and requesting physician toresolve any concerns you may have regarding previously reported patient results during this time period.(2) Do not use the Access Total T4 assay as the only marker for evaluating pregnant patients for thyroid disorders.(3) Please share this information with your laboratory staff and retain this notification as part ofyour laboratory Quality System documentation. If you have forwarded Access Total T4 reagent kits (Part Number 33800) to another laboratory, please provide a copy of this letter to them.(4) Please complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For questions regarding this notice, consignees were instructed to contact Beckman Coulter Customer Support Center.(1) Via our website, http://www.beckmancoulter.com/customersupport/support(2) Via phone, call 1-800-854-3633 in the United States and Canada(3) Outside the United States and Canada please contact your local Beckman Coulter |  | 151740 | 151,740 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Egypt, El Salvador, Ethiopia, France, Germany, Greece, Guatemala, Hong Kong, India, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Libyan Arab Jamahiriya, Liechtenstein, Malaysia, Mexico, Republic of Moldova, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Pakistan, Panama, Philippines, Poland, Puerto Rico, Reunion, Romania, Russian Federation, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, Yemen, and Zambia | Terminated | August 10, 2012 | 519 | 151740 | Not_Computer
Z-1906-2011 | 58175 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Penumbra brand Penumbra Coil System, Penumbra Coil 400, Product is manufactured and distributed by Penumbra Inc., Alameda, CA |  |  | March 04, 2011 | 2011 | Penumbra Inc. | The diameter of a component of the device is out of tolerance specification; the Delivery Pusher pull wire which secures the coil in place can prolapse and allow premature detachment of the coil. | Process change control | Penumbra sent an Urgent Voluntary Field Removal Notice dated March 3, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to immediately inspect all stock of Penumbra Coil 400 and identify product from the Lots.Remove product from these Lots from inventory and place it in quarantine for return to Penumbra Inc. (Do not use product from these Lots).Complete and return the attached product identification/ return form with in 5 business days via fax to +1 510-217-6414.Penumbra employees will contact customer directly to arrange return and replacement of product.For questions call Regulatory Affairs at 510-748-3224 or e-mail RAQA@prenumbrainc.com. |  | 2082 | 2082 | Worldwide Distribution - USA (nationwide) and the countries in the European Union (Denmark, Germany, Italy, Netherlands, Slovakia, Slovenia, Spain, United Kingdom and Switzerland) and South Africa. | Terminated | June 06, 2011 | 94 | 2082 | Not_Computer
Z-1970-2011 | 53195 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LANTIS oncology Information System, an accessory to the Linear Accelerator, as part of the COHERENCE RT Therapist, but also cleared under the ONCOR Avante-Garde and Expression, as well as the syngo RT Therapist system, as part of the ARTISTE And PRIMUS system. Manufactured by IMPAC Medical Inc., Sunnyvale, CA, device distributed by Siemens Medical Solutions Oncology Care System, Concord, CA |  |  | August 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc | When editing the note type of a previously entered note, then reverting to the original type, data may be lost. | Software design | Siemens sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to configure their system to lock all new notes automatically.To automatically lock new notes, open eChart Navigator, right mouse click in the notes section, select "configuration" and select "Automatically Lock New Notes."If the customer chooses not to configure their system to automatically lock new notes, follow at least one of the suggestions listed below:If the customer must edit a note type, use the append feature and document that they are making an edit to the note.Lock all new notes that are entered via eChart Notes. To lock a note manually, after creating the note, right mouse click and select "Lock,"Access and enter patient notes via eChart Notes until a LANTIS upgrade has been performed at your site.When entering patient notes via eChart Navigator, make sure that you are entering the notes in the correct section.\When entering patient notes via eChart Notes, make sure that the correct note type has been selected.Customers are to keep the Customer Safety Advisory Notice together with their LANTIS 8.30 User Manual where it should remain until the measures defined above are performed. For questions call 1-888-826-9702. |  | 30 | 30 affected units | Worldwide Distribution - USA including DE, WI, TX, TN, NJ, PA, NE, MA, and AL and the countries of South Africa, Australia, Ireland, Germany, Brazil, South Korea, India, and Japan. | Terminated | April 13, 2011 | 600 | 30 | Not_Computer
Z-1969-2011 | 56807 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio Rad brand D-10 Hemoglobin A1c Program Reorder Pack, which contains Update Kit D-10 A1c Program Diskette; |  |  | September 08, 2010 | 2010 | Bio-Rad Laboratories, Inc., Hercules, CA | The firm discovered that the Update Kit Floppy Diskette included in the reorder pack was mislabeled with an incorrect Lot Number for the Calibrator/Diluent Set as: S01819; The correct Lot Number for the Calibrator/Diluent Set is: S01704. | Other | The Bio-Rad Laboratories, Inc, sent a Medical Device Correction letter dated 9/8/10, to all affected customers via Federal Express addressing this mislabeling issue, a new Update Kit Floppy Diskette, Lot No. BA02210 has been produced, and Bio-Rad requests that you please destroy the Lot No. AA01815. If you have any questions please contact your regional Bio-Rad office. You may also call (510) 724-7000. |  |  |  | Nationwide Distribution including : PA, OH, MO, CA, TX, FL, MA, MD, KY, OH, MI, NJ, and WA. | Terminated | June 29, 2011 | 294 |  | Not_Computer
Z-1976-2011 | 57714 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | SoftPerm (synergicon A) Daily Wear Contact Lenses, Rx Only, Sterile. Contents: One sterile contact lens, rigid center/soft hydrophilic skirt (25% water), in normal saline solution buffered with sodium bicarbonate. Ciba Vision Group, Duluth GA 30097 USA |  |  | February 01, 2011 | 2011 | CIBA Vision Puerto Rico, Inc. | Out of specification values were obtained for sterility testing. The pH of the saline solution may be higher than the spec (up to pH 8.6 compared to a spec limit of pH 7.8). The osmolarity of the saline solution may be lower than specification (263 mOsm/kg compared to a spec limit of 270) | Other | The firm, CIBA Vision, sent two letters of notification to its customers. First letter "IMPORTANT PRODUCT INFORMATION" dated June 2010 and the second "URGENT: Voluntary Product Recall" dated February 1, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: 1) complete the requested information on the attached SoftPerm (synergicon A) Lenses Reply Form and drop the form in the mail within 10 business days and 2) If they have any of the recalled SoftPerm (synergicon A) lenses (shipped to you after March 1, 2009) still in their possession, return as follows: Package the product and affix the prepaid UPS return label and ship within 10 business days or mail to CIBA VISION Corporation, c/o Stericycle Pharmaceutical Services, 2670 Executive Drive, Ste. A, Indianapolis, IN 46241, Atth: SoftPerm Recall. CIBA ask customers to please not include other products, lots, kits, packs, or competitor's products with their shipment. In addition, CIBA will issue a credit to the customers account upon receipt of the returned product. Please feel free to contact Stericycle at 877-542-5928 or CIBA VISION Technical Consultation Department at 1-800-241-7468 should you have any questions or concerns. |  | 28124 | 28, 124 | Worldwide distribution: USA (nationwide) including PR; and countries including: Australia, Austria, Argentina, Bahrain, Belgium, Canada, Colombia, Czech Republic, Denmark, Dubai (United Arab Emirates), Finland, France, Great Britain, Greece, Hungary, Iceland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Netherlands, Oman, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, and Venezuela. | Terminated | November 30, 2011 | 302 | 28124 | Not_Computer
Z-1968-2011 | 58054 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Lube Jelly, 1 oz. water-base sterile lubricating jelly, a component of various Presource Custom Sterile Convenience Packs/Modules; Packaged by Cardinal Health Medical Products and Services, McGaw Park, IL 60085 U.S.A.; The affected lots of modules actually contain a 1 oz. tube of Non-Sterile Kendall Vaseline White Petroleum Jelly. |  |  | February 28, 2011 | 2011 | Cardinal Health | The Presource kits were assembled with non-sterile petroleum base lubricant instead of the labeled sterile water base lubricating jelly. | Employee error | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated February 28, 2011 to all affected consignees via UPS next day air, informing them that their Presource kit/modules listed in the attached spreadsheet contain a non-sterile petroleum-based lubricating jelly instead of the labeled water-based sterile lubricating jelly. Cardinal instructed the accounts to quarantine the pack/modules with the indicated work order/lot number on the enclosed spreadsheet and contact the appropriate Customer Service group listed in the letter for return instructions. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules being returned. If the product was further distributed, then those customers should also be notified. |  | 4236 | 4,236 modules | Nationwide Distribution | Terminated | April 19, 2012 | 416 | 4236 | Not_Computer
Z-1973-2011 | 58035 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Rapid Response H.U.T. (Head Up Tilt) Table. Model 1017. Powered electrophysiology tilt table. Table can tilt -15 degrees to 90 degrees (Trendelenburg). Medical Positioning Incorporated.800-593-3246. |  |  | February 01, 2010 | 2010 | Medical Positioning Inc | The tilt tables work intermittently and some have become stuck in an inclined or reclined position during procedures. | Other | An undated Urgent - Voluntary Medical Device Product Correction letter was issued on 2/1/2010 by certified mail with return receipt requested. The letter identified the model and serial numbers of the affected product. The consignees were advised to stop using the table immediately and remove it from service until further notice. Customers were to contact Mr. Joe Hurtig at 1-800-593-3246 or jhurtig@medicalpositioning.com to confirm receipt of the notice and to provide the firm with the correct contact information. Individualized correction packages including replacement parts and instructions were sent to each consignee. The instructions consisted of (1) Rapid Response HUT Motor Replacement Instructions; (2) Control Box Replacement Instructions; and (3) ILD External Power Supply Removal Instructions. |  | 58 | 58 total | Worldwide Distribution -- USA, including Puerto Rico, and Canada. | Terminated | June 15, 2011 | 499 | 58 | Not_Computer
Z-1972-2011 | 58022 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC Clinical Systems with Serial Number less than 3412 Including: UniCel DxC 600 Part #: A10405, |  |  | November 16, 2010 | 2010 | Beckman Coulter Inc. | The hose clamps in the DxC (with serial numbers less than 3412, labeled behind the lower right side instrument door) hydropneumatic system are sharp and pose a possible risk of injury. An operator could be cut by the hose clamps, which could in turn result in increased exposure to any waste products (biohazards) that might be present on the assembly. | Pending | The recall communication was initiated with Beckman Coulter forwarding a Product Corrective Action Letter (PCA) dated November 16, 2010 to the affected customers. The recall notice explained the reason for recall and asked that customers take the following actions:(1) Customers with a service contract will have hose clamp covers installed by their Field Service Engineer during the next scheduled Preventive Maintenance (PM). A kit can be ordered through their Customer Service Representative if consignees prefer not to wait for the next Preventive Maintenance visit and(2) Customers without a service contract can obtain a kit by contacting their local Customer Service Representative.Customers were also informed that Kit (PN A68986) containing hose clamp covers and installation instructions is available at no charge from their Customer Service Representatives. One kit is required for each DxC system.The following telephone numbers were given for customers to obtain the kits: (800) 526-3821 (United States) and (800) 463-7828 (Canada).Consignees were asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If any of the affected product was forwarded on to another laboratory they were asked to provide a copy of this letter to them.Questions are to be directed to Beckman Coulter's Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. |  | 1535 | 1,535 units | Worldwide Distribution--United States and countries of Angola, Australia, Belgium, Bulgaria, China, Columbia, Czech Republic, Denmark, Eritrea, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Mexico, Netherlands, Reunion, Serbia, South Korea, Spain, Slovakia, Swaziland, Sweden, Switzerland, Tunisia, Turkey, Taiwan, United Arab Emirates, and the United Kingdom. | Terminated | June 28, 2012 | 590 | 1535 | Not_Computer
Z-1977-2011 | 58016 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | There is a possibility that a specimen could be diluted if a user requests Diluent Dispense while specimens are being automatically processed in cassette presentation. The dilution may cause erroneous results.IMPACT:Erroneous results from a diluted specimen could impact clinical decision making in cases such as blood transfusion, chemotherapy or antibiotic administration- particularly for popu | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with PCA Response Form (via US Postal Service for US customers) dated October 20, 2010, to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter describes the product, problem and the following actions to be taken. The customers were instructed to remove all cassettes from the mix station prior to initiating the Diluent Dispense function to prevent the likelihood of a specimen being diluted; do not use pre-diluted samples for RETIC counts; verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and the Body Fluid results are properly validated at the LIS. This will be corrected in a future version; use the lip on the reagent drawer as shown in the picture above; avoid touching the drawer slides while replacing containers; clean all spills promptly to reduce the risk of exposure to blood borne pathogens; complete and return enclosed PRODUCT CORRECTIVE ACTION RESPONSE FORM via fax to 786-639-7500/7501/7502/7503/7503/7504; mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami, FL 33116-9015 or email to: Regulatory.Notifications@beckmancoulter.com, and share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If you have any questions regarding this PCA, contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including: Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1413 | 422 | Not_Computer
Z-1981-2011 | 57776 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | SmartBlock PainPump, Part numbers: P49524-S, P49534-S, P49624-S, and P49634-S. |  |  | March 17, 2010 | 2010 | Hsmg, Inc. D/b/a Smartinfuser Usa | Elastomeric tube in product's compression unit has the potential to burst following filling. | Other | Firm notified consignees by e-mail and telephone on 3/17/2010. The correspondence identified the affected product and lot numbers; and asked that customers specify which parts, lot numbers and in what quantities they still had in their possession. |  | 368 | 368 units | Nationwide Distribution | Terminated | April 25, 2011 | 404 | 1851 | Not_Computer
Z-1990-2011 | 58225 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Posterior Referencing Instrumentation for the NEXGENÂż COMPLETE KNEE SOLUTION MODULAR T-HANDLE, REF 00-5901-061-00, NONSTERILE, Zimmer UK, Ltd.,, Zimmer Warsaw,IN |  |  | March 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potential for breakdown of the AlTiN PVD (black) coating during surgery and the coating could fall into the surgical site. The firm has received 22 reports of the black coating coming off the instruments after cleaning and sterilization cycles and 5 resulted in a delay of surgery. | Device Design | The firm, Zimmer, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 14, 2011 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:Instructions for Handling of Affected Product:1) Stop using the devices and to quarantine immediately2) Carry out a physical count of all affect product in their territory and record data on Inventory Identification Form included with this letter-Note: The instruments must be returned even if the instrument is not exhibiting the black powder3)Ensure they have provided acknowledgement via "email return receipt" upon receiving this notification4) Fax copy of completed form to (574) 372-42655) Return the recalled product along with the completed Inventory Identification Form and the Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580A new or refurbished instrument will be provided as replacement for their returned product once the inventory is available.Instructions for Health Care Professional Notification1) Provide a list of all hospital that currently use or have used the affected product listed in notice by utilizing the User Facility Documentation Form and the Health Care Professional Documentation Form2) Fax a copy of the completed form(s) to: Zimmer, Inc. at (574) 372-4265Zimmer will provide a similar notification via FedEx to these hospitals making them aware of this recall and provide a "Dear Doctor" letter for their surgeons.Important: If you have distributed these affected lots further, please provide a copy of this letter to these customers when you implement these recall instructions.For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc., at 1-800-613-6131.ADDENDUM INFOThe information listed above was updated to include the additional 5 |  | 552 | 552 | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Germany, Japan, New Delphi and Singapore. Addendum: Finland, Switzerland, UK, Denmark, Spain, Portugal, France, Netherlands, Italy, South Africa, Norway, Spain, and Sweden. | Terminated | March 19, 2013 | 736 | 9462 | Not_Computer
Z-1980-2011 | 58288 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Floor Cabinet Part Number: 723335 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | Sharp edges on drawer slides within the DxH800 Floor Stand (Cabinet) may expose operators to physical injury while replacing reagents and waste containers.IMPACT:There is potential for exposure to blood borne pathogens such as HIV, HBV, and HCV. | Other | Beckman Coulter sent an Urgent Product Corrective Action letter dated October 20, 2010, with a PCA Response Form (via US PostalService) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. This letter informed the customers about the problems identified and provided an action/resolution to this issue. Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative |  | 422 | 422 units in total (175 in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | April 02, 2012 | 530 | 422 | Not_Computer
Z-1983-2011 | 58048 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Sky 1000/BHM Medical V10 Bariatric Ceiling Lift System; a patient transfer lift; ArjoHuntleigh, Addison, IL 60101; Model numbers LF20010, LF20011, LF20013, LF20014, LF20015, LF21015, LF21019, 9120001, 9120004, 9120005, 9120008, 9120009 and 9121009 |  |  | March 21, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The battery-powered control board in the ceiling lift motor may overheat, resulting in smoke and possibly flames and the lift becoming inoperable. | Device Design | ArjoHuntleigh sent an Urgent Field Safety Notice dated March 4, 2011, to all affected customers via UPS Ground on March 21, 2011. The notice informed the accounts of the potential for smoke and flames coming out of the Maxi Sky 1000/V10 bariatric ceiling lifts. The accounts were requested to immediately cease use of the affected lifts, powering them down by activating the emergency stop, and tag the lifts to prevent further use. If there is a clinical need to continue use of the lift before it is retrofitted or replaced, specific instructions were given to reduce the risk of a device malfunction. The accounts were requested to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 1246 | 1,246 units | Nationwide Distribution | Terminated | February 16, 2012 | 332 | 1246 | Hardware | Device Operation | Replace device | Remove or Replace
Z-1984-2011 | 55910 | No | GKA | Abnormal hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT nbs Sickle Program Resin Update CD-ROM; Model Number:250-3020, included in 250-3000; Distributed and Manufactured by Bio-Rad Laboratories, Inc. The VARIANT nbs Sickle Cell Program is intended as a qualitative screen for the presence of hemoglobins F, A, S, D, C and E in eluates of neonatal blood collected on filter paper by high-performance liquid chromatography(HPLC). The VARIANT nbs Sickle Cell Program is intended for Professional use only. |  |  | April 14, 2010 | 2010 | Bio-Rad Laboratories Inc | The affected resin Update CD-ROMs cause all user-defined settings in the Setup/Test/Data Setup sub screen and Set/Test/Pattern Setting sub screen to be overwritten when the Update Kit procedure is performed. | Software design (manufacturing process) | Bio-Rad Laboratories sent out the Urgent Medical Device Correction notice on 4/14/2011, by Federal Express. An Update was sent on May 13th, advising customers of a minor discrepancy in the parameters included on the affected Resin Update CD-ROMs. Customers were instructed to examine your inventory for the affected reorder pack lots. Please complete and return the Medical Device Correction Response FormQuestions regarding this recall please call (510) 724-6114. |  | 658 | 658 units | Worldwide Distribution -- USA , Brazil, Canada, Caribbean, France, Italy, and United Kingdom. | Terminated | August 08, 2011 | 481 | 658 | Other |  |  | Other
Z-1979-2011 | 58173 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER AC-T 5diff Autoloader Hematology Analyzer, Part Number: 175356, Serial |  |  | February 23, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter identified a problem with occasional unexpectedly high results for RBC, HCT, PLT and MPV without instrument generated messages, for the first run of a capped patient sample. Upon repeat analysis the results are correct. Patients could be affected as described below: (1) False high RBC and HCT may result in a delay of treatment,or the delivery of | Pending | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated February 23, 2011 with an enclosed customer response form to all customers who have the affected instruments via US Postal Service (in US). The letter described the product, problem and actions to be taken. The customers were instructed not to analyze specimens using the Autoloader. Instead, analyze their samples uncapped (open vial) using the Manual (Stat) Mode. The end of this notice includes a Quick Reference for RunningManual Samples. For more detailed instructions, please refer to your instrument's Instructions for Use (IFU), part number 624026, Section 8.4 Running Manual (Stat) Samples. An electronic version of the IFU can be found in the instrument's On-Line Help screen as well as the Beckman Coulter web site. For inconsistent data during patient's follow-up, please consider unusual/unanticipated high or elevated platelet levels as possible falsely elevated results due to this failure. An appropriate action for the need to look-back at previous results is at the discretion of the Laboratory Director, given the fact that laboratories have a variety of measures that are employed to ensure the validity of the results. Some of these provisions include delta checks, laboratory protocols for validating results and interaction with physicians.Additionally, the customers were also requested to complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM within ten (10) days via fax to 786-639-7500/7501/7502/7503/7504; mail to Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami, FL 33116-9015, or email to: Regulatory.notifications@beckmancoulter.com.For any questions concerning this notice, contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact your local Beckman Coulter representative. |  | 1211 | 1,211 units | Worldwide distribution: USA (nationwide) and countries of: Angola Mexico Australia, Azerbaijan, Bangladesh, Bermuda, Botswana, Bulgaria, China, Cote d'Ivoire, Finland, France, Germany, Ghana, Hong Kong, India, Italy, Mexico, Netherlands, New Zealand, Nigeria, Palestinian, Puerto Rico, Romania, Russian Federation,Saudi Arabia, Slovenia, South Africa, Spain, Switzerland,Taiwan, Tanzania, and Turkey. | Terminated | July 16, 2012 | 509 | 1211 | Not_Computer
Z-2000-2011 | 51504 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS 6.1 Oncology Information System with MEDOne Module, Linear Medical Accelerators; Product Numbers: 4503178, 5835991, 5851055; Product is manufactured by IMPAC Medical, Inc., Sunnyvale, CA. |  |  | March 06, 2009 | 2009 | Siemens Medical Solutions USA, Inc | The product has a software problem; it does not calculate the Creatinine Clearance correctly for patients less than two years old. | Software change control | On 3/6/2009, a Safety Advisory Letter was distributed to all affected users, with a description of the problem and identifying the affected product. Customers were instructed not to use the LANTIS to calculate Creatine Clearance for any patient under two years of age. Customers should direct their questions with regards to the letter to their local Siemens Service organization. A Siemens Representative will update the systems. |  | 17 | 17 | Worldwide Distribution | Terminated | April 15, 2011 | 770 | 17 | Software | Output/Calculation | Software update | Software Update
Z-1998-2011 | 56636 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Cyclosporine Sample Pretreatment Reagent, COBAS INTEGRA, IVD, Roche Diagnostics GmbH, Assembled for and Distributed by: Roche Diagnostics Corporation, Indianapolis, IN. |  |  | August 13, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The cyclosporine sample pretreatment reagent lot was assigned an incorrect expiration date. | Other | Roche Diagnostics Corporation sent an Urgent Medical Device Removal letter dated 8/13/2011, by UPS ( signature required ) to the consignees. The letter identified the product, the problem, and the action to be taken by the consignee.Consignees were instructed to discontinue use of affected product upon receipt of replacement product. Consignees were instructed to dispose of the product at their facilities. Therefore, no product returns are anticipated. Questions concerning this recall please call (317) 521-3911. |  | 22 | 22 kits | Nationwide Distribution -- including, IN, NJ, NY, TX, CA, PA, OH, and FL. | Terminated | May 04, 2011 | 264 | 22 | Not_Computer
Z-1999-2011 | 58202 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar SP Infusion Suite Database, Version 1.0; an Rx PC software application; Hospira, Inc., Lake Forest, IL 60045; list 13092-04-01 |  |  | March 23, 2011 | 2011 | Hospira Inc. | The software enables programming of the GemStar device for a Bolus Lower Lockout Limit outside of the hard limits entered into the GemStar SP Infusion Suite Database PC application. | Software design | The firm, Hospira, sent "URGENT DEVICE FIELD CORRECTION" letters dated March 23, 2011 to their customers on the same date. The letter described the product, problem and preventative action to be taken by the customers until the software can be upgraded. The customers were requested to contact Hospira Global Product Safety and Complaints at 1-800-441-4100 to report adverse events, and Hospira Advanced Knowledge Center at 1-800-241-4002 for technical assistance. The accounts were also requested to complete the attached Reply Form and return it to Hospira via fax at 1-262-577-6921.For additional information and technical assistance, contact Hospira Advanced knowledge Center at 1-800-241-4002, option 4 (available 24 hours a day/7 days per week). |  | 139 | 139 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Chile, Colombia and Singapore | Open, Classified |  |  | 139 | Software | Device Operation | Software update | Software Update
Z-1874-2011 | 58134 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i, DFP-8000D |  |  | December 21, 2010 | 2010 | Toshiba American Medical Systems Inc | Under some conditions during high doses procedures the cumulative air kerma display may reach its maximum capability and fail to display the accrued value. | Radiation Control for Health and Safety Act | An Important: Electronic Product Radiation Warning notification was sent to all consignees. The letters described the reason for the recall and identified the affected product. The system software will be modified to allow the shift of the decimal to one position to the right allowing for a display of 99999.9 mGy. This will allow for the retention of the resolution until such time as the system exceeds the 9999.99 mGy range. A Toshiba Service person will contact customers when the software is available. Installed systems with the current version of software ( version 4.23) will begin to receive the corrected software April 2011. Previous versions of the software will begin to receive corrected software June 2011. The software will be provided free of charge. Questions related to the letter should be directed to the Regulatory Affairs Department at 800-421-1968. |  | 220 | 220 units Nationwide | Nationwide Distribution | Terminated | March 18, 2013 | 818 | 220 | Computer
Z-2002-2011 | 58125 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sarns Flexible Weighted Sucker in K9 Pack (custom pack), catalog number 63635-01, Sterile and Non-pyrogenic. Sterilized by ethylene oxide. Terumo Cardiovascular Systems, Ashland, MA . |  |  | February 25, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received two reports of the Terumo Weighted Flexible Sucker becoming ensnared in cardiac tissue after it was inserted into a ventricle through a valve. The firm has determined that the products's instructions for use are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However this is not stated in the IFU. The firm is conducting | Device Design | SEE UPDATE BELOW IN ALL CAPS.The firm, Terumo Cardiovascular Systems (Terumo CVS), sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were informed that the product's Indications For Use (IFU) are inadequate; the product should not be used to drain a cardiac chamber or inserted through a valve. However, this is not stated in the IFU. In addition, the customers were instructed to read and understand the changes and clarifications to the product's IFU; review the Medical Device Correction notice; assure that all users are aware of this notice, and confirm receipt of this communication by faxing the attached Customer Response Form via fax to: Terumo CVS Customer Service at 800-292-6551 (US only).If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818. Customers Service hours: Monday-Friday, 8AM-6PM.ON 8/30/2011 THE FIRM ISSUE A LETTER ENTITLED "URGENT MEDICAL DEVICE RECALL" THAT STATED "AFTER FURTHER CONSIDERATION, TERUMO CVS DECIDED TO DISCONTINUE SALES OF THE WEIGHTED FLEXIBLE SUCKER AND IS ASKING USERS TO RETURN ALL UNUSED PRODUCT" TO THE FIRM. |  | 11690 | 11,690 | Worldwide distribution: USA (Nationwide) and countries including: BRAZIL, BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES, and MALAYSIA. | Terminated | December 10, 2012 | 654 | 11690 | Not_Computer
Z-2022-2011 | 58041 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CHEM 8+ Cartridge, List Number 03P91-25; 510 K053110 |  |  | February 15, 2011 | 2011 | Abbott Point Of Care Inc. | i-STAT CHEM 8+ cartridge may generate false results for sodium, chloride and hematocrit. | Other | The firm, Abbott Point of Care (APOC), sent an "Urgent Recall Notice" letter dated February 2011 with business reply cards attached via Fed Ex overnight to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of cartridges from the box numbers of identified lots; discontinue use as a precaution, if their cartridges are no longer traceable to a specific box; return all unused cartridges as instructed on the enclosed Business Reply Card; if they've provided another institution with the cartridges, provide the institution with a copy of this letter, and complete and return the business reply card to APOC.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-336-8020, Option 1 or your Abbott Point of Care representative. |  | 6575 | 6575 cartirdges | Nationwide distribution. | Terminated | July 10, 2012 | 511 | 6575 | Not_Computer
Z-2013-2011 | 58148 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Lubricating Jelly packaged as follows: |  |  | February 07, 2011 | 2011 | DeRoyal Processing Center | Possible contamination due to sterilization issue. Inspection of contract manufacturer found cGMP deficiencies. | Process control | The firm notified their consignees by a Re: DeRoyal Recall Of Triad Group Lubricating Jelly letter, dated Friday, March 18,2011. The letter discussed the reason for recall and the affected products. The letter asked customers to identify inventory using the attached spreadsheet and quarantine to prevent further use. Customers are to complete the spreadsheet regardless if they have effected product or not. The completed form can be faxed to 865-362-3716 or emailed to recalls@deroyal.com. All affected products should be return to DeRoyal at the address provided. Distributors are to forward a copy of the notice to their end-users. If there are questions or assistance is needed with the recall, Stephanie DeBusk should be contacted at 865-362-1037 or Sharon Cook at 865-362-6157. |  | 323440 | 323,440 units | Nationwide Distribution | Terminated | August 05, 2014 | 1275 | 323440 | Not_Computer
Z-2021-2011 | 58080 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Avanto, Espree, Trio A Tim System. |  |  | March 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. The potential exists for a cable connection to become overheated, leading to the emission of smoke and possible material damage. | Other | Siemens Medical Solutions Inc., issued a Customer Safety Advisory Notice dated November 24, 2010, to affected customers via Update Instructions MR029/10S. Affected customers will be notified via hand delivered letter by Siemens Service Engineers at the time of inspection. A future Update instruction will be released for installation of detectors in appropriate locations on systems to switch off the power electronics in the event of a malfunction in the power supply of the gradient coils.Further questions regarding this recall please call (610) 448-3237. |  | 1419 | 1419 | Product was distributed domestically to hospitals and imaging centers in the District of Columbia, Puerto Rico, AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY(every State except VT). | Terminated | December 09, 2014 | 1377 | 1419 | Not_Computer
Z-2023-2011 | 55577 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Siemens brand 550 TxT Patient Treatment Table, part number 7346534; Product is manufactured and distributed by Siemens AG, Medical Solutions, Roentgenstrasse 19-21, Kemnath, Germany 95478. |  |  | March 31, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Potential for causing overheating of the power supply during treatment, and potential for unexpected table descent. | Component design/selection | Siemens issued a Customer Information letter dated March 21, 2010. There were two safety issues. There were two occurrences of smoke caused by an overheating of the power supply during patient treatment. Also, there had been a report of an unexpected lowering of the 550TxT table in patient rescue mode. With the released kit that accompanies this letter, Siemens has implemented a minor modification on the 550TxT to include an additional fuse within the power supply circuit to address the first issue described above. They also included an additional modification to address the lowering of the table by the integration of additional relays. As soon as the "move enable" button is released, the table brake will be enabled which is controlled by the implemented relays. No further action is required for the continued operations of the 550TxT table.Further questions regarding this recall please call (925) 246-8407 |  | 326 | 326 units | Worldwide Distribution -- Including USA, states of NJ, WI, NY, and MA. | Terminated | March 01, 2012 | 701 | 326 | Battery | Physical Safety Hazards | Repair | Repair
Z-2025-2011 | 58276 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hospira LifeShield Latex-Free HEMA Y-Type Blood Set, Nonvented, 100 Inch with Clave and Secure Lock; a sterile Rx tubing set for the alternate administration of fluids from a blood bag and I.V. solution container; Made in Costa Rica; Hospira, Inc., Lake Forest, IL 60045 USA; List No. 12450-48 |  |  | March 28, 2011 | 2011 | Hospira Inc. | There is a potential for the air filter at the distal end of the set to become detached. This could result in the contamination of the sterile fluid path inside the tubing. | Other | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated March 28, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return the Reply Form via fax to 1-866-324-3734, even if you do not have the affected product; inform healthcare professionals in their organization of this recall; if they have distributed the product further, notify their accounts and have them fax the reply form to 1-866-324-3724; return the affected product to Stericycle using labels provided, and call Stericycle at 1-877-877-0481, if they have not received a return label. For medical inquiries, please call Hospira Medical Communications at 1-800-615-0187. Call Hospira Global Product Safety and Complaints at 1-800-441-4100 (8am-5pm CST, M-F) or by e-mail at (ProductComplaintsPP@hospira.com). |  | 47472 | 47,472 sets | Worldwide distribution: USA (nationwide) and countries including: Canada, Korea and Singapore. | Terminated | May 09, 2012 | 408 | 147646 | Not_Computer
Z-2047-2011 | 55154 | No | CFT | Prolactin (lactogen) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Active Prolactin ELISA, 96 Wells, DSL-10-4500 |  |  | August 03, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Prolactin ElLISA kit, lot number 890291 over estimates the expected dose values for patient serum samples due to a shift in the assigned calibrator values.Beckman Coulter recommends that patient results generated with the affected lot be evaluated in the context of other diagnostic tests and the clinical presentatio | Equipment maintenance | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION" Letter dated August 03, 2009 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:Action/Resolution:Discontinue use and discard remaining product with lot number 890291 of Active Prolactin reagents. Contact Beckman Coulter at 1-800-231-7970 for replacement product.Newer lots (lot number 991210 or higher) of the product have been realibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, they were told to contact Technical support at the phone number listed below.Consignees were asked to share this information with their laboratory staff and retain this letter for their Quality System documentation and asked to complete and return the enclosed response form within 10 days via fax to: 786-639-4000 so that Beckman can be assured that they received the notification.If you have any other questions regarding this notification, please contact DSL Technical Support Center at 1-800-231-7970 in the United States and Canada , or contact their local Beckman Coulter Representative. |  | 30 | 30 Kits | Nationwide distribution: USA including states of: FL, GA and NC. | Terminated | April 27, 2012 | 998 | 30 | Not_Computer
Z-2045-2011 | 57750 | [6 empty] | IMMAGEÂż Immunochemistry Systems Wash Solution Part Number: 447060 |  |  | August 24, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an ImmageÂż Systems Wash Solution Lot 71744 was manufactured by a Beckman European facility and not intended to be used in the U.S. However, this lot was inadvertently shipped to the U.S. and delivered to five customers. | Other | The firm, Beckman Coulter, contacted the customers who received the IMMAGE Immunochemistry Systems Wash Solution, lot #71744 by telephone. Beckman sent a follow-up "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 25, 2010, to all affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of ImmageÂż Immunochemistry Systems Wash Solution lot 71744; return the material to Beckman Coulter's Chino warehouse: Beckman Coulter Inc., 15989 Cypress Ave, Chino, CA 91710, Phone: 909-597-3967;package the material the way it arrived and use Beckmans internal Fedex account to ship the material; complete and return an enclosed FAX BACK RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulator Affairs, 250 S. Kraemer Blvd, Brea, California 92821, Mail code E2.Se.08; share this information with laboratory staff; retain the notification as part of their Quality System documentation, and provide a copy of the letter if they forwarded the affected products.If you have any questions regarding this letter, please call Beckmans Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada. |  | 9 | 9 units total in US | Nationwide distribution: USA including states of : AL, CA, MI and IN. | Terminated | June 26, 2012 | 672 | 9 | Not_Computer
Z-2044-2011 | 58174 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTERÂż HmX Hematology Analyzer |  |  | March 02, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a problem with Integrated Circuit Chips used in different boards throughout the Beckman Coulter LH500/HmX/HmX Autoloader Hematology analyzer. | Other | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated March 2, 2011, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the COULTERÂż LH 500 Series Analyzer, COULTERÂż HmX Hematology Analyzer, and COULTERÂż HmX Hematology Analyzer with Autoloader. The letter described the product, problem and actions to be taken. Customers were informed that the suspect Integrated Circuit Chips have been removed from all systems in inventory and manufacturing.The customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality Systemdocumentation. If customers have forwarded any of the affected product(s) listed on the PCA letter to another laboratory, please provide a copy of the letterto them.The customers were instructed to complete and return the enclosed Response Form within 10 days so Beckman are assured they have received this important communication and are taking appropriate action.The customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States and Canada or contact their local Beckman Coulter Representative. |  | 288 | 288 units total (158 in US) | Worldwide distribution: USA (nationwide) and countries of: Canada, France, Germany, Italy, Spain, United Kingdom, and Singapore. | Terminated | May 29, 2012 | 454 | 288 | Not_Computer
Z-2028-2011 | 57340 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman Disposable Twist Drill |  |  | November 30, 2010 | 2010 | Codman &amp; Shurtleff, Inc. | Packaging defect may compromise sterility of device | Package design/selection | The firm, Codman, sent an "Important Notice Urgent- Voluntary Product Recall Notification" letter dated November 30, 2010 to its customers via FedX. On January 7, 2011, Codman expanded the recall to four additional product codes.On January 31, 2011, the firm issued a supplemental recall letter and identified one additional product code: 80-11197. The letters described the product, problem and actions to be taken. The customers were instructed to identify the unused affected products in their inventory and follow the recall instructions provided: 1) If you have NO product to return, please advise Codman as soon as possible using the FACSIMILE form attached and fax to 508-828-6118 and 2) if you DO have product to return, please call 1-800-225-0460, Option 4, for instructions for return.Codman is presently working toward supplying replacement products as soon as possible. Please call Customer Service to check on the status of product availability at 1-800-225-0460.Should you have any questions regarding this recall, please call Customer Service at 1-800-225-0460 or your local Codman Representative. |  | 1549 | 1549 units | Worldwide distribution: USA including states of :AL, FL,GA, KY, MA, MI, MN, MS, OR, PA, RI, UT,WI and WV; and countries including: Argentina, Australia, Canada, Bahamas , Belgium, Dubai, Japan, Madrid, Mexico, New Zealand, Saudia Arabia, UAE, Uruguay, and United Kingdom. | Terminated | March 14, 2013 | 835 | 14792 | Not_Computer
Z-2035-2011 | 57933 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ***REF MCA5-15SB ***1.5 mm Fluted Ball, 5.6 cm. Use with Micro Curved Attachment (MCA).***ANSPACH***Sterile, Sterilized using Irradiation***Rx Only*** Made in the USA***Manufacturer: The Anspach Effort, Inc. 4500 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 327-6887/ +1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP11 1LA United Kingdom Tel: +441494 616126*** |  |  | September 29, 2010 | 2010 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL is recalling their Sterile Bone Cutting Burs that was manufactured with incomplete flutes which could cause degraded cutting performance and in some cases not cut at all. Anspach has not received any complaints related to this problem. | Other | The firm, Anspach, sent an "URGENT: Medical Device Product Removal" letter dated October 6, 2010, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Screen their inventory immediately for any of the part numbers listed in the letter. Do not include any product received after October 7, 2010.2) If they have any of the products listed, contact Anspach Customer Support at (800) 327-6887 to arrange for their immediate return and replacement.3) Complete and return the attached Customer Reply Form indicating how many of each part number is being returned and confirming their receipt of this letter, and fax to 1-800-327-6661. If they have distributed any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitate to contact Anspach Customer Support at (800) 327-6887. |  | 271 | 271 cutting burs affected) | Worldwide distribution: USA including states of: AZ, CA, FL, GA, IL, IN, LA, MA, and MI; and countries of: Australia, Belgium, England, France, Italy, Japan, South Africa, Spain and Sweden. | Terminated | August 12, 2011 | 317 | 271 | Not_Computer
Z-2040-2011 | 57888 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Orthoralix 8500 DDE Panoramic Dental X-Ray System; Extraoral x-ray unit with timer; Gendex Dental Systems, Des Plaines, IL 60018; part numbers 110-0199G2, 110-0199G4, 110-0234G1 and 110-0234G3. |  |  | February 04, 2011 | 2011 | Gendex Dental Systems | There is an anomaly in the Orthoralix firmware which can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obviou | Other | Gendex sent a recall letter to their accounts on February 4, 2011 via first class mail, stating that an anomaly in the Orthoralix firmware can cause a message corruption which will distort the patient image being transferred from the X-Ray system to the PC. This image distortion/corruption (distorted, fragmented, flipped, negative, reversed or had lines across them) is most likely to occur during periods of high network traffic and is significant enough that it is clear and obvious these images cannot be used for diagnostic purposes. The letter stated that resending the image from the X-Ray system to the PC will correct the issue present in certain images sent during periods of high network traffic. Instructions were provided to download an Appendix to Orthoralix DDE User Manual that explains how to resend the images. If that does not resolve the issue, users were requested to contact Gendex Technical Support at 1-800-323-8029. Customers were asked to complete and return the enclosed Acknowledgement Form by fax to 215-997-5665 or e-mail: regulatoryaffairs@gendex.com. For questions regarding this recall call 800-323-8029. |  | 6835 | Total 6,835 units of all products | Worldwide Distribution--United States, including Puerto Rico, Australia, Canada, Chile, Colombia, Germany, Hong Kong, India, Italy, Korea, Mexico, New Zealand, Peru, Singapore, Taiwan and Vietnam. | Terminated | February 15, 2012 | 376 | 6835 | Software | Display/Image | Instructions to resend images from system to PC | Safety Notice/Insructions
Z-2054-2011 | 58324 | No | MRA |  |  | Orthopedic | PMA | R3 TM CERAMIC ACETABULAR LINER, BIOLOXÂż FORTE, 32 MM O.D. 55 MM O.D., REF 71338950, STERILE R, QTY (1), Smith &amp; Nephew, Inc., Memphis, TN 38116, 2017/06 |  |  | March 11, 2011 | 2011 | Smith &amp; Nephew Inc | During the manufacturing process for several batches of R3 Ceramic Liners, the titanium rings were pressed onto the ceramic component with a higher force than allowed by manufacturing specifications. This has the potential to result in lower than expected strength for the liners. | Employee error | All affected Smith &amp; Nephew Distributors were notified via email and the hospitals were notified by letter delivered by Fed-Ex on 3/11/2011. They were instructed to immediately cease distribution, quarantine affected product for return, if product further distributed, identify customers and immediately notify them of recall.Please contact Mr. David Archer, Group Director, at 701.396.2121, ext. 5373, if you have questions. |  | 325 | 325 units | AK, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OR, PA, SC, TN, TX, WA, and WI | Terminated | February 08, 2012 | 334 | 2271 | Not_Computer
Z-2051-2011 | 58231 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 ultrasound system |  |  | March 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The Firm discovered a software error on the Accuson S2000 ultrasound systems affecting the measurement results in the patient report. | Software design | Siemens issued a Customer Safety Advisory Notification letter to all customers on March 15, 2011, regarding a software error on the ACUSON S2000 ultrasound systems affected measurement . They are in the process of releasing a software update that will correct this issue. Siemens recommends reviewing the patient report at the end of each study to confirm that the report contains only measurements relevant to the exam just performed. No patient injury has been reported. Further questions regarding this recall please call (650) 694-5398. |  | 1533 | 1533 units | Worldwide Distribution including USA, states of AZ, VA, CA, NE, IL, MN, TN, FL, NM, and DC | Terminated | June 18, 2012 | 461 | 1533 | Software | Output/Calculation | Software update | Software Update
Z-2061-2011 | 53053 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Premium Edition Systems with software versions 5.0.00 and 5.0.01 (model numbers 10348531, 10348532 and 10348533), when using the V5Ms transducer (model number 08263703) and the Plus option (10427795) |  |  | July 24, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Product Failure-- When the 64-Channel System V5M is used in the 128-Channel mode, an over current fault occurs which shuts the system down. | Software design | Siemens sent a Customer Safety Advisory Notification letter on July 24, 2009. The letter identified the product, the problem, and the action to be taken by the customer. A software update will be conducted in the field. Until that update occurs, customers were instructed to image without the CW doppler to avoid a system shut down.Customers were instructed to contact their service support person with any questions.For questions regarding this recall call 650-694-5993. |  | 70 | 70 units | Worldwide Distribution - USA including CA, CT, FL, MA, and TX and the countries of United Arab Emirates, Germany, Ecuador, Spain, France, Great Britian, Saudi Arabia, Sweden, Thailand, and South Africa. | Terminated | April 25, 2011 | 640 | 70 | Software | Device Operation | Software update | Software Update
Z-2065-2011 | 53209 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson X300 systems at software versions 1.00.00 to 1.0.10 and 2.0.00 to 2.0.02, model numbers 10037409, 10132987, and 10038837, manufactured by Siemens Medical Solutions ultrasound business unit, Mountain View, CA |  |  | July 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Software Error with potential for misinterpretation of data-- The previous patient's last capture clip image can appear in the next patient's study. | Software design | The firm, Siemens, sent a "CUSTOMER SAFETY ADVISORY NOTIFICATION" letter to its customers on July 30, 2009. The letter described the product, problem and actions to be taken. The customers were instructed they need to update any of the patient registration settings to reset the cine buffer so that no data is transferred between studies, if the system software is not being updated. Siemens informed the customers that a mandatory update software is being installed on their system free of charge.If you have any questions, contact (650) 969-9112. |  | 2312 | 2312 units | Worldwide distribution: USA (nationwide) and countries including: Albania, algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile,China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Kazakhstan, Kenya, Kosovo, Lativia, Lithuania, Luxembourg, New, Zealand, Nicaragua, Norway, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Viet Nam. | Terminated | April 25, 2011 | 634 | 2312 | Software | Display/Image | Software update | Software Update
Z-2060-2011 | 58261 | [6 empty] | LuxtecÂż MLX 300 Watt Xenon Light Source, |  |  | March 07, 2011 | 2011 | Integra LifeSciences Corp. | Malfunction of the MLX Control Board may cause loss of manual control and Lamp shut down. | Process change control | Integra notified all customers/distributors via written communication, dated 3/08/2011. On 3/09/11, Integra Sales Representatives and Service Technicians were notified via telephone conference call and provided copies of written communications. Customers were advised to immediately discontinue use of any affected product and to place in a secure area. Note: This defect only applies to MLX units that were repaired between 12/01/2009 and 1/31/2011. This defect does not occur in manufactured units that have not been serviced, nor does it apply to MLX units that were serviced outside of this specific timeframe. Customers with devices already known to require repair and/or servicing will be contacted directly to arrange for the return of their device(s). Additionally, sales specialists were instructed to examine each customer's inventory for those device(s) that may potentially have the jumper-wire and to perform a stand-by test to determine if repair and/or servicing are required. |  | 159 | 159 units | Nationwide, Canada, Australia, Germany, Netherlands, New Zealand, Norway, Singapore and Sweden | Terminated | March 26, 2014 | 1115 | 159 | Hardware | Device Operation | Replace | Remove or Replace
Z-2063-2011 | 58246 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03899; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The optical guidance platform may not be properly enforcing a 24 hour time limit between optical camera recalibrations. | Software change control | The firm, Varian, sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check the Calibration timeout on their optical guidance platform. If this exceeds 24 hour, please set the factory default of 16 hours to ensure that the camera will be calibrated daily. The instructions are located in the document. The customers were also instructed to download the most recent version of the Optical Guidance Platform Instructions for Use at www.MyVarian.com or request a copy from their local Varian Customer Support District or Regional Manager. In addition, the customers were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. Varian is preparing a software patch which will be applied to affected units to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager; USA and Canada -1888-VARIAN5 (888-827-4265) or Europe- +41 41 749 8844. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 474 | 599 | Software | Device Operation | Software update | Software Update
Z-2062-2011 | 58320 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | S8-3t Micro TEE Transducer is an Ultrasonic Diagnostic. The Part # is 989605379471. |  |  | March 27, 2011 | 2011 | Philips Medical Systems Nederlands | Users of the Philips model S8-3t Micro TEE Transducer have reported that significant image quality degradation may occur after a number of use/cleaning cycles. If users do not inspect the transducer for such damage prior to each use, imaging may not always be available, which could pose a risk for patients. | Other | Philips Healthcare began mailing Urgent Medical Device Correction letter on March 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers are being instructed on how to inspect the distal tip for defects that cause image quality degradation, and to plan for epicardial / transthoracic ultrasound imaging as an adjunct to the S8-3t. Customers are also reminded to carefully follow cleaning described under Handling the S8-3t Transducer, of Philips' iE33 System User Information. Customers can call Philips Customer Service at 800-722-9377 and reference to FCO 79500198 for any questions about this recall. |  | 155 | 155 units | Worldwide Distribution - USA AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, KY, LA, MA, MD, MI, MN, MS, NC, ND, NH, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Israel, Italy, Japan, Korea Rep, Netherlands, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, Ukraine, and United Kingdom. | Terminated | February 14, 2012 | 324 | 155 | Not_Computer
Z-2074-2011 | 57779 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCelÂż DxC 680i |  |  | October 22, 2010 | 2010 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter has confirmed that for User-Defined Chemistries (UDRs), the polychrome correction is not automatically updated when wavelengths are modified. | Environmental control | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased Synchron LX and UniCelÂżÂż DxC Clinical SystemsOperating Software (all versions). The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:-After modifying any UDR wavelength, they MUST reboot to apply the update of the polychrome correction. -If the modified UDR requires calibration, recalibrate after rebooting. - Complete and return the enclosed RESPONSE FORM within 10 days via fax to (714) 961-4234 or mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs, 250 S. Kraemer Blvd, PO Box 8000, Brea, California, 92282-8000; Mail Code E2.SE.08. In addition, the customers were instructed to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If customers forwarded any of the affected product listed above to another laboratory, please provide a copy of this letter to them.If customers have any questions regarding this Product Corrective Action letter, please call Beckman's Customer Support Center (Hotline) at (800) 854-3633 in the United States and Canada or www.beckmancoulter.com/customersupport/supportrequests. |  | 4879 | 4879 units total (3490 in US) | Worldwide distribution: USA (nationwide) and countries including:Angola, Australia, Belgium, Bulgaria, China, Colombia, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Eritrea, Finland, France, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Macao, Mexico, Namibia, Netherlands, Nez Zealand, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, South Africa, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1411 | 6289 | Computer
Z-2052-2011 | 57744 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific's LATITUDE Patient Management System: |  |  | January 10, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that data within certain printed reports obtained from the LATITUDE Patient Management System may not be accurate, in that, the Most Recent Presenting Electrogram (MRP-EGM) for some of the selected patients may be either missing or incorrect. | Software design | Consignees were sent a Boston Scientific "LATITUDEÂż Medical Records Information" letter dated January 2011. The letter was addressed to "Dear Doctor". The letter described the problem and the product involved in the recall.The letter also included the Rate of Occurrence, Action and Further Information sections. |  | 25836 | 25,836 | Worldwide Distribution -- USA, including DC and PR, and countries of AUSTRALIA, CANADA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | November 15, 2012 | 675 | 25836 | Other | Output/Calculation | Notification | Safety Notice/Insructions
Z-2068-2011 | 58042 | No | GKE | Osmotic fragility test. | Hematology | Hematology | 510(K) Exempt | Advanced Instrument Model 2020 Multi-Sample Osmometer |  |  | January 19, 2011 | 2011 | Advanced Instruments, Inc. | A poorly soldered power cord may pose a potential electrical shock hazard. | Employee error | Advanced Instruments Product Service Department contacted each customer on 1/19/11 via an Urgent Product Recall Notice. The letter identified the reason for recall along with the affected products. The letter provided instructions to safely examine the date code of the AC Power Inlet assembly on the affected product. If customers were unable to perform the inspection or had questions concerning the instructions, they were to contact Advanced Instruments Product Service Manager at 781-471-2147 or via email jorges@aicompanies.com. The attached response forms were to be completed and returned as instructed. |  | 3 | 3 units | Worldwide Distribution -- USA, including states of CA, GA IL, NJ, and MA and country of Belgium. | Terminated | December 01, 2011 | 316 | 3 | Battery | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-2066-2011 | 58371 | No | HEM | Obstetric-gynecologic ultrasonic imager. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | OB-Link release 2.0.7. distributed by Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098 |  |  | May 12, 2008 | 2008 | Digisonics, Inc | Data/measurements from a OB report sent from ultrasound systems populated the wrong OB-View report | Other | Digisonics visited consignee in April 2008 as follow-up to complaint. |  | 2 | 2 units | Nationwide Distribution including OK | Terminated | May 02, 2011 | 1085 | 2 | Other |  |  | Other
Z-2059-2011 | 58209 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice Computed Tomography Scanners, equipped with Software Version 1.1.1. Model #E989605657371. |  |  | March 10, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The Philips MX 16-slice CT scanners are subject to recall/field correction due to a problem encountered with the units when using the 'Combine Viewing' Option. The firm determined that inaccurate dimension measurements display when using the Combine Viewing Option. Moreover, if the dimension measurements from merged images created using this option are used, there is a potential for applying inac | Other | Philips Healthcare issued an 'URGENT- Field Safety Notice MX 16-slice' (dated 3/10/2011) concerning a software issue to customers. The notification informed the firm's users/customers of the potential for inaccurate dimension measurements produced by the MX 16-slice CT scanners when using the Combine Function Viewing Option. Moreover, the letter notified customers of the potential for patient misdiagnosis in the event that inaccurate dimension measurements are used by a physician for reference. The Field Safety Notice described the following: 1) How to identify the affected product; 2) What actions should to be taken by the customers with regard to avoiding the issue -discontinue use of the Combine Viewing Option and instead use the MPR Batch for diagnostic reviews until a planned software upgrade is installed by the recalling firm; and 3) The firm's plans to carry out the installation of a field change order kits to correct the issue.The planned software upgrade is expected to be completed on or before 6 months after the new software is released. If you need any further information or support concerning this issue, pleasecontact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center(1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 193 | 25 units were distributed domestically &amp; 168 system units were distributed both domestically &amp; internationally | Worldwide Distribution -- United States (AL, AR, CA, FL, GA, IL, IN, MS, NJ, NV, NY, OH, PA, SC and TX) including Puerto Rico and countries of Albania, Argentina, Australia, Belgium Brazil, Chile, Columbia, Dominican Republic, Ecuador, Germany, Greece, India, Indonesia, Italy, japan, Kenya, Lebanon, Malawi, The Netherlands, Philippines, Poland, Romania, Saudi Arabia, South Africa, Spain, Switzerland, Turkey, Uganda, United Arab Emirates and Venezuela | Terminated | March 01, 2013 | 722 | 193 | Software | Display/Image | Software update | Software Update
Z-2092-2011 | 58289 | No | FTN | Plastic surgery kit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Argon Medical Devices Arterial Line Kit/1, 18ga x 3", Part Number 498110, Lot numbers 99768048 and 99768520 |  |  | March 14, 2011 | 2011 | Argon Medical Devices, Inc | Packaging for sterile tray has the potential to have a hole which could compromise sterility for the products inside. | Other | Argon Medical Devices sent an Urgent Medical Device Recall letter dated March 14, 2011, by certified mail to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify, quarantine and return any product remaining at their facility utilizing the instruction on the 2011 Tray Recall Response Form. Customers were instructed to complete the response form included with the letter and return by fax or email as indicated as soon as possible. Customers were instructed to pass the notice on to all those who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. For questions regarding this recall call (903) 677-9375. |  | 480 | 480 units | Worldwide Distribution - USA including AL, AZ, CA, CO, DE, GA, HI, IL, IN, LA, MD, MN, ND, NE, NH, NJ, NY, OK, PA, SC, TX, VA, WA, and WI and the countries of Nationwide, Japan, Australia, Canada, Singapore, and the Netherlands | Terminated | April 24, 2012 | 407 | 5780 | Not_Computer
Z-2050-2011 | 58135 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative Susceptibility Card, AST-N184, REF 411 154, For In Vitro Diagnostic Use, 20 cards/box. The firm names on the label are bioMerieux, Inc., Durham, NC, and bioMerieux sa, Marcy-l'Etoile France. |  |  | April 04, 2011 | 2011 | Biomerieux Inc | The piperacillin/tazobactam drug on the antimicrobial susceptibility cards is providing false susceptible and false resistant results for E. coli, as well as an additional 7 organisms; and the same drug is providing false resistant results for K. pneumoniae. | Nonconforming Material/Component | BioMerieux sent an Urgent Product Correction Notice (EXPANDED RECALL), VITEK 2 Piperacillin/Tazobactam Test, Document #9300786," letter dated March 30, 2011, via Fed EX on April 4, 2011. to all affected customers. The letter describes the issues, how the issues impact the test results, describes the product changes that are going to take place, and provides instructions on required actions to be taken.The product changes will involve the eventual changeover to new cards that will be created which do not contain Piperacillin/Tazobactam (TZP2). The new cards will begin shipping by July 16, 2011, but in the meantime, the cards currently in possession of the customers can be used with specific required actions that need to be followed. Until the TZP redevelopment is completed and a resolution can be implemented, BioMerieux requires their customers to immediately take the following actions: The customers will be required to suppress the TZP2 resistant and susceptible results for Escherichia coli, Morganella morganii, Proteus mirabilis, Proteus vulgaris, Providencia rettgeri, Providencia stuartii, Pseudomonas aeruginosa, and Salmonella enterica and use an alternate method of testing. In addition, the customer is to suppress the TZP2 resistant result only for Klebsiella pneumonia and use an alternate method of testing.The letter provides instructions for configuring the VITEK 2 software to suppress the results from reporting.The instructions contained in the recall notification are to be applied to all current and future product lots that contain TZP2 until further notice.The customer is to complete the enclosed acknowledgement form and return it via FAX to acknowledge receipt of the notification and is requested to file the correction notice with their VITEK 2 Systems permanent records according to their internal policy or procedure.For questions or additional information customers were instructed to call 1-800-682-2666, option 3.In addit |  | 9000 | 3,000 boxes as of 3/31/11 | There was no U.S. distribution as the cards were for export only to Canada and Sweden. | Terminated | June 01, 2012 | 424 | 10479 | Computer
Z-2103-2011 | 58271 | No | OCJ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConfidenceÂż Introducer Needle Side Hole 13G x 4-Inch |  |  | March 22, 2011 | 2011 | DePuy Spine, Inc. | Mislabeled: Package labeled (11g x 6 inch) contains a 13g x 4-inch side hole needle | Packaging process control | The firm, DePuy Spine, sent an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated March 22, 2011 to its consignees/customers ( hospital sites and domestic sales distributors). A product reconciliation form was provided to assist in the retrieval of affected product. The letter described the product, problem and actions to be taken. The customers were instructed to return any needles that they may have to their local Distributor for DePuy Spine products and complete and return the Product Reconciliation Form via fax to 508-828-3762; call Customer Service at 877-379-4871 to receive a return goods authorization number and/or mail the form with the return product to DePuy Spine Inc., 50 Scotland Boulevard, Bridgewater, MA 02324, Attn: CONFIDENCE NEEDLE RECALL COORDINATOR, RGA #.If you have any questions or concerns with regard to this notice, please contact the DePuy Spine Regulatory Compliance Department at 508-828-2790. |  | 67 | 67 units | Worldwide distribution: USA including states of: AZ, IA, LA, TX, NE, NJ ,NY, PA, VA, and WA; and countries of: Canada, Australia, Belgium, France,Hungary, Italy, Malaysia, Solvenia, UK, South Africa, Switzerland, and UAE. | Terminated | February 20, 2013 | 701 | 282 | Not_Computer
Z-2082-2011 | 58291 | No | LSE | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | BioPlex 2200 EBV IgG Calibrator Set. Catalog #663-1200. |  |  | March 11, 2011 | 2011 | Bio-Rad Laboratories | Bio-Rad Laboratories confirmed a downward shift in control values for EBV NA-1 analyte when using the BioPlex 2200 EBV IgG Calibrator Set Lot #46704. | Other | The firm, Bio-Rad Laboratories, sent an "Voluntary Field Correction: BioPlex 2200 EBV IgG Calibrator Set, Lot# 46704 for EBV NA-1 Analyte" letter dated March 16, 2011 domestic customers via Fed-Ex and and to international subsidiaries via e-mail. The letter described the product, problem and actions to be taken. The consignees/customers were informed that Bio-Rad confirmed downward shift in control values for the EBV NA-1 analyte when using the BioPlex 2200 EBV IgG Calibrator Set, Lot # 46704. If this occurs, the customers were instructed to readjust the positive control mean using the procedure presented in the customer notification "BioPlex 2200 Quality Control Values (L60013402)". If the re-adjusted mean of the EBV NA-1 controls is within the manufacturer's limit, patient results may be reported. If the new assay positive control mean still does not fall within the manufacturer's limit, the customers were instructed to contact their local Bio-Rad Technical Support and do not report the EBV NA-1 values. The customers were also instructed to complete and return the CUSTOMER RESPONSE FORM via fax to 510-741-5775.If the BioPlex 2200 EBV VCA IgG and EBV EA-D IgG positive and negative control means are within the acceptable QC range, and subsequent passing QC runs are obtained, patient results may be reported for these analytes. Bio-Rad will be manufacturing a new lot of the BioPlex 2200 EBV IgG Calibrator Set which will bring the control values within the manufacturer's limit. It may be necessary to readjust the QC ranges after the consignees receive this new calibrator set lot.Bio-Rad will send their consignees/customers the replacement calibrators and calibrator CDs when available. The consignees/customers were also instructed to discard the calibrator set lot# 46704 in accordance with national and/or local regulations.Customers can call Clinical Diagnostics Group at 510-741-4431 or email to BPD-ProductSupport@bio-rad.com for any questions |  | 270 | 270 sets | Worldwide distribution: USA including the following states: AL, AZ, CA, CT, FL, GA, HI, IL, IN, LA, MA, MD, MI, MN, NC, NJ, NY, OH, OR, PA, TN, TX, VA, WA, and WI; and countries of: Australia, Czech Republic, France, Germany, Italy and Republic of Korea. | Terminated | December 22, 2011 | 286 | 270 | Not_Computer
Z-2067-2011 | 57912 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos UVAR XTS Procedural Kits |  |  | February 01, 2011 | 2011 | Therakos Inc | Therakos received an increased rate of complaints of centrifuge bowl leaks for certain lots of Therakos UVAR XTS Procedural Kits (Product Code XT 125). | Process control | Therakos notified all direct Therakos US and ex-US End Users and Distributor consignees via Federal Express courier with an Urgent Product Correction Notification, dated 2/1/2011. The letter identified the reason for recall along with the affected product. Customers were to immediately discontinue using and discard the identified affected lots of the product. Customers were to use alternative lots. In order to receive credit for the discarded product, customers were to complete the attached confirmation of receipt and destruction. If there were any questions, a question and answer section was provided but additional questions could be directed towards Therakos Customer Service Center at 1-800-828-6316, Option 2,1. |  | 4323 | 4,323 kits | Worldwide Distribution | Terminated | March 13, 2014 | 1136 | 4323 | Not_Computer
Z-2110-2011 | 58407 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | ClearviewÂż iFOB, ClearviewÂż - Clearview iFOB Specimen Collection (accessory) |  |  | March 21, 2011 | 2011 | Alere North Americal | Potential for false negatives test results and negative/faint positive controls | Other | Alere sent a Voluntary Product Recall letter dated March 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to indicate zero on the attached Recall Reply Form, sign, and fax to 1-877-877-0147, if they have no remaining product.For any remaining product customers were instructed to discontinue use and indicate the quantity of kits in the attached Recall Reply Form so that the firm can arrange for them to receive credit. Customers were instructed to destroy any affected units at their site. For any further distribution of product, customers will need to be contacted. Physicians and hospitals were advised to retest patients with negative results. Questions relating to this Voluntary Product Recall call 1-877-597-9586. |  | 1800 | 1800 units | Nationwide Distribution | Terminated | March 07, 2013 | 717 | 15543 | Not_Computer
Z-2140-2011 | 58342 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Interlok 67 MM fixed Cruciate Tibial Plate with locking bar CO-CR-MO/TI 6AL 4V ALLOY, REF 141232, Sterile, BIOMET ORTHOPEDICS, WARSAW. IN I-Beam Tibia Plates are knee joint replacement prostheses intended for application with or without bone cement. |  |  | March 03, 2011 | 2011 | Biomet, Inc. | The firm is recalling the product due to the product's packaging being labeled "I-Beam 75mm. Tibial Plate", but the package actually contains the "Interlok 67mm. CR Tibial Plate". | Packaging process control | Biomet Orthopedics Inc initiated an electronic communication with Direct Accounts on March 3, 2011 and instructed them to quarantine and return all remaining/unused devices to the firm. For questions regarding this recall call 574-372-3983. |  | 8 | 8 | Worldwide Distribution -USA including FL, IN, and NC and the countries of Japan and the Netherlands | Terminated | July 27, 2011 | 146 | 8 | Not_Computer
Z-2105-2011 | 58255 | No | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Inter x 5002 Device, Inter x Professional Sport Device, Inter x 1000 Personal Device, and Inter x Personal Sport packaged with Triad alcohol wipes. The alcohol wipes manufactured by the Triad Group were labeled under brand names Triad, Select, Select Medical Products, PSS Select or PSS World Medical Inc. |  |  | February 07, 2011 | 2011 | Neuro Resource Group Inc | Product package contains alcohol prep pads recalled by Triad Group because the prep pads have the potential to be contaminated with an objectionable organism. | Material/Component Contamination | Triad Group issued a press release on 2/7/2011, and notified consignees by e-mail 3/2/2011. Neuro Resource Group, Inc. ( NRG ) removed the recalled Triad wipes from all kits in house and obtained non recalled wipes from an alternate supplier. For replacement wipes customers may contact NRG customer service at 1-877-314-6500. |  | 57252 | 57,252 wipes | Worldwide Distribution | Terminated | September 23, 2011 | 228 | 57252 | Not_Computer
Z-2112-2011 | 58273 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Portable Blood Gas Analyzer System, Part Number: 393-090, 510 (k) #K092686. All units equipped with software versions: 2.4.1680.25 and all previous versions are subject to recall. |  |  | January 12, 2011 | 2011 | Radiometer America Inc | Following the receipt of two (2) reported customer complaints, the firm determined that the current operating software for their ABL 90 Analyzer does not prevent the further use of the sensor cassette when there is the possibility of iron rich plaque build-up in the reference electrode. If the device operator/user ignores the error message and continues to obtain biased measuring results using the | Software design | The firm, Radiometer America, sent two Urgent "Medical Device Field Correction" letters issued on February 4, and on February 24th, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to replace the sensor cassette. In order to prevent this situation from occurring, the firm informed the customers in the recall/correction notification that they intend to upgrade all ABL 90 Flex analyzers to a new software version. The new software will perform a check during a normal system FLUSH in order to allow the analyzer to detect the bias and will notify the customer immediately of this condition and will automatically disable the cassette thus preventing further use. The letter goes on to remind the device customers that until such time as the software upgrade is installed, they should make sure to adhere to the procedures stated in the Reference Manual which call for the replacement of the sensor cassette immediately upon receiving message #1310. The letter further instructs the customer to complete and return the attached FAX RETURN FORM via fax to 440-871-0463. The letter also states that the expected release date for the new software should be sometime in the second half of February 2011. If you have any questions regarding this letter, or the upgrade, please contact Technical Support staff at 1-800-736-0600 Opt 4. |  | 565 | 32 units were distributed in the U. S. and 533 units were distributed world-wide | Worldwide distribution: USA including states of CA, DC, PA, MN, NY, NJ, TX, OH, MO, WI, UT, and MI; and country of Canada. | Terminated | May 09, 2012 | 483 | 565 | Software | Output/Calculation | Replace sensor cassette/Software update | Software Update
Z-2118-2011 | 53013 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Antares 5.0 Base System, model number 10042832. Distributed by Siemens Medical Solutions USA, Inc., Business Unit Ultrasound, Mountain View, CA. |  |  | May 15, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Shock Hazard. The manufacturer failed to set the system power switch to "ON" during the line level leakage and high pot current draw testing. | Release of Material/Component prior to receiving test results | Siemens Medical Solutions USA, Inc. sent a Customer Safety Advisory Notification letter out on May 15, 2009. Service engineers were to visit each customer to conduct necessary testing.For questions regarding this recall call 650-694-5993. |  | 1 | 1 unit | Worldwide Distribution - USA including CA, IL, MD, NC, OH, and UT and the countries of Australia, China, Denmark, Spain, South Korea, Norway, Poland, Russia and Sweden. | Terminated | April 29, 2011 | 714 | 30 | Not_Computer
Z-2016-2011 | 58188 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Shear Valve Assembly used on CELL-DYN 3500SL Analyzer and CELL-DYN 3500SL Analyzer (refurb). |  |  | March 15, 2011 | 2011 | Abbott Laboratories | The Nylon Washer has an outer diameter which is undersized resulting in the inability to generate patient results. | Pending | Product Correction letters were sent on March 15, 2011. via Federal Express. The letter identified the affected product, explained the problem, the patient impact and the necessary action that were to be taken. Customers were to review the letter and inform their Laboratory Manager of the pending action.; complete and return the Customer Reply Form; and save a copy of the letter for their records. If there were any questions, customers were to contact Customer Support at 1-877-4ABBOTT or their local hematology customer support representative for customers outside of the US. |  | 174 | 174 instruments total, all varieties | Worldwide Distribution -- AZ, CA, FL, IA, ID, IN, MA, MI, MN, MS, ND, NY, OK, OR, SC, SD, TX, WI, WY and Germany, Singapore, Mexico, China, South Korea, Argentina, Japan, Brazil, Canada, Australia, and Chile. | Terminated | November 22, 2011 | 252 | 174 | Not_Computer
Z-2122-2011 | 58482 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6201, Single Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product Code # 2M8063 |  |  | March 24, 2011 | 2011 | Baxter Healthcare Corp. | Pumps submitted for service may not have been properly serviced before being returned to the customer. | Employee error | Baxter visited the direct accounts starting on 3/24/11 informing them that the affected Flo-Gard pumps may have been released from the preventative maintenance with incorrect accuracy testing. The affected pumps were located, taken out of service and placed in quarantine on the same date. A technician not involved with the previous testing visited both accounts on 3/25/11 and performed preventative maintenance service of the three affected devices, and were found to be within specification. |  | 3 | 3 units | Puerto Rico | Terminated | April 29, 2011 | 36 | 3 | Not_Computer
Z-2121-2011 | 58286 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | RETIC results may be invalid when the instrument user interface allows reflex processing of a pre-diluted specimen in the RETIC test panel. Beckman Coulter has validated pre-diluted specimen processing in CBC panels. RETIC test panel using pre-diluted samples has not been validated.IMPACT:The reticulocyte value is an indicator for assessing red cell production. Erroneous results could impact c | Other | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation.2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU).3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS.4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens.5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Beckman Coulter will make corrections in future version software.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including:Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | September 02, 2014 | 1413 | 422 | Software | Output/Calculation | Software update | Software Update
Z-2120-2011 | 58519 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic, Kappa 600 |  |  | April 04, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | A rare measurement lock-up condition that may (at a rate of approximately 1 in 18,000 devices) inappropriately trigger ERI/RRT in the Medtronic Dual Chamber Pacemakers: Kappa 600, 700, 800, 900, EnPulse, Adapta, Versa, Sensia. This is caused by a random lock-up of the measurement system hardware that may result in an incorrect battery voltage reading of zero. This issue does NOT impact battery lon | Component design/selection | The firm, Medtronic, sent an "Important: Medical Device Correction" letter dated April 2011 to its customers. The letter described the product, problem, and actions taken. Medtronic informed the customers that this issue does NOT impact battery longevity and does NOT require device explant. Currently, the device can be reset to normal operation by a Medtronic representative. Reset devices are no more likely to experience a recurrence of this issue. Medtronic is planning to release a programmer software update that will allow the clinician to reset the device. In addition, Medtronic has consolidated their product performance data with the vision of creating a primary reference tool for performance updates.If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 1989180 | 1,989,180 | Worldwide distribution: USA (nationwide) including Guam, Puerto Rico and DC; and countries including: Algeria, Andorra, Anguilla, Argentina, Armenia, Aruba, Australia, Austria, Bahamas, Bangladesh, Belgium, Bermuda, Bolivia, Brazil, Brunei, Darussalam, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Virgin Islands (British). | Terminated | November 27, 2012 | 603 | 1989180 | Hardware | Device Operation | Programmer software update | Software Update
Z-2113-2011 | 58367 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Doctors Review System, OB-Windows/OB-View software release numbers 4.1 through 4.1.3, Digisonics, Inc., 3701 Kirby Drive, Houston, TX 77098 |  |  | February 09, 2007 | 2007 | Digisonics, Inc | Study would on occasion have fetal heart rate and/or cervical length on the report that was not entered nor did it belong to that patient. Software displayed incorrect data for fetal heart rate and cervical length in OB ultrasound reports. | Other | The firm, Digisonics Inc, notified the affected customers of software release via e-mail on January 24, 2007 and follow-up e-mail on February 7, 2007. The e-mails described the product, problem and the actions taken. The firm was able to trap the error and after this report the developer was able to determine the cause and a new executable was sent and installed at the customers site.If you have any question, please call 800-940-3240 or e-mail: service@digison.net. |  | 128 | 128 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | May 02, 2011 | 1543 | 128 | Software | Output/Calculation | Software update | Software Update
Z-2119-2011 | 58230 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The 4D Integrated Treatment Console; Model Number: H51; Reference: 4DITC versions 8.6 and below; File Reference: CP-2007-0288; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 28, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly, in that the Multileaf Collimator (MLC) message displayed on the 4DITC on versions 8.6 and below, does NOT interlock beam delivery. | Software design | Varian Medical Systems sent an Urgent Medical Device Correction Letter on February 28, 2011, to all affected customers.Varian has created the procedure described in CTB-GE-579 to inform users on proper interfacing of the demountable BrainLAB m3 on the Varian Clinac. The procedure ( CTB-GE-579) is enclosed and posted to the customer site: http://www.MyVarian.com.This design fix in 4DITC v10.1 or later will be provided to all affected sites.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager.USA and Canada - 1-888-VARIANS (888-827-4265)Europe - + 41 41 749 8844 |  | 17 | 17 | Class 2 Recall - Worldwide Distribution -- USA including states of ; CA, NV, NJ, GA, IL, and TX and countries of Australasia, Belgium, Brazil, China P.R., Denmark, France, Germany, Germany ( BrachyTherapy) Hong Kong, India Pvt. Ltd, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland and United Kingdom. | Terminated | March 30, 2012 | 396 | 17 | Software | Physical Safety Hazards | Software update | Software Update
Z-2114-2011 | 58282 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | For Panels containing CBC: The RBC and HCT are improperly corrected and not flagged when there is a WBC Pump Failure and the WBC ~ 140x103/uL. For Body Fluid Panels:The RBC is improperly corrected and not flagged when there is a WBC Pump Failure. IMPACT:Erroneous RBC and/or HCT could affect the interpretation and management of patients at risk with various types of anemias. | Other | The recall communication was initiated on 10/20/2010, with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with PCA Response Form (via US Postal Service for US customers) to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter informs the customers about the problems identified and provides an action/resolution to this issue. When a WBC Pump Failure occurs, the WBC or TNC results will be flagged with 'R' and the WBC Pump Failure System message. The RBC and/or HCT results associated with these affected analyses also require review. This will be corrected in a future version of software.Beckman Coulter asked consignees to complete and return the enclosed response form within 10 days and share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Nationwide in the US and worldwide to: Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom | Terminated | September 02, 2014 | 1413 | 422 | Software | Alarm/Message | Software update | Software Update
Z-1895-2011 | 57948 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2 (part # 378600) is a microprocessor-controlled pipetting and diluting system. |  |  | January 18, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after erroneous, though potentially credible, results due to intermittent reagent level sense failures could be generated. During these failures, the instrument fails to produce a warning that a reagent vial is low or out of reagent. The low or absent levels of reagent used during these failures may lead to insufficient antibody being dispensed into a daughter tube there | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated January 18, 2011 to all their customers. The letter described the product, problem and actions to be taken. The customers were instructed to: 1) Follow PrepPlus 2 Instructions for Use (Part # 4277264-BA) for all reagent and specimen requirements prior to initiating a run on the PrepPlus 2 including:A) Ensuring all dead volume requirements for reagents as indicated in 7.11 and 7.12 of the Sample Processing sections and 12.2 of the Troubleshooting Guide. Appropriate volume of reagent in all vials per the planned number of tests to be run must contain at least the volume listed on the Worklist (plus an additional 10% of the full volume of the container to account for dead volume)".B) Ensuring specimen processing dead volume requirements are met as indicated in section B of Specimen Tube Size to ensure the minimum amount of specimen needed.C) Monitoring for any foaming/bubble accumulation in the reagent vial.2) A review of all clinically relevant results generated from samples prepared on the system be conducted. Abnormally dim or unexpected negative sample staining patterns should be reviewed to ensure correlation with clinical symptomatology and other diagnostic data. If the clinical information Is in accordance with the clinical data released, the validity of the results could be considered accurate and no retrospective action may be needed. Ifinconsistent data is observed, please adhere to good laboratory practices within your institution and consider data re-analysis and re-testing.The customers were also instructed to share this information with their laboratory staff and retain it as part of their laboratory Quality System documentation; if they have forwarded any of the affected product(s) listed above to another laboratory, provide a copy of the letter to them, and complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM via fax to: (786)639-7500/7501/750 |  | 369 | 369 units (140 in the U.S.) | Worldwide distribution: USA and countries including: Australia, Barbados, Belarus, Belgium, Canada, Croatia, France, Germany, Ireland, Italy, Japan, Malawi, Mozambique, Namibia, Poland, Puerto Rico, Russian Federation, Slovenia, South Africa, Spain, Swaziland, Sweden, Switzerland, Ukraine, United Kingdom, Zambia. | Terminated | September 02, 2014 | 1323 | 369 | Not_Computer
Z-1885-2011 | 57743 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | MDI(R), S.A.L.T. Airway Device, Supraglottic Airway Laryngopharyngeal Tube, Intended for use with 6.5mm -9.00mm ET tubes, Including SALT(TM) airway, ET tube clamp, Latex-free securing strap, Tongue blade, Lubricating jelly, NON STERILE, REF 79-580, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, Manufactured in the Dominican Republic of U.S. Materials |  |  | January 05, 2011 | 2011 | Microtek Medical Inc | Product contains Triad Lubricating Jelly which was recalled. | Other | The firm, Microtek Medical, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and place any S.A.L.T. Airway Device (79-580) in quarantine; if they have distributed any of the product, they should identify their customers and notify them at once of this recall; complete and return the enclosed URGENT: MEDICAL DEVICE RECALL-RETURN RESPONSE form via fax to: (662)244-3011 or email: tom.bonner@ecolab.com, and return any unused product inventory to Microtek Medical. The firm will provide the customers with a replacement product as soon as product becomes available.If you have any questions in regards to this recall, please feel free to contact the V.P.-Regulator Affairs at 1-888-844-0988 or email: Tom.bonner@ecolab.com. |  | 2230 | 223 cases of 10 units | Nationwide distribution: AZ, CA, CT, FL, IN, MS, NC, NH, PA, TX and WI. | Terminated | June 20, 2011 | 166 | 2230 | Not_Computer
Z-2046-2011 | 58240 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac High Energy Medical Linear Accelerator, a Trilogy Radiotherapy Delivery System; Model #s: H14, H18, H27, H29, HCX; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The wedge angle labeling on the wedge tray may not correctly match the wedge body. However, the wedge body is labeled with the correct wedge angle. | Labeling mix-ups | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 24, 2011, to all affected users. The letter described the product, problem and actions to be taken. The customers were instructed to IMMEDIATELY inspect all wedges to verify that the wedge angle label on the wedge body matches the wedge angle on the wedge tray; if any wedge is identified with mismatched labeling follow the written instructions provided in the letter; always measure wedge data at installation of new or replacement wedges; verify all treatment accessories of "matched" linear accelerators, and please advise the appropriate personnel working in your Radiotherapy Department of the content of this letter.Varian has enhanced their manufacturing quality control process to further ensure correct labeling of wedges. Also, Varian will replace any wedge that is identified as mislabeled, at the request of the customer. If you require further clarification, please feel free to contact your local Varian Customer Support Service Manager at: 1-888-VARIAN5 (888-827-4265) for USA and Canada; +14 41 749 8844 Europe; North American: support-americas@varian.com; Australia/New Zealand:support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America:soporte.al@varian.com. |  | 5541 | 5541 units | Worldwide distribution. | Terminated | November 05, 2012 | 620 | 5541 | Not_Computer
Z-2048-2011 | 58331 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Sheath Introducer Kit |  |  | March 25, 2011 | 2011 | Arrow International Inc | Labeling discrepancy - the outer corrugate label as well as a card inside the package state the product contains dry natural rubber - this is not duplicated on the unit packaging. | Other | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine any products with the listed lot numbers; complete and return the Immediate Attention Request form via fax to 800-343-2935 indicating how much product they have in stock and whether they need labels. In the event that an alternative approach is needed, the customers are to contact Arrow Customer Service for more information. Once the labels have been received and applied the product can be released from quarantine.An Arrow's Sales Representative will pay a personal visit to their customers and obtain the quantity and the amount of labels they need. If you have any other questions, feel free to contact your local sales representative or Customer Service at 866-396-2111. |  | 840 | 840 | Nationwide distribution: USA including states of: PA and NJ. | Terminated | January 29, 2014 | 1041 | 840 | Not_Computer
Z-0836-2011 | 56024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems FastPlan software, Versions 5.5 and 5.5.1, Model Number HZ3, manufactured by Varian Medical Systems, Palo Alto, CA |  |  | March 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to software Issue, an error occurs when performing cone planning using MRI, which may result in the patient receiving a higher radiation dose than expected, potentially to healthy tissue. | Device Design | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated.Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated.Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan.For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager.Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.comInternet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com |  | 218 | 218 units | Worldwide distribution - USA nationwide and the countries of Australia, Brazil, Canada, Chile, Colombia, Georgia, Germany, Hong Kong, India, Ireland, Japan, Macau, Malaysia, Russia, Singapore, Spain, Sri Lanka, and Taiwan | Terminated | February 24, 2012 | 703 | 218 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-2127-2011 | 58313 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Cirrus HD-OCT Model 4000 instruments, new and refurbished, distributed with, or upgraded to, version v 1.0 - 4.5, 5.0.0 or 5.1.0 |  |  | March 10, 2011 | 2011 | Carl Zeiss Meditec Inc | A software defect exists in versions 5.0.0 and 5.1.0 where when following a specific workflow, exam data may be saved under the name of the first patient | Software design | Carl Zeiss Meditec, Inc. sent an Field Correction Action letter dated March 9, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letters requested acknowledgement of the letter and a response and the fact that a follow up packet with revised software will be sent to all customers with V 5.0.0 or 5.1.0 software.For questions, please call 1-855-853-9461. |  | 6888 | 3465 US; 3423 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Algeria, Argentina, Australia, Bangladesh, Belgium, Boliva, Brazil, Canada, France, Chili, China, Columbia, Denmark, Ecuador,Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Japan, Lebanon, Malaysia, Mexico, New Zealand, Pakistan, Peru, Singapore,South Korea, Spain, Sri Lanka, Taiwan, Uruguay and Venezuela | Terminated | March 29, 2012 | 385 | 6888 | Software
Z-2125-2011 | 58424 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, RT Chart Application v 8.0 through 8.9 |  |  | March 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue was identified with the gantry rotation direction when creating opposing field within the RT Chart function of ARIA, or possibly with third party planning systems and delivering the plan at 4DITC. | Software design | Varian Medical Systems, Inc. Oncology Systems sent Urgent Medical Device Correction Field Safety Notice letters on March 9, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letters identifeid the product, the problem, and the action to be taken by the customer. Customers were instructed to:Manually enter the gantry angel 0.0 when the plan indicates angel 0.0. Manual entry is only allowed to one decimal place, so the range that triggers the issue cannot be created if the keyboard is ued to enter the gantry angel.Users with third party plan sources should verify whether gantry angles with two or more decimal places are supported. If so, customers may be affected by this issue, and should take steps suited to the device.In all cases, the display at the Treatment console is indicative of what direction the gantry will move - users are advised to be especially attentive to these values, particulary when non-zero couch rotation values are used.Users are advised to always run a simulation of the plan, including all plan motions, at the treatment unit prior to first patient treatment. For further clarification customers were insturcted to contact their local Varian Customer Support District Manager. For questions regarding this recall call 650-424-6471. |  | 612 | 612 units | Worldwide Distribution - USA including CA, DC, FL, GA, MD, MN, MS, NY, OH, and TN and the countries of Canada, Australia, Austria, Bahrain, Bagladesh, Belarus, Belgium, Brazil, Columbia, Costa Rica, Cypress, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Guadeloupe, Guatemala, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Luxemborg, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion, Romania, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tobago, Tunisia, Turkey, Ukraine, Emirates, and the United Kingdom. | Terminated | March 13, 2012 | 370 | 612 | Not_Computer
Z-2128-2011 | 58287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System Part Number: 629029 |  |  | October 20, 2010 | 2010 | Beckman Coulter Inc. | Body Fluids results may be incorrectly interpreted by the Laboratory Information System (LIS). Body Fluid results ~ 1,000 celis/uL are reported using the US numeric separators to demark thousandths places (comma, ",") and partial units (decimal, "."). An LIS that uses non-US formats could interpret the comma as a partial units separator, which could lead to reporting erroneously low body fluid r | Other | The firm, Beckman Coulter, issued an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 20, 2010, with PCA Response Form to all customers who purchased the UniCel DxH 800 Coulter Cellular Analysis System. This letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove all cassettes from the mix station prior to initiating the Diluent Dispense function. Do not initiate Diluent Dispense while specimens are being processed in cassette presentation.2)Do not use pre-diluted samples for RETIC counts. The instrument is validated for pre-diluted CBC's only, as indicated in the Instructions for Use (IFU).3)Verify that their information Technology personnel are aware that the Body Fluids data string is only transmitted to the LIS using the US numeric separators and that the Body Fluid results are properly validated at the LIS.4) Use the lip on the reagent drawer to move the drawer as shown. Avoid touching the drawer slides while replacing containers. Clean all spills promptly to reduce the risk of exposure to blood borne pathogens.5)Complete and return the enclosed PRODUCT CORRECTIVE RESPONSE FORM within 10 days, share this information with laboratory staff and retain this notification as part of laboratory Quality System documentation.Beckman Coulter will make corrections in future version software.Questions were directed to Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or consignees/customers local Beckman Coulter representative. |  | 422 | 422 units in total (175 in US) | Worldwide distribution: USA (nationwide) and countries including:Australia, Belgium, Canada, Colombia, Croatia, France, Germany, Greece, Hong Kong, India, Italy, Japan, Korea, Republic of, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1413 | 422 | Software | Output/Calculation | Software update | Software Update
Z-2123-2011 | 58574 | No | LTI |  |  | Gastroenterology/Urology | PMA | Lap-Band Access Port Needle 10-Pack, 20 G x 51 mm (2 in.), Product Code: 8-20302-10, Lot Number: 9L01N |  |  | March 21, 2011 | 2011 | Allergan | The recall was initiated because Allergan has determined that a number of Access Port Needle pouches from one specific lot number may have small gaps in a pouch seal. Consequently the product may no longer be sterile. | Pending | Allegan sent an Urgent Medical Device Recall letter dated March 21, 2011, to all affected customers. Allergan requested that upon receipt of letter, that the customer check their existing inventory of product for affected lot. Customers were requested to complete Business Reply Card, stating quantity of inventory of affected lot, if any, and return to Allergan Medical. Customer with remaining inventory were instructed to contact Allergan Medical Product Support at 1-800-624-4261, x5972 or 1-805-961-5972 for return and immediate replacement. . Customers will be issued credit for returned product.For questions regarding this recall call 805-319-0589. |  | 1888 | 1878 x 10-packs | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, MN, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Australia, Canada, Germany, Great Britian, and the Netherlands | Terminated | April 06, 2012 | 382 | 1888 | Not_Computer
Z-2124-2011 | 58470 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂżVentralex" Hernia Patch Medium Circle with Strap Diameter 6.4cm /2.5",\ |  |  | April 11, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Medium Ventralex Hernia Patch 6.4cm /2.5" packaged in a pre-printedcarton identified for a BardÂż Ventralex Small Circle 4.3cm /1.7" | Labeling mix-ups | Davol sent an Urgent: Medical Device Recall letter on April 11, 2011, via FedeX Priority Overnight to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and remove and return the recalled lot. Distributors were requested to notify their accounts.For questions contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 146 | 146 units | Nationwide Distribution - USA including AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MD, MI, MO, MT, NC, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, and WV. | Terminated | October 27, 2011 | 199 | 146 | Not_Computer
Z-2135-2011 | 58325 | No | HWD | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Low Profile Self-Tapping Bone Screw, 6.5 mm X 25 mm, Qty. 1, Biomet Orthopedics, 66 East Bell Drive, P.O. Box 587, Warsaw, IN 46581 USA. STERILE. |  |  | March 10, 2011 | 2011 | Biomet, Inc. | One lot of 30 mm screws was switched with one lot of 25 mm screws. | Mixed-up of materials/components | Biomet Orthopedics sent an Urgent Medical Device Recall Notice letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove the affected lots from their inventory, and return them to Biomet. Complete and enclosed FAX Back Response Form to 574-372-1683 prior to return of product. Distributors were instructed to notify all sub-accounts of the recall.Confirm receipt of notice by calling, 800-348-9500, extension 3755 or 3983.For questions regarding this recall call 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 100 | 100 | Worldwide Distribution - USA including CA, GA, IA, IL, IN, KS, KY, MD, MI, MN, MO, NC, ND, NJ, NY, OK, PA, SD, TX, VA, WA, and WI and the country of Japan. | Terminated | February 17, 2012 | 344 | 100 | Not_Computer
Z-2134-2011 | 58402 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Patient Monitor, Model 91387. All modules may be affected except Models 90470, 90478, 90496, &amp; 91496. |  |  | March 24, 2011 | 2011 | Spacelabs Healthcare, Llc | The circuit that supplies backup power to the modules in the Medical Patient Monitor, Model 91387, is not working. If there is a power interruption, the three-minute backup may not be operational and customers may lose any changes they have made to the alarm settings. | Other | SPACELABS Healthcare sent an Urgent Medical Device Correction letter dated April 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use their monitors until they can be upgraded. Customers were recommended to brief their staff about the possibility of losing changes to the alarm settings and to restore them after a power interruption. Field service engineers will contact customers to schedule a mutually agreeable time to come to their hospital and upgrade their affected patient monitors at no cost to the hospital. US Consignees can call 1-800-522-7025, Select 2 for Technical Support. Consignees outside of the US can call 1-425-657-7200 x5089 for any questions about this recall. |  | 4470 | 4470 units total (1894 units in US and 2579 units internationally) | Worldwide Distribution - USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, Republic OF CHINA, Thailand, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA.AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PORTUGAL, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, Republic OF CHINA, Thailand, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA. | Terminated | February 15, 2012 | 328 | 4470 | Hardware | Device Operation | Notification until upgrade | Repair
Z-2133-2011 | 58278 | No | OHR | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | Enforcement Discretion | Catheter Insertion Kit (containing, Under pad, PVP Pad, BZK Prep Pad, Sterile Lubricating Jelly, Vinyl Gloves and Poly Zipper Bag). |  |  | March 24, 2011 | 2011 | Apogee Medical, Inc | Kits contain PVP (Providone Iodine) prep pads that are being recalled by the Triad Group due to possible micro contamination, Elizabethkingia meningoseptica. | Other | The firm, Apogee Medical LLC, sent an "Urgent Medical Device Recall" letter dated March 23, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review their inventory; complete and return the enclosed Product Recall Response Form as soon as possible via fax to (919) 570-9611, mail to: Apogee Medical LLC, Youngsville Commerce Center, 90 Weathers Street, Youngsville, MC 27596, or e-mail to: dpeper@apogeemed.com; cease further distribution and return any product that is the subject of this recall to Apogee Medical LLC. All returned product must be labeled with a returned goods authorization label that has been provided with the letter. Apogee will replace returned product upon receipt of the Product Recall Response Form and identified product.In addition, if the customer distributed the product, they are to identify their customer(s) and notify them at once of this product recall. The notification should include a copy of the this recall notification letter and should instruct the customer to return the recalled product so that it can be returned to Apogee.Should you have any questions regarding this notification, please call the Manager Regulatory Affairs/Quality Assurance at (919) 435-5409. |  | 5503430 | 5,503,430 units | Nationwide distribution: USA including states of: AL, AZ, CA, FL, GA, IL, IN, KS, LA, MA, MO, NJ, NY, OH, OK, TN, TX, VA and WA. | Terminated | February 28, 2012 | 341 | 5503430 | Not_Computer
Z-2129-2011 | 58383 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | APD Disposable Sets and Extension Sets used with Baxter HomeChoice and HomeChoice PRO Automated Peritoneal Dialysis Systems; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | March 16, 2011 | 2011 | Baxter Healthcare Renal Div | Clinicians are being reminded of the current warning in The HomeChoice APD Systems Patient At-Home Guide, 07-19-61-244 (October 2,2009), which states on page 3-13 "To reduce the risk of bums, electrocution, fire, or injury to persons: Close supervision is necessary when this product is used by, on, or near children or those unable to care for themselves." Tubing entanglement around the neck may r | Other | Baxter sent Safety Alert letters dated March 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to consider the following safeguards for at-risk patients on HomeChoice Therapy:Have a caregiver present while the patient is performing therapy.Move the HomeChoice cycler to the foot of the bed, if possible, to minimize the likelihood of entanglement that may lead to asphyxiation.Install bed railings for patients at risk for seizures or falls from bed.Complete the attached customer reply form and fax it to Baxter at t847-270-5457.For questions regarding this recall call 888-736-2543, option 2. |  |  |  | Worldwide Distribution - USA including AZ, CA, CT, DC, FL, KY, MO, NC, NM, Ny, OH, PA, TN, TX, and VA and the countries of Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, South Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Qatar, Russia, South Africa, Saudi Arabia, Singapore, Slovania, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingom, and Venezuela. | Terminated | March 08, 2012 | 358 |  | Not_Computer
Z-2136-2011 | 58052 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | HealthPort BioAnologics Patient Electrode Sensor Test Kit, Part # SEN-0025-0, 25 kits/carton |  |  | February 08, 2011 | 2011 | Vermed, Inc. | Vermed kits contain a product component (Triad Alcohol Prep Pads) recalled due to potential Bacillus cereus contamination. | Other | Vermed sent consignee a recall notification letter on February 8, 2011, followed by a phone call and email. The consignee was instructed to respond with a Recall Reply Form via fax and to notify any customers to whom the affected product was further distributed. All recalled alcohol prep pads were to be returned to Triad Group. Firm is replacing Triad alcohol pads in kits with pads from another supplier and has removed Triad from its approved vendor list.For questions regarding this recall call 802-722-9976. |  | 1250000 | 2500 kits total; 500 from each lot except 1000 from Lot 203031 | OR | Terminated | July 25, 2011 | 167 | 1250000 | Not_Computer
Z-2138-2011 | 58449 | No | DTH |  |  | General &amp; Plastic Surgery | PMA | SharpointTM 4-0 Black Monofilament Non-Absorbable Nylon 18"/DS24mm Reverse Cutting Edged Needle. |  |  | March 30, 2011 | 2011 | Surgical Specialties Corp | Mislabeled - a number of units of the product labeled as 4-0 Black Mono Nylon suture product contain a 4-0 Black Braided Silk suture. | Mixed-up of materials/components | Surgical Specialites sent an Urgent Product Recall Notification letter on March 30, 2011, along with an attached Customer Acknowledgment Form to the affected consignees. It was reported that upon opening a box labeled as 4.0 Black Mono Nylon, it was observed that the package acutally contained 4-0 Black Braided Silk suture. Customers were to review their current inventory, segregate the affected lot number and / or contact their customer accounts that have purchased this specific item / lot code requesting a return of the product to customers' facility. To ensure customers have received this notification and to arrange for credit and the return of identified product, the Customer Acknowledgment Form should be completed and returned via fax to our Quality Assurance department at (610) 404-8488. |  | 2628 | 2,628 | Class 3 Recall - Worldwide Distribution-- USA and Australia. | Terminated | August 04, 2011 | 127 | 2628 | Not_Computer
Z-2137-2011 | 58166 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | NX is the radiographer's image identification and quality control tool. It offers a broad array of features developed especially to meet the needs of the radiographer. Standard tasks are completed quickly and effortlessly with a simple touch screen. In-room integration provides a more convenient workflow for the radiographer throughout the imaging process. NX communicates seamlessly with the hospital's picture archiving and communication system (PACS), radiology information system (RIS) and hospital information system (HIS). It is DICOM compliant and meets IHE guidelines. |  |  | March 04, 2011 | 2011 | AGFA Corp. | When a radiographer is working on two open studies on an NX workstation and, in specific situations, a problem of image/annotation mix-up can occur on the NX system. | Other | AGFA Corp. sent an "URGENT FIEL SAFETY NOTICE" letter via FED-EX to the customers on March 4, 2011. The letter describes the safety alert and mitigation. Customers were instructed to complete and return the Urgent Field Safety Notice Feedback Form by fax to 864-421-1664.For questions regarding this recall call 864-421-1815. |  | 775 | 775 units | Worldwide Distribution - USA including CA, GA, MS, ND, NM, NY, OK, PA, SC, SD, TX, and VA and the country of Canada | Terminated | March 27, 2012 | 389 | 775 | Not_Computer
Z-2139-2011 | 58326 | [6 empty] | ***REF IOSS413*** OSSEIOTITE Certain Implant 4 x 13mm*** Sterile using Radiation*** Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain*** |  |  | March 10, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i is recalling their product Osseotite Certain Implant Device IOSS413. The internal thread was not manufactured correctly. | Other | BIOMET 3i sent an Urgent Medical Device Recall Letter on March 14, 2011, to all affected customers via emial or fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told to contact Customer Service at 1-800-342-5454.Customers were instructed to respond with the attached Response Fax or email DomesticComplaints@Biomet.com.If customers still had implant(s) in their inventory, they were to return the implant(s) to BIOMET 3i using the following address:BIOMET 3iRegulatory Compliance Department4555 Riverside DrivePalm Beach Gardens, Florida 33410Incident No. CD117886A replacement implant(s) will be sent to them one the product is returned.For questions regarding this recall call 561-776-6906 (EST 8:00am to 5:00pm). |  | 217 | 217 | Worldwide Distribution - USA including AK, FL, GA, NC, NJ, NY, SC and the countries of Canada, Europe, Japan, and South America. | Terminated | January 25, 2012 | 321 | 217 | Not_Computer
Z-2149-2011 | 53211 | [6 empty] | 4Z1c Matrix Array Transducer, model number 10033682, used with ACUSON SC2000 ultrasound systems, model number 10433816, manufactured by Siemens Medical Solutions, Ultrasound Business Unit, Mountain View, CA. |  |  | July 10, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Burn risk-- due to coolant leaks, patients may suffer burns or blisters due to increased temperature. | Equipment maintenance | A customer safety advisory notification was sent out to customers on 7/10/09, and a mandatory software update will also be conducted. |  | 42 | 42 units | Nationwide. Foreign: Japan, Italy, South Korea, Singapore, Yemen, Germany, Australia, and Greece. | Terminated | May 05, 2011 | 664 | 42 | Other | Physical Safety Hazards | Software update | Software Update
Z-2143-2011 | 58480 | No | MJA | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BACTEC(tm) MGIT(tm) 5-Tube Spare AST Set Carrier, Bag of 3, Catalog 445943. ***Becton, Dickinson and Company |  |  | August 31, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic device components are mislabeled. | Other | BD Diagnostic Systems sent an Urgent Product Recall letter on August 31, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter advised customers to discard any carriers or labels that contain the incorrect barcode for replacement. Accounts were requested to complete an attached reply form to acknowledge their receipt of the notification and fax it to 410-316-4258.For questions regarding this recall call BD Customer Service at 1-800-675-0908. |  | 4123 | 4123 total labels | Worldwide Distribution - USA including CA, MA, MD, VA and the countries of Argentina, Colombia, Belgium, China, India, Hong Kong, Peru, Singapore, Thailand, and Trinidad. | Terminated | July 08, 2011 | 311 | 4123 | Not_Computer
Z-2145-2011 | 58305 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029. |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the DxH 800 may fail to meet the carryover specification for the WBC Differential (DIFF) when the WBC is greater than 70 X 10 to the 3rd power cells/micro L (70 X 10 to the 9th power cells/L). Beckman Coulter indicated the DIFF% carryover is unlikely to have clinical impact. | Other | Beckman Coulter sent a Product Corrective Action (PCA) letter with attached Recall Response Form via US Postal Service on February 11, 2011, to all customers who have the affected instruments. Consignees were informed of the reasons for recall, the products affected and that a resolution for these issues would be corrected in a future release(s) for the DxH 800. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with your laboratory staff and retain this notification as part of your Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, Consignees were told to contact Beckman Coulter, Customer Service at (800) 526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Class 2 Recall - Worldwide Distribution -- USA, and the countries of : Canada, Australia, Macao, Belgium, Malaysia, Colombia, Myanmar, Croatia, Netherlands, Czech Republic, Norway, France, Philippines, Germany, Puerto Rico, Greece, Qatar,Hong Kong, Russian Federation, Hungary, Saudi Arabia, India, Singapore, Spain, Israel, taly Sweden, Japan, Switzerland, Korea, Turkey and United Kingdom. | Terminated | September 02, 2014 | 1299 | 566 | Not_Computer
Z-2146-2011 | 58181 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Single Day Infusor 2 mL/h Portable Elastomeric Infusion System; an Rx sterile disposable drug delivery system; Baxter Healthcare Corporation, Deerfield, IL 60015; REF 2C1071KJP |  |  | March 15, 2011 | 2011 | Baxter Healthcare Corp. | Some of the Single Day Infusor System 2 mL/h have a foil over-pouch that was incorrectly labeled as the Half Day Infusor System 5 mL/h, product code 2C1073KJP, lot 10H054. | Labeling mix-ups | Baxter Healthcare Corp. sent Urgent Product Recall letters dated March 15, 2011, via first class mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The accounts were requested to examine all of their inventory locations for both product code 2C1071KJP and 2C1073KJP, lot 10H054, and return the affected lot for replacement or credit. The accounts were also requested to complete the attached customer reply form listing the number of held for return, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Center for One Baxter at 1-800-422-9837. |  | 3696 | 3,696 units | Worldwide Distribution - USA including AL and the countries of Mexico and Spain | Terminated | March 12, 2012 | 363 | 3696 | Not_Computer
Z-2148-2011 | 58307 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029 |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: The DxH 800 may fail to meet the carryover specification for NRBC% when the WBC is greater than 50 X 10 to the 3rd power cells/micro L (50 X 10 to the 9th power cells/L) Beckman Coulter indicated the NRBC% carryover is unlikely to have clinical impact. | Other | Beckman Coulter sent an Urgent Product Correction letter, dated February 11, 2011, via US Postal Service to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were informed the issues would be corrected in a future release(s) for the DxH 800. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain this notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA including AL, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Australia ,Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | September 02, 2014 | 1299 | 566 | Not_Computer
Z-2147-2011 | 53051 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X300 Systems with software versions 3.0.00 and 3.0.02, model numbers 10348531, 10348532 and 10348533, manufactured by Siemens Medical Solutions, Mountain View, CA. |  |  | July 24, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Mistreatment risk-- when a user defined formula is used to calculate a value in a vascular exam, data is switched laterally when it is internally transferred and displayed on the summary report within the system. | Software design | A field correction was initiated and a Customer Safety Advisory Notification was sent out on 7/29/09. The letter identified the affected product and stated that there was a potential for user defined calculations to be incorrect within the vascular package when using the listed software versions of the product. The letter also discussed the steps that customers could take to avoid the potential risk of the issue. A mandatory update software was developed and will be installed free of charge. Customers should have been contacted to schedule a time for the installation. Questions were to be directed to the local service support person. |  | 154 | 154 units | Worldwide Distribution -- Austria, Brazil, China, Columbia, Germany, Denmark, Egypt, Great Britain, Honduras, Ireland, India, Italy Japan, South Korea, Mexico, Malaysia, Peru, Poland, Portugal, Romania, Russia, Sweden, Singapore, Thailand, Turkey, Taiwan, United States, Uruguay, Venezuela, and South Africa. | Terminated | May 09, 2011 | 654 | 154 | Software | Treatment/Delivery/Therapy | Sofrware update | Software Update
Z-2152-2011 | 57044 | No | CBL | Carbon dioxide absorbent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allied Healthcare Products Inc. CarbO2Lime Carbon Dioxide Absorbent Granules for Canister Refill, Non-Hygroscopic Soda Lime USP-NF, Type: High Moisture (12% to 19%), Granule Size: 4 to 10 Mesh (4.8-2.0mm), packaged in 5-gal. pails, #55-01-0003; prepacked canisters (12 canisters/case), #55-01-0005; and replacement bags (12 bags/case), #55-01-0006, RX. The label shows the product is Manufactured by Allied Healthcare Products Inc., St. Louis, MO. |  |  | August 12, 2010 | 2010 | Allied Healthcare Products Inc | The granules do not absorb as long as expected in the anesthesia machines | Other | Allied Healthcare Product, Inc. issued recall letters dated August 10, 2011, via certified mail beginning 8/12/10 and/or via e-mail with a read receipt. Letters were issued for both distributors and end-use customers (such as medical facilities or pharmacies) explaining the reason for recall. Distributors were requested to conduct a subrecall using the enclosed customer letter for end users. A response form was included for both distributors and end users to complete to indicate the amount of product to be returned. The customer was instructed to contact the recalling firm for return instructions.For questions regarding this recall call 314-268-1661. |  | 317915 | 949 pails, 335/12-canister cases, and 706/12-refill bag cases | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MD, MS, NC, NE, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Canada and Grenada. | Terminated | March 28, 2012 | 594 | 317915 | Not_Computer
Z-1971-2011 | 57383 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Precision 500D X-ray systems with under-table collimator part number 5234960. |  |  | June 10, 2010 | 2010 | GE Healthcare, LLC | Use of high force when moving the image carriage/IDD into forward/park position can cause changes to the mechanical alignment of the collimator blades and allow the blades to open wider thanspecified, consequently allowing radiation outside the permitted field without warning to the user. Thisissue manifests with symptoms wherein the user cannot fully close down the lateral blades. | Radiation Control for Health and Safety Act | GE Healthcare sent an Importan Electronic Product Radiation Warning letter to customers. The letter discussed the product defect description and related hazards, affected product details, instructions, and product correction. The firm will send a GE Field Service Representative to visit each affected customer site to inspect and correct all potentially affected units in the field following the instructions in field modification instructions No: 10846.If you have any questions please call 800-437-1171. |  | 440 | 440 US | Nationwide Distribution | Terminated | September 15, 2014 | 1558 | 440 | Not_Computer
Z-2132-2011 | 58567 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LXi 725, Part Number: 476501. |  |  | April 05, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received reports of incorrect electrolyte results on Synchron LX systems due to a variety of maintenance-related hardware issues. | Equipment maintenance | The recall communication was initiated on 04/05/2011 with Beckman Coulter forwarding a Product Correction letter with attached customer response form to the affected customers via US Postal Service. The letter informed customers of the products affected, the issue, the impact, and the actions to be taken.Customers were instructed to follow the steps provided in order to address the issues related to microbial contamination; and to complete and return the enclosed response form within 10 days so that receipt of this important information could be confirmed. In addition, the letter requested customers share the information with their laboratory staff and retain the notification as part of your laboratory Quality System documentation.For customers who have any questions regarding this update, the letter indicated to contact Customer Support Center Call Center via the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/. They may also contact Beckman Coulter at (800) 854-3633 in the United States and Canada or through your local Beckman Coulter representative.Beckman Coulter will issue two additional recall notification ("UPDATE TO URGENT PRODUCT CORRECTIONSynchron LX20, LX20 PRO, LXi 725 Clinical Systems") beginning on 05/20/2011, one to Laboratory Directors and one to Physicians. The notices will be mailed via US Postal Service on Friday, May 20th. Customers who have subscribed to email notifications will receive an email containing the letters on Monday, May 23th. Both letter informs the Laboratory Directors and Physicians that Beckman Coulter has received reports of erroneous/aberrant sodium, potassium, chloride and calcium results that have been released from the laboratory. Electrolyte shifts can be low or high and may occur above, below or within the normal range.Laboratory Directors were instructed to immediately disseminate the Attachment to this letter (via notification/alert procedure appropriate to your |  | 1284 | 1284 units total for all products | Worldwide Distribution -- Australia, Belgium, Canada, China, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Hong Kong, India, Ireland, Italy, Japan, Republic of Korea, Kuwait, Mexico, Netherlands, Oman, Panama, Philippines, Puerto Rico, Saudi Arabia, Singapore, Slovakia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | February 06, 2013 | 673 | 1284 | Not_Computer
Z-2153-2011 | 58318 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Circumcision Clamp, models 1.1 cm and 1.3 cm, |  |  | January 26, 2011 | 2011 | Surgical Design Inc | Medical device used in surgical procedure is defective and may break during use. | Nonconforming Material/Component | Surgical Design, Inc. sent a Customer letter dated February 23, 2011, informing them of the affected product and providing instructions for its specific consignees. The instructions to the User is to please check your inventory to determine whether any Circumcision Clamps bought directly match those on your shelves. Please fill out the form and mail it to Surgical Design at the address on the form. If any clamps from the affected lots remain in your inventory, please ship them back to Surgical Design, noting the number being returned on the Reply form. These are the affected Lots.100210 (1.3cm)100211 (1.1cm) and100060409 (1.3cm)If you have any further questions, please call (703) 541-0196. |  | 6100 | 6100 clamps | Class 2 Recall -- Nationwide Distribution. | Terminated | June 21, 2013 | 877 | 6100 | Not_Computer
Z-2150-2011 | 58504 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON S2000, Model 10434730, ABVS ultrasound systems; |  |  | March 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | The weld connecting the ABVS arm to the system base has the potential to fail, causing the arm to fall off of the unit. | Process change control | Siemens Medical Solutions USA, Inc. initiated the recall via phone calls and e-mail on March 28, 2011, to all affected customers, informing them of the affected product and providing instructions on the recall. Three units S/N 250171, 250173, and 250177 were to be removed and shipped back to Siemens. Siemens Service personnel were to be dispatched to the customer sites to perform validated stress test on units S/N 250157, 250162, 250163, 250164, 250166, 250167, and 250169 at customer sites.For questions regarding this recall call 650-694-5993. |  | 10 | 10 | Worldwide Distribution - USA, Portugal, Germany, Russia, Estonia, Brazil, France, and Saudi Arabia. | Terminated | May 06, 2011 | 39 | 10 | Not_Computer
Z-2151-2011 | 58441 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Commander, LANTIS Commander 6.1, LANTIS 8.3 Commander, LANTIS System, LANTIS 6.1 Complete, LANTIS 6.1 Essential, LANTIS 8.3 Complete and LANTIS 8.3 Essential. |  |  | March 23, 2011 | 2011 | Siemens Medical Solutions USA, Inc | There is a potential safety issue in regard to data transfer between LANTIS OIS system and the hospital Information System (HIS) when the HL7 (ADT) interfact is used. | Other | Siemens sent an Urgent Medical Device Correction letter on March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the Safety Advisory Notice in their LANTIS owner manual and contact there Siemens Service Engineer or Application Specialist or contact Lantis4.healthcare@siemens.com.For questions regarding this recall 925-602-8157. |  | 47 | 47 | Nationwide Distribution. | Terminated | April 19, 2012 | 393 | 47 | Other
Z-2154-2011 | 58379 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Enzymatic Creatinine (ECREA) Flex(R) reagent cartridges, K1270 for the Dimension Vista(R) 500 System |  |  | June 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Results may be falsely depressed up to 0.4 mg/dL or falsely elevated up to 0.6 mg/dL across the assay range. | Device Design | Siemens sent an Urgent Field Safety Notice letter dated June 2010 to all Dimension Vista(R) ECREA customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to:Discontinue use of ECREA method and transition to the Dimension(R) Jaffe creatinine (CREA) method.To notify anyone to whom they may have distributed the product.Complete and return the Effectiveness Check Questionnaire attached to the letter by fax to (302) 631-8467.For questions regarding this recall call (302) 631-6299. |  | 1399 | 1,399 | Nationwide Distribution - including AL, AZ, CA, CO, DE, FL, GA, ID, IL, IA, MD, MA, MI, MT, NH, NJ, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI. | Terminated | December 19, 2013 | 1288 | 1399 | Not_Computer
Z-2156-2011 | 58323 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Brightview XCT, 3/4" |  |  | March 28, 2011 | 2011 | Philips Medical Systems | The current Headholder to support Bright View XCT brain imaging allowed for imaging the brain but doe allow for imaging of the associated shoulder/neck area of the body. | Device Design | Customer information sheets will be provided to all customers when the field change order commences in April 2011.Philips will implement a full change order for the installation of the new headholder to replace the old headholder, free of charge, to all affected Philips BrightView XCT SPECT/CT units.For questions regarding this recall call 408-463-3000. |  | 1 | 1 | Worldwide Distribution - USA including US distribution to OH, MA, FL, VT and KY and the countries of Norway, Saudi Arabia, Belgium, Brazil, Japan, China, Poland, Austria, United Kingdom, Spain, Indonesia, Malaysia, Taiwan, Germany, United Arab Emirates, Netherland, Switzerland, Ireland, Czech Republic and Canada. | Terminated | June 22, 2011 | 86 | 50 | Not_Computer
Z-2157-2011 | 58308 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029 |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: The handheld barcode scanner does not read barcode labels using NW 7 symbology.Impact: Specimens with NW7 labels cannot be read by the handheld scanner. | Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter with attached Recall Response Form sent via US Postal Service on February 11, 2011 to all customers who have the affected instruments. Consignees were informed of the reasons for recall, products affected and resolution for these issues would be corrected in a future release(s) for the DxH 800; but were also asked to take the following actions:If they use NW7 bar code symbiology, the labels can be read using one of the following methods:-At the single-tube presentation station (using the internal bar code scanner)-Enter the bar code information manually at the user interface of the single-tube station-Use cassette presentation.Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain this notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800526- 7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA (nationwide) and the countries of Australia ,Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | September 02, 2014 | 1299 | 566 | Not_Computer
Z-2179-2011 | 55466 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Inc. SSH-880CV Aplio Artida; Diagnostic Ultrasound System |  |  | April 20, 2010 | 2010 | Toshiba American Medical Systems Inc | There is a possibility of the system shutting down owing to the power supply unit causing the protection function to activate thus turning off the system prematurely. | Device Design | The recall was initiated on 4/20/10 with Toshiba America Medical Systems (TAMS) forwarding a Medical Device Correction with a Customer Reply Form (via US Postal Service for US Customers) to all customers who purchased the SSH-880CV Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid this problem until corrective action is taken, in case this problem occurs; please contact their Toshiba Representative. If emergency use is necessary, use the following procedure which allows the system to be used temporarily: 1. Turn OFF the breaker at the back of the system and confirm that the error code near the breaker disappears. 2. Turn ON the breaker again to supply power to the system. Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email to raffairs@tams.com. If customers had any questions they can call Paul Biggins at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 14 | 14 | USA, including Puerto Rico. | Terminated | May 25, 2011 | 400 | 14 | Other
Z-2178-2011 | 55952 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter Access Free T3 Calibrators, Part Number A13430 in vitro diagnostic. |  |  | December 17, 2009 | 2009 | Beckman Coulter Inc. | Beckman Coulter has confirmed reports that the Access Free T3 calibrator lot numbers 989902, 989903, 989904, 989905, 989906 and 989907 will show a decline in performance before reaching their stated expiration date.The affected calibrator lots will cause one or more quality control samples to produce out-of-range high result. | Nonconforming Material/Component | Beckman Coulter sent a letter dated December 29, 2009, providing the customers with an explanation of the reason for recall and asked them to take the following actions:- Discontinue use of the lots of reagent listed above.- Review your historical quality control to ensure that the assay performance of these lots has been within your laboratory specifications.NOTE: Out-of-range high quality control results may indicate invalid test results.Consignees were asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation; and complete and return the enclosed response form within 10 days so that Beckman can be assured that they received this important notification.Assistance or questions regarding this notification, were directed to TechnicalSupport at 1-800-854-3633 in the United States and Canada. Outside the United States andCanada, contact consignees were told to contact their local Beckman Coulter Representative. |  | 2320 | 2,320 kits | Distribution Nationwide and Canada. | Terminated | March 29, 2012 | 833 | 2320 | Not_Computer
Z-2169-2011 | 57446 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Statspin Express 4 Centrifuge |  |  | August 18, 2010 | 2010 | StatSpin, Inc., dba Iris Sample Processing | Rotors may fail due to wire spokes breaking | Device Design | Iris notified customers by letter in August 2010. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to install the new rotors and discard the older rotors with the wires. There was no need for customers to return the rotors, but the firm requested customers fax or e-mail the enclosed form to verify that they have installed the new rotors. For any questions customers should call 1-800-782-8774. |  | 2090 | 2090 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia Belgium, Czech Republic, Finland, France, Greece, Kuwait, Norway, Turkey,UK, and Norway. | Open, Classified |  |  | 2090 | Not_Computer
Z-2168-2011 | 57594 | [6 empty] | Philips FloTrak Elite modules: |  |  | January 04, 2011 | 2011 | Respironics Novametrix, LLC. | Software: two displayed respiratory parameters are incorrect due to errors in the system software: MValv (alveolar minute ventilation) and Vt/kg (tidal volume/patient weight). In both instances, the displayed parameters are higher than actual. | Software design | Respironics issued Field Action Notification Letter via Fed'x on 1/4/11, notifying customers of the issue and to instruct them on conditions for continued use and return of the device. The device will be updated with revised software to correct the calculation errors. Any questions or need additional information about this recall action, Contact Respironics US Customer Service at 1-800-345-6443 for customers residing in the USA or +011 (203)697-6488 for International customers. |  | 38 | 38 units | NationwideForeign: Belgium, China, Columbia, Egypt, England, India, Italy, Saudi Arabia, and Spain | Terminated | March 07, 2013 | 793 | 220 | Computer
Z-2164-2011 | 57472 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | S5-7606 Silicone Oil injection Cannula 6mm Long - 20G |  |  | December 17, 2010 | 2010 | Becton Dickinson and Company | Restricted or no fluid flow for Visitec High Viscosity Injector. | Component change control | Beaver Visitec (BVI) notified customers via telephone and by sending an Urgent Recall Customer Notification letter dated December 7, 2010. The letter identified the product, the problem, and the action to be taken by the customer. ,Customers were requested to return all unused product in their inventory for credit or refund and complete the attached Customer Response Form and return it by fax to 781-893-7957. Questions can be directed to Inside Sales Manager at 781-906-7952. |  | 1000 | 1000 units | Worldwide Distribution - USA (nationwide) and the countries of Israel, Jordan, Russia, Saudi Arabia, South Africa, Turkey, Iraq, Lebanon, Morocco, Tunisia, Syria, India, Argentina, Australia, Austria, New Zealand, China, UK, Germany, Russia, Belgium, Italy, France, Sweden, and Ireland. | Terminated | July 03, 2013 | 929 | 19129 | Not_Computer
Z-2172-2011 | 58234 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, |  |  | February 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because: Specimen tubes may fall out of a Type A cassette (PN A53551, kit of 20) when the analyzer inverts the cassette in the mixing station, or when an operator manually inverts a cassetteThe impact:- Specimen tubes may fall out of cassettes and be misplaced inside the instrument- Specimen tubes may break upon impact, causing possible biohazard exposure- Patien | Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter with attached Recall Response Form being sent via US Postal Service on February 11, 2011 to all customers who have the affected instruments. Customers were informed of the reasons for recall, products affected and resolution for these issues would be corrected in a future release(s) for the DxH 800; but were also asked to take the following actions:-Periodically inspect the mix station drip tray for misplaced tubes that may have fallen from the cassette.- Periodically examine the Type A cassettes to determine if the expandable grippers (clips) within the lavender-colored insert are stuck in the open position (tubes will be loose). Consignees were told to contact their Beckman Coulter representative for assistance with questionable grippers.- Clean any spills according to their laboratory standard operating procedures. Consignees were requested to complete and return the enclosed Response Form within ten (10) days and to share this information with their laboratory staff and retain the notification as part of their Quality System documentation. If consignees have transferred ownership or location of the analyzer(s) to another laboratory, they were requested to provide a copy of the letter to that party. For questions concerning this notice, consignees were told to contact Beckman Coulter Customer Service at 800-526- 7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 566 | 566 units | Worldwide Distribution - USA (Nationwide)and to the countries of Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Turkey, &amp; United Kingdom | Terminated | November 21, 2012 | 649 | 566 | Not_Computer
Z-2165-2011 | 58060 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 4000 |  |  | February 14, 2011 | 2011 | Instrumentation Laboratory Co. | Measurement (amperometric) spike can occur early in cartridge life on the glucose and lactate sensors during patient blood analysis, leading to erroneously highresults | Device Design | Instrumentation Laboratory notified accounts by a Field Safety NotifIcation Letter dated Feb 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use whichever of the following mitigation options works best for their institution. Both options will have the same effect to eliminate the potential impact of a rare spike on the glucose and lactate sensors. Customers were instructed to complete the enclosed Response Tracking Form and return immediately fax to 781-861-4207 or e-mail to ra-usa@ilww.com.Customers were advised to contact their local representative with any questions. |  | 4379 | 4379 units | Worldwide Distribution - USA (nationwide) and the countries of CANADA ISRAEL ITALY APAN KUWAIT MALAYSIA MEXICO NETHERLANDS NORWAY PANAMA POLAND PORTUGAL PUERTO RICO SAUDI ARABIA SLOVENIA SPAIN SWEDEN SWITZERLAND TAIWAN TURKEY UNITED ARAB EMIRATES UNITED KINGDOM and URUGUAY. | Terminated | May 23, 2012 | 464 | 4379 | Other |  |  | Other
Z-1678-2011 | 58047 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | 2.0 mL Cartridge/Syringe Piston component of Insulin Pump |  |  | February 25, 2011 | 2011 | Animas Corporation | Action is being undertaken to address the occurrence of adverse events associated with leaking cartridges. Firm became aware of this issue through customer complaints from pump users reporting the cartridges were leaking. | Material/Component Contamination | Animas Corporation sent an Urgent Recall letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Change their cartridge immediately using the instructions provided. Return all of the unused affected product using the enclosed box. Shipments must be returned using the United States Postal Service. For technical questions call Animas Customer Support at 1-855-254-5668 or call the toll free number at 1-8777-937-7867.For any questions about returning the product call 1-877-280-2339.Customers were instructed to notify the firm by logging on to the following website: www.animasnotification.com or call the Call Center at 1-877-280-2339 if they have no product to return. |  | 384180 | 384,180 | Worldwide Distribution - USA (nationwide) and France. | Terminated | April 03, 2014 | 1133 | 384180 | Not_Computer
Z-2182-2011 | 58466 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | KCI RotoProne Critical Care Therapy System buckles for securing patient to the surface during proning therapy, Model 209500. |  |  | February 18, 2011 | 2011 | KCI USA, Inc. | Buckles used to secure patient to surface during proning therapy can come open if not properly engaged by user. | Device Design | Firm sent a letter with the "Warning: Improper Insertion of the RotoProne Therapy System Buckle Can Lead To Unintentional Bed Exit" to all consignees on 2/23/2011. The letter identified the affected product and the reason for the recall. Customers were instructed to read the attached product warning regarding the proper method to attach the RotoProne buckles and to post it with the device to ensure continued proper use of the buckles. If customers had questions regarding the proper use of the buckles, they were to contact KCI Customer Technical Service at 1-800-275-4524. |  | 229 | 229 units | Worldwide Distribution - USA and Kuwait. | Terminated | October 24, 2011 | 248 | 229 | Not_Computer
Z-2176-2011 | 58252 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, version 2.6 and 2.6.1; Reference/FSCA Identifier: CP-03976; Model Number: HZ1; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The anomaly that has been identified with the Optical Guidance Platform (OGP) software may not be always displaying the correct transfer date on the patient file. | Software design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated March 3, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the Optical Guidance Platform (OGP) Instructions For Use on Loading a treatment plan; before positioning the patient, compare the plan details loaded into the OGP with the plan details printed from the treatment planning systems; verify that the patient information, isocenter, gantry and couch angles are identical as requested in the Instructions For Use, and advise the appropriate personnel working in their radiotherapy department of the content of this recall letter.Varian has made changes to the OGP Instructions For Use -Varian strongly recommends you download the most recent version at www.MyVarian.com or request a copy from your local Varian Customer Support District or Regional Manager. Varian also is preparing a software correction which will be applied to affected unites to correct this issue.If you require further clarification, please feel free to contact your local Varian Customer Support District or Regional Manager at the follow: USA and Canada - 1-888-827-4265; Europe +41 41 749 8844; E-mail-North America: support-americas@varian.com; Australia/New Zealand: support.anz@varian.com; Europe: support-emca@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com, and Latin America: soporte.al@varian.com; Internet: Oncology Systems customer site - www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 599 | 599 units | Worldwide distribution. | Terminated | June 19, 2012 | 474 | 599 | Software | Output/Calculation | Instructions for verification/Software update | Software Update
Z-2175-2011 | 58182 | No | NSX |  |  | General Hospital | Contact ODE | Philips IntelliVue Clinical Information Portfolio software as follows: |  |  | March 16, 2011 | 2011 | Philips Healthcare Inc. | Edits to an order can result in unintended scheduled interventions and a patient may receive additional interventions or medication doses that are not intended. | Software design | Philips Healthcare notified accounts on 3/16/11 via the Urgent Medical Device Correction Notification. Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel.All customers potentially affected by this issue will receive the software correction as part of the field action. |  | 387 | 387 units | Worldwide distribution: USA, Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Italy, Kenya, Latvia, Malaysia, Netherlands, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 387 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2180-2011 | 58395 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, DigiView and ERS software release 3.6.4.4 through 3.6.5.2+, 2008 Digisonics, Inc., 3701 Kirby Drive, Houston TX 77098, 800-940-3240. |  |  | January 20, 2009 | 2009 | Digisonics, Inc | Data from one patient's study had the potential for being put into another patient's report. | Device Design | Digisonics notified end users via e-mail January 20, 2009 and January 21, 2009, that Data from one patient's study was being put into another patient's report. The problem occurred when a System bug was found in DigiView Release 3.6.4.4, and higher versions in which under certain circumstances (specific workflow) a failure to switch to a patient's images occurs. Two software patch versions will be made available to sites on or after February 5, 2009, to correct the reported problem.If you have any further questions, please contact Digisonics Customer Support at (800 ) 940-3240. |  | 70 | 70 units | Worldwide Distribution -- USA including States of : OH, NC, LA, TX, CA, MO, NY, MD, MA, DC, IA, AZ, PA, MI, NC, FL, IN, SC, ID, MN, AR, OK, IL, CT, and countries of Canada, Lebanon, and New Zealand. | Terminated | May 31, 2011 | 861 | 70 | Software | Output/Calculation | Software update | Software Update
Z-2166-2011 | 58400 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DigiView software release 3.4.4.4 through 3.5.5.2. , 2008 Digisonics, Inc, 800-940-3240. |  |  | November 04, 2008 | 2008 | Digisonics, Inc | Failure of software to update some measurements when tracings or points are amended after measurement calculations have been performed. | Device Design | Digisonics, Inc., notified consignees by telephone and e-mail on 10/17/2008, informing them that a problem had been found in the Save Linked Images functionality of DigiView Release 3.6 (and all releases which start with 3.6 such as 3.6.1.). The problem is the failure of the software to update some measurements when tracings or points are amended after measurement calculations have been performed. A utility program, called Linked Image Measurement Verify was being supplied to all sites as a patch to the DigiView program. Release 3.6.6 was to be provided to all customers having previous 3.6 version. Problem was fixed in version 3.6.6 and 3.7.1.If you have any questions, please contact Digisonics Customer Support at (800) 940-3240. |  | 70 | 70 units total | Worldwide Distribution -- USA including States of : TX, OH, NC, MA, NY, IL, FL, OK, MO, ID, AZ, AR, MI CA, MD, and country of Canada.*** Due to the age of this recall, district is recommending no audit checks*** | Terminated | May 31, 2011 | 938 | 70 | Software | Output/Calculation | Software update | Software Update
Z-2181-2011 | 58590 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical System TrueBeam system for stereotactic radiosurgery and radiotherapy. Model number H19; |  |  | April 05, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Position sensor failure mode may result in an inaccurate position calculation. 1. The video returned by one of the two cameras inside the Spectra is all white or all black, and the Spectra stops tracking. Or 2. Certain video intensities are not available on one of the two cameras. The image has abnormally abrupt transitions from dark to light portions, without the normal shades of gray. | Nonconforming Material/Component | Varian Urgent Medical Device Correction, Urgent Field Safety Notice letters were sent to all affected facilities by certified mail on April 5, 2011. |  | 20 | 20 | Nationwide distribution. | Terminated | June 19, 2012 | 441 | 20 | I/O | Output/Calculation | Notification | Safety Notice/Insructions
Z-2170-2011 | 58168 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Beckman Coulter HbA1c APT (Hemoglobin A1c, Whole Blood Application), Part Number: OSR61177, Lot Numbers: 9390 &amp; 9472 |  |  | March 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a problem with two lots of HbA1c APT reagent OSR61177.When HbA1c APT reagent lots 9390 and 9472 are calibrated with HbA1c calibrator ODR3032,elevated recovery may be observed for patient samples. Falsely elevated HbA1c values may beinterpreted as increased risk for diabetes or poorly controlled diabetes. Consequent medicalacti | Pending | Beckman Coulter initiated the recall communication with a Product Corrective Action (PCA) letter with attached recall response form on March 1, 2011 to the affected customers. The letter explained the problem and identified the action to be taken. The PCA also instructed the consignees to:(1) Please discontinue the use of lots 9390 and 9472 and discard according to your local procedures.(2) Please contact your local Beckman Coulter support Organization for replacement reagent.(3) You may need to re-evaluate samples for HbA1c which were tested with either of the listed lots of reagent.In addition, customers were instructed to:(1) share the information with their laboratory staff and retain this notification as part of their Quality System documentation. If any of the affected product was forwarded to another laboratory, a copy of the letter should be provided to them.(2) complete and return the enclosed response form within 10 days so that the firm was assured the customer received the important communication. For any questions regarding this notice, customers were instructed to contact Customer Technical Support at 1-800-854-3633 in the U.S.A and at 00353656831111 in Europe. |  | 373 | 373 kits | Worldwide Distribution - USA (Nationwide) Austria, Croatia, France, Germany, Korea, Philippines, Russian Federation, Thailand, Venezuela, and Vietnam. | Terminated | June 28, 2012 | 485 | 373 | Not_Computer
Z-2177-2011 | 58380 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension(R) EZCR Flex(R) Reagent Cartridge. Catalog number DF270B. In vitro diagnostic test for the quantitative measurement of creatinine in human serum, plasma and urine on the Dimension Vista(R) System. |  |  | June 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Patient results may be falsely depressed or falsely elevated across the assay range. | Device Design | Siemens issued an Urgent Field Safety Notice dated June 2010 to all Dimension(R) EZCR customers; to notify them that Siemens confirmed an issue which many cause EZCR results to be falsely depressed or falsely elevated across the assay range. The letter asks users to discontinue use of this lab test kit. |  | 860 | 860 | Nationwide distribution. | Terminated | March 19, 2013 | 1013 | 860 | Not_Computer
Z-2171-2011 | 58216 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Ventricular Bolt Monitoring Kit W/Cranial Access, Drain, 1104HMC, Integra LifeSciences, 3498 West 2400 South, West Valley City, Utah 84119. |  |  | February 22, 2011 | 2011 | Integra LifeSciences Corp. | Outer box of Monitoring Kit was mislabeled with the incorrect part number. | Error in labeling | Integra LifeSciences Corp. sent an Urgent Product Recall Notification letter dated February 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine any affected product and arrange for return to Integra by contacting Integra Customer Service Returns at 1-800-654-2873 (choose option 2) to obtain a Return Merchandise Authorization number and place an order for replacement product. Customers were asked to complete the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified. For any questions regarding this recall call 609-936-2485. |  | 45 | 45 kits | Nationwide Distribution including CA, CT, GA, HI, IL, KY, MA, NC, NJ, NM. NV, OH, TX, WA, and WI. | Terminated | May 10, 2011 | 77 | 45 | Not_Computer
Z-2173-2011 | 58180 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NEXGEN COMPLETE KNEE SOLUTION FEMORAL AND PROVISIONAL IMPACTOR/EXTRACTOR, NONSTERILE, REF 00-5901-026-00, Zimmer UK, Ltd., Zimmer Warsaw, IN. |  |  | March 14, 2011 | 2011 | Zimmer Inc. | The firm is initiating a recall due to the potiential for the spring clip to break off during use and fall into the surgical site. The firm has received 24 complaints of the spring breaking, four of which were reported that the patient had to undergo additional x-rays to ensure the fractured clip had not been left in-vivo. | Device Design | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" notice dated March 14, 2011 to its Consignees/customers. The letter described the product, problem, and the actions to be taken. The customers were instructed to stop using the device and quarantine immediately; carry out a physical count of all affected product in their territory and record this data on the Inventory Return Certification Form and fax to (574) 372-4265; ensure they have provided acknowledgement via "email return receipt" upon receiving notification; return the recalled product along with the completed Inventory Return Certification Form and Sterilization Certification Form to: Zimmer Distribution Center, Attn: Product Service, 1777 West Center Street, Warsaw, IN 46580-Note: The instrument must be returned even if the spring clip is not missing from the instrument; and if they distributed these affected lots further, provide a copy of this letter to the those customers when they implement these recall instructions. The letter stated that a new or refurbished instrument with a redesigned spring clip will be provided as replacement for the recalled device. For shipping assistance, questions or assistance in notifying your accounts about the recall, please contact Zimmer, Inc. at 1-800-613-6131. |  | 802 | 802 devices | Worldwide distribution: USA (nationwide)including states of: HI, CA, NY, MI, IN, GA, FL, WI, MN, NJ, IL, VA, TX, NC, PA, MS, WA, OR, OH, KS, TN, NH, AZ, NV, UT and AK and countries of : Australia, Canada, China, Germany, India, Japan, Singapore and Sweden. | Terminated | August 28, 2013 | 898 | 829 | Not_Computer
Z-2184-2011 | 58583 | No | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IGFBP-3 ELISA Reagent Kit; Part Number" DSL-10-6600; Lot Numbers: 992743, 992287, 991916, 991530, 991432, 990220 &amp; 890830 |  |  | March 30, 2011 | 2011 | Beckman Coulter Inc. | The recall was was initiated because patient results with the IGFBP-3 ELISA kit may be elevated by 70% compared to stated expected ranges in the Instructions for Use (IFU). | Process control | Beckman Coulter sent a Product Corrective Action (PCA) letter with attached customer response form on March 30, 2011 to all customers who have the affected product, IGFBP- 3 Enzyme-Linked immunosorbent Assay ( ELISA) kit .This kit provides materials for the quantitative measurement of IGFBP-3 in serum. The letter informed the customers of the reason for recall, the products affected and actions to be taken.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. If they forwarded any of the affected product listed above to another laboratory, they were asked to provide a copy of the letter to them.Beckman further asked consignees to complete and return the enclosed Response Form within 10 days so they are assured consignees have received this important communication.Questions were directed to the Beckman Coulter Customer Support Center.- Website: http://www.beckmancoulter.com/customersupport/support- Phone, call 1-800-854-3633 in the United States and Canada- Outside the United States and Canada contact local Beckman Coulter representative. |  | 1212 | 1,212 units | Worldwide -- USA, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, France, Germany, Greece, Hong Kong, India, Italy, Korea, Republic of, Lebanon, Luxembourg, Morocco, Netherlands, New Zealand, Puerto Rico, Syrian Arab Republic, Taiwan, Thailand, Turkey &amp; United Kingdom | Terminated | December 02, 2011 | 247 | 1212 | Not_Computer
Z-2183-2011 | 58479 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | TEMPUS IC PROFESSIONAL Patient Monitor, Part No. 00-1002, Manufactured by: Remote Diagnostics Technologies Ltd., The Old Coach House, Farleigh Wallop, Hants, UK RG25 2HT. |  |  | March 25, 2011 | 2011 | Remote Diagnostic Technologies Ltd. | It was discovered that when a 12 Lead ECG is recorded and monitoring is quickly restarted, the ECG can cease functioning. | Component change control | Remote Diagnostic Technologies, Ltd notified its consignees of the problem and the work-around via letter with the subject line "ECG Issue - Field Action Required" on 03/25/2011. The letter stated that customers can avoid the issues by viewing the recorded ECG for at least 13 seconds. Two diagrams were attached with the letter depicting the conditions under which the issue can occur. Customers can continue to use the device normally in all other respects. The letter states that the firm is working on a software modification which will permanently fix the issue and prevent it from occurring. Customers are to contact the Regulatory Affairs and Operations Manager if they have any questions. |  | 19 | 19 units | Nationwide Distribution -- KY, MD, VA, and FL. | Terminated | September 16, 2011 | 175 | 19 | Software | Device Operation | Software update | Software Update
Z-2188-2011 | 57194 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial Blood parameter module Catalog number 500A , Terumo Cardiovascular Systems Corporations, Ann Arbor, MI. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The CDI 500 monitors may display inaccurate values. These errors include inaccurate potassium, CO2, hematocrit, O2 saturation, pH, hemoglobin and other miscellaneous inaccurate values. | Device Design | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 27, 2010 to its customers. The letter described the product, problem and actionst to be taken. The customers were instructed to reveiw this notice; assure that customers are ware of this notice; confirm receipt of this notification by faxing the attached Customer Response Form, and place the attached stickers in a visible place in the Operators' Manuals, such as the inside of the front cover, if additional stickers are needed, the customers should contact Terumo CVS Customer Service at 1-800-521-2818. The letter states that when a displayed value is significantly different from expectations based on the clinical situatioin, the user should verify the accuracy by independent means before treating. The letter included "Caution" stickers that were to be applied to the CDI system operatros manual.The letter also states "Please note that we are not recommending that you stop using your CDI monitor system."If you have any questions or concerns, please contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are from Monday-Friday, 8AM - 6PM EST. |  | 3260 | 3260 all products | Worldwide distribution: USA (nationwide) and countries of : Australia, Barbados, Belgium, Brazil, Columbia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Guatamala, Hong Kong, Honduras, Jordan, Japan, Libya, Mexico, Malaysia, Panama, PH, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad, Taiwan and Uruguay, Germany, Saudia Arabia, Paraguay, South Korea, Dominican Republic, Phillipines, South Africa and UAE. | Terminated | July 31, 2014 | 1368 | 3260 | Other | Display/Image | Instructions | Safety Notice/Insructions
Z-2190-2011 | 58517 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE / RT Therapist v4.1 system. |  |  | March 24, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When using the ARTISTE Linac and RT Therapist v4.1 along with an imported cropped reference image, the image maybe processed incorrectly, thus resulting in a mistreatment. | Pending | Distribution of the Urgent: Medical Device Correction Customer Advisory Notice began on March 24, 2011 via certified mail or by the Siemens Customer Service Group. The letter identified the issue and when it occurs. Users were asked to verify the isocenter position of the reference image in the Portal Review application before using a reference image imported from the treatment planning system. Siemens is currently working on a solution and will notify customers as soon as the update is available. The letter should be included with the Owner Manual of the system in the chapter for "Safety Advisory Letters" until an update is provided. |  | 228 | 228 active devices | Worldwide Distribution | Terminated | February 13, 2012 | 326 | 228 | Software | Display/Image | Software update | Software Update
Z-2191-2011 | 58219 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Forteo [teriparatide (rDNA origin) injection] black starter kits containing Triad alcohol pads. The starter kits look like insulated lunch boxes and contain alcohols pads, needles, instruction guides, and freezer gel packs. Note: the kits do not contain the drug, Forteo. |  |  | March 17, 2011 | 2011 | Eli Lilly and Company | Eli Lilly and Company announced 3/17/2011 that patients should not use the alcohol prep pads made by the Triad Group that are contained in the black starter kits for Forteo [teriparatide (rDNA origin) injection] in the United States.The Tri-ad Group is recalling the alcohol prep pads due to potential contamination with the bacteria, Ba-cillus cereus, which could result in life threatening infecti | Material/Component Contamination | The firm, Eli Lilly and Company, sent two letters entitled "URGENT MEDICAL DEVICE RECALL TRIAD GROUP ALCOHOL PREP PADS INCLUDED IN FORTEO STARTER KITS ALL ALCOHOL PREP PAD LOTS" to its customers. The first letter was dated March 18, 2011 and second follow-up letter was dated March 28, 2011. The letters described the product, problem and actions to be taken. The customers were instructed to examine their inventory of Forteo starter kits; if they have any Forteo starter kits with a black bag, remove and appropriately discard the Triad Group alcohol prep pads; discontinue providing the product with the starter kits; inform any patients that may have received Forteo starter kits containing the Triad Group alcohol prep pads to discontinue use; note: after removing the Triad Group alcohol prep pads, all other components of the Forteo starter kits are acceptable for use: travel bag, pen case, gel packs, needles and instructional materials, and complete and return the enclosed business reply card. "It is important to note that this recall does not affect or involve the Forteo delivery device."For questions pertaining to the recall, contact Triad Group at 262-538-2900. (Please refer to the Triad Group press release issued on January 5, 2011 on the FDA website.) If you have product specific medical questions related to Forteo please call 1-800-LILLYRX. |  | 47000 | approximately 47,000 kits | Worldwide distribution: USA (nationwide) and countries of: Czech Republic, Germany, Slovak Republic, and Switzerland. | Terminated | August 30, 2011 | 166 | 47000 | Not_Computer
Z-2192-2011 | 58451 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Nokor Admix Needle |  |  | March 01, 2011 | 2011 | Becton Dickinson &amp; Company | Incorrect unit package labels which reflect Reorder #305216 instead of the correct item, Reorder #305213, were used. | Packaging | BD sent an Urgent: Product Recall letter with return response forms to direct consignees on 03/01/11 and 03/02/11 via UPS. The letter identified the affected product and lot along with the reason for the recall. The letter asked customers to complete the enclosed Distributor Packing List and to follow the instructions provided on the Packing List for returning the affected product. The firm will issue a credit upon receipt of the affected product. Customers can contact BD Customer Service if they need assistance or have questions regarding the recall. |  | 161000 | 161,000 units | Worldwide Distribution -- USA, including the states of NJ, IL, LA, NM, OH, NV, and VA, and country of Canada. | Terminated | November 30, 2012 | 640 | 161000 | Not_Computer
Z-2193-2011 | 58421 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digisonics, Digiview and ERS software releases 3.6.44, 3.6.4.6, and 3.6.5, Digisonics, Inc., 3701 Kirby Drive, Houston, TX 77098, Tel: (800) 940-3240, Fax: (713) 529-7999 |  |  | February 06, 2009 | 2009 | Digisonics, Inc | Left ventricle volume was inaccurate when the volumes were measured from images having different scales. | Other | The firm, Digisonics, notified consignees/customers via e-mail entitled "Distribution of Notification of Biplane Error" dated February 6, 2009. The email described the product, problem and actions to be taken. The customers were instructed to immediately contact Digisonics if their site calculates biplane volumes-a utility will be provided which will run on their database and identify all studies in which biplane volumes are calculated. The customers were also instructed to share the following important information with all users of DigiView at their site. The email stated -until the software patch has been installed, biplane volumes should be calculated only from pairs of images having the same depth scale. Replacement DigiView/ERS software that prevents further occurrence of this problem will be provided if it appears their site could be impacted.The Distribution of New Software Release was sent on March 6, 2009.If you have any questions, please call Digisonics Customer Support at 800-344-4266 |  | 74 | 74 units | Worldwide distribution: USA (Nationwide) and countries of: Canada, Lebanon and New Zealand. | Terminated | May 31, 2011 | 844 | 74 | Software | Output/Calculation | Software update | Software Update
Z-2198-2011 | 58198 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Curad Flex-Fabric/Tissu Souple, 3/4 x 3 in., 100 sterile adhesive bandages, Made in China, Manufactured for Medline Industries, Mundelein IL 60060 |  |  | February 22, 2011 | 2011 | Jiangsu Wujin Hualian Health Dressing Co., Ltd. | The sterility of the product has the potential to be compromised due to weak sealing of the bandages. | Process control | Changzhou Hualian Health Dressing Co., Ltd notified sole consignee by e-mail and a Medical Device Recall letter, dated 3/18/2011. The letter identifies the affected product and the reason for the recall. The customer was to examine their inventory and quarantine the affected product. If the product had been further distributed, those customers were to be identified and notified of the recall. The enclosed response form was to be completed and returned to Medline Industries, Inc. Questions could be directed to Customer Service, Monday through Friday between 8:30 A.M. and 4:00 P.M. |  | 1092 | 1,092 cases | Nationwide Distribution -- IL. | Terminated | April 23, 2012 | 426 | 1092 | Not_Computer
Z-2195-2011 | 58438 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Medical Linear Accelerator, All Varian Clinac with Model Numbers: H14, H18, H26, H27, H28, H29, Hcx; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 28, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly whereby, following prolonged use, the screw fastener holding the wedge body to the tray may fail. | Process control | The firm, Varian Medical, sent an "Urgent Medical Device Correction" Letter dated March 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately inspect all hard wedges for missing or sheared screw fasteners, if any wedge is found to have one or more missing or sheared screw fasteners, remove the wedge from use and notify Varian immediately; inspect all hard wedges for loose screw fasteners, and retighten as necessary, and implement a program of monthly inspection of wedges for loose or missing hardware. The customers were also instructed to advise the appropriate personnel working in their Radiotherapy Department of the content of this letter.Varian Service shall immediately begin an interim action to replace screw fasteners on all hard wedges for Clinacs installed prior to January 2004. Varian Service will contact the customers to arrange for the inspection visit.Varian also stated that they shall continue their Investigation of this issue to further clarify the exact conditions under which the wedge screw fasteners may fail, and to analyze possible final solutions. Varian will notify the customers when final determinations are made.If you require further clarification, please feel free to contact your local Varian Customer Support Services Manager and/or USA and Canada at 1-888-4265; Europe: +41 41 749 8844; or email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.Helpdesk@varian.com; Japan: Japan.apps helpdesk@varian.com; Latin America: soporte.al@varian.com, a d Internet: Oncology Systems-www.myvarian.com or Varian Medical Systems - www.varian.com. |  | 5616 | 5616 units | Worldwide distribution: | Terminated | June 07, 2013 | 802 | 5616 | Not_Computer
Z-2197-2011 | 58167 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Manta Ray" Anterior Cervical Plate (ACP) System, |  |  | January 10, 2011 | 2011 | Integra LifeSciences Corp. | Based upon the results of the firm's literature review, they determined that the Package Insert (Information for Use) and information in the Surgical Technique should be revised to include additional warnings, precautions and possible adverse events. | Pending | The firm, Integra Spine, sent an "URGENT: FIELD CORRECTIVE ACTION" letter dated January 10, 2011 with a Field Corrective Action Acknowledgement and Return Form sent via Fed Ex/ priority mail to all customers including physicians and distributors. The letter describes the product, problem and actions to be taken. The customers were instructed to review the current inventory to identify product from the affected part numbers; replace the existing IFU (22-PI-01 rev C) with the revised IFU (22-PI-01 rev D) and the existing Surgical Technique (ST-22-001 rev A) with the revised Surgical Technique (ST-22-001 rev B), and complete and return the Field Corrective Action Acknowledgement and Return Form via fax to: 877-558-6227 or email.If you have any additional questions, contact the Manager Quality Systems Integra Spine at 330-475-8646. |  | 244 | 244 IFU | Nationwide distribution: USA including states of: AL, AZ, CA, CO, FL, GA, ID. IL, KS, LA, MD, MO, NH, NM, NV, NY, OH, OK, OR, PA, PR, SD, TN, TX, UT, VA, WI, and WV. | Terminated | October 27, 2011 | 290 | 244 | Not_Computer
Z-2196-2011 | 57951 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Aortic Perfusion Cannula with Duraflow Coating, REF: DARH201190TA, 20 Fr. x 28 cm, Sterile EO, Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | January 21, 2011 | 2011 | Edwards Lifesciences Llc | Bonding of the tip to the aortic cannula tubing may decline over time and may contribute to a drop in bond strength. Possible tip separation. | Other | The firm, Edwards Lifesciences, sent an "URGENT PRODUCT RECALL - ACTION REQUIRED" letter dated January 21, 2011 it customers via Fed-Ex. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory for any of the lot numbers; return all unused product from their stock to Edwards; contact their Customer Service Organization at (800) 424-3278 to obtain a Returned Goods Authorization number and replacement product; complete and return the Aortic Perfusion Cannula Recall Response form via fax to: Edwards at (949) 250-3489, and return unused product to Edwards Lifesciences, 12050 Lone Peak Drive, Draper, UT 84020. If you have questions that have not been answered by this letter, call Edwards Customer Service at (800) 424-3278 from 8:00AM - 4:30PM Pacific Time. |  | 20 | 20 | Nationwide distribution: CA only. | Terminated | May 13, 2011 | 112 | 20 | Not_Computer
Z-2194-2011 | 58490 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra NeuroSciences, Combo Kit, Lot 171237, Integra LifeSciences 3498 West 2400 South, West Valley City, UT 84119. |  |  | February 22, 2011 | 2011 | Integra LifeSciences Corp. | Outer packaging of Combo Kit was mislabeled with an incorrect lot number. | Labeling mix-ups | Integra LifeSciences Corp. sent a " Urgent Product Recall Notification" letter dated February 22, 2011, to the only affected consignee. It was reported the Combo Kits of SPOO98 / INS- 4500 ( Hermetic Large Style Ventricular Catheter set and drill set for NY) - Catalog # 31046 Lot Number 171237, Outer packages were labeled incorrectly. The customer was advised to return this, and they would be replaced with correctly labeled product. The issue is easily recognizable, and does not pose a safety risk to patients.An acknowledgement form was included to be returned to Integra. Further information is available at 609-936-2485. |  | 7 | 7 Kits | Nationwide Distribution -- including the state of New York. | Terminated | May 13, 2011 | 80 | 7 | Not_Computer
Z-2064-2011 | 58160 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Defibtech Lifeline AED and ReviveR AED, |  |  | April 01, 2011 | 2011 | Defibtech, LLC | Software with version 2.004 software or earlier may cause an affected AED to cancel shock during the charging process | Software design (manufacturing process) | The firm, Defibtech, sent an "URGENT MEDICAL DEVICE SAFETY INFORMATION AND CORRECTION" letter dated March 11, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Remove the pads from the storage pouch on the back of their DDU-100 series AED and inspect the 9-digit serial number label.2)Carefully note the 9-digit serial number on the back of the AED, then visit the Defibtech website at www.defibtech.com/fa11 and enter their serial number to determine if their AED is affected; email their serial number and contact information to fa@defibtech.com or call Defibtech directly at 1-877-453-4507 or 1-203-453-4501 and Defibtech will inform them if their unit is affected.3) If their AED is affected, confirm receipt of this notification with Defibtech as soon as possible by website and/or confirmation card. Complete and return the confirmation form via fax to 1-203-738-1072. Note: "it is very important for you to contact us either through the Defibtech website or return this card to confirm that you have received this notice. This information will expedite the process of upgrading your AEDs."If your DDU-100 series AED serial number DOES NOT fall within the listed range, your AED is not affected by this recall.A letter was also sent to the distributors and direct accounts that have received the affected product by first class mail with a prepaid, return written acknowledgement card. The field action consists of a software change to the AED. All affected AEDs operating with software version 2.004 or earlier are to remain in the field and in service throughout the field action process.Upon receipt of contact information from affected customers, Defibtech will send an upgrade kit to customers consisting of a software data card, instructions on how to perform the upgrade, and a pre-paid upgrade confirmation postcard that must be returned to Defibtech upon completion |  | 91472 | 91, 472 units | Worldwide distribution: USA (nationwide) and countries including: AUSTRALIA, AUSTRIA , BAHRAIN, BRAZIL, CANADA, CHILE, CHINA , CYPRUS, CZECH REPUBLIC, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA, REPUBLIC OF LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND URUGUAY. | Open, Classified |  |  | 91472 | Software | Device Operation | Software update | Software Update
Z-2200-2011 | 58505 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Optical Guidance Platform, Version 2.6 and 2.6.1, Model Number: HZl, Manufactured and Distributed by: Varian Medical Systems Inc., Palo Alto, CA. |  |  | March 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Optical Guidance Platform (OGP) Software v2.6 and v2.6.1 where the transfer of datasets from treatment planning systems other than the FastPlan system result in a lateral offset error. | Software design | On 3/30/2011, Varian distributed an Urgent Medical Device Correction Urgent Field Safety Notice to all affected users, with a description of the problem and identification of the affected product. The letter also informed customers of how they can determine if the problem affects them. The letter states that there is no work around for the issue unless they possess a FastPlan system. Varian is preparing a software patch to correct the issue and it will be applied to all OGP v2.6 and v2.6.1 systems. Customers are to advise the appropriate personnel of the content of the letter. If further clarification is needed, customers are to contact their local Varian Customer Support District or Regional Manager. |  | 598 | 598 units | Worldwide Distribution | Terminated | June 19, 2012 | 447 | 598 | Software | Output/Calculation | Software update | Software Update
Z-2199-2011 | 58596 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | The Operating Manual contains the Operating Instructions for AS 104 Blood Cell Separator and has a Software Version 4.71. The AS 104 Blood Cell Separator Device has a Catalog Number 9007031. |  |  | April 19, 2011 | 2011 | Fresenius Kabi, LLC | Operating Manual of the AS 104 Blood Cell Separator is being updated to eliminate any reference to Red Blood Cell (RBC), C4Y, P1Y and BMSC sets. This is a corrective action and recall expansion of Z-0616-2011. | PMA | Fresenius Kabi sent a "Urgent Notification" letter dated May 6, 2011, to the affected consignees they are recalling the Operating Manual for the Fresinius AS104-Blood Cell Separator.The Red Blood Cell (RBC) sets have been recalled according to recall number Z-0616-2011, and will permanently not be used with the AS 104 Blood Cell Separator Device. The RBC set was never approved by the FDA according to Warning Letter (Reference # GEN1001508) dated 31 January 2011. The customers are being instructed to complete the self addressed postage pre-paid " Product Response Card." Place a check mark next to the disposition choice used.Fresenius plans to send the revised Operating Manual and a notification letters to consignees at the end of April 2011.Consignees who have questions about this recall can call the firm at 1-425-242-2074. |  | 158 | 158 units | Nationwided Distribution -- AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MI, MN, NC, NH, NJ, NM, NY, OH, PA, SD, TN, TX, VT, and WI. | Terminated | December 06, 2011 | 231 | 158 | Not_Computer
Z-2212-2011 | 55207 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate TM Sample Processing System, P/N A20745 Basic Assembly - Automate |  |  | May 07, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an issue with the AutoMate Sample Processing Systems. If during the recovery from Error 81201001, a stuck sample carrier with a tube is moved backwards to the loading position, sample tube/input tray information will become mismatched for samples being loaded onto the AutoMate. | Other | Beckman Coulter sent an Urgent Product Corrective Action letter dated May 7, 2009, with a Fax Back Response Form (sent by US mail) to all customers who received the AutoMate Sample Processing System. The letter identified the product, the problem and a workaround to correct the problem. Customers were instructed to take the following steps:If the stuck sample carrier is moved backwards during recovery from this error, the sample tube/input tray information will become mismatched.Figure 2 in the PCA letter shows the AutoMate viewed from the rear. Sequence 1 to 3 depicts the error recovery issue. The sample is loaded, but cannot move along the track. As the sample attempts to move along the track, it is captured by the track stopper (not seen) below the track. Error 81201001 will be displayed at the console with the following description: The stopper (SL01) did not stop the sample tube carrier. If the stopper or an operator forces the sample backward along the line, it will cause the information mismatch.If you encounter this scenario, do not move the carrier backward to the loading position. See Figure 3 in the PCA letter.Recovery from Error 81201001:1. The normal recovery for error 81201001 is to move the sample carrier forward. If there is no room to do so in the Input module, extra steps must be taken for the carrier to move forward.a. Resume the module by performing error recovery on the Hardware Error Recovery screen at the console. b. Continue the console error recovery process until a carrier is released from BR01, thereby creating space for the stuck carrier to be pushed forward.c. Move the stuck sample carrier forwardd. Once again, resume the module by performing error recovery on the Hardware Error Recovery screen. (The error should not return again)Recovery if Symptoms are Observed: 1. If the incorrect processing described above is observed, the system must be stopped immediately and shut down. All samples m |  | 5 | 5 units total | Nationwide Distribution including VA KY and CA | Terminated | June 28, 2012 | 1148 | 5 | Not_Computer
Z-2204-2011 | 55542 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | CHVH-001A/1B Variable Helical Pitch (vHP) Software for TSX-101A: Aquilion 64; System, X-Ray, Tomography Computed |  |  | April 09, 2010 | 2010 | Toshiba American Medical Systems Inc | The recall was initiated by Toshiba America Medical Systems (TAMS) because the software associated with TSX-101A: Aquilion 64 can prevent scanning from being performed at the specified helical pitch (couch movement speed). | Software design | Toshiba America Medical Systems (TAMS) forwarded an Urgent: Medical Device Correction letter, dated 4/6/10, with attached Customer Reply Form, via US Postal Service to all customers who purchased the CHVH-001A/1B variable helical pitch (vHP) software for TSX-I0IA: Aquilion 64. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that a Toshiba service representative would contact them to schedule an appointment in order to install a corrective measure that will prevent the occurrence of the problem. Toshiba would like customers to set a scan range different from that for the previous scan for the second and subsequent vHP scans, if multiple vHP scans with the same scan conditions are performed in Auto Mode.Customers were instructed to complete and return the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com. Customers with any questions can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Representative at (800) 521-1968. |  | 35 | 35 | Nationwide Distribution -- MN, OH, MO, MI, NY, KS, MD, TX, NJ, IL, ID, TN, AZ, PA, HI, GA, MT, WA, FL, NC, WI, ME, &amp; CA. | Terminated | August 17, 2012 | 861 | 35 | Software | Device Operation | Software update | Software Update
Z-2203-2011 | 55543 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Software for SSH-880CV Aplio Artida; System Diagnostic Ultrasound |  |  | January 06, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) has become aware that the FLEX M software has reversed left and right annotation and post FLEX M measurement error, which may result in incorrect values. | Software design | The recall communication was initiated on 1/6/10 with Toshiba forwarding an Urgent: Medical Device Correction letter with attached Customer Notification Reply Form (via US Postal Service for US Customers) to all customers who purchased the FLEX-M mode on Aplio Artida Diagnostic Ultrasound System. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to avoid performing M image measurement using images saved in FLEX-M mode until the corrective action is completed. If customers need to perform measurement for images saved in FLEX-M mode, they were instructed to use manual calibration. Customers were instructed to note that even though the corrective action is taken, the existing saved images (saved images before the corrective action is taken) are not modified. Therefore, when referring to the measured values on the images which are saved in FLEX-M mode before the corrective action is taken, customers were instructed to carefully handle the measurements results because they may be incorrect.Customers were instructed to complete the attached form and fax it to the toll free number at the top of the form. The form can also be sent via e-mail to raffairs@tams.com.If customers have any questions regarding the medical device correction letter, they can contact the Director of Regulatory Affairs at (800) 421-1968 or their local Toshiba representative. |  | 28 | 28 | Nationwide Distribution -- Including states of IA, MT, PA, CA, OR, GA, LA, OH, TX, MA, OH, WI, and Puerto Rico. | Terminated | April 20, 2012 | 835 | 28 | Software | Output/Calculation | Software update | Software Update
Z-2201-2011 | 56389 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | Software CD is labeled in part: |  |  | July 19, 2010 | 2010 | Bio-Rad Laboratories | The EVOLIS Assay Protocol Files Software is configured in such a way that may result in a miscalculation of quantitative test results and/or an error in reporting positive/negative sample values. | Software design | Bio-Rad Laboratories sent an Urgent Product Correction letter dated July 19, 2010, to all customer sites via FedEx. There are two issues related to The EVOLIS Assay Protocol Files (APF) Software that could potentially affect patient results. Issue 1 involves quantitative sample test/calculations/results interpretation for assays with linear regression, point to point or 4PL data reduction. Issue 2 involves sample test results with " Result Report Flags''. A software patch to correct these two issues will be sent to each customer site. Bio-Rad Laboratories will work with each customer to complete documentation confirming that the software patch has been successfully installed on each instrument. If you have any questions or concerns regarding these two issues, please contact Technical Support at 1-800-224-6723, option 2 and then option 3. |  | 272 | 272 | Nationwide Distribution -- Including Puerto Rico. | Terminated | August 25, 2011 | 402 | 272 | Software | Output/Calculation | Software update | Software Update
Z-2214-2011 | 57591 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | BrightSpeed Excel Select, BrightSpeed Edge Select ,BrightSpeed Elite Select and BrightSpeed Elite Computed Tomography X-ray systems with software version 09BW35.11 or 09HW30.4. |  |  | January 19, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential set of circumstances that could cause X-ray continuation during an unexpected table stop on certain BrightSpeed systems (BrightSpeed Excel/Edge/Elite Select and BrightSpeed Elite Computed Tomography X-ray Systems ) if this event were to occur on your system, patient safety may be impacted. | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION GE Ref: 22932" letter dated January 19, 2011 to its consignees/customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Contact Information, and Additional Dosimetric Information.A second letter dated March 11, 2011 was sent to the consignee/customer letter consisting of a cover letter and an attachment. "This letter is to notify you of an error in the January 19, 2011, Urgent Medical Device Correction Letter, that you have already received (GE Ref: 22932). In the Safety Issue section of that letter it is written: "If this issue were to occur, it would result in the remainder of the exposure being delivered at the patient position where the table stopped up to the prescribed scan time or 30 seconds, whichever is shorter." The stated 30 seconds limitation is incorrect for your particular model. This information has been corrected on the attached Urgent Medical Device Correction letter, so please discard the original letter if it is still in your possession. "GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction.The customers were also instructed as follows: prior to GE's installation of a software update, continue to monitor patient scanning closely. Should the table stop unexpectedly and X-ray remains on, manually stop the scan, and ensure that all potential users in your facility are made aware of this safety notification and the recommended actions.If you have any questions, please contact GE immediately at the following: United States: 800-437-1171; Japan: 0120-055-919; Korea: 1544-6119; Australia/New Zealand: 800-659-465 and China: 800-810-8188. For other countries please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notifi |  | 479 | 479 | Worldwide distribution: USA (nationwide) including states: AL, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, OH, PR, RI, SC, SD, TN, TX, and WV; and countries including: URUGUAY, TURKEY, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, PUERTO RICO, PORTUGAL, POLAND, PARAGUAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LEBANON, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, EGYPT, COLOMBIA, CHINA, CHILE, BRAZIL, BELGIUM, BELARUS, BAHRAIN, AUSTRALIA, ANGOLA, ALGERIA. TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, PHILIPPINES, PARAGUAY, OMAN, NEW ZEALAND, MOROCCO, KUWAIT, KAZAKHSTAN, ISLAMIC REPUBLIC OF IRAN, ARGENTINA. | Terminated | July 10, 2012 | 538 | 479 | Software | Device Operation | Software update | Software Update
Z-2215-2011 | 57469 | No | MSW | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access Immunoassay Systems Thyroglobulin, |  |  | October 18, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because In addition to Access Tg Sample Diluent, an additional diluent, Access Wash Buffer II, was added to Thyroglobulin instructions for use (IFU) for evaluation of samples having Tg levels above the reportable range of the assay. When customers interchange between diluents there is an unacceptable shift in results. A mediandifference of 16% (ranging from 14% to 23%) b | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated October 18, 2010, with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access Immunoassay Systems Thyroglobulin. The letter describes the product, problem and actions to be taken. The customers were instructed to:1. Do not dilute elevated thyroglobulin samples with Wash Buffer II.2. Do not run thyroglobulin patient samples using the onboard dilution feature on the UniCel DxI instrument.3. Dilute elevated thyroglobulin samples only with Access Thyroglobulin Sample Diluent (REF 33866).4. Review your patient result archives. Manually dilute and repeat the analysis of all elevated thyroglobulin samples that had previously been diluted with Wash Buffer II, using Access Thyroglobulin Sample Diluent instead.5. A forthcoming version of the UniCel DxI Assay Protocol File (APF) will remove the Access Thyroglobulin onboard dilution assay (dTg) from the list of available onboard dilution assays.6. The Access Tg product Instructions for Use will be updated to eliminate Wash Buffer II as an acceptable diluent for elevated thyroglobulin samples. Customers should update procedure manuals and other documents that your laboratory may have created which list Wash Buffer II as an acceptable diluent for elevated thyroglobulin samples.In addition, customers were instructed to complete and return the enclosed PCA Response Form with in 10 days via fax to: 786-639-4000 or mail to: Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318, Attn: Regulatory Affairs. Customers with any technical questions regarding the Product Corrective Action were instructed to contact Beckman's Customer Support Center on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside the US and Canada, customers were instructed to contact their local Be |  | 6414 | 6,414 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | August 10, 2012 | 662 | 6414 | Not_Computer
Z-2208-2011 | 58037 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek GRYPHON T BR DS Anchor w/ORTHOCORD |  |  | February 24, 2011 | 2011 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Metal debris in the shaft component of the device | Manufacturing material removal | DePuy Mitek sent an Urgent Voluntary Product Recall Notification letter dated February 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to immediately check their inventory to determine if they have any affected product. Customers were asked to complete the enclosed business reply form and fax it to 508-828-3750 or email to mitekcomplaints@its.jnj.com whether or not they have any of the affected product remaining in their inventory. Customers were asked to utilize the return good authorization number provided by customer service and return all identified product to:DePuy MitekATTN: Recall Coordinator50 Scotland BoulevardBridgewater, MA 02324For any questions call 1-877-379-4871 or contact DePuy Mitek Customer Service directly at 1-800-382-4682, PROMPT 3. |  | 1126 | 1126 units | Worldwide Distribution - USA (nationwide) and the countries of Austrailia Austria Belgium Canada China Estonia Finland France Germany Great Britain Indonesia Ireland Italy Japan Lebanon Lithuania Netherlands Poland Singapore Slovenia South Africa Spain Sweden and Switzerland | Terminated | February 20, 2013 | 727 | 10139 | Not_Computer
Z-2219-2011 | 58429 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | PercuNav Field Generator Stand is a Computed tomography, X-Ray System. Part Number 453561431911. |  |  | April 01, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | The PercuNav Field Generator Stand could tip over during set up, during use, or during transport. | Device Design | On 4/4/11, Philips Healthcare began sending out the URGENT-Medical Device Correction notification letter to their consignees.Philips informed their customers of a potential tipping hazard of the PercuNav Field Generator Stand. The notification letter also provides instructions to customers to ensure the articulating Mounting Arm is positioned so the Field Generator is not further than the specified distance from the back of the clamp on the stand's central pole during set up or during use. Customers are instructed to apply a provided label on stand to warn of a potential for tipping due to excessive force. For transport, customers are instructed to remove the Field Generator from the Mounting Arm and place the Field Generator in the storage compartment on the system cart. The Mounting Arm should then be collapsed and folded into a 90 degree angle.Customers are informed that a Philips field service representative will contact them to schedule installation of a service kit on the stand, free of charge, once service kits are available.Customers can call Philips Customer Service at 1-866-767-7822 and reference to FCO 79500197 for any questions about this recall. |  | 65 | 65 units total (33 units distributed to US and 32 units to foreign consignees) | Worldwide Distribution -- USA, Australia, Austria, Belgium, China, France, Germany, Hong Kong, India, Korea, Netherlands, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | March 15, 2012 | 349 | 65 | Not_Computer
Z-2217-2011 | 58667 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Audio-Video Engineering PAS-210 Remote Active Speakers Kit used with the IntelliVue Information Center |  |  | May 03, 2011 | 2011 | Philips Healthcare Inc. | PAS-210 Kit Remote Active Speaker Assembly used with the IntelliVue Information CenterSpeaker cable connection may lead to intermittent or loss of audio alarm annunciation | Component design/selection | Philips Healthcare issued an Urgent - Medical Device Correction letter dated April 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the "Action to be taken by Customer/User" section of the instructions while they await the correction, which will likely come in the form of a replacement audio cable for all affected speaker assemblies.For further information or support concerning this issue, please contact Philips Healthcare Customer Care Service Center at 800-722-9377 # 3, # 1 and reference the letter or your local Philips representative in your area. |  | 15062 | 15062 units | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 25, 2017 | 2184 | 15062 | I/O | Alarm/Message | Replace cable | Remove or Replace
Z-2213-2011 | 58235 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTER LYSE S III diff Lytic Part Number: 8546796. |  |  | December 28, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated by Beckman Coulter because the COULTER LYSE III diff Lytic Reagent Lot number 101554F was released for distribution to customers between November 17 and December 1, 2010 prior to completion of QC testing. | Pending | Beckman Coulter sent a Product Corrective Action ( PCA) letter dated December 28, 2010, to all affected consignees with an attached PCA Response form via US Postal Service for US Customers) who purchased the Coulter Lyse S III diff Lytic Lot # 101554F. The letter provides customers with an explanation of the problem identified and an action to be taken. If customers currently had any of the lot mentioned above in their laboratory, they were asked to discontinue use and discard remaining product according to their laboratory chemical safety procedures. And to complete and return the enclosed response form within 10 days to acknowledge that they received this notice and are taking appropriate action.For the population at most risk, there is no serious consequences and only remote probability of injury and medically reversible consequences due by the small change for Hgb and WBC results.Consignees were also asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation.Technical questions regarding this Product Corrective Action were directed to Beckman Coulter Customer Service at( 800) -526-7694 in the United States and Canada. |  | 1296 | 1,296 units | Worldwide Distrubution-- USA ( nationwide) including states of : AL, AR, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, RI, SC, TN, TX, VA, WI, &amp; WV and country of Canada. | Terminated | June 04, 2012 | 524 | 1296 | Not_Computer
Z-2221-2011 | 58614 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Radiesse, Injectable Implant - 1.3cc Syringe, Sterile, Sterilized Using Steam, Use Before 2013-01, Ref 8046M7, Lot 1024417, Manufactured by: BioForm Medical Inc, 4133 Courtney Rd #10 Franksville, WI 53126. |  |  | April 06, 2011 | 2011 | Merz Aesthetics, Inc. | Merz Aesthetics, Inc. is recalling 1 lot (1034417) of Radiesse 1.3 cc Implant. The lot was supposed to be scrapped because it had failed to meet its Loss On Drying (LOD) specification. Unfortunately it had been marked as approved, released and shipped. | Employee error | The firm, Merz Aesthetics, Inc., contacted its consignee/customer via e-mail on April 6, 2011. The e-mail states the product, problem and action to be taken. The customer was instructed to follow up with the concerned and inform the status time to time to the Recall Coordinator. The product was confirmed to be quarantined and has been requested to be sent back.If you need any additional information regarding this issue or if you have any questions, contact the International Regulatory Affairs Specialist at 262-835-3300 ext. 3081. |  | 695 | 695 units | International only: Taiwan. | Terminated | November 17, 2011 | 225 | 695 | Not_Computer
Z-2069-2011 | 57473 | No | RGB |  |  |  |  | Lockheed Martin Aculight Capella-r 1850 Laser. This device has the Accession number 0612339. |  |  | November 19, 2010 | 2010 | Lockheed-Martin Aculight | It has been determined that there are certain software driven operating sequences involving off-on cycling of the key-switch that do not result in a reset of the Stim Enable button, and emissions can resume after a standard 5 second countdown without pressing this button a second time. | Radiation Control for Health and Safety Act | On 11/22/10, Lockheed Martin Aculight sent a Memo to their Customers and Users informing them of the "Stim Enable" button reset error. The defect involves the function of the "Stim Enable" button on the front panel of the Capella R-1850.In December 2010, the firm sent a second lletter with instructions on how to return the Capella R-1850 Laser back for repair.The firm's plan to correct the defect involved removal of existing software, installation of revised software, and documented verification testing, per written instructions, to confirm resolution of the Stim-Emit button failure to reset. Corrections were to be made at no cost to the customer. A Lockheed Martin Aculight customer service representative will contact you shortly and provide detailed information on how to return the Capella-R 1850 laser product in your possession for testing and repair on this defect.Should you have any questions, please call (425) 482-1100. |  | 17 | 17 units | Devices were distributed Universities, Institute, and government consignees in CT, IL, MA, MD, OH, TN, TX, and UT.One unit was distributed to Germany. | Terminated | August 23, 2011 | 277 | 17 | Software | Device Operation | Software update | Software Update
Z-2220-2011 | 58581 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, Client software versions 8.0 through 8.6; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | April 04, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly in Eclipse Treatment Planning System, versions 8.0 to 8.6, whereby Actual SSD (source surface distance) value may not correspond to the Planned SSD value after Planning Approval. | Software design | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 4, 2011, to its customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Set the Planning Approved status in Eclipse (Eclipse displays both the Actual and Planned SSD values during plan approval) and verify both SSD values prior to approving the plan.2)Verify the SSD values in the Eclipse Treatment Plan Report prior to approving the plan for treatment or proceeding with treatment delivery (this report lists both the Actual and Planned SSD values).3) Prior to approving the plan in RT Chart, verify both the Actual and Planned SSD values in the Parameters Workspace by selecting Field Parameters&gt;Geometry tab.4) Prior to treatment, they should verify intended SSD values during independent MU calculation.5)Re-familiarize themselves with the concepts of Actual SSD and Planned SSD by reviewing the listed information.Please advise the appropriate personnel working in your radiotherapy department of the content of this letter.Varian mitigated the product behavior in version 8.9.If you require further clarification, please feel free to contact the Varian Oncology Help Desk ; phone: USA and Canada - 1888-827-4265; Europe: +41 41 749 8844; email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com; Latin America: soporte.al@varian.com; Internet: Oncology Systems-www.myvarian.com; Varian Medical Systems -www.varian.com; or your local Varian Customer Support District or Regional Manager. |  | 5288 | 5288 units | Worldwide distribution. | Terminated | June 27, 2012 | 450 | 5288 | Software | Output/Calculation | Software update | Software Update
Z-2216-2011 | 58534 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS accelerators that utilize the RT Therapist v4.2 system in combination with syngo RT Oncologist 4.2. |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential safety issue that may result in mistreatment when customers are configured with the ARTISTE Linac RT Therapist v4.2 in combination with RT Oncologist v.4.2 and use the remote "Online 2D Image Review. | Software design | The firm, Siemens, sent an "Urgent: Medical Device Correction Customer Safety Advisory Notice" on March 31, 2011 via Siemens Customer Service Group or certified mail to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to not use the "Online 2D Image Review" on the Syngo RT Oncologist 4.2; to always calculate the offset directly on the RTT 4.2 before it is applied for the treatment with the patient on the treatment table; include this notice in their syngo RT Therapist 4.2 System Owner Manual chapter "Safety Advisory Letter" where it should remain until the measures defined in the notice are performed, and in the interest of safety, perform the preventative measures noted in this notice and inform all affected personnel immediately.Siemens is currently working on a solution for the incremental loading workflow in syngo RT Therapist 4.2 and will provide a software update as soon as possible. Siemens will notify the customers as soon as an update for their system will be available. UPDATE: The release of Safety Update UI TH010/11/S to distribute the software update for RTTv4.2.94O04 began on April 15, 2011.If you have any questions, contact Regulatory Affairs at 925-602-8157. |  | 45 | 45 active devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, OH, PA and UT; and countries including: Australia, Belgium, Croatia, Germany, Japan, South Africa, Sweden and United Kingdom. | Terminated | March 30, 2012 | 365 | 45 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2205-2011 | 58510 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT, Computed Tomography X-Ray System, including the Brilliance iCT SP system, Model Numbers: 728311 and 728306. The affected systems are equipped with the following software version: 3.2.0.19010. |  |  | March 07, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | A program software error in the algorithm produces image artifacts that could possibly lead to a misdiagnosis of patients who have small metal implants. | Other | A 'Field Safety Notice - Urgent - Medical Device Correction' notification was sent to the firm's customers via Certified mail on 3/8/2011. The notification letter explains the nature of the Metal Artifact Reduction/Metal Artifact Correction software error; the Brilliance CT X-Ray systems subject to correction and how to identify them; and the circumstances under which this error can occur. It also asks that the user stop running the MAR application on any of the affected Brilliance systems identified with the recalled serial numbers and software versions alluded to in the medical device correction notification. Lastly, the notification states that Philips Medical will provide a solution that will fix the issue and the notification lists contact names; telephone numbers; and E-mail addresses which the customer can contact for further information concerning this device issue, or customer support assistance. |  | 8 | 8 units | Nationwide Distribution -- DC, IL, IN, MI, MO, NY, OH, OR, PA, TX, and VT. | Terminated | February 22, 2013 | 718 | 23 | Computer
Z-2222-2011 | 58586 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague, Colleague CX and Colleague CXE Infusion Pumps, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8151, 2M8161 and 2M9161. Intended for use in a wide variety of patient environments for adult, pediatric and neonatal patients. |  |  | April 14, 2011 | 2011 | Baxter Healthcare Corp. | The Hi-Pot safety testing may not have been conducted on the affected pumps during service, prior to being returned to the end user. | Employee error | The firm, Baxter Healthcare Corporation, telephoned the customers on April 14, 2011, informing them that that the Hi-Pot test,and sent a follow up "URGENT DEVICE CORRECTION" letter dated April 27, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate the pumps and take them out of service as soon as possible (Baxter will then send a shipping container for retrieval, re-inspection and return of the device back to the customer at no cost.); and return their colleague pumps for full functional testing including the HiPot test to the Baxter Buffalo Grove Global Technical Services (GTS) Center at Baxter GTS, 900 Corporate Grove Drive, Buffalo Grove, IL 60089.Baxter will perform full functional testing, including the HiPot test, on affected Colleague pumps for which the customer has continuing medical necessity. Should Baxter identify any issues through repair or service prior to returning the Colleague pumps to the customer.If you have questions regarding this communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CT, Monday through Friday. |  | 117 | 117 pumps | Worldwide distribution: USA (nationwide) including states of :Alabama, Arizona, California, Connecticut, Florida, Georgia, Illinois, Kentucky, Louisiana, Massachusetts, Maryland, Mississippi, North Carolina, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Texas, Virginia, Washington and West Virginia, and country of: Canada. | Terminated | May 09, 2012 | 391 | 117 | Not_Computer
Z-2223-2011 | 58514 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | V-CATH Insertion Kits, Sterile; Product Code Number: 301-00; |  |  | April 05, 2011 | 2011 | Neo Medical, Inc. | Firm received recall notice from their contract manufacturer due to visible particulates in Bacteriostatic Sodium Chloride Injection, USP 0.9%, 30 mL vials that are included in insertion kits. | Material/Component Contamination | Neo Medical forwarded Urgent Recall Notices to Dealers and customers on April 5, 2011 via fax. Neo Medical Customer Care also contacted customers by phone on April 7, 2011, followed by Urgent Product Recall Notification letters were sent on April 11, 2011 along with a copy of the Recall Notification from Centurion. The letter requested that customers with affected lots immediately discontinue distribution and contact the end users instructing them to stop dispensing the product. Customers were to complete the accountability information even if inventory had been depleted and fax it to 510-353-9025. In order to return the affected product, customers should contact the Complaint Coordinator at 925-719-2197 or pjduvall@neomedicalinc.com. Replacement or credit would be issued for all product returned. |  | 1990 | 1990 | Worldwide Distribution -- USA, including states of MO, WA, PA, CA, OH, UT, TX and Canada. | Terminated | November 22, 2011 | 231 | 1990 | Not_Computer
Z-2218-2011 | 58549 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | AVOXimeter 1000, AVOXimeter 4000 |  |  | March 23, 2011 | 2011 | International Technidyne Corp. | An incorrect calibration code that affects the total hemoglobin measurement was applied to affected Avoximeter instruments. | Device Design | International Technidyne Corp. sent a letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to check their inventory to determine if they have any of the affected product. If they have any of the affected product, customers were instructed to stop using it and remove it from their inventory. Customers were also instructed to complete the attached Customer Account Tracking Form and return it by fax to (732-635-0144) or e-mail (techsupport@itcmed.com) or mail to:ITC Technical Support20 Corporate Place SouthPiscataway, New Jersey 08854For any questions call the Distributor or ITC Nexus Dx Technical Support at 800-631-5945 (US) or 732-548-5700 (International), ext 4707. |  | 17 | 17 | Worldwide Distribution - USA including TX, KS, OH, OR, FL, GA, NY, MA, and DC and the countries of Taiwan and Thailand | Terminated | October 25, 2012 | 582 | 17 | Other | Output/Calculation | Remove | Remove or Replace
Z-2202-2011 | 58179 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Theken L"POD" Titanium Vertebral Body Replacement System, Vu c"POD" PEEK-OptimaÂż Vertebral Body Replacement System, c"POD" Titanium Vertebral Body Replacement System |  |  | January 18, 2011 | 2011 | Integra LifeSciences Corp. | Certain Theken Spine Vertebral Body Replacement Systems were distributed (i) without certain components (spacers and set screws) that are necessary for their use as vertebral body replacements, and (ii) with a component (insertion tool) that is not relevant to the devices' intended use. | Other | The firm, Integra Spine, sent an "URGENT: RECALL NOTIFICATION" letter dated January 18, 2011, to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to review their current inventory and return all remaining kits; complete and return the RECALL ACKNOWLEDGMENT AND RETURN FORM via fax to: 877-558-6227 or email: melissa.niosi@integralife.com, and contact Integra Spine Customer Service at 866-942-8698 regarding availability of suitable replacement products.Integra Spine will be taking immediate action to restore your inventory and ensure minimal impact on availability for surgical procedures.Should you have any additional questions, please feel free to contact me at 609-936-2495 or email at jon.caparotta@integralife.com. |  | 93 | 93 Kits | Nationwide distribution: USA including states of: AL, AZ, CA, FL, ID, IL, MD, MI, NH, NM, NV, OH, OK, OR, PA, TX, UT and WI. | Terminated | October 27, 2011 | 282 | 93 | Not_Computer
Z-2229-2011 | 58675 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Kits and Trays containing alcohol swabstick/prep pad: |  |  | February 07, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | Custom surgical kits contained either one or more of the following possible contaminated recalled products: Triad lubricating jelly, alcohol prep pads, and alcohol swabs. | Other | DeRoyal notified their consignees by letter with the subject line "Re: DeRoyal Recall of Triad Medical Lubricating Jelly, Alcohol Prep Pads and Alcohol Swabs" on 02/07/2011 that kits in their possession contained recalled Triad products that were possibly contaminated. Customers were to use the attached spreadsheet to identify and quarantine the affected product. The spreadsheet should have been completed even if there was no affected product. The form should have been faxed back to 865-362-3716 or emailed to recalls@deroyal.com. Customers' DeRoyal Sales Representative would re-label the affected product in order to prevent further use. The DeRoyal Sales Representatives are also providing in-service training on effectuating this recall, including the importance of identifying, retrieving and discarding the recalled products as well as the documentation necessary to account for these products. Distributors were directed to recall down to the retail level. If there are questions or if further assistance is needed, customers can contact their DeRoyal Sales Representative or customer support department at 1-800-251-9864. |  | 340976 | 340,976 units total | Nationwide Distribution | Terminated | August 06, 2014 | 1276 | 340976 | Not_Computer
Z-2237-2011 | 58563 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Nasal Cannula (Adult) Salter Style with modified nasal prongs, 7" (2.1M) Q supply tube and CO2 sample tube with male luer connector 22mm (I.D) x 6mm (O.D) Adaptor. Adult Divided Cannula. |  |  | April 19, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | Complaint received that the product in the box was different from labeled product. | Labeling mix-ups | The firm, Salter Labs, sent an "URGENT PRODUCT RECALL" subtitled Adult Divided Cannula Product Code 4707 Lot 012411, letter dated April 19, 2011 via certified mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory; return any of the lot in their possession to Salter; contact Salter Customer Service at 1-800-235-4203 to arrange for the return of the product, and complete and return the Customer Reply Form via fax to 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 22150 | 22,150 pieces | Nationwide distribution. | Terminated | August 23, 2011 | 126 | 22150 | Not_Computer
Z-2231-2011 | 58569 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTERÂż Ac"T" 5diff Fix Reagent. |  |  | April 08, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that an issue has been identified with one of the raw materials used to manufacture the Coulter A T 5 diff PN 8547171. | Other | Beckman Coulter initiated a Product Corrective Action (PCA) letter on April 8, 2011, with attached PCA Response Form to all customers who purchased the COULTERÂż Ac"T" 5diff Fix Reagent. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discard the product according to their laboratory safety procedures and use the attached response form to request replacement product. All flagged results should be reviewed as outlined in the Instructions for Use manual and your laboratory protocol. A retrospective review of results is not required if all recommendations for review have been followed as shown in the Instructions for Use.Customers were instructed to share the information with their laboratory staff and to retain the notification as part of their Quality System documentation. If the product was further distirbuted to another laboratory, customers should provide them a copy of the notification letter.The raw material issue has been corrected beginning with COULTERÂż A T 5diff Fix lot 16902E.Customers were instructed to complete and return the enclosed Response Form within ten (10) days.Customers with any questions concerning this notice, were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 4662 | 4662 units total (2278 in US) | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Angola, Australia, Botswana, Bulgaria, Czech Republic, Estonia, Finland, France, Germany, Greece, Guyana, Italy, Japan, Kenya, Mexico, Morocco, Namibia, Netherlands, Nigeria, Philippines, Puerto Rico, Russian Federation, Slovakia, South Africa, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Zambia | Terminated | June 04, 2012 | 423 | 4662 | Not_Computer
Z-2227-2011 | 58584 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UnicelÂż DxH 800 CoulterÂż Cellular Analysis System. |  |  | April 14, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the auto prune will not remove patient results if the patient record includes demographics where the results are less than 30 days old. | Software design | The recall communication was initiated on 4/12/11 with Beckman Coulter forwarding an Urgent Product Correction (PC) letter with attached PC Response Form (via US Postal Service for US customers) to all customers who purchased the UniCelÂż DxH" 800 CoulterÂż Cellular Analysis System PN 629029. The letter provides customers with an explanation of the problem identified and an action to be taken. For Issue 1 &amp; 2, customers were instructed to perform manual database maintenance procedures to locate and delete unnecessary patient result and demographic information that extends beyond the life (in days) of the control lots. Ensure an alternative method of analysis is available during database maintenance.For Issue 3, customers were instructed to not attempt to transmit the files again after receiving an incomplete report. Instead, submit the missing file(s) by one of the following methods:- Submit your data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)- Fax printouts of the missing files to 786-639-3999- Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.- Include the IQAP number on any communication. A corrected report will be returned to the customers' laboratories.For issue 4, customers were instructed to submit their control data to IQAP by one of the following methods:-Submit their data at the electronic IQAP website (eIQAP) at(http://www.beckmancoulter.com/qap/index.jsp)-Fax printouts of the missing files to 786-639-3999-Email an electronic copy (.CSV) of the control file to IQAP.Reports@beckman.com.-Include your IQAP number on any communication.Customers were instructed to complete and return the enclosed Response Form within ten days. Customers were instructed to contact Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 673 | 673 units | Worldwide Distribution -- US, Canada, Australia, Belgium, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Republic of Korea, Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 23, 2012 | 405 | 673 | Software
Z-2225-2011 | 58333 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Universal Power Supply Units, Part Number DP0007 and DP0800 , for use with the Integra Padgett Dermatomes - Model S, SB, and S6, Device Listing #E208841, 510 (k) exempt. |  |  | March 11, 2011 | 2011 | Integra LifeSciences Corp. | Dermatome power supply units distributed by the recalling firm may not meet the required medical device electrical safety regulations which can result in an increased potential for electric shock to the device user and/or patient. | Other | Integra LifeSciences Corp. initiated their recall by contacting their Sales Representatives via telephone conference calls on 3/10/2011. Separate written recall communications, in the form of an 'URGENT PRODUCT RECALL NOTIFICATION' were provided via E-mail and FedEx to the firm's Sales Representatives on 3/11/2011; and to their hospital and distributor customers on 3/11/2011; 3/14/2011; and on 4/12/2011. The recall notification addressed to the non-Integra Sales Rep customers informs the customers of the problem with the power supply units (that they may not meet the required medical device electrical safety standards which may result in an increased potential for electrical shock). The notification goes on to ask that the customers examine their inventories and determine if they have any affected product on-hand; if so, the customers are instructed to discontinue use of the power supply unit and return the unit(s) to Integra LifeSciences Corporation immediately for reworking. The letter also provides photographs of the affected units which the customer can use to locate and verify the serial number(s) of the device (if applicable). In addition, the notification is accompanied by a 'RECALL ACKNOWLEDGMENT AND RETURN FORM' which the customer is asked to sign and complete in order to document receipt of the recall notification and whether or not the customer has any of the recalled power supply units available for return. The notification letter provided to the Integra Sales Representatives describes the device models affected; the reason for the recall; and asks that the Reps. examine their inventories for the presence of any of the affected power supply units. If found, the Reps are instructed not to use, or to further distribute any of the affected units in their possession, but to immediately make arrangements to return the unit(s) to Integra for rework. The notification supplied to the Sales Reps provides them with a synopsis of the firm's field action |  | 1047 | 892 units were distributed in the U.S &amp; 155 units were distributed in foreign countries | Worldwide Distribution - USA including FL, CO, IL, AR, MD, IN, CA, MT, AZ, NY, LA, TX, MA, MO, CT, MI, NE, OH, NC, WV, VA, PA, NJ, TN, UT, MN, OK, WI, GA, AL, WA, KS, OR, AK, DC, SC, NE, MS, RI, NV, IA, and ND and the country of Canada | Terminated | April 17, 2012 | 403 | 1047 | Not_Computer
Z-2232-2011 | 58464 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAKO PN 204588 Camera |  |  | October 25, 2010 | 2010 | Mako Surgical Corporation | MAKO Surgical Corporation recalled ONE infra-red camera MAKO PN 204854 which was a component of their RIO Robotic Arm Interactive Orthopedic System with SSN P7-03127, on November 22, 2010. This component can function improperly and cause the accuracy of the system to be outside of the specification listed by the manufacturer. The recall was initiated after the camera manufacture, NDI, Waterloo, On | Nonconforming Material/Component | The firm, MAKO, sent a "RECALL NOTICE" letter dated November 22, 2010, to its customer. The letter described the product, problem and action to be taken. The customer was instructed to complete and return the attached Acknowledgement Form to their MAKOplasty Specialist (MPS) via mail to: MAKO Surgical Corp., Attn: Jim Pomeroy, 2555 Davie Road, Fort Lauderdale, FL 33317. The letter stated that "It is important that the camera be replaced immediately, and MAKO is currently addressing the situation with the camera manufacturer." In addition, MAKO will replace the camera with a camera that is not affected by the recall and verify the customers system to meet the functional specifications.If you have any questions or concerns, please contact your MPS or Sr. Director of Quality Assurance at 954-927-2044 ext 604, and /or Sr. Director, Customer Support at 954-927-2044 ext. 442. |  | 1 | 1 | Nationwide distribution: NY only. | Terminated | November 09, 2011 | 380 | 1 | Hardware | Output/Calculation | Replace component | Remove or Replace
Z-2224-2011 | 58558 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac and TrueBeam, High Energy Linear Accelerator, Model Numbers: H14, H19, H29; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | March 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The coolant may leak. The resulting coolant leak presents a risk of electrical shock to any person working within the protective housing. | Process design | On 3/31/2011, all consignees were notified via an Urgent Medical Device Correction letter and on-site visit by the firm's sales representative, informing them of the affected products, the issue, and recommended actions. If non-Varian personnel service the device, then they should have been immediately notified of the content of the letter. Anyone who performs work on the device should exercise extreme care to identify any possible coolant leak and immediately locate and correct the source of the leak. Customers were to immediately notify Varian Service or their internal service personnel of any coolant leakage. Varian was to contact customers to arrange for inspection if one had not been performed. The appropriate personnel should have been advised of the content of the letter. If further clarification was needed, customers were to contact their local Varian Customer Support. |  | 718 | 718 units | Worldwide Distribution | Terminated | March 13, 2012 | 348 | 718 | Not_Computer
Z-2226-2011 | 58555 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON Systems Calibrator 1, Part Numbers: 468405, Lots prior to M005558 |  |  | March 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because standardization of the Synchron Systems Calibrator 1, used for both the Synchron LX and UniCel DxC systems, has been established to the new IFCC Reference Preparation for Plasma Proteins, ERM-DA470k (formerly CR470) beginning with lot number M005558.During the standardization to the new IFCC standard, ERM-DA470k, we confirmed that Complement C4 (C4) had been re | Pending | Beckman Coulter initiated a Product Corrective Action (PCA) letter on the week of March 14, 2011, to the affected customers. Though Beckman is recalling the LX Cal 1, the recall letter was sent to the C4 reagent customers, as these are the only customers who will be impacted by the problem noted in the PCA letter. The letter provided the customers with an explanation of the problem identified and an action to be taken.The letter recommended consignees take the following actions:(1)Notify their clients to review and evaluate the impact on patients whose Complement C4 recovered near the clinical decision points.(2) Re-evaluate their laboratory's reference ranges for C4.The letter further requested that consignees share this information with laboratory staff and retain this notification as part of your laboratory Quality System documentation.Questions regarding this Product Announcement letter, or requests for the correlation graphs should be directed to Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/Support Request or call (800) 854-3633 in the United States and Canada. |  | 96173 | 96,173 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Belgium, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Czech Republic, Denmark, Eritrea, France, French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Republic of Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Namibia, Netherlands, New Caledonia, New Zealand, Oman, Panama, Poland, Puerto Rico, Qatar Reunion, Russian Federation, Rwanda, Serbia, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Viet Nam | Terminated | May 04, 2012 | 420 | 96173 | Not_Computer
Z-2242-2011 | 58564 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs Nasal Cannula (Adult) Salter Style with 14' (4.3) supply tube. |  |  | April 19, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | One complaint received of product in box different from what was on label. | Other | Salter Laboratories sent a URGENT PRODUCT RECALL NOTIFICATION letter dated April 19, 2011, certified mail to all of the affected consignees, along with a RECALL EFFECTIVENESS CHECK FORM to acknowledge receipt of this communication.It was reported that a package labeled 1600-14, Adult Nasal Cannula actually contained a 25-foot length of tubing and has determined that additional product may have been incorrectly labeled in this manner.No illness or injury would result from this product mix-ups.Should you have any questions, please contact Salter Laboratories at 1-800-235-4203. |  | 3500 | 3500 pieces | Worldwide Distribution - USA (nationwide) and the country of South America. | Terminated | August 23, 2011 | 126 | 3500 | Not_Computer
Z-2240-2011 | 58487 | mattress, flotation therapy, non-powered | IKY | Nonpowered flotation therapy mattress. | General Hospital | General Hospital | 510(K) Exempt | Stryker, Impression, Non-Powered Mattress, Model 2980, Stryker Medical, Portage, MI. |  |  | April 29, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Affected mattresses may not conform to federal flammability standards. | Other | Stryker Medical sent an URGENT MEDICAL DEVICE RECALL letter dated April 29, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate all affected mattresses , remove the existing mattress cover, and replace it with an enclosed mattress cover. Customers were then to notify Stryker to confirm repairs. If the mattresses were loaned or sold, Stryker requested customers to notify the new user of the recall. Customers with questions or concerns, were asked to call 269-389-6689. |  | 213 | 213 | Worldwide Distribution - USA including California and the country of Canada. | Terminated | October 07, 2011 | 161 | 213 | Not_Computer
Z-2243-2011 | 58533 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE , ONCOR and PRIMUS digital accelerators with syngo RT Therapist v4.1 and 4.2 system; |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens found a potential safety issue related to patient treatment with table offsets on syngo RT Therapist (RTT) 4.1 and 4.2 which could result in dose being delivered to wrong location. | Software design | Siemens initiated an Urgent Medical Device Correction letter on March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:use only relataive table setup with all planned table values '0' if applicable.In cases where the treatment plan contains at least one table parameter other than '0' to consider the following:Use only one PVG plus resulting offset per fraction.Once an offset has been applied do not use 'apply offset' at RTT for any further offset.Select the PVG and treatment beams in the same auto-sequence. In casethis is not possible, the table position for each treatment beam, in the fraction, needs to be verified separately (especially for treatments with beams having different table posiitons. If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Scenario 2 - Manual change of table posiiton followed by a PVG calculated offset.Always deliver the PVG's at the planned table position. Avoid manual change of table posiiton prior to applicaiton of PVG.Select the PVG and treatment beams in the same auto-sequence. In case this is not possible, the table position for each treatment beam, in the fraction, need to be verified separately (especially for treatments with beams having different table positions).If an in-session resumption or individual download of beams has to be performed, make sure that the treatment table is in the correct position.In order to prevent any unintended table movement select manual table movement for all directions in the syngo RT Therapist service UI.Customers were informed that a software update would be provided as soon as possible.Customers were asked to include the Customer Sa |  | 164 | 164 active devices | Worldwide Distribution - USA including CA, DE, FL, KY, LA, MA, MI, MO, NC, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV and the countries of Australia, Belgium, Canada, Croatia, Egypt, France, Germany, Hungary, India, Ireland, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Republic of Korea, Russian Federation, Saudi Arabia, South Africa, Spain, Turkey, and the United Kingdom. | Terminated | March 30, 2012 | 365 | 164 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2241-2011 | 58582 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system, and has a Part number 989605858501. |  |  | February 09, 2011 | 2011 | Philips And Neusoft Medical Systems Co., Ltd. | Dimension measurement displays on a combined image were found inaccurate after the clinician merged multiple image series into one image. | Other | Neusoft Medical Systems sent an Urgent Device Correction Notification letter dated February 9, 2011, via certified mail to their distributor NEUISYS. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue use of combined images. Instead, use MPR Batch for diagnostic reviews until the planned upgrade is provided to the product.Neusoft Medical System and Philips Medical System released the mandatory Field Change Order (FCO) by March 31, 2011. All nine affected systems will be updated with aforementioned kits in six months after the release.Consignees were instructed to contact Service Support Department by e-mail to nms-service@neusoft.com |  | 9 | 9 units distributed in US | Nationwide Distribution including CA, GA, LA, NC, OH, TN and TX. | Terminated | August 24, 2011 | 196 | 9 | Other | Output/Calculation | Upgrade | Repair
Z-2239-2011 | 58685 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż Soft Mesh 4"x6" (10cm x 15cm) Flat, Product Code: 0117010. |  |  | April 25, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeling: complaints of product labeled as BardÂż Soft Mesh 4" x 6" containing BardÂż Soft Mesh 6" x 6" | Labeling mix-ups | Davol issued an URGENT: MEDICAL DEVICE RECALL notification letter to customers dated April 25, 2011, informing them the recall involving a single lot of the BARD Soft Mesh - Lot# HUUF0530. This recall has been initiated because this lot number may contain a different size patch than is indicated. Customers were instructed to immediately examine inventory and quarantine product subject to recall. Customers are asked to identify and notify their customers to whom recalled product may have been further distributed. A Recall Effectiveness Check reply form should be completed and returned to Davol via fax (401-825-8753). For directions on returning product or for questions regarding the recall, contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 198 | 198 units | Nationwide Distribution | Terminated | October 28, 2011 | 186 | 198 | Not_Computer
Z-2238-2011 | 58789 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Maquet Magnus Carbon-fibre table top; Device Part Number 1180.16xx (x represents a digit from 0-5 and F-for the US market) |  |  | April 15, 2011 | 2011 | Maquet Inc. | Maquet GmbH's quality monitoring detected that incorrect cover sheets were installed with some Magnus carbon fibre table tops. | Device Design | Maquet sent an Urgent Medical Device Corrective Action letter dated 4/15/2011, along with a Action Response Form, via Federal Express to their three customers informing them the incorrect cover sheets were installed with some of the affected "MAGNUS 1180.16F0 carbon-fibre table tops." The firm will have a Service Technician contact you to service the OR table, and will be exchanging the cover sheet free of charge. To the customers, please follow the instructions on the response form. At this time there have been no reports that injury has occurred.If you have additional questions, please contact yur local MAQUET Customer Service at 1-888-627-8383 |  | 3 | 3 | LA, MO, VA | Terminated | August 21, 2013 | 859 | 3 | Not_Computer
Z-2246-2011 | 58565 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 6.5 Client Software. |  |  | April 05, 2011 | 2011 | AGFA Corp. | Software issue that occurs after dictation of a study. | Other | Agfa Corp. sent an "URGENT SAFETY NOTICE" customer letter via Agfa email to customers on 4/14/2011, along with a "URGENT SAFETY NOTICE FEEDBACK FORM," The letter identified the product , the problem, and the action to be taken by the customer.It was discovered a software issue in the IMPAX 6.5. The upgrade has already been released to address the software problem.Customers were instructed to fax back the feedback form as soon as possible to (864) 421-1664.If you have any questions, please contact your local Agfa Healthcare organization. |  | 95 | 95 units | Nationwide Distribution -- including states of WI, and SC. | Terminated | September 15, 2011 | 163 | 95 | Software |  | Software update | Software Update
Z-2245-2011 | 58328 | [6 empty] | ***REF INT411***OSSEOTITE Tapered Certain Implant 4 x 11.5mm***Sterile Using Radiation***Rx Only***BIOMET 3i Dental lberica S.L.WTC Almeda Park, Ed, 1 Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain***Phone +34 934-705-500; Fax +34 933-717-849. |  |  | March 10, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i is recalling their product Osseotite NT Certain Dental Implant INT411 . The internal thread was not manufactured correctly. | Other | Biomet 3i, LLC sent an Urgent Medical Device Recall letter dated March 14, 2011, via email or fax to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A small percentage of the implant lots may have a condition affecting the internal thread. The condition, if present, will prevent the cover screw, temporary healing abutment, final restorative abutment, etc. from engaging the implant. Customers were instructed to contact Customer Service at 1-800-342-5454 if this condition is recognized upon attempting to assemble any of the screws with the implant.Customers were instructed to return the implant to the firm for replacement to the following address:BIOMET 3iRegulatory Compliance Department4555 Riverside DrivePalm Beach Gardens, Florida 33410Incident No. CC117886Customers were also instructed to respond with the attached Response Fax or email to DomesticComplaints@Biomet.com.For any questions regarding this recall call 561-776-6906. |  | 236 | 236 | Worldwide Distribution - USA (nationwide) and the countries of Europe, Japan, Asia, and South America. | Terminated | January 25, 2012 | 321 | 236 | Not_Computer
Z-2244-2011 | 58550 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT - FA Reagent Bottle |  |  | April 13, 2011 | 2011 | bioMerieux, Inc. | The charcoal suspension in a portion of the lot does not meet the intended density. | Other | BioMerieux sent an Urgent Product Removal Notice dated April 13, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Evaluate positive results adhering to their institution's protocols for determining false versus true positives. Negative bottles should be checked by smear and/or subculture at some point prior to discarding as negative. Any remaining inventory of BacT/ALERT FA culture bottle lot 1027481 should be destroyed in accordance with their facility's policies and procedures. An Acknowledgement and Product Replacement Form was included to be completed and returned to bioMerieux or credit or replacement of unused bottles. |  | 327700 | 327,700 bottles | Worldwide Distribution - USA (nationwide) and the countries of China, Tokyo, South Korea, Canada, Sweden, Austria, Poland, and Belgium | Terminated | April 03, 2012 | 356 | 327700 | Not_Computer
Z-2293-2011 | 58545 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics version 9.0, Siemens Medical Solutions, USA, Inc., Ann Arbor, MI. |  |  | March 24, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | If one or more graphs are in a table on the report, and if any of those graphs are deleted, it may lead to reports failing to save and/or marking the study as read might leave the study in an unread state or might mark the study as read but leave the report with a "Preliminary" watermark displayed. | Software design | The firm, SIEMENS, sent a "Customer Safety Advisory Notice" letter dated March 24, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to place graphs in sections and not tables in the report; avoid deleting graphs from the reports in run-time and include this safety advisory notice in their operator's manual where it should remain until the service patch is applied. SIEMENS will be issuing a service patch to fix the defect. The letter states the patch will be available in the next few weeks.If you have any questions, call +1 (734) 205-2400. |  | 29 | 29 devices | Nationwide distribution: USA including states of: CA, CT, FL, ID, MD, MI, MN, NJ, NY and PA. | Terminated | July 10, 2012 | 474 | 29 | Software | Device Operation | Software update | Software Update
Z-2233-2011 | 58373 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | MicroStream Filterline Infant/Neonatal products: Part # 006285 - FILTERLINE H SET INF/NEO 25 UN PHYSIO (Inf/Neo Airway Adapter). The adult/pediatric airway adapter and the adult/pediatric FilterLine and VitaLine sets are not affected by this recall. |  |  | March 25, 2011 | 2011 | Oridion Medical 1987 Ltd. | fine plastic strands on inner surface of the infant neonatal airway adapter may become dislodged and inhaled by the patient | Equipment maintenance | The firm, Oridion, sent an "Microstream Filterline Field Correction Notice" letter dated March 25, 2011 to its US and foreign customers via telephone, email with the documentation discussed in the call, and letter. The letter described the product, problem and actions to be taken. The customers were instructed to place all affected product held in their inventory on ship hold; gather these units in their inventory and at the customer sites and exchange with the new product; inspect filter line starter (sample) kits shipped with Capnostream 20 and Microcap patient monitors for affected product, if affected Infant/Neonatal filter lines are found, remove them from the starter (sample) kit; send a purchase order in order to obtain replacement product for units in their stock and at customers sites, and complete and return the Inventory Reporting Form Excel Spread Sheet within two weeks of receiving the email even if they or their customers do not have the affected inventory. Oridion shall conduct a removal of the products from the related production lots from the field and an immediate replacement with product from subsequent production lots that meet all requirements. This action will be conducted by Oridion's distributors or by Oridion directly in some cases. Oridion anticipates that new product will be available for shipment during the week of April 4th.Questions regarding the recall can be directed towards the Customer Service Coordinator at: 1-888-674-3466 x1234 or NA-FCA@oridion.com. |  | 350 | 350 units in US (14 boxes with 25 units ea) | Worldwide distribution. | Terminated | March 19, 2012 | 360 | 10525 | Not_Computer
Z-2291-2011 | 58420 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; |  |  | March 29, 2011 | 2011 | Baxter Healthcare Corp. | The LOGIX OE automated backup does not retain the three most recent backup files and instead will only retain the last three backups of the previous year. The software stores backup files in numeric order rather than chronological date order. | Software design | Baxter Healthcare Corp. sent an URGENT DEVICE CORRECTION letter dated March 29, 2011, via first class mail to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. It was discovered there is an issue with the LOGIX OE automated backup not retaining the three most recent backup files and instead will only retain the last three backups of the previous year. The accounts were requested to turn off the automated backup function on the LOGIX OE Software, and perform manual backups following the instructions included in the Urgent Device Correction letter, and retain a copy of the letter at each computer on which LOGIX OE Server is installed. The customer were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. |  | 84 | 84 units | Worldwide Distribution -- USA including Puerto Rico, and the country Canada. | Terminated | January 05, 2017 | 2109 | 84 | Software | Device Operation | Instructions on turning off auto backup | Safety Notice/Insructions
Z-2261-2011 | 58021 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Case Label: McKESSON Medi-Pak Uretheral Catheter Tray With Red Rubber Catheter Sterile Single Use Disposable 20 Trays Per Case Reorder No. 37-202 Marketed By McKesson Corporation McKesson Medical-Surgical, Richmond, VA 23228 Made in the U.S.A. |  |  | January 20, 2011 | 2011 | Nurse Assist, Inc | The trays contain Triad Group Sterile Lubricating Jelly packs recalled by Triad Group because of the possibility of inadequate sterilization. | Other | The firm, Nurse Assist Incorporated, sent a letter dated January 20, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory immediately; quarantine the product; identify their customers and notify them at once about the product recall (Nurse Assist will provide bright labels to affix to your affected inventory); provide their customers with labels as well, and complete and return the attached RECALL ACKNOWLEGEMENT form via fax to: 817-231-1500 or mail to: Nurse Assist, Inc., 3400 Northern Cross Blvd., Forth Worth, TX 76137.Only the product and lots identified on the attached list are affected.If you have any questions or to request labels, please call Nurse Assist Customer Service Monday through Friday, from 8:00 am to 5:00 pm Central Time at 1-800-649-6800. |  | 1504 | 1504 cases | Worldwide distribution: USA (nationwide) and countries including: Canada, Dubai and Saudi Arabia. | Terminated | October 25, 2012 | 644 | 124789 | Not_Computer
Z-2294-2011 | 58600 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll E Series BLS Defibrillator/Pacemaker/Monitor |  |  | April 21, 2011 | 2011 | ZOLL Medical Corporation | Zoll E Series BLS Defibrillator may fail to auto charge and deliver energy | Software design | The firm, Zoll, sent an "Urgent Medical Device Correction" letter dated April 22, 2011 via UPS to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately alert all E Series BLS users to this problem; instruct users to not touch patient and keep the patient still during ECG analysis, Instruct users that if they experience this problem, to press the Analyze button again, and contact ZOLL Technical Service for the software upgrade kit. The letter also states that affected units can remain in service until the software has been updated. If you have any questions, please contact 24/7 technical support at 1 (800) 348-9011 or +1 (978) 421-9460. |  | 874 | 874 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Finland, Germany, Malaysia and United Kingdom. | Terminated | January 23, 2014 | 1008 | 874 | Software |  | Software update | Software Update
Z-2292-2011 | 58601 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | WaveLight FS200 Femtosecond Laser System |  |  | April 18, 2011 | 2011 | Wavelight AG | The software calculating the patterns were found to has a rounding error for certain parameter ranges. These rounding errors could result in an incomplete cut in the donor or patient tissue. | Other | WaveLight AG sent an URGENT MEDICAL DEVICE CORRECTION letter April 18, 2011 to the affected customers. The letter identified the product, problem and the action to be taken by the customer. The letter advised the customers todiscontinue keratoplasty cuts with the WaveLight FS200 until the new version can be installed. Alcon will contact you to schedule a visit of a service engineer at your site, to disable the keratoplasty function in your device, until a new software version can be installed to fix this problem. If you have experienced a product problem or adverse event related to this medical device correction, please contact Alcon Medical Safety at 1-800-561-6466, option 3.Customers were requested to complete the attached reply form to acknowledge receipt of this notification and fax it to ( 817) 916-9087. |  | 55 | 55 units | Worldwide Distribution - USA including states of TX, KS, NC, CA, FL, OH, SC, MD, VA and NY and countries of Switzerland, Germany, Greece, United Kingdom, Ireland, Spain, Italy, Australia, India, Malaysia, France, Finland, Saudi Arabia, Czech Republic, Poland, China and South Korea and Iran. | Terminated | September 23, 2011 | 158 | 55 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2383-2011 | 58650 | No | KNY | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus ACMI, REF 25 BX, PTFE Coated 3 cm Flexible Tip Guidewire, .035" x 150 cm, Sterile EO, Rx Only, Gyrus ACMI, Inc. 136 Turnpike Road,Southborough, MA 01772-2104 USA |  |  | April 15, 2011 | 2011 | Gyrus Acmi, Incorporated | Gyrus ACMI, Inc. wants to bring a recall issue to their customers attention concerning one lot of their 25BX-PTFE Coated 3 cm Flexible Tip Guidewire. As a result of a manufacturing error in the production this lot of guidewires, some units of product were loaded in the dispenser backwards. In such cases, the stiff end of the guidewire would be dispensed from the product first and, if undetected, | Process control | The firm, Olympus ACMI, sent an "Urgent: Medical Device Recall" letter dated April 15, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) Inspect their stock to confirm if any of the affected lot of products remains in their possession2)Immediately cease any further use of the affected product, remove, quarantine and return it. 3) Call their Gyrus ACMI customer service representative at 1-888-542-7266 to obtain a Returned Goods Authorization 4) Return the product to: Attn: PTFE Guidewire Recall RGA#_________Gyrus ACMI LP, 6845 Wedgwood Court, Maple Grove, MN 553115) Complete and return the RECALL REPLY FORM via fax, even if they are returning no product, to: Regulator Affairs (508) 804-2624. The letter states that the customer will receive repeat mailings until they have faxed in the form. Additionally, If they have further distributed the product they are requested to identify and notify their customers at once.If you have any questions, please do not hesitate to contact customer service at 1-888-524-7266. |  | 145 | 145 | Nationwide distribution: USA including states of: CA, GA, HI, IN, KY, MT, NY, NC, PA, TX, UT, and WA. | Terminated | August 23, 2012 | 496 | 145 | Not_Computer
Z-2295-2011 | 58612 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50Âż Ventilator. Model #HT50 H1-B |  |  | January 13, 2009 | 2009 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical Instruments confirmed that the Newport HT50 ventilators that were shipped in December 2008 were without the label on the top of the ventilator that reminds users to keep HT50 ventilators plugged into an external power supply whenever possible to preserve the longterm capacity of the battery. | Other | Newport Medical Instruments, Inc. sent an FIELD CORRECTION letter dated January 13, 2009, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. The Newport HT50 ventilators that were shipped in December 2008, were without the label on the top of the ventilator that reminds users to keep HT50 ventilators plugged into an external power supply whenever possible to preserve the longterm capacity of the battery. The customer were also requested to complete the attached customer REWORK INSTRUCTIONS form and fax it to (714) 427-0489. If you have any further questions, please call (714) 427-5811 x 344 |  |  |  | Worldwide Distribution-- USA including states of OH, VA, WA, WI and country of Turkey. | Terminated | June 01, 2011 | 869 |  | Not_Computer
Z-2247-2011 | 58471 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Philips MicrostreamÂż Model M1923A FILTERLINE H SET INF/NEO 25UN PHILIPS |  |  | April 07, 2011 | 2011 | Philips Healthcare Inc. | Fine plastic strands on inner surface of the infant/ neonatal airway adapter may become dislodged and inhaled by the patient. | Equipment maintenance | The firm, Philips, sent a Field Safety Notice titled "URGENT MEDICAL DEVICE RECALL" letter dated April 2011 to its customers. Customer notifications outside the US will be managed by Philips representatives in each affected geography. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all products from affected lots; remove from inventory and dispose in accordance with local regulations (new products will be sent proactively to each affected customer). Direct accounts were instructed to notify additional customers to whom the product has been distributed. These customers are instructed to follow the instructions in the "Action to be taken by Customer/User" section of the Notice. Questions regarding the recall can be directed towards the customer's local Philips representative at 1-800-722-9377. |  | 25400 | 25,400 units (1016 boxes) | Worldwide distribution. | Terminated | March 14, 2013 | 707 | 25475 | Not_Computer
Z-2382-2011 | 58629 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator. |  |  | May 14, 2010 | 2010 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed that the Newport HT50 hard shell Emergency Preparedness (EP) Storage Cases must be relabeled to caution users to follow current labeling guidelines for charging the battery for 8 hours but not longer than 24 hours at any one time. | Other | Newport Medical sent an FIELD CORRECTION NOTICE to all affected customers dated May 14, 2010. The letter identified the product, the problem and the action needed to be taken by the customer. Newport Medical also provided a revised labeling for the Emergency Preparedness (EP) Storage Case and an instruction sheet describing where to apply the new labels. Customers were instructed with the information that all the EP storage cases must be relabeled to caution users to follow current labeling guidelines for charging the battery for 8 hours but not longer than 24 hours at any one time. Customers were instructed to complete the attached acknowledgement form and fax it to " Regulatory@NewportMedical.com " Customers with any questions should call Technical Support Department at (800) 451-3111 for US customers, and (714)427-5811 x500 for international customers. |  | 1113 | 1,113 EP cases | Worldwide Distribution -- USA including states of PA, TX, AK, GA, CA, MD, KY, FL, VA, LA, NC, OK and countries of Canada, Colombia, Armenia and Japan. | Terminated | June 01, 2011 | 383 | 1113 | Not_Computer
Z-2387-2011 | 58652 | No | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | DrugCheck, Manufactured by: Express Diagnostics Int;l Inc. 1550 Industrial Drive, Blue Earth MN 56013 USA, Model numbers : |  |  | March 18, 2011 | 2011 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc. has initiated a medical device recall on various model numbers of DrugCheck devices,This recall was due to incorrect product expiration dating after findings from an FDA inspection. Use of expired product may not work properly, and may give incorrect screening results. | Incorrect or no expiration date | The firm, Express Diagnostic Int'l, Inc., sent an "Urgent Recall of DrugCheck Product" letter dated April 19, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine the product; notify Express Diagnostic of the quantity even if no product is remaining; complete and return the DISTRIBUTOR/CUSTOMER RECALL CONFIRMATION FROM via fax to 1-507-526-2252 as soon as possible. In addition, if they have further distributed the product to identify their customers and notify them of the recall. Provided procedures to be followed for them to destroy the quarantined product. Express Diagnostics will replace the DrugCheck Drug Screen devices.If you have any questions, please contact the Quality Assurance Manager at 507-526-3951, ext. 117. |  | 6925 | 6925 devices | Worldwide distribution: USA (nationwide) including states of: FL, IA, MD, OK, and WA; and countries including: AUSTRALIA and ITALY. | Terminated | July 06, 2012 | 476 | 6925 | Not_Computer
Z-2386-2011 | 58431 | No | OHH | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | AC Power Adaptor (transformer), TENPAO Model U090100Q31; used with the Pump In Style Advanced family of breast pumps; Medela Inc., McHenry, IL 60051-0650; Made in China; Article #9207010. |  |  | April 04, 2011 | 2011 | Medela Inc | The power adaptor (transformer) used with the Pump In Style Advanced family of breast pumps have power cords that may easily become damaged if not properly handled. | Device Design | Medela Inc., sent an " IMPORTANT FIELD ACTION " letter dated April 4, 2011, to all affected customers via The United States Postal Service ( USPS) United Parcel Service ( UPS) and E-mail .The letter identified the product, the problem, and the action to be taken by the customer. It was discovered the power adaptor/transformer used with the Pump in Style Advanced family of breast pumps can develop damage in the form of twists, kinks, bends, bulges or exposed wires in the power cord. Use of the power adaptor with a damaged cord could result in a short in the cord resulting in sparks, smoke, burning and melting of the transformers. If the users have a damaged power adaptor or power adaptor cord, they were requested to discontinue use of the adaptor and visit Medela's website at www.poweradaptorhelp.com to request a replacement adaptor. Questions about the recall process were directed to Medela Customer Service at 1-855-367-0104 or e-mail at poweradaptor@medela.com. Distributor/retailers were requested to post the recall notice within their establishment visible to their customer base, and to notify their customers of the recall. Information on the recall was posted on Medela's website www.medelapoweradaptorhelp.com on April 4, 2011. |  | 478823 | 17,681 units alone, 461,142 units with products | Worldwide Distribution - Nationwide Distribution and Puerto Rico and countries of Canada, Bermuda, Cayman Islands, Hong Kong and Haiti.CHI-DO proposes phone call audit checks of 2 consumers, 2 retailers and 1 hospital account by each district. | Terminated | October 16, 2012 | 561 | 478823 | Battery
Z-2385-2011 | 58199 | No | JEY | Bone plate. | Dental | Dental | 510(k) | ***BIOMET MICROFIXATION BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***6 HOLE REGULAR DOUBLE Y PLATE***"1.5MM SYSTEM PLATE DOUBLE Y-SHAPE, REGULAR 0.6 MM TITANIUM***NON-STERILE PRODUCT*** |  |  | July 28, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling their 1.5 6 Hole Reg Double Y, Lots 785710, 753120, 693500 and 733600. This product is labeled as 01-7110-K (1.5 6 Hole Reg Double Y), however the package contains a 01-7112-K (1.5.6 Hold Long Double Y). | Labeling mix-ups | The firm, Biomet Microfixation, sent an "URGENT MEDICAL SAFETY ALERT NOTICE" letter dated July 28, 2010 to its customers/distributors. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their stock and ensure that the plates were properly distributed into their trays and that they have the sufficient quantity of plates that they need. In addition, they were instructed to complete and return the attached Inventory Reconciliation form via fax to: Attn: Customer Relations Specialist at 904-741-9425 and notify the recipient of this notice, if they have further distributed the product.If you have any questions, please contact the Customer Relations Specialist at 800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 180 | 180 | Worldwide distribution: USA (nationwide) including states of: FL, IN, KY, MT, NC, NE, NY, OH, TX, VA, and WI; and countries including: Argentina, Canada, Colombia, Czech Republic, Japan, Korea, Mexico, Norway, South Africa, Spain and Turkey. | Terminated | June 24, 2011 | 331 | 180 | Not_Computer
Z-2384-2011 | 58632 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator, Model # HT-50-H, HT50-HB, HT50-H1, and HT50-H1B |  |  | July 22, 2009 | 2009 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed that the HT50 Dual Pac Ventilator is not compatiable with the external battery pack (p/n BAT3300A) supplied with the previous single battery ventilator. The Dual Pac system has a slightly higher amperage draw than the single battery and as a result can cause the fuse in the external battery pack to blow making the battery inoperable. | Other | Newport Medical sent an IMPORTANT PRODUCT NOTICE on July 22, 2009, to all affected customers who purchased the Newport HT50 Dual Pac Ventilator with the external battery pack (p/n BAT3300A) supplied with the previous single battery ventilator. Newport Medical also sent an upgrade kit (FSK3301) which includes instructions for changing the fuse on the external battery. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to not use an external battery with the new Dual Pac Ventilator until they have completed the fuse replacement. Customers with any questions about this fuse replacement or require additional upgrade kits, were instructed to contact Newport Medical's Technical Support Department at (800) 451-3111 or (714) 427-5811 x500. |  | 11 | 11, 333 units total | Nationwide Distribution | Terminated | June 01, 2011 | 679 | 11 | Battery
Z-2380-2011 | 58346 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Material ULT2175, ULTIMATE LINE TRAY, CENTURION ULTIMATE LINE TRAY, REORDER ULT2175, STERILE, LATEX FREE, CENTURION MEDICAL PRODUCTS CORP. WILLIAMSTON MI 48895, USA 800.248.4058 www.centurionmp.com |  |  | March 31, 2011 | 2011 | Centurion Medical Products Corporation | This voluntary recall was initiated because the vials of Bacteriostatic Sodium Chloride Injection, USP, 0.9% utilized to make the products may contain visible particulates. | Material/Component Contamination | Centurion Medical Products Corporation issued press on April 1, 2011, and sent an Urgent Voluntary Recall letter dated March 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the kits immediately and instructions were provided for the return of the kits.Customers were also asked to complete a form to fax or email back to the firm regarding the number of kits the firm has in their possession.Customers were instructed to provide recall notification to other firm's who received the recalled kits from their direct account.For any questions regarding this recall call 631-924-4000, ext 142. |  | 7160 | 7160 | Nationwide Distribution | Terminated | September 13, 2011 | 166 | 71471 | Not_Computer
Z-2392-2011 | 58665 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | URI-DRAIN Medium Male External Catheter With Two-Sided Adhesive Strap and Skin Protective Wipe |  |  | April 29, 2011 | 2011 | Covidien LP | A component of the Uri-Drain device recalled the Skin-Prep (manufactured by H&P; Industries-Triad) a protective wipe due to bacterial contamination | Other | The firm, Covidien, sent an "URGENT PRODUCT RECALL" letter dated April 29, 2011 to its customers via US Postal Service. The letter described the product, problem and actions to be taken. The customers were instructed to immediately stop using the product, place in quarantine, and return any inventory. Distributors are requested to notify their customers. All customers are required to respond to the notice regardless of remaining inventory. Responses will be tracked for follow-up and effectiveness checks.For Customer Service questions regarding the recall please telephone: 1-800-962-9888 option 1 (weekdays, 8am to 6:30pm Eastern) and email to: sdfeedback@covidien.com. Should you have any medical questions concerning this recall, please call 508-261-6393. |  | 5760 | 5760 units | Worldwide distribution: USA (nationwide) and countries including: Canada and Bermuda. | Terminated | May 29, 2015 | 1491 | 62352 | Not_Computer
Z-2395-2011 | 58518 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ES0101, AEM Disposable Curved Scissors, 1/2", 35cm, Sterile R, Encision Inc., Boulder, CO 80301 |  |  | April 07, 2011 | 2011 | Encision, Inc. | Sterility may be compromised due to holes in packaging. | Package design/selection | Encision, Inc. sent an Important Safety Notice letter dated April 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to either carefully inspect any remaining pouches for readily visible holes in the Tyvek pouch, or return the affected lots to Encision for inspection.Customers were also asked to fill in and fax back the attached Confirmation to 303-444-2693. |  | 598 | 574 Ten-Count Boxes, 24 Kits | Worldwide Distribution - USA (nationwide) and the countries of Australia and New Zealand | Terminated | August 09, 2011 | 124 | 2052 | Not_Computer
Z-2401-2011 | 58568 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Xpert Infinity Software, V 4.1a. |  |  | April 01, 2011 | 2011 | Cepheid | A mismatch of Patient ID or Sample ID and the Host Order from the hosted LIS and its requested test results could occur when running in batch mode and cartridges are introduced in rapid succession. | Software design | Cepheid sent an Corrective Action Notification letter dated April 13, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cepheid recommended customers using 4.1a software on the INFINITY 48 System that is interfaced to a LIS with Host Ordering to make these alterations in operationof the system.After scanning the Patient ID (optional). Sample ID and Cartridge barcode, the host test order information should apear highlighted in the Order Test workspace.Once the highlighted host order information appears select SUBMIT.Warning. If the host test order informaiton is not highlighted, cancel the order by selecting CLOSE. Re scan Patient ID (optional). Sample ID and cartridge barcode and verify highlighted host test order before SUBMIT. Contact Cepheid Technical Support at 888-838-3222 in the US or 33 5 63 82 53 19 outside of the US for any questions regarding this Corrective Action Notice. |  | 19 | 16 US, 3 ROW | Worldwide Distribution - USA including CA, UT, TX, NJ, FL, NC, CO, IN, VT, MN and the countries of the UK, Denmark and Saudi Arabia. | Terminated | June 20, 2011 | 80 | 19 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-2403-2011 | 58636 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron UMPC is a component in the DataCaptor" Connectivity System. |  |  | August 25, 2010 | 2010 | Capsule Tech Inc. | Software: Data cached on the Neuron due to server unavailability is not being delivered to the server after server availability is restored.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes.Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power source | Software design | Capsule Tech initiated three corrective action notifications for the Neuron" UMPC units three failure modes as follows:a) Data cached on the Capsule Neuron" UMPC due to server unavailability is not being delivered to the server after server availability is restored:i) A customer advisory was provided to a subset of consignees, along with DataCaptor software version 6.6.5 on December 1, 2010.ii) A customer advisory was provided to the remaining consignees, along with DataCaptor software version 6.6.7 on January 27, 2011.If a compact flash failure occurs, the file system may be corrupted and result in system hangs, lock-ups or crashes:i) A customer advisory was provided to all consignees on March 23, 2011, via phone call, email and FedEx;.The correction will be provided as a software update. A formal recall notification letter will be provided to all consignees alerting them to the availability of the correction once it is released, via phone call, email and FedEx. c) Batteries can discharge completely in a Capsule Neuron" UMPC even though it is connected to power. Note: the Capsule Neuron" UMPC operates normally when connected to power.i) A customer advisory was provided to all consignees on April 23, 2010, via FedEx.For questions regarding this recall contact Capsule technical support team at (800) 260-9537 or e-mail support@capsuletech.com.. |  | 3144 | 3,144 Capsule Neuron UMPC units | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | March 20, 2012 | 573 | 3144 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2402-2011 | 58557 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cholestech LDX Lipid Profile"GLU test cassettes, Catalog Number 10-991, Lot # 245104 |  |  | February 18, 2011 | 2011 | Biosite Inc Dba Innovacon Inc. | The recall was initiated due to an issue related to an error in the encoding of calibration coefficients in the sample cassette that cause the Cholestech LDX analyzer to report glucose values as &gt;500 mg/dL regardless of the true value of the specimen. | Pending | Alere San Diego, Inc ( formerly known as Biosite Incorporated sent an URGENT MEDICAL DEVICE RECALL letter on February 18, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were requested to:(1) Discontinue use and/or sale of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 245104.(2) Discard in accordance with your local regulations, all units of the above listed product.(3) Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation.(4) If you have forwarded the product listed above to another laboratory, please provide a copy of this letter to them.(5) Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice, to indicate the number of boxes of test cassettes discarded from your inventory and replacement product will be sent to you.For questions about the information contained in this notification, customers were instructed to contact Alere Technical Services at 1-877-308-8289 (24 hours a day, 7 days a week) or via e-mail at service.cholestech@alere.com. |  | 5305 | 5,305 Boxes | Worldwide Distribution - USA ( nationwide ) including the countries of South America ( EL Salvador) , Europe ( Denmark, France, Norway and Switzerland), Bahrain, Kuwait, South Africa, and United Arab Emirates. | Terminated | April 02, 2012 | 409 | 5305 | Not_Computer
Z-2407-2011 | 58267 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above |  |  | May 27, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: For a specific clinical setup, the MLC leaf positions on the exported DRR images might not match the MLC leaf positions displayed in XiO. | Software design | The firm, Elekta, issued an "important Safety Notice" undated, to its customers. The notice indicates the problem and the clinical impact. There is no reasonable workaround for this issue. The notice states a future XiO Release will resolve the issue and the customers will be notified when this solution is available. A return postcard was included with the mailing to verify consignee received the notice. For sites not returning the postcard within one month, the recalling firm will contact them via phone or email to confirm receipt of the notice. The customers were instructed to distribute this notice to any and all users of the CMS software at their organization who are potentially affected by this issue.If you have any questions, please call 314-993-0003 or 800-878-4267. |  | 415 | 415 | Nationwide distribution, including Puerto Rico, to medical facilities. Distribution was also made to Military and Government consignees. Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, and United Kingdom. | Terminated | April 10, 2014 | 1414 | 415 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-2405-2011 | 58664 | No | EXF | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Drainage Bag, Model 600D, 600 mL, Single use, Ethylene Oxide sterilized, Remington Medical, Inc., 6830 Meadowridge Court, Alpharetta, GA 30005. |  |  | January 12, 2011 | 2011 | Remington Medical Inc. | There have been complaints of the Drainage Bags leaking from the outlet and the inlet port. | Other | Remington Medical Inc.sent an Urgent Medical Device Recall letter dated January 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine product subjec to recall. If they further distributed this product, they were to identify their customers and notify them at once of the product recall. Notification to their customers should include a copy of the recall notification letter. Customer were to receive a Return Authorization Form, reference RGA #600-D.For questions regarding this recall call 800-989-0057, ext. 213. |  | 89294 | 89294 units | Nationwide Distribution | Terminated | September 13, 2011 | 244 | 89294 | Not_Computer
Z-2406-2011 | 58619 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 Ventilator. Model #HT70, |  |  | September 20, 2010 | 2010 | Newport Medical Instruments Inc | The recall was initiated because Newport Medical has confirmed a limited number of power cords recently received from their supplier, Glob-Tek, have a slightly oversized plug connector which prevents thecord from locking into place. These cords are fully functional as supplied. However, the cord can be easily disconnected. If the cord is disconnected, the HT70 will alert the user with a power di | Other | Newport Medical sent an PRODUCT NOTICE LETTER, dated September 20, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to discard the original cord and only use the replacement cord to avoid unintentional disconnection of the power cord. Each affected Customer was sent a replacement cord.Customers with any questions regarding the Product Notice were instructed to contact the Technical Support Department at (800) 451-3111. |  | 22 | 22 units total | Worldwide Distribution--USA (nationwide) including the states of MA, PA, OH, IA and Puerto Rico and the countries of Colombia, UAE, Hong Kong, Hungary, India, Japan, Nepal, China, Philippines, Singapore and South Africa. | Terminated | June 03, 2011 | 256 | 22 | Battery
Z-2404-2011 | 58637 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC (Multi Leaf Collimator). Linear Accelerators that may be affected are: |  |  | April 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety risk related to the behavior of the 160 MLC when using the "Motion Stop" button followed by "Reset" several times during the treatment of one beam or one segment. | Other | Siemens sent a Consumer Safety Advisory and Urgent Medical Device Correction letter to all affected customers beginning on April 15, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Updated instructions for using the Motion Stop button with the 160-MLC were provided with the letter. Customers were requested to use the Motion Stop button and the corresponding reset procedure only when required. In the unlikely event that the Motion Stop button needs to be used repeatedly during treatment of the same beam or segment, customers were advised to perform an "in session resumption of the beam by selecting 'abort' at the syngo RT Therapist and then re-selecting the beam again for the remaining delivery. In this case, the beam will be re-loaded with the originally prescribed leaf positions. Please consider the already delivered dose for resumption.Customers were asked to include the Customer Advisory Notice in their LINAC System Owner Manual in chapter 'Safety Advisory Letters' where it should remain until the measures defined above were performed. Customers were asked to inform all affected personnel immediately. |  | 221 | 221 active devices | Worldwide Distribution - USA including CA, DE, FL, IL, KY, LA, MA, MI, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV. | Terminated | March 16, 2012 | 336 | 221 | Not_Computer
Z-2409-2011 | 58681 | No | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SureCath Set with Accessories, Coloplast Corp.-MN 06093, Sterile EO., Manufacturer Coloplast -A/S 3050 Denmark. |  |  | April 21, 2011 | 2011 | Coloplast Manufacturing US, LLC | Coloplast has initiated a recall on SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies because they contain Iodine prep pads which may expire before the overall product expiration date. SpeediCath with Insertion Supplies and SureCath Set with Insertion Supplies are both packaged with a set of accessories which includes Iodine prep pads. The Iodine prep pads included i | Incorrect or no expiration date | Coloplast Manufacturing US, LLC sent a Coloplast "Product Field Correction Notice" dated April 21, 2011. The letter was addressed to "Dear Valued Customer". The letter described the problem and the product involved in the recall. Customers were advised to provide the current inventory levels of the lots affected and the firm will be providing them with new iodine prep pads with sufficient shelf life. This product will include customer instructions, informing the customer to discard the original prep pads and use the new prep pads instead. If they prefer, they could themselves send the affected product back to Coloplast and they in return will attach new prep pads to the product for them.If they have further distributed the product, they advised consignees to contact their customers and to include the letter provided to them to inform them of the recall. |  | 438412 | 438,412 | Nationwide Distribution including AL, AZ, CA, CO, CT, DE, FL,GA, ID, IL, IA, MD, MA, MI, MN, MS, NJ, NY, NC, OH, OK, OR, PA,PR, SC, TN, TX VA, WA, WI. | Terminated | April 24, 2012 | 369 | 564759 | Not_Computer
Z-2411-2011 | 58785 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eleganza 3 Hospital Bed |  |  | May 02, 2011 | 2011 | Linet Americas Inc | Linet has been made aware that there has been breakage of the Mobi-lift handle on a few beds which may cause failure during patient assist use. | Device Design | The firm, Linet America, notified US customers only via phone on May 5, 2011. A follow up "service advisory information" letter dated May 11, 2011 was sent to all customers. The letter described the product, problem and action to be taken. The customers were instructed to remind/retrain staff to ensure that the Mobi-lift handles are always in the stowed position during patient transport and after use, and complete and return the attached Inspection procedure for each affected Eleganza 3 bed via mail to: Linet Americas, 10420-R Harris Oaks Blvd., Charlotte, NC 28269 or fax to: 704-248-5655. Linet Americas will be receiving a supply of new Mobi-lift design and will begin the process of executing service calls to replace the Morbi-lift handles with the new design.If you have any questions, contact the Director of Quality and Service Operations at 704-248-5650 or 980-406-6446. |  | 1601 | 1601 units | Worldwide distribution: USA (nationwide) including states of: KY, MA, NC, and NY; and country of: Canada. | Terminated | July 19, 2012 | 444 | 1601 | Not_Computer
Z-2408-2011 | 58205 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.0.2 and above. |  |  | September 10, 2010 | 2010 | Computerized Medical Systems Inc | When RT Plans are exported from XiO and RT Structure Sets and Images are exported from Monaco, the DICOM iso-center coordinates may not be accurate. | Software design | Elekta / Computerized Medical Systems Inc began mailing an Important Safety Notice to all affected customers in September 2010. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised not to mix treatment plan information from multiple planning systems. Customers were asked to return the postcard as proof of receipt and their understanding of the Safety Notice. For questions regarding this recall call 408-830-8000. |  | 1377 | 1377 | Worldwide Distribution - USA (nationwide and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 1306 | 1377 | Other |  |  | Other
Z-2412-2011 | 58735 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DuraspanTM Dialysis Catheters, Ultra and Long Term, 15.5 Fr, Sterile EO, Manufactured for r4 Vascular, Inc . 7550 Meridan Circle North, Suite 150, Maple Grove, MN 55369. |  |  | May 03, 2011 | 2011 | R4 Vascular Inc | r4 Vascular, Inc. is conducting a voluntary recall of specific lots of Duraspan Dialysis Catheters. Duraspan catheters are 15.5 Fr dual lumen chronic hemodialysis catheters. Over a six month period, seven patients at five hospitals experienced separation of the luer connector from the extension tube on the Duraspan catheter. Testing during failure analysis disclosed that these lots and other lot | Nonconforming Material/Component | R4 Vascular Inc sent an Urgent Medical Device Recall letter dated May 3, 2011, to all affected customers. The letter described the problem and the product being recalled. Consignees were instructed to stop distributing the product and to quarantine the recalled lots and if they had already distributed the affected product to seek its return from their customers. A template letter to physicians and healthcare facilities was enclosed for Distributors/Retailers to send or hand deliver to their customers.For questions regarding this recall call 1-704-905-2105. |  | 2489 | 2489 kits | Nationwide Distribution including CA, CO, CT, FL, GA, IL, LA, MD, MA, MN, MS, MO, NJ, NC, OH, PA, SC, TX, and WA. | Terminated | July 27, 2012 | 451 | 2489 | Not_Computer
Z-2415-2011 | 58677 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Pinnacle Destination Guiding Sheath, 6 French, 90 cm, Straight Tip, Tuohy-Borst Valve, Lot MD24, MD31, ME07, &amp; ME14. 510(k) K012812 |  |  | June 24, 2010 | 2010 | Terumo Medical Corporation | Molding defect in destination product that render the device unusable. | Equipment maintenance | Terumo Medical Corporation sent an Urgent Product Bulletin Recall letter dated June 24, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo requested customers to immediately discontinue use, and quarantine any available stock of the listed product codes. In order to facilitate the exchange of these devices, a Terumo Territory Manager contacted all customers by July 9th to confirm receipt of the Terumo Destination Bulletin and scheduled an appointment to visit each customer in person by July 31st to personally manage the exchange. For questions or for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 214 | 214 units | Worldwide Distribution - USA including AL, AK, AZ, TX, LA, MS, MN, MO, CA, GA, FL, HI, KS, IA, ID, IL, TN, NM, OR, OK, WA, and WI and the countries of Belgium, Brazil, and Panama | Terminated | July 07, 2011 | 378 | 258 | Not_Computer
Z-2413-2011 | 58591 | No | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ARCHITECT HAVAB-M Assay Reagent; a chemiluminescent microparticle immunoassay (CMIA) for the qualitative detection of IgM antibody to hepatitis A virus (IgM anti-HAV) in human adult and pediatric serum and plasma (dipotassium EDTA, lithium heparin, and sodium heparin) and neonatal serum; Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA, List 6L21-25 |  |  | April 20, 2011 | 2011 | Abbott Laboratories | The HAVAB-M reagent lot has reduced specificity and may cause an increase in false grayzone and/or false reactive results. | Nonconforming Material/Component | Abbott Laboratories issued a Product Recall letters dated April 20, 2011 to all affected customers that received ARCHITECT HAVAB-M reagent, lot 93794HN00. The letters informed the accounts of the increased grayzone and/or reactive results and were instructed to discontinue use of the lot and destroy any remaining inventory of the lot according to laboratory procedures. If the account did not have an alternate lot of reagent, they were told to continue using the recalled lot until their lab receives the replacement lot, and to perform supplemental testing with another HAVAB-M method to confirm the results for all specimens that are grayzone or reactive. The accounts were also requested to complete and fax back to Abbott the enclosed customer reply sheet, acknowledging the receipt and understanding of the recall. For further questions regarding this please contact Abbott at 1-877-422-688 |  | 788 | 788 kits | Nationwide Distribution including Puerto Rico | Terminated | February 10, 2012 | 296 | 788 | Not_Computer
Z-2420-2011 | 58658 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics OrthoPAT -Orthopedic Perioperative Autotransfusion System |  |  | April 29, 2011 | 2011 | Haemonetics Corporation | 2 Models Model 1:OrthoPAT devices with no centrifuge drain spill containment system may have potential for fluid contamination leading to potential spark or smoke may occur are recalledModel 2:OrthoPat equipped with a spill collection drainage system updated with revised Instructions for Use and Quick reference Guide Addendum | Device Design | The firm, Haemonetics, sent an "Urgent Field Recall" letter dated April 29, 2011 to all customers. The letter describes the product, problem and actions to be taken. This recall covers the 2 models. Model 1: OrthoPAT devices without drain ports (Serial numbers 00001 through 02279-T). All OrthoPAT devices without drain ports worldwide, will be removed from distribution and scrapped. Model 2: OrthoPAT devices with drain ports (serial numbers beginning with 02280-T and higher). All OrthoPAT devices with drain ports will remain in clinical use and will receive updated labeling.The customers were instructed to follow the enclosed directions to confirm their possession of any units of this model and arrange for their return to Stericycle; immediately cease use of devices not equipped with a spill collection drainage system ( to determine whether their OrthoPAT device is equipped with the spill collection drainage system, please refer to the attached illustrations/instructions); complete and return the IMPORTANT-CUSTOMER CONFIRMATION sheet via fax to: Stericycle at 1-866-792-5450, and RETURN the OrthoPAT devices NOT equipped with spill collection drainage system (for instructions on how to arrange return of OrthoPAT Devices, call a Stericycle Representative at 1-866-918-8736). A Haemonetics representative will be in contact with the customers facility within 14 working days to assist with the recall actions.The firm, Haemonetics, will update the OrthoPAT labeling. The labeling will have specific instructions to permanently deploy the centrifuge spill collection bag at all times. Device cleaning instructions in the event of a blood spill are also clarified and include specific instructions to unplug the device from the AC outlet in the event of a fluid spill on the device. If a device does experience a fluid spill, new instructions direct the user to return the device to Haemonetics for evaluation and repair. Operator Manuals will be replaced with a new revisi |  | 2976 | 2976 units | Worldwide distribution: USA (nationwide) and countries including: Austria, Australia, Belgium, Switzerland, Canada, China, Czech Republic, Germany, Spain, Finland, France, Great Britain, Greece, Guatemala, Ireland, Israel, India, Italy, Japan, Lithuania, Luxembourg, Latvia, Monaco, Netherlands, Norway, Panama, Poland, Puerto Rico, Russian Federation, Saudi Arabia, Singapore, Slovak Republic, EI Salvador, Turkey and Taiwan., | Terminated | April 21, 2014 | 1088 | 2976 | Not_Computer
Z-2467-2011 | 58703 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Hitachi HIVISION Preirus/Avius Ultrasound equipped with V03.02 Software - Doppler Trace Measurement program. 510 (k) Numbers: Preirus - K093466 (6/1/10), Avius - K102901 (10/27/10). |  |  | April 18, 2011 | 2011 | Hitachi Medical Systems America Inc | Software anomaly. The Doppler Trace Measurement feature on certain models of the firm's diagnostic ultrasound systems, may, under certain conditions whereby multiple measurements are done sequentially, record incorrect values. The recorded value is displayed to the user on a report; however depending upon the measurement area, displayed velocity values could be lower than at the true velocity lead | Software design | The firm, Hitachi, sent a "DEVICE CORRECTION NOTICE" letter dated April 15, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were inform of the nature of the software error in the Doppler Trace Measurement feature; the circumstances under which this error can occur; and the safety issue involved should incorrect values generated lead to a false diagnosis of stenosis in a patient. The letter provides the customer with a workaround solution in order to avoid this issue. The customer were also instructed to call the HMSA Applications Helpline at 877-871-2777 to reprogram their system to temporarily eliminate the noted safted issue and to immediately have all selections unchecked on the Doppler trace measurement result store setting set up screen. The notification letter also states that Hitachi Medical Systems America (HMSA) will be having their service personnel schedule a visit at each customer location and install a software upgrade on each affected system unit in order to directly address, and permanently resolve this issue. If you need any further information regarding this notice, please contact the Manager of Regulatory Affairs at 800-800-3106 x3720 or email to: thistlewaited@hitachimed.com. |  | 24 | 24 system units | Nationwide distribution: USA including states of: CA, MD, KY, FL, MN, NJ, IA, OH, TX, IL, and IA. | Terminated | February 19, 2013 | 673 | 24 | Software | Output/Calculation | Software update | Software Update
Z-2468-2011 | 58620 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator, Model Number: e360 |  |  | October 27, 2009 | 2009 | Newport Medical Instruments Inc | Cregarding e360 ventilators that were shipped with a different software version from what the customers ordered. | Mixed-up of materials/components | Newport Medical Instruments Inc. initiated a field correction on October 27, 2009, to address complaints received regarding e360 ventilators that were shipped with a different software version from what the customers ordered. Newport Medical Instruments provided the correct software to the Newport sales or service providers who were responsible for the units affected. They installed the correct software into the ventilators at the customer sites. No written Correction Notice was provided since the corrections were done at the customers' request.Should you have any questions or require additional information, please call (714) 427-5811 extension 344. |  | 18 | 18 | Worldwide Distribution--USA (nationwide) including states of CA, KY, and Silver City, New Mexico and country of Quito, Ecuador. | Terminated | June 08, 2011 | 589 | 18 | Software |  | Software update | Software Update
Z-2423-2011 | 58663 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Ringloc + Acetabular Shell limited hole finned 50 mm Size 23 porous coat, for uncemented use, sterile, REF 16-104150, Biomet Orthopedics, Warsaw, IN. |  |  | April 14, 2011 | 2011 | Biomet, Inc. | The firm has initiated this action following an investigation which found that the locking ring may be assembled incorrectly in the acetabular shell. If the locking ring is incorrectly assembled in the acetabular shell, the liner cannot be completely seated in the shell. | Employee error | Biomet Orthopedics sent an URGENT MEDICAL DEVICE RECALL NOTICE DATED April 14, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Immediately locate and remove the affected product.Follow the instructions on the enclosed "FAX Back Response Form" and fax a copy of the Response Form to 574-372-1683 prior to return of product. Customers were to confirm receipt of the notice by calling 800-348-9500, ext 3755 or 3756. For any questions regarding this recall call 574-371-3755 or 574-371-3756. |  | 762 | 762 all products | Worldwide Distribution - USA (nationwide) and the countries of Australia, Netherlands and Japan. | Terminated | September 16, 2011 | 155 | 2069 | Not_Computer
Z-2541-2011 | 58571 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bio-modular reverse shoulder humeral tray with lock ring, 44 mm 10 degree angle, sterile, single use, REF 11-113685, Biomet Orthopedics, Warsaw, IN. |  |  | April 22, 2011 | 2011 | Biomet, Inc. | The firm is initiating this recall following an investigation which identified the rare possibility that certain units of Comprehensive Reverse Shoulder Humeral Trays may contain a locking ring that is incorrectly assembled. | Employee error | The firm, Biomet Orthopedics, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 22, 2011 to their customers. The letter describes the product, problem, and actions to be taken. The customers were instructed to immediately locate the recalled product; discontinue use; return the product back to Biomet; confirm receipt of the notice by calling 800-348-9500 extension 1368 or 1325, and complete and return the attached form via fax to: 574-372-1683 or mail to: Biomet, 56 East Bell Drive, Warsaw, IN 46580.Customers with questions regarding this recall should call 800-348-9500 extension 1368 or 1325. |  | 1575 | 1,575 units all products | Worldwide distribution: USA (nationwide) including: Puerto Rico and countries including: Australia, Canada, Chile, Columbia, Costa Rica, Korea and Netherlands. | Terminated | February 21, 2013 | 671 | 1575 | Not_Computer
Z-2552-2011 | 58263 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV (Cardio Vascular) |  |  | March 21, 2011 | 2011 | AGFA Corp. | Customer error with the angiography (XA: X-Ray Angiography) measurement calibration functionality within the Cardiovascular Review Station resulted in improper calibration. | Other | The firm, AGFA Healthcare, sent an "URGENT SAFETY NOTICE" letter dated March 21, 2011 via FED-EX to its customers.. The letter describes product, problem, and actions to be taken. The customers were instructed to distribute this information within their facility to all individuals who need to be aware and to complete and return the URGENT SAFETY NOTICE FEEDBACK FORM via fax to: 864-421-1664 or email. AGFA is supplying the customers with a CRS User Manual Addendum to emphasize the proper steps required to achieve an accurate measurement when using the calibration tool within CRS. Mitigation will provide the customer/end-user with awareness of the correct XA measurement calibration technique and the potential consequences if proper technique is not followed. Additionally, AGFA will offer customer training on the use of the calibration tool. This training will be provided remotely via the web at no cost. Customers were ask to register before May 1, 2011 by contacting AGFA Customer Support Center toll-free at 877-777-2432. Training sessions are from May-June 2011.If you have any questions about this matter, please contact your local AGFA HealthCare organization or the AGFA Customer Support Center at 877-777-2432. |  | 536 | 536 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | November 29, 2011 | 253 | 536 | Not_Computer
Z-2419-2011 | 58405 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Atlantis Âż SR Pro 2, 40 MHz Coronary Imaging Catheter, 3.6F (1.18mm) x 135 cm, USE BY 2011-05, REF 39014, UPN H749390140, Sterile R Sterilized using irradiation, Made in USA 47900 Bayside Parkway, Fremont, CA 94538 |  |  | March 28, 2011 | 2011 | Boston Scientific Corporation | Catheter tip detachment that is not associated with elongation of the catheter typically caused by excessive pulling forces. EXPANDED: Boston Scientific is initiating a recall of all iCross Coronary Imaging Catheters because they have received and confirmed 8 occurrences of catheter tip detachments due to embrittlement. | Process design | Boston Scientific " Important Customer Notification" letters dated March 14, 2011 were hand delivered to customers by Boston Scientific sales representatives starting on March 28th, 2011. The sales representatives verbally reviewed information. A tracking log will be maintained document completion of the customer notifications.The letter described the product, problem and possible root cause as well as one corrective action, a change in sterilization method from Gamma to E-beam implementation.Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter dated May 26, 2011, to all affected customers. Customers were instructed to Identify and discontinue use of any affected product within their inventory and to segregate it immediately and return it to Boston Scientific |  | 30669 | 30,669 | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Japan, Trinidad, Tobago, and American Virgin Islands. | Terminated | July 30, 2012 | 490 | 96702 | Not_Computer
Z-2551-2011 | 58736 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte@ Sodium Bicarbonate Liquid Concentrate, Dialysate Concentrate for Hemodialysis Liquid |  |  | March 23, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Mislabeled: incorrect lot number printed on case labels | Employee error | Fresenius sent an CUSTOMER NOTIFICATION LETTER letter dated April 18, 2011, to all affected customers via Certified Mail, Return Receipt Requested. Fresenius also contacted affected customers by telephone April 18, 2011, The letter identified the product, the problem and the action needed to be taken by customer. Customers were instructed to examine their stock to determine whether they have the incorrect case label and are asked to remove thecorrectly labeled product from the incorrectly labeled case and discard the case.Acknowledgement forms were Fax-back forms for clinics and mail-back forms (including self-addressed, stamped envelopes), for home patients.Question or concerns should be directed to Fresenius Customer Service Team at 1- 800.323.5188. |  | 3190 | 3,190 cases | Nationwide Distribution--AL, CA, CT, FL, GA, NJ, NV, NY, OR, PA, and WA | Terminated | September 12, 2012 | 539 | 3190 | Not_Computer
Z-2553-2011 | 58730 | No | HFS | Nipple shield. | General Hospital | General Hospital | 510(K) Exempt | Contact Nipple Shield (Sterile) - 24mm (Standard); Individually wrapped silicone contact nipple shields for immediate use for latch-on difficulties while breastfeeding the baby under professional care only; Manufactured for Medela, Inc., P.O. Box 660, McHenry, IL 60051-0660, Made in Switzerland; article #67203S |  |  | May 18, 2011 | 2011 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | Packaging | The firm, Medela, sent an "Important Device Recall" letter dated May 9, 2011 to all of their customers/direct accounts. Medala also began telephoning or visiting the direct accounts/customers on May 18, 2011. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1) not to use the affected lots and place them in quarantine2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form inlcude with this letter4) complete and return the Product Inventory Form via fax to 800-995-7867 or email jim.hughes@medela.com, even if they have no affected product on hand5)If you have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. Each affected product will be replaced with (2) non-sterile nipple shields.5) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. Questions about the recall process were directed to Medela Customer Service at 1-888-835-5968 or e-mail at medelareturns@medela.com. Distributor/retailers were requested to notify their customers of the recall, using the enclosed customer letter. |  | 4320 | 4,320 units | Nationwide distribution: USA including states of: AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, MA, MD, ME,MI, MN, MS, MT, NC, ND, NH, NJ, NY, OH, OK, PA, SC, TX, UT, VA, VT, WA, WI, and WV. | Terminated | February 16, 2012 | 274 | 4320 | Not_Computer
Z-2388-2011 | 58794 | No | LKB |  |  | General Hospital | Enforcement Discretion | Churchill Medical Systems |  |  | April 19, 2011 | 2011 | Vygon Corporation | Skin-Prep wipes purchased by Churchill (a Vygon company) for inclusion into Dressing Kits and PICC Insertion Kits were recalled by the manufacturer due to potential bacterial contamination. | Nonconforming Material/Component | A recall letter printed on Vygon letterhead dated April 20, 2011 and specific to each customer was sent via overnight courier mail. The letter described the affected product, issue and actions to take. Customers were instructed to check their inventory for the recalled products, cease use and distribution, and quarantine all affected product. Each customer received a custom copy of the letter with the specific product codes and lots that were shipped to them described in the Recall Acknowledgement and Inventory Return Form. Distributors were provided with instructions on how to handle product that was further distributed by them. The firm issued an updated Recall Notice dated May 31, 2011 to two consignees effected by the expanded recall.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-mail at customerservice@vygonus.com. |  | 7900 | 7,900 units | United States (FL, MA, MD, and VA). | Terminated | June 18, 2012 | 426 | 15400 | Not_Computer
Z-2414-2011 | 58661 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Datascope/ Maquet IABPs. System 98/98XT Intra-Aortic Balloon Pump; CS100/CS100i Intra-Aortic Balloon Pump; CS300 Intra-Aortic Balloon Pump; |  |  | March 16, 2011 | 2011 | Maquet Datascope Corp - Cardiac Assist Division | Several models of the Intra-Aortic Balloon Pumps are being recalled due to a potential problem with the fan's assembly which may cause the power supply to detect an overheating event and shut down. | Device Design | Maquet sent an Urgent Product Field Correction letter dated March 16, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to note the WARNING in the IABP Operating Instruction. Customers were advised to refer to the IAB Instructions for Use, Section D (Manually inflating and Deflating a Catheter). Section D of the IAB instruction for Use reiterates that an IAB should not remain inactive for more than 30 minutes and also provides detailed instruction on how to keep an IAB active during IABP failure. Customers were advised that their Service Representatiave would arrange to replace their supply which would contain a new fan assembly. Upon completion of the power supply replacement, customers were asked to sign a service repair order to verify satisfactory completion of the work. For technical questions contact Technical Support Department at 1-800-777-4222 and press 4.For scheduling a service visit by a Field Service Representatiave, contact the firm's US Call Center at 1-800-777-4222 and press 3.Maquet issued a press release on June 6, 2011 advising consumers of the issue and that their recall was considered a class one. |  | 885 | 885 units | Worldwide Distribution | Terminated | October 25, 2012 | 589 | 885 | Not_Computer
Z-2557-2011 | 58455 | joint, hip, external limb component | ISL | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Accolade TMZF Plus Hip Stem: |  |  | March 01, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Accolade TMZF Plus Hip Stems are being recalled due to grit blast media which was observed in the drive hole. | Employee error | Stryker Orthopaedics sent an Urgent Product Recall letter dated March 1, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax back the attached Product Recall Acknowledgement Form within 5 days to 201-831-6069.Customers were also instructed to:Contact the hospitals in their territory that have the affected product to arrange return of the product.Return the affected product via and Inter-Org Transfer to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey 07430. Attach the fluorescent orange PRODUCT REMEDIATION sticker to their return, indicating the Product Remediation # RA 2009-459. Mark the outer box with the words "Product Recall."For any questions call (201) 972-2100. |  | 44 | total of all units is 44 | Nationwide Distribution | Terminated | October 10, 2012 | 589 | 44 | Not_Computer
Z-2561-2011 | 57760 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.03.00 and above. |  |  | January 13, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: For VMAT plans, the measured dose of radiation delivered is lower than the software calculated dose by about 5% in Monaco Release 2.03.00 and 2.04.00. | Software design | Computerized Medical Systems, Inc. sent a "USER NOTICE" on January 13, 2011, to all affected users. This notice identified the product, the problem, and the action needed to be taken by the customer. A return postcard is included for the customer to confirm receipt of the user notification. The firm intends for correction in the dose calulation engine to appear in the Monaco Release 3.00 in March 2011. Any further questions please call (403) 830-8023. |  | 94 | 94 | Worldwide Distribution-- USA (nationwide) and countries of Austria, Australia, Belarus, Canada, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, Poland, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand and UK. | Terminated | April 03, 2014 | 1176 | 94 | Software | Output/Calculation | Software update | Software Update
Z-2562-2011 | 57845 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning Workstation. Focal Release 4.51.00 and above. |  |  | November 09, 2010 | 2010 | Computerized Medical Systems Inc | Focal: If a plan involving Dynamic MLC Delivery is sent from XiO to Focal and called up there in Plan Review, the angle of the gantry is set to zero for all beams. If this plan is then sent from Focal to the R&V; System, the gantry angles all remain at zero which could lead to patient mistreatment. | Software design | Elekta CMS Software sent a USER NOTICE dated November 5, 2010, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends to correct the software so it does not modify anything about the plan sent from XiO.Customers were instructed to contact their local customer support office with any questions. |  | 155 | 155 | Worldwide Distribution - USA (nationwide) and the countries of Algeria, Australia, Belarus, Belgium, Brazil, Canada, Chile, China, Egypt, Estonia, Finland, France, Germany, Hong Kong, India, Italy, Latvia, Libya, Lithuania, Mexico, Morocco, Netherlands, New Zealand, Philippines, Poland, Romania, Singapore, Slovenia, South Africa, Spain, Suriname, Switzerland, Taiwan, Turkey, Turkmenistan, and the United Kingdom. | Terminated | April 02, 2014 | 1240 | 155 | Software | Output/Calculation | Software update | Software Update
Z-2559-2011 | 58226 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.34.00 through 4.50.00. |  |  | June 12, 2009 | 2009 | Computerized Medical Systems Inc | Focal Software: When a patient is defined as prone and the shifts are manually entered and the LOCK SHIFT box is checked the X coordinate (left/right) and the Z coordinate (anterior/posterior) are reversed (left to right, anterior to posterior). This could potentially lead to incorrect isocenter definitions from prone patients and thus a mislocation of dose. | Software design | Computerized Medical System sent a CUSTOMER ADVISORY dated June 12, 2009, to all affected users. This advisory identified the product, the problem, and the action needed to be taken by the customer. There were no patients mistreated as a result of this issueThe firm will send out a User Notice to all sites still runningaffected software to remind them that the issue resolution is available.For further questions, please call (403 ) 830-8023 |  | 462 | 462 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and the countries of Algeria, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Egypt, Estonia, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, Nicaragua, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, and Venezuela. | Terminated | April 09, 2014 | 1762 | 462 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-2560-2011 | 58229 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 1.0.0 through 1.0.2. |  |  | June 30, 2008 | 2008 | Computerized Medical Systems Inc | Monaco: If the user selects Monte Carlo Photon as the Secondary Algorithm (to be used for final dose calculation through segments), and selects SKIP Forward button at any time during stage 2 optimization, the system may only use the pencil beam algorithm for final computation of dose even though Monte Carlo was selected. Other than specific output to the console window, there will be no other in | Software design | Computerized Medical Systems, Inc. sent a "MONACO CUSTOMER ADVISORY" letter dated June 2008 to all affected customers. The letter describes the product, problem, and the action to be taken. The customer was instructed to use caution when using the Skip Forward button in stage 2 according to written instructions provided in the letter.This defect exists in Monaco Release 1.0.0 through 1.0.2 and will be resolved as soon as possible. Note: The issue was resolved in Monaco Release 2.0.1 and above.Questions should be directed to the manufacturer's representative telephone number at 408-830-8023. |  | 47 | 47 | Worldwide Distribution-- USA (nationwide) including the states of IL, MI, MO, OR, SC, TX, VA, and Puerto Rico and the countries of Canada, China, Cyprus, Estonia, France, Germany, Greece, Italy, Japan, Netherlands, Philippines, Poland, Portugal, Slovenia, South Korea, Spain, and United Kingdom. | Terminated | April 04, 2013 | 1739 | 47 | Software | Output/Calculation | Instructions/Update | Safety Notice/Insructions
Z-2558-2011 | 58272 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTI LION E-Plus Cassette, Model BP7507, Opti Medical (25 Cassettes with Samples Probes per box). |  |  | March 15, 2011 | 2011 | OPTI Medical Systems, Inc | The values reported by the chloride sensors may report erroneous values below specification on aqueous controls and potentially with patient samples of whole blood and serum/plasma. | Nonconforming Material/Component | OPTIMedical Systems sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 15, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were notified by letter via e-mail and fax. Customers were instructed to discontinue use of the product and to return any unused product to OPTI Medical Systems, Inc. Customers outside the U.S. were instructed to destroy unused product and certify the destruction on the enclosed Recall Response Card. The Response Card was to be completed and returned to OPTI Medical Customer Service via fax. OPTI Medical Systems will replace the cassettes they have with product from another lot.Customers were instructed to contact the Technical Support Department at 1-800-490-6784 for questions relating to this notice. |  | 86 | 86 boxes (25 units (per box) | Worldwide Distribution-- USA (nationwide) including states of CA and VA, and countries of India, Lebanon, Philippines, Sweden, Venezuela and Vietnam. | Terminated | August 08, 2011 | 146 | 86 | Not_Computer
Z-2556-2011 | 58762 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | 804159, Terumo Coronary Ostia cannula, 14 Fr, 10 (25 cm) long |  |  | May 16, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The presence of adhesive and/or skiving of plastic material observed inside the distal tip of the cannula, which could lead to particulate and/or decrease in the internal diameter. The plastic skiving may detach and enter the patient, embolizing the coronary vasculature. The potential risk due to decrease in the internal tip diameter is inadequate delivery of cardioplegia and/or prolonged procedu | Process control | Terumo Cardiovascular Systems Corporation sent an URGENT MEDICAL DEVICE REMOVAL notice dated May 16, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken. The letter described the problem, the potiential hazard, and the action to be taken by the consignee. Consignees were instructed to discontinue use of the product and return all product to Terumo. The letter provided a list of alternate products to be used instead of the recalled products. Consignees with questions should call 1-800-521-2818 Monday through Friday 8AM to 6PM. |  | 960 | 960 | Worldwide Distribution - USA (nationwide) and the countries of BELGIUM, BRAZIL, DOMINICAN REPUBLIC, UNITED ARAB EMIRATES (UAE), SINGAPORE, MALAYSIA, AND SOUTH AFRICA. | Terminated | March 21, 2012 | 310 | 6400 | Not_Computer
Z-2563-2011 | 58676 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface to Data Innovations Instrument Manager software; GE Healthcare IT Solutions, Barrington, IL 60010. |  |  | May 03, 2011 | 2011 | Ge Healthcare It | The GE Centricity Laboratory Instrument Interface to Data Innovations is processing a preliminary result instead of the final result from the Process System Manager when there is a test rerun conducted. | Software design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 3, 2011 to all affected customers. The letter describes the product, problem, and action to be taken by the customer. The letter provided the users with specific safety instructions to follow until the software was updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 14 | 14 units | Worldwide Distribution-- USA (nationwide) including states of Georgia, New York, Ohio, Oregon, and Tennessee and countries of Canada, England, and Qatar. | Terminated | October 04, 2011 | 154 | 14 | Software | Device Operation | Software update | Software Update
Z-2565-2011 | 58747 | No | JPF | Prothrombin-proconvertin test and thrombotest. | Hematology | Hematology | 510(K) Exempt | Factor II (Prothrombin) G20210A Kit Light Cycler Instrument; Manufactured in Germany for Roche Molecular Systems, Inc. Branchburg, NJ 08876 USA DisFor use with Light Cycler 2.0 instrument; |  |  | March 31, 2011 | 2011 | Roche Molecular Systems, Inc. | Roche Molecular Systems has decided to discontinue marketing the Factor II (Prothrombin) G20210A Kit and the Factor V Leiden Kit for use with the Light Cycler 2.0 Instrument in the US Market | Other | Roche sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 31, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. The customer was instructed to discontinue use of the affected product by August 31, 2011.Complete the attached faxback form ( 4293-00-0311 ) and fax it to 1-888-345-5359.For questions or technical support, please contact Roche Molecular at 1-800-526-1247, |  | 13864 | 13,864 units | Nationwide Distribution | Terminated | August 13, 2012 | 501 | 30890 | Not_Computer
Z-2564-2011 | 58258 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Monaco Release 2.0.1 and above |  |  | May 31, 2010 | 2010 | Computerized Medical Systems Inc | Monaco: due to a software defect, the Radiological Depth (cm) and SSD to DRP (cm) parameters are reported incorrectly on the Beam Summary Report. This occurs for all Monaco IMRT plans. | Software design | Computerized Medical System sent a "USER NOTICE" letter dated June 24, 2010, to all affected users. This notice identified the product, the problem, and the actions to be taken. There is no current workaround. All customers will be mailed a postcard to be returned as proof of receipt and understanding of the notification. Consignees will be contact via telephone or e-mail after June 30, 2010 if the postcard is not returned This notice is to be distributed to any and all users of CMS software at your organization who are potentially affected. The notice states the issue will be resolved in Monaco Release 2.04, but no date for this release was given.For further questions, please call (403) 830-8023 |  | 62 | 62 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Austria, Belgium, China, Finland, France, Germany, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, South Korea, Spain, Switzerland, Taiwan, Thailand and United Kingdom. | Terminated | April 10, 2014 | 1410 | 62 | Software | Output/Calculation | Software update | Software Update
Z-2569-2011 | 57843 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation. Monaco Release 2.04.00 and above. |  |  | October 27, 2010 | 2010 | Computerized Medical Systems Inc | Monaco: In XiO IMRT (Dynamic MLC Delivery) software, the radiation beam angle is set to zero. If different beam angles are selected via Monaco software, the beam angles will remain at zero. | Software design | Computerized Medical Systems, Inc. sent a User Notice to the affected consignees on October 27, 2010, explaining the reason for correction, the clinical impact, and the workaround method. (VERIFY USE NOTICE DATE) A return postcard was included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 11 | 11 | Worldwide Distribution - USA including CA, IL, MN, NM, OH, NY, TX, and WA and the Netherlands. | Terminated | April 02, 2014 | 1253 | 11 | Software | Output/Calculation | Software update | Software Update
Z-2567-2011 | 58227 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.1 through 4.40.00. |  |  | December 09, 2008 | 2008 | Computerized Medical Systems Inc | XiO Software: Under certain conditions, when the port shape is altered, dose is not recalculated. The port-change flag is not being set correctly when the port shape is regenerated. The dose distribution will not match the blocking displayed and the mis-match might not be obvious. Monitor Unit calculations will also be based on an incorrect dose distribution. This problem has existed since Re | Software design | Computerized Medical Systems, Inc. sent a "XiO CUSTOMER ADVISORY" letter dated December 8, 2008 to all affected customers. The advisory describes the product, problem, and the actions to be taken. The customers were instructed to set the prompt to "Yes" to avoid the problem from occurring according to written instructions provided in the advisory. This problem has existed since Release 4.3.1 and is expected to be resolved with XiO Release 4.50.00. Questions should be directed to the manufacturer's representative at 408-830-8023. |  | 784 | 784 | Worldwide Distribution-- USA (nationwide) including Puerto Rico and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Tajikistan, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1947 | 784 | Software | Output/Calculation | Software update | Software Update
Z-2572-2011 | 58233 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.1 through 4.34.02. |  |  | March 31, 2008 | 2008 | Computerized Medical Systems Inc | Focal: In absolute dose mode a beam with a port may be calculated with open normalization, with the result that the displayed dose reflects open normalization while the monitor units and QA plan reflect blocked normalization. This is particularly important when these beams also include compensators. The problem is detectable by looking at the delivered dose to the weight point in the patient as | Software design | Computerized Medical Systems Inc sent a "CUSTOMER ADVISORY" letter dated March 2008 to all affected customers. The letter describes the product, the problem and the actions to be taken. The firm provided a workaround for customers that states "Alternatively the plan can be imported directly into XIO". A copy of the Advisory was shipped with software within the affected range to those requesting a software upgrade. The issue was resolved in Focal Release 4.40.00 introduced in June 2008.The firm will send out a User Notice to all sites still running affected software to remind them that the issue resolution is available.For further questions, please call (403) 830-8023. |  | 307 | 307 | Worldwide Distribution -- USA (nationwide) and the countries of Albania, Argentina, Australia, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Thailand, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | April 09, 2014 | 2200 | 307 | Software | Display/Image | Software update | Software Update
Z-2574-2011 | 58214 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 1.5.0 and above. |  |  | May 14, 2008 | 2008 | Computerized Medical Systems Inc | In XiO, when Telepathy is exited, re-entered, and the temporary plan re-called, the OLD beam weight will be displayed but the dose distribution will reflect the modified beam weight. | Software design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter to all affected customers in May 2008. The letter identified the product, the problem and the action needed to be taken by the customer.The issue was resolved in XiO Release 4.50.00 which has been available since June 2009.A User Notice was sent in March 2011, to all the customers reminding them of the availability of the correction. This notice was sent hardcopy via US Mail but there was no return receipt.For further questions, please call ( 408) 380-8023 |  | 936 | 936 | Worldwide Distribution -- USA (nationwide) including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Israel, Italy, Japan, Libya, Lithuania, Malaysia, Mexico, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 2155 | 936 | Software | Display/Image | Software update | Software Update
Z-2568-2011 | 58215 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 through 4.50.00. |  |  | May 13, 2009 | 2009 | Computerized Medical Systems Inc | In XiO, when the beams are renumbered, it is possible to get into a state where the dose displayed for a beam does not match the beam placement on the patient. This can occur when dose is calculated, beams are renumbered, and the dose is re-summed only. | Software design | Computerized Medical Systems, Inc. sent an CUSTOMER ADVISORY letter dated May 2009, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.Computerized Medical Systems, Inc have created "patches" to resolve the problem. A User Notice was sent in March 2011, to all users still using the affected software notifying them of the updated software available to correct the problem. This letter was sent by US Mail but no return receipt was included.For further questions, please call (408) 380-8023. |  | 1116 | 1116 | Worldwide Distribution -- (USA) nationwide Distribution, including Puerto Rico, and countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia, Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1792 | 1116 | Software | Display/Image | Software update | Software Update
Z-2571-2011 | 58657 | No | JFP | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ca++ Membrane Units, Model #942-046, and Model #942-060 used on ABL505/555, EML105, ABL600, ABL700 Series, and ABL800 Analyzers. The membranes are packaged in a plastic box which contains four (4) membranes per box. |  |  | April 07, 2011 | 2011 | Radiometer America Inc | Based on internal testing and user/customer complaint reports received, the firm determined that wrinkles may arise on the outer membrane of the calcium membrane units upon installation. Retention of fluids between the wrinkles can cause a positive bias resulting in calcium measuring errors. This situation poses a risk to the safety of patients, particularly when the measurement results are used | Pending | The firm, Radiometer America, sent a "MEDICAL DEVICE FIELD CORRECTION" letter dated April 14, 2011, to all affected customers. The letter describes the product, problem and the actions to be taken. The letter provides the customers with a series of steps to be taken when replacing the Calcium Membrane (Until further notice the above procedure must be carried out when replacing the Calcium membrane). The letter also provides quality control instructions for verifying calcium performance. Additionally, the customer is required to manually adjust the upper and lower QC limits on the analyzer in relation to the limits printed on the insert sheet in order to detect the bias during quality control measurements (This must be done for every lot of QC ampoules used, and for any new lot of ampoules installed).The letter asks that each customer complete and return the attached form via fax to 440-871-0463, in order to verify that the notification letter was received. Questions should be directed to the firm's Technical Support staff at 800-736-0600, Opt. 4. |  | 30323 | 30323 boxes | Worldwide Distribution--USA (nationwide) including the states of AK, AL, AR, AZ, BM, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV, and the countries of: Algir, argentina, Austria, Austrailia, Bahrain, Belarus, Belgium, Brazil, Brunei Darussalam, China, Columbia, Cuba, Greece, Denmark, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greenland, Guatemala, Hungary, India, Iraq, Ireland, Isreal, Italy, Jamaica, Japan, Kasakhstan, Kenya, Lebanon, Lithuania Korea, Malaysia, Malta, mexico, Moldavia, Nepal, New Zealand, Norway, Netherlands, Peru, Philipines, Poland, Portugal, Qatar, Oman, Rumania, Russia, Saudi Arabia, Singapore, Serbia and Kosvo, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Syria, Thailand, The Netherlands, Tunesia, Turkey, UK, Ukraine, United Arabia Emirates, Urguary, Usbekistan, and Venezuela. | Terminated | February 19, 2012 | 318 | 30323 | Not_Computer
Z-2579-2011 | 58197 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.3.0 and above |  |  | October 03, 2010 | 2010 | Computerized Medical Systems Inc | XiO: When the beam template is recalled, XiO prematurely initiates a dose calculation, re-snaps the MLC or Port to the structure of the current patient , but then fails to re-calculate dose for the new MLC or Port position. If no subsequent change is made to force a dose recalculation, XiO will display the dose from the original MLC/Port shape dose instead of the dose from the re-snapped MLC/Port | Software design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 14, 2010 to its affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to distribute this notice to any all users of CMS software at their organization who are potentially affected by this issue, and complete and return the attached postcard to confirm receipt of the Safety Notice. The notice also states that this problem has existed since the XiO Release 4.3.0 and is resolved in the XiO release 4.26.00. Elekta will notify the customers when this solution is available.If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 1208 | 1208 | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 03, 2014 | 1278 | 1208 | Software | Device Operation | Software update | Software Update
Z-2578-2011 | 58200 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation treatment Planning System, XiO Release 4.2.0 and above |  |  | September 17, 2010 | 2010 | Computerized Medical Systems Inc | When a dialog box is displayed and the dose status is changed to "Dose not current", beam edits can be made in the graphics area using the mouse (rotate collimator, reposition collimator, move isocenter, move weight point, etc.). Then if CANCEL is selected on the dialog box, any edits made on the dialog box are undone, the dose is left unchanged, but the beam edits made in the graphics area with t | Software design | The firm, Elekta CMS Software, sent an "IMPORTANT SAFETY NOTICE" dated September 6, 2010 to its affected customers. The notice describes the product, problem and actions to be taken (workaround). The customers were instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue. The notice also states that this problem has existed since XiO Release 4.2.0 and will be resolved in XiO Release 4.70.00. Elekta will notify the sites when this solution is available. If you have any questions, please call 314-993-0003 or Toll Free: 800-878-4267. |  | 1293 | 1293 | Worldwide distribution: USA (nationwide) including Puerto Rico, and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, .Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 08, 2014 | 1299 | 1293 | Software | Device Operation | Software update | Software Update
Z-2577-2011 | 58210 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.0 through 4.40.00 |  |  | April 08, 2008 | 2008 | Computerized Medical Systems Inc | XiO: When the patient is positioned with their feet toward the CT gantry, the coordinates exported to the LAP IsoMark laser positioning system are not correct. | Software design | The firm, CMS, Inc., sent a "XiO Customer Advisory" letter dated April 8, 2008 to all affected customers. The letter describes the product, the problem, and the actions to be taken ( workaround). The letter also states the problem will be resolved in a future release. No release number or date was given.If you have any questions, please call 408-380-8023. |  | 75 | 75 | Worldwide distribution: USA (nationwide) and countries of: China, France, Germany, Japan, Mongolia, Netherlands, Nicaragua, Poland, Russia, South Korea, Switzerland, and Turkey. | Terminated | April 08, 2014 | 2191 | 75 | Software | Output/Calculation | Software update | Software Update
Z-2529-2011 | 58910 | Yes | NIP |  |  | Cardiovascular | PMA | Boston Scientific Innova Over-the-Wire Self-Expanding Stent System, Catalog No. 39181-07403. 7 mm x 40 mm x 130 cm (stent diameter x stent length x delivery system length). Made in USA, Two Scimed Place, Maple Grove, MN 55311 USA. The Innova Self-Expanding Stent System is comprised of two components: the implantable endoprosthesis and the stent delivery system. The stent is a laser cut self-expanding stent composed of a nickel titanium alloy (Nitinol). On both the proximal and distal ends of the stent, radiopaque markers made of tantalum increase visibility of the stent to aid in placement. The stent is constrained within a 6F (2.1 mm maximum OD) delivery system. The delivery system is a triaxial design with an outer shaft to stabilize the stent delivery system, a middle shaft to protect and constrain the stent, and an inner shaft to provide a guidewire lumen. The delivery system is compatible with 0.035 in (0.89 mm) guidewires. When ready to be implanted, the stent is deployed by retracting the middle shaft of the delivery system. A radiopaque marker at the distal end of the delivery system aids in visibility during deployment. As the stent is exposed to body temperature, it expands to appose the vessel wall. |  |  | May 19, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a recall of the Innova Self-Expanding Stent System because they have received 6 complaints involving no deployment/partial deployment of the Innova Stent. Potential health hazard events resulting from this type of failure include increased procedure time, vessel wall injury, and/or emergency surgery to remove the partially deployed stent. This recall d | Device Design | Boston Scientific sent an Urgent Medical Device Recall Removal - Immediate Action Required letter, dated May 13, 2011, to all affected customers via registered mail. The letter explained the issue, identified affected product, stated that further distribution and or use of the devices should cease immediately, and asked that devices in inventory to be segregated and immediately returned to Boston Scientific. An Account Reply Verification Tracking Form was also asked to be returned. Boston Scientific Sales Representative were assist in the return process by providing a Returned Goods Authorization (RGA) number.Customers were instructed to contact their local Sales Representative for any questions. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Product removal in US is being managed through SuperNOVA Clinical Trial suspension in accordance with IDE #G100291. OUS distribution to include: Austria, Finland, France, Germany, Italy, Netherlands, Spain, and Switzerland. | Terminated | February 04, 2014 | 992 |  | Not_Computer
Z-2582-2011 | 57807 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System. XiO Release 4.50.00 and above |  |  | November 16, 2010 | 2010 | Computerized Medical Systems Inc | XiO: When a new electron beam is added to a plan, the calculation Algorithm is set to Pencil Beam, and an Aperture is added to the beam, the user can select Port-Port Properties and edit the Material Thickness values. If the algorithm is then changed to Monte Carlo, the calculation will proceed and the user is not warned the Material/Thickness values just entered are not used in the dose calculati | Other | Elekta CMS Software drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was distributed to all affected sites in November 2010. The issue will be resolved in the XiO Release 4.70.00, expected to be available in July 2011. |  | 556 | 556 | Worldwide Distribution - USA (nationwide, Washington, DC and Puerto Rico and the countries of Albania, Algeria, Australia, Austria, Bahamas, Belarus, Belgium, Brazil, Canada, Chile, China,Croatia, Cyrus, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Phillipines,Poland, Portugal, Romoania, Russia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Suriname, Switzerland, Taiwan, Thailand, Turkey, Turmenistan, Ukraine, United Kingdom, and Vietnam | Terminated | April 03, 2014 | 1234 | 556 | Software | Treatment/Delivery/Therapy | Software Patch | Software Update
Z-2580-2011 | 58787 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | MTD 4000 - Mettler Traction Decompression System, Model: ME 4000 |  |  | April 14, 2011 | 2011 | Mettler Electronics Corp | The recall was initiated because Mettler has confirmed the possible failure of internal mechanical component and software detection for potential failure regarding Electronics Traction Device - MTD4000. The firm are initiating the recall because component failure might result in patient injury. Use of the device should cease immediately. | Device Design | Mettler Electronics sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 14, 2011 to all affected consignees. The letter identified the product, the problem and the actions to be taken. Customers were instructed to immediately cease use of the recalled device, examine their inventory and quarantine product subject to recall. In addition, consignees were told if they have further distributed the affected product to please identify their customers and notify them at once of this product recall. The letter asks that that each includes a "Recall Return Response Form" that consignees are required to complete and return. If you have any questions please contact Mettler Electronics Corp 714-533-2221 x331, Monday through Friday, between the hours of 8:30 A.M. and 4:00 P.M. Pacific Time. |  | 221 | 221 | Worldwide Distribution --USA (nationwide) and countries of Taiwan, Ireland and Haiti. | Terminated | July 10, 2012 | 453 | 221 | Hardware |  | Remove | Remove or Replace
Z-2590-2011 | 58717 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ENDO-FUSE(TM) FUSION ROD, 7 mm X 70 mm IMPLANT MATERIAL: STERILE R, Ti6A14V, CPTi, REF: 500005970, LOT 088605422, Wright Medical Technology, Inc , 5877 Airline Road, Arlington, TN 38002 |  |  | April 20, 2011 | 2011 | Wright Medical Technology Inc | Two lots- one lot of Endo-Fuse Fusion Rod, 7x50mm was commingled one lot of Endo-Fuse(R) Fusion Rod, 7x70mm. Not all parts in each lot are affected. The commingled rod sizes might not be detected prior to surgery. | Process control | The firm, Wright Medical, sent an "URGENT: Medical Device Voluntary Recall" letter dated May 6, 2011 to their customers via FedEx . The letter describes the product, problem, and actions to be taken. The customers were instructed to examine their inventory for the affected products; immediately return any affected product to Wright Technology as soon as possible (contact Customer Service at 800-238-7117 for return instructions and replacement inventory), and immediately complete and return the attached response form via fax to: 901-867-7401.If you have any questions regarding this matter, please contact 800-874-5630. |  | 121 | 121 | Nationwide distribution: USA including states of: AL, AZ, CA, CO, FL, ID, IN, IL, LA, MI, MN, MO, NC, NE, NH, NY, OH, PA, SC, TN, TX, VA, VT, WA, and WI. | Terminated | December 27, 2012 | 617 | 186 | Not_Computer
Z-2586-2011 | 58784 | No | MZP |  |  | Microbiology | PMA | COBASÂż AmpliPrep / COBASÂż TaqManÂż HCV Test, US-IVD |  |  | April 14, 2011 | 2011 | Roche Molecular Systems, Inc. | The COBAS Ampliprep/COBAS TaqMan HCV Test has been shown to under-quantitate a subset of genotype 4 patient specimens by approximately 1.0-1.5 log 10 in the absence of any sequence mismatches. | Other | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 14, 2011 and faxback form to their customers UPS ground. The letter described the product, problem and actions to be taken. The customers were instructed to review prior test results if there have been indications of discrepancies between test results generated with COBAS AmpliPrep/COBAS TaqMan HCV Test as compared to other relevant clinical and laboratory findings; complete and return the attached Product Advisory Notice: Affiliate Feedback Fax-Form via fax to: 1-866-792-5445, and file this letter for future reference, in particular for use with their 2010 CAP Survey Results to explain any under-quantitation of sample HVL 2-01.Please contact Roche Molecular Diagnostics Technical Support at 1-800-526-1247 if you have any questions about the information contained in this letter. |  | 50993 | 50,993 units | USA (nationwide) and countries of: Australia, Austria, Azerbaijan, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Egypt, Finland, France, Germany, Ghana, Greece, Hong Kong, India, Israel, Italy, Japan, Kuwait, Luxembourg, Malaysia, Mongolia, Netherlands, New Zealand, Norway, Poland, Romania, Saudia Arabia, Sinagpore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and United Kingdom. | Terminated | August 27, 2012 | 501 | 339392 | Not_Computer
Z-2584-2011 | 59037 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VS3 Vital Signs Monitor Models: |  |  | June 10, 2011 | 2011 | Philips Healthcare Inc. | Speakers on the SureSigns Patient Monitors and Vital Signs, failure may be identified by the display of a Speaker Malfunc message in the alarm message area or an AUDIO FAILED message in the main screen of the device or by the absence of audible sound | Device Design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 2011 to all affected customers. The letter identifies the products, problem, and actions to be taken.The letter states that customers may continue to use the affected products while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP. Additionally, the letter states that if the customer experience this INOP or there is no sound from the SureSigns Monitor, the customer is to remove the device from use and contact their local Philips service representative.The letter explains that Philips is completing the design for a replacement speaker assembly, which is expected to be available within 8 weeks. When available, Philips will replace the affected speaker assemblies free of charge.Customers are instructed to contact their local Philips Customer Care Center at 1-800-722-9377 if further information or support is required. |  | 5225 | 5225 units | Worldwide Distribution--USA (nationwide) and countries of ARAB JAMAHIRI, AUSTRIA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, CANADA ,CHILE, COLOMBIA, CROATIA, CYPRUS, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, KUWAIT, LEBANON, LIBYAN, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND, OMAN, PHILIPPINES, POLAND, PORTUGAL, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA ,SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | August 15, 2013 | 797 | 5740 | I/O |  | Replace | Remove or Replace
Z-2591-2011 | 58853 | No | OUB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MiraDry System consisting of the DTS3000 Console, MiraDry Handpiece |  |  | April 24, 2011 | 2011 | Miramar Labs, Inc. | Reports of skin burns on axilla after treatment with the miraDry System. | Device Design | The firm,Miramar, has a service provider in Japan that assists in dealing with customers. The service providers were provided a copy of the investigation report and communicated directly with the customers to set up appointments for the field correction. The customers were personally visited by a representative from Miramars service provider in Japan. All devices had field correction performed on April 26-28, 2011. |  | 8 | 8 units | International only: Japan | Terminated | June 17, 2011 | 54 | 8 | Not_Computer
Z-2588-2011 | 58808 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Kinamed SuperCable Grip and Plate System Driver, 3.5mm Hex, Catalog Number: 35-860-2060 |  |  | August 18, 2010 | 2010 | Kinamed Inc | The recall was initiated due to inadequate mechanical strength of these screwdrivers' hexagonal tipfeature, causing it to strip during final tightening. | Nonconforming Material/Component | The firm, Kinamed Inc., sent a recall letter dated August 18, 2010 to consignees/customers via FedEx. The letter described the product, problem and actions to be taken. The customers were instructed to Immediately examine their stock and return any Kinamed SuperCable Screwdriver, 3.5mm Hex bearing these lot numbers, along with the completed attached Effectiveness Check Form and shipment under RMA #74567 to the attention of: "Director of Quality Assurance"Kinamed, Inc.820 Flynn RoadCarnarillo, CA 93012-8701Note: The Kinamed sales agent shipped and had the product replaced for the customers.If you have any questions, please contact (303)-681-2208, (805) 384-2748 or(800) 827-5775 |  | 8 | 8 units | Worldwide distribution: USA (nationwide) state of: CO and country of: Columbia. | Terminated | October 25, 2011 | 433 | 8 | Not_Computer
Z-2585-2011 | 58837 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA NXT Ultrasonic Tissue Ablation System Service Module |  |  | May 02, 2011 | 2011 | Integra Limited | The insulated covering of a wire cable that powers the vacuum pump in the CUSA NXT Service Module can become abraded over time resulting in the exposure of bare wire. If this bare wire were to come in contact with the metal surface of the pump, it could result in a short circuit and a loss of aspiration. | Device Design | Integra sent a "FIELD ACTION NOTIFICATION" dated May 2, 2011 to all affected customers. The letter identified the product, the problem and the actions to be taken. Customers were informed of a Field Corrective Action that is being initiated. All devices (customer owned, Integra Loaner Pool, and Sales demos) will undergo rework of the vacuum pump assembly. The rework will be scheduled, at customer sites by Integra Field Service Technicians or returned to a designated Integra Service Center for rework. A loaner will be presented to customer accounts upon request. The letter asks that each customer complete and return the enclosed Recall Acknowledgement and Return Form to Integra, maintaining a copy of the letter and form for your records. For any additional information please call 609-936-2485. |  | 105 | 105 units | Worldwide -- USA (nationwide) and countries of France, Germany, Hungary, Ireland, Italy, Netherlands, Spain, Sweden and Taiwan | Terminated | March 26, 2014 | 1059 | 105 | Hardware | Treatment/Delivery/Therapy | Rework | Repair
Z-2602-2011 | 58763 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encompass Hip Stem, Plasma Spray, 12/14 Taper, Standard, REF 108-XXXX, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | May 10, 2011 | 2011 | Ortho Development Corporation | Tests on PVC packaging material resulted in alert levels for cytotoxicity. | Package design/selection | Ortho Development Corporation notified consignees by letter on May 11, 2011. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to return all remaining affected units to Ortho. Replacement product was to be shipped by May 31. |  | 113 | 113 units | Nationwide Distribution - USA including CA, FL, and TX. | Terminated | November 09, 2011 | 183 | 113 | Not_Computer
Z-2598-2011 | 57801 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning Workstation, Monaco Release 2.03.00 and above |  |  | January 13, 2011 | 2011 | Computerized Medical Systems Inc | Monaco: When the isocenter location is edited in the IMRT Activity screen, the shift coordinates are not updated in the IMRT Plan report. | Software design | Computerized Medical Systems Inc drafted a User Notice which explains the reason for correction, the clinical impact, and the workaround method. The User Notice was sent January 13, 2011. A return postcard is included for the customer to confirm receipt of the User Notice. The firm intends for this issue to be corrected in Monaco Release 3.00 in March 2011. |  | 137 | 137 | Worldwide Distribution - USA (Nationwide distribution to AK, AZ, CA, CO, FL, IL, IN, MA, MI, MN, MO, NC, NM, NY, OH, OR, PA, SC, TX, VA, WA, and WI and the countries of Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, India, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Phillipines, Poland, Portugal, Singapore, Slovenia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and the United Kindom. | Terminated | April 02, 2014 | 1175 | 137 | Software | Display/Image | Software update | Software Update
Z-2599-2011 | 58250 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.1.1 through 4.40.00. |  |  | February 06, 2009 | 2009 | Computerized Medical Systems Inc | Focal Software: if individual MLC leaf position edits are made, the original MLC aperture is retained. A subsequent aperture edit should not affect the MLC leaf positions however it was found that instead the leaves will then snap (unintentionally) back to the original aperture. Also, if the patient is sent to XiO after the MLC is edited, an isocenter shift made in XiO will also cause the MLCs to | Software design | Elekta CMS Software sent an "ADVISORY NOTICE" dated February 6, 2009 to its affected customers. The notice identified the product, the problem, and the actions to be taken. Customers were informed that the workaround method available is to (edit the MLC pattern using the aperture first and the individual MLC leaf edits second). The firm states that this issue will be resolved in Focal version 4.50.00 and above. A "User Notice" will be sent to all users still using the affected software notifying them of the updated software available to correct the problem. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 760 | 760 | Worldwide Distribution -- USA (nationwide) including the state of: Puerto Rico and countries of: Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Hungary, India, Israel, Italy, Japan, Malaysia, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sudan, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, United Kingdom, Uruguay and Venezuela. | Terminated | April 08, 2014 | 1887 | 760 | Software | Output/Calculation | Software update | Software Update
Z-2597-2011 | 58502 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Peloris II Rapid Tissue Processor; Leica Biosystems Melbourne Pty Ltd, Mount Waverly, 3149 VIC, Australia; catalog number 26.0005. |  |  | April 11, 2011 | 2011 | Leica Microsystems, Inc. | The processor potentially has a faulty impellor which could lead to reagent carry-over from one processing step to another, | Nonconforming Material/Component | Leica Microsystems sent a MEDICAL DEVICE CORRECTION letter dated April 11, 2011, to the affected customers via first class mail. The letter identified the product, the problem, and the action needed to be taken by the customer.The Customers were instructed to discontinue use of the device until their local Leica Microsystems representative contact them to arrange the replacement of the faulty impellor. The accounts were asked to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 22 | 22 units | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Canada, Germany Japan, Portugal and United Kingdom. | Terminated | June 20, 2011 | 70 | 22 | Not_Computer
Z-2595-2011 | 58260 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.0.2 and above |  |  | May 28, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: when Digitally Reconstructed Radiographs (DRRs) are generated in XiO, and DICOM is used to export the DRRs (File-Print-DRR), the exported images are associated with the Beams, but NOT the Plan. The Plan association had always been done manually by the user. This is functionality that has never been provided on XiO. The problem occurs because XiO is not storing a persistent DICOM | Software design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available.No date was given for the mailing or the solution. |  | 1292 | 1292 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) and the countries of Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 10, 2014 | 1413 | 1292 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-2592-2011 | 58269 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.50.00 and above |  |  | May 10, 2010 | 2010 | Computerized Medical Systems Inc | Focal: for XiO plans calculated in relative dose mode, the Focal Plan Review application will display the DVH and maximum dose correctly. However, the isodose display is not scaled correctly. If the user evaluated the plan based only on the isodose display, they might misinterpret the actual dose delivered to the patient. | Software design | Elekta / Computerized Medical Systems Inc sent an undated Safety Notice to consignees beginning on May 21, 2010. The notice identified the product, the problem, and a workaround. |  | 308 | 308 | Worldwide Distribution - USA (Nationwide including Puerto Rico) Foreign distribution was made to Albania, Australia, Austria, Belgium, Canada, Chile, China, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, and Turkmenistan. | Terminated | April 02, 2014 | 1423 | 308 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-2604-2011 | 58270 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 4.50.00 and above |  |  | May 19, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: when a treatment machine is defined in Source File Maintenance (SFM), the MLC Model is set to Elekta, and beams using this treatment machine are added to a new plan in Teletherapy, XiO automatically adds an MLC to the beam. This is because the Elekta linac will always have an MLC in place when the plan is actually delivered. | Software design | A Safety Notice, dated 5/20/10, was sent to Elekta / Computerized Medical Systems Inc clients in May 2010. The notice indicates the problem, the clinical impact. The notice states a future XiO Release will resolve the issue. A return postcard is included with the mailing to verify consignee received the notice. For sites not returning the postcard by 5/28/10, the recalling firm will contact them via phone or email to confirm receipt of the notice. The notice also informed the consignees they would be contacted when a solution was available. |  | 57 | 57 | Nationwide distribution to medical facilities including Puerto Rico. Distribution was also made to two Military and Government consignees. Foreign distribution was made to Australia, Belgium, Brazil, Canada, Colombia, France, India, Ireland, Israel, Italy, Malaysia, Netherlands, Philippines, Poland, Spain, and United Kingdom. | Terminated | April 02, 2014 | 1414 | 57 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2573-2011 | 58811 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The Ultraview SL 91517 Capnography module is a mainstream and sidestream gas analyzer designed to measure the concentration of carbon dioxide in a gas mixture, to aid in determining the patient's ventilatory, circulatory, and metabolic status. |  |  | May 12, 2011 | 2011 | Spacelabs Healthcare, Llc | Capnography Module, Model 91517, resets 20 seconds after power failure, instead of after three minutes as specified. This reset clears the RAM and the module returns to default settings. | Software design | Spacelabs Medical sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 12, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. Customers were recommended to set all their Capnography Module Configuration Manager alarm settings to "ON". The letter also recommended that customers brief their staff about the possibility of losing changes to the default settings and to restore them after a power interruption.A field service engineer will be contacting customers to set up a date and time when affected modules can be upgraded at no cost to the customer. For questions about this corrective action program, call Spacelabs Medical at 1-800-522-7025, Select 2 for Technical Support. For outside the US, customers can call 1-425-657-7200, x5089. |  | 5883 | 5883 units | Worldwide Distribution--USA (nationwide) and countries of ALGERIA, ARGENTINA, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA,COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FRANCE,GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, ISRAEL, ITALY,JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEPAL,NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PALESTINE, PANAMA,PERU, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND,SYRIAN ARAB REPUBLIC, TAIWAN, REPUBLIC OF CHINA, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | December 02, 2011 | 204 | 5883 | Other | Device Operation | Upgrade | Repair
Z-2594-2011 | 58253 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.40.00 and above. |  |  | January 19, 2009 | 2009 | Computerized Medical Systems Inc | Focal: the couch position can be moved between scans of the same CT study. When the images are imported into Focal, the software is not checking the DICOM Image Position (Patient) tag and therefore is not aligning the images sets if there are different couch positions indicated. A shift between image sets will exist of the same amount that the CT couch was moved. | Software design | Elekta CMS Software sent an "ADVISORY NOTICE" dated January 19, 2009 to its affected customers. The notice identified the product, the problem, and the acton to be taken. Customers were informed that the workaround method available is to (use the image fusion feature already in Focal to relign the scans). The firm states that this issue is resolved in Focal version 4.62.00, which was released in March 2011. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 31 | 31 | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, FL, IL, IN, KS, MO, NC, OH, OK, PA, TN, and WA. and countries of Australia, Austria, Canada, India, Netherlands, Portugal, Singapore, Spain, and United Kingdom. | Terminated | April 10, 2014 | 1907 | 31 | Software | Display/Image | Software update | Software Update
Z-2593-2011 | 58262 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO Release 2.1.0 and above |  |  | May 27, 2010 | 2010 | Computerized Medical Systems Inc | XiO Software: The Patient Scan position is being confused with the treatment orientation of the patient. The Patient Scan position is determined from the data entered in PFM. Subsequently, when a New Teletherapy Plan is created for a patient in Teletherapy, the user is prompted with Head/Foot Gantry on the text page. This prompt is asking for the orientation of the patient when they are treated. | Software design | Computerized Medical Systems Inc sent an undated Safety Notice to all affected customers. The notice indicated the problem, the clinical impact, and the workaround. The notice stated the sites will be notified when a solution is available. |  | 1400 | 1400 | Worldwide Distribution - USA (Nationwide distribution, including Washington, DC and Puerto Rico) Foreign distribution was made to Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahamas, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Libya, Lithuania, Malaysia, Malta, Mexico, Mongolia, Morocco, Netherlands, New Zealand, Nicaragua, North Korea, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Turkey, Turkmenistan, Ukraine, United Kingdom, Uruguay, Venezuela, Yemen, and Zimbabwe. | Terminated | April 09, 2014 | 1413 | 1400 | Other |  | Instructions | Safety Notice/Insructions
Z-2596-2011 | 58256 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Radiation Treatment Planning System, Focal Release 4.3.0 through 4.34.01 |  |  | November 30, 2007 | 2007 | Computerized Medical Systems Inc | Focal: when a plan is sent to Focal from XiO and the structure properties are edited in Focal CT Sim activity, these changes result in changes to the MLC or field size for the plan when they should not. XiO will recalculate dose when modified plan is returned to XiO however the user may not notice that it has changed form the plant they had approved earlier. Editing the structure properties (col | Software design | Computerized Medical Systems Inc sent a Customer Advisory, dated 11/30/07, was sent to consignees December 2007 to notify them of the issue and how to avoid the workaround. The advisory listed the problem, clinical impact, and a workaround. The firm will be sending each site a reminder that the issues is resolved. |  | 233 | 233 | Worldwide Distribution - USA (Nationwide AL, AR, CA, DE, FL, GA, IL, IN, MA, MD, MI, MOP, NE, NY, OH, PA, RI, TX, VA, WA, WI, and WV and the countries of Albania, Argentina, Belgium, Bosnia and Herzegovina, Brazil, China, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, India, Italy, Japan, Netherlands, Peru, Portugal, Qatar, Romania, Russia, Serbia, South Africa, South Korea, Spain, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | April 10, 2014 | 2323 | 233 | Software | Output/Calculation | Software update | Software Update
Z-2608-2011 | 58906 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | BBL" OXACILLIN SCREEN AGAR, catalog #221952, box of 10 plates. |  |  | April 20, 2011 | 2011 | Becton Dickinson &amp; Co. | Oxacillin Screen Agar may not have been produced according to specification and could result in false antibiotic susceptibility test results in patient specimens. | Nonconforming Material/Component | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated April 2011 to all affected customers. The letter identifies the product, problem and actions to be taken by the customers.The letter instructed customers to discontinue use of the affected lot number and discard any remaining packages. Customers were requested to complete and return an attached response form via fax at 410-316-4258. The letter states that all discarded product will be replaced.Questions concerning the recall process were directed to BD Customer Service Department at 1-800-675-0908. |  | 5440 | 5440 plates | Worldwide Distribution--USA (nationwide) and the countries of Canada, Columbia, Jamaica, South Korea, and Taiwan. | Terminated | August 15, 2011 | 117 | 5440 | Not_Computer
Z-2601-2011 | 59000 | No | LDR | Infusion pump. | General Hospital | General Hospital | 510(k) | "***DOSI-FLOW 1 I.V. EXTENSION SET WITH I.V. FLOW REGULATOR REGULADOR DE CAUDAL PARA SETS DE INFUSION I.V.***REF 259-0250***50 UNITS***STERILE OE***LATEX-FREE***LEVENTON LEVENTON, S.A.U. C/Newton, 18-24 08635 Sant Esteve Sesrovires BARCELONA-SPAIN***Distributed by: Victus www.victus.com 4918 S.W. 74th Court Miami, FL 33155***" |  |  | March 11, 2011 | 2011 | Leventon S. A. U. | Leventon, S.A.U. Barcelona, Spain is recalling DOSI-FLOW REGULATOR, Part Number 259-0250, Lot Numbers 101444L and 101494L. Product has potential to leak. | Device Design | The firm, Leventon, S.A.U. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to the affected customer. The letter identifies the product, problem, and actions to be taken by the customer.The letter instructs the customer to examine their inventory and quarantine any product subject to the recall. They were also instructed to identify their customers who have received the product, notify them of the recall, send a copy of the recall notification letter and request that the customers return any unused product. The customers were also instructed to complete and return the attached RECALL RESPONSE FORM within 5 working days via fax to 00 34 93 6532556 or e-mail to dsalvatierra@leventon.es.Questions concerning this recall should be directed to the QA Manager at 00 34 93 8176316 or e- mail to dsalvatierra@leventon.es. |  | 1000 | 1000 cases (50,000 units) | Nationwide Distribution: FL only. | Terminated | June 28, 2011 | 109 | 1000 | Not_Computer
Z-2606-2011 | 58833 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems |  |  | May 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that Access Ostase Calibrator and QC vials from the affected lots have shown an increased rate of fracturing and may leak the contents in the box during shipment to customers.Operator injury may occur due to contact with broken glass, exposure to infectious agents, and/or exposure by skin contact to agents that may cause sensitizatio | Package design/selection | Beckman Coulter sent an PRODUCT CORRECTIVE ACTION (PCA) letter on May 10, 2011 to all affected customers who purchased Access Ostase Calibrators and Access Ostase QC. The letter identified the customers with an explanation of the problem and the potential impact should the problem occur and the action needed to be taken. Customer were instructed to:(1) Inspect all Access Ostase Calibrator and QC vials and packaging for signs of failure prior to use, including:-Cracking, particularly along the bottom of the vial,-Absence of reagent and/or leaking of contents into the box(2) Human source material used in the preparation of Access Ostase Calibrator and QC material has been tested and found non-reactive for Hepatitis B (HBV), Hepatitis C (HCV), and Human Immunodeficiency Virus (HIV-1 and HIV-2). No known test method can offer complete assurance that infectious agents are absent. Access Ostase Calibrator and QC vials should be handled in accordance with Good Laboratory Practices, and should be handled as if capable of transmitting infectious disease.(3) Skin contact with Access Ostase Calibrator and QC material may cause sensitization. Wear suitable gloves when handling vials. If contact with skin occurs, wash immediately with plenty of soap and water.Customers were told to share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected products listed above to another laboratory, please provide a copy of the letter to them.In addition, customers were requested to complete and return the enclosed Response Form within 10 days so we are assured you have received this important communication.For replacement of defective vials:(1) In the United States, please contact Client Services at 1-800-526-3821, option 1.(2) In Canada, please contact Customer Service at 1-800-463-7828.(3) Outside of the United States and Canada, contac |  | 702 | 702 kits | Nationwide in US and worldwide to: Algeria,Argentina,Australia,Austria,Belgium,Canada,China,Colombia,Czech Republic,France,Georgia,Germany,Greece,Hong Kong,Hungary,India,Italy,Japan,Jordan,Lebanon,Liechtenstein,Mexico,Morocco,New Zealand,Panama,Reunion,Romania,Russian Federation,Saudi Arabia,Serbia,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | July 11, 2012 | 428 | 891 | Not_Computer
Z-2570-2011 | 58896 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHECK PERFORMA STRIP NO. 05987288031, 100 BANDELETTES REACTIVES, ROCHE DIAGNOSTICS GMBH, WWW.ACCUCHEK.COM glucose monitoring test strips. |  |  | May 18, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Erroneous results: Internal investigations revealed that the test strips exhibited a performance issue that can cause test strip errors or a bias in test result that exceed the tolerance range according to EN ISO 15197. | Component design/selection | Direct accounts were notified by a Roche Diagnostics press release on 5/18/2011 that included the product's name, lot number, and information to withdraw the remaining inventory of lot 470049 from storage |  | 27019 | 27,019 units | No US distribution, in France Only. | Terminated | October 25, 2011 | 160 | 27019 | Not_Computer
Z-2607-2011 | 56787 | No | LGW |  |  | Neurology | PMA | Eon Mini Neurostimulation (IPG) System Model 3788 |  |  | May 24, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | Defective battery was identified as cause for reports of inability to communicate or recharge the product. | Process control | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers.St. Jude Medical recommends no action be taken if the product serial numbers do not match the numbers listed in the letter. For product remaining in unimplanted inventory, customers are instructed to contact their SJM representative so that product may be returned and replacements be provided. For implanted product that matches the serial numbers, the Medical Advisory Board recommend not to unnecessarily explant the devices if they are functioning as intended. Replacement devices will be provided at no additional cost to the customer. A Product Reconciliation Form was attached to the letter for customers to complete and return any unused product to the St. Jude Medical Neuromodulation Division at 6901 Preston Road., Plano, TX 75024. Questions regarding this recall are directed to the customer's St. Jude Medical Neuromodulation Division Representative. |  | 17141 | 17,141 | Worldwide Distribution-- USA (nationwide) including PUERTO RICO and countries of AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, FINLAND,FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA,SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Terminated | July 21, 2015 | 1519 | 17141 | Battery | Device Operation | Replace | Remove or Replace
Z-2575-2011 | 58995 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Integra Pain Management Disposable, Convenience Kit |  |  | May 19, 2011 | 2011 | Integra LifeSciences Corp. | Kits contain protective wipes that may be contaminated with Bacillus cereus. | Nonconforming Material/Component | Integra LifeSciences consignees were notified by Urgent Medical Device Recall letter on May 20, 2011 and told to quarantine any remaining product and to notify customers to return any remaining product to them. Further information is available at (800) 241-2210. |  | 520 | 520 kits | UT | Terminated | August 01, 2011 | 74 | 560 | Not_Computer
Z-2610-2011 | 58779 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI TM 101 Monitor, CDI TM 101 Hematocrit/Oxygen Saturation Monitoring System, |  |  | April 28, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that the presence of electromagnetic interference (EMI) in the operating room may affect the accuracy of the valuesdisplayed by the CDI"101 Hematocrit/Oxygen Saturation Monitoring System: hematocrit (HCT), hemoglobin (Hgb) and Oxygen Saturation. | Device Design | Terumo Cardiovascular Systems, Corp issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 28, 2011, to all affected customers. The letter identifed the product, the problem and the action needed to be taken by the customer. After a solution has been verified, Terumo CVS will implement a correction for all affected units. Customers were instructed to review this Medical Device Correction notice, Assure that all users are aware of this notice. Confirm receipt of this and to fax back the Customer Response form.For further question please contact Terumo CVS Customer Service at 1-800-521-2818, Monday - Friday, 8 AM - 6 PM EST. |  | 622 | 622 | Worldwide Distribution -- USA ( nationwide) and the countries of Colombia, Vietnam, India, Indonesia, Pakistan, Canada, Chile, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Singapore, Taiwan, Belgium, Thailand, Vietnam and India | Terminated | November 06, 2012 | 558 | 622 | Other | Output/Calculation |  | Other
Z-2625-2011 | 58822 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 12-lead ECG Cable Assembly contains the following cable and attachments: |  |  | May 11, 2011 | 2011 | Physio Control, Inc. | The 12-lead ECG cables have experienced premature failures which cause the shielding and conductors to break internally creating ECG noise or "lead off" indications. | Device Design | Physio-Control, Inc. sent an "URGENT-MEDICAL DEVICE ACCESSORY INFORMATION" letter dated May 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The customers were advised to follow the instructions:-Use standard hard paddles or therapy electrodes for ECG monitoring if monitoring cannot be performed using the ECG cable. Refer to Troubleshooting Tips for ECG Monitoring in the LIFEPAK 12or LIFEPAK 15 Operating Instructions-Monitoring chapter.-Disconnect the ECG cable from the monitor after use and coil it loosely for storage. Please refer to the Storage Instructions for 12-lead ECG Cable.-Inspect cables daily for bulges which may indicate damaged internal wires. If bulges are visible, replace the ECG cable.-Perform periodic 6 month function testing: Refer to Maintenance and Testing Schedule in the LIFEPAK 12 or LIFEPAK 15 Operating Instructions- Maintaining the Equipment chapters for instructions on how to perform a patient ECG cable function check to confirm proper operation ofthe cable.Customers can visit Physio-Control's website at www.physio-control-notices.com/12LeadECG for additional information or todownload copies of current LIFEPAK 12 defibrillator/monitor or LIFEPAK 15 monitor/defibrillator Operating Instructions and Operator's checklists.Customers can also call Physio-Control's Technical Support Team at 1-800-442-1142, option 5, from 6:00 A.M. to 4:00 P.M. (Pacific), Monday-Friday. |  | 21825 | 21,825 units distributed worldwide | Worldwide Distribution--USA (nationwide) including Puerto Rico and countries of Australia, Brazil, Chile, Germany, Hong Kong, Mexico, The Netherlands, and New Caledonia. | Terminated | August 12, 2011 | 93 | 21825 | I/O
Z-2612-2011 | 59072 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | 45cm Laparoscopic Instruments Sterilization tray. |  |  | June 08, 2011 | 2011 | Stryker Endoscopy | During testing and review, Stryker Endoscopy discovered that the sterilization parameters for the 132 degree C gravity steam sterilization cycle in the IFU may not consistently sterilize all instruments in the 45cm Laparoscopic Instruments Tray. | Nonconforming Material/Component | Stryker Endoscopy sent a "URGENT: DEVICE CORRECTION" letter dated June 8, 2011 to affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to discard all old Instructions for Use sheets for their 45cm Laparoscopic Tray and download the new IFU from the URL link:http://stryker.com/45cm tray. Additionally, customers were instructed to complete and return an attached "acknowledgement of receipt" form and fax to 408-754-8378 or scan and e-mail to trays@stryker.com.Questions can be answered by calling 1-800-624-4422 or by e-mail at trays@stryker.com. |  | 603 | 603 | Worldwide Distribution--USA (nationwide) and countries of Brazil, Canada, Chili, Germany , Korea, Mexico, Philippines, Poland, Portugal, Singapore, Sweden, Switzerland, and Taiwan. | Terminated | October 25, 2011 | 139 | 603 | Not_Computer
Z-2609-2011 | 58767 | Yes | MRM |  |  | Cardiovascular | PMA | Paradym DR Model 8550 and Paradym VR Model 8250, Sterile EO, Manufactured in Italy by Sorin CRM S. r.l., Via Crescentino s.n., 13040 Saluggia(VC) Italy |  |  | May 02, 2011 | 2011 | ela Medical Llc | On February 28, 2011, the company discovered that since the beginning of the US commercial distribution of Paradym DR 8550 and Paradym VR 8250 ICDs in October 2010, the devices included an incorrect version of the Patient Booklet.The company determined that there is no risk of adverse event associated with this labeling error. | Error in labeling | Sorin CRM USA instructed its field representatives on April 27, 2011 via a conference call to attach the correct Patient Booklet (ref. N582,) as an addendum to the device's external packaging, along with a cover letter "Quality Note" dated May 2, 2011 explaining the discrepancy. The field representatives were fedexed the information on May 2, 2011.The letter instructed the Field Representative to attach the correct patient booklet and an addendum, to the outside of the device box.Field Representatives were asked to complete the acknowledgement form by identifying S/N's in which they attached the pouch and their signature of completion and fax the form to 866-500-6096 or e-mail a scanned copy to sorin.ptusa.sorin.com.For questions regarding this recall call 877-663-7674." For affected units that were already implanted as of April 27, 2011, the patient booklet given to the patient may include information written by the physician or hospital staff that is related to the implanted system (ICD and leads), physician references and patient medications. For this reason and given the fact that the company determined that there is no risk for the patient associated with this labeling error, the company decided to not replace the patient booklet for this population. |  | 275 | 275 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, FL, GA, IA, MA, MI, MO, MS, NJ, NY, NC, OH, OK, PA, SC,A TN,T X, and WA. | Terminated | April 26, 2012 | 360 | 275 | Not Computer
Z-2624-2011 | 58877 | No | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Product Brand Name: Coated Bio-Eye Orbital Implant w/Conformer. |  |  | May 20, 2011 | 2011 | Integrated Orbital Implants Inc | The recall was initiated because Integrated Orbital Implants has confirmed certain lots of the Coated Bio-Eye were manufactured without expiration dating visible on the packaging. All unlabeled product is currently beyond its expiration date. | Other | Integrated Orbital Implants Inc initiated a recall communication on 5/20/2011 with Integrated Orbital Implants (IOI) forwarding a Urgent Safety Notification letter to all their customers who had purchased the Coated Bio-Eye implants. The letter asked the customers to look for evidence of a printed expiration date on Coated Bio-eye package. If any Coated BioEye packages were not labeled with an expiration date the letter instructed customers to immediately sequester the products and not to use the products in surgery and contact IOI by telephone at (858) 6779990 or 1 (800) 4246537 to request an RMA number and shipping instructions.Also stated in the letter, IOI will replace Coated BioEye implants without expiration dates free of charge upon receipt of the implants labeled withoutexpiration dates. |  | 4237 | 4237 | Worldwide Distribution - USA (nationwide) and the countries of nationwide and worldwide to: Japan, Greece, Germany, Romania, UAE, Australia, Canada, Philippines, Korea, Taiwan, Chile, Denmark, UK, Italy, Pakistan, Turkey, Philippines, India, South Africa, and Singapore. | Terminated | January 11, 2012 | 236 | 4237 | Not_Computer
Z-2623-2011 | 58834 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Large Bore 4-Way Stopcock, Volume 0.65 mL, Injection Site, Rotating Male Luer Lock; a sterile, nonpyrogenic fluid pathway; product code 2N5601. Intended use: For the administration of drugs and solutions to a patient's vascular system through a vascular access device. |  |  | May 17, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for pinholes and tears in the individual packaging of the stopcocks, compromising the sterility of the product. | Package design/selection | Baxter Urgent Product Recall letters dated May 17, 2011 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that potential pinholes and tears in the individual packaging could compromise the sterility of the stopcock. The accounts were requested to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/17/11, and instructed to notify their customers and ministries of health of the recall. |  | 4800 | 4,800 units | California and internationally to Japan and Australia | Terminated | June 18, 2012 | 398 | 4800 | Not_Computer
Z-2611-2011 | 59004 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Pathfinder Holter Analyzer is a Programmable Diagnostic Computer. |  |  | May 24, 2011 | 2011 | Spacelabs Healthcare, Llc | A potential hazard associated with the use of Spacelabs Healthcare Pathfinder Holter Analyzer software. The second 24 hours of a 48 hours EVO recording will not be analyzed. | Software design | Spacelabs Healthcare sent a "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 3, 2011 to all customers. The letter describes the product, problem, and the actions to be taken by the customers. Spacelabs Healthcare recommended that until software upgrades can be done, the customers should weigh the benefits versus the risks when deciding whether or not to continue using the product. Spacelabs Healthcare also recommended that customers advise their staff that the second 24 hours of a 48 hour EVO recording would not be analyzed.Field service engineers will be contacting customers to set up a time and date to install the revised version of the Pathfinder software. Spacelabs will upgrade Pathfinder software regardless of whether the consignees are using EVO recorders to ensure this hazard does not happen in the future.Consignees can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the US, 44 1992 507740 for any questions about this recall. |  | 58 | 58 units total (5 units in US and 53 units internationally) | Worldwide Distribution--USA (nationwide) including states of CA and NC and countries of Czech Republic, Denmark, Hong Kong, Israel, The Netherlands, and United Kingdom. | Terminated | October 26, 2011 | 155 | 58 | Software | Output/Calculation | Software update | Software Update
Z-2616-2011 | 58838 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Procedure Kit Part number 70367-07 Mission St. Joe's - Asheville X-Coated Low Prime Pack, Sterile, Terumo Cardiovascular Systems, Ashland, MA. The pressure monitoring kit is indicated for sterile single use with the reusable pressure transducer. |  |  | May 03, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has recieved complaints regarding 35 pressure monitoring kits in which the membrane in the kit separated form the pressure dome housing.The firm determined that the defect was caused by a change in the pressure dome's manufacturing process. | Other | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE REMOVAL notice dated May 3, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were sent 10 replacement pressure monitoring kits with the removal notice as immediate replacements. The firm will send additional replacements to customers that have more than 10 units of affected product as noted on the Customers' returned reponse form. Customers that received the custom packs were provided stickers to identify the recalled packs. The customers were instructed not to open the packs to remove the affected pressure monitoring kit prior to clinical use, but to replace the pressure domes at the time of use.For more information please contact Terumo Cardiovascular Customer Service at 800- 521-2818. |  | 1515 | 1515 all custom packs | Worldwide Distribution -- USA (nationwide )including the states AR, CA, FL, GA, MO, MT, NC, ND, OK, OR, PA, TN, TX, UT, VA, WA and WY, and the countries of Argentina, Belgium, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Guatamala, Honduras, Hong Kong, Indonesia, Kuwait, Malaysia, Mexico, Pakistan, Panama, Philippines, Singapore, South Korea, Thailand, United Arab Emirates, Uruguay and Vietnam | Terminated | July 18, 2012 | 442 | 21010 | Not_Computer
Z-2581-2011 | 58854 | No | DXQ | Blood pressure cuff. | Cardiovascular | Cardiovascular | 510(k) | Cuffable Blood Pressure Cuffs; Manufactured for Vital Signs, Inc., a GE Healthcare Company 20 Campus Road, Totowa, NJ 07512 Made in Mexico. |  |  | April 26, 2011 | 2011 | GE Healthcare, LLC | BP cuffs may not properly inflate due to a leak | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 26, 2011 to all affected customers. The letter identified the product, the problem and the action to be taken. Customers were instructed to discontinue use of the product, isolate all affected products, return only unused products and distribute this letter to any other healthcare institutions that are potentially affected by this issue. The letter asks that each customer complete and return the enclosed confirmation form via fax to 800-535-7923, in order to confirm receipt of the notification letter. For any questions, contact your local Vital Signs Account Manager or the International Customer Service at 1-800-932-0760. |  | 2684996 | 2,684,996 units | Worldwide Distribution -- USA (nationwide) including Puerto Rico and countries of: Argentina, Australia, Belgium, Canada, Chile, Ecuador, Finland, Germany, Hong Kong, Israel, Italy, Latvia, Lebanon, Nicaragua, Panama, Peru, Saudi Arabia, Slovenia, UK, Uruguay and Venezuela | Terminated | May 28, 2013 | 763 | 2684996 | Not_Computer
Z-2633-2011 | 55276 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tubing Assembly #142, V2, V12 Mixer/Wash a component of |  |  | July 23, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because the Wash Concentrate Tube Assembly 142 may fail to deliver the appropriate amount of wash concentrate on some SYNCHRON DxC Systems due to an accumulation of crystals.The impact of this issue would include:(1) Quality Control and patient results can potentially be affected.(2) If system has been calibrating properly and Quality Control is recovering within es | Pending | Beckman Coulter sent out an "Urgent Product Corrective Action" letter dated July 23, 2009 and mailed the week of July 27, 2009 to all affected customers. The letter provides the cusotmers with explanation of the problem identifed and asks them to take the following actions:" If consignees are not able to resolve these issues through routine troubleshooting they were told to contact BCI Customer Technical Support at (800) 854-3633 in the United States and Canada." This issue will be corrected with an upcoming software version and tubing modification." Complete and return the enclosed response form within 10 days so that Beckman Coulter may maintain their records.Consignees were also asked to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation.Questions regarding this Product Corrective Action letter were directed to their Customer Technical Support Center (Hotline) at (800) 854-3633 in the United States and Canada. |  | 2392 | 2392 units | Worldwide Distribution - USA (Nationwide) and Canada | Terminated | July 02, 2012 | 1075 | 2392 | Computer
Z-2622-2011 | 58228 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Interplant Radiation Treatment Planning System, Interplant Release 3.4.0 |  |  | December 01, 2008 | 2008 | Computerized Medical Systems Inc | Interplant Software: If the trajectory needle with a non-zero retraction is adjusted in any of the planning views (Transverse, 3D) and the program is closed and reopened, the Needle Definitions page in the Dose Plan Summary (Plan View) displays the needle with a 0.00 retraction even when the needle is displayed correctly in the planning view. If the Eliminate Leading Spacers function is applied t | Software design | Computerized Medical Systems, Inc. sent a "CUSTOMER ADVISORY" letter dated December 1, 2008, to all affected customers. The letter identified the product, the problem, and the actions to be taken. Customers were instructed to discontinue use of the product. There is no workaround at this time. This issue was resolved in Interplant Release 3.4.1 software. All possibly affected users have been shipped the Interplant Release 3.4.1 software in May 2010. If you have any questions, please call 314-993-0003 or toll free 800-878-4267. |  | 41 | 41 | Worldwide Distribution -- USA (nationwide) including states of: CA, KS, KY, MA, MD, ME, MO, NC, NE, NY, OH, PA, TX and VA and countries of: Australia, Brazil, Chile, China, Germany and Japan. | Terminated | April 04, 2013 | 1585 | 41 | Software | Display/Image | Software update | Software Update
Z-2627-2011 | 58873 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BACTEC" FX Instrument-Top, catalog #441385. |  |  | August 10, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic instrumentation may exhibit random operational errors that could delay effective patient treatment. | Software design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated August 2010 to all affected customers. The letter describes the product, problem and action to be taken.The letter described the type of error messages the customers may experience. All affected units will be scheduled for on site software update as well as a disk drive replacement. A BD field service engineer will contact the customers to schedule the on site visit.Customers experiencing issues prior to the scheduled visit or have questions should contact the BD Diagnostics Service Department at 800-638-8663. |  | 907 | 907units | Worldwide Distribution--USA (nationwide) including countries of Australia, Belgium, Canada, China, Colombia, New Zealand, Hong Kong, Japan, Korea, Philippines, Singapore, Taiwan, and Thailand. | Terminated | July 05, 2011 | 329 | 907 | Software
Z-2629-2011 | 58793 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo-s Ventilator Systems; Article number 66 40 440. |  |  | April 20, 2011 | 2011 | Maquet Inc. | Through customer complaints, Maquet has been informed that SERVO-i ventilators have stopped to ventilate and have generated three different technical error codes. | Other | Maquet Inc. sent an URGENT DEVICE SAFETY ALERT letter dated April 20, 2011, to all affected customers via Federal Express. The letter identified the product, the problem and the action needed to be taken by the customer.The letter informs the user of the issue and provides additional instruction and how the existing manual should be updated. Customers are asked to fill out the attached response form and fax it to 1-973-807-9210. The following warning will be included in all user's manuals delivered with new ventilators:'Warning ! The Servo-i Ventilator System is not intended to be used during radiotherapy, since this may cause system malfunction.Warning! The Servo-s Ventilator System is not intended to be used during radiotherapy, since this may cause system malfunction.'Additionally the warning stated above will be included in the format of a sticker with the yearly preventive maintenance kit for Servo-i and SErvo-s to update exiting User's manuals for the installed base of Servo-i and SErvo-s ventilators. A maintenance report form will also be included in the kit which serves as an acknowledgement from the user that the sticker has been added to an affected ventilator unit's operating manual.For further questions, please call (973) 709-7660. |  | 17881 | 17039 units worldwide; 842 units US | Worldwide Distribution - USA (nationwide) | Terminated | September 17, 2013 | 881 | 75503 | Not_Computer
Z-2631-2011 | 58859 | [6 empty] | 1500T14 Cardiac Ablation System - Generator |  |  | May 09, 2011 | 2011 | St Jude Medical | The recall was initiated because St. Jude Medical has confirmed its T14 and T12 RF Generatormay malfunction upon encountering an unexpected electric current (e.g. static electricity).The resulting 'malfunction results in a failure of the generator to respond to key pressesproperly, and may cause unexpected energy delivery or inability to discontinue energydelivery with the start / stop but | Other | The firm, St. Jude Medical (SJM), sent a "Product Correction" letter dated May 9, 2011 to it's affected customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the T14/T12 generators and remote controls until the units have been properly serviced; complete and return the attached field action form, and utilize the units in accordance with the Operators Manual. In case of system malfunction, the customers were instructed to shut off the RF power delivery by turning off the generator via the rocker switch, dislodging the power cord, or disconnecting of the patient cable or indifferent electrode connection. An SJM sales representative will contact the customers in the near future to facilitate repair or exchange of the impacted unit.If you have questions regarding this notice, contact the Marketing Manager, AF at (905-286-4015), Marketing Director, AF/CRM, at (905-286-4107); Product Manager (Tokyo) 03-6255-5740, and/or your SJM sales representative. |  | 10 | 10 | International distribution only: Canada and Japan. | Terminated | April 03, 2012 | 330 | 168 | Hardware | Device Operation | Repair/Replace | Repair
Z-2634-2011 | 58887 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | ***Processed at: Regeneration TECHNOLOGIES, INC. Alachua, FL, USA 32615***Distributed by: SpinalGraft Technologies, LLC 4340 Swinnea Road STE 39 Memphis, TN 38118***OSTEOFIL DBM Paste, 5cc Tissue ID: 321933 Reorder: 002405 Lot #(Donor ID): 101050353 Expiration: 02/21/2013***STERILE***" |  |  | April 09, 2011 | 2011 | RTI Biologics, Inc. | RTI Biologics issued a recall of Osteofil DBM Paste, 5cc that was distributed on 03/04/2008 to one customer. The product was labeled as 5.0cc but should have been labeled as 0.5cc. All customers were notified of this recall on 04/09/2011. | Labeling mix-ups | RTI Biologics sent a "VOLUNTARY RECALL" letter dated April 9, 2008 to the one distributor, Spinal Graft Technologies. LLC. The letter describes the product, problem, and action to be taken by the distributor.The letter states that some of the product was issued a return authorization and requests that the distributor recover the remaining product from distribution. The recovered product will be issued a return authorization. RTI states the grafts can be returned in a non frozen state.Recall effectiveness will be determined by return of the product. |  | 26 | 26 units | Domestic Distribution--TN only. | Terminated | July 24, 2012 | 472 | 26 | Not_Computer
Z-2645-2011 | 58206 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ERGO++ Stereotactic Radiation Treatment Planning System, ERGO++ Release 1.6.3 and 1.6.3.1 |  |  | December 11, 2009 | 2009 | Computerized Medical Systems Inc | ERGO Release 1.6.3 is overestimating the MU values. | Software design | Computerized Medical Systems, Inc. sent an "USER NOTIFICATION" dated December 11, 2009, to all affected customers. The notice identified the product, problem and suggested a workaround to minimize the risk. The notice stated that this workaround does NOT substitute the need for regular Quality Assurance. Customers are instructed to review page 39 of user manual for further explanation of warnings. Customers were asked to complete the Notice for Customer section and return to 3D Line Medical System. This issue has been resolved in Release 1.7.0 and above. No follow-up notice will be distributed to customers. For further information please contact: CMS Software (314) 812-4460 (St Louis Office) |  | 45 | 45 | Nationwide distribution to medical facilities in CA, MO, NC, ND, NV, NY, OH, OK, TX, and WA. Foreign distribution was made to Brazil, Cuba, Czech Republic, France, Germany, India, Italy, Japan, Malaysia, Poland, Portugal, Serbia, and Spain. | Terminated | April 01, 2014 | 1572 | 45 | Software | Output/Calculation | Software update | Software Update
Z-2635-2011 | 58813 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BEAMVIEW IT, camera imaging system, Part Number: 5410389 &amp; 5852301; |  |  | November 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc - Radiation Oncology | The firm has become aware of a potential issue that may result in a service engineer receiving an electrical shock should the connector cable housing become unfastened from the chassis. | Other | The firm, SIEMENS, sent a "Customer Information" letter to its customers on November 4, 2010. The letter describes the product, problem and actions to be taken. The customers were instructed to include this letter in their BEAMVIEW TI System Owner Manual chapter "Safety Advisory Letters" where it should remain. Siemens has developed an updated kit to install an additional grounding cable on the camera's central control unit to eliminate the problem. The updated kit was delivered with the Customer Information letter.If you have any questions, contact the Service Manager- Regulatory Affairs at 925-246-8407. |  |  |  | Worldwide distribution: | Terminated | November 14, 2011 | 375 |  | Not_Computer
Z-2654-2011 | 57906 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | First Breath Âż Adult Face Tent with Tubing 188cm (5ft.) (Non-latex) |  |  | February 11, 2011 | 2011 | Smiths Medical | A sticky substance on the inside and/ or outside of the Mask | Nonconforming Material/Component | Smiths Medical notified consignees by letter on 2/14/11 titled Urgent Medical Device Correction Notice via U.S. Mail (USA only) This Notice includes an Urgent Medical Device Correction Notice Confirmation Form that provides instructions to consignees to return unused affected product to Smiths Medical and to complete the effectivity tracking form. On 3/17/11 Smiths Medical issued an additional recall letter to extend the recall to 3 additional lots not included in the original recall letter dated 2/14/11. Also, the firm issued a recall letter on 3/17/11 to include the Value Pack (VP1302) General Anesthesia Kit to include 7 kit lots. Contact Recall Coordinator 1-800-348-6064 prompt 4, ext 2273 |  | 600 | 600 units | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Armenia, Brazil, Canada, France, Hungary, Italy, Jordan, Japan, Libya, Poland, Russian Federation, and Saudi Arabia. | Open, Classified |  |  | 88653 | Not_Computer
Z-2657-2011 | 58828 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Silicone CTS (Close to Shaft) Cuff Adjustable Neck Flange Tracheostomy Tube (Wire Reinforced), Arcadia Medical Silicone Air Cuff Adjustable Tracheostomy Tube, Sterile, Latex-Free, |  |  | May 08, 2011 | 2011 | Arcadia Medical Corporation | The product does not contain a punched hole above the cuff access port (a lumen that enables suctioning of secretions proximal to the inflated cuff) as intended by the devices design. This eliminates the user's ability to suction secretions through this suction port, necessitating that the user revert to standard of care methods for suctioning secretions from around the cuffed trach tube shaft. | Process control | Arcadia Medical Corporation sent a "MEDICAL DEVICE RECALL" letter on May 9, 2011 and email on May 8, 2011 to all affected costumers. The letter included; description of the product name, afffected lot numbers, description of the problem, instructions to discontinue use of the recalled devices and to return them to Arcadia Medical for replacement . The letter instructed the customers to email Arcadia Medical with the number of devices in their inventory. For additional information contact Arcadia Medical Corporation at 219-779-9431. |  | 118 | 118 | Worldwide Distribution --USA including the state of Ohio, and the countries of Denmark, Italy, Germany, Turkey, Slovakia, Isreal, France, Sweden, and Romania | Terminated | August 30, 2011 | 114 | 118 | Not_Computer
Z-2661-2011 | 59064 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | 0200086, Triad Alcohol Prep Pads, 100 box |  |  | February 28, 2011 | 2011 | Alere San Diego | The recall of the Triad Group alcohol prep products is due to potential contamination of these products with the bacteria Bacillus cereus, which could lead to life-threatening infections. | Other | Alere sent a CUSTOMER NOTIFICATION letter on February 28, 2011 to all affected custumers.The letter included; description of product, affected lots, statement of potential for contamination which could lead to life-threatening infections. Alere asked its customers to discard the Triad Alcohol Prep Pads provided and complete and FAX the enclosed Verification Form within 10 days to confirm receipt of notice and to indicate the number of discarded/required replacement alcohol prep pads.The letter instructed customers to call Alere Technical Services for additional information at 888-246-7483. |  | 4829 | Estimated 4829 total units | Worldwide Distribution--(USA) Nationwide and countries of Canada, Australia, Japan and Puerto Rico | Terminated | February 29, 2012 | 366 | 4829 | Not_Computer
Z-2658-2011 | 58875 | No | MKZ | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | BD ProbeTec" CT/GC/AC Reagent Pack, catalog #440450, labeled in part ***Becton Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152*** |  |  | September 29, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent packs may be underfilled and contribute to invalid quality control testing or false positive patient results. | Process control | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated September 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to discontinue use of the product and discard any remaining product. BD will replace any discarded products. Additionally, customers were requested to complete and return an attached form whether or not they have any inventory remaining, to confirm receipt of the notification. The letter instructed the customers to fax the completed form to: Attention: Regulatory Compliance at 410-316-4258.For assistance regarding replacements, contact BD Customer Service Department at 1-800-675-0908. All other inquiries should contact BD Technical Services Department at 1-800-638-8663. |  | 130 | 130 kits | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CO, CT, FL, GA, KY, LA, ME, MN, MO, MS, MT, NE, OH, PA, RI, TX, and WA and the countries of Australia, Canada, Europe, and Japan. | Terminated | June 24, 2011 | 268 | 130 | Not_Computer
Z-2663-2011 | 58442 | No | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | Infusable and Infusa Scan Single Patient Use Pressure Infusors; Vital Signs Inc. |  |  | March 04, 2011 | 2011 | GE Healthcare, LLC | Infusable and InfusaScan Pressure Infusors may leak and not apply desired pressure to their IV solution bag and subsequently the rate of infusion may be inadequate. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 25, 2011 with an attached return response form to its customers via Fed Ex. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) DO NOT use infusable or InfusaScan pressure infusors (listed lot numbers).2) Isolate all affected product and contact Vital Signs Customer Service at +1-973-956-5300 to arrange for the return and replacement of the infusor.3) If they have forwarded any of the affected product to any other healthcare institution, forward a copy of this letter to those institutions.4)Return only unused product.In addition, the customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions, and to complete and return the attached confirmation form via fax to +1-973-956-5440.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 954888 | 954,888 units | Worldwide distribution: USA (nationwide) states including: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV, WY; and countries of: Argentina, Australia, Chile, Colombia, Canada, Denmark, Dominican Republic, Finland, Germany, Greece, Hong Kong/Macao, India, Israel, Italy, Korea, Lebanon, Lativia, Mexico, Malaysia, Nicaragua, Netherlands, New Zealand, Paraguay, Philippines, Portugal Panama, Puerto Rico, Peru, Russia, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Taiwan, Turkey, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | October 31, 2012 | 607 | 954888 | Not_Computer
Z-2660-2011 | 58864 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Condyle Kit with Hexalobular, REF 114700 DISCOVERY ELBOW SYSTEM, HUMERAL CONDYLE SET - HEXALOBULAR, CO-CR-MO / YI-GAL-4V ALLOY, IF USED FOR REVISION SURGERY, INSTRUMENTS 414926 &amp; 414923 MAY BE REQUIRED LOT 183300, 1 SET, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE, P.O. BOX 587 WARSAW, IN 46581 USA EXPIRY DATE: 2021-03 |  |  | April 29, 2011 | 2011 | Biomet, Inc. | The firm initiated this recall after becoming aware that a Discovery Condyle Kit with Hexalobula contained two male condyles instead of a male and female condyle. | Packaging process control | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated May 25, 2011 to all affected customers. The notice describes the product, problem, and actions to be taken by the customers.The notice instructed customers to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Additionally, customers were instructed to ship the package priority mail with a copy of the response form.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday 8 AM to 5 PM. |  | 10 | 10 KITS | Worldwide Distribution--USA (nationwide) including the states of AR and SD and the country of Japan. | Terminated | August 29, 2011 | 122 | 10 | Not_Computer
Z-2662-2011 | 58186 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify &amp; Record System, for MOSAIQ |  |  | March 10, 2011 | 2011 | Impac Medical Systems Inc | There's a problem in the database conversion program that is used when upgrading from the Multi-ACCÂżS to MOSAIQ or from one version of MOSAIQ to another. If the gantry start angle is incorrect prior to pressing beam-on, the patient will be mistreated. | Software design (manufacturing process) | The firm, Elekta, sent a "USER NOTICE" to its customers with an attached confirmation of receipt form. The notice describes the product, problem and actions to be taken. The customers were instructed to follow the Interim mitigation instructions noted in the notice -verify the Arc start angle of the Simple ARC with MLC treatment field definition prior to delivery and if the customer identifes a field with an errant gantry start angle, to cease using that plan immediately and consult with the original plan for that field and make the correction to the gantry start angle. In addition, the customers were instructed to complete and return the Confirmation of Receipt via fax to: North America and Asia at 702-992-5002 or Europe and Africa +44 (0)1293 654401, Attn: Elekta OIS Su[pport and/or scan and email to: support@impact.com, europe.support@impact.com or by mail; and distribute this notice to any all users of IMPAC software at their organization who are potentially affected by this issue.For questions or requests for upgrade, email support@impac.com (North America/Asia) or europe.support@impac.com (Europe) or call USA 1-800-488-4672. |  | 840 | 840 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | September 16, 2011 | 190 | 840 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-2659-2011 | 58770 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Precision Ice One-Patient-Use Control Unit and Wrap System (Model # 4BB02BK). |  |  | April 18, 2011 | 2011 | Pro Trainers' Choice Company | The control panel of the Precision Ice One Patient Use Control Unit and Wrap System has a overheating issue. The control panel printed circuit boards were missing a jumper connection across resistors R28 A, B, C, and D. | Other | Pro Trainers' Choice sent a Mandate-Medical Device Recall on April 18, 2011 and April 20, 2011 to all affected costumers.The letter included a description of the product, affected model and lots. It instructed the costumers to discontinue using the Precision Ice, Model 4BBK02BK until the device is returned and corrected by Pro Trainers' Choice. The device should be returned to PO Box 27, Kingston, WA 98346.The customers were instructed to notify their sub-account customers of this recall if the devices were further distributed.Customers can call Pro Trainers' Choice Company at 360-297-3902 or visit the website www.precisionice.com for any questions about this recall. |  | 336 | 336 units | Nationwide distribution including CA, CO, FL, IN, KS, NY, OR, TX, and WI. | Terminated | March 01, 2012 | 318 | 336 | Hardware | Physical Safety Hazards | Return | Remove or Replace
Z-2646-2011 | 58894 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Smith &amp; Nephew Hip Arthroscopy Repair Instrument Tray |  |  | May 19, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Hip Arthroscopy Repair Instrument Tray shipped with an incorrect Instructions For Use (IFU) | Labeling mix-ups | Smith &amp; Nephew sent an "URGENT FIELD CORRECTION NOTIFICATION" letter dated May 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to confirm receipt of the corrected IFU document (10600729 rev B.) and communicate to appropriate facility individuals. Once appropriate individuals are informed, customers were instructed to complete and return information requested at the bottom of the letter. The completed information should be returned to Smith&amp; Nephew via fax at 978-749-1654 or by mail at Smith &amp; Nephew Endoscopy, 150 Minuteman Road, Andover, MA 01810.Questions may be answered by contacting the Manager, QA and Compliance at 978-749-1397. |  | 11 | 11 units | Worldwide Distribution--No USA Distribution; countries of Italy, Portugal, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 356 | 11 | Not_Computer
Z-2670-2011 | 58716 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | SCREENERS Dip Drug Test, Catalog Number: SCD-7MB300 |  |  | April 26, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech's internal review of regulatory files found that the two affected lots were manufactured without proper regulatory clearance. | Process control | Ameditech, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 18, 2011 to the affected customer. The letter describes the product, problem, and actions to be taken by the customer.The letter requested the customer to:(1) Cease distribution of the affected lot and provide remaining inventory count to Ameditech.(2) Destroy any remaining inventory of these lots and document destruction per your site requirements.(3) Ensure that all users of the device receive a copy of the Urgent Medical Device Recall letter.A recall verification form was attached to the letter to be completed and returned by the customer via fax at 1-858-677-0243.Any questions about the information contained in the recall notice should be directed to Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 6000 | 6000 test cards (120 kits) | Nationwide Distribution | Terminated | August 11, 2011 | 107 | 6000 | Not_Computer
Z-2665-2011 | 58884 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | HALO Catheter with Auto 10 Technology, Part Number D-1160-43-S, Catalog Number |  |  | May 23, 2011 | 2011 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster, Inc. has recently become aware that 20 electrodes catheters with Auto ID Technology are not being recognized when connected to the CARTO 3 System. | Pending | Biosense Webster, Inc. sent a Customer Notification Letter, dated May 24, 2011, with an Acknowledgement Form and Instructions for Manually Entering the Catheter Identification in the CARTO 3 System to all US and Outside US (OUS) customers who purchased these products. For customers in the US, two letters and forms will be sent via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively. Instructions for manual entry of the catheter identification in the CARTO 3 System are being included in the communication to customers. These instructions will allow customers to input the identity of the catheter into the CARTO 3 System and be able to utilize the catheter.Customers were instructed to complete and return the attached Acknowledgement Form in accordance with the instructions listed on the form. Customers were instructed to contact their BWI representative for any questions. |  | 431 | 431units (298 in US) | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Japan, and Singapore | Terminated | October 12, 2012 | 508 | 1531 | Not_Computer
Z-2672-2011 | 58966 | No | FAQ | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | RUSCH Urinary Drainage Bag, 2000 ml |  |  | June 02, 2011 | 2011 | Teleflex Medical | Sterile integrity cannot be guaranteed. | Other | Teleflex Medical sent an " URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 2, 2011. The letter identified the product, the problem and the actions needed to be taken. The customers were instructed to immediately discontinue distribution, quarantine any affected products, communicate this recall to any of their customers using the provided customer letter, and complete and fax the enclosed Recall Acknowledgement Form to Customer Service at 1-866-804-9881. Once the fax is received instructions for product return will be provided to customers. In the event that an alternative approach is needed, customers are instructed to contact Teleflex Medical Customer Service at 1- 866-246-6990. For further questions contact your sales representative or Customer Service at 1-866-246-6990. The recall was expanded and an additional letter was sent out on July 20, 2011 to the affected consignees with the above instructions.*****RECALL EXPANDED***** Product 2. Teleflex sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated July 20, 2011, to all affected customers.The letter identified the product, the problem, and the action needed to be taken by the customer. On July 19, 2011 this recall was expanded to include additional product. Please see above strategy information for details |  | 14488 | 14,488 eaches | Nationwide Distribution: AZ, CA, FL, GA, IL, KS, KY, LA, MA, MS, NC, NJ, NM, NY, OH, OK, PA, TN, TX, VA, WI including Puerto Rico. | Terminated | April 05, 2013 | 673 | 28978 | Not_Computer
Z-2671-2011 | 58718 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | B. Braun CelsiteÂż Venous Access Ports , Distributed by Aesculpa Inc. |  |  | April 13, 2011 | 2011 | B. Braun Interventional Systems | The affected products contain the incorrect winged SurecanÂż needle accessory component in certain CelsiteÂż Access Port package configurations. The winged SurecanÂż needle is used to deliver medications and fluids into implanted CelsiteÂż Access Ports.Two winged SurecanÂż needle sizes are provided in CelsiteÂż U.S. package configurations: 20G (PN 4058631) or 22G (PN 16006593A). A 20G or 22G needle | Process design | B.Braun Interventional Systems Inc, (BIS) sent a "Urgent Medical Device Recall" letter on April 16, 2011 to all affected customers.The letter described the problem and the product involved in the recall. Informed consignees that their local B.Braun Interventional System Inc Product Specialists will contact them with the specifics for return/replacement of the product upon return of qualified removed product. They also requested that a CelsiteÂż Access Port Patient Implant Form be filled out for each patient implanted with one of the CelsiteÂż Access Ports from the affected lots. This Form to be provided to them by their local BIS Product Specialist.For additional information contact 800 523-9676 x4459. |  | 83 | 83 | NY, ND, OK ,PA | Terminated | July 06, 2012 | 450 | 83 | Not_Computer
Z-2673-2011 | 58878 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Instrumentation Laboratory GEM Premier 3500 Blood Gas Analyzer |  |  | May 11, 2011 | 2011 | Instrumentation Laboratory Co. | Software upgrade: misuse of the barcode gun by repeatedly clicking can cause the instrument to attach test results to an incorrect patient ID. | Software design | Instrumentation Laboratories sent a "FIELD SAFETY NOTIFICATION" letter dated April 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.Included with the letter was an upgrade package containing Software V7.2.2 disk along with installation instructions. Also included with the upgrade kit was a Upgrade Tracking Form to be completed and returned via fax at 781-861-4207 or via e-mail at ra-usa@ilww.com.Customers with questions were directed to contact their local representative. |  | 840 | 840 units | Worldwide Distribution-- USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE , INDONESIA, ISRAEL, ITALY, JAPAN , KOREA, REPUBLIC OF CHINA, MALAYSIA, NETHERLANDS , PHILIPPINES, POLAND, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, THAILAND ,TURKEY, and UNITED KINGDOM. | Terminated | March 15, 2013 | 674 | 840 | Software | Output/Calculation | Software update | Software Update
Z-2668-2011 | 58832 | No | MIP |  |  | Gastroenterology/Urology | PMA | AMS ActiconÂż Neosphincter Pump, Product Number 72402287, Sterilization Method Steam, American Medical systems Inc., 10700 Bren Road West, Minnetonka, MN 55343. |  |  | May 09, 2011 | 2011 | American Medical Systems, Inc. | American Medical Systems is recalling all unexpired AMS 800 and AMS Acticon Neosphincter Control Pumps.Based upon a review of our product test procedures, we are unable to confirm that all Control Pumps have met our requirements. Test errors could potentially result in a failure to identify a malfunction in the Control Pump activation mechanism. This malfunction may lead to inadvertent activa | Nonconforming Material/Component | American Medical System sent an "URGENT MEDICAL DEVICE RECALL letter dated May 9, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customerCustomers were instructed to return any existing stock . Please contact American Medical System at 1 (800) 328-3881 to obtain a Return Authorization number. Customer Service will also provide a FedEx shipping number for your use to return affected product.For further questions please contact American Medical Systems at 1- (800)328-3881. |  | 225 | 225 | Worldwide Distribution--USA (nationwide) including Washington D.C. and Puerto Rico and the countries FRANCE, GERMANY, CANADA, ISRAEL, SOUTH AFRICA, LEBANON, AUSTRIA, ARGENTINA, SAUDI ARABIA, SPAIN, BELGIUM NETHERLANDS, PERU, VENEZUELA, SRI LANKA, PORTUGAL, TUKEY, BRAZIL AUSTRALIA, CZECH REPUBLIC, TUNISIA. LUXWMBOURG, QATAR, SOLOMON ISLANDS, BOSNIA &amp; HERZEGOVIA, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, ECUADOR, ESTONIA, FINLAND, GREECE, HONG KONG, HUNGARY, ICELAND, INDIA, IRELAND, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PANAMA, POLND, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVAKIA, SLOVENIA, SWEDEN, SWITZERLAND, TAIWAN PROVIDENCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM and URUGUAY. | Terminated | July 06, 2012 | 424 | 15691 | Not_Computer
Z-2685-2011 | 56408 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel DxI 800 Access Immunoassay System with Unicel DXI system software version 4.3 |  |  | July 06, 2010 | 2010 | Beckman Coulter Inc. | Test results may have incorrect ranges. If a default sample type setting on the tests screen is changed while running Unicel DXI system software version 4.3, the new setting is not saved to the system database. The changed sample type setting will reset to the previously saved setting when power is interrupted to the system, or if an operator reboots the system. If the default sample type i | Software design | Beckman Coulter forwarded on 7/6/10 a Product Corrective Action letter with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the UniCel Dxl 600 and Dxl 800 Access Immunoassay systems. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to:A revised version of UniCel DxI System Software that corrects this issue is currently in development. Until the new software version is available, Beckman Coulter has provided an alternate method for changing the default sample type setting. See the appendix included with this letter for additional information.Refer to the attached Questions and Answers document for additional details and instructions, and to determine if their laboratory is running UniCel DxI Sytem Software version 4.3. Customers may wish to confirm that proper ranges and units of measure are being reported for any Access assays with sample type settings that were changed while running software version 4.3.Customers were asked to share this information with their laboratory staff, and retain the notification as part of their laboratory Quality Ststem documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that consignees received the notification.If they need assistance or have any questions regarding this notification, customers were told to contact Technical Support at 1-800-854-3633 in the United States or Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 626 | 626 units | Nationwide in the US and to Canada | Terminated | May 07, 2012 | 671 | 794 | Software | Output/Calculation | Software update | Software Update
Z-2676-2011 | 58807 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Jet-Ventilator Catheter for Adults, 14G, STERILE, NO LATEX, REF 30-02-918-1, VBM Medizintechnik GmbH, Einsteinstrasse 1 7212 Sul z a.N. Germany |  |  | May 09, 2011 | 2011 | VBM Medical Inc. | Product labeled with incorrectly as being appropriate for use in adults when the intended population for use of the product is children and infants. | Error in labeling | VBM Medical Inc sent an "URGENT FIELD SAFETY NOTICE" dated May 9, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The reference and lot numbers, details of the error, along with the incorrect and correct box label, and instructions to the customers to discard the box with the incorrect label or send the complete product back to VBM Medical Inc. for re-packaging. For any further or detailed questions, please contact VBM Medical Inc. at (317) 776-1800.On 7/12/11 upon receiving the recall classification letter, the firm sent out an additional letter to its direct accounts informing them of the recall's classification and reiterating the instructions listed above. |  | 29 | 29 | Worldwide Distribution -- USA ( nationwide ) including the country Montreal QC | Terminated | March 14, 2012 | 310 | 417 | Not_Computer
Z-2677-2011 | 59005 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Remel Spectra MRSA, Ref R01821, 10 plates per package, and Ref R01822, 100 plates per package. The firm name on the label is Remel, Lenexa, KS. |  |  | May 11, 2011 | 2011 | Remel Inc | The product failed to prevent growth of methicillin-sensitive Staphylococcus aureus (MSSA) | Process control | Remel, Inc. sent an "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" letter, dated May 10, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to review their inventory and discard all remaining product lots. Additionally, customers are to complete the Product Inventory Checklist and indicate whether or not they desire replacement of the product or credit. A self-addressed , stamped envelope was enclosed in the letter to use to return the checklist. Questions can be answered by contacting the Technical Services Department at 800-255-6730 or 913-888-0939. |  | 48807 | 187/10-packs and 261/100-packs | Nationwide Distribution--including Puerto Rico. | Terminated | June 18, 2013 | 769 | 48807 | Not_Computer
Z-2693-2011 | 58975 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | THD155428- 15.5F X 28cm Titan HD(TM) Catheter Kit |  |  | May 19, 2011 | 2011 | Medical Components, Inc dba MedComp | Product was packaged with the incorrect valved peelable introducer. The catheter sized is 15.5F and the introducer packaged in the kit is a 15F. The kit should contain a 16F introducer. | Mixed-up of materials/components | Medcomp sent a "PRODUCT ALERT" letter dated May 19, 2011 to all affected customers. The letter described the product, problem, and actions to be taken. The letter is requesting the return of all affected product. Customers are instructed to contact their customer service representative for an RGA (returned good authorization) number. If you have any questions or concerns contact Medcomp at 215-256-4201 ext. 2225. |  | 90 | 90 | Worldwide Distribution -- USA (nationwide) including states of AZ, CA, CO and RI and countries of Australia, Brazil, and Netherlands | Terminated |  |  | 90 | Not_Computer
Z-2692-2011 | 58560 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software versions |  |  | April 13, 2011 | 2011 | Eos Imaging | Print true size, print requests from the EOS acquisition station or the sterEOS image review station may, in certain cases, be affected by a different enlargement factorduring printing. T | Software design | EOS imaging sent a "FIELD SAFETY NOTICE" letter dated April 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to wait for the printing queue to empty before sending a second "print true queue" request with different print settings. Customers should then click the button shown on the instruction sheet to view the print queue.EOS imaging will provide a software update to correct all affected systems at no cost to the customers. An EOS representative will contact customers to schedule a time to upgrade the software.Questions can be answered by calling 33 1 55 25 63 90. |  | 15 | 5 EOS Systems, 10 sterEOS workstations | Worldwide Distribution--USA (nationwide) including the states of CA, IA, MO, VA, and WI, and countries of Canada and Europe. | Terminated | January 18, 2012 | 280 | 15 | Software | Output/Calculation | Software update | Software Update
Z-2687-2011 | 59019 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | The product is the "FASTx Sternal Intraosseous Device". |  |  | November 05, 2010 | 2010 | Pyng Medical Corporation | Device failed to operate; the insertion of the infusion tube of the FASTx Sternal Intraosseous Device was not achieved in live subjects. | Other | Pyng Medical Corp. sent a "FSN/RECALL FOR FASTx STERNAL INTRAOSSEOUS DEVICE, P/N FASTx -H" letter dated November 5, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are to cease distribution and use of all devices. All devices remaining in inventory should be quarantined. Customers are asked to complete the information at the bottom of the letter and return the product by FEDEx. Questions regarding this notice should contact the Director of Operations at 1-604-303-7964 ext. 211. |  | 402 | 402 devices | Worldwide Distribution--USA (nationwide) and the countries of Australia, Austria, Canada, England, Italy, Mexico, Poland, and Sweden. | Terminated | August 10, 2011 | 278 | 402 | Not_Computer
Z-2688-2011 | 58444 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LiNA Loop Endoscopic Monopolar Loop for LSH (Laparoscopic Supracervical Hysterectomy). Large loop, 4mm connector, 1 X REF: EL-200-4 and 10 X REF: EL-200-4, Manufactured by LiNA Medical ApS, Surgery Division. |  |  | December 01, 2009 | 2009 | LINA Medical ApS | Potential for the LiNA Loop to break during operating procedure. | Labeling design | Lina Medical Aps notified customers by newsletter in December 2009 followed by a letter to distributors including warnings added to instructions for use in January 2010. |  | 200 | 200 pieces | GA, MN, PA, WA, TN, MA, IN, Ontario Canada | Terminated | February 23, 2012 | 814 | 2720 | Not_Computer
Z-2678-2011 | 58532 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Systems, CDI blood parameter monitoring system 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe, Catalog AVHCT. |  |  | April 21, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Inaccurate readings may result after Methylene Blue, an intravascular dye, is administered to the patient. The presence of Methylene Blue may interfere with the readings from the shunt sensor for the CDI 500 system in a way that would cause inaccurate pH values. Because the pH value is used to calculate other values, these parameters could also be affected by the presence of Methylene Blue: K | Device Design | Terumo sent URGENT MEDICAL DEVICE CORRECTION letters to Chief of Perfusion, Department of Cardiothoracic Surgery or Director of Operating Room Services, dated 4/21/2011. The letter explained the reason for the correction and the potential hazard. The letter stated that Terumo was updating the Operators Manual for all CDI systems and would provide the manuals to all users when available. Consignees with questions should call 1-800-521-2818. Monday- Friday 8 am- 6 pm. |  | 3767 | 3767 units (all CDI 500 products) | Worldwide distribution, including USA, Australia, Barbados, BELGIUM, BRAZIL, CANADA, Chile, Colombia, Costa Rica, Dominican Republic, ECUADOR, El Salvador, France, Germany, Guatemala, Holland, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Korea, Libya, MALAYSIA, MEXICO, PAKISTAN, PANAMA, Paraguay, Saudi Arabia, SINGAPORE, South Africa, Spain, TAIWAN, THAILAND, The Phillipines, TRINIDAD, UNITED ARAB EMIRATES, Uruguay, Venezuela and Vietnam. | Terminated | August 06, 2012 | 473 | 5582 | Software | Device Operation | Update manual | Other
Z-2691-2011 | 59039 | No | MND | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Disposable Hemostasis Clip, DCH-7-230, Rx only, Single Use only, Cook Medical |  |  | June 07, 2011 | 2011 | Cook Endoscopy | The disposable clip may not deploy after being positioned inside the patient. | Device Design | Wilson-Cook Medical, Inc. (Cook Endoscopy) sent an "URGENT: PRODUCT RECALL" notice dated June 7, 2011 and a "PRODUCT RECALL EFFECTIVENESS SURVEY' dated June 21, 2011 to all affected consignees. The notice described the product, problem, and actions to be taken. Medical users were instructed to quarantine any affected products and return to Cook Medical. The distributors were instructed to contact their customers and coordinate the return of the affected products to Cook Medical. The recalled notice and survey should be completed and returned to Cook Medical via fax to 812-339-7316 or scanned and emailed to CustomerRelationsNA@cookmedical.com. If you have any questions regarding this matter, please contact the Customer Relations Department at toll free-1-800-457-4500, press 4 then enter extension 2146 or by email at CustomerRelationsNA@cookmedical.com . |  | 307 | 307 units | (USA) Nationwide Distribution including the states of: CA, HI, MA, ME, NC and NY. | Terminated | May 15, 2012 | 343 | 307 | Not_Computer
Z-2674-2011 | 58909 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, CaridianBCT, Lakewood, CO 80215. |  |  | December 09, 2010 | 2010 | Caridian BCT, Incorporated | An additional air detection system is being added to all existing systems to detect air in the system's return line. | Software design | CaridianBCT sent a "MEDICAL DEVICE RECALL" letter dated December 2010 to all customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to continue using their systems as long as they followed the instructions in the Operator's Manual. An Acknowledgement of Receipt form was included with the letter to be completed and faxed to 1-303-876-9277 or e-mailed to Regulatory.Affairs@caridianbct.com. A CaridianBCT representative will schedule a visit to facilities to upgrade their systems within the next 6 months.Further information is available by contacting the CaridianBCT Support Center at 877-339-4228 or 303-231-4357. |  | 594 | 594 units | Worldwide Distribution--USA (nationwide) and countries of Australia, Belgium, Canada, China, Hong Kong, Japan, Malaysia, Mexico, Pakistan, Saudi Arabia, Singapore, South Korea, Taiwan, Thailand, and Vietnam. | Terminated | September 12, 2011 | 277 | 594 | Software
Z-2708-2011 | 58913 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Solution Set for Epidural Use, 110" (2.8 m), Male Luer Lock Adapter; a sterile, single use, nonpyrogenic fluid pathway for use with 6200 and 6300 series Flo-Gard Volumetric Infusion Pumps; Baxter Healthcare corporation, Deerfield, IL; product code 2C7554s |  |  | May 31, 2011 | 2011 | Baxter Healthcare Corp. | The lot of Baxter Solution Set for Epidural Use is being recalled because the position of the slide clamp and roller clamp may have been reversed during assembly. The reversed orientation of the slide clamp and roller clamp will cause the set to be incorrectly loaded into an infusion pump, resulting in fluid flow from the patient to the pump, rather than from the pump to the patient. | Employee error | Baxter Healthcare sent Urgent Product Recall letters dated May 31, 2011, via UPS overnight delivery to all affected customers. The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected inventory from their facility and return it to Baxter after calling Baxter Healthcare Center for Service at 1-888-229-0001 to obtain the returned goods authorization The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 5/31/11, and instructed to notify their customers and ministries of health of the recall. *** The recall was expanded to include two additional lots of product code 2C7554s on July 11, 2011 via letter of the same date. The recall strategy remains the same as that of May 31, 2011. *** |  | 13920 | 13,920 units | Worldwide Distribution - USA including Alabama, Arizona, California, Illinois, Kentucky, Louisiana, New Mexico, North Carolina, Ohio, Oregon, Tennessee and Washington, and internationally to Canada and South Africa. | Terminated | June 18, 2012 | 384 | 13920 | Not_Computer
Z-2703-2011 | 58936 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***LINVATEC SRS SHOULDER RESTORATION SYSTEM CrossFT 5.5mm CrossFT Suture Anchor with Two#2 (5 metric) Hi-Fi Sutures***SRS.Linvatec.com***ConMed Linvatec Biomaterials, Ltd. Hemiankatu 6-8L 33720 Tampere, FINLAND***STERILE EO***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA Made in USA***REF CFP-5502***" |  |  | January 26, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL is recalling CrossFT Suture Anchor with Two #2 HI-Fi sutures, Lot 211581, Product Number CFP-5502. There is a possibility that the product may have been incorrectly assembled. The disposable driver may have the incorrect color handle (red versus black). | Employee error | ConMed Linvatec sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated January 27, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers were instructed to check their inventory for affected lots and return for credit. A Reply Form and FedEx Air bill was enclosed with the letter for customers to complete and return with the product.Questions regarding the recall can be answered by contacting the Customer Service Department via e-mail at Custserv1@linvatec.com. |  | 45 | 45 each | Worldwide Distribution--USA (nationwide) and countries of Canada and South Africa. | Terminated | July 18, 2011 | 173 | 45 | Not_Computer
Z-2698-2011 | 58702 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Oxford Femoral Knee System REF 154600, LOT 2272736, WARNING DO NOT RESTERILIZE, STERILE UNLESS PACK DAMAGED, BIOMET UK Ltd, (510K# K011138) |  |  | April 22, 2011 | 2011 | Biomet U.K., Ltd. | This recall was initiated due to a customer complaint which reported that while peeling off the Tyvek lid on the blister pack, very little effort was required to remove it and this could lead to contamination of the device. | Packaging | Biomet sent a "FIELD SAFETY NOTICE" dated April 21, 2011, to all affected accounts. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, locate and remove from inventory and quarantine pending return to their Bioment distributor. Customers were also instructed to complete and return the attached "Fax Back Response Form" to their local Biomet distributor.For questions customers should contact their local Biomet Contact. |  | 12 | 12 devices | Outside the US to Austria, Belgium, and Germany | Terminated | August 03, 2011 | 103 | 12 | Not_Computer
Z-2701-2011 | 58145 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | CODMAN ENTERPRISEÂż Vascular Reconstruction Device and Delivery System 4.5mm x 28mm |  |  | March 21, 2011 | 2011 | Codman &amp; Shurtleff, Inc. | Mislabeled: One of the stick on stickers on the sterile pouch inside of the product box indicates the stent size to be 22mm in length, when in fact the product is 28mm in length. | Labeling design | Codman &amp; Shurtleff, Inc.sent an "Urgent Medical Device Correction Notice" on March 21, 2011.The letter recommend the incorrect label be removed and discarded. There are 3 additional peel-off labels, 2 on the outer box and one on the pouch. All 3 of these contain the correct printed size information (28 mm) that can be used for applying to the inventory sheet, patient record or patient card, as needed. It is NOT required to return any of this product to Codman.Customer Service at 1-800-225-0460 or contact your local Codman Neurovascular Representative |  | 380 | 380 units | Worldwide - Nationwide (USA), Canada, Korea. | Terminated | February 17, 2016 | 1794 | 380 | Not_Computer
Z-2705-2011 | 58376 | No | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Sequoia final driver closure top retention, non-sterile, REF 3384-2, Zimmer Spine, Austin, Texas, The device is used to apply final torque to closure tops of the implant. |  |  | April 04, 2011 | 2011 | Zimmer Inc. | A rounded or partially rounded tip on the screwdriver may not provide the required locking torque to lock the closure top to the polyaxial screw. When the hex tip of the screwdriver is not fully engaged and/or the screwdriver is used off axis, the hex tip of the screwdriver can become rounded or partially rounded. | Device Design | The firm, ZIMMER, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 4, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that each instrument tray had both final drivers (with and without retention feature); contact their customer service representative if they do not have both versions; ensure proper seating of the drive by moving the driver side to side, and complete and return the Distributor Certification Form via fax to: Zimmer Spine at 512-258-0995. The letter states the firm will be distributing the version of the screwdriver with no retention feature beginning in March 2011 and distributing the new surgical technique in April 2011. The firm is including an additional final driver in the instrument tray for final tightening of the closure tops. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance &amp; Regulatory Affairs at 800-655-2614. |  | 505 | 505 devices | Worldwide distribution: USA (nationwide) including the states of: AL, AZ, CA, CT, CO, FL, GA, LA, MI, MN, MO, NC, NY, OH, OK, PA, TN, TX, WA and WI; and countries of: Czech Republic, Denmark, France, Singapore, Spain, Switzerland, and United Kingdom. | Terminated | August 28, 2013 | 877 | 505 | Not_Computer
Z-2706-2011 | 59028 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Brand Name |  |  | May 20, 2011 | 2011 | Staar Surgical Co. | The recall was initiated by Staar Surgical due to an error in the Directions for Use (DFU) for Staar Surgical's Collamer Ultraviolet Absorbing Posterior Chamber Single Piece Foldable Intraocular Lens (Collamer IP IOL), alsoknown as the nanoFLEX IOL. The Device Description erroneously indicates that the lens could beimplanted in the ciliary sulcus. | Other | Staar Surgical send out an URGENT - Medical Device Correction Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to acknowledge receipt of the letter by signing, dating, and faxing the form to Staar Surgical. For any questions call (800) 292-7902, ext 2345. |  | 16750 | 16,750 devices distributed to the field | Worldwide Distribution - USA (nationwide) and the countries of South Africa, Slovakia, Mexico, Hong Kong, and Sri Lanka | Terminated | March 04, 2014 | 1019 | 16750 | Not_Computer
Z-2707-2011 | 58576 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MaxPlus bifuse extension set (minibore) needleless connectors, |  |  | April 20, 2011 | 2011 | Carefusion 303, Inc. | Carefusion has initiated a product recall for the MaxPlus bifuse extension set (minibore) needleless connectors Model MP9232-C as a result of a potential for occlusion. The product may be occluded at the bifused component. | Device Design | CareFusion initiated on May 3, 2011, two distinctive "URGENT MEDICAL DEVICE RECALL NOTIFICATION" via mail delivery, to all affected customers and distributors. The notices identified the product, the problem, and the actions to be taken. Customers were instructed to 1.) Immediately complete and return the enclosed pre-addressed and postage paid Recall Response Card to CareFusion. Note on the card your distributor, the MaxPlus bifuse extension set lot numbers and quantities that will be returned.2.) All recalled product should be returned directly to the distributor from whom it was purchased. Customers will receive their return goods number through their distributor or CareFusion. 3.) Also, enclosed is a Frequently Asked Question (FAQ) Letter to address any questions or concerns related to the use of the MaxPlus" bifuse extension sets.The second notification was specifically addressed to distributors. The distributors were provided specific instructions for product returned. For customers who may have questions, please contact your distributor directly, or call the CareFusion Support Center at 1-800- 562-6018. The Support Center will be in operation from 7 a.m. - 4 p.m. PST. For distributors who may have any questions or require additional materials, please contact the CareFusion Support Center at 1800- 562-6018. Additionally, instructions were given to report adverse reactions experienced via the FDA's Med Watch Program. |  | 26150 | 26,150 units | Nationwide Distribution -- USA including states of: CA, IA, IN, and TX | Terminated | November 28, 2011 | 222 | 26150 | Not_Computer
Z-2695-2011 | 58592 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | RUSCH CrystalClear Tracheostomy Tube, Cuffed, Sterile Pack |  |  | April 04, 2011 | 2011 | Teleflex Medical | Recalled due to air leakage from the main ventilation lumen through air passageway, resulting in potential injury requiring medical intervention. | Other | Teleflex Medical sent an URGENT FIELD SAFETY NOTICE dated January 4, 2011, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock for product, cease use and distribution of stock from the affected lot number and quarantine all products from the affected lot number immediately. Customers were instructed to complete the field safety corrective action acknowledgement and stock status form and fax to their local Teleflex Medical Customer Service Sales Rep.If the customers did have stock from the affected lot number, they were to contact their local Teleflex Medical Customer Service/Sales Rep or distributor as indicated on the attached field safety corrective action acknowledgement and stock status form for a Return Autorization Number.Distributors were to communicate the field safety corrective action notice to their customers and forward the completed Acknoweldgement Form to orders.intl@teleflexmedical.com or fax to +353 1 437 0773.For questions regarding this recall call +353 906 460 838 |  | 328727 | Total 328,727 for all units | Worldwide Distribution - Arabia, Asia, Austria, Belgium, Czech Republic, Denmark, Estonia, Germany, Greece, Ireland, Italy, Spain, Saudi Kuwait, Portugal, Russia Federation, Sweden, Slovenia and South America. | Terminated | March 02, 2012 | 333 | 328729 | Not_Computer
Z-2699-2011 | 58810 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | LIFECATH S PICC |  |  | April 26, 2011 | 2011 | Vygon Corporation | Peripherally inserted central catheter has the potential to leak or break at the hub while in use, which could lead to a loss of blood or catheter migration. | Other | Vygon Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 26, 2011. The letter identified the affected products, problem, and actions to be taken. The customers were instructed to check inventory and quarantine all affected products; immediately cease use and distribution of all affected products; return any affected products using the instruction from Customer Service involving product pickup (contact Customer Service at 1-800-473-5414 for a Return Merchandise Authorization number); and immediately complete and return the attached form via fax to 215-390-2019. Distributors were provided with instructions on how to handle products that were further distributed.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-,ail at customerservice@yygonus.com. |  | 570 | 570 | Nationwide Distribution--USA including states of: AZ, CA, FL, MO, NE, NJ, NY, PA, TX, UT, VA, and WA. | Terminated | June 01, 2012 | 402 | 1170 | Not_Computer
Z-2702-2011 | 58836 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Radiation Treatment Planning System, model H48; |  |  | May 03, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Analytical Anisotropy Algorithm versions when modeling the second source, the collimator jaw positions are read from the first control point of the field. This can cause the delivered dose to be higher than planned locally by up to 10%. | Software design | Varian Medical Systems, Inc. Oncology Systems sent a "Urgent Medical Device Correction" letters on May 3, 2011***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc. at 888- 827-4265. |  | 10649 | 5201 units ***12-12-12 AMENDED TO: 5448 units*** | Worldwide distribution; Nationwide (USA) and the countries of Vietnam, Venezuela, United Kingdom, United Arab Emirates, Ukraine, Turkey, Tunisia, Trinidad and Tobago, Thailand, Taiwan, Syria, Switzerland, Sweden, Sri Lanka, Spain, South Korea, South Africa, Slovenia, Singapore, Saudi Arabia, Russia, Romania, Qatar, Portugal. Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Malaysia, Macau, Lithuania, Libya, Latvia, Kuwait, Kazakjstan, Japan, Italy, Israel, Ireland, Iraq, Indonesia, India, Hungary, Hong Kong,, Guam, Greece, Germany, France, Finland, Egypt, Ecuador, Denmark, Czech Republic, Colombia, China, Canada, Brazil, Belgium, Belarus, Bangladesh, Bahrain, Austria, Australia, Argentenia, Algeria and Canada. | Terminated | July 09, 2014 | 1163 | 10649 | Software | Treatment/Delivery/Therapy | Instructions/Software Update | Software Update
Z-2704-2011 | 58857 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Broselow/Hinkle Pediatric Emergency System, Vital Signs Inc., 20 Campus Road, Totowa, NJ 07512 USA |  |  | April 26, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | Broselow Pediatric Emergency System Kit may contain oxygen masks that may contain a potentially sticky substance on the inside and/or outside of the mask. | Other | GE issued an URGENT MEDICAL DEVICE CORRECTION letter dated April 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the affected product, isolate all affected product, complete the attached confirmation form and fax back per the instructions and return any unused product.For any product forwarded to any other healthcare institution, customers were instructed to forward a copy of the URGENT MEDICAL DEVICE CORRECTION letter to those institutions. For any questions contact Vital Signs Customer Service at 1-800-932-0760. |  | 33002 | 33,002 units | IL | Terminated | January 14, 2013 | 629 | 33002 | Not_Computer
Z-2694-2011 | 58912 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Tobii C-Series Lithium batteries Type TB8A, TB8B, TB12A used with the Tobii C8 and Tobii C12 Speech Generating Communication Devices. |  |  | May 31, 2011 | 2011 | Tobii Assistive Technology, Inc. | Lithium Batteries may overheat and become combustible | Environmental control | Tobii Technology AB sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were advised to stop using the batteries and to remove them from their devices. They were instructed to dispose of the batteries in an environmentally responsible manner as hazardous materials. Tobii will replace the affected batteries. Customers were asked to complete and return an acknowledgement form via fax at 781-461-8213 or e-mail at recall@tobiiati.com.Contact the firm by E-mail: recall@tobiiati.com or telephone: 1-800-793-9227. |  | 2430 | 2430 units | USA (nationwide) | Terminated | October 01, 2012 | 489 | 2430 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-2711-2011 | 58809 | No | ETA | Total ossicular replacement prosthesis. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Kurz, Heinz Kurz Gmb, Medizintecnik, Tuebinger Strasse 3, 7214 Dusslingen, Germany, Surgical Implant, Clip - Piston MVP Titanium, REF: 1006 711, Lot 2 09 13 06. |  |  | April 27, 2011 | 2011 | Heinz Kurz GmbH Medizintechnik | Outer package labels product .75 mm longer than actual size. | Labeling False and Misleading | Kurz Medical, Inc. sent an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" letter dated May 9, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to return unused product for replacement by using an enclosed FedEx label. If the product has been implanted , the letter requests the customers send the date of surgery and size of the implant. |  | 40 | 32, only 8 consigned with U. S. hospitals | Nationwide Distribution--USA (nationwide) including the states of MN, NC, and WA. | Terminated | November 23, 2011 | 210 | 40 | Not_Computer
Z-2709-2011 | 59043 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product Brand Name: Versafitcup Double Mobility Liner Inserter |  |  | May 23, 2011 | 2011 | Medacta Usa Inc | Medacta International has recently realized that the base of the Liner Inserter Ref 01.26.10.0018Lot 105549 of the Versafitcup Double Mobility system could become rusty after the washingprocedures. This is due to an incorrect raw material used by one of our suppliers during theproduction phase of this lot. For this reason, Medacta International has made the decision toremove all the pieces | Labeling False and Misleading | Medacta USA sent an "URGENT SAFETY NOTIFICATION-VERSAFITCUP DM LINER INSERTER" letter dated May 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected device, immediately sequester it and not use the device in surgery. The letter instructs customers to pack the affected device for shipment to Medacta USA under Return Material Authorization 1590. Customer will receive a replacement within a few days.Questions regarding this recall are directed to the Director of Regulatory and Quality at 805-437-7085 ext. 26. |  | 27 | Worldwide: 18 pieces. In the USA: 9 pieces | Worldwide Distribution--USA (nationwide) including the states of CA, ID, IL, and OH, and 9 products to foreign countries. | Terminated | July 12, 2011 | 50 | 27 | Not_Computer
Z-2713-2011 | 58544 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Comprehensive Shoulder Instruments Fracture Stem Inserter w/Version Control, REF 31-406901, Biomet Orthopedics Warsaw, IN, Instrument used to introduce the stem into the canal. |  |  | April 07, 2011 | 2011 | Biomet, Inc. | When the surgeon attempts to remove the inserter from the stem, the inserter may become stuck not allowing it to disengage from the stem. The stem may have to be removed from the patient in order to disengage the inserter from the stem. If the stem is removed from the patient, it could potentially cause a significant delay in the procedure or damage/fracture of the humerus which may require inter | Nonconforming Material/Component | The firm, Biomet Orthopedics, sent two "URGENT MEDICAL DEVICE RECALL NOTICE" letters one dated April 7, 2011 and the second letter dated April 12, 2011 which included an additional lot number, to its customers. The letters described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of the noted product; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the Fax Response Form within three (3) business days via fax to 574-372-1683; use priority carrier for their shipment; If the product was further distributed, they MUST notify and give letter to Hospital personnel, and confirm receipt of this notice by calling, 800-348-9500, extension 3755 or 3756.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8AM to 5PM. |  | 469 | 469 both products | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Australia, Canada, Chile, Columbia, Costa Rica, Italy, Japan, Korea, Netherlands, New Zealand, Spain and Sweden. | Terminated | April 08, 2013 | 732 | 469 | Not_Computer
Z-2710-2011 | 58695 | No | MEQ |  |  | Gastroenterology/Urology | PMA | The following kits contain the UROLOGIX, Coolant Bag, CA1111B, Sterile EO, Manufactured for: UROLOGIX, Inc., 14405 Twenty-first Avenue North, Minneapolis, MN 55447. |  |  | April 28, 2011 | 2011 | Urologix, Inc. | Urologix is recalling certain lots of TargisÂż System Coolant Bag devices distributed over a two week period. The devices are packaged into disposable Procedure Kits containing a microwave catheter, a Coolant Bag and a RTU PlusÂż balloon. The reason for this Recall involves three lots of Coolant Bags that contain an increased risk of leakage of water at the Sensor Module. This leak is detectabl | Nonconforming Material/Component | Urologix, Inc. sent an "IMPORTANT RECALL INFORMATION" letter dated April 28, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs the customers to identify the affected kit lots, remove the Coolant Bag pouch(s) and to complete and return a Return Material Authorization form with the product(s). Additionally, an enclosed Customer Acknowledgement Form should be completed and faxed to 763-475-1443.Customers with questions should contact the Director of Regulatory and Quality at 763-475-1400. |  | 119 | 119 | Nationwide Distribution--including the states of AL, CA, CO, FL, GA, MA, MI, MN, NJ, NY, OR, PA, TN, WV, and including Puerto Rico. | Terminated | April 26, 2012 | 364 | 119 | Not_Computer
Z-2714-2011 | 56848 | No | LNX |  |  | General Hospital | 510(k) | SoftReports versions 1.1.6.x and 1.1.7.x |  |  | June 23, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer is issuing a correction for SoftReports Versions 1.1.6.x released 02/2008 and 1.1.7.x released 06/2008. Potential hazard - when using SoftReports Layout Designer, results displayed in the report may not match the patient information in the report header. | Software design | SCC Soft Computers sent a "RISK TO HEALTH NOTIFICATION TASK" letter dated June 9, 2010 to affected customers. The letter identified the product, problem, and actions to be taken by the customers.Customers were offered an upgrade or hot fix. The letter requested acknowledgement of receipt of the task notification and permission to install the hot fix. |  | 49 | 49 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 26, 2012 | 764 | 49 | Software | Display/Image | Software update | Software Update
Z-2715-2011 | 59174 | No | FPB | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pall Total Nutrient Admixture Filter, 1.2 micron filter |  |  | June 15, 2011 | 2011 | Pall Life Sciences Puerto Rico Llc | Incorrect filter media used (0.2 um filter used instead of the required 1.2 um) within the TNA1 filter used with total parenteral nutrition mixtures. | Labeling mix-ups | Pall Life Sciences LLC, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 15, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructed customers to maintain the product in quarantine. If distributed, the letter request that customers notify the user to discontinue use. Additionally, customers were asked to complete and return an attached Recall Response Form. The letter also included instructions on how to return the product.Please contact the Customer Service Department at 516-801-9923, if you have questions. |  | 15575 | 15,575 units | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, IL, KS, LA, MA, MO, MS, NC, NY, OK, SD, TX, WA, and WY, and Puerto Rico and the countries of Colombia and Costa Rica. | Terminated | January 10, 2013 | 575 | 15575 | Not_Computer
Z-2720-2011 | 58971 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heater Model 333, Cardiopulmonary bypass temperature controller. |  |  | May 02, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The probe jack contacts (terminals) on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with each other. If this situation occurs, the service technician, user, and/or patient could be exposed to electric shock during servicing and/or use. | Device Design | Cincinnati Sub-Zero Products Inc issued an "Urgeant Medical Device Field Action" letter on May 2, 2011 to all affected customers.The letter includes; description of product and problem, affected lots and instructs customers to discontinue its use. The product is to be returned to the Service Department for an update and the customer is instructed to complete and mail Field Action Form.For additional contact the firm at 1-800-989-7373 or (513) 772-8810. |  | 116 | 116 units | Worldwide Distribution and countries of The Netherlands, Kuwait, India, Saudi Arabia, Taiwan and Australia. (USA) Nationwide including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, LA, MA, MD, MI, MIN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, and WI. | Terminated | March 07, 2013 | 675 | 744 | Not_Computer
Z-2731-2011 | 58733 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DYONICS POWER Shaver Blades, 4.5mm Long Curved Full Radius Convex, package of 3. |  |  | May 06, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Product has insufficient weld between the shaft tube and tip, could result in the tip separating from the shaft tube during use. | Process control | Smith &amp; Nephew notified customers by an "URGENT - PRODUCT RECALL 1st NOTIFICATION" letter dated May 6, 2011. The letter describes the product, problem, and actions to be taken. The letter asked customers to call 800-343-5717 (option 3) or send an email to endo.andreturn@smith-nephew.com for a Return Authorization (RA) Number to return product, obtain credit or order replacement blades. Customers are to complete the form and return to the address provided being sure to reference the RA number. For any questions regarding this recall customers were instructed to call 508-337-3731. |  | 744 | 744 units | Worldwide Distribution -- USA (nationwide) and countries of: Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Germany, Hong Kong, Italy, Japan, Mexico, Netherlands, Turkey, and United Kingdom. | Terminated | May 09, 2012 | 369 | 2598 | Not_Computer
Z-2716-2011 | 58855 | No | MYA | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | ultraView Universal DAB Detection Kit, Part number: 760-500 (05269806001) |  |  | May 16, 2011 | 2011 | Ventana Medical Systems Inc | The recall was initiated because Ventana Medical Systems, Inc. (Ventana), hasidentified the presence of precipitate in some DAB chromogen dispensers in two lots of ultraViewUniversal DAB Detection Kit (part number 760-500, lot A05885 and lot B00988) and one lot of iViewDAB Detection Kit (part number 760-091, lot B01057A.) | Pending | The firm, Ventana, sent two letters dated May 16, 2011, an "URGENT Medical Device Recall" notice to customers who received the affected lots and an "URGENT Medical Device Correction" to customers of unaffected lots. The letters described the product, problem and actions to be taken. The customers were instructed to discontinue use of any kits from the lots identified above; follow their internal standard operating procedures for assessing potential impact to reported patient results, If their laboratory had already used kits from these lots; conduct visual inspections of any DAB chromogen dispensers in use as part of their daily maintenance routine; inspect their current inventory and contact their local customer support team (800-227-2155) to replace any unused or partially used kits from the affected lots.Ventana highly recommends use of same-slide controls, especially when testing for targets of therapy, to mitigate any impact on patient results.If you find a dispenser containing precipitate in lots not identified in this letter or have any questions and concerns, please notify/contact your local customer support team (800-227-2155) . |  | 3990 | 3990 units (2981 in US) | Worldwide distribution: USA (Nationwide) and countries of Canada, Germany and Israel. | Terminated | February 10, 2012 | 270 | 8030 | Not_Computer
Z-2719-2011 | 58465 | No | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PERPOS PLS SYSTEM, CATALOG NO. 9045-02 |  |  | March 23, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to K-wires with a distal tip diameter that may be greater than specified causing an interference fit with the canulated screw. Removing the K-wire during a procedure may dislocate the implant resulting in the implant loosening or being pulled out or position.Use of this product may result in a delay of the surgical procedure and/or pull-out of the facet screw | Pending | Interventional Spine sent an RECALL NOTIFICATION letter dated March 23, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. The customer was instructed to take the following actions: Immediately examine inventory and quarantine product subject to recall. In addition, if this product was further distributed, please notify the Vice President Quality Assurance at once. Consignee were asked to also complete and return the enclosed response form as soon as possible. Consignee with questions were directed to the Consignees' Regional Director or the Vice President Quality Assurance.For further questions, please call (949) 472-0006 |  | 864 | 864 units total for all products | Nationwide Distribution including the states of to NC, NY, TN &amp; VA | Terminated | February 14, 2013 | 694 | 864 | Not_Computer
Z-2734-2011 | 58678 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Foundation Knee System Instrumentation, Tibial Broach Assembly. |  |  | April 21, 2011 | 2011 | Encore Medical, Lp | Strike plate may dislodge from the broach stem handle during impaction of the device. | Process control | Encore Medical, Lp (d.b.a. DJO Global) sent out a "Urgent Field Safety Notice" dated April 20, 2011 to all affected customers. The firm also e-mailed the consignees as well. The letter described the product, affected lots, reason for recall and provided recommendations including; asking customers contact firm for a replacement, provided a "Field Safety Response Form" to fill out and return. The firm also e-mailed the consignees as well. For additional information please contact the Director of Commercial Logistics at (512) 834-6330 or the International Customer Service at 512 834-6275. |  | 15 | 15 units | Louisiana and Arkansas and the countries of Germany, Italy, United Kingdom and Japan. | Terminated | April 06, 2012 | 351 | 15 | Not_Computer
Z-2733-2011 | 58686 | No | MXC | Medical magnetic tape recorder. | Cardiovascular | Cardiovascular | 510(k) | SJM Confirm, Implantable Cardiac Monitor (ICM) Models: DM2100, DM2102 (IDE). |  |  | March 31, 2011 | 2011 | St Jude Medical CRMD | The recall was initiated because St. Jude Medical made new software available for the SJM Confirm DM2100 implantable cardiac monitor (ICM) that incorporates enhancements to the devices sensing algorithm. These enhancements are intended to provide improved sensitivity and specificity for diagnosing arrhythmias. The user is presented with the option to upgrade new software into the SJM Confirm ICM | Software change control | St. Jude Cardiac Rhythm Management Division sent an "Important Product Correction" letter on March 25, 2011 to all affected customers.The letter explains the reason for the recall and provided recommendations including a software upgrade by SJM Field Clinical Engineer or Technical Services. Customers were recommended to fill out and return a "Physician Acknowledgement Form".For additional information please contact your local St.Jude Medical Sales Representative, Field Clinical Engineer or the Technical Services Department at 800-722-3774 |  | 83 | 83 units | Worldwide distribution: (USA) Nationwide including the States of CA, WA, FL, WV, NY, MD and LA, and the country of Germany | Terminated | July 19, 2011 | 110 | 83 | Software | Output/Calculation | Software update | Software Update
Z-2748-2011 | 58525 | No | KMF | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(k) | smith&nephew; NO-STING SKIN-PREP No-Sting Protective Wipes Effective protection between tape and skin Reduces risk of tape stripping Improves tape, film and appliance adhesion No-sting, non-irritating Reduces friction:Packaged 50 wipes per box. |  |  | April 04, 2011 | 2011 | Smith &amp; Nephew Inc. | Smith &amp; Nephew, Inc., St. Petersburg, FL is recalling multiple lot numbers of the following products: REMOVE Universal Adhesive Remover Wipes, UNI-SOLVE Adhesive Remover Wipes, SKIN-PREP Protective Wipes, PERI-PREP Protective Wipes, and NO-STING SKIN-PREP Protective Wipes. These products were manufactured by H&P; Industries dba The Triad Group who has initiated a recall of products manufactured i | Material/Component Contamination | Smith &amp; Nephew Inc. sent an "URGENT DEVICE CORRECTION" letter dated April 8, 2011 to all customers who received the recalled products. The letter described the product, problem and the appropriate actions to be taken. Customers are instructed to immediately discontinue using the recalled products, examine their inventory, and follow the return or disposition instructions provided. The letter instructed customers to seek an alternative product for use. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. If you have any questions contact the Recall Hotline Center at 1-888-613-0271. |  | 108490 | 108490 cartons (471800 wipes); 192,945(kits containing wipes) | Worldwide Distribution -- USA (nationwide) including the states of: NC, PA, NM, OR, VA, TN, NY, OR, AL, IL, MN, TX, and Puerto Rico and the countries of: Australia, Canada, England, France, Germany, Hong Kong, Japan, Mexico, New Zealand, Singapore, South Africa, Taiwan, and United Kingdom | Terminated | July 18, 2011 | 105 | 55600840 | Not_Computer
Z-2747-2011 | 58845 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Quinton BETA-CAP Adapter; a Peritoneal Dialysis accessory to the PD catheter used to connect the catheter to the PD transfer set; Covidien product code 8814-661001, Baxter product code 5K4560 |  |  | May 19, 2011 | 2011 | Baxter Healthcare Renal Div | A change made in the Quinton BETA-CAP adapter used to connect the patient PD catheter to the Baxter transfer set has resulted in a reduction in thread engagement which may lead to an increased risk of separation and the potential to develop peritonitis. | Component change control | Baxter sent an URGENT PRODUCT RECALL letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The letter described the problem of loose connections when using the Quinton BETA-CAP with Baxter transfer sets. Baxter requested the dialysis centers to contact their patients currently using a Quinton BETA-CAP adapter to schedule an office visit with their attending physician as soon as practical tohave the connection inspected or replaced. If they choose to continue use of the Quinton BETA-CAP adapter, Baxter recommends the clinician exercise extra caution and vigilance to ensure the connection between the Quinton BETA-CAP adapter and the Baxter Transfer Set is secured and tight. The clinicians were instructed to also check their inventory for the Quinton Beta Cap adapter 8814-661001 purchased directly from Baxter, and contact Baxter at 888-229-0001 to arrange for the return and credit of the unused product. Baxter recommends that the Covidien BETA CAP adapter no longer be used and to instead exclusively use the Baxter Titanium Adapter (5C4129) which has sufficient thread geometry for maintaining a secure connection with the Baxter Transfer Set. The customers were also instructed to complete and fax the enclosed customer reply form to (847) 270-5457, indicating the quantity of Quinton Beta Cap adapters being returned and acknowledging the understanding and dissemination of the information in the letter. |  | 13942 | 13,942 units | Class II Recall -- Nationwide Distribution | Terminated | May 21, 2012 | 368 | 13942 | Not_Computer
Z-2744-2011 | 58919 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Cambridge Endoscopic Autonomy Laparo-Angle 5mm Metzenbaum Scissors with Monopolar Cautery, 45 cm |  |  | May 20, 2011 | 2011 | Cambridge Endoscopic Devices, Incorporated | PET shrink sleeve insulation material dislodgingfrom the distal tips of Autonomy Laparo-Angle (ALA) Metzenbaum Scissors | Process control | Cambridge Endoscopic Devices, Inc. sent an "IMPORTANT MEDICAL DEVICE RECALL" letter dated May 20, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are instructed to return unused inventory using RGA Number 301. Returned devices will be replaced.Questions regarding this recall, call 508-405-0790 for assistance. |  | 255 | 255 units | Worldwide Distribution--USA (nationwide) and the countries of Armenia , Australia, Austria, Chile, China, Denmark, Hong Kong, India, Ireland , Italy, Japan, Saudi Arabia, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | July 24, 2012 | 431 | 902 | Not_Computer
Z-2741-2011 | 58655 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | JASMINE Full Body (Patient) Lift, electrically-powered device with a 500 lb. weight capacity. |  |  | April 01, 2011 | 2011 | Invacare Corporation | Component Defect. The rear mast connector on certain patient lifts distributed by the firm, may break. Failure of the rear mast connector, while in use would result in a patient being dropped, possibly resulting in serious injury or death to the patient or care provider. | Other | INVACARE contacted their customers via telephone on 4/8/2011 and 4/11/2011, and informed them of this recall. During these telephone conversations, the customers were asked to immediately cease use of the JASMINE patient lifts and to quarantine all product on-site until further notice from Invacare. The firm followed up on the telephone recall notification with a recall letter dated: April 18, 2011. The recall letter was sent via registered mail or Federal Express so as to verify customer's receipt of the recall notification. The recall letter refers the customer/care provider back to the telephone conversations. and notifies the customer of the possibility of failure of the rear mast connector which could result in a patient's possibly being dropped from the patient lift. The letter asks that the customer remove the suspect patient lift(s) from service and quarantine the device in order to prevent further use until such time as an Invacare Rep. contacts the customer by telephone and makes arrangements to pick-up the affect product. The letter is accompanied by a Response Form and a Product Tracking Sheet which lists all of the JASMINE devices shipped to the customer's facility. The customer is asked to complete the attached Response Form and to fax the form back to Invacare at the telephone number provided on the bottom of the form. The signed Response Form provides information to Invacare documenting the customers receipt of the form and understanding of the instructions contained therein, and also; the number of affected devices remaining on-hand at the customer's location, and the number of devices further distributed by the customer. In those instances in which the products have been further distributed, the customer is asked to 'carry out the instructions of the recalling firm and extend the recall to their sub-account customers. |  | 274 | 274 units | Worldwide Distribution - USA including DC, FL, GA, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NJ, NV, NY, OH, OK, PA, RI, SC, TN, VA, WA, and WI and the countries of Canada, Ireland, United Kingdom, Norway, Austria, Germany, and Australia. | Terminated | February 28, 2013 | 699 | 274 | Not_Computer
Z-2745-2011 | 58908 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HALO Ultra Ablation Catheter: |  |  | May 25, 2011 | 2011 | Barrx Medical, Inc. | Complaints of malfunction of the HALO ULTRA Ablation Catheter related to the device pivot mechanism's inability to return the electrode cap to a flat neutral position. | Component design/selection | BARRX Medical, Inc. sent a "PRODUCT RECALL NOTIFICATION" letter dated May 25, 2011 via Federal Express to all of their customers. The letter described the product, problem and actions to be taken. The customers were instructed to discontinue use of the recalled product. BÂżRRX Medical Territory Managers will be in contact with each facility and arrange for pick up of the recalled product. The "Steps for Voluntary Recall" form that is enclosed with each letter provides customers with complete instructions on segregating the product, completing the Tracking/Verification Form, obtaining Returned Good Authorization (RGA) number and preparing product for pick up. Customers are to complete and return the Tracking/Verification Form by fax to 408-738-1741 which should include amount of product on hand, contact name and hospital/distributor information. Questions should be directed to BÂżRRX Medical Customer Service at 888-662-2779. |  | 936 | 936 devices | Worldwide Distribution -- USA (nationwide) and the countries of the United Kingdom and China | Terminated | September 07, 2011 | 105 | 936 | Not_Computer
Z-2746-2011 | 58907 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Wire Caddy, Stryker Instruments Wire Caddy, REF 505-505, STERILE, QTY 1, MANUFACTURED FOR STRYKER INSTRUMENTS, KALAMAZOO, MI USA 49001 (269) 323-7700 (800) 253-3210. |  |  | May 27, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | The towel clamps packaged with the wire caddy have been found to have blood stains on them which could result in infection or blood borne pathogen transmission. | Material/Component Contamination | Stryker Instruments sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter to all affected customers. The letter described the product, problem, and actions to be taken by the customers.The letter instructs customers to quarantine affected product and return the recalled product using the shipper supplied to them by Stryker. Replacements will be issued to customers upon receipt of the product.For questions regarding this recall please call 1-800-800-4236 ext. 3808. |  | 522 | 87 boxes, 435 untis | Nationwide Distribution-- including the states of Alabama, Arizona, California, Florida, Illinois, Kansas, Maryland, Michigan, Missouri, Pennsylvania, Texas and Virginia. | Terminated | August 03, 2011 | 68 | 522 | Not_Computer
Z-2723-2011 | 58536 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Pathfinder NXT Closure top driver - short, REF 3566-2, Zimmer Inc., Warsaw, IN |  |  | May 20, 2011 | 2011 | Zimmer Inc. | Zimmer Spine, Inc. is initiating a voluntary recall regarding an increased risk of rounding or partial rounding of the hex tip on the 3566-1 (long) and 3566-2 (short) Closure Top Driver used for insertion, provisional tightening, and final locking of Closure Tops during the screw and rod construct assembly of the Pathfinder NXT Pedicle Screw System.When the hex tip of the Closure Top Driver is | Device Design | Zimmer, Inc. issued an "URGENT MEDICAL DEVICE RECALL" letters dated April 2011, to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed that (THE GOLD CLOSURE TOP DRIVERS 3566-1 and 3566-2 ARE NOT TO BE USED FOR FINAL TIGHTENING); an alternate instrument for final construct tightening will be included in the instrument set 3570-1 short and 3570-2 long (These instruments are silver instead of gold for easy identification and do not include a retention feature for holding the Closure Top); to ensure final drivers are added to each PathFinder NXT set and inspect Closure Top Drivers for rounded hex wear; contact their customer service representative if they do not have both versions; and complete and return all attached forms via fax to: Zimmer Spine at 512-258-0995. The letter states that the firm will revise the surgical technique to show the updated uses of the screwdrivers. Please keep Zimmer Spine informed of any adverse events associated with the device or any other Zimmer spine product by calling 800-326-0635. For any further questions or concerns contact the Vice President Quality Assurance &amp; Regulatory Affairs at 800-655-2614. |  | 135 | 135 devices | Worldwide Distribution -- USA (nationwide) including states of: AZ, CA, WI, MI, KS, OK, TX, GA, TN, FL, NJ, CT, PA and the countries of: Australia and Italy. | Terminated | October 07, 2011 | 140 | 269 | Not_Computer
Z-2739-2011 | 58562 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reverse Shoulder Prosthesis System RSP Humeral Socket Shell DJO 9800 Metric Blvd. Austin, TX 78758 STERILE R Model numbers: 508-00-000 and 508-00-008. |  |  | April 18, 2011 | 2011 | Encore Medical, Lp | The device has the potential to not appropriately mate with the humeral stem. | Nonconforming Material/Component | DJO Surgical sent an "URGENT FIELD SAFETY NOTICE" letter dated April 18, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.The letter instructed customers to quarantine affected products and to contact Customer Service to obtain replacement products and a RMA.Questions are directed to the Director of Commercial Logistics at 512-834-6330. |  | 81 | 81 units | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, LA, MA, MD, ME, NM, NY, OK, PA, SC, TN, TX, UT, and WA and the countries of Germany, Italy, and United Kingdom. | Terminated | January 25, 2012 | 282 | 81 | Not_Computer
Z-2738-2011 | 58547 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Large Roller Pump 6" Diameter, Catalog 801041, Terumo Cardiovascular Systems, Corp., Ann Arbor, MI |  |  | March 31, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The firm has received multiple reports of the lower roller pump tube clamp mechanism failing to function in a small population of devices, making it difficult to remove or insert the tubing. If the failure occurs, the clamp dial will no longer be able to control opening/closing of the tube clamp mechanism. | Device Design | Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated March 21, 2011. The letter identified the product, the problem and the action needed to be taken by the customer. An additional URGENT MEDICAL DEVICE RECALL correction notification and cover letter was sent to all consignees on XX/XX/XX to explain the reason for the expansion of the recall and will provide an update to the timing of implementation of the field correction.Expansion now includes devices manufactured between 2006 and March 2012.The correction strategy is to issue the customers a letter and addendum to the Operator's Manual that specifies what the user should do in the event of a tube clamp failure for any reason. The letter states that the firm is still investigating the issue and the correction. The letter states that in the coming months when a resolution is identified, consignees may be contacted to make arrangements for a upgrade to the system. There is no specific timeline for the correction to be made.If you have any further questions, please call a Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours Monday - Friday, 8 Am - 6 Pm EST. |  | 111 | 111 devices. | Worldwide Distribution -- USA (nationwide) and the countries of Argentina, Australia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, Indonesia, Japan, Jordan, Korea, Kuwait, Malaysia, Mexico, Pakistan, Republic of Georgia, Russia, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, The Phillipines, Turkey, United Arab Emirates and Vietnam. | Terminated | August 28, 2013 | 881 | 3679 | Not_Computer
Z-2740-2011 | 58792 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NucliSens EasyQ Enterovirus serotype 68 v1.1: Catalog number: 200300 |  |  | May 10, 2011 | 2011 | bioMerieux, Inc. | The NucliSENS EasyQ Enterovirus Reagents are unable to detect Enterovirus serotype 68, in respiratory samples. | Device Design | BioMerieux Vitek, Inc. sent an "Urgent Product Correction Notice" letter on/about May 10, 2011.The letter included; a description of the product, problem, possible impact to the patient and recommendations. Recommendations included instructing customers to discontinue the use of the kit for the detection of Enterovirus 68 and Rhinovirus 87 serotypes and noted the corrected IFU will be available the first week of April 2011. An Acknowledgement form was attached to be returned via fax.For additional questions please contact your local BioMerieux Clinical Customer representative at (800) 682-2666, option 3, option 2. |  | 42 | 42 kits | (USA) Nationwide distribution including the states of MD, MI, MN, MT, NY and PA | Terminated | October 06, 2011 | 149 | 42 | Not_Computer
Z-2742-2011 | 58847 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG Hemostasis System Level II Control: |  |  | March 04, 2011 | 2011 | Haemoscope Division of Haemonetics Corp | Haemoscope received a report that during the process of Validation, Level II QC was running with an out-of-range low MA (Maximum Amplitude). | Other | Haemoscope, a Division of Haemonetics Corp, sent a letter dated March 2, 2011 to all of the affected customers. The letters informed the customers that Haemoscope had received a higher than normal inquiry rate for the Level II Quality Control Kits, part 07-008, lots HMO196 and HMO 197. The accounts were requested to return their stocks of those two lots for further testing and evaluation and to complete the enclosed acknowledgement form, confirming receipt of the letter and the number of kit boxes being returned, and fax it back to Haemoscope at 847-588-0455. The returned kits will be replaced at no charge.For additional information please contact Haemoscope at 847-588-0453, ext-217 . |  | 568 | 568 kits | Worldwide Distribution: (USA) Nationwide and the countries of Canada, Czech Republic and the Philippines | Terminated | September 25, 2012 | 571 | 568 | Not_Computer
Z-2760-2011 | 58335 | [6 empty] | J-Vac Drain Adaptor Manufactured for: Ethicon, Inc. Somerville, New Jersey, 08876-0151. |  |  | March 24, 2011 | 2011 | Ethicon, Inc. | There is a potential for the sterile barrier to be compromised in multiple lots of Blake silicone drains, Blake silicone drain kits, Blake cardio connectors, J-VAC reservoirs, J-VAC drain adaptors. | Packaging process control | Ethicon, Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL" letter dated March 24&amp; 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the products and to return the products to Stericycle. The letter includes shipping labels for product return. Ethicon will send replacement product. Customers are asked to return a completed Business Reply Card to Stericycle regardless of whether or not they have the products subject to this recall.Customers with questions regarding this recall are directed to call 1-866-664-1401. |  | 846990 | 846, 990 | Worldwide Distribution-- USA (nationwide) and Puerto Rico and the countries of Armenia, Australia, Austria, Azerbaijan, Columbia, Cyprus, Czech Republic, Denmark, French Guinea, French Polynesian, Germany,Israel, Italy, Japan, Kuwait , Netherlands, New Caledonia, New Zealand, Norway, Singapore, Slovenia, South Africa, Spain, Belarus, Belgium, Brazil, Canada, Chile, Dominica Republic, Egypt, Estonia, Finland, France, Greece, Guadeloupe, Hong Kong, Hungary, Ireland, Latvia, Lithuania, Malaysia, Martinique, Mozambique, Portugal, Reunion, Russia, Saudia Arabia, Sweden, Switzerland, Ukraine, and United Kingdom. | Terminated | January 10, 2014 | 1023 | 846990 | Not_Computer
Z-2753-2011 | 59024 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bariatric Target Ring (Nitinol) (a thin, circular nitinol ring) Model # K-9536-N, Lot #08-001, Lot #08-002 &amp; Lot #08-003. One ring per plastic pouch, three rings are then placed into a plastic box with lid. This box is then placed into a solid white box. Item is shipped and sold non-sterile. |  |  | May 24, 2011 | 2011 | Kapp Surgical Instrument Inc | During an FDA Inspection on 3/21/11 - 3/22/11 at Kapp Surgical an uncontrolled instruction for use was discovered which contained sterilization parameters which differed from the instructions contained in the device master record. Kapp Surgical decided to conduct the recall because we cannot be sure that incorrect instructions were not shipped. | Labeling mix-ups | Kapp Surgical Instrument, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to check their inventory and quarantine any product affected by this recall. The firm will make arrangements for returning the products. Replacements will be made by Kapp Surgical.Questions about this recall are directed to the QA Manager at 800-282-5277. |  | 222 | 222 units (rings) | Product was distributed to the following states: OH, PA &amp; SC. Product was also shipped to Saudia Arabia. | Terminated | February 19, 2013 | 637 | 222 | Not_Computer
Z-2727-2011 | 59110 | [6 empty] | Stainless Steel, Progressive Lower, .016: Arch Wires. |  |  | January 07, 2008 | 2008 | Highland Metals | Boxes of stainless steel orthodontic wires were mislabeled as size.018" | Employee error | Highland Metals sent out "Highland Metals Advisory Notice Mislabeld Products Alert" and "Advisory Notice" letters from January 11, 2008 to April 18, 2008 to all affected costumers.The letter included: product name and affected lots. It advised them of the mislabeling and instructed them to return the product for replacement.For additional information contact Highland Metals at (408) 271-2955 |  | 2200 | 22 packs of 100 | Worldwide distribution including the countries of Colombia, France, Scotland, Germany, Australia, Mexico, England, United Kingdom, Thailand; and (USA) Nationwide including the states of OK, CA, NJ, AZ, MD, LA, IL, PR, FL, NE, IN, TX, NH, NY, CO and PA. | Terminated | July 07, 2011 | 1277 | 14600 | Not_Computer
Z-2755-2011 | 58850 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ElevateÂż System with InteProÂż Lite", Anterior and Apical Prolapse Repair System, REF720093-01, Sterile EO, Manufactured by:American Medical Systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USA |  |  | May 09, 2011 | 2011 | American Medical Systems, Inc. | This is an important notification regarding the ElevateÂż Anterior and Apical Prolapse Repair Systems, Part Number 720093-01, manufactured between August 3, 2010 and January 17, 2011. We have recently received complaints stating that the eyelets were missing from the center graft of some units within the scope of this issue.The complaint rate to-date for the affected product is 0.045%, however th | Process control | American Medical Systems (AMS) sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 9, 2011 to all customers who received the recalled product. The letter identified the product, problem, and emphasize that the recalled product should not be implanted. The customers were instructed to examine their inventory for the recalled product and call AMS Customer Service at 1-800-328-3881 (select option 1) for an exchange of the product. For assistance in evaluating inventory the customers are instructed to contact their local AMS sales representative. |  | 8563 | 8563 | Worldwide Distribution -- USA (nationwide) including DC and Puerto Rico and countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, FRANCE, GERMANY, PORTUGAL SPAIN, UNITED KINGDOM, ARGENTINA, COLOMBIA, ESTONIA, FINLAND, GREECE, IRELAND, ISRAEL, ITALY, LEBANON, NETHERLANDS, NEW ZEALAND, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SWEDEN, SWITZERLAND, TAIWAN REPUBLIC OF CHINA, VENEZUELA | Terminated | June 28, 2012 | 416 | 8563 | Not_Computer
Z-2761-2011 | 55227 | No | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active Cortisol EIA, 96 Wells, DSL-10-2000 |  |  | July 20, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the listed lot numbers of Active Control ELISA reagent will result in falsely high values for patient samples due to a shift in the calibration of the assigned standard values.Beckman Coulter recommends that patient results, generated with the affected lots, be evaluated in the context of other diagnostic tests and the clinica | Other | Beckman Coulter sent a Product Corrective Action (PCA) letter dated July 13, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use until further notice and discard remaining affected product. Customers were instructed to contact Beckman Coulter at 1-800-231-7970.Newer lots (lot number 991069 or higher) of the product have been recalibrated to restore the performance as described in the Instructions for Use (IFU). For method correlation information, please contact Technical Support at the phone number listed below.Consignees were asked to share this information with laboratory staff and retain the PCA letter for their Quality System documentation, and complete and return the enclosed Response Form within 10 days so that the firm can be assured that they received the notification.Questions regarding this notification were directed to the DSL Technical Support Center at 1-800-231-7970 in the United State and Canada, or their local Beckman Coulter Representative. |  | 1369 | 1369 Kits | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | April 27, 2012 | 1012 | 1369 | Not_Computer
Z-2764-2011 | 58976 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerator; |  |  | May 18, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Coolant System input water supply manifold may leak in some Clinac Linear Accelerators, posing a risk of electrical shock to any person working within the protective housing. | Other | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 18, 2011, via certified mail, return receipt requested. The letter identified the product, problem and recommended actions to be taken. The letter states that customers will be contacted by Varian Service to arrange for replacement of the recalled product. For additional information contact Varian Medical Systems, Inc at 888- 827-4265. |  | 93 | 93 | Worldwide Distribution -- USA (nationwide) and countries of: Australia, Columbia, Denmark, France, Ireland, Italy, Japan, Morocco, Poland, Portugal, Saudi Arabia, South Korea, Spain, and United Kingdom. | Terminated | March 13, 2012 | 300 | 93 | Not_Computer
Z-2763-2011 | 58974 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA, versions 8.0 and 8.1; |  |  | May 17, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An issue has been detected when using the RT Chart function in ARIA versions 8.0 and 8.1. When users select the "Convert to Setup Field" option, a Set-up field is created, that removes the DMLC plan components in order to create an open collimator field for setup purposes. | Software design | Varian Medical Systems, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 17, 2011, via Certified Mail to US; Federal Express to Canada. Foreign distribution will be handled by Varian Managers. The letter contained recommended actions. For additional information contact Varian Medical Systems, Inc at 888- 827-4265.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** |  | 357 | 357 | Worldwide Distribution -- USA (nationwide) and countries of Algeria, Australia, Austria, Belgium, Brazil, Chile, China, Canada, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Iraq, Ireland, Italy, Japan Jordan, Kuwait, Malaysia, Morocco, Netherlands, Oman, Pakistan, Panama, Portugal, Romania, Russia, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Tunsia, Turkey, Ukraine, United Arab Emirates, and United Kingdom, | Terminated | March 15, 2012 | 303 | 357 | Software
Z-2762-2011 | 58830 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving Communication System |  |  | May 10, 2011 | 2011 | Siemens Medical Solutions, USA, Inc | For systems upgraded to syngo Dynamics version 9.0, if a legacy report generated by version 3.x or earlier is opened in the version 9.0 portal, the legacy report may be deleted and replaced with a new report from the version 9.0 reporting system or a blank report. | Software design | Siemens Medical Solutions, USA, Inc sent a letter entitled "Customer Safety Advisory Notice" to all direct accounts on May 10, 2011. The letter stated the name and version of the software, problem noted, a work-around to be implemented until a service patch is created to fix the problem, and that the letter should be placed in the software's operator's manual to be seen by all who utilize the software. |  | 38 | 38 | Nationwide Distribution | Terminated | July 09, 2012 | 426 | 38 | Software | Output/Calculation | Software update | Software Update
Z-2765-2011 | 58566 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost 94, Diagnost 96, Diagnost 97 Universal X-Ray System |  |  | April 18, 2011 | 2011 | Philips Healthcare Inc. | Updated Instructions For Use: During an examination where the patient is positioned on a -rotatable- chair with knees pointed to the rear of the system, the patients lower legs or ankles may be injured/broken | Device Design | Philips Healthcare notified accounts by letter on April 18, 2011. The letter advised users of the problem and provided updated instructions for proper placement of the patient during the examination. An Updated Instruction For Use was provided.For further information and support customers were instructed to contact the Philips representative. |  | 52 | 52 units | Nationwide Distribution | Terminated | February 15, 2013 | 669 | 52 | Not_Computer
Z-2766-2011 | 58994 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam Linear Accelerators (aka Trilogy Mx) |  |  | May 19, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Imaging arms of the TrueBeam Accelerator may have loose encoder pulleys that could lead to inaccurate readout of arm geometry. | Device Design | Varian Medical Systems, Inc. Oncology Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letters on May 19, 2011, via certified mail, return receipt requested in the US. The letter identified the product, problem and recommended actions to be taken. For additional information contact Varian Customer Support at 1-888- 827-4265. |  | 77 | 77 | Worldwide Distribution -- USA (nationwide) and countries of: Canada, Denmark, France, Germany, Isreal, Italy, Netherlands, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 13, 2012 | 299 | 77 | Not_Computer
Z-2767-2011 | 58981 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 System, VITEK 2 60 110V (Catalog #27225) and VITEK 2 60 220V (Catalog #27226). |  |  | June 02, 2011 | 2011 | Biomerieux Inc | Cards may not load into the incubator correctly causing them to be loaded in the wrong order which could result in the results being linked to the wrong accession number and reporting results erroneously. | Device Design | Biomerieux, Inc. sent an "URGENT PRODUCT CORRECTION NOTICE" dated June 2, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers.The letter explains how to recognize the issue and instructs customers to take certain steps to resolve the problem. Additionally, an Acknowledgement Form is attached to the letter for customers to complete and return via fax to 919-314-4539. A firmware update will be applied by Biomerieux, Inc. technical support representatives.Questions regarding this recall are directed to the Customer Service Department at 1-800-682-2666, option 3. |  | 627 | 627 instruments | Worldwide Distribution-- USA (nationwide) and the countries of Austria, Australia, Belgium, Brazil, Canada, China, Colombia, Germany, France, India, Israel, Italy, Japan, Korea, Mexico, The Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Arab Emirate, and United Kingdom. | Terminated | June 06, 2013 | 735 | 1734 | Not_Computer
Z-2769-2011 | 58993 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The SIU (System Interface Unit) is a component in the Tool Connection Unit of the PercuNav system. It is part of the position sensor system used to calculate the position of the instrument. |  |  | June 01, 2011 | 2011 | Philips Medical Systems North America Co. Phillips | PercuNav Tool Connection Unit may contain an electronic component that does not meet specification, which could cause registration difficulties and/or inaccurate instrument tracking. | Nonconforming Material/Component | Philips Healthcare sent an "MEDICAL DEVICE CORRECTION" letter dated May 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The notification letter instructs customers to follow all directions in the Instructions for Use. A Philips Healthcare representative will schedule an appointment to replace the Tool Connection Unit at no charge to the customers.If further information is needed concerning this recall please call Philips Customer Service at 866-767-7822. |  | 10 | 10 units (6 in US and 4 internationally). This was later changed to 7 US and 3 International) all serial codes remain the same | Worldwide Distribution--USA (nationwide) including the states of CA, GA, OH, WA, and Washington DC, and the countries of Australia, India, Netherlands, and South Korea. | Terminated | February 15, 2012 | 259 | 10 | Not_Computer
Z-2781-2011 | 54621 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | DICOM Query/Retrieve, version 1.1.4. Individual part # 9733808, or as part of a system order, DICOM Q/R Software Kit, part # 9733933. DICOM Query/Retrieve software is installed on Medtronic StealthStation systems. |  |  | June 02, 2011 | 2011 | Medtronic Navigation, Inc. | Software allows incorrect patient or patient history to be displayed. | Software design | Medtronic Navigation sent a "FIELD SAFETY NOTICE" letter dated June 2, 2011 to all affected customers. The letter identified the product, problem and description of the corrective action. The letter states that Medtronic field representatives will visit each consumer/user and upgrade the version of DICOM Q/R to 1.1.5. Further information is available by contacting Medtronic Navigation Technical Services at 1-800-595-9709. |  | 90 | 90 | Worldwide distribution - USA (nationwide) -- including states of: TX, NC, ID, FL, OH, UT, CA, TN, IL, AZ, NE, LA, PA, OR, AR, MI, NJ, MD and countries of: Canada, Denmark, Finland, Great Britain, Hungary, Ireland, Japan, Korea, Netherlands, Poland, South Africa, Spain, and Taiwan. | Terminated | July 12, 2011 | 40 | 90 | Software | Display/Image | Software update | Software Update
Z-2782-2011 | 57012 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Site Rite Vision Ultrasound System, REF 9770032, |  |  | June 21, 2011 | 2011 | Bard Access Systems | Purchasing erroneously ordered the incorrect fan using the distributor's part number instead of using the specified manufacturer part number. | Process control | Bard Access Systems, Inc sent out an "Urgent" Medical Device Recall (User Level)" dated June 2011 to all affected customersThe letter included: product name, description of problem and requested the return of units bearing "B50" on the fan. Bard Access Systems, Inc. will replace the incorrect fan with the correct fan. For additional information call 1-724-779-3000. |  | 86 | 86 units | Nationwide (USA) distribution including the states of AL, AR, AZ, CA, CO, FL, IA, IN, MD, MI, MO, MS, NC, NJ, NM, NY, OH, PA, RI, SC, TX, UT, VA, VT, and WA. | Terminated | July 11, 2012 | 386 | 86 | Not_Computer
Z-2788-2011 | 59006 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; |  |  | May 20, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Eclipse Monte Carlo (eMC) algorithm where not-equidistant data points in the open beam profile used during configuration may lead to inaccurate results. | Software design | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail dated May 20, 2011, to all affected customers. ***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product from use and advise users to manually initiate the configuration process for resampling to be applied when a previous calculation model is copied. Customers were asked to sign and return the notice to confirm receipt. For questions call USA and Canada 1-888-827-4265For Europe call +41 41 749 8844 |  | 10356 | 10,356 | Worldwide Distribution - USA (nationwide) | Terminated | July 13, 2012 | 420 | 10356 | Software | Output/Calculation | Remove | Remove or Replace
Z-2780-2011 | 58969 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; Release 6.5 through 10 |  |  | May 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, user editing of the MLC Leaf positions may not automatically invalidate calculated dose distributions. | Software design | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter included the description of the product, problem and recommendations and includes the release of a service pack in the future. For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265) |  | 10664 | 10,097***12-12-12 AMENDED TO: 10664 units*** | Worldwide distribution: Nationwide (USA) and the country of Canada. | Terminated | August 15, 2014 | 1190 | 10664 | Software
Z-2787-2011 | 58897 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens YSIO - solid state x-ray imager |  |  | April 28, 2011 | 2011 | Siemens Medical Solutions USA, Inc | There is a potential hazard to staff or patients when lowering the table height on the Ysio system. If staff or patients lower extremities are placed underneath the table and the user lowers the table top, it is possible that parts of the table may collide with extremities and cause injuries. | Other | Siemens sent a Safety Advisory Notice to all affected customers. The Safety Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to place the Safety Notice in the System's instruction for use. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. Customers were instructed to complete the Acknowledgement of Receipt Form and return to the firm. |  | 141 | 141 | Nationwide Distribution including Washington, DC and the following States: AR, CA, CO, FL, ID, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Terminated | April 10, 2014 | 1078 | 141 | Not_Computer
Z-2771-2011 | 55164 | No | LIG | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxC 880i, 860i,680i and 660i Synchron Access Clinical Systems |  |  | May 07, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of inaccurate Access Intrinsic Factor Ab test results displayed on, or printed from, the UniCel Dxl portion of an integrated workstation listed above. The problem is limited to numerical Intrinsic Factor Ab results produced on integrated systems running UniCel Dxl system software version 4.0, 4.1, or 4.2. The DxC portion of your inte | Software design | Beckman Coulter, Inc. sent out an "Urgent: Product Corrective Action" letter dated May 19, 2009 to all affected customers.The letter explained the reason for recall and provided recommendations including: -System software currently under development will correct this problem, and will be made available soon. Until then, please refer to UniCel DxC reports, user interface information, and LIS transmissions, which contain accurate Access Intrinsic Factor Ab test results. -Refer to the UniCel DxC system documentation for information on displaying and printing results from the DxC portion of an integrated workstation, or from your laboratory information system (LIS) -Past results for these assays only need to be reviewed if they were reported from the Dxl portion of an integrated workstation -Complete and return the enclosed response formFor additional information please contact the Customer Technical Support at 1-800-854-3633, or your local Beckman Coulter representative. |  | 5 | 5 units | Worldwide Distribution: (USA) Nationwide including KY; and the country of Canada. | Terminated | May 07, 2012 | 1096 | 5 | Software | Output/Calculation | Software update | Software Update
Z-2783-2011 | 59115 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Bard PowerLine 6 Fr. Dual-Lumen Polyurethane Catheter with SureCuff Ingrowth Cuff, REF 0700610, Sterile, Contents...1 Each - PTFE Introducer and Dilator, 6.5 FR. x 10 cm..., Assembled in Mexico. |  |  | June 20, 2011 | 2011 | Bard Access Systems | Kit label indicates the wrong dimension for an enclosed introducer. | Labeling design | Bard Access Systems, Inc. notified customers by letter dated June 16, 2011. The letter notified customers of the mislabeling issue and offered a replacement product if desired. For additional information contact Bard Access Systems at 1-800-443-3393. |  | 6485 | 6,485 units | Worldwide distribution - USA (nationwide) including the states of: HI, KY, TX, VT and the country of Canada. | Terminated | March 01, 2012 | 255 | 6485 | Not_Computer
Z-2775-2011 | 58720 | No | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter - Manual Inhibin A ELISA Reagent Kit; Part Number DSL-10-28100t-4 |  |  | March 31, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because patient results are falsely elevated by approximately 30% because the reference standard of the Manual Inhibin A ELISA assay has shifted from the assays original calibration.Erroneously elevated results may lead to unnecessary ultrasound testing to confirmgestational age, patient psychological stress until diagnostic testing is completed, and/ormedical decis | Process design | Beckman Coulter sent an "Urgent Product Correction" letter dated March 31, 2011 to all affected customers.The letter included an explanation of the problem and recommendations including; discontinue use of the kit and a Response Form to fill out and return.For additional information contact the Beckman Coulter Customer Support Center at 1-800-854-3633 or your local Beckman Coulterrepresentative. |  | 5521 | 5521 units total for all products | Worldwide distribution: (USA) Nationwide and the countries of Armenia, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Jordan, Republic of Korea, Lebanon, Malaysia, Mexico, Morocco, Netherlands, Peru, Portugal, Romania, Russian Federation, Serbia, Slovenia, Spain, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | July 11, 2012 | 468 | 5521 | Not_Computer
Z-2786-2011 | 58885 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Systems Symphony, Sonata and Symphony a TIM System |  |  | March 31, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm has discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. | Nonconforming Material/Component | Siemens initated their recall of the device on March 31, 2001, by sending a Customer Safety Advisory Notice/Letter to affected consignees. The letter identified the product, the problem, and the action to be taken by the customer.Customers were notified that detectors will be installed in appropriate locations on the system, which will switch off the system's power electronics in the event of a malfunction in the power supply of the gradient coils, preventing damage to the system and its environment. |  | 440 | 440 | Nationwide Distribution including AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and Puerto Rico. | Terminated | January 08, 2014 | 1014 | 440 | Not_Computer
Z-2778-2011 | 58970 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac iX, 23EX, CX, DHXC, Novalis Tx, Trilogy and TrueBeam Linear Accelerators Drive Stands |  |  | May 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Some drive stand bolts may fail to meet mechanical specifications. | Nonconforming Material/Component | Varian Medical Systems, Inc sent a "Urgent Medical Device Correction Urgent Field Safety Notice" letter dated May 13, 2011 to all affected customers.The letter included a description of product, problem and recommendations. This includes performing daily QA to detect isocenter shift and the use of independent patient position system. Customers are advised to discontinue its use if a shift is detected and to contact Varian Medical Systems for a replacement stand bolts and adjustment if necessary.For additional information please contact the Varian Oncology Help Desk at 888-VARIAN5 (888- 827-4265) |  | 36 | 36 | Worldwide distribution: (USA) Nationwide including the states of PA. NJ, NY, MA, VA, TX, KS, FL, OH, SC, MN, CA, TN and NV; and the countries of Canada, Colombia, France, Italy, Japan, Netherlands and United Kingdom. | Terminated | July 13, 2012 | 427 | 36 | Not_Computer
Z-2770-2011 | 59003 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, REF 5515, 7F, Sterile EO, Rx ONLY, 1 x Catheter and 5 x Catheter, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | June 09, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions, Inc investigation of a recent Device Experience Report has made them aware of a potential problem with their 7F Langston dual lumen pigtail catheter (Model 5515) with lot number 548264.Recently it was reported that pouches containing Langston catheters (Model 5515) were missing a seal,posing a risk of product contamination. It is possible that a compromised sterile barrier | Packaging process control | Vascular Solutions, Inc. sent an "URGENT MEDICAL DEVICE FIELD ACTION- LOT SPECIFIC" letter dated June 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructed customers to check their inventory for the affected product and provided information on how to examine the product. Customers are to complete and return a Field Action Customer Inventory Form via fax at 763-656-4250 or e-mail at www.vascularsolutions.com. VSI Customer Service will contact customers to set up return shipment of affected product. All affected product will be replaced upon receipt of the shipped product.Customers should contact their local Sales Representative for questions concerning this recall. |  | 23 | 23 (5 pack) boxes | Worldwide Distribution--USA (nationwide) including the states of CA, FL, IL, MA, MI, MO, NM, NC, OK, and WA and the countries of Austria and Canada. | Terminated | April 05, 2012 | 301 | 23 | Not_Computer
Z-2777-2011 | 59135 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Steerable Coronary Sinus Catheter, Model SCS01, 9 Fr x 20.4", Sterile R, Edwards Lifesciences Port Access Systems. |  |  | June 15, 2011 | 2011 | Edwards Lifesciences, LLC | Possibility of separation of the introducer sheath from the hub on Coronary Sinus Catheters. | Employee error | Edwards Lifesciences, LLC sent an Urgent Product Recall - Action Required letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusotmers were instructed to check their inventory for any affected product and return all unused product from their stock to Edwards at:Edwards Lifesciences12060 Lone Peak PkwyDraper, UT 84020Attention: Suzanne Carpenter, RGA #Customers were instructed to contact Customer Service to obtain a Returned Goods Authorization number and replacement product . Once customers verified their inventory, they were to complete the attached form and fax it to Edwards at (949) 250-3489. For any questions customers were instructed to call Edwards Customer Service at (800) 424-3278. |  | 97 | 97 units | Nationwide Distribution. | Terminated | June 01, 2012 | 352 | 2895 | Not_Computer
Z-2789-2011 | 59055 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR/PRIMUS Linear Accelerator systems with RTT v2.x AND an ARTISTE Linear Accelerator/RT Therapist v4.1 system with the LANTIS OIS. |  |  | May 25, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential safety issue of possible data loss for the transfer of the same patients treatment data between the LANTIS OIS and the RTT v2.x software and the RT Therapist version 4.1 system | Software design | Siemens Customers Service group began providing Update Instructions and installing the software update on affected machines beginning May 25, 2011. A customer information letter was also provided. |  | 82 | 45 customers with affected configuration in US, 37 customers ROW. | Worldwide Distribution - USA including distribution to LA, CA, PA, MA, NE, NJ, DE, WI, NC, NY, WI, MD, FL, KY and to Poland, Netherlands, Ireland, G=Belgium, Germany, New Zealand, P.R. China, Saudi Arabia, Australia, Malaysia, Egypt, Canada, and Spain. | Terminated | March 01, 2012 | 281 | 82 | Software | Device Operation | Software update | Software Update
Z-2779-2011 | 59100 | No | KBW | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Labeled in part:"***BIOMET MICROFIXATION***BIOMET MICROFIXATION***1520 TRADEPORT DRIVE***JACKSONVILLE, FL 33218 USA***www.biometfixation.com***2.0 MM X 7.0 MM TITANIUM***#01-6577***LOT XXXXXX***QTY 00001***X-DRIVE FOSSA SCREW 5/PKG***2.0 MM X 7.0 MM***". |  |  | June 25, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling part number 01-6577, 2.0x7 mm Fossa X-drive screws, 5 pack, lot # 233270. This product is being recalled due to the possibility that the pack may have contained another part, 01-6581 2.0x11 mm Fossa x-drive screws, 5 pack. | Mixed-up of materials/components | Biomet Microfixation, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 25, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to complete and return an Inventory Reconciliation form via fax to 904-741-9425. The form requests that customers indicate whether they want to keep the product or return for credit.Customers can contact the Customer Relations Specialist at 1-800-874-7711 or 904-741-4400 ext. 468 for questions regarding this recall. |  | 25 | 25 units | Product was distributed to one distributor in Argentina. | Terminated | November 09, 2011 | 502 | 25 | Not_Computer
Z-2794-2011 | 59057 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker 4.0 mm. Angled Aggressive 12 Degree Blade, Precision Cutters, Stryker Instruments, Kalamazoo, MI 49001 USA, (269) 323-7700 (800) 253-3210, Sterile, Single Use, Rx ONLY |  |  | June 13, 2011 | 2011 | Stryker Instruments Division of Stryker Corporation | Recalled device was packaged with the incorrect angled blade. The labeling indicates a 12 degree angle but the product has a 40 degree angle. | Labeling mix-ups | Stryker Instruments Division of Stryker CorporationOn 6/13/11 sent an 'URGENT MEDICAL DEVICE RECALL NOTIIFICATION" letter to all affected customers that included the device name, lot number, dates of distribution, reason for the recall, risks associated with use of the device, a photo of the device, instructions to quarantine and return the devices, and a form to fax back to the firm to confirm receipt and adherence to the recall letter. |  | 95 | 95 Units | Worldwide Distribution - USA including CA, LA,, MI, NJ, OH, OR, AR, IN, and Internationallly to Australia, Duisburg, Newbury and Hong Kong. | Terminated | August 03, 2011 | 51 | 95 | Not_Computer
Z-2790-2011 | 59035 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCI, Humidifier Adaptor, Rx only, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | June 06, 2011 | 2011 | Teleflex Medical | The packaging may not have been properly sealed, and therefore the sterility of the product cannot be guaranteed. | Packaging | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 6, 2011 to all affected customers. The letter described the product, problem, and the action to be taken.The letter instructs customers to discontinue use and quarantine the affected product. Customers must complete and return a Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service. A representative will contact customers with instructions on how to return the affected product. Customers with no affected product are also instructed to complete and return the form to verify receipt of the notification.Contact Customer Service at 1-866-246-6990 if you have questions regarding this recall. and a record of returns. A third letter was addressed to the Risk Manager with the above instructions. |  | 218880 | 228, 960 eaches | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, KY, MI, MN, MO, MS, NM, NE, NY, OH, OK, PA, TX, VA, WI and WM and the countries of Japan, Philippines, and Thailand. | Terminated | May 01, 2014 | 1060 | 218880 | Not_Computer
Z-2792-2011 | 58979 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Abbott Vascular COPILOT Kit, COPILOT Bleedback Control Valve |  |  | January 07, 2011 | 2011 | Abbott Vascular | Five lots of Guide Catheters and Accessory products were processed through Ethylene Oxide (ETO) sterilization with additional plastic material covering the pallet, which is not consistent with validated packaging configuration for ETO sterilization. | Process control | Abbott Vascular sent an Urgent Device Recall letter, dated January 7, 2011, noting the products being recalled and the reasons for recall. The firm also provided the following reconciliation instructions to its territory manager:-Help customer identify aflected product.- CaIl Abbott Vascular Customer Service (800) 221-9902 to receive RGA number. Record RGA number below.- Fax this completed form to (951) 914-3826.- Return a copy of this completed form with the returned product |  | 389 | 389 units total | Nationwide distribution, and Ireland. | Terminated | June 26, 2012 | 536 | 389 | Not_Computer
Z-2795-2011 | 58824 | No | MSB | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | XVIVO Perfusion AB (Part of Vitrolife), THAM Solution, 3.3 mmol/ml, 50 mL, REF 99019, Sterile, Used to increase the pH of the Perfadex used for perfusion of donor lungs. |  |  | May 12, 2011 | 2011 | Vitrolife, Inc. | Precipitation observed in some bottles of Tham buffer used for adjusting pH of Perfadex prior to perfusion of donor lungs. | Process control | XVIVO Perfusion, part of the Vitrolife Group, sent out a "Partial recall of Perfadex with THAM" letter on May 13, 2011 to all affected customers.Customers were instructed to to quarantine all Perfadex with THAM. Replacement THAM was to be shipped by Vitrilife within a few days and the exchanged THAM was to immediately sent to Vitrolife, Englewood, CO.For additional Information please call (303) 762-1933 |  | 1457 | 1457 bottles | Worldwide distribution including USA and Europe. | Terminated | July 15, 2011 | 64 | 1457 | Not_Computer
Z-2799-2011 | 58768 | No | NSQ |  |  | Dental | Enforcement Discretion | "***REF IWTH56***Certain EP HEALING ABUTMENT 5mm(D) X 5mm(P) X 6mm(H).*** Healing abutment*** STERILE.***Rx Only.*** Made in USA*** Implant Innovations, Inc. Global Headquarters 4555 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 342-5454 www.3i.online.com. Outside US 561-776-6700***Implant Innovations Siberica S.L., WTC Almeda Park, Ed. 1 Planta 1a Pl. De la Pau s/n 08940-Cornella de Llobregat (Barcelona) Spain. Phone: +34 934 705500. Fax +34 933 717 849. Lot Number: 602691.***" |  |  | February 23, 2010 | 2010 | Biomet 3i, LLC | Biomet 3i of Palm Beach Gardens, FL is recalling their Certain EP Healing Abutment (IWTH56). The Healing Abutment dimensions did not match outside label and laser etching on part. The product was actually 8mm but the package indicated 6mm. | Error in labeling | The firm, BIOMET 3i, initially contacted each US customer by telephone using a prepared phone script. Following telephone contact, the firm followed up each call with a faxed "Urgent: Medical Device Recall" letter dated March 1, 2010. International customers were notified first by email then by fax. The letter and email described the product, problem and actions to be taken. The customers were instructed to return the healing abutment(s) for replacement if they have not already opened and used the product; respond with the attached response fax and return the product(s) to the: Regulatory Affairs Department, BIOMET 3i, 4555 Riverside Drive, Palm Beach Gardens, FL 33410, Incident No. IC95863. BIOMET will send a replacement healing abutment(s) once the returned product is received.If you have any questions or concerns, please contact the Recall Coordinator at 561-776-6906. |  | 49 | 49 | Worldwide distribution: USA (nationwide) and countries of: Australia, Europe and Japan. | Terminated | November 09, 2011 | 624 | 49 | Not_Computer
Z-2798-2011 | 58863 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Terumo(tm) Hypodermic Needle, sterile, individually packaged into a "blister-pack" flexible plastic bottom-web thermo-formed to create a pocket, heat-sealed with paper top-web,1,000 needles per case. |  |  | June 08, 2010 | 2010 | Terumo Medical Corporation | Hypodermic needle packaging defect could compromise sterility. Affected lot numbers are: MC1725, MC2025, MC2025 | Packaging | The firm, Terumo Medical Corporation sent an "URGENT PRODUCT BULLETIN: RECALL OF TERUMO PRODUCT" letter dated June 8, 2010. Consignees were contacted via telephone on June 7 thru 9, 2010 to confirm receipt of the Product Recall Bulletin. The letter describes the product, problem and the actions to be taken. The letter instructed customers to immediately discontinue use and quarantine any available stock of the listed product codes. The letter stated that distributors must not distribute any of the listed product codes, quarantine any available stock and send the recall bulletin immediately to each customer that received the recalled products to inform them of this issue. Additionally, Terumo Medical Corporation sent a "Terumo Customer Instructions for Returning Recalled Product and Customer Confirmation of Product Recall Notification" letters dated June 8, 2010 to all affected facilities. Customers were instructed to fill-out the requested information and return to Terumo Medical Corporation (TMC) Product Recall Coordinator. Consignees were contacted via telephone on June 8 and 9, 2010 to confirm receipt of the recall information and compliance with the requested action. For questions or need for further information, call (800) 283-7866 and ask to speak with the Product Recall Coordinator. |  | 453000 | 453,000 needles | Nationwide Distribution (USA) -- AZ, CA, CO, FL, GA, IL, KS, KY, MA, MO, NC, NJ, NY, OH, PA, TN, TX, WA, UT, and VA | Terminated | July 14, 2011 | 401 | 453000 | Not_Computer
Z-2796-2011 | 58806 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague 3 and 3 CX Infusion Pumps, Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2 M8153 and 2M8163. |  |  | May 04, 2011 | 2011 | Baxter Healthcare Corp. | The Air In Line (AIL) test was not properly performed by the service technician at the Service Center. | Employee error | Baxter telephoned all seven affected customers beginning May 4, 2011, informing them that that a post repair test was not performed correctly on their Colleague pumps during the last service. As a result, the device Air Sensor may not function as designed and could lead to undetected air in the IV line. Baxter wants to ensure that the devices are properly tested, safe and effective. Baxter requested that they locate the pumps and remove them from use as soon as possible. Baxter will coordinate with the affected customers, the return of the device for the necessary testing to be completed at their facility and return of the device back to the customer at no cost.Please call ( 800 ) 437-5176 for a complaint that includes any indication of the failure of the product to meet a customer's expectations for quality or to meet performance specifications. |  | 12 | 12 pumps | Nationwide Distribution including the states of NY, Oh, KY,SC, PA, FL and Puerto Rico. | Terminated | January 10, 2017 | 2078 | 12 | Not_Computer
Z-2797-2011 | 58916 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | Sysmex Stromatolyser-IM; a lytic reagent for use with the XE-2100, XE-2100L and XE-5000 Series Sysmex Automated Hematology Analyzers; the reagent solution is supplied in a 10-liter cubetainer (cardboard cube with a 10-liter plastic liner); Sysmex America, Inc., Mundelein, IL 60060; catalog number SIM-220A. |  |  | June 23, 2009 | 2009 | Sysmex America, Inc. | The affected lots of Stromatolyser-IM lysing reagent was found to be contaminated with Pseudomonas bacteria, which resulted in rising IMI channel background counts, increased false "Platelet Clump" flags and a strong sulfur-like odor when used on the XE-2100, XE-2100L and XE-5000 analyzers. | Environmental control | Sysmex America, Inc. sent a "MEDICAL DEVICE RECALL" letter dated June 2009 to all affected customers. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to discontinue use of the affected SIM lots, and to remove affected lots from their analyzers as soon as possible to avoid long-term performance problems on the analyzer. A Stromatolyser -IM (220-A) Voluntary Recall form was attached for customers to complete and return via fax at 800-218-2711. The form will be used to place orders to replace the affected product. Any questions concerning the information contained in the recall letter are directed to the Sysmex Technical Assistance Center at 1-888-879-7639.The firm sent a second notice to non-responding accounts on September 14, 2009. A copy of the June 26, 2009 recall letter and the fax back form were included with the letter. The customers were informed that Sysmex had not received their completed fax back form and were requested to respond to confirm the notification was received. |  | 4005 | 4,005 cubetainers | Worldwide Distribution--USA (nationwide) and the countries of Brazil, Canada and Honduras. | Terminated | July 13, 2011 | 750 | 4005 | Not_Computer
Z-2801-2011 | 59207 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | iSymm Intraocular Lens (Model FC-60AD) |  |  | June 17, 2011 | 2011 | Hoya Surgical Optics, Inc. | The recall was initiated because Hoya Surgical Optics has confirmed that the iSymm Intraocular Lens (Model FC-60AD) and iSert Intraocular Lens (Model PC-60AD) have been sold and marketed with out an FDA approval. | PMA | Hoya Surgical Optics Inc. sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated June 22, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken.The letter instructs customers to quarantine affected product and return to HSO at their earliest convenience. Shipping instructions are included in the letter. Customers are to complete an Effectiveness Check Questionnaire and fax it to 1-909-680-3986.Questions regarding the recall process are directed to Hoya Customer Service at 1-866-750-5870 or e-mail at customercare@hoyasurgopt.com. |  | 61617 | 61,617 units total | Nationwide Distribution | Terminated | November 29, 2012 | 531 | 61617 | Not_Computer
Z-2800-2011 | 59018 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Model 134, WarmAir Convective Warming Device, 115V, Part #86180, Device Listing #D111865, The WarmAir 134 is used in conjunction with a single-patient use blanket to warm a patient's body temperature. The device draws in room air through a filter by means of an electric blower. It heats the air to the desired temperature by controlling the air temperature entering a flexible connecting hose, and conveys it to the disposable blanket (i.e. FilteredFlow Blankets, Warming Tube) in contact with the patient, in order to warm the patient. |  |  | May 19, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | A component defect. The recalling firm determined that one of the connectors attached to the firm's WarmAir 134 device's Printed Circuit Board (PCB) can overheat. Should this event occur, the excess heat build up within the device may result in damage to the device's external case and/or melting of some of the internal components of the unit. | Nonconforming Material/Component | CSZ Cincinnati Sub-Zero sent an "URGENT MEDICAL DEVICE RECALL FIELD ACTION" notice to all affected customers. The letter identifies the product, problem, and actions to be taken.CSZ instructs customers to complete and return an enclosed recall response form and to state whether or not their facility is in possession of the device. The completed form should be returned via fax at 513-772-9119 or via e-mail to WA134Recall2011@cszinc.com. Additionally, customers are instructed to discontinue use of the device and return to their Biomedical Engineering and/or Service Department for device exchange and return.Questions regarding this notice are directed to the WarmAir 134 field action (recall) administrator at 1-800-989-7373. |  | 463 | 463 units were distributed | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, LA, MA, MI, MO, NC, NJ, NM, NV, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and the country of Canada. | Terminated | February 26, 2013 | 649 | 463 | Hardware | Physical Safety Hazards | Return | Remove or Replace
Z-2725-2011 | 56016 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IQecg IQdevices SDK, Part Numbers: 4-100-1403, Software Version 8.3.2 |  |  | May 18, 2010 | 2010 | Brentwood Medical Technology Corp | The recall was initiated because when a user of the IQmanager, IQdevices SDK, and IQiC (Centricity) software versions listed above accesses the QT dialogue box to either accept or edit the QT value, the QTc value is recalculated using a Hodges QTc equation, regardless of patient's age. The software does this without notifying the user of the change from use of the Bazett equation, which is utiliz | Software design | The firm, Midmark, sent an "URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter on May 19, 2010 with an attached Corrective Return Material Authorization (RMA) form via US Postal Service Certified Mail to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their IQecg, IQmanager, or lQiC Centricity software to determine the software version currently in use. If it is a listed product, it is subject to this field corrective action. A sample software label was attached to aid the consignees/customers in locating and identifying the software CD. For customers who have already installed the software, the version number can be identified in the "About" page.Consignees/customers were instructed to complete the enclosed Field RMA form and return it with the affected software to:Quality Assurance DepartmentMidmark Diagnostics Group1125 West 190 S t.Gardena, CA 90248In addition, if they further distributed this product, identify customers and notify them at once of this product field corrective action.Replacement software was to be shipped to the consignees/customers upon receipt of the returned software CD. If they require a UPS return shipping label or an on-line download of replacement software, contact Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or 5108. If you have any questions, call Midmark Diagnostics Group Support Services at 800-624-8950 extension 5145 or QA/RA Manager, Midmark Diagnostics Group 310-516-5116. |  | 8 | 8 units | Worldwide distribution: USA (Nationwide) and countries of: Algeria, Australia, Canada, France and Saudi Arabia. | Terminated | August 22, 2011 | 461 | 1621 | Computer
Z-2833-2011 | 59096 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific, WALLGRAFTÂżÂż OVER-THE-WIRE, Tracheobronchial Endoprosthesis, Made in IRELAND, Ballybrit Business Park, Galway, IRELAND, STERILE, EO. |  |  | May 25, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating an important field correction related to WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis products. This letter is intended to clarify and re-emphasize that U.S. WALLSTENT", WALLSTENT" RP, and WALLGRAFTÂż Tracheobronchial Endoprosthesis stents are NOT currently cleared by the FDA for use in benign tracheobronchial strictures. | Labeling False and Misleading | The firm, Boston Scientific, sent an "Urgent Medical Device Field Correction" letter dated May 25, 2011 to its customers. The letter described the product, the problem and actions to be taken. The letter states that no product is being recalled and asks consignees not to return the product. The customers were instructed to read the enclosed information and post it in a way to ensure it is easily accessible to all users of this device; complete and return the Reply Verification Tracking Form via fax to: CV Field Action Center, 1-866-213-1806, and if they are a distributors to forward this notification to their customers. If you have any questions or would like assistance with this field correction, please contact your local Sales Representative and/or call 763-494-1133. |  | 23218 | 23,218 FOR ALL PRODUCTS | Worldwide distribution. | Terminated | August 10, 2012 | 443 | 23218 | Not_Computer
Z-2835-2011 | 59078 | [6 empty] | OneTouch Delica Lancing Device; |  |  | May 27, 2011 | 2011 | Lifescan Inc | Individuals attempting to remove a lancet from the OneTouch Delica lancing device without following instructions for use may experience an accidental needle stick. | Other | Lifescan sent an Urgent Medical Device Correction letter dated May 27, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told that no action on their part is needed and were asked not to return any of the affected product. Customers were instructed to continue to distribute the product. Healthcare professionals and pharmacies were asked to share the information with their patients who use the OneTouch Delica Lancing Device. Patients with questions should contact Customer Care Line at 1-800-574-6138. For questions regarding this recall call 1-800-600-7226. |  | 1032515 | 1,032,515 in US | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 05, 2012 | 223 | 1032515 | Not_Computer
Z-2818-2011 | 59104 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Syneture Sofsilk 3/0 BK 12X18" PRE-CUT WAX-RACE |  |  | June 03, 2011 | 2011 | Covidien LP | Potential for the sterile barrier to be breached compromising the sterility of the product | Equipment maintenance | Covidien sent an URGENT MEDICAL DEVICE RECALL letter dated June 3, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify and quarantine all affected inventory. Distributors who further distributed the affected product were instructed to forward the letter to their customers who received the affected product. Customers were instructed to contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722- 8772, option 1, to obtain a Return Material Authorization (RMA) prior to returning the affected product. Customers were also asked to complete the attached product recall form in its entirety and fax the form to 203-822-6009 and include a copy with products being returned. If customers did not have any of the affected product they were to simply return the recall form via fax or email indicating they had zero units. For any questions regarding this recall customers were to contact their Covidien Representative or Covidien Quality Assurance at 203-492-5232. |  | 960 | 960 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, Denmark, France, Germany, Italy, Japan, Korea, and Singapore | Terminated | February 10, 2014 | 983 | 93600 | Not_Computer
Z-2834-2011 | 59099 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | LuxStar Dental Light have the following Agency model or Catalog Model numbers as ML1000, OL1000, CL1000, WL1000, CB1000, TL1000, and UL1000. |  |  | May 27, 2011 | 2011 | Dental Equipment LLC | Lens heat shield (birdcage) of the Luxstar Dental Light could fall and injure a patient. | Device Design | Dental Equipment, LLC sent an "URGENT MEDICAL DEVICE RECALL" notification letter via USPS 1st class to US consignees and via Federal Express or Fax to foreign consignees on June 24, 2011. The letter informed customers about the product, problem and tether upgrade corrective action. A Field Correction Acknowledgement Form and the Field Correction Completion Form was attached for customers to complete and return via the fax number 503-537-3617. The letter states that customers are to review their records in order to identify where the affected devices are located and which are still in service. Customers can call the Customer Service at 800-304-5332 and order the self-install Upgrade Kits with the Part Number 30792 at no charge. Customers are instructed to complete the Rework Correction Form once all of the lights have been updated and fax or mail to Marcus, attention Bonnie Dundas, at (503) 537-3617. |  | 13226 | 13,226 units distributed worldwide | Worldwide Distribution -- USA (nationwide) including states of: CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, and Puerto Rico and the countries of: Australia, Belgium, Canada, France, Germany, Ireland, Italy, Mexico, Norway, Poland, Russia, United Kingdom, Vietnam, and Trinidad. | Terminated | October 24, 2012 | 516 | 13226 | Not_Computer
Z-2831-2011 | 59184 | sheath, corrugated rubber, for non-indwelling catheter, non-sterile | NNY | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Rusch, Sure Seal Golden Drain, One Piece Urinary Incontinence Device, Teleflex Medical. |  |  | June 27, 2011 | 2011 | Teleflex Medical | The Kit contains Skin-Prep protective wipes that were manufactured by the Triad Group and are being recalled by Smith &amp; Nephew, due to possible bacterial contamination. | Other | Teleflex Medical sent an "Urgent Medical Device Recall Notification" dated June 27, 2011 to all affected customers.The firm provided recommendations including all affected product be discontinued and quarantined. A Recall Acknowledgement Form was provided to be completed and returned. If the product was further distributed, they were urged to contact their customers and communicate the recall.For additional information please contact Customer Service at 866-246-6990 or your local sales representative. |  | 6880 | 6880 units (total for both products) | Nationwide Distribution (USA) including the states of FL, GA, OH and SC. | Terminated | March 20, 2012 | 267 | 6880 | Not_Computer
Z-2837-2011 | 59263 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Non-DEHP I.V. Catheter Extension Set, 6.6 (17 cm), Volume 0.8 mL, Injection Site, Male Luer Lock Adapter; an Rx sterile, single use nonpyrogenic fluid pathway; 200 sets per case; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 2N3375 |  |  | July 06, 2011 | 2011 | Baxter Healthcare Corp. | Some of the individual extension set packages were improperly sealed, compromising the sterility of the product. | Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated July 6, 2011 via first class mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from their facility; call Baxter Healthcare's Center for Service at 1-888-229-0001 between 7:00 am and 6:00 pm Central Time, Monday through Friday to return affected product and receive credit; complete and return the attached CUSTOMER REPLY FORM to Baxter via fax at 1-847-270-5457 or scan and email to fca@baxter.com and if you are a dealer, wholesaler or distributor/reseller, notify your customers of this recall. If you have questions regarding the communication, call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 9800 | 9,800 sets | Worldwide distribution: USA (nationwide) including: AL, LA, MO, SD, TX and WV; and country of: Canada. | Terminated | April 18, 2012 | 287 | 9800 | Not_Computer
Z-2838-2011 | 58769 | No | NSP |  |  | Dental | Enforcement Discretion | "***3i Incise Custom-Made Label. QP157 Approved 07/09/10*** 3i inc se Custom-made Dental Framework***Biomet 3i 4555 Riverside Drive Palm Beach Gardens, FL 33410 800-342-5454 www.biomet3i.com Please consult instructions for use and recommendation***For Dental Lab use ONLY Non-sterile/For single Use***" |  |  | December 01, 2010 | 2010 | Biomet 3i, LLC | On December 1, 2010 Biomet 3i recalled their product Incise Zirconia Coping (CBZR0102). The Coping contains a manufacturing defect. Coping is oversized. All affected direct accounts were notified by letter on December 14, 2010. | Nonconforming Material/Component | Biomet 3i sent an "MEDICAL DEVICE RECALL" letter dated December 14, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken.The letter instructs customers to complete and fax the Product Recall Response form to 561-514-6316. In addition, a shipping label is enclosed with the letter for returning the affected product.Customers with questions and concerns should contact Biomet 3i recall coordinator at 561-776-6906. |  | 16 | 16 | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | July 18, 2011 | 229 | 16 | Not_Computer
Z-2840-2011 | 59155 | [6 empty] | ENCORE(R) UNDERGLOVE, Latex, Powder-Free, Smooth Surgical Gloves, REF 2018455, Size 5 1/2, Sterile, Manufactured for Ansell Healthcare, 1635 Industrial Road, Dothan, AL 36303, Made in Sri Lanka |  |  | June 10, 2011 | 2011 | Ansell Healthcare Products LLC | Gloves do not meet the tensile requirements of ASTM D 3577 reducing the shelf life of the product from 5 years to 2 years. | Other | Ansell Healthcare Products LLC sent an Urgent Voluntary Recall letter dated June 9, 2011, to all affected customers on June 10 and 13, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue use and distribution of any of the affected product and return them to the firm. If the customer distributed gloves further, customers were advised to forward a copy of the notification letter along with their tracking number to their customers. Customers were instructed to complete and return the attached Response Form to Ansell by fax at 1-800-722-8155. Returned product would be replaced or credited to the customer. |  | 9580 | 9580 pairs | Nationwide Distribution | Terminated | January 10, 2013 | 580 | 1176176 | Not_Computer
Z-2839-2011 | 57518 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos Cellex Procedural Kits. |  |  | December 02, 2010 | 2010 | Therakos Inc | Cellex Procedural kits may have minor tears in their Tyvek covers. | Employee error | Therakos, Inc. sent an "Important Product Notification" dated December 2, 2010 to all affected customers.The letter included description of problem, affected lot and recommendations. Customers were asked to dispose any product with affected lot and to contact them for for replacement, and to complete and return the attached Fax reply form.For additional information contact Therakos Customer Technical Services at 877-865-6850. |  | 110 | 110 units | Nationwide (USA) distribution including the state of MI, MO, WI, CA, OH, NH, MA, and OR. | Terminated | February 19, 2015 | 1540 | 110 | Not_Computer
Z-2849-2011 | 58980 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | "***REF: 922070 TMJ SYSTEM ARTHROSIMPLICITY KIT***STERILE PRODUCT***STERILE EO***DISTRIBUTED BY BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" |  |  | October 08, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling TMJ System OnPoint Scope Procedure Kit REF 24-3050; TMJ System ArthroSimplicity Kit REF 922070 &amp; 24-4050; TMJ System InnerVue Scope Procedure Kit REF 922090. The OnPoint Scope Kits and ArthroSimplicity Kits contain Chloraprep Swabsticks that do not contain any warnings regarding the use of the product on eyes, ears or mouth. | Pending | Biomet Microfixation (formerly W. Lorenz) sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter to the consignees/customers on October 7, 2010. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to fill out the attached Inventory Reconciliation sheet and fax to 904-741-9425 and if the consignees/customers further distributed this product, they should notify their customers and provide them with a copy of this safety alert (the letter must be given to the personnel responsible for receiving field action notices). If you have any questions, please contact the Customer Relations Specialist at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 82 | 82 | Worldwide distribution: USA (nationwide) including states of: AL, AR, CA, CO, CT, DC, CR FL,GA, HI, IA, IL, IN, KS, KY, LA, MA, MD,ME, MI MN, MO, MS, MT, NC, ND, NJ, NY, OH, O K, OR, PA, RI, SC, TN, TX, UT, VA and WA; and countries of: Argentina, Bahrain, Canada, Colombia, Costa Rica, and Finland. | Terminated | April 19, 2012 | 559 | 828 | Not_Computer
Z-2853-2011 | 59270 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Code: 2M8161. |  |  | July 05, 2011 | 2011 | Baxter Healthcare Corp. | While replacing a defective pump head mechanism and performing calibration, Baxter's Service center may have failed to apply GLYPTOL (a red insulating varnish) on a potentiometer that helps maintain pump accuracy for detecting an air bolus. | Process control | Baxter Healthcare Corporation telephoned its customers on July 1, 2011 and sent a follow up "URGENT PRODUCT RECALL" letter dated July 6, 2011, informing them that the Glyptol (a red insulating varnish) may not have been applied on the Air-In-Line (AIL) circuit board. The letter described the product, problem and actions to be taken. The customers were instructed to locate the affected pumps, remove from service and return to Baxter for inspection as soon as possible. If GLYPTOL is found to be missing, recalibration will be performed and GLYPTOL will be applied per requirement before returning the pumps back to the customer at no cost. The customer was instructed to complete and sign the attached customer reply form. The letter states that if the COLLEAGUE Pumps were distributed to other facilities, please forward communication to the pump owner as appropriate. For any questions contact The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time, Monday through Friday. |  | 2 | 2 units | Nationwide (USA) Distribution -- including the state of Pennsylvania. | Terminated | January 20, 2012 | 199 | 2 | Not_Computer
Z-2858-2011 | 59051 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Runthrough NS Tapered PTCA Guidewire, item code number TW*DP418FH |  |  | June 03, 2011 | 2011 | Terumo Medical Corporation | Baloon catheter guidewires may become uncoiled during angioplasty procedures. | Device Design | Terumo Medical Corporation initiated contacting consignees on June 3, 2011 via an in-person visit to hand deliver an "URGENT PRODUCT BULLETIN: TERUMO PRODUCT SAFETY ALERT" letter dated June 3, 2011. The letter emphasized to consignees the importance of following the current instructions and recommendations provided to minimize the potential for damage during use and removal of the recalled product. For further questions, please contact your Terumo Territory Manager or call (800) 283-7866 and ask to speak with the Runthrough NS Guidewire Product Manager, at extension 4129. |  | 70 | 70 units | Nationwide (USA) Distribution -- including the states of: CT, GA, and SC. | Terminated | July 27, 2011 | 54 | 70 | Not_Computer
Z-2803-2011 | 58996 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GEM PREMIER 4000 CARTRIDGE K+ REPORTING: |  |  | May 31, 2011 | 2011 | Instrumentation Laboratory Co. | K+ results on the GEM Premier 4000 are too low when compared to a reference analyzer, withbiases exceeding allowable error claim of aÂż 0.5 moUL | Nonconforming Material/Component | The firm, Instrumentation Labs, issued a Revised "Urgent Field Safety Notification" letter dated July 14, 2011, to GEM Premier 4000 customers. The letter describes the product, problem, updated issue description and actions to be taken. The customers were instructed to follow the instructions to Disable K+ Parameter ,provided in the letter, to eliminate the potential of erroneous K+ result reporting and to complete and return the MANDATORY RESPONSE TRACKING FORM immediately via fax at 781-861-4207 or email to ra-usa@ilwww.com. The letter also noted an Important Change to Customer Actions from June Urgent Field Safety Notification.Prior strategy :Instrumentation Laboratories (IL) notified U.S and Canadian customers via the Urgent Field Safety Notification mailed USPS between June 2 - 3,2011. International Affiliates / Dealers notified through e-mail on June 1, 2010. A response (tracking) form was included. A Worldwide Regulatory Notice was e-mailed on May 31.2011.to IL country affiliates and distributors, notifying them of the issue, providing a risk summary and the support documentation for the notification. The customer notification will be handled per their local regulatory requirements.If you have questions regarding this notification, please contact IL at 800-678-0710, press 2 at prompt. |  | 2369 | 2369 | Worldwide distribution: USA (nationwide) and countries of: CANADA, ARGENTINA, AUSTRALIA ,AUSTRIA, BELGIUM, BRAZIL, CHILE, CHINA, CROATIA (local name: Hrvatska), CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, MALAYSIA, MEXICO, NETHERLANDS, NORWAY, PANAMA, POLAND, PORTUGAL, PUERTO RICO, SAUDI ARABIA, SLOVENIA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VIET NAM. | Terminated | November 27, 2012 | 546 | 28616 | Not_Computer
Z-2859-2011 | 59103 | No | IZQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiomat Illumena Syringe, 150mL Linden Luer Syringe with Handi-Fil Straw (p/n 900103), in a sealed molded plastic tray. |  |  | May 20, 2011 | 2011 | Mallinckrodt Inc | Syringes packaged and labeled as the Illumena 150mL Linden Luer Syringe with Handi-Fil Straw, product 900103, contained the Illumena150mL Syringe with Handi-Fil Straw, product 900101. | Other | COVIDIEN sent an Urgent Device Recall letter dated May 20, 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use the attached business response form and report the quantity of lot number 0253304 currently in their stock. If the customer did not have any of the affected lot they were to indicate by checking the appropriate box on the form. When completed, customers were asked to fax to (314) 654-8206 or email to productmonitoringhazelwood@covidien.com. All product from lot number 0253304 was to be returned to Covidien. Once the business notification form is received a Return Goods Authorization for the quantity of product reported will be sent to the customer. Customers would receive credit for all return product from affected lot.For any questions regarding this recall customers were asked to contact the Product Monitoring department at 800-778-7898. For questions about credit for the product, call Customer Service at 888-744-1414, menu option 1, and then menu option 2. |  | 550000 | 11,000 units (200 cases/50 units per case) | Worldwide Distribution - USA including CA, FL, IL, IN, MI, MO, NY, OH, TN &amp; TX. and the country of Australia. | Terminated | February 19, 2013 | 641 | 550000 | Not_Computer
Z-2836-2011 | 59161 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Kimberly-Clark, Three Quarter Drape, Non sterile/Disposable, |  |  | June 06, 2011 | 2011 | Kimberly-Clark Corporation | Metal shavings were found within the folded portion of the surgical drape. | Process control | Kimberly-Clark Corporation sent an "Urgent: Voluntary Product Recall" letter dated June 6, 2011 to all affected customers and a subsequent letter dated June 22, 2011.The letter notified them of the recall and provided recommendations. They were instructed to evaluate their facilities unused inventory of the affected drapes to determine if any of the impacted product lots remain with their facility. If present, they were to quarantine and discontinue use. A Customer Product Recall Response Form was included to be completed and faxed to KC. They will be contacted by a representative regarding the return and replenishment of the product. If the product was further distributed, the consignee was advised to contact the end user, make them aware of the recall and to retrieve any of the affected product in inventory. For additional information please contact Kimberly-Clark Regulatory Affairs at 770-587-7400. |  | 223 | 223 cases (300 units per case) | Nationwide (USA) distribution, including the states of CA, FL, GA, IA, IL, MO, NC, NY, OH, TN, TX, VA and WI. | Terminated | October 14, 2011 | 130 | 223 | Not_Computer
Z-2855-2011 | 59291 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MatriStemÂż Plastic Surgery Matrix XS |  |  | July 07, 2011 | 2011 | ACell, Inc | Surgical mesh for implantation may have elevated endotoxin levels. | Nonconforming Material/Component | ACELL, Inc. sent an "Urgent: Medical Device Safety Alert" letter dated July 07, 2011 to all affected customers.The letter notified customers (physicians) of the possibility of elevated endotoxin levels in the product and affected lots. They were instructed to monitor their patients receiving the affected implants.For additional information please contact the Director of Regulatory Affairs &amp; Quality Assurance at (410) 953-8508. |  | 30 | 30 units | Nationwide (USA) Distribution including the states of CA, CT, FL, ID, IL, MD, MI, MS, NC, NJ, NV, NY, OH, OR, PA, TX and VA | Terminated | October 21, 2011 | 106 | 30 | Not_Computer
Z-2856-2011 | 59259 | No | LGB | Gonococcal antibody test (GAT). | Microbiology | Microbiology | PMA | PALL Medical LeukoGuardÂż LGB Arterial Line Leukocyte Reduction Filter with Bypass Loop, STERILE. CATALOG/MODEL/Reorder Number: LGB. CLASSIFICATION NAME: Filter, Blood, Cardiopulmonary Bypass Arterial Line. COMMON/USUAL NAME: Arterial Line Filter with Bypass Loop. -- Manufactured by: PALL NEWQUAY, CORNWALL, TR9 6TT, UK -- Device Listing # D000745. --- Designed to reduce the levels of circulating leukocytes and remove microemboli greater than 40 microns in size from the perfusate during extracorporeal circulation including: gas emboli, fat emboli and aggregates composed of platelets, red blood cells and other debris. |  |  | May 26, 2011 | 2011 | Pall Corporation | Potential for the housing of the arterial line Leukocyte Reduction Filter to separate before or during the priming procedure. | Process control | Pall Medical sent an "Important Customer Notification" letter dated May 26, 2011 to all affected customers.The letter included a description of the problem, affected lots and provided recommendations including the return of the product.For additional information, please contact the firm at (516) 484-0263. |  | 232 | 232 units | Nationwide (USA) Distribution including the states of CA, LA, MI, TX and WI. | Terminated | October 03, 2012 | 496 | 232 | Not_Computer
Z-2854-2011 | 59189 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VirtuoSaph Plus Endoscopic Vessel Harvesting System , REF VSP550, Sterile, Terumo Cardiovascular System |  |  | June 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems received four (4) complaints from a single customer about tunnel breathing/inadequate insufflation with the VirtuoSaph Plus Endoscopic Vessel Harvesting System during cases in a limited production launch situation. Teumo's investigation identified issues with the devices's capability to maintain pressure within the tunnel. | Device Design | Terumo Cardiovascular Systems contacted the one customer to this device on June 10, 2010 by telephone.There was one consignee. The firm contacted this one consignee by phone on 6/10/2010. There were 2 contacts at this one consignee. The recalling firm contacted both of the contacts on 6/10/2010. The consignee was asked to dispose of remaking inventory of devices. The consignee said had disposed of the product.For information on this recall call (734) 741-6173 |  | 35 | 35 units | Nationwide (USA) Distribution including the state of New York. | Terminated | July 27, 2011 | 47 | 35 | Not_Computer
Z-2158-2011 | 58419 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Healthcare DISCOVERY MODEL NM/CT 670 |  |  | January 26, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a missing laser aperture warning label (i.e. Laser Aperture - DoNot Stare Into Beam) affecting certain Discovery Systems. | Radiation Control for Health and Safety Act | GE Healthcare will bring affected systems into compliance. A GE Healthcare Service Representative will provide a product label correction to address this issue. This will be carried out via a Field Modification Instruction (FMVI 408311 planned to be released in June 2011 with an estimated completion date of August 2011.For questions call 262-513-4122. |  | 4 | 4 | Nationwide Distribution | Terminated | August 31, 2012 | 583 | 4 | Not_Computer
Z-2865-2011 | 59201 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Prime 5th Wheel Stryker Medical Stretchers Model 1105 |  |  | July 19, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | The firm has determined that the brake rod assemblies of various stretchers that were built with a 5th wheel were manufactured with roll pins that did not meet the materials specifications. Some roll pins may have fractured during the assembly process and could fall out of the brake rod assemblies, rendering the brake or steer systems inoperable from certain pedals or causing a reduction in overa | Nonconforming Material/Component | Stryker Medical sent an "Urgent Medical Device Correction" dated July 19 ,2011 to all affected customers.The letter included; the name and model numbers of the stretchers, the reason for the correction and dates of manufacture. It contained the instructions to check the braking system and if found to be inoperable, to remove it from service, and if operable to check prior to each use. Also, information regarding contacting the recipients of any loaned or sold stretchers to relay the information in the letter and Medwatch reporting information. A self addressed postcard to confirm receipt of the letter was also included.For additional information call (269) 389-6604. |  | 80 | 80 | Worldwide Distribution: Nationwide (USA) including the states of NY, MI, LA and CA ,and the countries of Canada, China, Spain, Japan, Korea, United Kingdom, South Africa. | Terminated | August 29, 2013 | 772 | 112 | Not_Computer
Z-2873-2011 | 59002 | [6 empty] | "***REF: 91-6704***THINFLAP SYSTEM CROSS-DRIVE, SELF-DRILLING SCREW 1.5X4MM (.059X.159") (WXL)***TITANIUM***NON-STERILE PRODUCT***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA*** www.biometmicrofixation.com***QTY 00001***". |  |  | May 17, 2010 | 2010 | Biomet Microfixation, Inc. | On May 17, 2010, Biomet Microfixation issued a Safety Alert on their Thinflap System Cross Drive, Self Drilling Screw, 1.5x4mm, Model #: 91-6704, Lot#: 173770. The product is labeled incorrectly. Some packages of 91-6704 lot 173770 are labeled as cross drive screws, but contain center drive screws. All affected customers were notified of the recall on May 17, 2010. | Labeling mix-ups | Biomet Microfixation notified distributors and customers by letter, email and telephone calls. Distributors and customers were instructed to fill out and return the Inventory Reconciliation sheet, via Fax to 904-741-9425. Distributors were also instructed to notify their customers if the product was distributed further.For questions regarding this recall customers were instructed to call 1-800-874-7711 or 904-741-4400, ext. 468. |  | 927 | 927 | Worldwide Distribution - USA including AL, CA, GA, HI, IL, MO, NC, NY, TX and the country of Japan. | Terminated | October 12, 2011 | 513 | 927 | Not_Computer
Z-2867-2011 | 59077 | [6 empty] | Philips brand BrightView Gamma Camera System, |  |  | June 03, 2011 | 2011 | Philips Medical Systems | The gap between the two parts of the table may permit pinches to occur on either patients skin or other body parts; it has the potential for serious injury if a patients finger is caught between the pallet and patient table resulting in broken finger bone. | Device Design | Philips Healthcare issued a Customer Information letter (Field Change Order (#88200402)) dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instruction for use for their system, which states "Monitor the patient and the system vigilantly during any system motions to make sure that ...accessories, equipment, and all ojbects such as hair, jewelry, or neckties-remain clear of any moving camera and imaging table parts. For additional information, customers were istructed to see the BrightView SPECT instructions for Use, section 2.2 - Safety Warnings and Precautions and Safety and Information Labels. Philips Healthcare will notify all affected customers and will install a sticker through an Action for Performance - Proactive Field Change Order (FCO) to resolve the issue. For customers in North America who require further information or support contact Customer Care Solutions Center at 1-800-722-9377, Select Option 5. In all other countries the local Philips Healthcare office should be contacted. |  |  |  | Worldwide Distribution - USA (nationwide) and Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, China, Chile, Columbia, Czech Republic, Denmark, Dominican Republic, Egypt, France, Greece, India, Iran, Israel, Italy, Japan, Lithuania, Malaysia, Mexico, Netherlands, Peru, Portugal, Russia, Saudi Arabia, Singapore, South Korea, Spain, Switxerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, and the United Kingdom. | Terminated | November 06, 2012 | 522 |  | Not_Computer
Z-2866-2011 | 59167 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Avanta Fluid Management Injection System, Catalog numbers: AVA 500 TABL, AVA 500 PEDL |  |  | June 02, 2011 | 2011 | Medrad Inc | Event was initially intended to be a product enhancement of the MEDRAD Avanta Fluid Management Injection System. Upgrades were initiated in October, 2007. Observance of a possible increase in complaints and adverse event reporting resulted in the firm issuing "Important Product Use Reminder" to remind users of proper procedures for using the system and the need for diligence in reducing air embo | Device Design | Medrad Inc., sent two types of "MEDICAL DEVICE FIELD CORRECTION" notifications on June 2, 2011 to all affected customers. 1. To those customers who Avanta Systems have already undergone market withdrawal and thus have been upgraded, the notification provided customer awareness that prior action should be considered a field correction. The letter stated that NO action is required since their system(s) has already been upgraded. If you have any questions, concerning these activities contact the Director of Quality Systems at (724) 940-6905. 2. To those customers whose Avanta Systems have not yet been upgraded through market withdrawal, the notification informed customers that the recalling firm is undertaking a Field Correction to upgrade their system(s). The letter included a step by step guide that will allow customers to verify that their system(s) has NOT been upgraded. Customers are instructed to complete and fax this form to MEDRAD at (412) 406-0941. If you have any questions, concerning these activities contact the Product Manager at (412) 767-2400, extension 8401. |  | 842 | 842 | Worldwide Distribution -- (USA) (nationwide) including the state of Puerto Ricoand the countries of: AE, AL, AT, BE, BG, BR, CA, CH, CO, CY, DE, DK, DO, ES, FI, FR, GB, HK, HR, HU, IE, IL, IS, IT, JO, JP, KR, KW, LY, MA, MT, MX, MY, NL, PA, PH, PK, PT, QA, RS, RU, SA, SE, SG, SI, TR, VI, and ZA. | Terminated | June 05, 2012 | 369 | 842 | Software
Z-2878-2011 | 59021 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ATLANTIS (R) Self-Tapping Variable Angle Screw, Rx only, Medtronic Sofamor Danek USA, 1800 Pyramid Place, Memphis, TN 38132 |  |  | June 28, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are longer then size etched on the screw. | Other | Medtronic sent a Voluntary Recall letter dated June 28, 2011, directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included. The letter identified the product, the problem, and the action to be taken by the affected customers. If appropriate, personal follow up with Sales Representatives will be made to assure that the recall communication is understood by Risk Managers and to facilitate return of product and the associated Recall response documentation. Foreign consignees will be notified by email on June 29, 2011.For any questions or concerns regarding this recall customers were to contact their Medtronic Sales Representative or the Global Quality Department at 800-876-3133, ext 3173. |  | 101 | 101 units | Worldwide Distribution - USA (nationwide) and the countries of MN, AK, FL, PA, ME, GA HI, CA, TX, OH, IN, MO, WI, MI, NY, IN, LA, MD, MA, NC, OR, NE WA, ID, CO, OR, UT, IL, WV, Canada, Singapore, Netherlands, Brazil, Columbia, Hong Kong, and Mexico. | Terminated | December 24, 2012 | 545 | 460 | Not_Computer
Z-2880-2011 | 59178 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | eSwallow Multi-Program Dysphagia Therapy Test unit |  |  | June 22, 2011 | 2011 | eSwallow USA LLC | It was discovered that the firm was distributing medical devices without FDA approval. | Other | eSwallow USA LLC sent an URGENT MEDICAL DEVICE RECALL letter dated June 22 2011, via US mail and email notification to each distributors. The letter identified the product, the problem, and the action to be taken. Distributors were instructed to immediately examine inventory, quarantine all product subject to recall, and identify customers that received these products and notify them at once of the recall. The letter advised distributors to forward a copy of the recall notification letter to their customers. The letter states that the distributors are to return all inventories to the eSwallow Corporate Office eSWALLOWusa, 106 Hidden Drive, Scottsboro, Al 35769. The recall should be carried out to the customer/user level. Customers are instructed to complete and return the attached Response Form as soon as possible. If you have any questions, call (256) 571-0443. |  | 123 | 123 units | Nationwide (USA) Distribution - including states of AL, AZ, CA, GA, FL, LA, MS, and VA. | Terminated | July 03, 2012 | 377 | 123 | Not_Computer
Z-2879-2011 | 59146 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Arthrex Reciprocating Saw Attachment |  |  | May 24, 2011 | 2011 | W &amp; H DentalWerk Buermoos GMBH | Reciprocating saw attachment V 600 Sr (REF: 17306210), screws can become loose or come off (the attachment) and fall into the surgical wound and may not be noticed due to its small size. | Equipment maintenance | Madrimed GmbH (Distributor partner) of W&H; Dentalmechanik GmbH sent a letter dated May 24, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were ask to advise their users to check the adapters thoroughly before and after each application, until the conversion of all attachments have been completed and to send all products that need to be converted to the nearest service centers.W&H; decided to convert all products that might be affected in their service centres and equip them with an additional security device.If you have any questions, call +43627462360 or email: regulatory@wh.com. |  | 194 | 194 units | Worldwide distribution: USA (nationwide) and countries of: Austia, Australia, Canada, Czech Republic, Denmark, France, Germany, Italy, Kuwait, Malaysia, Netherlands, Russia, Slovakia, South Africa, Spain, Taiwan, Turkey and United Kingdom. | Terminated | October 09, 2015 | 1599 | 194 | Not_Computer
Z-2882-2011 | 57968 | No | JKC | Dehydroepiandrosterone (free and sulfate) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DHEA RIA Reagent |  |  | September 09, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that DHEA measurements using DHEA RIA (REF DSL-8900) kit lots identified may overstate actual DHEA levels in patient samples. | Process control | Beckman Coulter sent an "URGENT: PRODUCT CORRECTION ACTION" letter dated November 15, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue using the affected products and discard them according to local regulations. Customers were asked to review patient DHEA results ran on the affected lots and interpret them in light of additional tests (such as LH, FSH, testosterone, DHEA-S, cortisol, 17-hydroxyprogesterone, and estradiol) and the patient's clinical status. A response form was enclosed for customers to complete and return.Contact the Beckman Coulter Technical Support Center at 1-800-854-3633 in the United States, or their local Beckman Coulter Representative for questions regarding this notification. |  | 3542 | 3,542 units (1181 in US) | Worldwide Distribution--USA (nationwide) and the countries of Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, France, Germany, Hungary, India, Ireland, Italy, Japan, Mexico, Netherlands, Poland, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 04, 2012 | 603 | 3542 | Not_Computer
Z-2885-2011 | 59109 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. |  |  | June 06, 2011 | 2011 | Physio Control, Inc. | The potential reliability issue is the result of a design characteristic of a SRAM (Static Random AccessMemory) electronic component on the main printed circuit board assembly. | Component design/selection | Physio-Control, Inc. sent an "URGENT- ACTION REQUIRED" letter dated June 2011 to all affected customers. The letter describes the product, problem, and the action to be taken by the customer.The letter recommends that users who store their device(s) in hard-shelled carrying cases remove them from the case(s)immediately and discontinue the use of the case(s). Additionally, the letter recommends that customers perform regular maintenance in accordance with the Operating Instructions. A Confirmation Sheet was enclosed with the letter for customers to complete and return via fax at 1-866-448-9567. Customers can call the Technical Support Team at 1-800-442-1142, Option 5. |  | 12197 | 12197 units distributed worldwide (5169 units outside US; 7028 units in the US) | Worldwide Distribution--USA (nationwide) and the countries of Argentina, Australia, Bahamas, Brazil, Canada, Chile, Columbia, Germany, Hong Kong, Iran, Japan, Mexico, Netherlands, Peru, and Uruguay. | Terminated | February 23, 2012 | 262 | 12197 | Hardware |  | Instructions | Safety Notice/Insructions
Z-2884-2011 | 58921 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | LPCCC2: Low Profile Abutment Non-Hexed Castable Cylinder by Biomet 3i |  |  | March 14, 2011 | 2011 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling the Model LPCCC2 Low Profile Abutment, Non-Hexed Castable Cylinder, Lots 2010101036 and 2010101037  Exp. 10/2015. Package labeling may not correctly describe the package contents. Product was distributed in Europe only. | Labeling mix-ups | Biomet 3i sent a "Medical Device Advisory Notice" letter dated February 17, 2011 to all affected customers.The letter described the problem, instructed customers to returned any unused product and to confirm receipt of the notification.For additional information, please contact Customer Service at 800-342-5454. |  | 232 | 232 unit | Worldwide distribution: http://webapps.cdrh.fda.gov/division-tracking/findTrackable.do?docNum=RCL110656&reviewType;=RCL | Terminated | July 27, 2011 | 135 | 232 | Not_Computer
Z-2881-2011 | 59101 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Calloway LabsMulti-Drug 9 Panel Assay |  |  | June 10, 2011 | 2011 | Alere San Diego | Alere is initiating a voluntary recall on the listed lots of MULTI-DRUG 9 PANEL INTEGRATED CUP manufactured by Alere for sole use by Calloway Labs due to a mistake that occurred while processing the order, these products were incorrectlylabeled as CLIA Waived, when in fact the devices are classified as Moderately Complex. Therefore these devices are misbranded. | Vendor change control | Alere San Diego sent an "Urgent Medical Device Recall" letter dated June 9, 2011 to all affected customers.The letter included: products affected, reason for recall and instructed them to return all affected product. An acknowledgement form was also included for the customer to return to Alere.For questions contact Technical Services at 888.246.7483, option 2 or by e-mail at InnovaconComplaints@alere.com. |  | 2400 | 2400 kits (60,000 devices) | (USA) Nationwide Distribution including the states of MO and MA. | Terminated | April 04, 2012 | 299 | 2400 | Not_Computer
Z-2887-2011 | 59016 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent, Part Number: 447070, Lot Number: M908398 &amp; M911529. |  |  | May 31, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because the listed lots of Rheumatoid Factor (RF) Reagent might produce false positive results. A false positive result might lead to additional unnecessary testing. | Pending | Beckman Coulter sent an "URGENT FIELD SAFETY NOTICE/PRODUCT CORRECTION" letter dated May 31, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the affected products. Customers were asked to complete and return an attached FAX BACK Response Form to 714-961-4234. Replacement products may be obtained by calling Order Assistance at 800-526-3821 in the United States, (800) 463-7828 in Canada or contacting your local Beckman Coulter representative.For any questions regarding this notice, contact the Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/Support or call (800) 854-3633 in the United States and Canada; or contact your local sales representative. |  | 6302 | 6302 units (850 in the US) | Worldwide Distribution--USA (nationwide) and the countries of Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, France, Germany, Greece, Hong Kong, India, Ireland, Italy, Japan, Kuwait, Macao, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, and United Kingdom. | Terminated | June 03, 2013 | 734 | 6302 | Not_Computer
Z-2857-2011 | 59288 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Global Focus Marketing and Distribution, Silencer S2200 Centrifuge |  |  | July 15, 2011 | 2011 | Global Focus Marketing &amp; Develop | Potential for bucket and insert not seated on the rotor to cause the lid to break and buckets to be ejected. | Nonconforming Material/Component | Global Focus Marketing and Distribution, Ltd issued a press release on 07/19/11 and a URGENT: DEVICE RECALL letter was mailed to the firm's customers on 07/15/11 and 07/21/11. Discontinue use. For questions, contact: (800) 323-4306 or silencer@gfmd.com |  | 859 | 859 units | Worldwide to USA, United Kingdom, Ireland, Canada, and China. | Terminated | November 06, 2013 | 845 | 859 | Not_Computer
Z-2889-2011 | 59346 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Dual Program Reorder Pack; |  |  | July 11, 2011 | 2011 | Bio-Rad Laboratories Inc | Reports of calibration failure and late retention time using Elution Buffer AA10434 and AA10435. | Other | Bio-Rad Laboratories Inc sent "URGENT: MEDICAL DEVICE CORRECTION" letters to all affected D-10 Dual customers and Bio-Rad subsidiaries on July 11, 2011 via FedEx or fax. The letters identified the product, the problem, and the action needed to be taken. The letters ask that inventory be examined and affected reorder pack lots be quarantined and destroyed according to local, state and federal regulations. All D-10 Dual customers were instructed to review their chromatograms generated from these lots. If the calibration has passed with good chromatography, the controls are within specification, and the F, A1c, and A2 peaks elute in their specified retention time windows for the patient samples, the customer may continue to use and report results with the affected buffer lots. The letters include a chart showing Good Patient Chromatogram with Correct Retention times and Failed Calibrator Chromatogram with late retention times. If calibration fails and/or late retention times cause one or more peaks to be misidentified, affected reorder packs should be destroyed. The correction letters and response form should be translated, if necessary, and distributed to all affected customers. Tracking documentation should be compiled for all customers who receive the Medical Device Correction letter. The Medical Device Correction Response Forms should be completed and returned by fax to Bio-Rad CSD Regulatory Affairs at 510-741-3954 or a PDF copy to Jackie_Buckley@bio-rad.com |  | 418 | 418 | Worldwide -- USA (nationwide) including the states of: AR, CA, CT, FL, HI, IN, MD, MN, MO, MS, NJ, NY, OH, PA, RI, TX. and the countries of: Brazil, Canada, China, England, France, India, Italy, Mexico, New Zealand, and S. Korea. | Terminated | March 12, 2012 | 245 | 418 | Not_Computer
Z-2888-2011 | 59023 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | WarmAir Model 135 Convective Warming Device, |  |  | June 01, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Due to its design the circuit breaker on the device could trip causing it to shut down and not provide heated air to the blanket. | Device Design | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 12, 2011 to all affected customers.The letter includes: description of product, problem, and information to request a repair tubing kit free of charge. It also asks customers to complete and return a "Recall Response Form".For additional information e-mail WA135-2011@cszinc.com or call (800) 989-7373 or (513) 772-8810. |  | 5474 | 5474 units | Worldwide Distribution: (USA) Nationwide, and the countries of Australia, Brazil, Bulgaria, Canada, Chile, Ecuador, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sri Lanka, Taiwan, Thailand, The Netherlands, Turkey, U.A.E., UK, Uruguay, Venezuela, and Vietnam. | Terminated | August 02, 2013 | 793 | 5474 | Hardware |  | Return | Remove or Replace
Z-2890-2011 | 59036 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC(tm) MGI (tm) System Software, Catalog 445922, Lot V5.02A. |  |  | March 28, 2011 | 2011 | Becton Dickinson &amp; Co. | Software contained three additional set configurations which caused the system to reinitialized the AST set carrier defaults. | Software design (manufacturing process) | BD Diagnostic Systems verbally notified all affect customers via the telephone on March 28, 2011. The notification identified the product, the problem, and the action to be taken by the customer. An "URGENT PRODUCT RECALL" follow-up letter was mailed to all customers on April 8, 2011 via UPS. Customers were instructed to complete and fax the attached reply form to acknowledge their receipt of the notification to Regulatory Compliance, 410-316-4256. The letter instructs customers to refer to the modified release notes and IMPORTANT NOTICES outlined in the attached notification letter. If further assistance is needed regarding replacements and all other inquires, please contact BD Technical Services and Support at 1-800-638-8663. |  | 421 | 421 disks | Nationwide Distribution (USA) including states of: AK, AL, AR, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | July 29, 2011 | 123 | 421 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-2893-2011 | 59088 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac Linear Accelerators; |  |  | May 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The throat cover on the High Energy Clinac may detach if not properly installed and possibly strike a patient. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 24, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter provides directions on how to re-install the gantry center throat cover.Customers can contact their Varian Customer Support Manager at 1-888-827-4265 for further clarification. |  | 278 | 278 | Worldwide Distribution-including USA (nationwide) | Terminated | March 14, 2012 | 295 | 278 | Not_Computer
Z-2860-2011 | 59148 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow ON-Q with Select-A-Flow Pumps, |  |  | April 15, 2011 | 2011 | I-Flow Corporation | The recall was initiated because I Flow became aware that a small percentage of ON-Q Pump with Select A Flow may develop a minor leak with a flow rate controller. There may be consequences for the patient, such as under delivery of pain medication or the need to prematurely remove the pump, which would result in loss of the pain treatment provided by the pump. | Process control | I-Flow Corporation sent an "Important Voluntary Market Withdrawal Notice" dated May 12, 2011 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to quarantine and return all impacted products within their facilities. Customers were instructed to respond with in 5 days of receiving the Market Withdrawal Notice and fax it to (949) 206-2663. Customers with additional questions, were instructed to call I-Flow Customer service at (800) 448-3569 (8-5pm, Pacific Standard time) or I-Flow 24 hour technical support line (800) 444-2728. |  | 96892 | 96,892 units total (95,637 units in US) | Nationwide (USA) Distribution and worldwide to: Australia, Canada, Cyprus, Italy, Germany, Finland, Greece, Israel, South Africa, Sweden, UAE, and UK. | Terminated | August 09, 2012 | 482 | 96892 | Not_Computer
Z-2892-2011 | 59377 | [6 empty] | Aquamantys 2.3 Bipolar Sealer |  |  | July 18, 2011 | 2011 | Salient Surgical Technologies, Inc | Insulating material on the shaft of the device may separate and expose the electrical conductor and result in burning the patient skin | Nonconforming Material/Component | Salient Surgical Technologies, Inc. sent an "URGENT MEDICAL DEVICE RECALL" notification letter dated July 18, 2011 to all affected customers and distributors. The letter describes the product, problem and the actions to be taken. The letter requested that customers and distributors immediately identify and discontinue use of all Aquamantys 2.3 Bipolar Sealer devices in inventory and quarantine them until they can be returned to Salient. The letter states that a Salient sales representative will arrange return of the recalled product to Salient. The letter provides information for customers about the replacement of the recalled product. If you have any questions, please contact Customer Service, at 866-777-9400, or your sales representative. |  | 48963 | 48, 963 units | Worldwide Distribution - USA (nationwide) and countries ofChina, Denmark, France, Germany, Greece, Italy, Japan, , South Africa, Spain, and Turkey. | Terminated | March 09, 2012 | 235 | 48963 | Not_Computer
Z-2891-2011 | 59272 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DeRoyal (R) Incision and Drainage Tray, RES 47-763.02, Rx Only, STERILE EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA |  |  | April 28, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | Custom surgical kits contain recalled Triad PVP prep pads | Other | DeRoyal sent a Customer letter dated July 6, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer.The customer was instructed to reference the attached spreadsheet to identify your accounts kits/trays and the corresponding lot numbers affected by this recall.Please complete and return the Acknowledgement Form by fax ( 865-362-3716) confirming you have been notified of this action.For further questions, please call ( 865) 362-2334. |  | 11856 | 11,856 kits/trays | Nationwide Distribution--including the states of CA, FL,GA, HI, IL, IN, LA, MA, MO, MS, NY, PA, and RI. | Terminated | August 07, 2014 | 1197 | 11856 | Not_Computer
Z-2897-2011 | 57856 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio Artida Ultrasound Diagnostic System SSH-880CV. |  |  | October 01, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is initiating a field correction on the Aplio Artida Ultrasound Diagnostic System because the panel operation for the Aplio Artida may become disabled (i.e., system may hang up) when processing to display patient information (heart rate, time, VCP counter, ect.). | Software design | Toshiba America Medical Systems, Inc. sent out an "Urgent Medical Device Correction" dated December 21, 2010 to all affected customers.The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Toshiba will issue a field modification to correct the software, and if the problem occurs, they were instructed to shutdown the system holding down the power supply button for 5 seconds for 5 seconds and then restart the system. A form to send back to Toshiba is included. For questions on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 46 | 46 units | Worldwide Distribution: (USA) Nationwide distribution including the states of MD, OH, PA, NY, MA, GA, AL, TX, IL, WI, IA, MN, KY, MT, OR, LA, and CA; and the country of Puerto Rico. | Terminated | April 20, 2012 | 567 | 46 | Software | Device Operation | Software update | Software Update
Z-2894-2011 | 59113 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow ASK-20608-CMC StimuCath Continuous Nerve Block Kit. |  |  | June 17, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging were found to have an insufficient seal. Therefore, packaging may have been compromised, and sterility of the product cannot be guaranteed. | Packaging | Arrow International, Inc. sent an "Urgent Medical Device Recall" letter dated June 17, 2011 to all affected customers.The letter describes the problem and provides recommendations including; quarantine affected product, return of product, and send an acknowledgement form back to manufacturer.For additional information call 800 233-3187. |  | 95 | 95 kits | Nationwide Distribution including the state of North Carolina. | Terminated | October 01, 2012 | 472 | 95 | Not_Computer
Z-2895-2011 | 58031 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion One: Aquilion TSX-301A, 64 System, X-Ray, Tomography. |  |  | January 07, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction for the TSX-301A and TSX-301B: Aquilion 64 System, X-Ray, Tomography because communication between reconstruction units in the system console may be broken due to failure of the optical conversion connector, resulting in a break of communication between units. | Software design | Toshiba America Medical Systems, Inc. sent an "Urgent Medical Device Correction" letter dated January 4, 2011 to all affected customers.The letter provides the customers with an explanation of the problem and the firms plan to install preventive measures and provided them with a Customer Reply form to send Toshiba. A representative will contact the customers for an appointment to install the measures to their CT system. For question on this recall call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 6 | 6 | (USA) Nationwide Distribution including the states of WI, MO, TN, IA, AL, AZ, TX, and GA. | Terminated | August 17, 2012 | 588 | 11 | Computer
Z-2848-2011 | 59047 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | Vital Signs- A GE Healthcare Company - Hygroscopic Condenser Humidifier (HCH)/Anesthesia Breathing Circuit; GE Medical Systems SCS 283 rue de la Miniere 78530 BUC, France p +33.1.30.70.40.40; Vital Signs Colorado Inc 11039 E. Lansing Circle Englewood, CO 80112 P +1.973.956.5300. |  |  | April 29, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | The Hygroscopic Condenser Humidifier (HCH) device may have an occlusion that can prevent proper flow of the medical gases or oxygen, possibly resulting in insufficient oxygen delivered to the patient. | Device Design | An "Urgent Medical Device Correction" letter with the return confirmation form was sent to affected customers. Phone calls were also placed to customers. Consignees were informed of the affected device, issue, and actions to take. Customers were instructed to isolate all affected product, arrange return by contacting Customer Service at 1-800-932-0760, and return only unused product. Customers were asked to complete the attached confirmation form and fax back as per the instructions. For questions, contact your local Vital Signs Account Manager or call Customer service at 1-800-932-0760. |  | 222009 | 222,009 | Worldwide Distribution including the United States. | Terminated | August 06, 2014 | 1195 | 222009 | Not_Computer
Z-2883-2011 | 59212 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Arrow NextStep Antegrade Chronic Hemodialysis Catheter. |  |  | June 28, 2011 | 2011 | Arrow International Inc | Firm has received complaints indicating breakage and/or separation of the stylet within the Arrow(R) NextStep(R) Antegrade Catheters. | Device Design | An "Urgent Medical Device Recall" letter dated June 28, 2011 was issued to consignees. The letter requested that customers check their stock, cease use and distribution, and quarantine all affected product. Consignees were further directed to complete the "Recall Acknowledgement &amp; Stock Status Form", return the form to their sales representative, and return affected product as instructed on the letter. For questions or concerns, call 1-800-233-3187. |  | 41 | 41 | United States (states of CA, DE, FL, MI, NC, and TN). | Terminated | April 16, 2012 | 293 | 41 | Not_Computer
Z-2900-2011 | 59046 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | AxioBionics Wearable Therapy BioShorts that includes 2-16 BioGel Velcro Electrodes in sizes: 2"x2", 2"x2.5", 2"x3", 2"x3.5", 2"x4", 2"x5", 2"x6", 2"x8", 1.5"x3", 1.5"x5", 1.5"x6", and 1"x1.5". |  |  | May 27, 2011 | 2011 | Axiobionics, LLC | Premature Device Failure : Failure of the electrode sites to conduct and provide muscle stimulation due to silver fabric component failure. | Nonconforming Material/Component | Axiobionics, LLC is notifying all affected customers by phone. Three attempts will be made to contact the customer by phone. Contact by telephone will be made to prescribing physician to determine if Axiobionics has latest patient contact information. If information is different, new contact information will be recorded and 3 additional attempts will be made with the new information.If the patient cannot be reached by telephone, a certified letter will be sent to the patient's address. If contacting by mail, the letter template will be provided to the FDA recall office . All patients should have been contacted by telephone, certified mail, or in person by June 10, 2011.All patients affected by the field correction will be asked the following questions to determine if their electrodes are functional and resistances are not exceeding the output capabilities of the stimulator they are using. 1. Are you using your garment?2. Is your garment delivering the stimulation required for your therapy?3. If the patient is using an NP4 stimulator: What are the resistance readings on each channel?4. If the patient answers No to question 2 or states that the resistance is reading greater than 200 ohms on at least one channel in question 3, the garments will be returned to Axiobionics for repair.Garments that are not functioning correctly will be repaired first. Functioning garments will be tested by immersion into a salt bath under the given protocol to determine if the silver fabric is properly coated.Questions regarding this correction should be directed to the Axiobionics representative at 800-552-3539. |  | 9 | 9 | Nationwide Distribution--USA (nationwide) including the states of AZ, GA, IL, IN, MI, MO, NY, TX, and WY. | Terminated | August 29, 2013 | 825 | 106 | Not_Computer
Z-2913-2011 | 58920 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***BIOMET MICROFIXATION***CENTER DRIVE STAINLESS STEEL***2.0MM SYSTEM SCREWDRIVER BLADE***Non-Sterile Product***BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" |  |  | March 15, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation, Jacksonville, FL is recalling Center Drive Blade, 2.0, Model Number 01-9047, Lot # 064980. Product is non-conforming (blades do not retain screws). | Nonconforming Material/Component | The firm, Biomet Microfixation, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" undated letter to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the affected products and immediately return them to Biomet Microfixation, Attn: Global Complaint Coordinator, 1520 Tradeport Drive, Jacksonville, FL 32218-2480; complete and return the Inventory Reconciliation sheet via Fax to the attention of: Global Complaint Coordinator at 904-741-9425, and notify their customers if product was further distributed. If you have any questions, contact the Global Complaint Coordinator at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 35 | 35 | Worldwide distribution: USA (nationwide) and countries of: Canada, Israel and South Africa. | Terminated | November 09, 2011 | 604 | 35 | Not_Computer
Z-2902-2011 | 59136 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Ultraview Spacelabs Medical Multigas Analyzer, Model 91518, V1.00.14. |  |  | June 08, 2011 | 2011 | Spacelabs Healthcare, Llc | Spacelabs received a report of the Model 91518 Multigas Analyzer trending low levels of an anesthetic agent that is not present during a particular case. The error is limited to only the trend table data. | Software design | Spacelabs Healthcare, Llc. sent a "Customer Bulletin" dated June 20, 2011 to all affected customers. The customers were notified of the trend problem and are advised to follow the instructions in the Perioperative section and the ADMIT/DISCHARGE section of the users manual.For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support or 1-425-657-7200 x5089. |  | 481 | 481 units worldwide (333 units internationally and 148 units in the US) | Worldwide Distribution: (USA) Nationwide distribution, and the countries of ARGENTINA, BAHRAIN, BOLIVIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, ECUADOR, EGYPT, FRANCE, INDIA, MEXICO, NICARAGUA, PARAGUAY, PERU, POLAND, Slovakia, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, CHINA, THAILAND, UNITED ARAB EMIRATES, U.K., URUGUAY, VENEZUELA, and VIETNAM. | Terminated | September 21, 2011 | 105 | 481 | Software
Z-2915-2011 | 59210 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore computer tomograph software version 3.6. Catalog #882476, Gantry S/N 9001-9023. |  |  | August 05, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Two non-safety related defects have been found on the Philips GEMINI TF Big Bore running version 3.6; if these defects were to recur, it could impact the performance or reliability of the system. Philips is issuing a software update to all Philips GEMINI TF Big Bore customers to correct these defects. | Software design (manufacturing process) | Philips Medical Systems (Cleveland) Inc. sent an "GEMINI TF BIG BORE CPE CUSTOMER LETTER" dated October 20, 2011 to all affected customers. The letter describes the product, problem, and updates that are available to the customers. The letter provides additional information about acquisition workflows for SUV Calibrations and Validations.For support concerning this upgrade, contact the Customer Care Solutions Center at 1-800-722-9377. |  | 23 | 23 units | Worldwide Distribution-USA (nationwide) including the states of CO, LA, MD, MI, MO, MN, PA, TX, VA, VT and WI and the countries of Belgium, Germany, Italy, Saudi Arabia and Switzerland. | Terminated | September 21, 2012 | 1143 | 23 | Software | Output/Calculation | Software update | Software Update
Z-2914-2011 | 59166 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical Brake Hubs component of the Philips MX8000 Series Dual v. Exp. Computed Tomography System, Serial Numbers: 8008-8413, 80001-800022, Model #728130, 510 (k) #K010817. |  |  | June 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips Medical Systems decided to conduct a recall/field correction of the patient support vertical brake hub component in all of their CT MX8000 Dual v. Exp Computed Tomography X-Ray System scanners due to a component defect which could cause an unexpected downward motion of the table (couch) while the patient is on the table. This event could occur because the short key connecting the vertical | Device Design | Philips Medical Systems (Cleveland) Inc. issued an "URGENT - Medical Device Correction - Field Safety Notice" letter dated June 1, 2011 to all affected customers, via Certified Mail for tracking purposes. The notification identified the product, the problem involved, and the action(s) to be taken by the customers with regard to this device issue. The notification instructs the customers to monitor their units and if the recalled product malfunctions, to immediately notify their Phillips service representative. The letter also states that a Philips Field Service Engineer will visit each customer location and perform a brake repair or replacement on their system to resolve the issue. For further information or support, the customers are instructed to contact their local Philips representative or local Philips Healthcare office. For North America and Canada the customers are instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompt(s). |  | 412 | 412 system units were distributed. | Worldwide Distribution -- USA (nationwide) including the states of: AR, AZ, CA, CO, FL, FA, HI, IL, IN, KY, MA, MI, MO, MS, NC, NV, NY, OH, OK, SC, SD, TN, TX, VA, VT, WA and Puerto Rico and the country of Canada. | Terminated | August 07, 2013 | 796 | 412 | Not_Computer
Z-2903-2011 | 58698 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | "***REF MCA5-15SB QTY: 1***1.5 mm Fluted Ball, 5.6 cm. Use with Micro Curved Attachment (MCA).***Sterilized using Irradiation***RX Only***ANSPACH Manufacture The Anspach Effort, Inc. 4500 Riverside Drive |  |  | October 13, 2010 | 2010 | The Anspach Effort, Inc. | On July 30th, 2010 the Anspach Effort, Inc. initiated a voluntary recall on their Anspach Single Use, Sterile Bone Cutting Burrs. A complaint was recieved stating that the product had failed during a procedure. | Nonconforming Material/Component | The Anspach Effort, Inc. sent an "URGENT: MEDICAL DEVICE PRODUCT REMOVAL" letter dated October 20, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customer to screen their inventory for the affected product and not to include any product received after October 18, 2010. Customers are instructed to contact Anspach Customer Support at 800-327-6887 to arrange for return and replacement of product. Customers are to complete and return a Customer Reply Form via fax at 800-327-6661.Questions regarding this notice should be directed to Anspach Customer Support at 800-327-6887.Customer notification began on 10/20/2010 in writing (Fed Ex) with instructions to return the required devices. Anspach utilized their sales representatives to facilitate the product replacement and return activities. |  | 46 | 46 | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, France, Italy, Japan, South Africa, Spain, Sweden, and United Kingdom. | Terminated | November 07, 2011 | 390 | 1816 | Not_Computer
Z-2754-2011 | 59111 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 6F Zuma Guide Catheter with AL 1.5 Curve with side holes |  |  | June 08, 2011 | 2011 | Medtronic Vascular | Mislabeled; Product labeled as 6F Zuma Guide with AL 1.5 curve with side holes may contain a 5F Zuma Catheter with SAL 1.0 curve | Labeling Change Control | The firm, Medtronic, sent an "Urgent Medical Device Recall Notice(for US)" letter dated June 7, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to segregate the product so that it will not be used. A Medtronic Sales Representative will visit the customers to account for all units they received and initiate a credit as appropriate.If you have any questions, please call (978) 739-0371 or (978) 739-3259. |  | 5 | 5 units | Nationwide distribution: USA including states of: CA, FL, and MA. | Terminated | April 17, 2012 | 314 | 5 | Not_Computer
Z-2916-2011 | 59441 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Vital Sign elance Monitor, Model 93300 |  |  | July 06, 2011 | 2011 | Spacelabs Healthcare, Llc | Model 93300 Vital Signs Monitor may stop operating while in use, and when this occurs, the monitor may also fail to power ON. The resulting situation may cause a loss of monitoring for the patient. | Process design | Spacelabs Healthcare, Llc sent an "Urgent-Medical Device Correction" letter dated 18 July 2011 to all affected customers. The letter notified the customers of the problem and recommended customers make their staff aware. It also informed them a Field Service Engineer will contact them to check, and if necessary make adjustments.Customers can contact Spacelabs Healthcare at 800-522-7025 and select 2 for Technical Support or outside the U.S. at +86512-8717-1888, extension 304. |  | 1970 | 1970 units | Worldwide Distribution: (USA) Nationwide distribution including the state of WI, and the countries of Puerto Rico, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Chile, China, Colombia, Cyprus, Dominican Republic, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guatemala, Honduras, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Kenya, Korea, Kuwait, Lebanon, Malaysia, Napel, Netherlands, Nicaragua, Oman, Palestine, Panama, Paraguay, Peru, Philippines, Republic of Macedonia, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Tunisia, Turkey, UAE, United Kingdom, Venezuela, and Vietnam. | Terminated | January 25, 2012 | 203 | 1970 | Other | Device Operation | Notification | Safety Notice/Insructions
Z-2928-2011 | 58684 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | remel Haemophilus Test Medium (Agar), 100mm plates, 10/pk, Ref. #R01503. The firm name on the label is Remel, Lenexa, KS. |  |  | April 21, 2011 | 2011 | Remel Inc | Organism strains may fail to consistently produce confluent growth | Other | The firm, Remel, Inc., sent two "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" letters one dated April 20, 2011 to: Microbiology Lab Supervisors and the second dated April 25, 2011 to:Distributors via regular mail. The letters described the product, problem and actions to be taken. The customers were instructed to to review their inventory for the listed product lot(s); discard all remaining units of the product lots; and complete and return the Product Inventory Checklist by May 9, 2011 regardless if they have affected product in the enclosed self-addressed, stamped envelope. The distributors were further instructed to notify their customers if they further distributed the product. If you have inquiries concerning this advisory, please contact our Technical Services Department at 800-255-6730 or 913-888-0939. |  | 223 | 223/10-pack boxes | Nationwide distribution. | Terminated | June 21, 2013 | 792 | 761 | Not_Computer
Z-2919-2011 | 59102 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, XiO versions equal to or greater than 2.4.0 and less than 4.63.0 |  |  | June 09, 2011 | 2011 | Computerized Medical Systems Inc | The XiO Software currently prompts for entry of CT to ED data for a studyset. No indication is given that the Electron Density data should only be used for Photon and Electron calculations and either Relative Stopping Power or Mass Density should be used for Proton calculations (depending on the algorithm). | Software design | The firm, Eletka, Inc., issued an "Important Safety Notice" dated May 18, 2011 via mail to all affected sites. The notice explains the issue and identifies a work around to the problem until the new release is available. The customers were instructed to complete and return the enclosed "Confirmation of Receipt" form via FAX to 314-993-1175, attn: QA; e-mail or mail: CMS Software, Attn: QA, 13723 Riverport Drive, Suite 100, Maryland Heights, MO 63043.The notice states that "this problem has existed since XiO Release 2.4.0 and has been resolved in XiO Release 4.63.0 by adding warning messages and explanations to the XiO On-Line Help. Your site will be notified when this release is available."If you have any question, call 314-993-0003, 800-878-4267 or 408-830-8023. |  | 7 | 7 | Worldwide distribution: USA (nationwide) including: IL, IN, MA, NJ and OK; and country of Germany. | Terminated | April 04, 2013 | 665 | 7 | Other | Display/Image | Software update | Software Update
Z-2918-2011 | 58978 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Aruba Wireless Network Access Points, Model 65 utilized by the Acuity Central Monitoring System. |  |  | May 16, 2011 | 2011 | Welch Allyn Protocol, Inc | Aruba Wireless Network Access Points, Model Number 65, experienced a higher than expected incidence of dropouts due to external RF interference. | Software design | Welch Allyn Protocol, Inc. sent a letter dated May 16, 2011 to affected customer.The letter stated there were concerns over drop-outs in the system recommended the system at this hospital to revert back tothe Symbol wireless infrastructure. For questions call Welch Allyn at 315-685-4100. |  | 12 | 12 units were affected | Worldwide distribution: (USA) Nationwide distribution including the state of California; and the countries of Singapore; Philippines, France, Italy, Malaysia, Poland, UK, England, UAE, Germany, Spain, Ireland, Portugal, New Zealand, Hong Kong, Thailand, Australia, Canada, and the Netherlands. | Terminated | August 12, 2011 | 88 | 12 | Other
Z-2931-2011 | 58852 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Pinnacle Pelvic Floor Repair Kit--Anterior/Apical STERILE, and Pinnacle Pelvic Floor Repair Kit--Posterior STERILE. |  |  | May 10, 2011 | 2011 | Boston Scientific Corporation | The device may exhibit low tensile strength between the needle and suture and lead to needle detachment during mesh leg placement. | Device Design | Boston Scientific Corporation sent on May 10, 2011 an "URGENT MEDICAL DEVICE RECALL-IMMEDIATE ACTION REQUIRED" letter to all affected customers.The letter included; reason for recall, instructions for the quarantine, discontinued use, and return of the devices, and information for Medwatch reporting. The letter also includes a form that is to be faxed to the firm.For information call (508) 683-4678. |  | 540 | 540 | Worldwide Distribution; (USA) Nationwide distribution, and the countries of Puerto Rico, France, Great Britain, Ireland, Italy, Spain, Sweden, Canada, and Australia. | Terminated | November 15, 2011 | 189 | 540 | Not_Computer
Z-2926-2011 | 59214 | No | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | 1125 Prime Zoom Stryker Medical Stretcher |  |  | July 19, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | Stryker Medical has determined that some Prime Brake Pad Assembly components were manufactured with brake shoes that did not meet specifications. Accordingly some brake shoes used within the Prime Brake Pad Assembly may wear prematurely, over the life of the product. Reduced braking force may result in a patient fall or caregiver muscle strain, either of which may require medical intervention. | Nonconforming Material/Component | Stryker Medical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 18, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to locate the affected stretchers and to test the brakes. If the brakes are found to be inoperable, customers are instructed to remove the stretchers from service and contact the firm at 800-STRYKER, option 3. Additionally, customers are instructed not to put the stretchers back in service until they are repaired by the firm's field service representative. For units with properly functioning brakes, the letter also instructs the user to test the brakes before each use until the unit has been serviced. A post card was enclosed to confirm receipt of the letter. Questions or concerns may be addressed by calling Stryker at 269-389-6604. |  | 7 | 7 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, FL, GA, HI, IL, MA, MN, MO, NH, NJ, NV, NY, PA, RI, TX, UT, and VA and the country of THE NETHERLANDS. | Terminated | August 28, 2013 | 771 | 164 | Not_Computer
Z-2932-2011 | 59356 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH, GB: Female Catheter, All Purpose PVC, 08 Fr. 18 cm, Distributed by Teleflex Medical, Research Triangle Park, NC |  |  | June 30, 2011 | 2011 | Teleflex Medical | Pin holes were detected in the packaging therefore sterility of the device can not be guaranteed. | Other | Teleflex Medical sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 30, 2011 to all affected consignees. The letter described the product, problem and actions to be taken. The letter instructs the customers to discontinue use and quarantine any of the affected products. Customers were asked to complete the attached Recall Acknowledgement Form and return as per the instructions. A separate letter was sent to distributors instructing them to communicate the recall to any of their customers who may have receive the affected products, using the recall letter and form as a template. For questions, contact your local sales representative or call Customer service at 1-866-246-6990 |  | 14378890 | 14,378,890 units | Worldwide Distribution and (USA) Nationwide Distribution. | Terminated | April 29, 2014 | 1034 | 14378902 | Not_Computer
Z-2930-2011 | 59300 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical da Vinci Si Surgical System (IS3000); |  |  | July 05, 2011 | 2011 | Intuitive Surgical, Inc. | Potential failure of the retention component of the Master Tool Manipulator (MTM) of the Intuitive Surgical da Vinci Si Surgical System which could cause uncontrolled movement. | Device Design | Intuitive Surgical sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated July 1, 2011 to all affected customers. Customers were asked to complete the attached response form and return as per the instructions. The letter identifies the product, problem, and actions to be taken by the customers. The letter reinforce to customers to follow the product labeling WARNING. The letter states that all affected customers sites will be receiving a call from a Customer Support staff to schedule a retrofit of the component. A letter was also sent to non affected customers to inform them of the urgent medical device correction and actions to be taken. If you have any questions, please contact Customer Service at 1-800-876-1310 or for Outside United States call +41 21 821 2020. |  | 183 | 183 | Worldwide Distribution -- USA (nationwide) and the countries of: Belgium, Czech Republic, Denmark, France, Germany, Greece, India, Ireland, Israel, Italy, Netherlands, Norway, Philippines, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Switzerland, Turkey and United Kingdom. | Terminated | September 02, 2011 | 59 | 183 | Not_Computer
Z-2920-2011 | 59419 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Solution Set with Duo-Vent Spike; 109" (2.8 m), Male Luer Lock Adapter 10 drops/mL. |  |  | July 21, 2011 | 2011 | Baxter Healthcare Corp. | The Solution Sets are being recalled because the position of the slide clamp and roller clamp may have been reversed during assembly. The reversed orientation of the slide clamp and roller clamp will cause the set to be incorrectly loaded into an infusion pump, resulting in fluid flow from the patient to the pump, rather than from the pump to the patient. | Employee error | Baxter Healthcare Corporation sent a "Urgent Product Recall" letter dated July 21, 2011 to all affected customers.The letter included a description of the problem and requested they contact Baxter at 1-888-229-0001 to return all affected product. They were asked to complete and return the attached customer reply form, and to contact and notify their customers.For questions contact Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 4752 | 4,752 units | (USA) Nationwide Distribution including the states of Alabama, California, Iowa, Illinois, Indiana, Massachusetts, Maryland, Michigan, North Carolina, New Hampshire, Ohio, Pennsylvania, Tennessee and Virginia. | Terminated | June 18, 2012 | 333 | 4752 | Not_Computer
Z-2917-2011 | 59301 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Automated External Defibrillator (non-wearable). Model number 9200, 9200N, 9200DN, 9200RD, 9210N, 9210RD, and CardioVive 92530. |  |  | June 29, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) intent to end service support and discontinue accessory sales of Powerheart G2 Automated External Defibrillators. Customers are recommended to choose a replacement AED prior to the current AED displaying any "Service Required" self test messages. | Nonconforming Material/Component | Cardiac Science Corporation (CSC) sent two "9200 SERIES CARDIAC SCIENCE AED NOTICE" letters dated August 2010 and March 2011 to all affected customers. The letters identifies the product, problem, and actions to be taken by the customers.Customers are encouraged to upgrade their devices through special offers by calling CSC at 800-426-0337, Option 2.For questions about the end of factory support for the affected devices, call the Technical Support team at 800-426-0337, Option 1.*****In August 2011, Cardiac Science sent out another updated recall notification letter to their consignees. This letter stated that once a device has entered the "Service Required" state, please contact the firm at 800.426.0337 Option 2. Per prior notifications, the 9200 series AEDs have reached the end of service life and can no longer be repaired. Information about these notifications may be found as an announcement in the Services and Support section of Cardiac Science website.Cardiac Science advised the customers to upgrade to a Powerheart AED G3 (this device supports the 2010 AHA/ERC Guidelines). It is stronglyrecommended that customer to choose a replacement AED prior to the current AED displaying any "Service Required" self test messages. For information about special offers to help you upgrade please call 800.426.0337. Option 2. |  | 48700 | 48,700 units (41396 units in US; 7304 units Internationally) | Worldwide Distribution-USA (nationwide) and the countries of AE, AT, AU, AW, BE, BR, CA, CH, CL, CN, CO, CY, CZ, DE, OK, EG, ES, FI, FR, GB, HK, HR, HU, 10, IE, IL, IN, IR, IS, IT, JP, KR, KW, KY, LB, LT, LV, MT, MX, NG, NL, NO, NZ, PE, PF, PL, PR, PT, QA, RO, RU, SA, SG, SK, SY, TH, TR, TT, TW, VG, ZA, and ZW. | Terminated | December 20, 2011 | 174 | 48700 | Not_Computer
Z-2945-2011 | 59058 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard PS Open Box Femoral 70 MM Left Interlock, " FOR CEMENTED USE ONLY" CO-CR-MO ALLOY. |  |  | July 06, 2011 | 2011 | Biomet, Inc. | The recalled device was manufactured with two distal pegs but package labeling indicates the device is pegless. | Employee error | Biomet Orthopedics sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter to all affected customers.The letter included: reason for the recall, risks associated with use of the device, and alternate instructions for use of the recalled device. Customers are instructed to quarantine and return the devices, and a return form.For information or questions on this recall call Biomet at (574) 371-3756. |  | 8 | 8 | Worldwide Distribution: (USA) Nationwide including the states of NJ, PA, LA, SD, TX, and KS; and the country of The Netherlands. | Terminated | September 26, 2012 | 448 | 8 | Not_Computer
Z-2948-2011 | 59252 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT Gamma Camera, 3/8". |  |  | June 15, 2011 | 2011 | Philips Medical Systems | Non-conformances with BrightView X/XCT systems running version 2.0.1 including data transfer, XCT acquisition, Pixel Mask for image reconstruction and table moving incorrectly. | Software design | Philips Medical Systems, Inc. set a "Customer Information" letter dated July 8, 2011 to all affected customers.The letter informed them of an upcoming software updated (version 2.5).For questions please contact 800-722-9377 (option 5). |  | 213 | 213 units, all systems | Worldwide Distribution: (USA) Nationwide distribution and the countries of Greece, Australia, Belgium, China, Czech Republic, France, Germany, Indonesia, Israel, Japan, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Puerto Rico, Saudi Arabia, Slovakia, South Korea, Spain, Switzerland, Taiwan U.A.E., and the UK. | Terminated | February 15, 2013 | 611 | 213 | Software
Z-2950-2011 | 59454 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | CAPIOXÂż Flexible Venous Reservoirs, 1500 mL, product codes: CXFVR1500X and CXFVRI500XC, used singly or packaged as components in cardiovascular procedure trays labeled in part ***Terumo Caerdiovascular Systems Corporation, Elkton, MD 21921*** |  |  | July 21, 2011 | 2011 | Terumo Cardiovascular Corporation | Reservoir blood bag used in cardiovascular procedures may leak. | Nonconforming Material/Component | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated July 20, 2011, to all affected consignee. The letter described the problem, the potential hazard, and the actions to be taken. The letter instructed customers to identify affected inventory and segregate from non-affected product. The recalling firm advises users to not use affected product and return for replacement. The customers are instructed to call the recalling firm Customer Service for a Returned Goods Authorization (RGA) number. Customers were asked to complete the attached Response Form and return as per the instructions. Consignees with questions should call 1-800-521-2818 Monday through Friday 8AM to 6PM EST. |  | 253 | 253 reservoirs | (USA) Nationwide Distribution including states of: FL, MI, MT, NY, TN and VA. | Terminated | February 10, 2012 | 204 | 253 | Not_Computer
Z-2946-2011 | 59368 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, ProntoÂż V4, Extraction Catheter, Sterilized using ethylene oxide, Rx Only. The following sizes/model numbers are affected: 5.5F( 4005), 6F(4006, 7F(4007, 8F(4008) are involved in the recall. |  |  | July 11, 2011 | 2011 | Vascular Solutions, Inc. | Vascular Solutions , Inc recently discovered that there was a manufacturing defect with Pronto V4 extraction catheters (Models 4005, 4006, 4007, 4008). Use of this defective product could lead to a tip separation. This maycause injury and/or procedural delay, resulting in a deteriorating hemodynamic condition of the patientwhich may require medical intervention. | Process control | Vascular Solutions, Inc. sent an "URGENT: MEDICAL DEVICE FIELD ACTION" letter dated July 8, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify the affected product and remove from current inventory. Included with the letter was a Field Action Customer Inventory Form for customers to complete and return. VSI will contact customers for replacement product upon receipt of the form and returned product.Contact your local representative for questions concerning this notice. |  | 1086 | 1086 | Nationwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, FL, ID, IL, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK ,OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | April 03, 2012 | 267 | 1086 | Not_Computer
Z-2952-2011 | 59382 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Turbo-Booster Laser Guide Catheter, 8 Fr, REF 518-063, Sterile EO, Spectranetics Corporation, 96 Talamine Court, Colorado Springs, CO 80907. |  |  | July 18, 2011 | 2011 | Spectranetics Corporation | Packaging integrity may be compromised, thus sterility can not be assured. | Package design/selection | Spectranetics sent an "PRODUCT RECALL NOTICE: EU SHIPMENTS OF TURBO-BOOSTER LASER GUIDE CATHETERS" letter dated July 15, 2011 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and set aside affected product. Customers are to contact their sales representative for information on how to return the affected product and receive replacement product.Further information is available at +011.49.719.447.2539. |  | 123 | 123 units | Foreign Distribution-including the countries of Belgium, Cyprus, Germany, Greece, Italy, Spain, Switzerland, and United Kingdom. | Terminated | August 12, 2011 | 25 | 180 | Not_Computer
Z-2953-2011 | 58043 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron JR. Microcoagulation Systems Prothrombin Time (PT). |  |  | May 10, 2011 | 2011 | International Technidyne Corp. | Certain lots of Hemochron Jr. PT Cuvettes (J201) demonstrate a higher bias than historically observed when compared to laboratory reference instruments. | Device Design | International Technidyne Corp. (ITC) sent "Urgent Medical Device Recall" letters dated May 10, May 25 and June 09, 2011 to all affected customers.The letter included a description of the problem, asked customers fill out the attached form, and return affected product for replacement.For questions, please contact ITC Technical Support at (800) 631-5945. |  | 10931 | 6,218 boxes in US ; 4713 boxes outside US | Worldwide Distribution: (USA) Nationwide Distribution including the countries of; Austria, Australia, Canada, Switzerland, Germany, Denmark, Estonia, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Korea, Liechtenstien, Libya, Netherlands, Norway, Oman,Portugal, Qatar, Saudi Arabia, Sweden, Turkey, Taiwan, and South Africa. | Terminated | November 02, 2012 | 542 | 10931 | Not_Computer
Z-2956-2011 | 59379 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT -Cardiovascular Perioperative Autotransfusion System |  |  | July 27, 2011 | 2011 | Haemonetics Corporation | Updated Operations Manual for the deployment of the spill collection system and cleaning | Component design/selection | Haemonetics sent a "FIELD NOTIFICATION: UPDATE OF CardioPAT OPERATIONS MANUAL FOR DEPLOYMENT OF SPILL COLLECTION BAG AND CLEANING" letter dated July 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter requests that customers complete an enclosed Customer Confirmation sheet indicating that the affected devices are in their possession and to return the sheet via fax to Haemonetics at 781-987-9252. Additionally, instructions for deployment of the spill collection drainage system, revised Operations Manual and Quick Reference Guide Addendum are enclosed with the letter.Contact Haemonetics at 1-800-537-2802 for questions regarding this notice.. |  | 316 | 316 units | Worldwide Distribution--USA (nationwide) and the countries of Canada, France, Germany, Great Britain, Italy, Japan, Spain, and The Netherlands. | Terminated | June 09, 2014 | 1048 | 316 | Not_Computer
Z-2955-2011 | 59392 | No | IKD | Electrode cable. | Physical Medicine | Physical Medicine | 510(K) Exempt | Supreme Electrophysiology Extension Cable, REF 401984, For use with Hexapolar Catheters, Sterile EO, Rx only, St Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55345. |  |  | June 01, 2011 | 2011 | St. Jude Medical | St. Jude Medical is recalling a single lot (3334664) of Supreme Electrophysiology Extension Cable which were packaged with a Quadripolar cable instead of a hexapolar Cable. St. Jude Medical has determined that the incorrectly labeled product poses no safety risk since the extension cable can be safely used with a corresponding Quadripolar catheter but cannot be incorrectly connected to a Hexap | Labeling mix-ups | St. Jude Medical sent an "IMPORTANT MEDICAL DEVICE FIELD ADVISORY NOTICE" letter dated June 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate the affected product and to complete and return an attached SJM AFD Reconciliation Form. The form will be used to replace product and as documentation that the customers received the notice. Fax the completed form to 877-497-031.Contact the Director of Product Marketing at 512-329-1054 for questions regarding this notice. |  | 32 | 32 | Nationwide Distribution-USA (nationwide) including the states of FL, NC, NY, and OR. | Terminated | August 16, 2012 | 442 | 32 | Not_Computer
Z-2959-2011 | 59119 | No | FJS | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Aspira Drainage Kit, Product Code 4991502. |  |  | May 21, 2011 | 2011 | Bard Access Systems | Labeling for drainage kit contained incomplete information. | Error in labeling | The firm, BARD, sent a letter entitled "Labeling Discrepancy on Aspira Drainage Accessory Kits" dated June 12, 2011 to its customers. The letter provides the customers with the information missing from the labeling and/or IFU stating that the product can be safely used for peritoneal and pleural applications.If you have any questions please do not hesitate to contact BARD at 1-800-290-1689. If there are any clinical questions please contact Medical Services and Support at 1-866-893-2691. |  | 39290 | 39,290 units | Nationwide distribution. | Terminated | November 02, 2011 | 165 | 68530 | Not_Computer
Z-2957-2011 | 59352 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | HOLDEX Single-Use Holder PP closed pack. |  |  | June 29, 2011 | 2011 | Greiner Bio-One North America, Inc. | At removal of the tube from the holder, the needle may dislodge and blood spill can occur. | Other | Greiner Bio-One North America, Inc. sent a "Recall HOLDEX Single Use Holder PP" letter dated June 29, 2011 to all affected customers.The letter instructed customers stop distributing the product immediately. A Product Disposition form was provided to be completed and returned to Greiner Bio-One at which time upon receipt, arrangements for product return to Greiner Bio-One will be made for replacement or credit. They were requested to provide a list of their customers to Greiner Bio-One that received the affected item.For additional information call Greiner Bio-One at (888) 266-3883. |  | 142400 | 142,400 pieces | Nationwide Distribution including the states of CA, GA, IL, KS, MA, IL, PA and VA. | Terminated | November 03, 2011 | 127 | 142400 | Not_Computer
Z-2958-2011 | 59449 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015. Product Code: 2M8161. |  |  | July 22, 2011 | 2011 | Baxter Healthcare Corp. | Baxter's Service Center tested and returned to the user the Colleague infusion pump that did not meet PSIG (Pounds-force per Square Inch) specifications | Other | Baxter Healthcare Corp. notified the consignee by telephoned on July 22, 2011, with the call documented per the approved call script. A follow-up "URGENT PRODUCT RECALL" notification was sent dated July 25, 2011, reiterating the telephone communication informing them that the COLLEAGUE pump may not have met certain PSIG (Pounds-force per Square Inch) specifications after testing by the Baxter Global Technical Service Center. Both forms of notification identified the affected product, problem description, hazard involved, and actions needed to be taken. The consignee was instructed to locate the device, remove it from service and return it to Baxter by contacting Baxter Global Technical Services at 1-800-843-7867 option 1, to obtain a prepaid UPS label and shipping box. Upon receipt of the returned unit, Baxter will perform the required testing and address any issues through repair or service. The notication instructs the consignee to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter as per the instructions provided. If the recalled product have been distributed to another facility, or if you are a dealers, wholesalers and distributors, the letter request that this communication be forward to the pump owner as appropriate. For any questions regarding this communication contact the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central time. |  | 1 | 1 unit | (USA) Nationwide Distribution including the state of Minnesota. | Terminated | February 14, 2012 | 207 | 1 | Not_Computer
Z-2965-2011 | 59330 | No | LZS |  |  | Ophthalmic | PMA | STAR S4 IR Excimer Laser System; Software Version 5.00 |  |  | July 18, 2011 | 2011 | AMO Manufacturing USA, LLC | Reports of lower success rates with Iris Registration (IR) image capture with the STAR S4 IR Excimer Laser System v 5.3 that with previous version. | Software design | Abbott Medical Optics (AMO) sent an "ADVISORY NOTICE" dated July 18, 2011 to all customers using the AMO WAVESCAN Wavefront System. The letter identifies the product, problem, and actions to be taken by the customers.The letter instructs customers to create a separate patient file for each eye on the WAVESCAN System (STAR S4 IR Excimer Laser Systems,software versions 5.00-5.30) according to the directions in attachment A. Additionally, customers are to complete and return an attached form to AMO via fax at the regional fax number.Contact an AMO representative for questions regarding this notice. |  | 673 | 673 total v 5.00 - 5.22 | Worldwide Distribution-USA (nationwide) including Puerto Rico and countries of Africa, Asian Pacific, Canada, Europe, Japan, Latin America, and the Middle East. | Terminated | September 10, 2013 | 785 | 673 | Other |  | Instructions | Safety Notice/Insructions
Z-2964-2011 | 59426 | No | HJB | Corneal radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lenstar LS 900 EyeSuite Biometry version i2.000, Four (4) Protective Cardboard Containers Per Each System. |  |  | June 20, 2011 | 2011 | Haag-Streit USA Inc | A patient record may be linked to an incorrect/inconsistent set of calculation results under specific circumstances when using EyeSuite Software version i2.000. | Software change control | Haag-Streit USA, Inc. sent an "Urgent Field Safety Notices" dated June 20, 2011 to all affected customers.The letter included a description of the affected product and problem. The firm instructed their customers to make sure all new and current templates are saved and upgrade to EyeSuite version i4.000. Customers are asked to complete and return the Software Update Verification Form.For additional information contact the Haag-Streit at (513) 398-3937. |  | 216 | 216 | Nationwide (USA) Distribution. | Terminated | March 08, 2013 | 627 | 216 | Software | Output/Calculation | Software update | Software Update
Z-2963-2011 | 59349 | [6 empty] | Heat Activated Niti Arches, Natural, Upper, .019x 25", Arch Wires. |  |  | July 12, 2011 | 2011 | Highland Metals | Heat Activated Niti, Natural, Upper, .019 x 25" Arch wires mixed with size 021x0.25" Arch Wires. | Mixed-up of materials/components | Highland Metals, Inc. notified their affected customers by e-mail and followed with an "Important Alert" letter dated July 12, 2011.The letters informed customers of the mix up of wires. Customers were asked to contact Highland Metals at 800-368-6484 to schedule a call tag to return any remaining product. Upon receipt credit will be issued.For additional information contact Highland Metals at (800) 448-7971 or (408) 271-2955. |  | 1000 | 10 packages of 100 | Worldwide Distribution; Nationwide (USA) distribution including the states of NV, CT, and FL; and the countries of Australia and Turkey. | Terminated | September 14, 2011 | 64 | 1000 | Not_Computer
Z-2981-2011 | 55941 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Aplio XG model SSA-790A |  |  | June 02, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) is initiating a field correction on the the Aplio XG model SSA-790A because it did not meet the requirement of clause 15b, IEC60601-1 which requires that the residual voltage shall not exceed 60V after 1 second from taking off the AC plug. | Device Design | Toshiba America Medical Systems (TAMS) sent an "URGENT:MEDICAL DEVICE CORRECTION' letter dated June 2, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers were instructed to be careful not to touch the contacts or prongs on the AC plug within two seconds after unplugging the system from a wall socket. This precaution is to be adhered to until a Toshiba Service Representative replaces the power supply. Customers are to complete and return a Customer Reply Form and fax to 877-349-3054.Customers with questions regarding the notification should contact the Director of Regulatory Affairs at (800) 421-1968 or call their local Toshiba Representative at (800) 521-1968. |  | 78 | 78 | Nationwide Distribution | Terminated | April 20, 2012 | 688 | 78 | Not_Computer
Z-2977-2011 | 59107 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE/RT Therapist v4.2 and/or Oncologist v4.2 systems. |  |  | June 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety rick related to transferring reference images for offset calculation between a syngo RT Therapist v 4.2 or a syngo RT Oncologist 4.2 and other versions of the COHERENCE/syn go RT Therapist or Oncologist or vice versa. An incorrect offset may be applied for patient positioning and result in dose being delivered to the wrong location. | Software design | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION/CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter provides preventive measure instructions for customers to use until an update for the system is available. |  | 121 | 121 devices | Worldwide Distribution-USA (nationwide) including the states of CA, FL, NC, NY, OH, PA, SD, UT, and WI and countries of Australia, Belgium, Croatia, Czech Republic, Germany, India, Ireland, Japan, Poland, Saudi Arabia, South Africa, Spain, Sweden, and United Kingdom. | Terminated | March 30, 2012 | 301 | 121 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2980-2011 | 58918 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia Radiation Therapy Pallet. |  |  | June 22, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | There a potential patient safety issue when using the Symbia Radiation Therapy Pallet. When a patient is not iso-centered before starting a scan with the Symbia T Series Systems and a non-circular orbit is selected, it is possible that one of the detectors may come into contact with the patient. | Device Design | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated June 06, 2011 to all affected customers.The letter informs customers of the potential safety issue if the patient is not iso-centered before starting a scan and a non-circular orbit is selected and provided recommendations including ensure the patient is iso-centered before starting a scan. If the patient cannot be positioned at the iso-center of the RTP, instructions were provided. The customers are asked to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter.For questions or assistance contact Siemens Medical Solutions USA, Inc. at (800) 888-7436 (USA). |  | 18 | 18 pallets | Worldwide Distribution: Nationwide (USA) including the states of Florida and Texas; and the countries of Canada, China, Italy, Mexico, Poland, Portugal and Russia. | Terminated | November 27, 2012 | 524 | 18 | Not_Computer
Z-2978-2011 | 59087 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac, Trilogy. Trilogy Tx and Novalis Tx Linear Accelerators. |  |  | May 26, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | On Varian Linear Accelerators where the customer may calibrate the collimator angle position readout in reverse, switching the 90 degree and 270 degree positions. | Device Design | Varian Medical Systems sent an "Urgent Medical Device Correction-Urgent Field Safety Notice" letter dated May 26, 2011 to all affected customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** The letters identifies the product, describes the problem and provides recommendations. Varian recommends maintenance only be performed by qualified personnel with appropriate training, always verify collimator calibration and anytime a collimator posthion readout calibration is performed, and always verify that the collimator angle is not calibrated in reverse by performing the appropriate checks. For questions contact Varian at (888) 827-4265. |  | 5685 | 5685 | Worldwide Distribution: Nationwide (USA) Distribution and the countries of; Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Belorussia, Bolivia, Bosnia and Herzegoniva, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Rep., Denmark, D.R. Ecuador, Egypt, Estonia, Finland, France, French Poly., Georgia, Germany, Greece, Guam, Guatemala, H.K., Honduras, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Martinique, Mexico, Moldova, Monte Negro, Morocco, Nepal, Netherlands, N.Z., Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, P.R., Qatar, Reunion, Romania, Russia, S.A.,Serbia, Singapore, Slovakia, South Africa, S.K., Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad/Tobago, Tunisia, Turkey, Ukrain, U.A.E., U.K., Uruguay, Venezuela, Vietnam, V.I., and Zimbabwe. | Terminated | July 13, 2012 | 414 | 5685 | Not_Computer
Z-2979-2011 | 59170 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera/Skyra System - Nuclear magnetic resonance imaging |  |  | June 20, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Firm discovered during product monitoring that on a few systems the gradient cable connections did not meet their specifications after installation. Locally, the cable connections overheated, making servicing necessary.The potential exists for a cable connection to become overheated, leading to the emission of smoke and possible material damage. | Nonconforming Material/Component | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" letter dated June 20, 2011 to all affected customers. The letter describes the product, problem, and actions that will be taken.The letter informs customers that an inspection of the gradient cable connections and an emergency shutdown mechanism will be installed by a field service representative during a visit to all affected locations. |  | 26 | 26 | Nationwide Distribution-USA (nationwide) including the states of AK, FL, GA, IL, MA, MI, MN, MT, NC, NJ, NY, OH, TX, UT, VA, WA, and WV. | Terminated | January 14, 2014 | 939 | 26 | Not_Computer
Z-2994-2011 | 58744 | No | LNR |  |  | Gastroenterology/Urology | PMA | XT125 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT; XT001 PHOTOPHERESIS PROCEDURAL KIT FOR USE WITH THE UVAR XTS INSTRUMENT |  |  | March 31, 2011 | 2011 | Therakos, Inc. | Increased complaints were received of tubing leaks at the junction between the tubing and Fluid Logic Module plastic ports of the kits. | Device Design | The firm, Therakos Photopheresis, sent a "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated March 31, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately discontinue using the affected products; deface and discard their remaining inventory of affected procedural kits, following their internal procedures for destroying and discarding of medical waste, and complete and return the attached Confirmation of Receipt Form no later than 08-April-2011 via fax to: Safety, Risk Management and Surveillance at 1-888-557-3759 or 1-585-453-4110. If you have any additional technical questions, please contact Therakos Customer Service Center at 1-800-828-6316. If you have technical questions regarding this communication, please call Therakos Technical Support at 1-877-865-6850. |  | 13239 | 13239 kits | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD,TN, TX, UT, VA, VT, WA, and WI; and countries of: Argentina, Austria, Australia, Belgium, Canada, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Korea, Kingdom of Saudi Arabia, Jordan, Netherlands, Norway, Philippines, Poland, Portugal, Russia, Singapore, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | March 18, 2015 | 1448 | 13239 | Not_Computer
Z-2990-2011 | 59042 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model DV5x Series CPAP with Heated Humidifier |  |  | March 11, 2011 | 2011 | Sunrise Medical HHG Inc. dba DeVilbiss Healthcare | Some units were manufactured with a defective humidifier chamber o-ring. Unit could leak, allowing water to escape during use, causing damage to furniture. | Nonconforming Material/Component | The firm, DeVilbiss Healthcare, issued a "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 11, 2011 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to verify their inventory of CPAPs with Heated Humidifiers and isolate any units that are included in the list; replace the humidifier chamber o-rings with the new ones provided; discard the original o-ring removed from the humidifier chamber; complete and return the URGENT: MEDICAL DEVICE CORRECTION Confirmation letter via mail to Quality Manager, DeVilbiss Healthcare LLC, 100 DeVilbiss Drive, Somerset, PA, 15501; affirm that all units have been corrected and that all end users have been sent replacement parts, when all of the CPAPs on the serial number list have been corrected, and return a copy of the sign serial number sheet as confirmation.If you have any questions regarding the information provided in this communication, please contact DeVilbiss Healthcare Customer Service at 1-800-388-1988 from 8:00AM-5:00PM EST. |  | 210 | 210 | Worldwide distribution; USA (nationwide) including states of AL, AR, CA, FL, IL, IN, IA, KS, KY, MD, MA, MI, MS, NE, NJ, NY, OH, PA, WA, WV, and WI; and countries of Europe. | Terminated |  |  | 210 | Not_Computer
Z-2993-2011 | 59401 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0mm diameter, length 320 mm. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Needle tip may separate from the needle shaft such that the tip may detach in the patient upon withdrawal. | Device Design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated July 13, 2011 to all possibly affected customers. This notification supplements and modifies the previous "Customer Technical Bulletins" originally sent in September 2008. The initial Customer Technical Bulletins informed customers of the possibility that needle tips cans crack and separate. It also reminded users that the product has a limited life and to check needles prior to use as per instructions. Current correction notice letter instruct customers to stop using any plastic needles in inventory with lot numbers C01 to D01 and those with lot numbers starting with H or I. The letter also request that the recalled products be returned to Varian BrachyTherapy for replacement. The letters include information as to where the lot number is located on the device. All appropriate personnel should be notified of the correction notice. Contact information is provided for all geographical areas. USA contact for questions or concerns is 1-800-360-7909. |  | 601 | 601 - all sizes. | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Austria, Bangladesh, Belgium, Brazil, Canada, Columbia, Cyprus, Czech Republic Denmark, Dominican Republic, Estonia, France, Germany, Greece, Hungary, India, Japan, Kazahastan, Kuwait, Lebanon, Macedonia, Mexico, Moldova, Nepal New Zealand, Norway, Pakistan, Philippines, Poland, Russia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain Sweden, Switzerland, Taiwan, Thailand, Trinidad, Tobago, Turkey, Ukraine, United Arab Emirates Uruguay, United Kingdom, and Venezuela. | Terminated | June 20, 2012 | 343 | 601 | Not_Computer
Z-2989-2011 | 59362 | No | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | INTRABEAM Shielding Material Flat 20 x 20 cm, |  |  | July 01, 2011 | 2011 | Carl Zeiss Meditec, Inc. | Carl Zeiss Meditec received a report that the INTRABEAM 20 x 20 cm sterile shield's packaging looked wet inside. Investigation found the substance to be Triethylene Glycol, | Process control | Carl Zeiss Meditec, Inc. sent an "Important Field Correction Action" letter dated July 13, 2011 to all affected customers.The letter included description of affected product and problem. Customers are asked to notify the firm if they have experienced any problems with the product and to return any remaining product for replacement upon completion of required testing. Customers are asked to fill out and return an Acknowledgement Form. For additional information contact Carl Zeiss Meditec at (925) 570-4844. |  | 80 | 8 boxes of 10 sterile packages. | Nationwide (USA) Distribution including the states of FL, CA, NY, VA and AL. | Terminated | October 13, 2011 | 104 | 80 | Not_Computer
Z-2983-2011 | 59347 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | LANTIS Oncology Information System (OIS) |  |  | June 30, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue exists for data transfer between LANTIS OIS system (containing PRIMEVIEW or PRIMEVIEW Data Gateway) and the Sequencer V&R; system via a third party interface is used. | Software design | Siemens sent an "Update Instructions/Customer Safety Advisory Notice" to all affected users, beginning June 30, 2011.The letter includes description of the product and problem and provides updated instructions to ensure data is stored in the system.If you have questions contact Siemens at (925) 602-8157. |  | 59 | 59 devices | Worldwide Distribution: Nationwide (USA) including the states of MI, DE, WI, FL, CA, TN, NJ, PA, NE, IL, WV , MA and VA; and the countries of Poland, France, Germany, Canada, Netherlands, India, Ireland, Egypt, P.R. china, Australia, Saudi Arabia, Turkey, Republic Korea, Belgium, Hungary, New Zealand, Italy, Malaysia, and Spain. | Terminated | March 09, 2012 | 253 | 59 | Other | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-2987-2011 | 59367 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Direct Drive Grasper, Model number C4140, 5MM X 45CM DIRECT DRIVE |  |  | June 17, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed specific lot numbers of the DirectDriveÂż Grasper (model numbers C4130 and C4140) due to a potential weakness in the locking trigger, which could result in the trigger fracturing and the jaw locking in the closed position. | Device Design | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 22, to all affected consignees. The letter described the product, problem and actions to be taken. Consignees were instructed to check their inventory for the recalled products and returned immediately as per the instructions. The letter instructs consignees to complete the attached Voluntary Recall Notification Confirmation Form and fax it to Regulatory Affairs at 949-713-8832 or email it to recall60342148@appliedmedical.com. For questions pertaining to product return contact Chris Jacobson, Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. For regulatory questions contact Susan Fehrenbacher at (949) 713-8041 or by email at sfehrenbacker@appliedmedical.com. |  | 38864 | 38,864 units total (24,868 units in US) | Worldwide Distribution - USA (nationwide) including the state of Puerto Rico and the countries of: Australia, Bahrain, Canada, Chile, Colombia, Ecuador, Europe, India, Israel, Jordon, Kuwait, Lebanon, Libya, Malaysia, Mexico, Peru, Qatar, Singapore, , South Africa, South Korea, Saudi Arabia. Taiwan, Thailand, Turkey, and Venezuela. | Terminated | March 26, 2012 | 283 | 38864 | Not_Computer
Z-2985-2011 | 59221 | [6 empty] | ATLANTIS(R) Anterior Cervical Plate System Screw Caddy; Self Drilling Fixed Screw Caddy (976-516). |  |  | July 06, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Lines on the screw gauge may not placed on the caddy properly which could result in improper measurement of screws. | Process control | Medtronic Sofamor Danek USA, Inc. sent "Voluntary Recall - Medtronic" letters on July, 2011 to all affected customers.The letter included; description of product and problem, and asked customers to stop using the product. Customers are asked to contact their Sales Representatives for replacement.For question on this recall call Medtronics Global Quality Department at (800) 876-3133. |  | 191 | 191 units | Worldwide Distribution: Nationwide distribution including the states of MN, OK, IL, AK, CA, PA, UT, LA, SC, NY, MI, GA, WI, KY, AL, MO, AZ, TN, TX, MA, NE, OH, FL, CO, MT, CT, VA, NC, NH, WA, WV, WY, IN, IA, KS, MT, ME, DE, DC, NJ, MD, MS, NV, HI, ID, ND, and AR; and the country of Uruguay. | Terminated | April 26, 2013 | 660 | 1064 | Not_Computer
Z-2982-2011 | 59386 | [6 empty] | Biopsy Trays (CT, Puncture &amp; Prep), Packaged individually in header bags, 10 per case. |  |  | May 13, 2011 | 2011 | North Coast Medi-Tek Inc | During routine bioburden testing, unusually high results were found for a Bioburden Test done on 4/19/11. Bioburden Level reported as &gt;150,000 CFU. Our Alert Level is 17,000 CFU and the Action Level is 23,000 CFU. | Other | North Coast Medi - Tek Inc. sent a URGENT PRODUCT RECALL letter dated May 13, 2011 to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The customers were instructed to stop shipment of all remaining product and return to recalling firm. They were also requested to contact their customers and notify them of the recall and retieve all remaining product.For further questions please call (440) 974-0750. |  | 2190 | 2,190 trays | Nationwide Distribution CA, CO, FL, GA, ID, IL, MA, NC, NE, NY, OH, OK, PA, TN, NC and TX | Terminated | February 16, 2012 | 279 | 2190 | Not_Computer
Z-2976-2011 | 59296 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper Systems Models: |  |  | July 08, 2011 | 2011 | Philips Healthcare Inc. | System freezes from 20 seconds up to one minute. No fluoroscopy exposure and no geometry movements during this time period. | Software design | Philips Health Care North America sent a FIELD SAFETY CORRECTION ACTION letter dated July 4, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.This FCD will be identified as FCD72200200 and will be free of charge. This FCD consists a CD with the correct firmware and a installation procedure to implement the FCD correctly.Please contact your local Philips representative. For further follow up please call (978 ) 687-1501. |  | 1423 | 1423 units | Nationwide Distribution | Terminated | May 14, 2014 | 1041 | 1423 | Software | Device Operation | Firmware update | Software Update
Z-3001-2011 | 58322 | [6 empty] | Micro-Fuel Cell Oxygen Sensor, Class RI5 |  |  | November 29, 2010 | 2010 | Teledyne Instruments - Analytical Instruments | The recall was initiated because during a test on a batch of sensors on the company shipping dock, Teledyne Instruments discovered electrolyte loss from the sealing cylinder on a few sensors. This test was done pursuant to a complaint from one customer of label discoloration caused because of the electrolyte loss. The riskof a leak form the oxygen sensor while in use is the disruption of therape | Process design | Teledyne Instruments sent emails and "Urgent--Device Field Corrections/Removal" letters in December 2010 to all affected customers.Th letter includes description of device and problem. Customers are advised to check them before use for any signs of electrolyte loss and if discoloration is noted to discontinue its use. and provides recommendation. Teledyne-desplaines to generate a product field corrections/removal log that lists names/addresses of the consignees, quantity of sensors shipped, sensor serial numbers and the date of shipment.For additional information or questions contact Teledyne Instruments at (626) 934-1504. |  | 2813 | 2813 | Worldwide Distribution: Nationwide (US) distribution including the states of CA, TN and WI; and the countries of , Canada, Brazil, Chile, Saudi Arabia, and the U.K. | Terminated | August 22, 2011 | 266 | 2813 | Not_Computer
Z-3011-2011 | 59112 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker EMS Stair - Pro Evacuation Chair Model 6252 |  |  | June 28, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | This recall was initiated as the firm ascertained the supplier may have reworked some of the red release bars used during the time frame of 22 December 2010 through 31 January 2011. The rework subjected the parts to excessive heat and as a result, the bars can bend during the action of squeezing the release bar against the black cross tube to release the track. The red release bars may bend to | Nonconforming Material/Component | The firm, Stryker Medical, sent an "Urgent-Medical Device Recall" letter dated June 28, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate the listed product; inspect each unit prior to use; return the enclosed post card to confirm receipt of this notification; if they have loaned or sold any of the chairs listed, forward a copy this letter to the new users and advise Stryker of new users names and locations, and if they have disposed of any of the affected chairs and they are no longer in use, advise Stryker of their obsolescence by providing them with the serial numbers. Stryker stated in the letter that, their goal is the timely upgrade of all affected units within the next 90 days.If you have any questions or concerns, please contact us at 800-STRYKER, option 2. Please reference RA-2011-027. Normal business hours are Monday-Friday 8am-5pm EST. |  | 214 | 214 | Worldwide distribution: USA (nationwide) and countries of: Canada, China, France, Germany, India, Netherlands, and Sweden. | Terminated | August 29, 2013 | 793 | 434 | Not_Computer
Z-3003-2011 | 59122 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE Linear Accelerator with RT Therapist v4.1 system. |  |  | June 03, 2011 | 2011 | Siemens Medical Solutions USA, Inc | A potential safety issue was discovered when a pause for an imaging segment is inserted (where an imaging segment does not exist) with the ARTISTE and RT Therapist 4.1 system will continue to the next image. | Software design | Siemens sent an "Urgent: Medical Device Correction/ Customer Safety Advisory Notice" letter to all affected customers starting June 3, 2011. The letter describes the product, problem, and instructions to prevent mistreatment. The letter states that Siemens is working on an updated version of RTT 4.1 that will fix the issues. Users will be notified as soon as the update for the system is available. The customers were instructed to include Customer Safety Advisory notice in their syngo RT Therapist 4.1 System Owner Manual chapter "Safety Advisory Letters" where it should remain. |  | 223 | 223 active devices | Worldwide Distribution - USA (nationwide) including the states of: DE, FL, KY, LA, MA, MI, MO, NE, NJ, OH, PA, TN, WI, UT, WV and Puerto Rico and the countries of: Australia, Belgium, Brazil, Canada, China, Egypt, France, Germany, Hungary, India, Ireland, Italy, Lebanon, Malaysia, Netherlands, New Zealand, Philippines, Poland, Republic of Korea, Russian Fed, Saudi Arabia, South Africa, Norway, Spain, Turkey and United Kingdom. | Terminated | February 17, 2012 | 259 | 223 | Software | Device Operation | Software update | Software Update
Z-2996-2011 | 59404 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Titanium Fletcher-style Applicator Set - Defined Geometry - CT Compatible with Titanium intrauterine probe with stopper length 60 mm, 30 degree angle. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential of weld failure of the Titanium intrauterine probe at the cervical stopper. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to return all affected product to Varian Brachy Therapy by using the contact information provided in the letter. Replacement product will be provided free of charge to the customers.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 437 | 437 - all sizes | Worldwide Distribution. | Terminated | June 20, 2012 | 343 | 437 | Not_Computer
Z-3002-2011 | 59378 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System; |  |  | June 30, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In affected Eclipse versions, the treatment field energy may change from the current selection to a previous selection without use intervention and without invalidating the MU or calculated dose distribution. | Software design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 30, 2011 to all affected consignees. The letter identifies the affected product, problem, and actions to be taken. The letter states that Varian will contact all affected customers to apply a database configuration script to the database and is preparing a correction for Eclipse Version 10 to address the issue. A Varian representative will contact each site to schedule correction when it is available. The letter also provides RECOMMENDED USER ACTIONS to follow. The Letters go on to reiterate caution and warnings from Eclipse instructions for use. All appropriate personnel should be notified of the correction notice. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. For further clarification, contact your local Varian Customer Support District or Regional Manager. |  | 2436 | 2436 | Worldwide Distribution. | Terminated | May 09, 2014 | 1044 | 2436 | Software | Treatment/Delivery/Therapy | Upgrade | Repair
Z-3008-2011 | 59445 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Deerfield, IL; product code 2M8161 |  |  | July 22, 2011 | 2011 | Baxter Healthcare Corp. | A review of Baxter service records indicates that the Baxter Colleague CX Infusion Pump has not met certain voltage specifications. This issue could potentially cause a delay in initiation of therapy. | Other | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump, and promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com.If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time. |  | 1 | 1 pump | Nationwide distribution: Missouri only. | Terminated | October 19, 2011 | 89 | 1 | Hardware | Treatment/Delivery/Therapy | Return | Remove or Replace
Z-3007-2011 | 59448 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pump. Baxter Healthcare Corporation, Deerfield, IL. Product Code: 2M8161. |  |  | July 22, 2011 | 2011 | Baxter Healthcare Corp. | The Colleague infusion pump may not have met certain Air in Line (AIL) specifications after testing by the Baxter Global Technical Service Center. This issue could potentially cause a delay in initiation of therapy. | Process design | The firm, Baxter, called by telephone on July 22, 2011, and sent a follow-up "Urgent Product Recall" letter dated July 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and take the affected pump out of service; call Baxter Global Technical Services at 1-800-THE-PUMP (1-800-843-7867), Option 1, between the hours of 6:00 AM and 7:00 PM Central Time, Monday through Friday, if arrangements have not yet been made to return the pump; promptly complete and return the CUSTOMER REPLY FORM via fax to 847-270-5457, or scan and email to fca@baxter.com, and forward this communication to the pump owner as appropriate, if they distributed the pump to another facility or if they are dealers, wholesalers, or distributors/resellers.The letter states "Upon receipt of the returned device, Baxter will perform required testing on the affected COLLEAGUE pump for which you have indicated there is continuing medical necessity. Should Baxter find any issues during testing, Baxter will address those issues through repair or service prior to returning the pump to you at no charge".If you have any questions regarding this communication, please call the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm Central Time. |  | 1 | 1 Pump | Nationwide distribution: MN only. | Terminated | November 09, 2011 | 110 | 1 | Not_Computer
Z-3009-2011 | 59275 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis U mobile x-ray system |  |  | June 29, 2011 | 2011 | Siemens Medical Solutions USA, Inc | When the heat unit of the angiography system reaches or exceeds 65%, the cool down time display will appear with an acoustic signal. If users ignore the warning and continue the radiation, the tube may be autoprotected when it reaches its temperature limitation causing the X-Ray to lock until the tube cools down. Following a tube auto-protection, an error message "Door is Open" or "Err 421/240: | Labeling False and Misleading | Siemens sent a "Customer Safety Advisory Notice" letters to all affected users dated June 13, 2011. The letter identified the product, the problem, and the actions to be taken. Siemens included in the notice an addendum to the instruction for use to each system and request customers to take special attention. Customers were instructed to place the Safety Notice in the System's instruction for use. All appropriate personnel should be notified of the correction notice. If product was distributed further, customers were to forward the safety notice to the new owner and inform the firm of the new owner. |  | 18 | 18 | Nationwide Distribution (USA) including the states of: CA, CT, FL, KS, MN, MO, NJ, NY, NC, PA, and UT. | Terminated | January 24, 2014 | 940 | 18 | Not_Computer
Z-3005-2011 | 59416 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Surgical Clamp Inserts, Model#G-8650, 86MM Insert Soft/Fibra 15/Bx |  |  | March 07, 2011 | 2011 | Applied Medical Resources Corp | The recall was initiated because Applied Medical has confirmed that Surgical Clamp Inserts, Model Numbers: G-8600, G-8650, G-8655, has the potential packaging nonconformance which could result in a compromise to the sterile barrier. | Other | The firm, Applied Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 11, 2011 to its customers. The letter provides customers with an explanation of the problem identified and actions to be taken. The customers were instructed to check their inventory; complete and return the attached Voluntary Recall Notification Confirmation Form via fax to Regulatory Affairs at 949-713-8832, and return the recalled Standard 86mm Inserts immediately to Applied Medical, Attn: RGA# 60308783, 22872 Avenida Empresa, Rancho Santa Margarita, CA 92688.Customers with product return questions, were instructed to contact the Sales Operations Administrator, at (949) 713-8652 or by email at cjacobson@appliedmedical.com. Customers with regulatory questions, were instructed to contact the Director, Regulatory Affairs, at (949) 713-8041 or by e-mail at sfehrenbacher@appliedmedical.com. |  | 20247 | 20,247 units total | Worldwide distribution: USA (nationwide) including states of: Arizona, California, Colorado, Florida, Georgia, Hawaii, Iowa, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, and Michigan; and countries of: Japan, Israel, and Singapore. | Terminated | March 26, 2012 | 385 | 20247 | Not_Computer
Z-3018-2011 | 57583 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | i-CAT Classic (X-Ray Tube head and Power Supply) |  |  | December 15, 2010 | 2010 | Imaging Sciences International, LLC | Recalling firm received a customer complaint regarding a defective X-ray tube head and power supply | Other | The firm, Imaging Sciences, sent an "Urgent Medical Device Field Correction" letter dated December 15, 2011 via first class USPS to its customer. The letter noted the product and described the problem and action to be taken. The customer was instructed to complete and return the ACKNOWLEDGEMENT FORM via fax to 215-997-5665 or email: darin.oppenheimer@imagingsciences.com. Imaging Sciences stated in the letter that an X-ray tube head and power supply will be provided and installed by one of their technicians. Should you have any questions regarding this information, call 215-622-1545. |  | 1 | 1 | X-ray tube and power supply were distributed to one consignee: Dr. Alan D'Allessandro, 2500 W. Higgins Road, Suite 665, Hoffman Estates, IL 60169 | Terminated | March 01, 2013 | 807 | 1 | Not_Computer
Z-3014-2011 | 59463 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Metha(R) Hip System Cobalt Chrome (CoCr) Metha Modular Neck Adapters used with the Metha Modular Short Hip Stems. |  |  | July 19, 2011 | 2011 | Aesculap Inc | Aesculap Implant Systems received a notification from the manufacturer, Aesculap, AG of a fractured CoCr neck adapter. | Unknown/Undetermined by firm | AESCULAP Implant Systems sent an "Important Recall Notification" letter dated to all affected customers.The informed customers of a complaint received where fractured CoCR adapter was found and to return any affected product to the firm. The reason for the fracture is currently under investigation. While the firm is trying to determine the root cause of the fracture, they will be offering the Metha non-modular short hip implants to their customers. Customers are asked to return an Inventory Sheet to AESCULAP.For additional information on this recall call (610) 984-9239 or (610) 984-9300. |  | 149 | 149 | Nationwide (USA) Distribution including the states of AL, FL, GA, MD, NM, NY, OK, and TN. | Terminated | March 19, 2013 | 609 | 149 | Not_Computer
Z-3012-2011 | 59059 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration, GE Healthcare 3000 North Grandview, Waukesha, WI 53188 |  |  | May 31, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that patient mismatch could occur on some Seno Advantage 2.1 and 2.2 Workstations with RIS CCOW integration. This issue may impact patient safety.When Seno Advantage workstation is synchronized with RIS using CCOW (an HL7 standard protocol designed to enable disparate applications to synchronize in real-time), the following issues could occur: 1.Two different | Software design | GE Healthcare sent a URGENT MEDICAL DEVICE CORRECTION letter to all affected customers dated May 31, 2011. The letter identified the product , the problem and the action needed to be taken by the customer.GE Healthcare will upgrade all affected workstations at no cost. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative |  | 391 | 391 | Worldwide Distribution--USA (nationwide) including the states of: AZ, CA, CO, CT, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NY, NC, OK, OR, PA, SC, TN,TX, and WI, and the countries of UNITED KINGDOM, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, PORTUGAL NORWAY, NEW ZEALAND, LUXEMBOURG, LEBANON, KOREA, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, DENMARK, CANADA, BELGIUM AUSTRIA, AUSTRALIA FINLAND. | Terminated | June 29, 2012 | 395 | 391 | Other |  |  | Other
Z-3017-2011 | 59517 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas 2020 Insulin Infusion Pump. |  |  | July 29, 2011 | 2011 | Animas Corporation | Incorrect component (1nF capacitor) was used to produce 191 printed circuit boards used in six pump lots of which 380 pumps were distributed. This issue may result in the pump displaying a "Replace Battery" alarm not related to the condition of the battery. | Mixed-up of materials/components | Animas Corporation sent an "Urgent: Insulin Pump Recall" letter dated July, 29 20011 to all affected customers.The letter included affected product and problem. They requested anyone who received the pump to send it back to Animas and a replacement pump will be sent at no charge.For additional information contact the Customer Support Department at (877) 937-7867. |  | 380 | 380 | Worldwide Distribution: Nationwide (USA) distribution including the states of CA, CO, MA, and MD; and the countries of Australia, Finland, France, Israel, Italy, Sweden and the U.K. | Terminated | December 07, 2011 | 131 | 380 | Hardware | Alarm/Message | Replace device | Remove or Replace
Z-3013-2011 | 59388 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Flexible Applicator probe; 3.2mm diameter, braced PVDF. |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Certain GammaMed Flexible applicator probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave. | Process control | Varian Medical Systems, Inc. sent out Customer Technical bulletins in August 2008 and Urgent Medical Device Correction/Urgent Field Safety Notices on July 13, 2011 to all affected customers.The letter included description of device and problem. The current notice supplements and modifies previous instructions and ask that they stop using any flexible applicator probes in inventory. They are to be returned to Varian Brachy Therapy for replacement. Letter go on to list possible scenarios if the length deviation is not detected and advised all personnel working in the radiotherapy department should be notified of the recall. For additional information contact (800) 360-7909. |  | 1343 | 1343 | Worldwide Distribution: Nationwide (USA) and including the countries of France, Germany, India, Trinidad and Tobago, United Kingdom, Canada, Nepal, Ukraine, Czech Republic, Bangladesh, Thailand, Turkey, Brazil, Estonia, Switzerland, Russia, Greece, Philippines, Spain, Belgium, Kuwait, Romania, Morocco, Mexico, Columbia, Austria, Poland, Kazahstan, Vietnam, Cyprus, New Zealand, Dominican Republic, Norway, Uzebestan, Hungary, Lebanon, South Africa, South Korea, Slovenia, Venezuela, Serbia, Tunisia, Pakistan, Macedonia, Ireland, Azerbajian, Ukraine, Morroco, Italy, Singapore, Nambia, Turkey, Sudan, and Taiwan. | Terminated | June 20, 2012 | 343 | 1343 | Not_Computer
Z-3016-2011 | 59376 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | AirLife Nebulizer Heater, Catalog Code 2M8021 |  |  | July 13, 2011 | 2011 | Carefusion 211 Inc | The recall was initiated because CareFusion's Investigation has confirmed the Nebulizer Heater can, over time, develop a failure mode that results in an internal electrical short. The short can result in smoke and sparks being emitted from the device. In the presence of oxygen enriched environments and a fuel source, this failure represents a potentialrisk of flame. CareFusion is voluntarily rec | Other | Carefusion sent a URGENT PRODUCT RECALL letter dated July 15, 2011, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to examine their inventory for the affected product as indicated on the sample shown in the enclosed Attachment A-1 and remove any affected heaters present in their facility. Confirm the total number of affected units in their possession and enter the quantity and respective serial number on the enclosed Customer Response Card (Attachment A-2). Customers were instructed to contact CareFusion's Technical Support Department at 800-554-8933, (6:30 am to 5:00 pm PDST), to obtain an RMA (returned materials authorization) number for the return of the affected product. Upon receipt of their response, CareFusion will ship to their customers new replacement Nebulizer Heaters in exchange for the affected units.Customers were instructed to acknowledge receipt of this communication by completing the enclosed Customer Response Card and faxing or sending a copy by return mail to the address indicated on the Customer Response Card, (Attachment A-2). |  | 22670 | 22,670 units total | Nationwide Distribution | Terminated | April 30, 2012 | 292 | 22670 | Hardware | Physical Safety Hazards | Remove | Remove or Replace
Z-3015-2011 | 59398 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm, Models: FSIII, FSIIIC and Officemate |  |  | July 21, 2011 | 2011 | Hologic, Inc. | The Flex arm connector to C-arm can break and the C-arm that is supported by the flex arm may fall down | Nonconforming Material/Component | The firm, Hologic, Inc., sent an "Urgent: Medical Device Recall" letter dated July 25, 2011 via FedEx to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to visually check the the Flex arm/C-arm on their unit for any sign of damage or cracking, if any damage or cracks are observed, remove the unit from service until they receive the retrofit kit; and to complete and return the Confirmation of Notification form via fax to Hologic at 781-280-0662 or mail using the enclosed pre-addressed, pre-paid return envelope. After the customer confirms that they have the unit in their possession, Hologic will expedite the user installable repair kit. The kits will be available starting the first week of August. The kits will also be dispatched via FedEx. Confirmation of installation will be requested for all dispatched kits.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. |  | 540 | 540 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Spain, and Viet Nam. | Terminated | June 19, 2012 | 334 | 540 | Not Computer
Z-3019-2011 | 57543 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | CIVCO Pro-Lok Platform, REF MTSBRT002. Firm name on label is CIVCO, Orange City, Iowa |  |  | December 30, 2010 | 2010 | Civco Medical Instruments Inc | Systems sold before December 2009 displayed a scale of 0-40cm going from left to right, while systems after December 2009 reversed this and were displayed as 40-0cm from left to right. | Nonconforming Material/Component | CIVCO issued a "FIELD SAFETY NOTIFICATION" letter dated 12/20/2010 beginning 12/30/2010, via FedEx or UPS to distributors and customers. The notification identifies the affected device, problem, and field safety corrective action. The letter states that if customers have more than one system purchased from different time periods that they should examine their systems and assess if there is any mismatch. The recalling firm will rework the system to correct the difference and send it back to the customer. To take advantage of this correction customers are advised to contact the recalling firm within thirty days of receiving the safety notification. The letter includes a response form to be completed and returned as per instructions, to indicate that the notice has been read and understood by all operators of the equipment. The letter to distributors requested that a copy of the notification letter be forward to all customers who have received this component. For assistance contact your CIVCO insides sales representative by calling CIVCO Medical Solutions, Radiation Oncology, at 800-842-8688. |  | 20 | 20 platforms | Worldwide Distribution - USA ( nationwide) including the state of PA. and countries of: Canada, China, England, Israel, Italy, Japan, Korea, The Netherlands, Russia, Spain, Sweden, Switzerland and Turkey. | Terminated | June 28, 2013 | 911 | 134 | Not_Computer
Z-3024-2011 | 56839 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical elance Vital Signs Monitor, Model 93300. |  |  | September 23, 2010 | 2010 | Spacelabs Healthcare, Llc | The monitor has a 3 second delay in the waveforms and respiration alarm; the delay is not disclosed in the IFU. | Error in labeling | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated September 24, 2010 to all affected customers. The letter describes the product, problem, and the action to be taken.The letter includes the revised Operator's Manual as a replacement for customers current manual.For Technical Support call Spacelabs Medical at 1-800-522-7025 and select 2. |  | 2517 | 2517 | Worldwide Distribution-USA (Puerto Rico) and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, KENYA, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MACEDONIA, MALAYSIA, NEPAL, NETHERLANDS, NICARAGUA, OMAN, PALESTINE, PANAMA, PARAGUAY, PERU, PUERTO RICO, PHILIPPINES, QATAR, ROMANIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | August 17, 2011 | 328 | 2517 | Not_Computer
Z-3023-2011 | 59235 | No | CGI | Estriol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AccessÂż Immunoassay Systems |  |  | December 08, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that a single lot of Access Unconjugated Estriol reagent (PIN 33570) may produce suppressed patient results and an increased incidence of IND (Indeterminate) flags. These issues have been associated with reagent lot 913348. This reagent met all testing specifications prior to release. | Other | Beckman Coulter sent a PRODUCT CORRECTIVE ACTION (PCA) letter dated December 8, 2009, to all affected customers. The letter was sent via US Postal Service for all US customers with attached PCA response form. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to discontinue use of Access Unconjugated Estriol reagent lot number 913348. " No further orders of Access Unconjugated Estriol reagent lot number 913348 will be shipped." Abnormally low Access Unconjugated Estriol results from samples tested previously with the affected lot should be reviewed in the context of other diagnostic tests." Access Unconjugated Estriol results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Rerun any tests that produce a questionable result.In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any technical questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at (800) 854-3633 in the US and Canada. Outside of Canada and the US, customers were instructed to contact their local Beckman Coulter Representative. |  | 2786 | 2786 units total (2776 in the US) | Worldwide Distribution -- USA (nationwide) including the states of CA, CT, MA, NC, NM and TX and the country of Canada. | Terminated | April 16, 2012 | 860 | 2786 | Not_Computer
Z-3025-2011 | 59661 | No | HDC | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GYNDisposables Inc, IUD INSERTION KIT 935K, STERILE, SINGLE USE, Tenaculum 356T, Uterine Sound 906S, Scissors, Povidone-Iodine Solution, (3) Swabsticks, Paper Towel, STERILE EO, REF 935K, QTY 1, |  |  | July 29, 2011 | 2011 | GYN Disposables Inc | Lack of sterility assurance. | Process design | GYN Disposables Inc. sent out an "Urgent: Medical Device Recall" letter dated July 29, 2011 to all affected customersThe letter includes: product affected, description of problem and asks customers quarantine all affected product. They can contact GYN Disposables to have the product picked up and to receive credit. A Response Form is also included.For questions on this recall call (888) 372-2203, extension 101. |  | 3987 | 3987 units | Nationwide Distribution. | Terminated | October 29, 2012 | 458 | 3987 | Not_Computer
Z-3022-2011 | 59394 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, 4D Structure Merge Tool Option in Eclipse Release 8.5 -10 (Client Build 8.2.2x through 10.0.34) with 4D Planning; Model Number: H48; |  |  | July 07, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified in the 4D Structure Merge option of certain versions of Varian's Eclipse Treatment Planning System. The treatment field energy may change from the current selection to a previous selection without user intervention and without invalidating the MU or calculated dose distribution. | Software design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/FIELD SAFETY NOTICE" letter dated July 7, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to cease using the 4D Structure Merge tool option in the Eclipse until further notice. Additionally, it is recommended that customers ensure that a qualified physician reviews the accuracy and placement of all patient structures used for plan development and evaluation and to conduct standard quality assurance procedures prior to patient treatment. Customers will be contacted by a Customer Service representative when the technical correction for this issue is available.Contact your local Varian Customer Support District or Regional Manager if you have questions regarding this notice. |  |  |  | Worldwide Distribution | Terminated | April 04, 2014 | 1002 |  | Computer
Z-3028-2011 | 59218 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | BD Vacutainer Push Button Blood Collection Set; REF 367324 Rx only, 23 G x 3/4" x 12 " (0.6 x 19 mm x 305 mm); |  |  | April 22, 2011 | 2011 | Becton Dickinson &amp; Company | BD Vacutainer Push Button Blood Collection Set can potentially undergo cannula separation from the hub. | Device Design | BD Diagnostics sent a "Lot Specific Product Recall" letter dated April 22, 2011 to all affected customers.The letter asked customers return any affected product and provides instructions for its return. The letter ask customers return the attached Packaging List with the shipment.For question on this recall contact BD Diagnostics at (201) 847-5117. |  | 1744000 | 1,744,000 | Nationwide (USA) Distribution. | Terminated | June 06, 2012 | 411 | 1778960 | Not_Computer
Z-3027-2011 | 59239 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AccessÂż Immunoassay Systems Total T4 Calibrators, |  |  | August 26, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the stability claims for the Access Total T4 calibrators may not be met. This issue is associated with calibrator lots 917074 and 008534. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 13, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue use of the affected lot numbers and review historical quality control performance. Additionally, a response form was attached to the letter for customers to complete and return. Alternate lots will be provided for current and future replenishment orders or requests.Customers with questions regarding this notification can call Technical Support at (800) 854-3633 or contact their local Beckman Coulter representative. |  | 2495 | 2495 units total (2492 in US) | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | July 11, 2012 | 685 | 2495 | Not_Computer
Z-3026-2011 | 59408 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | AlterG Anti-Gravity Treadmill; |  |  | July 15, 2011 | 2011 | Alter-G, Incorporated | Potential installation error on AlterG Anti-Gravity Treadmill that could cause greater than expected unweighting. | Device Design | Alter-G, Inc. notified all affected customers by telephone, e-mail and followed by a "Medical Device Customer Advisory Notification" letter dated July 15, 2011. Customers were provided information on the installation error, the potential harm to patients, and the actions to be taken. The letter instructs customers to discontinue use of the Alter G until the recalling firm can evaluate the machine and correct any installation error. The letter states that the recalling firm is planning to visit affected customer's facility as quickly as possible and will contact customers to schedule this visit. Questions or concerns should be directed to David Tjon at 510-270-5369. |  | 16 | 16 potentially affected units | Nationwide Distribution (USA ) - including the states of: OK and TX. | Terminated | August 16, 2011 | 32 | 16 | Not_Computer
Z-3032-2011 | 59456 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Patient Monitor is a Physiological, Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367, 91369, 91370, 91387, &amp; 91393 are sold with with Perioperative Mode (option D), which includes Start case/End case functionality. |  |  | July 08, 2011 | 2011 | Spacelabs Healthcare, Llc | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally. | Device Design | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated July 25, 2011 along with the Perioperative section of the Operations Manual to all consignees in the US. This letter includes recalled model 91369, 91370, 91387, &amp; 91393 (option D). On July 27, 2011, Spacelabs sent out an updated recall notification letter to include one additional model 91367 to all consignees in the US. The letter identifies the products, problem, and actions to be taken. On August 2, 2011, a recall notification letter (translated as necessary) and the Perioperative section of the Operations Manual was emailed to all international subsidiaries and distributors of record. The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular, they should know how to confirm the state of the alarm tones prior to starting a case. The letter states that the firm is in the process of upgrading the software and will be contacting customers to update monitors at no cost. If customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 2812 | total 2812 units distributed worldwide | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, Bahrain, Bolivia, Brazil, Canada, Caymen Island, Chile, China, Colombia, Cyprus, Ecuador, Finland, France, Germany, Guam, India, Indonesia, Israel, Italy, Jordan, Mexico, Morocco, New Zealand, Nicaragua, Pakistan, Paraguay, Peru, Philippines, Poland, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, Uruguay, and Venezuela. | Terminated | May 23, 2012 | 320 | 2812 | Software | Alarm/Message | Software update | Software Update
Z-3030-2011 | 59488 | No | JPD | Hemoglobin A[bdi2] assay. | Hematology | Hematology | 510(k) | Hemoglobin F &amp; A2 Calibrator, |  |  | May 25, 2011 | 2011 | Tosoh Bioscience Inc | The value after reconstitution appeared to be lower than the assigned value on the label. The measured value for Lot BCAL05R was 5.3%, whereas the value on the label is 5.7%. The criteria of the value is +/- 0.3%. | Process control | TOSOH Bioscience sent an "Urgent Recall Notification" letter dated July, 25, 2011 to all affected customers.The letter described the problem and asks customers to immediately discontinue use of the product and return all remaining product back to the firm. Customers can contact TOSOH for replacement.For additional information or question on this recall please contact TOSOH Technical Support at (800) 248-6769. |  | 24 | 6 boxes/4/2 mL btls | Worldwide Distribution: Nationwide (USA) distribution including the states of FL and IL; and the country of Canada. | Terminated | February 24, 2012 | 275 | 24 | Not_Computer
Z-3036-2011 | 56806 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Ultraview Spacelabs Compact monitor, models 91367, 91369, and 91370, with touch screen display and wireless option. |  |  | September 17, 2010 | 2010 | Spacelabs Healthcare, Llc | Monitor may reset repeatedly which can result in a loss of patient monitoring for several minutes. Cause is a software problem. | Software design | Spacelabs Medical sent an "URGENT - MEDICAL DEVICE CORRECTION" notification letter to U.S. consignees by mail on September 17, 2010 and to international consignees by e-mail on September 24, 2010. The letter identified the recalled product, warns customers of the potential hazard and identifies some ways to mitigate the risk if they decide to continue use. The letter states to correct this situation, the recalling firm Field Service Engineers will be contact customers to schedule a time to install a free software update. If customers have any questions about this corrective action program, please contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support, or outside the United States 1-425-657-7200 x5089. |  | 579 | 579 total (517 to US acc'ts; 62 to international acct's. | Worldwide Distribution - USA (nationwide) and the countries of: International distribution to: Canada, China, France, Germany, Italy, and the United Kingdom. | Terminated | August 22, 2011 | 339 | 579 | Software | Device Operation | Software update | Software Update
Z-3034-2011 | 58381 | No | MRW |  |  | Orthopedic | 510(k) | Catalog # 9024-00 or LSW-S-3040 for 4.5mm BONE-LOK PLS Implant |  |  | March 15, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to Pull Pins that may disengage from the Facet Screw during the compression step, requiring that compression be completed with a device driver rather than the compression tool.Use of this product may result in less than optimal compression leading to pseudoarthrosis, non-union or facet pain. | Other | Interventional Spine, Inc. sent a recall letter dated March 15, 2011, informing consignees of the product, the problem, and the action to be taken. Consignees were asked to immediately examine their inventory and quarantine product subject to recall. In addition, if they distributed the product further they were asked to notify Jane Metcalf, the Vice President of Quality Assurance at once.Interventional Spine further informed consignees that the recall was being carried out to the individual consignee level. And their assistance was appreciated and necessary to prevent use of this product.They were then asked to complete and return the enclosed response form as soon as possible. Questions were directed to call their regional directors.For questions regarding this recall call 949-472-0006. |  | 407 | 407 units | Nationwide Distribution including CO, CT, FL, KY, LA, MA, MO, NC, NY, OK, OR, TN, TX, VA &amp; WI | Terminated | February 11, 2013 | 699 | 960 | Not_Computer
Z-3035-2011 | 59365 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging versions 11.5, 11.5.1, 11.5.2, 11.6, 11.7 and 11.8. |  |  | February 02, 2011 | 2011 | Mckesson Information Solutions LLC | Images for one patient may appear in the study of another patient. | Other | Mckesson Medical Imaging Group sent a URGENT FIELD SAFETY NOTICE letter dated June 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.On February 2, 2011, the McKesson Medical Imaging company initiated a process for applying an image display software update to the HMI systems at the hospital site that initiated the complaint. They were advised to contact McKesson Customer Support at 1-800-663-2533, to indicate whether they would like to receive the software update. They were advised to distribute the notification to all those within their organization who need to be aware and to alert other parties affiliated with their organization that are affected by this advisory notice. |  | 438 | 438 devices | Worldwide Distribution--USA (nationwide) Australia, Canada, Guam and the United Kingdom. | Terminated | October 04, 2012 | 610 | 438 | Software | Output/Calculation | Software update | Software Update
Z-3038-2011 | 59465 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Duo-Decapolar Electrophysiology Catheter, Manufacturing Part Number D-1171-35-S, Catalog Number D728260RT |  |  | April 29, 2011 | 2011 | Biosense Webster, Inc. | The recall was expanded because between February and April 2011, Biosense Webster, Inc. notified customers that a voluntary Customer Notification had been Initiated due to the inaccessibility of the electronic version of the Instructions for Use (e-IFU) for the "10 ft. Interface Cable, 10 Pin - Shielded Tip Pins to 12 Pin - Blue", Catalog Number 0128709,through the J&J; Gateway (i.e., e-IFU websi | Labeling Change Control | Biosense Webster sent an "URGENT FIELD SAFETY NOTICE" letter dated April 29, 2011 to all affected customers. The letter describes the product, problem, and the action taken by the firm.The letter notifies and advises customers that the updated IFUs for ISMUS Catheter with Auto ID Technology and the Duo-Decapolar catheters have been released and posted on the e-IFU website as of April 22, 2011 and April 21, 2011. Customers were instructed to complete the Acknowledgement Form and return the form to Biosense. Customers with questions or require additionial information were instructed to contact their Biosense Webster representative. |  | 1530 | 1530 units | Nationwide Distribution | Terminated | October 12, 2012 | 532 | 1855 | Computer
Z-3042-2011 | 59525 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Equate Muti-Purpose Solution, For Soft Contact Lenses, Sterile, 12 Fl oz. (355 mL), Distributed by Wal-Mart Stores, Inc., Bentonville, AR 72716 |  |  | June 30, 2011 | 2011 | Bausch &amp; Lomb Inc | Muti-purpose solution bottles, lot GC1107 were packaged in a carton that was marked with an incorrect lot code and expiration date. | Packaging | Bausch &amp; Lomb Inc. notified the affected customer via teleconference and e-mail on June 30, 2011. The e-mail describes the product, problem, and actions to be taken by the customer.The customer was instructed to hold and stop shipments of any inventory of the affected product. An official recall was initiated with the customer on July 8, 2011 via e-mail and teleconference.Customer was asked to review their inventory and store returns, and send any of the affected product to Bausch &amp; Lomb Returns Dept., 130 Commerce Drive, Greenville, SC 29615. |  | 1200 | 1200 bottles | Distribution-in the state of AK only | Terminated | November 16, 2011 | 139 | 1200 | Not_Computer
Z-3041-2011 | 59520 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | AmscoÂż V-PRO 1 and AmscoÂż V-PRO 1 Plus Low Temperature Sterilization Systems Model #MLR, Serial # 033250701 - 031581115. |  |  | August 01, 2011 | 2011 | Steris Corporation | STERIS learned that AmscoÂż V-PRO 1 and V-PRO 1 Plus Low Temperature Sterilization Systems users may experience premature wear of the vacuum pump and filtering assembly components. Users may also experience incomplete aeration of the device's chamber following a failed leak test. | Other | Steris Corporation sent an URGENT VOLUNTARY FIELD CORRECTION NOTICE dated August 5, 2011, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.A STERIS Service Representative will contact your facility to arrange for STERIS to upgrade your V-PRO System(s). You can continue to use your system prior to the installation of the system upgrades. In the interim period should you experience any alarms or aborted cycles, the load mustbe handled in accordance with the AmscoÂż V-PROTM 1 and V-PROTM 1Plus Low Temperature Sterilization Systems Operators Manual and your internal hospital procedures.For further information or if you have questions regarding this matter, please contact Field Service Dispatch at 1-800-333-8828. |  | 663 | 663 units | Worldwide Distribution -- USA (nationwide) including the states of: AK, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY, and countries of: Australia, Bangladesh, Belgium, Belarus, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Egypt, Finland, France, Germany, Hungary, India, Iraq, Ireland, Italy, Japan, Kuwait, Lebanon, Mexico, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, South Korea, Spain, Switzerland, Turkey, Vietnam, United Arab Emirates, and United Kingdom. | Terminated | March 27, 2013 | 604 | 663 | Not_Computer
Z-3039-2011 | 59391 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Calcium Flex Reagent is used on the Dimension clinical chemistry system as an in vitro diagnostic test intended for the quantitive determination of calcium activity in human serum, plasma and urine. |  |  | July 19, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostic has confirmed a high frequency of "Abnormal assay" flags and imprecision. Test results across the assay range may be falsely depressed or elevated, as much as, 1.93 mg/dL at a concentration of 9.26 mg/dL. | Other | An Urgent Field Safety Notice datey July 2011 was issued to all affected customers. Customers are being advised to immediately discontinue use and discard any remaining inventory of Dimension(R) CA Flex(R0 reagent cartridge - lot EA2153. Replacement order for the affected lot is in process, and will be sent to impacted customers free of charge.Customers were instructed to forward the Field Safety Notice to anyone whom they may have distributed the affected product.For questions regarding this recall call 302-631-6299. |  | 2500 | 2500 cartons | Worldwide Distribution to Guam, Puerto Rico, Chile, Costa Rica, Venezuela, South Korea, Vietnam and domestically to AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MN, MI, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, WA, WV, WI, and WY. | Terminated | December 27, 2013 | 892 | 2500 | Not_Computer
Z-3044-2011 | 59375 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TSRH(R) Spinal System SCREW, REF 837-645, Size: 6.5mm x 45mm, Mat'l: TITANIUM ALLOY, Rx only, QTY: 1 EA, Medtronic Sofamor Danek USA. |  |  | July 12, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | Screws are 5 mm longer than the size etched on the screw. | Error in labeling | On 07/12/2011, the recalling firm, Medtronic notified the Sales Managers and Sales Force by conference call and email regarding the imminent notification and follow-up activities. A "VOLUNTARY RECALL" notification letter was sent to Risk Managers and Sales Representatives on 07/13/2011 with a delivery date of 07/14/2011. The letter states to immediately remove the recalled product from use and return to Medtronic. The letter includes a questionnaires form to be completed and returned as per the instructions. If appropriate, personal follow up with Sales Representatives will be made to assure that Risk Managers understand the recall communication and to facilitate return of product and the associated questionnaire documentation. For questions, contact your Medtronic Sales Representative or the Global Quality Department at 800-876-3133, extension 3173. |  | 17 | 17 units | Nationwide Distribution (USA) including the states of: OH, NY, and WI | Terminated | July 20, 2012 | 374 | 17 | Not_Computer
Z-3040-2011 | 59688 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | enFlow IV Fluid Warmer Part number 980105VS; |  |  | June 30, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | enFlow IV Fluid Warmer may have an issue which results in excessive external surface temperature on the underside of the warmer. | Device Design | GE Healthcare sent UMDC Letters/UMDC Confirmation Letters to customers on June 30, 2011. GE Healthcare sent UMDC Letters/UMDC Confirmation Letters on October 24, 2011 reflecting a new recall strategy. |  | 1307 | 1307 units | Worldwide Distribution - USA (nationwide) and the country of Vietnam | Terminated | June 14, 2013 | 715 | 1307 | Not_Computer
Z-3051-2011 | 56251 | No | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | "BIO-RAD Kallestad Mouse Stomach/Kidney 8 Well Slides, Bio-Rad Laboratories, STORE FROZEN" |  |  | June 30, 2010 | 2010 | Bio-Rad Laboratories | KALLESTAD Mouse Stomach/Kidney (MSK) test kits contained Crithidia slides instead of the MSK slides. | Employee error | Bio-Rad Laboratories notified affected customers by phone on 6/30/2010 and followed up with a 2nd call on 7/13/2010.The call informed customers of the problem with the kit and requested they destroy the Crithidia slide packs. Bio-Rad send out replacement slide packs. For information or questions on this recall call Bio-Rad at (425) 498-1709. |  | 7840 | 224 kits were distributed; 35 of those kits were determined to contain the wrong slides. | Worldwide Distribution: (USA) Nationwide distribution including the states of CT, OK; and the country of Puerto Rico. | Terminated | August 22, 2011 | 418 | 7840 | Not_Computer
Z-3049-2011 | 59466 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Churchill Medical Latex-Free IV Start Kit |  |  | July 15, 2011 | 2011 | Churchill Medical Systems, Inc. | Outer Kit labeled expiration date longer than the expiration of the shortest expired product | Employee error | Churchhill Medical Systems, Inc A Vygon Company issued an URGENT MEDICAL DEVICE RECALL letter was via overnight courier mail on July 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The product was requested to be returned. Distributors were provided with instructions to communicate the recall to customers who received the affected product by providing a copy of the notification and a copy of the Recall Acknowledgement and Inventory Return Form. The form should be completed in its entirety, signed and returned to the Distributor. Contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email to customerservice @vygonus.com |  | 130750 | 130,750 units | Nationwide Distribution | Terminated | March 27, 2013 | 621 | 130750 | Not_Computer
Z-3045-2011 | 59713 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VerifyÂż SixCess Indicator Strips, Model #: PCC005, PCC006, PCC007, PCC008, PCC031, PCC033, PCC034, PCC037 &amp; PCC038, Mfr. By: Albert Browne, LTD, Chancery House 190 Waterside Rd. Hamilton Industrial Park, Leicester, Leicester, LE5QZ, United Kingdom. |  |  | March 16, 2011 | 2011 | Steris Corporation | During an inspection the firm was notified that language contained in labeling and promotional literature directs Customers that the products can be used to replace biological indicators and/or release loads containing implants. | Labeling False and Misleading | STERIS sent a Customer Notification letter dated March 16, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review and follow the Challenge Packs' revised instructions for use enclosed as Attachment A (270-4) and Attachment B (275F 3).For any questions regarding this recall customers were instructed to contact their STERIS Account Manager or call 440-392-7421. |  | 2500 | 2,500 (one per customer) | Worldwide Distribution - USA (nationwide) and the countries of Lebanon and Qatar. | Terminated | March 19, 2013 | 734 | 2500 | Not_Computer
Z-3050-2011 | 59329 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew, Tendon Stripper, Slotted, 5.0 mm I.D. x |  |  | July 07, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Shaft tip (casting) does not meet the SNE 416specification, which could result in rusting of the Tendon Strippers | Nonconforming Material/Component | Smith &amp; Nephew sent an Urgent - Product Recall Notification letter dated July 7, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the form included in the letter and return to Hoangthi.le@smith-nephew.com or fax to 508-261-3620. Each affiliate must then notify each facility affected by the recall. To return product and obtain credit or order replacement product customers should contact the Smith &amp; Nephew Endoscopy Returns at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to, endo.andreturns@smith-nephew.com. For questions, contact Hoangthi by phone at 508-337-3731. |  | 570 | 570 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Chile, China, Columbia, Denmark, Dubai, Finland, France, Germany, Greece, Hong Kong, India, Italy, Israel, Japan, Korea, Malaysia, Mexico, Netherland, Norway, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, United Kingdom, and Venezuela | Terminated | July 23, 2015 | 1477 | 570 | Not_Computer
Z-3057-2011 | 59576 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20: |  |  | June 20, 2011 | 2011 | Maquet Inc. | HL 20 Consoles for 4-pump and 5-pump configuration can malfunction if an excessive amount of liquid is used causing circuit board malfunction. | Device Design | Maquet sent an URGENT PRODUCT RECALL NOTIFICATION letter dated June 17, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge receipt of the notification by completing and sending back the attached response form whether or not they have any affected product, fax to 973-807-9210 no later than June 30, 2011. For any questions customers were instructed to contact thier local MAQUET representative. |  | 229 | 229 units | Nationwide Distribution | Terminated | October 16, 2012 | 484 | 229 | Not_Computer
Z-3048-2011 | 59568 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Gamma X XL Portable Patient Monitor. |  |  | July 29, 2011 | 2011 | Draeger Medical Systems, Inc. | Monitor keys may spontaneously become inoperative or active. This may cause a membrane switch panel malfunction, which could result in the monitor discharging a patient automatically. | Device Design | Draeger Medical sent an "Urgent Medical Device Recall Notification" letters dated July 2011 to all affected customers.The letter included product and problem affected. The letter informed the customer a Draeger representative will contact them to upgrade the system. Until the solution is implemented, Draeger recommends the Delta and/or Gamma XXL may be used. If the affected monitors become inoperative or activate without user interaction, Draeger advises removing the monitor from service and contacting a Draeger representative.For questions or for information on this recall please call Draeger Service at (800) 543-5047 (press 1 at the prompt, then 2, then 32349). |  | 29 | 29 | Worldwide Distribution - (USA) Nationwide distribution including the states of AK, FL, LA, MN, MS, MO, NJ, NC, OH, OK, PA, TN, TX, NC, NH, NJ, NY, WV, and WI; and the countries of Australia, Austria, Argentina, Bahrain, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, France, Germany, Georgia, Greece, Hungary, Italy, Ireland, India, Kazakhstan, Latvia, Macedonia, Malaysia, Mexico, Netherlands, Norway, Russian Fed., Romania, Saudi Arabia, Slovakia, Spain, South Africa, Switzerland, United Arab Emirates, United Kingdom, Uruguay, Vietnam, and Yemen. | Terminated | August 16, 2012 | 384 | 61 | Not_Computer
Z-3056-2011 | 59476 | No | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | The Resolution II Clip, 155 cm size and 235 cm size, Boxes of 1, 10, and 20 clips consists of a pre-loaded, radiopaque single use clip on a flexible delivery system. |  |  | July 19, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific has received reports where, following clip deployment, the alignment of the clip and delivery system does not allow for the withdrawal of the delivery system away from the clip, resulting in tissue trauma, increased bleeding and the clip being pulled off tissue as the delivery system is withdrawn. | Other | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL letter dated August 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Medical Device Recall to their customers. Customers were instructed to:Immediately discontinue use of any affected product.Complete and return the Reply Verification Tracking Form by fax to 508-683-5578.Package and return any affected product to:Boston Scientific CorporationDistribution CenterAttn: QA Returns500 Commander Shea BlvdQuincy, MA 02171For questions regarding this recall call 508-683-5578. |  | 1302 | 14, 093 boxes of both 155 cm and 235 cm size clips | Worldwide Distribution - including USA (nationwide) | Terminated | June 22, 2012 | 339 | 1302 | Not_Computer
Z-3055-2011 | 59402 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Flexible Probe with Blocking Washer (part of Segmented Cylinder Set) |  |  | July 13, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Blocking washers on the Flexible Probe with Blocking washers may shift and cause potential for treatment length error. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to locate old design applicator probes and guiding tubes and to cease use of the affected products. Information in the letter included instructions on how to return the products to Varian BrachyTherapy.Customers should contact their local Varian Customer Support District or Regional Manager for questions regarding this notice. |  | 1019 | 1019 total | Worldwide Distribution | Terminated | June 20, 2012 | 343 | 1019 | Not_Computer
Z-3052-2011 | 59226 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 Multileaf Collimator |  |  | June 06, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | When using the HD120 Multi Leaf Collimator for beam shaping, and under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceeds the expected transmission dose. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter describes various jaw position recommendations to be followed when planning and treating dynamic fields on the HD120 MLC. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice. |  | 106 | 173***12-20-12 AMENDED TO: 106 units*** | Worldwide Distribution. | Terminated | December 23, 2011 | 200 | 106 | Not_Computer
Z-3059-2011 | 59524 | No | MJJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Adjustable Vent Valve, TruChoice Alliant Healthcare Products REFERENCE#213-030NS QUANTITY: 100 DESCRIPTION: ADJUSTABLE VENT VALVE, NON-STERILE; and 3T Medical Systems Catalog# 030NS Adjustable One Way Vent Valve, Quantity: 100, NON-STERILE |  |  | July 21, 2011 | 2011 | Alliant Enterprises LLC | The firm is initiating this recall due to the receipt of 2 complaints of malfunction of the joint holding the 2 halves of the vent valve together. | Other | Alliant Healthcare sent an URGENT: MEDICAL DEVICE RECALL letter on 7/21/2011 and 7/25/2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately review the recall notification and quarantine any affected product. If product was further distributed, customers were to forward the letter with the Recall Return Response Form immediately to all affected locations and indicate which locations on the Recall Return Response Form. Customers were also instructed to discontinue use immediately, dispose of the affected product per their normal disposal protocol, and fax the Recall Return Response Form to Alliant Healthcare Products at 269-629-5770.Upon receipt of the Recall Return Response Form, a substitute vent valve (non-adjustable only) will be sent to their location for each affected product for which disposal has been properly documented in writing. For questions call 269-629-0300, ext 226. |  | 14900 | 149 Boxes of 100 | Worldwide Distribution - USA including Michigan, Tennessee, New York, and California and the countries of Chile and Korea | Terminated | April 23, 2012 | 277 | 41475 | Not_Computer
Z-3054-2011 | 59276 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | @ Home Strep Test |  |  | July 06, 2011 | 2011 | Jay Bird Investments Inc | Home Strep Test does not qualify as a Class I device and needs pre-market approval. | PMA | Jay Bird Investments, Inc. sent a notification letter dated July 6, 2011 to customers. The letter describes the product, problem and the actions to be taken. The customers were instructed to notify retailers and remove the product from the shelves. For questions regarding disposition of the product please contact Myron Bouwhuis at 801-721-3528. |  | 7541 | 7,541 | Nationwide Distribution (USA) including the states of OR and UTAH. | Terminated | June 20, 2012 | 350 | 7541 | Not_Computer
Z-3063-2011 | 59523 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | Sysmex Stromatolyser-IM; a lytic reagent for use with the XE-2100, XE-2100L and XE-5000 Series Sysmex Automated Hematology Analyzers; the reagent solution is supplied in a 10-liter cubetainer (cardboard cube with a 10-liter plastic liner); Sysmex America, Inc., Mundelein, IL 60060; catalog number SIM-220A. |  |  | April 11, 2011 | 2011 | Sysmex America, Inc. | The affected lots of Stromatolyser-IM lysing reagent were found to be contaminated with Pseudomonas bacteria, which resulted in rising IMI channel background counts, increased false "Platelet Clump" flags and a strong sulfur-like odor when used on the XE-2100, XE-2100L and XE-5000 analyzers. | Process change control | Sysmex America, Inc. sent a "PRODUCT NOTIFICATION" letter dated April 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to discontinue use of the affected SIM lots, and to remove affected lots from their analyzers as soon as possible to avoid long-term performance problems on the analyzer. A Stromatolyser -IM (220-A) Voluntary Recall form was attached for customers to complete and return via fax to Sysmex Customer Service at 800-689-8296. The form will be used to place orders to replace the affected product. Any questions concerning the information contained in the notification are directed to the Sysmex Technical Assistance Center at 1-888-879-7639. |  | 1887 | 1,887 cubetainers | Worldwide Distribution--USA (nationwide and Puerto Rico) and the countries of Canada, Colombia, Ecuador, Mexico, Netherland Antilles and Peru. | Terminated | March 14, 2012 | 338 | 1887 | Not_Computer
Z-3053-2011 | 59693 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Katzen" Infusion Wire, 0.035 in x 180 cm, |  |  | August 01, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is voluntarily recalling one lot/batch of Katzen" Infusion Wires because we have identified that the label on the carton and the inner pouch indicated an incorrect device length. The labels indicate a device length of 180cm while the actual device length is 145cm. | Labeling mix-ups | Boston Scientific sent an "Urgent Medical Device Recall Removal" letter dated August 1, 2011 to all affected customers.The letter described the problem and product being recalled and advised customers to segregate the affected product immediately and return it to Boston Scientific. Boston Scientific requested the completion and return of the Reply Verification Tracking Form.For additional information on this recall call Boston Scientific Quality Systems at (763) 494-1133. |  | 33 | 33 | (USA) Nationwide distribution including the states of AZ, AR, CA, CO, FL, GA, IL, IN, NV, NY, NC, OH, PA, RI, and TX. | Terminated | August 10, 2012 | 375 | 33 | Not_Computer
Z-3074-2011 | 59680 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Habib, 4x Kit, Bipolar Resection Device RITA (Radiofrequency Ablation Devices). |  |  | August 30, 2010 | 2010 | Angiodynamics, INC | The labeling contains errors which result in the products being misbranded | Labeling False and Misleading | AngioDynamics, Inc. sent out a "Medical Device Recall" letter dated August 30, 2010 to all affected customers.The letter included; products affected, problem and asked them to return the product back to AngioDynamics. Customers were asked to fill out and return the included Recall Reply Form. For question on this recall contact AngioDynamics at (800) 772-6446. |  | 132 | 132 units | Worldwide Distribution: (USA) Nationwide distribution and the countries of Austria, China, Denmark, Egypt, France Germany, India, Italy, Latvia, Netherlands, Poland, Russian Federation, South Africa, Switzerland and the United Kingdom | Terminated | March 12, 2012 | 560 | 769 | Not_Computer
Z-3070-2011 | 59588 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Entroy Pool Lift; a patient lift used in conjunction with a chair or stretcher to lift them in/out of a swimming/therapy pool; Arjo Hospital Equipment AB, Eslov, Sweden; Models GAB1000-01-xx, GAB1000-11-xx and GAC100x |  |  | August 04, 2011 | 2011 | Arjo, Inc. dba ArjoHuntleigh | The Entroy Pool Lift boom may rapidly lower un-commanded to the bottom of its stroke, with a potential for injury to the resident or caregiver. | Device Design | ArjoHuntleigh sent an Urgent Field Correction Notice I.D. 110701AHE to the affected end user via UPS 2nd Day on August 4, 2011. The notice informed the account of the potential for lift boom to rapidly lower un-commanded to the bottom of its stroke. The account was requested to immediately cease use of the affected lift, and disconnect the power to the lift to prevent further use until corrected. The account was requested to complete and sign the Customer Response Form, and return it to ArjoHuntleigh to schedule a service call to correct the lift. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 54485. |  | 311 | 1 unit in the U.S., 311 units in the rest of the world | Worldwide Distribution - USA including CT | Terminated | November 28, 2011 | 116 | 311 | Not_Computer
Z-3069-2011 | 59560 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical 15x100 mm Kii Optical Access Systems, Non-Threaded 6/Box, Model C0R39 |  |  | December 18, 2009 | 2009 | Applied Medical Resources Corp | The recall was expanded on December 22, 2009 to include four models: C0R36, C0R37, C0R38, and C0R39. Initially, Applied Medical conducted a voluntary recall on June 9, 2009 of two lot numbers, 1084426 and 1085506, of the 15 mm Kii non-threaded optical access system, model C0R36, due to a potential weakness in the cannula tip, possibly resulting in a tip breakage. | Component design/selection | Applied Medical Resources Corp. notified customers via an overnight "URGENT: MEDICAL DEVICE RECALL" letter dated December 22, 2009. The letter informed customers of the recalled products, problem, and actions to be taken. The letter asked that the listed recalled products be returned immediately. The customers were asked to check their inventory, complete the attached Voluntary Recall Notification Confirmation Form and fax as per the instructions. The recall letter provided customers with a contact and telephone number for product return questions (1-949-713-8652) and a separate contact for regulatory questions (1-949-713-8369). |  | 13791 | 13,791 units total | Worldwide Distribution - Nationwide in the USA and to countries of: Australia, Austria, Belgium, Canada, Chile, Columbia, Cyprus, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Ireland, Italy, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Saudi Arabia, Singapore, South Africa, and Switzerland. | Terminated | September 06, 2011 | 627 | 13791 | Not_Computer
Z-3075-2011 | 58489 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, DASH 3000, 4000, 5000 Patient Monitors |  |  | July 05, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue with the Dash 3000/4000/5000 patient monitors configured with the 802.11b wireless LAN option that may impact patient safety. The Dash monitor may exhibit reboots on some networks when the 802.11b wireless LAN option is used to connect to the hospital wireless LAN infrastructure. This reboot results in a loss of patient monitoring for approxima | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. GE recommended actions that should be taken by a qualified service representative. For any questions regarding the medical device correction or identification of affected items contact Technical Support at 1-800-558-7044 or customers should contact their local Service Representative.UPDATE: A second letter dated 12/13/11 was sent on 1/20/12 to consignees. The revised letter provides an update to the safety instructions and product correction. |  | 8029 | 8029 devices | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI and the countries of UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SWEDEN, SPAIN, SLOVENIA, SINGAPORE, SAUDI A ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, KUWAIT, JAPAN, ITALY, ISRAEL, INDIA, GERMANY, FRANCE, FINLAND, ESTONIA, ECUADOR, DENMARK, CANDA, BELGIUM AUSTRIA, AUSTRALIA, and ALGERIA. | Terminated | September 11, 2012 | 434 | 8029 | Other | Device Operation | Instructions | Safety Notice/Insructions
Z-3085-2011 | 59469 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Surgical ReNew Laproscopic Long Fenestrated Grasper Forceps Tip |  |  | July 18, 2011 | 2011 | Microline Surgical, Inc. | Mislabeled: ReNew product number 3262 (Long Fenestrated Grasper) was packaged and labeled asproduct number 3642 (Super-Atrau Fenestrated Grasper). ReNew product number 3642 (Super-Atraumatic Fenestrated Grasper) was packaged andlabeled as 3262 (Long Fenestrated Grasper | Packaging process control | Microline Surgical sent a RECALL NOTICE letter dated July 18, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine product subject to the recall. If customers further distributed the product, customers should identify their accounts and notify them immediately of the product recall. Customers were instructed to contact Microline Customer Service to obtain a Return Authorization Number for return of the recalled item. Customers were asked to complete and return the enclosed Recall Return Response Form as soon as possible. For any questions customers were instructed to call the Regulatory Affairs Manager at 978-867-1726. |  | 170 | 170 units | Worldwide Distribution - USA including FL, MI, MN, TN and WI and the countried of Canada, The Netherlands, Thailand, UK, Norway, and Italy. | Terminated | September 09, 2011 | 53 | 470 | Not_Computer
Z-2600-2011 | 58804 | No | REO |  |  |  |  | Prism Laser 7642 |  |  | April 27, 2011 | 2011 | American Stress Technologies, Inc. | It was discovered The Prism 7642 was not certified at the time of delivery. The product lacks the required labeling and the product was sold and delivered without the submission of the Initial Product Report. The product lacks User In formation | Radiation Control for Health and Safety Act | American Stress Technologies sent a customer notification letter dated April 1, 2011, informing them of the action they have taken to correct this issue. (1) ASTI conducted laser safety tests and checks on May 3,2011 to verify the Prism was compliant with FDA regulations. This was confirmed in a June 17, 2011, electronic mail message.(2) ASTI placed the required labels described in the initial product report on the Prism on May 3, 2011.(3) ASTI delivered updated user manuals to UNO on May 3, 2011, that describe potential laser hazards and the required labeling in the same way as described in the initial product report.(4) ASTI completed the above corrective actions at no cost to UNO as confirmed in a June 20, 2011, electronic mail message.If you have any further questions please call (724) 410-1030 |  | 1 | 1 | Nationwide Distribution | Open, Classified |  |  | 1 | Not_Computer
Z-3079-2011 | 58874 | [6 empty] | Elekta Synergy |  |  | May 06, 2011 | 2011 | Elekta, Inc. | A small number of the Modular Reverse Diodes (D1A), a component in the equipment room of the Elekta Linear Accelerator, have ruptured. | Other | ELEKTA sent an IMPORTANT NOTICE (FCO 20000103022), dated May 6, 2011. The notice identified the product, the problem, and the aciton to be taken by the customer. Elekta recommended that service personnel always wear protective safety glasses, when they do work on the modulator assembly with the covers removed. Elekta also recommended that users make sure that the modulator cover is installed when the digital accelerator is in operation. Elekta will release a Field Change Order when a corrective action is identified. |  | 368 | 368 | Nationwide and Puerto Rico | Terminated | October 02, 2012 | 515 | 368 | Not_Computer
Z-3043-2011 | 59317 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMedÂż II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump, catheter, and catheter accessories. |  |  | July 05, 2011 | 2011 | Medtronic, Inc. - Neuromodulation | Medtronic is updating information regarding the potential for reduced battery performance that can lead to sudden loss of therapy in a small percentage of Medtronic Model 8647 SynchroMed II pumps that was communicated with Healthcare providers in July 2009. The purpose of the current communication is to provide updated information regarding the scope and occurrence of this issue and to emphasize | Component design/selection | Medtronic Neuromodulation notified physicians with an Urgent Medical Device Correction letter beginning July 05, 2011. A press was issued July 08, 2011. The letter described the issue, severity, recommendations if they notice issue, and patient management recommendations. UPDATE: in November 2011, MDT began to exchange unused pumps in inventory with pumps that contained the new battery. A letter dated November 2011 was left behind at locations in which the MDT representative swapped out the devices. |  | 139653 | 139,653 (105,002 US, 34,61 OUS) | Worldwide including USA, Puerto Rico, Algeria, Aruba, Australia, Austria, Belarus, Belgium, Brazil, Chile. Colombia. Costa Rica. Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Iran Iraq, Ireland, Israel Italy, Jordan, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Panama, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | August 24, 2012 | 416 | 139653 | Battery | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-3084-2011 | 59577 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Maquet HL Console for Heart Lung Machine HL20 |  |  | June 20, 2011 | 2011 | Maquet Inc. | Maquet HL 20 Twin Pump Modules may unintentionally increase speeds for a 1.5-2 second period of time before returning to its original speed. | Device Design | Maquet sent a PRODUCT RECALL NOTIFICATION LETTER/return response forms, dated June 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product until a software update was performed by their trained technicians. Customers were also asked to forward the information contained in the notification letter to all users and to all appropriate Biomedical Staff. Customers were instructed to acknowledge receipt of the notification letter by completing and faxing back the attached response form to 973-807-9210 whether or not they have any of the affected product. For questions customers were instructed to contact their local MAQUET representative. For questions regarding this recall call 973-709-7994. |  | 299 | 299 units | Nationwide Distribution | Terminated | October 16, 2012 | 484 | 299 | Software | Device Operation | Software update | Software Update
Z-3077-2011 | 59370 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure LC 1.0 (software version 3.0.11). |  |  | July 21, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Roche has detected a problem with the MagNA Pure LC 1.0 (software version 3.0.11) and the MagNA Pure LC 2.0 (software versions 1.1.23 and 1.1.24) instruments. Under certain circumstances, only a fraction of each sample is processed which may lead to false negative results, and an impact to patient health if used in a homebrew diagnostic application. | Other | Roche Diagnostics Operations, Inc. sent a letter dated July 21, 2011 entitled "URGENT MEDICAL DEVICE CORRECTION" to all affected customers.The letter included the name product, software versions, and reason for recall and informed them the problem is under investigation. In the meantime, customers are asked to wait 30 seconds after the finalization of the decontamination cycle before initiating the next operation, in order to avoid this issue.For questions on this recall please call Roche at (800) 428-5074. |  | 422 | 422 | Nationwide Distribution; including the islands of Puerto Rico and Hawaii. | Terminated | July 02, 2012 | 347 | 570 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-3083-2011 | 59483 | No | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008K@home |  |  | May 04, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Upgrade software to implement new on screen graphics and instructions, the addition of a bracket on the side of the machine to temporarily hold the bloodlines during set up in order to conform to approved 510(k) | Device Design | Fresenius Medical NA issued a formal field correction notification dated May 4 ,2011, sent by certified mail with return receipt requested. The letter identified the product the problem and the action needed to be taken by the customer. The customer was notified of the retrofit that would be done to their machines. The patients' facilities were made aware of the changes that would be made to each of their patient's 2008K@home machines on May 4, 2011. On May 13, 2011, all facilities with active 2008K@home machines were sent a notification letter of the retrofit.Please return the receipt for confirmation that you have received this notice.Please call (781) 699-2053, to schedule a certified technician to conduct the machine retrofit. |  | 705 | 705 units | Nationwide Distribution.Canada - FMCNA is not conducting the recall to Canada, as Health Canada did not require the same modifications to the machines | Terminated | August 14, 2014 | 1198 | 705 | Software |  | Software update | Software Update
Z-3081-2011 | 59462 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Mini MaxLock Extreme Plating System Case &amp; Tray, Model Number(s): MXM-901, |  |  | July 06, 2011 | 2011 | Orthohelix Surgical Designs Inc | The surgical tray does not meet our Instruction for Use (IFU) for dry time requirement. | Nonconforming Material/Component | Orthohelix Surgical Designs Inc sent a RECALL NOTIFICATION letter dated July 6, 2011 and an UPDATED URGENT RECALL 2ND NOTIFICATION letter dated July 13, 2011 to all affected customers. The letter identified the recalled product, the problem, and the Field Correction actions to be taken. Customers were instructed to remove the blue rubber mat in the utility bay of the tray base immediately, tag the mat with the tray number, and return it to recalling firm immediately. The letter states once customers have removed the blue rubber mat from the tray base, they may continue to use the tray and system normally. No further actions or precautions are needed. Distributors were also asked to notify their customers of the recall. If you have any questions, please call 330-247-1444 or 330-247-1445. |  | 61 | 61 trays | USA Nationwide Distribution including the states of: AR, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SD, TN, TX, UT, VA, WA, WI &amp; WY. | Terminated | September 10, 2012 | 432 | 61 | Not_Computer
Z-3080-2011 | 59400 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 754/754M Disposable MRI Ventilator Circuit |  |  | May 27, 2011 | 2011 | Impact Instrumentation, Inc. | Model 754/754 M Disposable MRI Ventilator Circuits are being distributed without market clearence (510k). | Labeling False and Misleading | Impact Instrumentation Inc. sent an "Urgent Product Recall" dated May 27, 2011 to all affected customers.The letter informed customers the product is not a cleared product and to return the product back to Impact Inst. for credit. A notification form is included in the letter to be return to the firm.For information on this recall contact the firms Customer Service Department at (800) 969-0750, 8:30 am to 5:00 pm EST. |  | 492 | 492 units | Nationwide Distribution; including the states of CA, FL, IL, NJ, OH, TX, UT, MN, MO, MS, OH, OR, TX, AR, and PA. | Terminated | May 23, 2013 | 727 | 492 | Hardware |  |  | Other
Z-3082-2011 | 59399 | No | JDB | Elbow joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Latitude", Tige Humerale, Humeral Stem, Medium 77 mm, Humeral Stem Right, For Cemented use only, Sterile R, REF 0030402, Tornier In. Edina, MN 55435 USA. |  |  | June 23, 2011 | 2011 | Tornier, Inc | Tornier has discovered that some models of the Tornier latitude Elbow Prosthesis (Small and Medium Humeral stems) were manufactured with a humeral screw that may not function as designed. | Process change control | Tornier, Inc. sent an "Urgent Product Correction" letter dated June 23, 2011. The letter was addressed to the doctors. The letter described the product and the problem. Doctors were insructed to monitor patients that have received the units for evidence of humeral screw loosening.For questions customers were instructed to contact the Customer Service Department at 1-888-494-7950. |  | 3 | 3 | Nationwide Distribution including FL and MO | Terminated | July 17, 2012 | 390 | 3 | Not_Computer
Z-3091-2011 | 57806 | No | LGW |  |  | Neurology | PMA | Software version 3.3 for Rapid Programmer devices, Model 3832. |  |  | June 02, 2009 | 2009 | Advanced Neuromodulation Systems | Product programmed with unapproved software was distributed. | Software change control | St. Jude Medical Neuromodulation contacted all affected customers by telephone between June 2, 2009 and June 5, 2009. Customers were given information about the product, problem and action to be taken.All affected devices were returned to the firm as instructed. |  | 14 | 14 units | Distributed in New Hampshire and Puerto Rico. | Terminated | September 27, 2012 | 1213 | 14 | Software |  | Software update | Software Update
Z-3092-2011 | 59754 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Evident MWA Generator, REF VTGEN, Output: 60 W / 50 Ohm, Valleylab, a division of Tyco Healthcare Group LP Boulder, CO 80301. |  |  | February 06, 2009 | 2009 | Covidien, PLC | Generator may deliver full power when a lower power is selected. | Component design/selection | Covidien sales reps began notifying affected customers on February 6, 2009, followed by a letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return the affected product to:CovidienEnergy Based Devices5920 Longbow RoadBoulder, Colorado 80301Customers were asked to label the exterior of the carton with their Return Authorization number. Covidien EbD would address the issue and return the unit to them.For any questions customers were asked to call Covidien EbD's Service Department toll-free at 1-800-255-8522, x 4923 or contact their local sales representative for additional information. |  | 357 | 357 units | Worldwide Distribution - USA (nationwide) and the countries of Belgium, Denmark, France, Germany, Hong Kong, Hungary, Israel, Italy, Netherlands, Poland, Portugal, Romania, Singapore, Slovakia, Spain, Switzerland, and the UK. | Terminated | August 30, 2011 | 935 | 357 | Not_Computer
Z-3100-2011 | 59351 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, Interpretation of a RapidArc Treatment Plan as a 3D Conformal Arc Treatment Plan; version v8.6 or 8.8 in combination with treatment planning systems; Model Number: H51; Reference/PSCA Identifier: CP-03063; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 27, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | In certain configurations, the 4DITC can interpret a RapidArc or VMAT treatment field as 3D Conformal Arc Treatment Plan. If a RapidArc or VMAT treatment field is incorrectly interpreted as Conformal Arc, the field will be delivered with a constant dose rate and gantry speed for the entire duration of the delivery, which can lead to delivered dose being different than planned. | Software design | Varian sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 27, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that Varian Service will contact customers to arrange to perform update to 4D Integrated Treatment Console (4DITC) software. For further clarification, contact your local Varian Customer Support District or Regional Manager. The letter also provides Varian Oncology Help desk contact information 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  | 645 | 645 | Worldwide Distribution - USA (nationwide) | Terminated | December 05, 2013 | 892 | 645 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-3093-2011 | 59781 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Encision AEM Suction Irrigation Electrode, L-Diamond, |  |  | August 24, 2011 | 2011 | Encision, Inc. | Electrode tips have been reported breaking off during use or cleaning. | Device Design | Encision Inc. sent out an "Urgent Product Recall" letter dated August 24, 2011 to all affected customers.The letter included: affected product, problem and instructed customers to discontinue use and return all affected electrodes to Encision. An enclosed reply form was to be faxed back to Encision indicating status of firm's inventory.For information on this recall call Encision at (303) 339-6917. |  | 169 | 169 units | Worldwide Distribution: Nationwide distribution, and the country of New Zealand. | Terminated | September 04, 2012 | 377 | 1078 | Not_Computer
Z-3099-2011 | 59238 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PFC Sigma Round Dome Patella 38 MM 3 Peg |  |  | June 28, 2011 | 2011 | DePuy Orthopaedics, Inc. | DePuy initiated a voluntary recall effective immediately, for one lot of P.F.C. SIGMA RD Dome Patella 3 Peg (38mm) due to packaging labels that did not correctly reflect the size of the enclosed patellar component. In the recalled lot, the packaging is labeled for a P.F.C. SIGMA 38mm 3 peg Round Dome Patella, but the product enclosed in the package is a P.F.C. SIGMA 35mm 3 peg Oval Dome Patella. | Employee error | The firm, Depuy , sent an email that included a letter entitled "URGENT INFORMATION - RECALL NOTICE" dated June 28, 2011 to DePuy Area customers ( Directors, Distributors, Office Managers, Hospital Representatives, and Doctors). The letter described the product, problem and actions to be taken. The letter also included implications of implanting the recalled device and patient communication information. The customers were instructed to immediately cease further implantation/use or distribution of the product, immediately return any recalled devices in their inventory, notify their customer of the recall and complete, sign and return the Reconciliation Form(s) via fax to DePuy at 574-372-7567 within seven business days.The letter also states that DePuy is not recommending prophylactic revision or additional follow up in the absence of symptoms. Additionally, DePuy recommends that surgeons continue to follow patients as per their standard care.Questions regarding the recall may be directed to the following DePuy representatives--questions about recall information call the Manager of Customer Quality at 574-372-7333; Salesforce questions, call the Group Product Director, 574-372-5046 or Associate Product Director at 574-372-7154 and for Clinical/Surgeon questions, contact DePuy's Scientific Information Office at 1-888-554-2482. |  | 29 | 29 | Nationwide distribution: CA, MI, NJ and WI. | Terminated | August 28, 2013 | 792 | 29 | Not_Computer
Z-3090-2011 | 59662 | [6 empty] | Bausch &amp; Lomb Stellaris 1.8mm Premium Vacuum Phaco Pack with MICS Needle, Ref: BL5114, packaged 6 units/case, RX, Sterile. |  |  | July 27, 2011 | 2011 | Bausch &amp; Lomb Inc | The aspiration channel on the affected vacuum cassettes may be blocked due to a molding malfunction | Nonconforming Material/Component | Bausch &amp; Lomb Inc. notified all affected customers by direct telephone contact on July 27, 2011 and a follow-up FAX or e-mail dated July 27, 2011The letter included; affected product, problem, and requests customers return the product. An acknowledgement form was included to complete and return with the product.For question on this recall contact Bausch &amp; Lomb at (949) 521-7895. |  | 6 | 6/6-unit cases | Nationwide Distribution. | Terminated | June 21, 2013 | 695 | 246 | Not_Computer
Z-3102-2011 | 57590 | No | EZO | Urethrotome. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | OLYMPUS Pediatric Knives and Resection Electrodes |  |  | December 22, 2010 | 2010 | Olympus America Inc. | The device sterile package seal shows evidence of being breached or was breached | Packaging process control | Olympus America Inc. sent an "URGENT PRODUCT RECALL" letter dated December 22, 2010 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory for the affected product and to destroy or discard, if found. A questionnaire was attached to the letter for customers to complete and return via fax to Olympus at 484-896-7128.Contact the firm at 484-896-5688 for questions regarding this recall. |  | 1305 | 1305 devices | USA (nationwide) including the states of AL, AR, CA, CO, CT, FL, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OR, TX, VA, VT, and WI. | Terminated | September 28, 2012 | 646 | 1305 | Not_Computer
Z-3101-2011 | 59262 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | On-Board Imager (OBI) 1.3, 1.4, 1.5; and Trilogy Mx, TrueBeam 1.0, Offline Review 1.0 - 2.0; |  |  | June 21, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified a software anomaly in the 3D Match environment, regarding a slight mismatch between CT image and Structures in On-Board Imager w/s. | Nonconforming Material/Component | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 21, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter recommends several user actions. The firm is developing a correction for the issue and will contact customers when the correction is available.Questions regarding this notice may be addressed by contacting the Varian Oncology Help Desk at 1-888-827-4265. |  | 2268 | 2268 | Worldwide Distribution | Terminated | March 15, 2012 | 268 | 2268 | Software | Display/Image | Correction | Repair
Z-3103-2011 | 59217 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrachyVision, versions 8.2, 8.5, 8.6 and 8.9. |  |  | August 11, 2010 | 2010 | Varian Medical Systems, Inc. | Medical device for brachytherapy treatment may cause users to enter incorrect syntax and lead to potential misinterpretation of patient treatment plan. | Software design | Varian Medical Systems, Inc. sent a "Medical Device Recall" letter dated August 05, 2010 to all affected customers.The letter included description of product affected and problem. Customers were advised to review template files to ensure that there are no characters immediately after the string in the customized dose quality parameters. Device users were requested to return a response form.For information on this recall contact Varian Technical Support at (800) 360-7909. |  | 1252 | 1252 units | Worldwide and Nationwide Distribution. | Terminated | November 25, 2011 | 471 | 1252 | Software |  | Instructions | Safety Notice/Insructions
Z-3113-2011 | 59632 | No | JXN | Propoxyphene test system. | Toxicology | Toxicology | 510(k) | DrugCheckÂż, Distributed by Express Diagnostics Int'l Inc., 1550 Industrial Drive, Blue Earth, MN 56013. Following cup Model numbers involved in the recall: 60670, 60700, 60713, 60717, 60870, 60923, 61012, 61019, 61023, 61024, 61040, 61042, 61057, 61102, 61105, 61203, 61204, 61206, 61208, 61301, 61303, 61401, 70603, 70809, 71032, 71043, 71055, 71065, 71202, 71204, 71205, 71401, 71402, 60718-3, 60720-6, 60925-A, 60927-3, 60930-A, 60980-3A, 61002-3, 61010-4, 61023-A, 61028-3A, 61039-4, 61040-A, 61055-A, 61056-3, 61057-6, 61100-3, 61102-3,61104-3,61105-5 CRL,61120-3,61202-4,61204-4,61204-A,61206-3A,61207-4,61302-3, 61303-A, 61401-A, 70660-3, 70716-3, 71045-A, 71055-6, 71102-3A, 71204-4, 71206-3A. |  |  | October 04, 2010 | 2010 | Express Diagnostics Int'l., Inc. | Express Diagnostics Int'l, Inc (EDI), is recalling DrugCheckÂż Drug Screen Cups containing buprenorphine (BUP) and propoxyphene (PPX) because they must be labeled as "forensic use only", meaning they cannot be distributed or used for treatment or diagnostic purposed in the United States. | Labeling False and Misleading | The firm, Express Diagnostics Int'l, Inc., sent a "RECALL NOTICE" letter dated October 4, 2010 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to notify their customers of the labeling change, and complete and return the DISTRIBUTOR/CUSTOMER CONFIRMATION FORM via fax to Express Diagnostics Int'l, Inc., at 1-507-526-2252. The letter also noted that "Further shipment of product containing BUP and/or PPX cannot continue until Express Diagnostics Int'l, Inc. receives the signed form."If you have any questions, call 507-526-3951. |  | 196750 | 196,750 | Nationwide distribution:AZ, CA, CT, GA, MD, MA, MN, MO, NH, NJ, OR, SC, and WA. | Terminated | July 06, 2012 | 641 | 196750 | Not_Computer
Z-3106-2011 | 59229 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Depuy Mitek Ratchet Handle with Quick Connect |  |  | June 27, 2011 | 2011 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Quick Disconnect Ratchet Handle may slip up the driver shaft or disengage leaving the shaft in the screw head. | Process change control | Depuy Mitek issued Original Customer Notification Letter on 6/27/11 and an Updated letter to clarify the lot number information sent 6/29/11. Hospitals requested to immediately check all inventories for affected product and return for repalcement product. Also notify surgeons aware of this potential issue and potential risk of a delay in surgery. questions or concerns with regard to this program, please contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-800-382-4682, PROMPT 3. |  | 75 | 75 units | Nationwide distribution. | Terminated | January 19, 2012 | 206 | 75 | Not_Computer
Z-3108-2011 | 59439 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | HiÂżArt versions 4.0.x, and HD versions 1.0.x, H-0000-0003. |  |  | July 13, 2011 | 2011 | Accuray Incorporated | As a result of an internal review, TomoTherapy has identified an issue with HiÂżÂżArt versions 4.0.x, and HD versions 1.0.x that we would like to bring to your attention.During DICOM export of plan level images with a non square exported Field of View (FOV), an anomaly in the process of squaring the plan level image may cause the image to shift with respect to ROIs and dose. When the anomaly occu | Software design | TomoTherapy sent a "Urgent Medical Device Correction" letter dated July 13, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This issue will be remedied with HiÂżÂżArt software version 4.1 and HD software version 1.1 and later.If you have any questions, please contact the TomoTherapy Technical Solutions Center by e-mail "support@tomotherapy.com" . Please refer to the letter for what phone number you should call depending on where you are. |  | 346 | 346 | Worldwide Distribution-- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, MT, NE, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WV, WI and WY and the countries United Arab Emirates United Kingdom Turkey Switzerland Sweden Spain Saudi Arabia Poland Netherlands Italy Germany France Belgium Taiwan Singapore Malaysia Korea Japan India Hong Kong China Australia Mexico El Salvador and Canada | Terminated | March 18, 2014 | 979 | 346 | Software | Display/Image | Software update | Software Update
Z-3104-2011 | 59567 | No | DIC | Alcohol test system. | Toxicology | Toxicology | 510(k) | Drug Check, Distributed by Express Diagnostics Intl Inc., 1550 Industrial Drive, Blue Earth, MN 56013. |  |  | November 01, 2010 | 2010 | Express Diagnostics Int'l., Inc. | Express Diagnostics Intl'l Inc is recalling DrugCheck devices because they were informed via letter from the alcohol test strip manufacturer on October 27, 2010 that the product was not 510 K cleared for in vitro diagnostic use and could not be labeled as such in the US. | Error in labeling | Express Diagnostics Intl., Inc. sent an Updated Recall Notice dated November 1, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have open orders of the affected product to contact customer service about their order. Customers were asked to complete and fax the form on page 3 of the document to 507-526-2252. Further shipment of products could not continue until Express Diagnostics received the signed form.For questions regarding this recall call 507-526-3951. |  | 56650 | 56,650 | Nationwide Distribution including AZ, CA,K CT, FL, GA, IL, IA, LA, MD, MA, MN, MS, MO,K MT, NE, NJ, NY, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY. | Terminated | July 06, 2012 | 613 | 56650 | Not_Computer
Z-3111-2011 | 59297 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips Medical: |  |  | July 08, 2011 | 2011 | Philips Healthcare Inc. | Under certain circumstances, edits to an order can result in unintended scheduled interventions on the IntelliVue Clinical Information Portfolio. A patient may receive interventions or additional medication doses that are not intended. | Software design | Philips Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice. Philips has developed a software upgrade to address this problem in the short term. This upgrade will be available immediately. This software upgrade will be performed by Philips service personnel. While customers await the correction of their devices, customers are instructed to follow the instructions for use for the device, which includes the directions to:1. Always use the medication administration record (MAR) in determining patient care.2. Always review orders and pending administrations in the administration record after any changes to an order are stored.For the longer term, Philips is developing safeguards intended to minimize the potential for similar errors resulting from order changes and upgrade to be available in September 2011.If you need any further information or support concerning this issue, please contact your local Philips representative. |  | 454 | 454 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 454 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-3105-2011 | 59572 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Treatment version 6.6.5052. |  |  | July 20, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Potential for Varian Treatment application failing to load the physical wedge filter for a treatment field on a Siemens Mevatron linear accelerator. Delivery of a treatment field without the planned wedge would deliver more dose than intended and would change the shape of the dose distribution for that field. | Other | Varian Medical Systems sent out "Urgent Medical Device Correction/Urgent Field Safety Notice" letters dated July 20, 2011 to all affected customers.The letter identified product, problem and provided recommended user actions when setting up and running the system. It also asked customers to notify appropriate personal provided instructions.For information or questions on this recall please call Varian at (888) VARIAN5 (888- 827-4265). |  | 118 | 118 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, FL, GA, IN, MA, MD, MI, MN, MO, NJ, NY, PA, SC, TN, VA and WV; and the countries of Canada, Austria, Belgium, Brazil, Denmark, Finland, France, Germany, India, Ireland, Italy, Lithuania, Poland, Portugal, Romania, South Africa, South Korea, Spain, Sweden, Switzerland and UK. | Terminated | December 05, 2013 | 869 | 118 | Other | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-3098-2011 | 59350 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand 4D Integrated Treatment Console (4DITC) Radiation Therapy, version 8.1 SP2 to 10.0 SP1; Model Number: H51; Reference/PSCA Identifier: CP-03940; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 23, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has scheduled for an upgrade of the device from current version to 10.2. There are circumstances in which the 4DITC may be involved in a failure to correctly record treatment history information. | Software design | Varian Medical Systems, Inc. sent an "Urgent Medical Device Correction/Urgent Field Safety Notice" letter dated June 23, 2011 to all affected customers. The letter identified the product, problem and recommended user actions to follow. Customers were instructed to notify all appropriate personnel of the correction notice. The letter states that the recalling firm will issue a Customer Technical bulletin to inform users of the mandatory upgrade and related information. For future reference, this document will be posted to the Varian customer support website: http://www.MyVarian.com. Any questions regarding this Product Notification letter, contact Varian Medical Systems Help desk at 1-888-827-4265 for USA and Canada and +41417498844 for Europe. |  |  |  | Worldwide Distribution - USA (nationwide) | Terminated | March 15, 2012 | 266 |  | Software | Device Operation | Software update | Software Update
Z-3107-2011 | 59470 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Li-Ion Docking Station; a lithium ion battery powered docking station for use with GemStar Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; List Numbers 13075-04-01, 13075-04-03, 13075-04-05, 13075-04-07, 13075-24-01, 13075-24-03, 13075-24-05 and 13075-24-07, Note: list numbers with xxxxx-04-xx are for U.S., Canada, Japan, Netherlands and Singapore, while those with xxxxx-24-xx are for international distribution only. |  |  | August 01, 2011 | 2011 | Hospira Inc. | Hospira has received reports of sparking, smoking, charring and electrical shock when using the GemStar Docking Station due to fluid ingress into the docking station. | Device Design | Hospira Inc. sent an "URGENT DEVICE FIELD CORRECTION" letter dated August 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The requests that customers do not hang or place fluid containers over the docking station, not spray fluid directly on the docking station, and to clean the docking station using a cloth dampened with an apporoved cleaning solution. A reply form was attached to the letter for customers to complete and return via fax at 262-577-6921.Contact Hospira at 1-800-441-4100 for questions concerning this notice.*** Hospira sent an update letter dated September 26, 2012 to the accounts, informing them of the availability of the improved design replacement docking stations, and requesting them to complete and return the enclosed reply via fax or e-mail to Hospira, indicating the number of docking stations to be returned for replacement at no charge if the old docking stations are returned within 45 days of receipt of the replacement stations. *** |  | 18000 | 18,000 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Belgium, Canada, Chile, France, Germany, Hong Kong, Ireland, Italy, Japan, Jordan, the Netherlands, Portugal, Saudi Arabia, Singapore, Spain, Sweden and Switzerland. | Open, Classified |  |  | 18000 | Not_Computer
Z-3114-2011 | 59777 | No | LZC | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | MICRO-TOUCH(R), S (indicating size small), NITRILE, Powder Free, Medical Examination Gloves, REF 6034301, 200 Pieces, Manufactured for: Ansell Healthcare - Made in Malaysia |  |  | August 04, 2011 | 2011 | Ansell Healthcare Products LLC | Gloves were inadvertently released for distribution prior to completion of testing required by the U.S. Food and Drug Administration. | Release of Material/Component prior to receiving test results | Ansell account managers were notified on August 4, 2011 to contact (by phone, fax, e-mail) their respective accounts and notify them of the recall situation. The Ansell account managers were provided with a copy of the recall notification to provide to their accounts. The recall notification includes a request that if these gloves have been further distributed to other customers, please advise them of the recall situation and have them return any outstanding stocks. |  | 3700 | 370 cases of 10 | Nationwide distribution: NC, GA, IL, MO, LA, and PA | Terminated | November 27, 2012 | 481 | 14420 | Not_Computer
Z-3112-2011 | 59345 | No | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | xTAG CYP2D6 Kit v3 IVD For use with Luminex 100/200 Instrument Luminex Molecular Diagnostics, Inc. 439 University Ave. Toronto, Ontario Canada Catalog # I030B0373 |  |  | May 24, 2011 | 2011 | Luminex Corporation | Inaccurate test results, due to possible defective primers, which could lead to incorrect dosage on manufactured drug assays. | Other | Luminex sent a "Customer Advisory" letter dated May 24, 2011, by email and courier to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to destroy any remaining product and to submit a return postcard to confirm that the consignee had been notified of the recall and had destroyed the product.For further questions please call Luminex Technical Support at 1-877-785-2323 ( US ) and 1-512-381-4397 ( Outside US ) Monday through Friday. |  | 45 | 45 kits | Worldwide Distribution -- USA ( nationwide) and the countries of Belgium, UK and Canada. | Terminated | March 23, 2012 | 304 | 45 | Not_Computer
Z-3110-2011 | 59436 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque antibody CD99 (H036-1.1) |  |  | July 12, 2011 | 2011 | Cell Marque Corporation | Incorrect label on product bottle, although labeling on outer box is correct. | Labeling mix-ups | The firm, Cell Marque, notified customers by telephone on July 12, 2011. The customers were instructed to return the product. Cell Marque immediately issued replacement of the product.If you have any further question or require additional documentation, contact the VP of Regulatory and Quality Affairs at (916) 746-8900. |  | 2 | 2 units | Worldwide distribution: USA (nationwide) include state of : CA and country of: Great Britain. | Terminated | August 31, 2011 | 50 | 2 | Not_Computer
Z-3109-2011 | 59366 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 91200, Juggerknot system, Soft Anchors, -Two implants size 1, Biomet Sports Medicine, STERILE. |  |  | July 21, 2011 | 2011 | Biomet, Inc. | The device packaging was punctured by the device, resulting in loss of implant sterility. | Package design/selection | Biomet Sports Medicine notified distributor accounts by telephone on July 08, 2011 with instructions to quarantine all lots of the recalled product. Urgent Medical Device Recall Notice letters dated July 21, 2011 were sent to all affected customers. The letter identified the recalled product, problem, and actions to be taken. The customers were instructed to immediately locate and discontinue use of the recalled product. The distributors are advised to communicate this notice to their customers if the recalled products were further distributed. The letter also states that distributors are responsible for the return of the recalled product. Customers were asked to complete a response form and fax to 574-372-1683, within 3 business days. Customers were instructed to provide phone confirmation at 800-348-9500, extension 3755 or 3756, when the recall notice is received. Further questions should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8 AM to 5 PM. |  | 4925 | 4925 | Worldwide Distribution - USA (nationwide) and the countries of: Columbia, Europe and Korea. | Terminated | April 26, 2013 | 645 | 4925 | Not_Computer
Z-3118-2011 | 59570 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance Vision Single Chamber Washer/Disinfector, Model # 1802983, |  |  | May 09, 2011 | 2011 | Steris Corporation | STERIS learned that Vision SC Washer users were experiencing nuisance alarms and/or aborted cycles that are adversely affecting the reliability and operation of their unit. | Pending | The firm, STERIS sent an "URGENT FIELD CORRECTION NOTICE" letter dated May 9, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed that they can continue using their Vision SC Washer(s) prior to installation of the system upgrades. The letter states that STERIS is correcting the issues by implementing a series of quality improvements to the customers Vision SC Washer(s) to ensure satisfaction. Additionally, a STERIS Service Representative will contact the customers facility to arrange for STERIS to upgrade their Vision SC Washer(s).For further information or if you have questions regarding the servicing of your unit, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 404 | 404 Units | Worldwide distribution: USA (nationwide) including states of: L, AR, AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI and WY; and countries of: Australia, Canada, Hong Kong and South Korea. | Terminated | March 01, 2013 | 662 | 404 | Not_Computer
Z-3125-2011 | 58705 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System |  |  | May 09, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the probe wash collar tubing may become stretched or trapped at the Sample Aspiration Module (SAM) during routine use of single-tube presentation station. When the tension on the tubing is severe, the tubing may not be able to return to its home state and may pull the wash collar out of alignment with the aspirate probe. The misal | Pending | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated May 11, 2011 with an attached Product Correction Response Form to all affected customers. The letter provides the customers with an explanation of the problem identified, potential impact should the problem occur and actions to be taken. The letter states that this issue will be corrected in a future hardware release. Customers were instructed to share this information with laboratory staff and retain this notification as part of Quality System documentation. If ownership or location of the analyzer(s) has been transferred to another laboratory, customers were instructed to provide a copy of the letter to that party. Customers were also asked to complete and return the enclosed Response Form within ten (10) days, so Beckman can be assured they have received this important communication. Questions concerning this notice, were directed to the Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their local Beckman Coulter representative. |  | 680 | 680 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Belgium, Canada, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Republic of Macao, Malaysia, Myanmar, Netherlands, Norway, Philippines, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | June 18, 2012 | 406 | 680 | Not_Computer
Z-3121-2011 | 59232 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, 1.5T Optima MR450w |  |  | July 29, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of two potential safety issues associated with MR products (Optima MR450w, Discovery MR450, Discovery MR750, Brivo MR355, Optima MR360) that may impact patient safety.1) Inline post-processed (i.e. automated post-processing) image locations are not correctly annotated under some conditions. 2) Reference images displayed in Functool FiberTrak do not match with | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated July 27, 2011 to customers with affected systems. The letter described the recalled product, safety issues and recommended actions to be taken. The letter provided immediate mitigation (Safety Instuctions) that address each issue. The letter states that the recalling firm will contact customers to arrange to provide a software upgrade to all affected systems at no cost. For any questions concerning this notification contact your local service representative. |  | 159 | 159 | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, LA,ME, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WV, WT, WY.and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CROAT IA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TUKEYand UNITED KINGDOM. | Terminated | August 07, 2013 | 740 | 770 | Computer
Z-3124-2011 | 59752 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Model # P6500, S/N: 400001 through 405005; Mfr. By: STERIS Corporation, Mentor, OH 44060. |  |  | August 11, 2011 | 2011 | Steris Corporation | STERIS learned that SYSTEM 1E users may experience aborted cycles due to a failure within the UV light system. | Other | STERIS sent an "Urgent Field Correction Notice" letter dated July 2011 to all affected customers. The letter describes the product, problem and actions to be taken. The letter states that STERIS Service Representative will contact customers to schedule the repair of the recalled product. For further information or questions contact STERIS Field Service Dispact at 1-800-333-8828. |  | 2235 | 2,235 units | Worldwide Distribution - USA (nationwide) and the country of Bahrain. | Terminated | March 19, 2013 | 586 | 2235 | Not_Computer
Z-3119-2011 | 59691 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Operations and Technical Manuals for the firm's WarmAir Model 135 Hyperthermia System (115V, 100V, and 230/240V), 510 (k) #K101148, Serial Numbers: 994-1350001 through 104-1356894. The WarmAir 135 Hyperthermia System is a patient temperature management device which provides forced air warmed by the controller to ... a blanket that is placed over or around adult, pediatric or neonatal patients in order to warm them. The heated air is blown through a connecting flexible hose to a disposable blanket to provide patient therapy by the means of warmed air. The system can be used in long-term care facilities, surgical facilities, hospitals including the Post-anesthesia Care Unit (PACU), Intensive Care Unit (ICU), Surgical Intensive Care Unit (SICU), Emergency Room (ER), Operating Room (OR), medical and surgical floors, or any other department or hospital facility requiring patient temperature management. |  |  | July 20, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Cincinnati Sub-Zero is conducting a field correction due to changes that have been made to the WarmAir Model 135 devices Operation and Technical Manuals in order to comply with FDA device labeling regulations. | Labeling design | Cincinnati Sub-Zero Products, Inc. sent an 'URGENT Medical Device Recall Field Action Notice" dated July 27, 2011, to all affected user/customers. The letter includes: description of device; the problem with the suspect manuals, and provides contact information for the customers to use to request new versions of the manuals. Upon receipt of the new manuals, the customers are asked to immediately collect and destroy/discard all of the old manuals at their facility and to replace them with the new versions. After the manuals have been discarded and replaced, the customers are asked to complete and return the attached "Recall Response Form" via fax to (513) 772-9119 or scan and email to WA135manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, OH 45241.For additional information the customers are provided with an e-mail address: WA135-2011@cszinc.com as well as a toll-free telephone number and fax number to reach the firm and obtain additional information as necessary. The contact numbers are: 1 (800) 989-7373, or (513) 772-8810, and/or fax # (513) 772-9119.. |  | 6561 | 6561 units were distributed. | Nationwide distribution: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | July 05, 2013 | 716 | 6561 | Not_Computer
Z-3031-2011 | 59335 | No | JKD | C-peptides of proinsulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT C-Peptide Reagent Kit; a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of C-Peptide in human serum, plasma and urine on the ARCHITECT i System; Produced by Biokit S.A., 08186 Barcelona, Spain for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA and Abbott, 65205 Wiesbaden, Germany; list 3L53-25 |  |  | July 11, 2011 | 2011 | Abbott Laboratories | C-Peptide lots 03610K000 and 01711A000 have the potential to generate falsely elevated results with certain patient samples and with non-Abbott Controls. | Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated July 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the lots and destroy any remaining inventory of the lots according to laboratory procedures. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051.Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. |  | 3049 | 3,049 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Korea, Spain, Switzerland, Taiwan, Trinidad &amp; Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | February 10, 2012 | 214 | 3049 | Not_Computer
Z-3128-2011 | 59685 | No | NDR | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Inhibin A ELISA, Part Number: DSL-10-28100-4 |  |  | November 10, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed the microbial contamination in the conjugate diluent bottles included in the Manual Inhibin A Elisa kit (Part Number DSL-10-28100-4), kit lot number 110761. | Other | Beckman Coulter initiated an "URGENT PRODUCT CORRECTION" on June 15, 2011, via a telephone script to the only one customer impacted by the recalled product. The telephone script provided an explanation of the problem identified and the actions to be taken. Beckman Coulter recommended that the customer discontinue use of the recalled product as a preventative measure. Beckman will send a representative to the customer's facility to replace the recalled product in all remaining kits of the affected lot. The customer was instructed to share this information with laboratory staff. For any questions concerning this information, contact Beckman Coulter at 952-368-7874. |  | 342 | 342 units | (USA) Nationwide Distribution in the state of CA | Terminated | July 11, 2012 | 609 | 342 | Not_Computer
Z-3129-2011 | 59601 | No | LJC | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | VIDAS Chlamydia test kits |  |  | July 08, 2011 | 2011 | bioMerieux, Inc. | Complaints were received of high rates of false positive results, leading to unnecessary antibiotic therapy being prescribed to patients. | Other | bioMerieux sent an "URGENT PRODUCT REMOVAL NOTICE" letter dated July 8, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to discontinue using and to destroy any of the affected product. The firm advises customers to confirm all positive results obtained with the affected lot number with another method prior to reporting results. An Acknowledgement and Product Replacement form was attached for customers to complete and return to the firm.Customers should contact their local bioMerieux Customer Service representative for questions regarding this notice. |  | 283 | 283 units | Nationwide Distribution- including CA, DC, FL, IL, IL, IN, MD, MI, MN, MO, MT, NC, NE, NJ, NY, PA, UT, VA and WI. | Terminated | October 23, 2012 | 473 | 283 | Not_Computer
Z-3130-2011 | 58613 | No | MIH |  |  | Cardiovascular | PMA | Cook Medical Zenith Flex AAA Endovascular Graft Main Body, REF TFFB-24-82-ZT, REF G48407, 18 Fr ID/ 7.1 mm OD, Sterile |  |  | April 15, 2011 | 2011 | Cook, Inc. | The firm initiated this action in response to information that there may be potential damage in the braided suture which is used to attach the external stents to the graft material. | Process control | The firm sent URGENT PRODUCT RECALL notices dated 4/15/2011 to Risk Managers/Recall Administration (Chief of Vascular Surgery and or Interventional Radiology). The letters explained the problem with the devices and the action to be taken. Consignees were instructed to return the recalled devices to Cook Medical for credit or replacement. Customers with questions or concerns can contact the firm at 800-457-2235. |  | 4 | 4 units | US, Canada, Japan | Terminated | September 13, 2011 | 151 | 71 | Not_Computer
Z-3159-2011 | 59659 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II TURBO Link Hemoglobin A1c Program Reorder Pack. |  |  | July 28, 2011 | 2011 | Bio-Rad Laboratories, Inc. | The Elution Buffer B, lot no. GA 10430, was shipped in a 2.0 L bottle instead of the standard 2.5 L bottle. | Employee error | The firm, BIO-RAD, sent a "Medical Device Correction" letter dated July 28, 2011 to all consignees/customers via Fed Ex on August 8, 2011. The letter described the product, problem and actions to be taken. The customers were informed that they could continue to use and report results with the affected buffer lot and the stability studies will continue until lot expiration date. The consignees were instructed to distribute this letter to all customers that have received the affected lot and to distribute the MEDICAL DEVICE CORRECTION RESPONSE FORM only to those customers that request replacement Elution Buffer B. In addition, the customers were to complete and return the MEDICAL DEVICE CORRECTION RESPONSE FORM via fax to (510) 741-3954, Attention: BIO-Rad CSD Regulatory Affairs Department.If you have any questions, please contact your regional Bio-Rad office or 510-724-7000. |  | 40 | 40 | Nationwide distribution: MD, MI, MN, NY, OH, SC, VA and WI. | Terminated | November 01, 2011 | 96 | 40 | Not_Computer
Z-3120-2011 | 59692 | No | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom dialysis prep kits labeled as follows: |  |  | March 22, 2011 | 2011 | Lee Medical International Inc | Firm manufactured custom dialysis kits with Povidone Iodine (PVP) Prep Pads which were recalled by H &amp; P industries. Those pads were subsequently recalled due to possible bacterial contamination. | Material/Component Contamination | The firm, Lee Medical International Inc., notified their consignees by email and telephone on 03/22/2011. Consignees were instructed to pull product from stock, quarantine for re-processing, and return or destroy product. Lee Medical arranged for replacement product from alternate manufacturer to be sent to customers.For questions call Production Manager or Operations Manager at 504-734-9336 or email melissal@eleemedical.com or chriss@eleemedical.com. |  | 19625 | 19,625 units | Nationwide Distribution (USA) - including states of: AZ, CA, FL, GA, MD, NJ, NM, NY, PA, and TX. | Terminated | January 30, 2012 | 314 | 19625 | Not_Computer
Z-3160-2011 | 59737 | No | NSX |  |  | General Hospital | Contact ODE | Picis ED PulseCheck-Software, transmission &amp; storage, patient data |  |  | August 01, 2011 | 2011 | Picis Inc. | Software does not display appropriate allergy interaction warning | Software design | Picis issued notification on 08/01/11 to affected clients, explaining the potential for the error, with instructions to monitor documented patient allergies closely when prescribing medication orders in ED PulseCheck, using Muitum drug database content. Picis will track the progress on acknowledgement of notification letter receipt through the implementation of software upgrade, for all potentially affected Customers.For questions regarding this recall call 781-557-3000. |  | 62 | 62 | Nationwide Distribution | Terminated | June 27, 2013 | 696 | 62 | Software | Alarm/Message | Software update | Software Update
Z-3162-2011 | 58473 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ratcheting Torque Limiting Handle, Catalog Number 52-1012, Orthofix Spinal Implants, Firebird Spinal Fixation System, Firebird Spinal Fixation System DDD Instrument Case. |  |  | April 01, 2011 | 2011 | Orthofix, Inc. | Upon reaching torque limit, there is a possibility the product may jam or exhibit a ratchet direction switch from forward to reverse. | Device Design | Orthofix, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 1, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers.The letter instructs customers to identify and remove the affected product from use. Customers were advised to use the combination of the Ratching R Handle and the Set Screw Driver if they experience problems with the Ratching Torque Limiting Handle. A Tracking and Verification form was attached for customers to complete and return to the Orthofix Inc. Regulatory Affairs Department via fax at 214-937-3322.Contact your Account Services Representative at 888-298-5700 for questions regarding the removal and return of the affected product. |  | 287 | 287 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia and Germany. | Terminated | March 23, 2012 | 357 | 287 | Not_Computer
Z-3161-2011 | 59305 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TiLite ZR and ZRA Series 2 are mechanical wheelchairs. Devices are sold under brand name TiLite. |  |  | July 05, 2011 | 2011 | TiSport, Llc | The Pivot Stud Bolt comprising part of the backrest assembly of TiLite ZR Series 2 and ZRA Series 2 mechanical wheelchairs have shown a tendency to fail due to the head of the bolt being sheared off. | Component design/selection | An Urgent Medical Device Correction letter, Product Correction Kit , and the Installation Instructions were sent to domestic suppliers, international suppliers, Veterans Administration hospitals and international distributors by US Priority Mail and by FedEx three days service.The letter identified the affected product and the reason for the recall. Customers were instructed not to send the product to their end user until the correction was made or to identify their end user and notify the end user of the product correction. Replacement parts and instructions to correct the issue were enclosed with the letter. Customers were to destroy the removed parts once they had completed the correction. Customers were to also return the Medical Device Correct Report that was provided, by fax to 509-586-2413, or email to ProdCorrection@tilite.com.Customers can call Technical Support at 800-545-2266, ext. 239 with any questions about recall. |  | 1037 | 1,037 units | Worldwide Distribution | Terminated | July 27, 2012 | 388 | 1037 | Not_Computer
Z-3172-2011 | 59453 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Wound Dressing Kit - Tier I, 10 MM Flat Drain; product code 0877061. |  |  | July 27, 2011 | 2011 | Medela Inc | Medela determined that some of the individual kits or bags of lot #370118 of "0877061 Wound Dressing Set, Tier I, 10 MM Flat Drain", were labeled incorrectly as "0877067 Wound Dressing Set, Tier III 15 Fr Round Channel Drain". The outside label on the case is correctly labeled as "0877061 Wound Dressing Set, Tier 1,10 MM Flat Drain". | Error in labeling | Medela Inc. sent an "Urgent Device Recall" letter dated July 26, 2011 to all affected customers.The letter described the product, problem and actions to be taken. The customers were instructed to do the following: 1) not to use the affected lot and place them in quarantine 2) stop sale and distribution and set aside their existing inventory of the affected product 3) conduct a physical count and record the data on the Product Inventory Form included with this letter 4) complete and return the Product Inventory Form via fax to 800-995-7867 or e-mail suction@medela.com, even if they have no affected product on hand 5)if they have the affected product in stock, Medela Customer Service will ship substitute product to the customer at no cost. 6) and to record the number of units on hand for return and replacement on the enclosed Product Inventory Form. The completed form should be returned via fax or e-mail to Medela even if the account has no inventory of the affected lots. If you have any questions about this recall please call (877) 735-1626 or e-mail at suction@medela.com. |  | 3000 | 3,000 kits | Nationwide Distribution including the states of California, Florida, Georgia, Illinois, Kansas, Kentucky, Massachusetts, Michigan, Mississippi, New Jersey, New Mexico, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee and Texas. | Terminated | February 16, 2012 | 204 | 3000 | Not_Computer
Z-3164-2011 | 59798 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DUET TRS 45 3.5mm Universal Straight Single Use Loading Unit Surgical Stapler with Implantable Staples |  |  | August 15, 2011 | 2011 | Covidien LP | Sterility of the device is compromised due to a breach in the sterile packaging | Storage | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated August 15, 2011, to its customers via federal express. The letter describes the product, problem, and actions to be taken. The customers were instructed to immediately identify and quarantine all affected inventory (Note: Units from the affected lots may have been incorporated into Covidien BEST PRACTICE KITS. If you have purchased BEST PRACTICE KITS, please evaluate whether these kits contain units from the affected lots.); complete and return the DUET TRS Universal Loading Unit Recalled Product Return form via fax to (203) 822-6009; and return all inventory from the affected lots (note: All affected units must be returned) with the complete form to: Field Returns Department, located at 195 McDermott Road, North Haven, CT 06473. Questions or concerns, contact Covidien Representative or Covidien Quality Assurance at (203) 492-5232. |  | 4362 | 4362 units | Worldwide distribution: USA (nationwide) and countries of: Asia, Australia, Canada, Europe, Japan, Latin America, and Pacific. | Open, Classified |  |  | 457461 | Not_Computer
Z-3173-2011 | 59830 | [6 empty] | BardPort Titanium Implanted Port. |  |  | August 30, 2011 | 2011 | Bard Access Systems | Implanted port has discrepant product information on package labeling. | Labeling Change Control | Bard Access System, Inc. sent a "Label Discrepancy Notification dated August 30, 2011 to all affected customers.The letter included affected product and informed them of the labeling discrepancy. Customers could either use the device or return it to Bard for replacement. For information on this recall call (801) 595-5568. |  | 1375 | 1375 units | Worldwide Distribution: Nationwide Distribution including the states of; AL, AR, AZ, CA, CT, DC, FL, GA, HI, IL, KY, LA, MA, MD, MI, MO, NJ, NV, NY, OK, PA, RI, SC, TN, TX, UT, VA, WI, and WV; and the countries of Canada and Belgium. | Terminated | March 30, 2012 | 213 | 1375 | Not_Computer
Z-3175-2011 | 58425 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Kyphon brand KyphoPak Express Tray First Fracture, 15/2 with KyphX Osteo Introducer System; Catalog Number: KPE1003; Product is manufactured and distributed by Medtronic Spine LLC (formerly known as Kyphon Inc), Sunnyvale, CA |  |  | March 23, 2011 | 2011 | Medtronic Spine LLC (formerly Kyphon Inc.) | Device was shipped past the labeled expiration date. | Employee error | On 3/23/2011, the firm initiated the recall and sent Urgent Notice: Device Recall letters identifying the affected product and explaining the reason for the recall. The letter asks customers not to use the affected product with the lot number 0001653390. The firm will replace any affected product at no cost to the customer. A Medtronic Spine Sales representative will assist with the segregation and return of the affected product. They will also assist the customer with contacting Medtronic Spine Customer Service to return the affected product. Included with the letter was a reconciliation form with instructions regarding the affected product. Questions regarding the letter should be directed to Customer Service at 866-959-7466. |  | 44 | 44 | Nationwide Distribution | Terminated | September 09, 2011 | 170 | 44 | Not_Computer
Z-3183-2011 | 59421 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System Communicator, Model 6482. |  |  | May 25, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has identified two occurrences where the LATITUDE system did not process a data payload, which is the information uploaded to the LATITUDE system from a remote interrogation of the implanted device. These occurrences resulted in non-displayed payloads because the data is not posted to the website. If the payload is not displayed, the pertinent data will not be available on the | Device Design | Boston Scientific notified all affected customers by telephone on May 25, 2011. The customers were informed that there was an interrogation for one of their patients that was not able to be displayed on the website. Boston Scientific detected that a payload failure had occurred, but could not conclude what information was contained within the payload. Each patient has had previous and subsequent payloads that have processed and displayed on the LATITUDE website. |  | 2 | 2 | Nationwide Distribution-in the states of Florida and Texas. | Terminated | December 04, 2012 | 559 | 2 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-3176-2011 | 59455 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Sarns Level Sensor II Pads. |  |  | August 12, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Level Sensor II mounting pads may not stick to the venous reservoir and may detach either at case setup or during a case. | Incorrect or no expiration date | Terumo Cardiovascular Systems Corp. sent an "Urgent Medical Device Removal" notice dated August 10, 2011 to all affected customers.The notice instructs customers to discard all remaining level sensor pads that do not have an expiration date on the package, and notify Terumo of the number of affected pads via a response form. Terumo will replace recalled product.For questions on this recall call Terumo Customer Service at (800) 521-2818, 8 am-6 pm EST. |  | 1531440 | 1,531,440 devices | Worldwide Distribution -- Nationwide Distribution, and the island of Puerto Rico. | Terminated | June 20, 2012 | 313 | 1531440 | Not_Computer
Z-3181-2011 | 59677 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Leica Bond Ready-to-Use Primary Antibody CDX2 (AMT28); an in vitro diagnostic immunohistochemistry primary antibody kit; Catalog No. PA0535; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. |  |  | July 26, 2011 | 2011 | Leica Microsystems, Inc. | There was an issue with the controlled temperature storage unit where the products were stored prior to shipment to the customers. | Employee error | Leica Microsystems, Inc. notified all affected customers by telephone on July 26, 2011. The product, problem, and actions to be taken by the customers were discussed. Customers were instructed to discontinue use of the product and to return the product. All affected product has been returned to the firm. |  | 3 | 3 units | Nationwide Distribution-in the states of Indiana and Kentucky. | Terminated | September 09, 2011 | 45 | 4 | Not_Computer
Z-3177-2011 | 59759 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monomid Non-Absorbable Suture are blue, monofilament synthetic absorbable suture, with different sizes and different needle sizes. |  |  | June 29, 2011 | 2011 | C P Medical Inc. | Unapproved dye was used in manufacturing the non-absorbable blue nylon suture material. | Process control | CP Medical sent a "RECALL NOTICE" letter dated June 29, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the affected products in the human market and to return all their existing stock. A Recall Checklist was enclosed for customers to complete and return.Contact Customer Service at 1-800-950-2763 for a Return Good Authorization and shipping account number. The firm will pay for shipping cost and replace product at no cost to customers. |  | 776 | 776 dozens boxes | Worldwide Distribution-USA (nationwide) and the countries of Italy, Romania, South Africa, and Taiwan. | Terminated | May 24, 2012 | 330 | 776 | Not_Computer
Z-3174-2011 | 59564 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Coloplast Corp.-MN 06093, Country of origin: DK ( SenSura 19011 Post-Op, ostomy bag with window, 100 mm , Sterilized using ethylene oxide; SenSura 19010 Post-Op ,ostomy bag with window, 70 mm , Sterilized using ethylene oxide; SenSura 19001 Drainable bag with soft tap 1 piece/Opaque with inspection windows, 10-76 mm Sterilized using ethylene oxide, Manufacturer: Colopast A/S 3050 Denmark.) |  |  | July 27, 2011 | 2011 | Coloplast Manufacturing US, LLC | On July 19, 2011 Coloplast became aware of the failure or routine sterility testing for one lot of SenSuera and Assura Post-Op Ostomy Bags. Based on these results it was determined that all affected lots be recalled and quarantined. This affects lots manufactured and distributed from January 2010 until July 19, 2011. | Process control | Coloplast sent a "RECALL OF STERILE SENSURA AND ASSURE POST-OP POUCHES" letter dated July 27, 2011 to all affected customers. An enclosed Field Safety Notice described the product, problem, and actions to be taken by the customers.Contact your Customer Service representative at 800-533-0464 for questions regarding this recall. |  | 4020 | 4020 | Nationwide Distribution--USA (nationwide) including the states of AL, AZ, CA, FL, GA, LA, MA, MD, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, RI, TN, TX, VA, WA, and WI. | Terminated | June 05, 2012 | 314 | 4020 | Not_Computer
Z-3179-2011 | 59656 | No | KQT | Gastroenterology-urology evacuator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Fluid Aspiration Set; a sterile set of silicone tubing and connectors for use with the Olympus SP-2 Fluid Irrigation &amp; Aspiration Pump; 10 sets per box; Manufactured for: Olympus America Inc., Two Corporate Center Drive, Melville, NY 11747-3157; catalog #72-00114-0 |  |  | July 25, 2011 | 2011 | Northgate Technologies, Inc. | The Red Tag was assembled to the wrong end of the silicone tubing. The Red Tag indicates how the fluid aspiration tubing set is loaded into the Olympus SP-2 Fluid Irrigation and Aspiration Pump. | Employee error | The firm, Northgate Technologies, sent an "URGENT: MEDICAL DEVICE RECALL" letters dated August 18, 2011 to the Olympus America customers who ordered the Aspiration Tubing Set, 72-00114-0. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall for subsequent return to Northgate Technologies; if they further distributed this product, they were to identify and notify their customers of the recall, and complete and return the enclosed Return Response Form via fax to (847) 608-9405 or mail to: Northgate Technologies Inc., 1591 Scottsdale Court, Elgin, IL 60123. If you have any questions, please call (224) 856-2222 or e-mail to: ckurek@northgate-tech.com. |  | 10680 | 10,680 sets | Nationwide distribution: USA including states of: CA, FL, ID, IL, MI, NY, PA, SD and WV. | Terminated | December 14, 2011 | 142 | 10680 | Not_Computer
Z-3185-2011 | 55319 | No | MHZ |  |  | Neurology | Contact ODE | The 91220 mCare 300 Vital Signs monitor has a 10.4 inch resistive touchscreen display. |  |  | April 14, 2010 | 2010 | Spacelabs Healthcare, Incorporated | There have been reports that the Spacelabs Medical Vital Signs Monitor, Model 91220, will not power up with AC power and will no longer charge the batteries. | Component design/selection | SpaceLabs Healthcare sent an URGENT - MEDICAL DEVICE CORRECTION letter dated April 15, 2010, to all affected customers. International consignees were notified by e-mail on April 22, 2010, with the customer notification letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to weigh the benefits versus the risks when deciding whether or not to continue to use the device until their monitors can be upgraded. Customers were asked to brief their staff about the possibility of the monitor failure to help prevent an adverse event from happening. |  | 614 | 614 units | Worldwide Distribution - USA (nationwide) and the countries of UNITED ARAB EMIRATES, ARGENTINA, AUSTRALIA, BELGIUM, BAHRAIN, BRAZIL, CANADA, CHINA, COLOMBIA, GERMANY, ECUADOR, EGYPT, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HONG KONG, JORDAN, REPUBLIC OF KOREA, KUWAIT, MEXICO, MALAYSIA, OMAN, POLAND, PUERTO RICO, PARAGUAY, ROMANIA, SAUDI ARABIA, SINGAPORE, SYRIAN ARAB REPUBLIC, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, and URUGUAY. | Terminated | September 12, 2011 | 516 | 614 | Not_Computer
Z-3187-2011 | 59764 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATHÂż ETS Compact Flex (240 mm shaft length) Linear Cutters, Model Number(s): SCW45; CTS45; SCB45; CTS45NK; 6CB45 and SCG45. Mfr. By: Ethicon Endo-Surgery, LLC, Guaynabo, PR 00969 |  |  | July 29, 2011 | 2011 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENDOPATHÂż ETS Compact Flex 45mm articulating linear cutters because the possibility exists that damage to the blister pack may have compromised the sterility of the device. | Packaging | Ethicon Endo-Surgery Inc sent an URGENT DEVICE RECALL letter dated August 11, 2011 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. Fill out the Business Reply Form and fax it to 1-866-853-2802 if they have any of the affected product. To ensure adequate identification of the return product at their facility customers should complete the Customer Purchase Order Field on the Business Reply Form. For any questions regarding this recall call 1-866-708-2330 and reference Event 2471. For clinical or product support call 1-800-USE-ENDO, option 6, or contact your local sales representative. |  | 27582 | 27,582 instruments | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, CZECH REPUBLIC, GREECE, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, POLAND, PUERTO RICO, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, THAILAND and UNITED ARAB EMIRATES | Terminated | February 15, 2013 | 567 | 27582 | Not_Computer
Z-3186-2011 | 59477 | No | DRA | Steerable catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic FlexCath, model 3FC12, Steerable Sheath. The FlexCath Steerable Sheath is a percutaneous catheter introducer fitted with a valve to allow for introduction, withdrawal, and swapping of catheters and wires while preventing air ingress and minimizing blood loss. A side-port with stopcock is integrated to allow continuous drip infusion, injection through the center lumen, flushing, aspiration, blood sampling and pressure monitoring. Rx only. Sterilized using ethylene oxide. Manufacturer: Medtronic CryoCath LP, Kirkland, Quebec H9H 5H3, Canada. |  |  | July 15, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic CryoCath would like to inform you of an observation related to a potential leak in the hemostatic valve of the Flex Cath 12 Steerable Sheath, Model 3FC12. There wave been reports of blood and saline solution leaking out of, and air ingress into the hemostatic valve. This includes reported instances of the introduction of air bubbles through the valve during aspiration of the side port | Pending | Medtronic Sales Reps began on July 15, 21011 hand delivering an "IMPORTANT MEDICAL DEVICE iNFORMATION" letter dated July 2011 to consignees. The letter identified the product, the problem, and the actions needed to be taken. The letter stated the rate of occurrence, described the potential clinical implications, and reinforced the importance of following the information found in the instructions for use as consignees continue to use the device. The letter also stated that Medtronic CryoCath is working urgently on a number of fronts to mitigate this observation and will continue to communicate with consignees as appropriate.For further information, please contact Medtronic VP of Quality at (763) 526-0576. |  | 6712 | 6712 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, GA, IA, IL, IN, KY, LA, MA, MI, MO, MN, MS, NC, ND, NE, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI. and the countries of: Australia, Austria, Belgium, Canada, Czech, Denmark, France, Germany, Greece, Hong Kong, Hungary, Ireland, Italy, Netherlands, Norway, Poland, Qatar, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | June 07, 2012 | 328 | 6712 | Not_Computer
Z-3188-2011 | 58478 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Orthovita(R) Cortoss(TM) Delivery Gun |  |  | March 31, 2011 | 2011 | Orthovita, Inc. | Recalling firm has discovered through internal testing that the Cortoss(TM) Delivery Gun may abrade its packaging material under certain transportation and handling conditions. This abrasion may cause small holes or tears in the packaging material, which may not be immediately visible to the user. Although t the Cortoss(TM) Delivery Gun is provided in double pouches and there have been no report | Packaging | Orthovita sent an "URGENT FIELD SAFETY NOTICE" letter dated March 31, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to Identify, quarantine and discontinue use of the affected product. The letter included an acknowledgement form to be completed and returned as per instructions.. The letter states that the recalling firm will contact customers to arrange for the return or destruction of any affected product. A pre-paid shipping label is provided for product return. The letter also states that the recalling firm will provide replacement product at no cost. For any questions regarding this notice call 484-323-8860. |  | 3174 | 3174 | Product was distributed within the US to medical facilities in AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NY, NC, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, and WI. Government accounts include Bethesda Naval Center, 8901 Rockville Pike, Bethesda, MD; Lackland Air Force Base, 2200 Berquist Drive, Suite 1, Lackland AFB, TX; Hunter Holmes McGuire VA Hospital 1201 Broad Rock Blvd., Richmond, VA; Madigan Army Medical Center, Madigan Army Medical Center, Tacoma, WA; Ralph H. Johnson VA Medical Center, 109 Bee Street, Charleston, SC; VA Medical Center, 1481 W. 10th St., Indianapolis, IN; VA Hospital of Palo Alto, 3801 Miranda Ave., Palo Alto, CA; VA Loma Linda Healthcare System, 11201 Benton Street, Loma Linda, CA; and VA Medical Center/NY Harbor Healthcare, NY, NY. | Terminated | September 25, 2012 | 544 | 3174 | Not_Computer
Z-3178-2011 | 59331 | No | LPB |  |  | Cardiovascular | PMA | STOCKERT 70 Radio Frequency Generator, Catalog Number 57001 |  |  | July 07, 2011 | 2011 | Biosense Webster, Inc. | Biosense Webster has initiated a voluntary field notification to all customers that have a Stockert 70 Radio Frequency Generator, to communicate undetected software anomalies in the form of software release note in order to amend the labeling and thereby enhance the safety of the product. | Software design (manufacturing process) | The recall communication was initiated on 7/8/11 and 7/11/11 with Biosense Webster forwarding A Customer Notification Letter, Acknowledgement Form and Software Release Note (Attachments 3, 4 and 1) via FedEx to all U.S. customers who purchased the Stockert 70 Radio Frequency Generator. Two letters and forms were sent to U.S. customers via express mail, addressed to the EP/Cardiology Lab and Risk Management Office, respectively.The Release Note provides tips for the avoidance of potential events when working with the EP-SH UTTLE, STOCKERT 70, STOCKERT J50 or STOCKERT J70 RF generators with SW versions 1.035 / 1.035J /1.035J70 to 1.037/1.037J70. |  | 3397 | 3,397 units total (1,217 in U.S.) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United States, Uruguay, and Vietnam. | Terminated | October 12, 2012 | 463 | 3397 | Software | Physical Safety Hazards | Notification | Safety Notice/Insructions
Z-3200-2011 | 59175 | [6 empty] | Philips GEMINI TF64 Diagnostic Imaging Systems, Model Number: 882471, 510(k) #K052640, Serial Numbers: 7502, 7504, 7506, 7514, 7083 and 7164. The recalled GEMINI TF64 system units equipped with software versions: 3.5.1.1 and 3.5.2.1 are subject to recall/software correction. |  |  | May 31, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | Software design | The firm , Philips Healthcare, issued a 'Customer Information Letter' dated August 9, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were not required to perform any action. A Philips' Field Service Engineer will visit each customer site and install the necessary software upgrade and Philips will track the progress of the software corrections as they take place.For customers in North America, if you need any further information or support concerning this update, please contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5 for "All Imaging Systems". Enter your site ID number. If you do not know your site ID, please hold and then select option 5 for "Nuclear Medicine" and finally select option 2 for "PET" support. In all other countries, the local Philips Healthcare office should be contacted at 440-483-3000. |  | 6 | 6 units | Worldwide distribution: USA including states of: FL, MA, NE, NC, OH, TX, WA and WV; and countries of: France, Germany, Italy, Kuwait, The Netherlands, Spain and Turkey. | Terminated | February 22, 2013 | 633 | 19 | Computer
Z-3191-2011 | 58985 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Sterile Bone Cutting Burrs model QD8-4OVD labeled in part:***QD8-4OVD***QTY: 1***4 mm x 8 mm Diamond Oval. Use with***QD8 Attachments.***Use by 2016-03***ANSPACH***Manufacturer***Made in USA***The Anspach Effort, Inc.***4500 Riverside Drive***Palm Beach Gardens, FL 33410, USA***Tel: (800) 327 6887 / + 1 561 627 1080***EC REP***Anspach Europe Ltd.***Aston Court, Kingsmead Business Park***Frederick Place***High Wycombe, HP 11 1LA***United Kingdom***Tel: +441494 616126***Rx Only***LOT E173049563***. |  |  | May 09, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Inc., Palm Beach Gardens, FL, is recalling their Anspach Single Use, Sterile Bone Cutting Burrs, catalog numbers, IMRI-S-8NS-M, M-6P, QD8-4OVD, S-5B-4, and SIL-7-9, lot #'s C323023789, C403026551, C473028567, C403026563, D153034013, D173034385, and C363024789, because the bar code which defines the expiration date may be incorrect. This could cause the bar code reader system to | Employee error | The Anspach Effort, Inc. sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated May 27, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers.The letter instructs customers to check their inventory and to segregate any affected product found. A Customer Response Form was attached for customers to complete and return to the firm via fax at 1-800-327-6661.Contact Anspach Customer Support at 800-327-6887 for questions regarding this recall. |  | 17 | 17 | Worldwide Distribution--USA (nationwide) including the states of Alabama, California, and Wisconsin and the countries of Columbia, France, Germany, Great Britain, Italy, Panama, Poland, Saudi Arabia, South Africa, Spain, and Switzerland. | Terminated | February 15, 2012 | 282 | 114 | Not_Computer
Z-3195-2011 | 59690 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed |  |  | June 22, 2011 | 2011 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on their TSX-301B: AQUILION Premium; System, X-Ray, Tomography, Computed,TSX-301A: AQUILION One; System, X-Ray, Tomography, Computed,TSX-20IA: AQUlLlON LB; System, X-Ray, Tomography, Computed, and TSX-1O1A: AQUILION 16; System, X-Ray, Tomography, Computed due to several software issues that cause the scanners to function incorr | Software design | Toshiba America Medical Systems (TAMS) sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 22, 2011, to all affected customers. Customers were instructed that a revised software to prevent occurrence of the problem would be installed on their system. Their Toshiba Service Representative would contact them for an appointment to install the software to their system when it became available.Customers were requested to take temporary measures until corrective measures were taken.1) In dual scanoscopy, do not press the short key between the completion of the first scanoscopy and the start of the second scanoscopy.2) In dual scanoscopy, do not click the (Scano Skip) button immediately before the completion of the first scanoscopy.3) When the SplitCase function is set to ON, do not select two or more orders simultaneously from the appointment list.4) When an image comment is to be preset by making a selection from the preset memo, do not select any image comment with more than 10 characters. Customers were asked to share the information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at the facility.Customers were asked to complete and return the attached form and fax it to 877-349-3054.For any questions call (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 155 | 155 units total | Nationwide Distribution | Terminated | August 11, 2014 | 1146 | 155 | Software | Display/Image | Software update | Software Update
Z-3184-2011 | 59797 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Straumann coDiagnostiX, version 8.0, used to guide a clinician in the proper implant selection, oral placement and the proper selection of an oral drill. |  |  | August 04, 2011 | 2011 | Straumann Manufacturing Inc. | Software version 8.0 has a software anomaly, in which the length is not adjusted accordingly when it is changed after the surgical sleeve position planning step has been completed. | Software design | Staumann USA notified accounts by an Urgent Field Safety Notice letter dated August 4, 2011 and delivered via Federal Express Delivery Service.The letter identified the affected product and explained the reason for the recall. Customers were asked to review the letter, take appropriate actions, including completing and faxing the enclosed. Customer Confirmation Form. Straumann released a corrected software upgrade (coDiagnostiX 8.0.2.3672), which is available for download when users go online and activate current software. If the user does not have internet access a replacement CD can be ordered from Technical Support at 866 - 531 -7365. Questions should be directed to Technical Support Monday through Friday from 8:00 am to 8:00 pm Eastern Time. |  | 272 | 272 | Nationwide Distribution | Terminated | May 23, 2012 | 293 | 272 | Software | Device Operation | Software update | Software Update
Z-3076-2011 | 59760 | No | RBS |  |  |  |  | Hamamatsu model L10951-35 Micro Focus x-ray generator. |  |  | July 07, 2011 | 2011 | Hamamatsu Corporation | It was discovered about a potential issue of X-ray leakage with our Model # LI1 0951-35 Micro Focus X-Ray generator. | Radiation Control for Health and Safety Act | Hamamatsu Corporation sent a letter dated July 19, 2011, to the manufacturer and their subcontractor informing them of the affected product and what they propose to do. They will repair the units both in the field and those returned to our plant by installing additional lead shielding in the units. The lead shielding will be fixed to the unit by soldering the addition to the existing lead shield (see attachment 2). We expect to complete the repair of all units in the field by October 1, 2011. All new units to be shipped from our plant will be repaired before shipment.For further questions, please call (908) 231-0960 X 2300 |  | 8 | 8 | Worldwide Distribution -- USA ( nationwide ) the countries of Aichi, India and Japan. | Open, Classified |  |  | 8 | Not_Computer
Z-3206-2011 | 57889 | No | DPZ | Ear oximeter. | Cardiovascular | Anesthesiology | 510(k) | DATEX-OHMEDA TRUSAT PULSE OXIMETER AND ACCESSORIES, MODEL 3500 (GE Healthcare) |  |  | April 08, 2011 | 2011 | GE Healthcare, LLC | The plastic enclosure on the external power supply of the TruSat Pulse Oximeter is prone to breakage, and overheating. The issue may impact patient or user safety. Potential for system or components to start on fire or explode (general case). | Device Design | On 4/8/11 GE Healthcare sent consignees an "Urgent Medical Device Correction" letter dated March 30, 2011. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering &amp; Home Healthcare Provider.The letter listed the Safety Issues, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 31846 | 31,846 | Worldwide distribution: USA, ZAMBIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UGANDA, TURKEY, TUNISIA, TRINIDAD &amp; Tobago, THAILAND, UNITED REPUBLIC OF TANZANIA, PROVINCE OF CHINA TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SUDAN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLANDS, NEPAL, MOZAMBIQUE, MEXICO, MALTA, MALAYSIA, MACEDONIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KENYA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GUADELOUPE, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, KOREA, EL SALVADOR, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE , CANADA, BRITISH INDIAN OCEAN TERRITORY, BRAZIL, BOLIVIA, BENNIN, BELGIUM, BANGLADESH, BAHAMAS, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Terminated | March 04, 2016 | 1792 | 31846 | Not_Computer
Z-3213-2011 | 58820 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Crystal Metal Miller Disposable Blades (sizes 0 art no: 50676 and 00 part no 50647 |  |  | March 10, 2011 | 2011 | Penlon, Ltd. (Fmly East Healthcare) | Disposable laryngoscope blades incorrectly labeled with wrong size. | Other | Penlon Ltd ( frmly East Healthcare ) sent an " Urgent Field Safety Notice" dated March 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed to check their stock with expiry dates between August 2010 to July 2015, and ensure that the sizes written on the blades, inner plastic bags, and the carton all match.Please complete the reply form below and send to the address provided. If you have any stock where the sizes on the bags, blades, and carton do not match, quarantine the blades and contact Penlon using the information below.For further questions please call 44 1235 547093. |  | 5 | 5 cartons (20 blades each) | Worldwide Distribution -- USA (nationwide) including the states of CA, VA and WI., and the countries of Belgium, Ireland, France, Netherlands, Norway, Portugal, Switzerland, UK, Hong Kong, Greece, Bahrain and China. | Terminated | September 21, 2012 | 561 | 5 | Not_Computer
Z-3211-2011 | 59438 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Victory 10, 3 Wheel (SC610) |  |  | June 22, 2011 | 2011 | Pride Mobility Products Corp | Contract manufacturer error. Component part(s) (fuses) specifications were not adhered to as outlined in Pride's Engineering drawings. | Nonconforming Material/Component | Pride Mobility Products Corp. sent a "RECALL OF VICTORY 3 AND 4 WHEEL UNITS; REPLACEMENT OF FUSES" letter dated June 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.Customers are asked in the letter to locate their affected Victory unit(s) and inform users of the recall. A letter has been provided by the firm to distribute to users alerting them to the recall. Additionally, customers are to set up an appointment with users to replace the fuses. Customers are instructed to contact Pride's Upgrade department at 866-222-6212 to order the kit and for questions regarding this recall. . |  |  |  | Worldwide Distribution-USA (nationwide) including the District of Columbia and Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and countries of Australia, Canada, France, and United Kingdom. | Terminated | November 20, 2014 | 1247 |  | Not_Computer
Z-3202-2011 | 59182 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | Cutting Burr labeled in part:"***S-1504TD QTY:1***Twist Drill for 1.5 mm x 4 mm Screw. Use with SHORT Attachments.***ANSPACH***Manufacturer Made in USA***The Anspach Effort, Inc***4500 Riverside Drive***Palm Beach Gardens, FL 33410, USA***Tel (800) 327 6887 / +1 561 627 1080***Anspach Europe Ltd.***Aston Court, Kingsmead Business Park***Frederick Place***High Wycombe HP11 1LA***United Kingdom***Tel: +441494616126***Sterilized using Irradiation***Do Not Re-use***Keep Dry***Caution; consult accompanying documents***Sterile unless damaged or opened***Rx Only***". |  |  | March 17, 2011 | 2011 | The Anspach Effort, Inc. | The Anspach Effort, Palm Beach Gardens, FL, is recalling their Single Use, Sterile Bone Cutting Burrs, Catalog #, S-1504TD, Lot #s D363039140. D323038110, D293037349 and Catalog # TURQ-434-4, Lot #s D343038845, D323038131, due to the burr being 2.0 mm longer than stated on the label. As a result, a hole drilled using this burr would be 2 mm deeper than the screw. | Mixed-up of materials/components | Anspach Effort, Palm Beach Gardens, FL sent a Customer letter dated March 31, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to screen their inventory, contact Anspach Customer Support at (800) 327-6887 if any of the recalled product was found, and complete and return the attached reply form. If the product had been further distributed, they were asked that the recall information be forwarded as appropriate. |  | 599 | 599 | Worldwide Distribution -- USA (nationwide) including the states of CA, CO, IN, KY, MD and TX. and the countries of Germany, Japan, Taiwan, United Kingdom and Sweden. | Terminated | October 11, 2011 | 208 | 628 | Not_Computer
Z-3210-2011 | 59826 | No | IRP | Powered inflatable tube massager. | Physical Medicine | Physical Medicine | 510(k) | Game Ready Professional Therapy System, |  |  | August 23, 2011 | 2011 | CoolSystems, Inc. | It was discovered that some Game Ready Professional Therapy System units that were to have been quarantined pending final inspection and testing were inadvertently moved to finished goods inventory and shipped. | Employee error | The firm, CoolSystems, Inc., contacted customers by phone on August 23, 2011. The customers were informed that some Game Ready Professional Therapy System units have been quarantined pending final inspection and testing were inadvertently moved to finished goods inventory and shipped. Instructions were given to immediately cease fulfilling orders from that inventory. Loaners were shipped to the customers overnight. The customers were instructed to replace the shipped device with the loaner that was sent and return it to the company ASAP. If you have any questions, contact the Vice President, Quality Assurance and Regulatory/Clinical Affairs at 510-868-5378. |  | 26 | 26 units | Nationwide distribution: CA, MA, MO, NH, OH, TX, UT, and WA. | Terminated | November 14, 2011 | 83 | 26 | Not_Computer
Z-3205-2011 | 59473 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | PCR Eleva Radiological Image Processing System with software R.1.2 |  |  | July 27, 2011 | 2011 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be interpreted as a RIGHT marker for side indication if the image is mirrored again within a PACS system. This could potentially lead to a misdiagnosis | Software design | Philips Healthcare Inc. sent out "Urgent --Medical Device Correction" notices dated June 10, 2011 to all affected customers.The letter described the product and problem. Customers were informed the "burned-in Mirror-Icon" may be misinterpreted as a Right marker. Customers are advised to preferably use lead markers during each X-ray acquisition or alternatively electronic annotations on each image for side indication. The presentation of these markers within each image will clearly show the difference between the intended Left/Right indication and the little mirror ICON. Philips will install improved software that prevents the misinterpretation of the symbol and will schedule the installation of the updated software.For questions on this recall call Phillips Healthcare Inc. at 978-687-1501. |  | 197 | 197 units | Worldwide Distribution: Nationwide distribution; and including the countries of Albania, Angola, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Croatia, Denmark, Egypt, France, Germany, Ghana, Greece, Greece, Haiti, Hungary, Indonesia, Iraq, Israel, Italy, Kazakhstan, Kenya, Latvia, Libya, Liechtenstein, Lithuania, Malawi, Malta, Martinique, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Portugal, Puerto Rico, Romania, Russia, Saudi-Arabia, Serbia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Tanzania, Thailand, Tschech. Rep, Turkey, Turkmenistan, UAE, Uganda, and United Kingdom. | Terminated | January 12, 2017 | 1996 | 197 | Software | Display/Image | Software update | Software Update
Z-3208-2011 | 59681 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, 510 (k) #K060937, Model #728306, Serial Numbers: 100019, 100023, 100087, 100103, and 100170. The recalled system units are equipped with operating software version; 3.2.0.19011. |  |  | June 15, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. Philips Medical's Field Service Engineers are installing software upgrades to certain of their Brilliance iCT and Brilliance ICT SP CT X-ray system units in order to correct non-conformances evidenced in these systems due to software errors. | Other | Philips Medical ( Cleveland) Inc. indicated in the Correction and Removal Report dated June 21, 2011, that a Philips Field Service Engineer will be contacting each site in order to schedule a time to implement the Field Change Order. After which time, the Field Service Engineer will visit each customer site to install the software upgrade and the firm will be tracking the progress of this corrective action.For further questions please call (440) 483-7000. |  | 5 | 5 units | Worldwide Distribution -- USA (nationwide) including the states of DC, IL, IN, OH, PA and VT and the countries of Belgium and Israel. | Terminated | February 22, 2013 | 618 | 8 | Software | Device Operation | Software update | Software Update
Z-3204-2011 | 58474 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Inc., Millennium and HD-120 MultiLeaf Collimators (MLC), a component of the Clinac, All Varian Clinac using the collimator cover P/N 1106900-xx, Model Numbers: |  |  | March 17, 2011 | 2011 | Varian Medical Systems, Inc. | An improperly installed collimator cover may detach and fall from the machine, potentially striking the patient or nearby personnel. | Component design/selection | On 3/17/2011, Varian Medical Systems sent an Urgent Medical Device Correction Letter to all affected users, with a description of the problem and user corrective action steps. |  | 3791 | 3791 units | Worldwide distribution including USA, Australia, Argentina, Italy, Morocco, South Korea, Belgium, Canada, Netherlands, Denmark, Japan, Hong Kong, Malaysia, Hungary, New Zealand, Sweden, Taiwan, Switzerland, Romania, Brazil, Slovakia, India, Colombia, UK, Ireland, Spain, Singapore, Martinique, Reunion, Venezuela, Iceland, Portugal, Turkey, Israel, Bahrain, Macau, Chile, Saudi Arabia, Latvia, China, Finland, Kazakhstan, Thailand, France, Poland, and Germany. | Terminated | June 20, 2012 | 461 | 3791 | Not_Computer
Z-3194-2011 | 59586 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console v 8.6 with ARIA RT Chart v8.5, all of 4DITC v 8.3; |  |  | July 21, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The 4D Integrated Treatment Console may erratically fail to load the planned physical wedge accessory without warning or notification when receiving an approved plan from the oncology information system. | Pending | Urgent Medical Device Correction/Field Safety Notification letters were sent to all US consignees by Certified Mail on July 21, 2011., as well as Varian Sales, Marketing and Service Organizations. Varian International Customer Support managers will notify international customers by appropriate methods. The letter informed customers of the issue and the corrective action steps. Letters request that all appropriate personnel working in the radiotherapy department be notified of the contents of the letter. A new release version is being created to correct the issue. All sites will be notified when it is available.Questions or concerns should be directed to 1-888-827-4265. Contact information for international customers is also provided as well as a list of all affected serial numbers. |  | 1125 | 1125 | Worldwide Distribution -- USA, China, Austria, Belgium, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Germany, Guadeloupe, Ireland, Italy, Latvia, Lithuania, Martinique, Netherlands, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Turkey, United Kingdom, Singapore, Algeria, Bangladesh, Belarus, Egypt, Guam, Hong Kong, India, Indonesia, Iraq, Israel, Kenya, Kuwait, Macau, Malaysia, Philippines, Qatar, Russia, South Africa, South Korea; Taiwan, Thailand, Tunisia, Ukraine, United Arab Emirates, Japan, Brazil, Chile, Columbia, Dominican Republic, Ecuador, Mexico, Peru, Venezuela, New Zealand, and Australia. | Terminated | August 01, 2013 | 742 | 1125 | Not_Computer
Z-3214-2011 | 59528 | Yes | DWB | Roller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns 8000 Cardioplegia Monitor, Catalog Number 16414; also known as PROG COMP BD AMPRO 420 CPG 8K, Catalog Number 816373 |  |  | August 29, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | The Cardioplegia Monitor for the Sarns Modular Perfusion System 8000 may fail to track cardioplegia delivery volume. | Component change control | Terumo Cardiovascular Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 29, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter advises customers to make sure all users are aware of the notice. Customers are instructed to complete and return an attached Customer Response Form. Terumo will replace the software on all monitors of the affected product.Customers with questions or concerns should call Terumo CVS Customer Service 1-800-521-2818. |  | 85 | 85 | Worldwide Distribution--USA (nationwide) including the states of AL, AR, CA, CO, GA, IL, IN, MS, NC, OH, TN, VA, and WI and countries of Belgium, Columbia, Hong Kong, India, Singapore, and Taiwan. | Terminated | March 08, 2013 | 557 | 85 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-3207-2011 | 59290 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes T1 Matrix Locking Cap of the Synthes Matrix System, Part Number 04.632.000. |  |  | March 11, 2011 | 2011 | Synthes USA (HQ), Inc. | Labeling Correction -- The locking cap was not able to be loosened, resulting in the breakage of the screwdriver or the rod to be cut and the assembly to be removed (screw, cap, and rod). This poses a potential risk of significant prolongation of OR time. | Labeling Change Control | The firm initiated this action on March 11, 2011 by issuing an Urgent: Medical Device Labeling Correction letter to all consignees. The letter identified the affected product and stated the reason for the correction. It also instructed the consignee to review and attach the supplementary technique page to any MATRIX Spine System Technique Guides that they may have had in their facility. Customers are to complete the Verification Section of the letter and return it to the firm via mail, fax, or e-mail. If customers have questions, they can call 610-719-5063 or their Synthes Spine Sales Consultant. |  |  |  | Nationwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI. | Terminated | December 18, 2014 | 1378 |  | Not_Computer
Z-3221-2011 | 59519 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101: Aquilion 32/64; System, X-Ray, Tomography, Computed. |  |  | December 08, 2010 | 2010 | Toshiba American Medical Systems Inc | There is the potential that the current may not be controlled correctly in examinations using the ECG modulation function, which adjusts the x-ray output in synchronization with the ECG waveform. | Software design | Toshiba America Medical Systems (TAMS) issued an Urgent: Medical Device Correction notification to consignees via USPS return receipt mail. The letter identified the affected product, the problem, and corrective action. The firm plans to install revised firmware to prevent the occurrence of the problem. A Toshiba Service Representative will contact customers to schedule an appointment for installation. The firm asked customers to stop using the Power Reduction option during ECG modulated scanning until the new firmware has been installed. Customers are to share the information in the letter with all users, reviewing radiologist, and the clinical engineering or Biomedical group at their facility. Customers are also to complete and return the attached reply form via fax or e-mail to TAMS as proof of notification. Questions pertaining to the letter should be directed to 800-421-1968 or customers' local Toshiba Representative at 800-521-1968. |  | 528 | 528 units | Nationwide Distribution. | Terminated | March 18, 2013 | 831 | 528 | Software | Output/Calculation | Firmware update | Software Update
Z-3217-2011 | 59359 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD MAX DNA Extraction Kit DNA-2 (Whole Blood) Catalog Number 437502. Kit contains 24/2D barcode sample preparation reagent tubes *** For use with the BD MAX System ***to extract DNA from Gram-negative bacteria and viruses, which may be present in clinical specimens. The specimen is mixed with BD MAX Sample Preparation Reagent and processed using the BD MAX System. |  |  | May 24, 2011 | 2011 | Becton Dickinson &amp; Co. | Sample results maybe switched between time it was scanned with handheld barcode reader and time it was placed in the sample rack, associating sample results to wrong patient. | Nonconforming Material/Component | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated May 2011, to all affected customers. The letter identified the products, the problem, and the action to be taken by the customers. The letter advised customers to discontinue use of the Sample Preparation Reagent tubes in the referenced kits and discard any remaining products. Customers were requested to complete an attached reply form to acknowledge their receipt of the notification and fax it to (410) 316-4258. For further assistance regarding replacements contact BD Customer Service Dept. 1-800-675-0908. For all other inquiries, contact BD Technical Services Dept. 1-800-638-8663. |  | 9 | 9 kits | Worldwide Distribution-USA (nationwide) including the states of CA, KY, MD, MI, NC, NY, TX, and WI and the country Belgium. | Terminated | October 12, 2011 | 141 | 392 | Not_Computer
Z-3254-2011 | 59770 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Circumcision trays and packs listed as follows: |  |  | May 05, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | The firm distributed surgical kits which contained circumcision clamps that were subsequently recalled by Surgical Design. | Other | DeRoyal Industries, Inc. sent a recall notification letter dated 05/05/2011. A second notice was sent on 07/26/2011 to affected customers who did not respond to the initial notice. Questions regarding this recall may contact the Corporate Attorney at (865) 362-1037. |  | 390 | 390 units | USA (nationwide) including the states of AR, AZ, FL, GA, LA, MI, NJ, NM, and OH. | Terminated | January 30, 2013 | 636 | 490 | Not_Computer
Z-3255-2011 | 59683 | No | LPB |  |  | Cardiovascular | PMA | 7 F Therapy Ablation Catheter, |  |  | February 10, 2011 | 2011 | St Jude Medical | The recall was initiated because St. Jude Medical has confirmed that the Therapy Ablation Catheter, 4 mm Tip Thermistor Quadripolar, were packaged with the incorrect Instructions for Use (IFU). To ensure accuracy of device labeling, St. Jude Medical is requesting that the incorrect IFU be replaced with the enclose, correct version. | Process control | St. Jude Medical sent an "important Medical Device Safety Information" letter dated February 22, 2011 to all affected customers.The letter includes: affected problem, reconciliation form and corrective action. Customers are asked to remove the incorrect Instructions for Use (IFU) and use the IFU provided in the letter, and to fill out and return the reconciliation form.For questions on this recall please contact the firm at (512) 329-1054 or (651) 756-2869. |  | 179 | 179 units total | Nationwide Distribution including the states of TN, MS, IN, MO, FL, WI, CA, and MT. | Terminated | April 03, 2012 | 418 | 179 | Computer
Z-3215-2011 | 59743 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific 25ga (0.52mm) Expect Endoscopic Ultrasound Aspiration Needle REF 5002 Use By 1999-12 LOT 13615651C1 STERILE EO |  |  | August 11, 2011 | 2011 | Boston Scientific Corporation | The firm has determined that a single lot of 25 ga Expect EUS/FNA devices were mislabeled. The carton label incorrectly described the devices as 22 ga. The inner pouch label incorrectly describes the device as 25 ga. | Other | Boston Scientific Corporation sent a "URGENT MEDICAL DEVICE RECALL letter dated August 11, 2010, to all affected customers. The letter identified the product, the problem and the action needed to be taken by customer.Customers were instructed to immediately remove the affected lot of Recall products from your inventory ( whether in Labs, Central Supply, Shipping and Receiving or any other location). Also, segregate this product in a secure location for return to Boston Scientific. Complete and return the Reply Verification Tracking Form ( even if you do not have any product to return) following the directions on this page and Reply Verification Tracking Form. Fax the Form to (508) 683-5578. Further questions please call (508) 683-4678. |  | 30 | 30 units | Nationwide Distribution including the states of VT, FL, SC, NC, WA, MA, and MN. | Terminated | November 10, 2011 | 91 | 30 | Not_Computer
Z-3222-2011 | 59704 | No | MJJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Custom Procedure Kits, REV OA |  |  | August 11, 2011 | 2011 | Terumo Cardiovascular Systems Corp | 3-T Medical Systems one way vent valves, included in the Custom Terumo Cardiovascular kits, may separate at the bond that holds the two halves | Nonconforming Material/Component | Terumo Cardiovascular Systems initiated contact on August 12, 2011 by phone, fax, email and Urgent: Medical Device Recall mail to all affected customers.The firm advised customers of the affected product, problem and required recall actions. Customers can remove whole customized packs that are not in urgent medical need and return packs. Or the User may opt to use a sterile LH 130 vent valve from Terumo Cardiovascular Systems. A response form was included to return back to the firm.For questions on this recall contact Terumo Customer Service at (734) 525-9580. |  | 34 | 34 units | Nationwide Distribution, including the states of FL and MI. | Terminated | March 20, 2012 | 222 | 4233 | Not_Computer
Z-3259-2011 | 59794 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS Angiographic Catheter, Soft-Vu Cobra (2) BRAIDED, 5 F x 90 cm x .035 in., Catalog No./REF 10709105, STERILE, MANUFACTURED IN USA, AngioDynamics, Inc. 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | June 07, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The lot of Soft-Vu Catheters was labeled 90 cm in length; however the product was cut to a 65 cm length. | Employee error | The firm ANGIODYNAMICS, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2011 to the consignees via Certified Mail, Return Receipt Requested. The letter described the product, problem and action to be taken. The customers were instructed not to use the product and to return it to AngioDynamics, Inc. in Queensbury, NY. The letter also instructed the consignees to (1.) Identify and segregate the recalled lot that is in your possession, (2.) Complete the enclosed Soft-Vu Angiographic Catheter Recall Reply Form and fax it to the attention of the Soft-Vu Angiographic Catheter Recall Coordinator at 518-798-1360. It is important that even if you do not have any product remaining in your possession that you fill out the attached form noting zero quantity to be returned and fax the form to ANGIODYNAMICS. (3.) Ship the recalled product to ANGIODYNAMICS. Replacement product will be shipped upon receipt and confirmation of the returned product. Reference Return Authorization Number XXXXXXXXXX on the outside of the shipping box and include a copy of the Soft-Vu Angiographic Catheter Recall Reply Form with your shipment. If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call the Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. You may also e-mail your questions to customerservice@angiodynamics.com. |  | 30 | Domestic - 25 units; Foreign - 5 units | Worldwide distribution: USA (nationwide) including states of: California, Tennessee, and Wisconsin, and country of: Spain. | Terminated | October 18, 2012 | 499 | 30 | Not_Computer
Z-3266-2011 | 59820 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS UNI*FUSE INFUSION SYSTEM WITH COOPER WIRE, Catalog No./REF 12401818, CATHETER SIZE 4F, CATHETER LENGTH 90 cm, INFUSION LENGTH 10 cm, OCCLUSION WIRE SIZE .035 in., STERILE, AngioDynamics Inc., 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | January 28, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The particular lots of UniFuse Catheters were packaged with the wrong occluding ball guidewires. | Employee error | URGENT MEDICAL DEVICE RECALL LETTERS (dated 1/28/11) and Recall Reply Forms were sent to the consignees via Certified Mail, Return Receipt Requested. The letters identified the scope of the recall, the reason for the recall, the status, and action to be taken. Customers were to identify and segregate the recalled lot in their possession. They were also to complete and fax back the Reply Form. The recalled product should be returned along with a copy of the Reply Form. Replacement product would be shipped upon receipt and confirmation of the returned product. Questions or concerns should be directed towards customers' local sales representative or call Julie Blair, Customer Service Manager at 1-800-772-6446. |  | 23 | 23 units | Worldwide Distribution -- USA and The Netherlands. | Terminated | October 18, 2012 | 629 | 47 | Not_Computer
Z-3263-2011 | 59726 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | COULTER DxH Cell Lyse, Part number 628018(1L) and 628019 (5L). |  |  | July 11, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter identified unexpected high particle counts in the lots of DxH Cell Lyse listed. These lots may generate high background counts;could result in a falsely elevated Total Nucleated Count (TNC) of up to 50 cells/ÂżL; or may also result in recovery of TNC values that were higher than expected for Level 1 of the Body Fluid Control. | Pending | The recall communication was initiated on 08/01/2011 with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with attached PCA Response Form (via US Postal Service for US customers) to all customers, who purchased Coulter DxH Cell Lyse. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to contact Beckman Coulter for product replacement. In addition, customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with questions regarding this Product Corrective Action were instructed to call Beckman Coulter Customer Service at 800-526-7694 in the United States or Canada, or contact their Beckman Coulter representative. |  | 4913 | 4913 units total (2245 in US) | Worldwide Distribution -- Australia, Belgium, Brunei Darussalam, Canada, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Macao, Malaysia, Netherlands, Norway, Philippines, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and the United States. | Terminated | June 04, 2012 | 329 | 4913 | Not_Computer
Z-3262-2011 | 59831 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS PROFILER BALLOON DILATATION CATHETER, Catalog No./REF 16801105, BALLOON DIAMETER 5MM, BALLOON LENGTH 4 CM, CATHETER LENGTH 40 CM, STERILE, MANUFACTURED IN USA AngioDynamics 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | February 25, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The particular lots of Profiler PTA Balloon Catheters were labeled incorrectly. Lot 532109 contains product that is actually 5 mm x 4 cm x 40 cm and Lot 532794 contains product that is actually 4 mm x 2 cm x 40 cm. | Employee error | Angiodynamics, Inc sent a "Urgent Medical Device Recall letter" dated February 25, 2011 to all affected customers via Certified Mail Return Receipt Requested. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed the following steps1. Identify and segregate the recalled lot that is in your possession.2. Complete the enclosed Profiler PTA Balloon Catheter Recall Reply Form and fax it to the attention of the Profiler PTA Balloon Catheter Recall Coordinator at 518-798-1360.The form lists the catalog numbers, lot numbers, and quantity our records indicate your facility has received.For further questions please call 1-800-772-6446. |  | 14 | 14 units | Nationwide Distribution including the states of CA, GA, ID, MI, OH and PA. | Terminated | October 18, 2012 | 601 | 29 | Not_Computer
Z-3268-2011 | 59467 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Contour, Contour TS Blood Glucose Test Strips and Bonus Packs |  |  | August 12, 2011 | 2011 | Bayer Healthcare, LLC | All 10- or 25- count glucose test strip vials, either loose vials or found in patient starter kits of Contour, Contour TS and Contour" USB blood glucose meters are being recalled a packaging issue which, under certain conditions, can cause some test strips packaged in some small-count vials to operate outside of US FDA-cleared performance specifications and produce a low-bias reading during blood | Packaging | Bayer sent an "URGENT DIABETES TEST STRIP RECALL" letter dated August 12, 2011, to its customers with the name of the test strips, sizes recalled, reason for the recall, potential injury, and instructions to discontinue use and return the test strips to Bayer. The recall letter also included a recall response form that was to be faxed back to Bayer at 1-800-876-2243. For questions call Customer Logistics at 1-800-348-2637. |  | 197342756237020 | 3,617,287 vials/ 54,555,460 test strips | Nationwide Distribution and Puerto Rico | Terminated | July 02, 2012 | 325 | 197342756237020 | Not_Computer
Z-3265-2011 | 59303 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart Automated External Defibrillator (AED), G3 is indicated for emergency treatment of victims exhibiting symptoms of sudden cardiac arrest who are unresponsive and not breathing. Post-resuscitation, if the victim is breathing, the AED should be left attached to allow for acquisition and detection of the ECG rhythm. If a shockable ventricular tachyarrhythmia recurs, the device will charge automatically and be able to deliver therapy. |  |  | June 30, 2011 | 2011 | Cardiac Science Corporation | Cardiac Science Corporation (CSC) has identified two errors, 0x48 and 0x99, that have repeatedly occurred during a device self-test in the G3 installed base over the last 6 years. These self test errors cause the device to enter into a service required state.CSC remind owners/operators of the Powerheart AED devices to verify the "Rescue Readiness" of their equipment and follow maintenance pro | Device Design | Cardiac Science began mailing out the recall notification letter on August 1, 2011, to all their US consignees via USPS first class mail. Cardiac Science's distribution partners will receive a notification letter and be provided with a template letter to send to their customers directly.CSC remind owners/operators of the Powerheart AED devices to check the status of of the AED every day to ensure it is Rescue Ready, and perform the regularly schedule maintenance as outlined in the Powerheart AED "Operator and Service Manual" provided with every AED.Customers can call Cardiac Science at 800-426-0337 or e-mail at aed205@cardiacscience.com. |  | 195739 | 195,739 units in US | Worldwide Distribution - US (nationwide)This action applies only to domestic | Terminated | February 13, 2014 | 959 | 195739 | Not_Computer
Z-3271-2011 | 59865 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VASOVIEW 7xB Endoscopic Vessel Harvesting System; |  |  | December 17, 2009 | 2009 | Maquet Cardiovascular, LLC | During routine production quality monitoring process, it was found that UV curing adhesive was not used in a limited number of devices. | Nonconforming Material/Component | MAQUET Cardiovascular, LLC, sent an "Urgent Device Removal/Immediate Action Required" letter dated December 17, 2009, to all affected customers. The letter identified the recalled products, the problem, and the actions needed to be taken. The letters asked customers to examine their stocks (inventory) immediately to determine if they had any affected lot numbers and to discontinue dispensing (distributing) the lot. The letter states that customers are to complete the attached Field Action Response form regardless of whether or not the affected products are located in their inventory. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. The letter instructed all users to return all unused affected products to MAQUET. For customer's convenience, all affected products that are returned may be exchanged with other unaffected VASOVIEW Endoscopic Vessel Harvesting System by overnight delivery. Questions should be directed to MAQUET Cardiovascular sales representative or Customer Service at 1-888-880-2874. |  | 177 | 177 units | Nationwide Distribution (USA) including the states of: AL, AZ, CA, CT, FL, KS, KY, IL, IN, MI, MO, ND, NY, NV, OH, OK, OR, PA, SC, TN, TX, UT, WA, and WI. | Terminated | September 21, 2011 | 643 | 390 | Not_Computer
Z-3269-2011 | 59664 | No | ODX | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CombiSetÂż True Flow" 2008K@home Bloodlines |  |  | August 11, 2011 | 2011 | Fresenius Medical Care Holdings, Inc. | Blood lines coiling tabs #1 and #5 on the arterial line are not in the correct positions | Device Design | Fresenius Medical Care sent out an "Urgent" recall notice dated August 11, 2011 to all affected customers .The letter included description of the product name, problem and asked them to examine their stock immediately and quarantine bloodlines from the recalled lots. Customers were asked to return the product to Fresenius.For information on this recall call the Fresenius Customer Service Team at 1-800-323-5188. |  | 118 | 118 cases | Nationwide Distribution. | Completed |  |  | 118 | Not_Computer
Z-3273-2011 | 57865 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Synthes Bone Marrow Aspiration System, Sterile, 11 gauge, 15cm needle (without side holes) |  |  | January 26, 2011 | 2011 | Synthes USA (HQ), Inc. | Firm became aware that some systems were released to stock without going through the sterilization process. | Nonconforming Material/Component | Synthes (USA) sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. Affected product removed from inventory will be replaced by Synthes. Contact the Synthes Spine Sale Consultant at 1-800-620-7025 ext.6883 for questions concerning this recall. |  | 96 | 96 parts | Nationwide Distribution-including the states of AR, CO, FL, GA, ID, LA, MA, MD, MO, MT, NY, NC, OH, SC, SD, TN, TX, and VA. | Terminated | April 16, 2015 | 1541 | 96 | Not_Computer
Z-3274-2011 | 59791 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker SmartPump Extension Tubing (Pneumatic tourniquet). |  |  | August 25, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | Three complaints have been reported indicating that the product inside the package did not match the label description. The units had quick connectors instead of luer lock connectors as stated in the label description. | Mixed-up of materials/components | Stryker Instruments sent out "Urgent Medical Device Recall" notices dated August 22, 2011 to all affected customers.The letter included the name of the recalled product, lot numbers, and the reason for the recall. Customers were instructed to send back any product to Stryker that is within the affected lots. The letter further stated Stryker will quarantine the product and replace any product that is returned. For information on this recall contact Stryker Instruments at (269) 389-4534. |  | 5920000003 | 5920-000-003 (30 bx) and 5920-000-005 (30 bx) | Worldwide Distribution: Nationwide Distribution including the states of HAWAII, NORTH DAKOTA, NEBRASKA, PENNSYLVANIA, and IOWA, and the country of FRANCE | Terminated | October 12, 2011 | 48 | 5920000003 | Not_Computer
Z-3277-2011 | 53897 | No | NSX |  |  | General Hospital | Contact ODE | CareSuite" - Critical Care Manager, PACU Manager and Anesthesia Manager Software, transmission &amp; storage, patient data. |  |  | August 12, 2009 | 2009 | Picis Inc. | Electronic health record software applications malfunctioned in a manner which failed to display the discontinued status for medication orders, as entered by the attending physician | Software design | Picis Inc. sent a customer notification letter dated August 11, 2009 to all affected customers. The letter describes the product, problem, and action to be taken. Customers informed of the potential patient safety error were advised to make the appropriate configuration change, within current software versions 7.1.4 or later. Contact your local Picis support representative for questions regarding this notice. |  | 8576 | 8576 UNITS | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Brussels, Belgium, Denmark, Finland, France, Germany, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and UK. | Terminated | September 27, 2011 | 776 | 8576 | Software | Display/Image | Instructions on changing configurations | Safety Notice/Insructions
Z-3275-2011 | 59908 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OPMIÂż PenteroÂż, OPMIÂż PenteroÂż C, OPMIÂż PenteroÂż 900, |  |  | September 12, 2011 | 2011 | Carl Zeiss Meditec, Inc. | Some third party surgical drapes have optically active protective lenses that could affect the ZEISS surgical microscope position information that is passed on to third party navigation systems. | Pending | The firm, Zeiss, sent an "IMPORTANT SAFETY ALERT FOR ZEISS SURGICAL MICROSCOPES" letter dated September 7, 2011, to all affected customers. The letter described the product, problem and action to be taken. The customers were instructed to use surgical drapes compatible with ZEISS surgical microscopes and always follow the instructions for use of the third party navigation system manufacturer for calibration, calibration verification and operation. In addition, the firm supplied customers with a supplement to their Zeiss surgical microscope user documentation and requested that they read the information and keep a copy with their original instructions also the customers were instructed to complete and return the Acknowledgement of Receipt form within 5 business days via fax to 1-866-809-6044 or USPS to Stericycle, Inc., Attn: Event 2480, 2670 Executive Dr. Suite A, Indianapolis, IN 46241.Should you have any questions about this Safety Alert, please contact your Carl Zeiss Customer Care representative at 1-888-871-7120. |  | 1677 | 1677 consignees | Nationwide distribution: USA only. | Terminated | February 24, 2012 | 165 | 1677 | Not_Computer
Z-3276-2011 | 59734 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LOGIX Order Entry (OE) Software, all versions, catalog 2M8410; Baxter Healthcare, Deerfield, IL 60015; The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. |  |  | August 15, 2011 | 2011 | Baxter Healthcare Corp. | There is the potential for an order entry error during use of the LOGIX Order Entry (OE) software. When adding ingredients to a TPN order, the user can accept the pre-populated "per units" of "Bag" or select other "per units". A dosing error may occur if the user does not select the prescribed "per units" for the added ingredient. | Software design | Baxter Healthcare Corp. sent an "Urgent Device Correction" letters dated August 15, 2011 which included LOGIX OE screen illustrations of a representative neonate and adult patient, via first class mail to all affected customers. The letter identified the product, problem and actions to be taken. The customers were instructed to follow safe compounding practice, which includes verifying that the ingredient per units entered match the TPN prescription prior to preparation. Baxter requests that customers communicate this issue to all personnel responsible for LOGIX order entry and order authorization at their facility. Customers are instructed to retain a copy of the letter at each computer on which LOGIX OE software is installed. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Baxter's Medical Information Services at 1-800-422-2751. |  | 95 | 95 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the country of Canada. | Terminated | January 05, 2017 | 1970 | 95 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-3258-2011 | 59825 | No | OAB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ANGIODYNAMICS NANOKNIFE IRE GENERATOR, Catalog/REF No. 20300101, ANGIODYNAMICS, 603 Queensbury Ave., Queensbury, NY 12804 |  |  | April 04, 2011 | 2011 | Angiodynamics, Inc. | Some units of Nanoknife have the potential to deliver a shortened pulse width without indicating an error. The affected units contain a "GOMe" circuit board which may cause the generator to drift from the calibration points set at the time of manufacture. | Component design/selection | AngioDynamics, Inc. sent a "NOTIFICATION OF MEDICAL DEVICE CORRECTION" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and action to be taken. The letter instructs customers to review the measurement graphs (displayed on generator) after each delivery of energy, and to contact AngioDynamics Customer Service at 1-800-772-6446 if shortened pulses are detected. Contact your local sales representative or call Customer Service for any additional concerns or questions regarding this notice. |  | 21 | 21 units | Worldwide Distribution--USA (nationwide) and the countries of Australia, Germany, Israel, Italy, and Spain, | Terminated | January 29, 2016 | 1761 | 21 | Hardware | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-3260-2011 | 59877 | No | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom Dialysis prep kits labeled as follows: |  |  | August 15, 2011 | 2011 | Lee Medical International Inc | The firm used alcohol triple swabstick and prep pads manufactured by H &amp; P Industries in the manufacture of their custom dialysis kits. Those swabsticks and pads were subsequently recalled due to possible bacterial contamination. | Nonconforming Material/Component | The firm, Lee Medical International Inc., notified their consignees by email and telephone on August 17, 2011. Consignees were instructed to remove the affected product from stock, quarantine for re-processing, and return or destroy product. Lee Medical arranged for replacement product from alternate manufacturer to be sent to customers. For questions call Production Manager or Operations Manager at 504-734-9336 or email melissal@eleemedical.com or chriss@eleemedical.com. |  | 4075 | 4,075 units | Nationwide Distribution (USA) including the states of: CA, FL, GA, IA, LA, MD, NY and TX. | Terminated | January 30, 2012 | 168 | 4075 | Not_Computer
Z-3278-2011 | 59862 | No | CJE | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Alkaline Phosphatase, List Number 7D55-21 &amp; 7D55-31. |  |  | August 26, 2011 | 2011 | Abbott Laboratories, Inc | The reagent has the potential to be contaminated with mold, the Cladosporium fungal species. | Pending | The firm decided to recall and notified consignees by a Product Recall - Immediate Action Required letter on 08/26/11. The letter identified the affected product, the reason for the recall, the patient impact, and necessary actions to be taken by the customer. Customers are to determine if they are using or have the affected product in their inventory. Customers are to discontinue use and destroy any remaining product of the specified lots in accordance with their facility's policies and procedures. Alternative lots of the product should be ordered and the enclosed Customer Reply form should be completed and returned. The letter also states that customers may want to decontaminate their systems by following the steps in Attachment A. If the product from the affected lots were further distributed, a copy of the letter should be provided to those customers as well. Questions regarding the letter should be directed to Customer Service at 1-877-4ABBOTT for US customers, and customers outside of the US should contact their local area Customer Service. |  | 6269 | 6269 units | Worldwide Distribution | Terminated | August 17, 2012 | 357 | 6269 | Not_Computer
Z-3282-2011 | 59527 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ADVANTAGE WINDOWS (Workstation) WITH FUNCTOOL OPTION |  |  | August 18, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware that image orientation of reference series displayed in Functool Diffusion Tensor/FiberTrak protocol on Advantage Workstation may not match with the orientation annotations displayed in the viewport. This issue may impact patient safety. | Software design | GE Healthcare sent "Urgent Medical Device Correction" letters dated August 18, 2011 to all affected customers.The letter included name of product affected and described problem. Recommended actions where provided to customers and they were notified a GE Healthcare representative will contact them for a software system upgrade.For question of this recall contact the GE Service Representative at (800) 437-1171. |  | 697 | 697 | Worldwide Distribution: Nationwide distribution including the states of AZ, CA, CT, DC, FL, ID, IL, KY, LA, MD, MI, MN, NJ, NY, OR, PA, PR, TN, TX, VA, WA, and WI; and the countries of VIETNAM, MEXICO, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SWEDEN, ST HELENA, SPAIN, SLOVENIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂżDÂżRATION, ROMANIA, REUNION, REPUBLIC OF CHINA, QATAR, PORTUGAL POLAND, PARAGUAY, NETHERLANDS, MOROCCO, MALAYSIA, MACEDONIA, LUXEMBOURG, LIBYA, KUWAIT, KOREA, JAPAN, ITALY, ISRAEL, INDONESIA, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, CZECH REPUBLIC, CROATIA, COSTA RICA, COLOMBIA, CHILE, CANADA, BELARUS, BRAZIL, BOSNIA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | August 09, 2012 | 357 | 697 | Software | Display/Image | Software update | Software Update
Z-3280-2011 | 58556 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Dual Blue-Release Adult Walkers - Wheeled, and Blue Release Junior Walkers with 5 inch Fixed Wheels, Model Numbers: 6240-5F and 6240-JR5F. The recalled product is packaged four (4) units per shipping box. |  |  | March 02, 2011 | 2011 | Invacare Corporation | The firm is conducting a recall/correction of their adult and junior wheeled walkers due to a component defect. The firm determined that the caster bolt on the wheeled extension was too short to adequately maintain the Nylex fastener. Consequently, over time, this condition incurs the potential for the caster to loosen and detach, causing the wheels to fall off the walker. This scenario would res | Device Design | The firm, Invacare Corporation notified their customers of this recall/corrective action via an 'URGENT RECALL INFORMATION' letter dated April 7, 2011. The letter was sent to the customers by registered mail or Federal Express. The letter informs the customer(s) of the component defect involved and the inherent possibility that over time, a walker wheel(s) could fall off resulting in instability of the product and potential for serious injury to the user of the device. The letter provides a list of the affected units sold to each facility on an enclosed tracking sheet. It also asks that the customer: 1) Examine their stock for the presence of any of the affected walkers and to quarantine any such products found; 2) Complete an attached recall response form indicating the customers receipt and understanding of the instructions provided in the letter and to fax the completed form to the telephone number listed on the bottom of the form. The letter goes on to state that once the faxed back form is received by Invacare, the company will send the customer replacement wheels kits for all affected walkers purchased; 3) the customer is asked upon receipt of the replacement wheel kits to replace the wheels of any units still in stock and discard all of the old wheel parts removed; 4) Lastly, the customer is asked to contact their sub-account customers and notify them of the recall. If the customer have any questions concerning the instructions, call the recall support line at 1-866-567-7899. |  | 12764 | 12,764 walkers were distributed. | Worldwide Distribution - USA (nationwide: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, including Puerto Rico and countries of Canada, Columbia, and Israel. | Terminated | July 02, 2013 | 853 | 12764 | Not_Computer
Z-3281-2011 | 59383 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #78231, Serial Number: 9592. |  |  | June 02, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Component defect. Philips Medical decided to conduct a recall field /correction (rework) of the mounting key locking threaded insert of the Detector Module system (DMS) bolt on one (1) Brilliance 64 CT X-ray unit. The firm initiated this action following Philip's receipt of the customer's complaint reporting that the retaining inserts were pulling out of the rotor during the removal of the UDMS | Other | Philips sent a Field Change Order (FCO) 72800519 on June 13, 2011 to the affected customer. A Field Service Engineer visited the customer site and installed a hardware upgrade to the affected Brilliance 64 device system (Serial Number: 9592). The correction action was completed on August 9, 2011.For questions regarding this recall call 1-440-483-7000. |  | 1 | 1 unit | Nationwide Distribution including PA | Terminated | March 15, 2013 | 652 | 1 | Not_Computer
Z-3279-2011 | 59503 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Trabecular Metal Modular Acetabular System is a modular acetabular cup system consisting of a shell with a preassembled locking ring, a liner, and optional screws. The Shells are available in a variety of sizes and are designed to be used with Trilogy Acetabular System liners. The shell implants are used to replace the hip socket during total hip arthroplasty. |  |  | September 01, 2011 | 2011 | Zimmer Inc. | Lot # 61132698 and Lot 001122467 (added 9/ 18/2012) may have a manufacturing non conformance that could prevent the devices from functioning as intended. Specifically, devices in the affected lot may have a locking ring groove depth that prevents the corresponding liner from snapping into the shell. The groove provides the locking ring room to displace when inserting a liner into the shell. Th | Nonconforming Material/Component | Zimmer, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 27, 2011 to all affected customers of lot 61132698 and 9/18/2012 to only affected customer of lot 001122467. The letters describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to stop using and to quarantine the affected product. An Inventory Return Certification Form was included with the letter for customers to complete and return to Zimmer, Inc. via fax at (574) 372-4265. Replacement product will be provided for all returned product. Contact Zimmer, Inc. at 1-800-613-6131 for questions regarding this recall. |  | 62255181 | 24 devices lot 61132689; 1 device lot 001122467 | Worldwide Distribution-USA (nationwide) including the states of CA, NC,NY, PA, WA, and WI and countries of Australia, Germany, The Netherlands, and Switzerland. | Terminated | August 29, 2013 | 728 | 62255181 | Not_Computer
Z-3283-2011 | 59686 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, LATITUDEÂż Patient Management System Communicator, Model 6476, Distributed by Boston Scientific, 4100 Hamline Avenue North, St Paul, MN 55112-5798 |  |  | August 01, 2011 | 2011 | Boston Scientific CRM Corp | A limited number of Model 6476, 6482, and 6468 (Not available in the US) LATITUDE Communicators are unable to authenticate, and thus cannot connect to the LATITUDE System. Customers are instead receiving Latitude Temporarily Unavailable (LTU) message. | Other | Boston Scientific CRM Corporation sent a " Dear Patient", letter dated August 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A new communicator is being shipped separately to you with a corrected security certificate that will allow it to be activated and function normally.1. Please activate this new equipment at your convenience to begin home monitoring. Activation instructions are enclosed with the communicator.2. We also ask that you return the original communicator previously delivered to you:" Place the original equipment in the shipping box used for the new equipment." Pack newspaper or other filler around the equipment to prevent damage during shipping, and seal the box." Enter your address in the FROM section of the prepaid Federal Express label, and apply the return label to the outside of the box directly over the original label." Call Federal Express at 1.800.463.3339 to schedule a pick-up.We apologize for this inconvenience. If you have further questions, please call 1.866.484.3268. |  | 470 | 470 | Worldwide Distribution -- (USA) nationwide including the states of AL, AK, AZ, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE,NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY., and the countries of United Kingdom, France | Terminated | November 06, 2012 | 463 | 470 | Other | Device Operation |  | Other
Z-3284-2011 | 59953 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; the pumps were sold under the following configurations: |  |  | September 16, 2011 | 2011 | Hospira Inc. | Hospira has received reports of incorrect seating of the regulator closer on the Plum A+ Infusion Pumps. If the regulator closer does not close the flow regulator actuator when the pump door is opened, and if the clinician has not engaged the clamp prior to opening the cassette door, unrestricted flow may occur. | Nonconforming Material/Component | Hospira, Inc. sent an "URGENT DEVICE RECALL" letter dated September 16, 2011 to all affected customers. The letter describes the product, problem and actions to be taken by the customers. The letter instructs customers on how to prevent occurrences of continuous recycling/rebooting during start-up. A reply form was attached to the letter for customers to complete and return via fax to 1-888-345-5358. Contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance or information regarding this notice. |  | 265105 | 265,105 pumps | Worldwide Distribution-USA (nationwide) including Puerto Rico and the U. S. Virgin Islands and countries of Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open, Classified |  |  | 290217 | Not_Computer
Z-3288-2011 | 59312 | No | BZK | Monitoring spirometer. | Anesthesiology | Anesthesiology | 510(k) | FloTrak Elite modules |  |  | July 29, 2011 | 2011 | Respironics Novametrix, LLC. | Pressure lines maybe occluded and display incorrect tidal volumes. | Manufacturing material removal | Respironics Novametrix, LLC on July 29, 2011 notified accounts via e-mail or tracked receipt, an Urgent - Medical Device Recall letter to notify customers of the issue and to instruct them on conditions for continued use. Monitors will need to be returned for functional performance testing at their facility. Customers can contact Respironics US Customer Service at 1-877-387-3311 for customers residing in the USA or +011 (203)697-6348 for International customers, if they have any questions or need additional information. |  | 41 | 41 units | Worldwide Distribution -- USA, Belgium, Canada, Chile, Columbia, Egypt, India, Italy, Spain, Saudi Arabia Thailand, GB, and Russia. | Terminated | May 23, 2012 | 299 | 295 | Not_Computer
Z-3289-2011 | 59724 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Comprehensive Reverse Humeral Fracture Inserter |  |  | August 03, 2011 | 2011 | Biomet, Inc. | Recall initiated in response to a Biomet Canada report, which noted that while testing the instruments, the pin used to secure the implant on the device did not engage on the effected lots and were distributed. | Nonconforming Material/Component | On 8/3/11 Biomet Orthopedics sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter to all affected customers.The letter included: reason for the recall, risks associated with use of the device, and alternate instructions for use of the recalled device. Customers are instructed to locate the affected products, discontinue use, return the devices, and a fax a copy of the response form.For information or questions on this recall call Biomet at (574) 372-1570. |  | 6 | 6 units (internationally) | International Distribution Only: Canada, Australia, and Columbia. | Terminated | October 26, 2011 | 84 | 6 | Not_Computer
Z-3126-2011 | 58719 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Part Number: 81600 |  |  | April 15, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because certain Access immunoassays can produce inaccurate results when performed within the published UniCel DxI Immunoassay Systems* room temperature operating specification of 18 degrees C to 32 degrees C (64.4 degrees F to 89.6 degrees F). An increase in room temperature causes assay results for some assays to decrease, while results for other assays increase with an i | Other | Beckman Coulter sent an "URGENT FIELD SAFETY NOTICE" letter dated April 18, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter includes information on the new room temperature operating range for all UniCel DxI Immunoassay Systems and Access Immunoassay Systems during system operation. Customers were instructed to review their historical room temperature records to determine if their Access Immunoassay System and UniCel DxI Immunoassay System were operating outside the revised temperature range. An attached Response Form was included for customers to complete and return.Please contact Beckman Coulter Customer Support Center for questions regarding this notice at 1-800-854-3633. |  | 11266 | 11,266 units | Worldwide Distribution--USA (nationwide) and Puerto Rico and the countries of Algeria, Angola, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Lebanon, Liechtenstein, Malaysia, Mayette, Mexico, Morocco, Nambia, Netherlands, New Zealand, Norway, Philippines, Poland,, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, United States, Venezuela and Vietnam. | Terminated | December 21, 2012 | 616 | 14648 | Not_Computer
Z-3163-2011 | 59427 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoMobile, Hi-Art System, H-0000-0003, TomoTherapy 1240 Deming Way, Madison, WI 53717 |  |  | July 07, 2011 | 2011 | Accuray Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the TomoMobile Hi-Art System.TomoTherapy has discovered that when attempting to open the TomoMobile shielding doors, while the door hinges are in the locked position, the hinges may fail allowing the door to disengage from the shielding. | Device Design | The firm, TomoTherapy, sent an "Urgent Medical Device Correction" letter dated July 7, 2011 to its consignees/customers. The letter described the Issue, Product Affected, Recommended Action (advised consignees to not attempt to open the shielding doors when the hinges are in the locked position. The Hinges should always be in the unlocked position when opening the doors) and Resolution. A TomoTherapy field service engineer or an authorized service representative will contact the customers when a resolution is available.If you have any questions, please contact the TomoTherapy Technical Solutions Center by email at support@tomotherapy.com or phone at: North America: 1-866-368-4807; Belgium: 0800 38783; France: 0805 631 565; Germany: 0800 000 7541; Italy: 800 986 399; Netherlands: 0800 020 1364; Spain: 800 300049; Switzerland: 0800 001927; United Kingdom: 0808 238 6035; China/CNC (Northern): 10 800 712 1701; China (Southern): 10 800 120 1701; Hong Kong: 800 967912; Japan: 0044 22 132374; Singapore: 800 1204 683; South Korea: 0079 81 4800 7204, all other locations: +1 608 824 2900 or +32 2 400 44 44. |  | 4 | 4 | Worldwide distribution: USA including states of: AZ, CA and OK; and country of: United Arab Emirates | Terminated | June 29, 2012 | 358 | 4 | Not_Computer
Z-0001-2012 | 57858 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Probe, Tactile, Left, ASM, Long, |  |  | January 20, 2011 | 2011 | Medtronic Navigation, Inc | Distal end of tactile probe may break off during use. | Labeling design | Medtronic sent an "Urgent Product Safety Notice" to all affected customers dated Jan 20, 2011.The letter included affected product and updated instructions for use. On-site training was given for the use of the probes by Medtronic Representatives.For information on this recall call Medtronics at (720) 890-3409. |  | 6 | 6 probes | Worldwide Distribution: Nationwide distribution including the states of CO and WI; and the country of South Korea | Terminated | October 03, 2011 | 256 | 6 | Not_Computer
Z-0007-2012 | 59847 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES Safety Infusion System, Model 621-100ES, 621-200ES, 621-300ES, and 621-400ES |  |  | July 20, 2010 | 2010 | B Braun Medical, Inc | The Outlook ES Dose Guard Care Area can inadvertently be exited when the pump enters the KVO (Keep Vein Open) state. The same condition can occur when a sequence of HOLD-HOLD key strokes are performed on the pump key panel. | Software design | The firm, B. Braun Medical Inc., sent an "URGENT: FIELD CORRECTION" letter dated September 8, 2011 to its customers. The letter described the product, problem and actions to be taken. The firm informed the customers that a B. Braun representative will be contact ing them to schedule a field software upgrade of the affected pumps. The customers were instructed to complete and return the FIELD CORRECTION INFORMATION FORM via mail using the enclosed envelop and contact B. Braun Customer Service for shipping instructions at 1-800-627-7867 if product is to be returned.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Technical Product Support at 1-800-627-7867 (800-627-PUMP). |  | 11525 | 11,525 pumps | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL, GA, IA, IL, KS, KY, LA, MA, MI, MO, NC, NJ, NY, OH, PA, SC, TX, VA, WA, WI, and WV; and country of: Canada. | Terminated | January 07, 2015 | 1632 | 11525 | Software | Device Operation | Software update | Software Update
Z-0003-2012 | 59979 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Cochlear Nucleus CI512 Cochlear Implant, REF Z209051, New with Nucleus 5, Sterile EO, Made in Australia, Cochlear Ltd 14 Mars Road, Lane Cove 2066, Australia. |  |  | September 16, 2011 | 2011 | Cochlear Americas Inc. | Cochlear implant may shut down and cease to function. | Component design/selection | Cochlear Americas, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 16, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory and quarantine the affected product. A recall response form was attached to the letter for customers to complete and return. Contact Cochlear Clinic Customer Service at 877-883-3101 for questions concerning this notice. |  | 33645 | 33,645 units (global); 25,516 (implanted); 1,741 remaining in U.S. Market for implantation. | Worldwide Distribution-USA (nationwide) and countries in Europe and Australia. | Terminated | July 10, 2012 | 298 | 33645 | Not_Computer
Z-0006-2012 | 59139 | No | LYJ |  |  | Neurology | PMA | Model 104 VNS Therapy Demipulse Duo Generator Cyberonics, Inc. Houston, Texas, Model 104 Generators are provided in a sterile single use packages as part of the VNS Therapy System. Product labeling may be found at http://us.cyberonics.com/en/vns-therapy-for-epilepsy/healthcare-professionals/vns-therapy/manuals-page. |  |  | May 06, 2011 | 2011 | Cyberonics, Inc | An investigation was initiated based on a report from the field in which an Intensive Follow-up Indicator (IFI) message was unexpectedly received by a medical professional when using Model 250 version 8.0 software to interrogate a patient's Model 103 Generator. | Software design (manufacturing process) | The firm, Cyberonics Inc., initiated containment actions on March 17, 2011 to prevent additional affected devices from being shipped. The firm initiated contact with implanting facilities in the United States on May 6, 2011 and international facilities and distributors on May 10, 2011 to remove product from the field that has not yet been implanted. Additionally, a "Field Safety Alert" letter was sent to all affected medical professionals. The letter identifies the affected product, the issue, and the action(s) to be taken by the physicians. Consignees are instructed to complete and return an attached Effectiveness Check Form. If assistance is needed physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com |  | 13 | 13 units | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, ID, IL, KS, OH, AR, MA, MI, MN, MO, NC, NH, NJ, NM, NY, OH, OR, PA, TN, TX, UT, WA, WI and WV. Products were also distributed to WASHINGTON DC and Puerto Rico and the countries of: Austria, Belgium, Finland, Germany, Ireland, Italy, Netherlands, Norway, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | December 04, 2012 | 578 | 105 | Software
Z-0008-2012 | 59213 | No | MRW |  |  | Orthopedic | 510(k) | OSTEOMED SPINE Facet Fixation System PrimaLOK FF Threaded Implant Driver Shaft OsteoMed L.P. 3885 Arapaho Road, Addison, Texas 75001 |  |  | June 07, 2011 | 2011 | Osteomed, Lp | Two of the Implant Driver Assembly tips were reported to break during surgery. | Device Design | OsteoMed LP sent an Urgent --Product Recall letter dated June 7, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately for affected inventory and promptly return any affected inventory to:OsteoMed LPAttan: Rebecca Ellis3885 Arapaho RoadAddison, Texas 75001If the customer further distributed any of the affected product customers were instructed to immediately contact their accounts and advise them of the recall situation and have them return any devices they may have to OsteoMed at the above address.Customers were asked to complete and return the enclosed response immediately in the enclosed, postage paid envelope even if they do not have the device. For any questions please call (800) 456-7779. |  | 20 | 20 units | Nationwide Distribution including PA, TX, CA, IL IA, FL, CT, and MO | Terminated | December 04, 2012 | 546 | 20 | Not_Computer
Z-0004-2012 | 58823 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, CADDÂż-Solis ambulatory infusion pumps, Models 2100, REF 21-2101-0200-51 and 2110, REF 21-2111-0100-51, Rx Only. |  |  | May 10, 2011 | 2011 | Smiths Medical ASD, Inc. | Reports of non-delivery/ significant under-delivery of fluids when the CADDÂż Medication Cassette Reservoir is not correctly latched to the pump.EXPANDED: On 7/14/11, Smith Medical expanded this recall to include all consignees with affected pumps. | Nonconforming Material/Component | Smiths Medical sent a customer's bulletin notice dated April 11, 2011, only to the customers who reported this issue. Smiths Medical will work directly with the five Consignees who have reported complaints for this issue.The customer was instructed to read the attached Customer Information Bulletin. If you have additional questions on how to correctly latch a cassette to the pump, than please visit the following website: https://www.smiths-medical-education.com.Smiths Medical will also provide replacement pumps, as requested. This also will be considered an extension of Smiths Medical's investigation into this issue and to further validate the effectivity of the corrective and preventative actions in actual use.For further questions please call (651) 628-7280.EXPANDED:Consignees were sent a Smith Medical " Urgent Safety Alert" letter dated June, 30, 2011. The letter was addressed to Risk/ Safety Managers, Clinicians/ Biomedical Engineering, Nursing, Pharmacy, Pain Management and Anesthesia Professionals, Medical Equipment Rental Companies,Distributors and other users of these devices. The letter described the problem and the product. It provided "Advice on Action to be Taken by the User". This included to read the attached Customer information Bulletin and to complete and return the attached Confirmation Form. |  | 4536 | 4536 (3792 USA, 744 OUS) | Worldwide Distribution -- (USA) nationwide including the states of AR, CA, MO, NY and WI., and the countries of BELGIUM, CANADA, SWITZERLAND, GERMANY, DENMARK, FRANCE, UNITED KINGDOM, IRELAND, ITALY, LUXEMBOURG, MALAYSIA, NETHERLANDS, NEW ZEALAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE. *** EXPANDED*** AL, AK, AZ, CO, DC, FL, GA, ID, IL, IN, KY, LA, MD,MA, JI, MN, MT, NE, NH, NJ, NM, NC, OH, OK, OR, PA, RI, TN, TX, UT, VT, VA, WA, WY. UNITED ARAB EMIRATES, AUSTRIA, AUSTRALIA. | Terminated | July 19, 2012 | 436 | 4536 | Not_Computer
Z-0011-2012 | 59956 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list 11005 |  |  | September 16, 2011 | 2011 | Hospira Inc. | Hospira has received reports of continuous recycling and/or rebooting of Plum A+ infusion pumps when the backlight intensity and/or display contrast settings for the LED display has been adjusted from the original default setting. This condition could cause a delay in therapy if the clinician is unable to start the device due to the continuous recycling and/or rebooting. | Software design | The firm, Hospira sent "URGENT DEVICE RECALL" letters dated September 16, 2011 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to Hospira via fax at 1-888-345-5358. The letter states that a software upgrade to address the issue is currently being developed. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for technical assistance. *** Follow-up letters dated June 26, 2012 were sent to the accounts listing the device modules with software versions v13.2, v13.3, v13.40 and v13.5 that will require the software upgrade. *** |  | 120 | 120 pumps | Worldwide Distribution - USA (nationwide) including Puerto Rico and the US Virgin Islands, and countries of: Argentina, Australia, Bahamas, Barbados, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Denmark, Dominican Republic, Ecuador, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United kingdom, United Arab Emirates, Uruguay and Vietnam. | Open, Classified |  |  | 290217 | Software
Z-0013-2012 | 59952 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Natura Durahesive ConvaTec Moldable Technology Post-Operative /Surgical System (57 mm) |  |  | June 22, 2011 | 2011 | Convatec Inc. | Post-operative kits 57 mm may contain 70 mm pouches. | Labeling mix-ups | The firm, ConvaTec, sent a "MEDICAL DEVICE RECALL" letter dated June 22, 2011 via UPS 2nd day air to all customers and sales representatives that received product. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop distributing and quarantine the recalled lot; perform a count of recalled product currently in inventory; complete and return the Recall Response Form via fax to (908) 904-2120 or mail to 200 Headquarters Park Drive, Skillman, New Jersey 08558; return the Recall Response Form even if no recalled product is in inventory; call the toll free number 1-800-582-6514 for return instructions and RGA number-(recall product must be returned in a separate container from any other product returns); mark all return shipments "Recalled Product" and provide supplied RGA number, and forward this letter to customers, if they distributed the product.Should you have any questions, call 1-800-582-6514. |  | 283 | 283 units | Nationwide distribution: AL, AZ, CA, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NV, NY, OH, OK, OR, PA, RI, TN, TX and UT. | Terminated | January 11, 2012 | 203 | 283 | Not_Computer
Z-0012-2012 | 59943 | No | ECI | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Biomet Microfixation outer packaging labeled in part:"REF: SP-2411***RUBBER BAND, LATEX FREE SMALL, HEAVY, 300/PACKAGE 4MM, (3/16" INNER DIAMETER) NON LATEX ELASTICS LOT 307850***DISTRIBUTED BY: BIOMET MICROFIXATION 1520 Tradeport Drive, Jacksonville, FL 32218 USA www.biometmicrofixation.com***." |  |  | August 18, 2011 | 2011 | Biomet Microfixation, Inc. | On 09/18/2011, Biomet Microfixation, Jacksonville, FL, recalled their SP-2411, Rubber Band, Lot 307850 because it was labeled and sold as "LATEX FREE" although it contains "LATEX". The firm's HHE identified this as a potential high health risk to the patients or end users with latex allergies. | Error in labeling | Biomet Microfixation, Jacksonville, FL sent a Urgent Medical Device letter dated August 19, 2011, to all affected customers. They were also notified by phone, and followed-up by e-mails starting on 08/18/2011. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed If you have any of the affected products in your facility, please remove it and return it immediately to Biomet Microfixation for credit. In order to return these products, please utilize the Return Goods Authorization (RGA) number provided on the reconciliation form.Please follow the instructions on the form and fill out the attached Inventory Reconciliation sheet and FAX to (904) 741-9425. If you have any further questions, please call 1-800-874-7711 or (904) -741-4400, extension 9468, Monday through Friday, 8am-4pm ET. |  | 22 | 22 packs | Worldwide Distribution-- USA (nationwide) including the states of FL, KY, MN, NJ, SD, and VA., and the country of Australia. | Terminated | February 06, 2012 | 172 | 22 | Not_Computer
Z-0016-2012 | 59854 | No | OFD | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | DeRoyal (R) Heat/Moist Exchange, REF 83-000042, Rx Only, NON-STERILE, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, USA: |  |  | August 23, 2011 | 2011 | Deroyal Industries, Inc. Lafollette | The firm manufactured kits containing a HCH device, manufactured by Vital Signs, that was subsequently recalled because the product may have an occlusion that could prevent proper flow of medical gases or oxygen to the patient. | Nonconforming Material/Component | The firm, DeRoyal, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated September 13, 2011 to its customers. The letter describes product, problem, and actions to be taken. Customers are instructed to identify inventory being recalled and place in quarantine to prevent further use. Constomers were asked to return affected product they have in inventory. DeRoyal will replace this product free of charge. Customers were asked to complete the attached Recall Response Spreadsheet and return as per the instructions. Constomers were asked to return affected product they have in inventory. DeRoyal will replace this product free of charge. Distributors were instructed to perform a sub-recall to their customers.Customers with questions should call 865-362-1037 between 9am and 5pm EST or email recalls@deroyal.com. |  | 200 | 200 units | Nationawide Distribution USA - including the states of: CA, FL, GA, NJ, and WA. | Terminated | January 30, 2013 | 526 | 1140 | Not_Computer
Z-0020-2012 | 59868 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II TURBO Hemoglobin A1c Program Reorder Pack, 1600 tests, |  |  | September 12, 2011 | 2011 | Bio-Rad Laboratories, Inc. | A previous recall related to reports of calibration failure and late retention times using Elution Buffer led to an examination of other buffers manufactured during the same time period and a decision was made to field correct these products until a definitive root cause and corrective action have been found and implemented. | Process control | Bio-Rad sent an Urgent Medical Device Correction letter on September 12, 2011, to all affected subsidiaries and customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discard all affected product in accordance with all local, state and federal disposal procedures. If customers distributed the product further they should direct their customers to discard all unused product.Customers were asked to please translate (if necessary) the attached "Customer Medical Device Correction Response Form," customers should add their local contact information and distribute to all affected customers in their region by mail or fax.Compile tracking documentation (e.g. Excel spreadsheet) for all customers who received the Medical Device Correction Letter.Complete the "SUBSIDIARY MEDICAL DEVICE CORRECTION RESPONSE FORM" and fax to Bio-Rad Regulatory Affairs at (510) 741-3954, or e-mail PDF to CSD_RA/Hercules/US/BIO-RAD.Customers would be issued credits for the affected products based on their request. For questions call (510) 741-3954. |  | 82 | 82 | Nationwide and to New Zealand, Canada, France, England, China, India, Mexico, Italy, Brazil, Australia, Singapore, Thailand, SriLanka, hong Kong, Sweden, Germany, Peru, Jamaica, Uruguay, Columbia, Argentina, Chile, Paraguay and Korea. | Terminated | May 11, 2012 | 242 | 287668114 | Not_Computer
Z-0030-2012 | 59957 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Minutex D-Dimer ; Trinity Biotech plc, |  |  | June 13, 2011 | 2011 | Tcoag Us, Incorporated | There is a performance issue with Minutex D-Dimer Kit which may cause false negative results for patient samples containing levels of greater than 250ng/mL D-Dimer. | Nonconforming Material/Component | The firm,Tcoag US, Incorporated, sent an "Urgent Device Recall" letter dated June 13, 2011 via UPS to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the products subject to recall; complete and return the attached Product Recall Notification Response Form by June 22, 2011 via fax at 973-775-3085; scrap the products consistent with their laboratory policy and if they have further distributed these products, notify their customers to quarantine and scrap the recalled product.For questions or technical assistance, please call the Tcoag hotline at 1-888-291-0415, option 2 (7.30 am to 6.00 pm Eastern Time). |  | 4409 | 4409 Units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | December 11, 2012 | 547 | 4409 | Not_Computer
Z-0032-2012 | 59853 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | MDC PACS - release R2.3 SP1 . |  |  | September 01, 2011 | 2011 | Philips Healthcare Inc. | Clinical measurements on true size printouts may be inaccurate | Software design | Philips Medical notified all affected customers with a "Field Safety Notice" dated August 2011.The letter provides instructions to the customer for how to avoid the problem. Customers are informed that until the software upgrade is recieved from Philip, do not use the true-size printing feature if the system is configured for more than one DICOM printer. Customers will be instructed that they will receive a software upgrade when it is available. For questions on this recall contact Philips representative at (800) 722-9377. |  | 20 | 20 units | Worldwide Distribution: Nationwide distribution including Puerto Rico; and the countries of Argentina, Brazil, China, France, Greece, India, Israel Mexico, and Turkey. | Terminated | March 13, 2017 | 2020 | 20 | Software | Output/Calculation | Software update | Software Update
Z-0029-2012 | 59787 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OLYMPUS ViziShot EBUS Aspiration Needle |  |  | August 25, 2011 | 2011 | Olympus America Inc. | Firm has received complaints from users of the ability to wipe material from the needle's stylet and the presence of material in patient cytology samples obtained by the needle following endobronchial ultrasound fine needle aspiration of a lesion. Olympus Medical Systems Corporation (OMSC) in Japan has confirmed that metal particles can occur on the stylet due to friction between the stylet and n | Device Design | OLYMPUS sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A revised Instruction Manual with the new WARNINGS will be sent to the customer's attention.Customers were instructed to examine their inventory to determine if they have the affected product.To share the information on the new WARNINGS with all of the appropriate staff at their facility and any other organization where the product has been transferred.Complete and return the enclosed questionnaire even if they do not have affected product. Customers were instructed to fax the completed questionnaire to 484-896-7128.For any question call 484-896-5688. |  | 22498 | 22,498 | Nationwide Distribution including the District of Columbia and the following States: AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | September 28, 2012 | 400 | 22498 | Not_Computer
Z-0031-2012 | 59746 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) Cuvettes |  |  | August 08, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed a low frequency defect in the molding of the Cuvettes that may result in Cuvette Ring and Cuvette Loader jams. It has been determined that the issue is related to a molding defect with the Cuvette flange or "wing" breaking, resulting in an error condition. Therefore, patient results are not impacted. | Nonconforming Material/Component | Siemens Healthcare Diagnostics sent a Urgent Field Safety Notice letter dated August 2011, to all affected customers. The letter identifed the product the problem and the action needed to be taken by customer.Customers were instructed to discard all cuvette cartridges of the identified lots. Customers were also informed that until a replacement lot is available, if an error is encountered, they may contact the Technical Solutions Center for assistance at (800) 441-9245. For replacement product , please call the Siemens Atlanta Customer Service Center at (800) 241-0420. |  | 1786 | 1,786 | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and Puerto Rico., and the countries of Austria, Belgium, France, Germany, Italy, The Netherlands, Portugal, Spain, Switzerland, Australia, Japan, South Korea, and Canada. | Terminated | December 24, 2013 | 869 | 1786 | Not_Computer
Z-0028-2012 | 59535 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2, Part Number: 378600, Serial Numbers: 1001008698, 1001008699, and 10018700. |  |  | August 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed an issue with the Coulter PrePlus 2. The reagent level sensing is impaired and will give no warning when a reagent vial is low or out of reagent. Due to the size difference in the vertical mounts, the reagent rack is not aligned optimally for the probe to aspirate or level sense the antibody. This misalignment yields impaired reagen | Mixed-up of materials/components | Beckman Coulter initiated an URGENT: PRODUCT CORRECTIVE ACTION (PCA) letter dated August 24, 2010, with attached PCA Response Form to all affected customers who purchased the Coulter PrepPlus 2. The letter provided the customers with an explanation of the problem identified and an action to be taken. The customers were instructed to not use the PrepPlus2 until their Beckman Coulter representative corrects the problem. A review of all results generated from samples prepared from the system should be conducted. Any unexpected dim fluorescence and/or decreased number stained cells should be evaluated for re-analysis, including preparing a new sample. Customers were also instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding this PCA, were instructed to call Beckman Coulter Customer service at (800) 526-7694. Outside of USA, customers were instructed to contact their local Beckman Coulter Representative. |  | 391 | 391 units total (193 in US) | Worldwide distribution. | Terminated | April 04, 2012 | 601 | 391 | Not_Computer
Z-0036-2012 | 59937 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System, Synchrony Respiratory Tracking System (Robocouch Patient Positioning System), a Radiation Therapy Device. Model Numbers: 026430, 028186. |  |  | September 12, 2011 | 2011 | Accuray Incorporated | A report and investigation determined that the set screws, which were part of the retrofit of a previous recall were incorrectly engaged at one facility which allowed rotation and subsequent descent of the assembly. | Process control | Urgent Device Correction Letters were sent to all sites affected by the retrofit on September 12, 2011 via Fed Ex. The letter identified the affected product and the description of the potential problem. Customers were asked to inspect their Synchrony System by following the instructions provided. If the lower extension feels loose, customers should contact Accuray Customer Support immediately to schedule service of the system. The letter also states that all affected systems will be inspected and corrected, if necessary by Accuray to prevent the potential problem. Accuray Customer Support will contact customers to schedule the inspection. Questions or concerns regarding the issue should be directed to Accuray Customer Support at 1-877-8667 for US customers and 1-408-716-4700 for customers outside of the US, or customersupport@accuray.com. |  | 236 | 236 total devices | Worldwide Distribution | Terminated | February 27, 2012 | 168 | 236 | Not_Computer
Z-0034-2012 | 59756 | No | FZS | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Axle Interspinous Fusion System Torque Driver Handle, Model #: X060-0323, Lot # 525941, S/Ns 001-017, Lot # 525942, S/Ns 001-050, Mfr. By: Bradshaw Medical Inc., Kenosha, WI 53144 |  |  | July 15, 2011 | 2011 | X Spine Systems Inc | The firm learned that the Torque Driver Handles were not actuating at a proper torque. It was determined that the handles were speed determinant and could output a higher torque if the handles were turned at a very slow speed. | Component design/selection | X-spine Systems, Inc. sent an URGENT DEVICE RECALL letter dated September 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their stock immediately to determine if they have any affected product. If so, customers were instructed once the set screw is tightened onto the locking plate, and resistance is detected in the handle, it will take less than 2 seconds to hear the handle "click." The "click" indicates the torque limit was reached. If further resistance is felt after 2 seconds, the surgeon should release the handle from tension, and try again.If the product was distributed further, they should contact their accounts, and advise them of this recall.The firm will send replacement handles.For any questions call (937) 847-8400, ext 115. |  | 39 | 39 Torque Driver Handles | Nationwide Distribution including CA, CO, FL, IA, IN, LA, NY, PA, TN, TX and UT. | Terminated | March 01, 2013 | 595 | 39 | Not_Computer
Z-0033-2012 | 59879 | No | FJP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sodium Bicarbonate Mixing and Delivery System Software v 7.1. Model # MD-380, Serial Numbers: 5103031102, 5101101101, 5102161007, 5103221007, 5104221007, 5105101008, 5109131010, 5111231011, 5114301012, 5102131101 &amp; 5113091012. Manufactured by: Isopure Corp, Simpsonville, KY |  |  | May 19, 2011 | 2011 | Isopure Corp | Software problem causes transfer line between mixing and storage tank to not be rinsed after disinfection, leaving residual chlorine in the line. | Software design | Isopure sent a letter dated May 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Isopure devised an immediate solution which must be implemented at their faclity in conjunction with the disinfection procedure. Customers were asked to inform all appropriate personnel, clinical and technical, who may be involved in operating the system during the normal disinfection schedule.On 5/19/2011, the firm sent letters to their customers via US mail and email. They also telephoned their customers to notify them of the problem. |  | 11 | 11 units | Nationwide Distribution including AR, CT, GA, MA, ME, NJ, PA, TN and WV. | Terminated | February 19, 2013 | 642 | 11 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-0043-2012 | 58746 | stains, hematology | KQC | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Harleco Giemsa Stain 25 gram glass bottle, Item #228-12 Manufactured for EMD Chemicals Inc., Gibbstown, NJ 08027 1-800-222-0342 |  |  | April 04, 2011 | 2011 | EMD Chemicals Inc. | Stains may yield lower then expected absorbance levels resulting in under stained hematology specimens. | Employee error | EMD Chemicals Inc. sent an "IMPORTANT PRODUCT RECALL" letter dated April 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. An attached response form is included for customers to complete and return via fax to EMD Chemicals at 856-423-6314. Any questions concerning this recall notice should be directed to EMD Chemicals Technical Service Department at 1-800-222-0342. |  | 36 | 36 units | USA Distribution in the states of PA and NY. | Terminated | October 13, 2011 | 192 | 36 | Not_Computer
Z-0037-2012 | 59898 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Paceart Integration Software, version 1.00 Model 6472. |  |  | August 12, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific has determined that under certain conditions, LATITUDE Paceart Integration (LPI) softwaremay not properly transfer data resulting in the incorrect summary of programmed parameter values in the ICD Summary tab display and Current Programming Tachy summary section of the Paceart report. | Software design | Consignees were sent on 8/12/11 a Boston Scientific letter dated August 2011. The letter was addressed to "Dear Doctor". The letter described the product and problem. It also gave Background Information, Description of Potential Data Translation Error, and Additional Information and Action information.A letter, including an HRS form, was also sent to all US LATITUDE Clinics that have active LPI software to notify them of the need to install the corrected LPI Software. |  | 398 | 398 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT DE, DC, FL, GA, IL,IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OJ, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA,WV, WI, and WY. | Terminated | November 06, 2012 | 452 | 398 | Software | Device Operation | Software update | Software Update
Z-0042-2012 | 59891 | No | MSQ | Campylobacter fetus serological reagents. | Microbiology | Microbiology | 510(k) | IDkit:HP" Breath Test for Detection of H. pylori, Manufactured by: Exalenz Bioscience Ltd 4 Ha' maayan St. Modiin, Israel 71700, Distributed by: Exalenz Bioscience Inc., 1313 N. Market St., Sute 5100, Wilmington, DE 19808 USA |  |  | June 07, 2011 | 2011 | Exalenz Bioscience Ltd | The product is in violation of the FD&C; Act in that it was illegally marketed without proper clearance (no 510(k) submitted). | No Marketing Application | The firm, Exalenz Bioscience Ltd. notified consignees located in the United States of the recall by telephone, electronic mail, or written correspondence between June 9, 2011 and June 30, 2011. The letter informed consignees that Exalenz will immediately cease marketing the IDkit:HP TWO test in the United States. Consignees were instructed to return shipment of the IDkit:HP TWO test kits that they have in their possession to the companys US warehouse. For question call the VP of North American Sales at 704-770-6162 or VP of Marketing at 617-794-6936. |  | 1240 | 1240 units | Nationwide Distribution (USA) - including the states of: FL, IL, LA, MA, MS, NY, TN, TX and VA. | Terminated | July 06, 2012 | 395 | 1240 | Not_Computer
Z-0039-2012 | 59316 | microtome, rotary | IDO | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | Thermo Scientific Microm: Rotary Microtome: REF HM355S-3 |  |  | August 12, 2011 | 2011 | Richard-Allan Scientific Company | The firm has received two complaints of finger lacerations in operators of this device. | Device Design | ThermoFisher Scientific sent "URGENT MEDICAL DEVICE CORRECTION, RE: Thermo Scientific HM355S-3 User Safety Advisory Notice" to all affected customers.The letter states the manufacturer identified a need to enhance the safety instructions and the description of the safety devices in the user manual. It mentions the HM3555S-3 User Manual has been modified to include more specific safe work practices and a new Quick -Reference Card summarizing the manual's safety enhancements and that the changes have been included on a CD that is enclosed with the letter. The firm requests confirmation of receipt of the new User Manual and Quick-Reference Card by October 31, 2011 to be sent by FAX to (269) 372-2674For questions on this recall call the firms Technical Support representatives at 1-800-522-7270 08:00 AM- 4:30 PM. |  | 786 | 786 | Worldwide Distribution: Nationwide distribution; including the country of Canada. | Terminated | July 25, 2013 | 713 | 786 | Not_Computer
Z-0038-2012 | 59888 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) EXL LOCI Cardiac Troponin I Calibrator (RC621) |  |  | August 31, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Internal testing has confirmed that recovery of levels 2 and 3 of LOCl Cardiac Troponin I Calibrator is lower than expected bottle value. Following calibration with LOCl Cardiac Troponin I Calibrator, lot 1AD045 an upward shift in patient and QC results may be observed. | Other | Siemens sent an Urgent Field Safety Notice dated August 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and discard any remaining inventory. Replacement orders would be processed and sent free of charge. Customers were also instructed to complete the attached FIELD CORRECTION EFFECTIVENESS CHECK form and fax it to (302) 631- 8457.Customers were instructed to forward the Safety Notice to anyone to whom they may have distributed the affected product. For any technical questions customers were to contact the Siemens Solutions Center at 800-441-9250. |  | 998 | 998 | Worldwide Distribution - USA (nationwide) including AL, AK, AZ, AR, CA, CO, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of GUAM, Austria, Canada, China, Czech Republic, Egypt, France, Germany, Italy, Poland, Portugal, Puerto Rico, Spain and Switzerland. | Terminated | May 20, 2013 | 628 | 998 | Not_Computer
Z-0040-2012 | 59712 | No | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Prostrate Seeding Applicator Needle. Product code PSS1820AT |  |  | August 12, 2011 | 2011 | Worldwide Medical Technologies, LLC | Length of the cannula hub is longer and may not permit a fit with the Applicator | Nonconforming Material/Component | BrachySciences, a division of Biocompatibles, Inc., sent out "Urgent Recall Device Recall" letters Dated August 12, 2011 to all affected customers.The letter included affected product and problem description. Customers are asked to quarantine and return unused product. A Product Recall &amp; Retrieval Form was included to send back to the firm.For questions on this recall contact the firm at (203) 262-0571, extension 105. |  | 1872 | 1872 units | Worldwide Distribution: Nationwide distribution including the states of AR, CO, DE. FL, KY, MD, MI, MN, NY, TX, and VA; and the countries of Belgium, and Germany. | Terminated | September 19, 2012 | 404 | 1872 | Not_Computer
Z-0041-2012 | 59418 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | ECG Cable Part Number: 01.57.471017 for use with the SE-1 series, SE-3 series and SE-601 series Smart Electrocardiograph |  |  | September 10, 2010 | 2010 | Edan Instruments, Inc. | The 3 leads of the cable, RL, LA and LL were labeled in wrong sequence. The left arm (LA) is mislabeled as right leg (RL). The left leg (LL) is mislabeled as left arm (LA). The right leg (RL) is mislabeled as left leg (LL). There have been no reports of injury as a result of this issue. | Vendor change control | Edan Instruments, Inc. sent an Urgent - Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that wrongly labeled cable may lead to incorrect ECG waveforms during ECG test, customers were reminded to review previous diagnoses made from this cable to mitigate any inappropriate recordings or treatment made.Customers were instructed to replace the affected ECG Cables once they receive the new ECG Cable, and collect all the affected ECG Cables and then destroy and dispose of them according to local regulations. Questions or concerns were directed to the following numbers: Oceania and Asia: +86 (755) 2689 8322North America, South America:+86 (755) 2689 9914Europe: +86 (755) 2669 4684 |  | 615 | 615 Ccables | Worldwide Distribution - USA (nationwide) and the countries of Dominica, Ecuador, Colombia, Honduras, Peru, Mexico, El Salvador, Guatemala, Venezuela, Uruguay, Chile, Panama, Nicaragua, Ecuador, Australia, Philippines, Malaysia, Indonesia, Chinese Taipei, &amp; Spain | Terminated | October 13, 2011 | 398 | 615 | Not_Computer
Z-0045-2012 | 58461 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PERPOS PLS System, Single Use; CATALOG #'s 9045-01 &amp; 9045-02. |  |  | February 03, 2011 | 2011 | Interventional Spine Inc | This recall has been initiated due to drivers that may contain a burr in the distal tip preventing full seating of the driver onto the head of the facet screw, resulting in difficulty screwing the facet screw into the bone.Use of this product may result in a delay of the surgical procedure | Nonconforming Material/Component | Interventional Spine, Inc. sent an "Urgent Recall" letter dated March 15, 2011 to all affected customers.The letter included; affected product, reason for recall and recommendations.Customers were asked to immediately examine inventory, quarantine affected product and to notify International Spine, Inc if any product remain.For information on this recall contact the firm at (949) 472-0006. |  | 2620 | 2,620 units | Word wide Distribution: (USA) Nationwide including the states of CO, CT, FL, KY, LA, MA, NC, NY, OK, OR, TN, TX, VA, and WI; and the countries of Italy, Spain, Turkey, and the U.K. | Terminated | February 05, 2013 | 733 | 2692 | Not_Computer
Z-0055-2012 | 59526 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 8F Safe Sheath Introducer, Catalog number: V18; |  |  | August 01, 2011 | 2011 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | Device Design | AngioDynamics, Inc contacted affected customers initially via telephone and a follow up "Notification of Medical Device Correction" letter was sent dated August 1, 2011. The letter informed them of the problem and provided instructions. Customers were instructed to utilize the 8F Introducer Sheath provided within the 8F Smart Port CT Implantable Port during port implantation. The 8F Valved Sheath (V18) may continue to be utilized during the placement of the 7.5F Smart Port Implantable Port (for example: CT75STSD; CT75STSD-NF). For question on this recall contact AngioDynamics, Inc at (800) 772-6446. |  | 1095 | 1095 total units | (USA) Nationwide distribution including the states of CA, FL, GA, IA, IL, IN, KY, MD, MI, MO, NC, NE, NJ, NY, OH, OK, OR, TN, UT, WA and WI. | Terminated | March 07, 2012 | 219 | 1095 | Not_Computer
Z-0048-2012 | 59768 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000rt System software; an automated system for performing fluorescence-based PCR that results in quantitative and qualitative detection of nucleic acid sequences; Abbott Molecular Inc., Des Plaines, IL 60018 USA; catalog #09K15-01 |  |  | September 08, 2011 | 2011 | Abbott Molecular | The Abbott m2000rt System software has software bug which can affect customers who manually enter a Test Order onto the m2000rt and sorts a column on the calibration or control grids of the Test Order screen by double clicking the column header. This causes a mismatch between specific concentration values and the designated calibrators and controls in the final results, which could cause an inval | Software design | Abbott Molecular Inc. sent a Field Correction Recall letter dated September 8, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to your m2000rt software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2012.Please complete the enclosed reply sheet and fax it to (847) 775-6728 at your earliest convenience to acknowledge receipt. We appreciate your assistance in this matter. |  | 926 | 926 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Belgium, Benin, Bolivia, Bosnia and Herzegovia, Brazil, Bulgaria, Burkina Faso, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Dem. Rep. of the Congo, Denmark, Egypt, Estonia, Ethiopia, Finland, France, Gabon, Gaza &amp; Jericho, Germany, Greece, Guatemala, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Libya, Lithuania, Malawi, Malaysia, Mali, Mauritius, Mexico, Moldova, Morocco, Mozambique, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tajikistan, Tanzania, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Venezuela, Vietnam, Zambia and Zimbabwe | Terminated | March 08, 2016 | 1643 | 926 | Software | Output/Calculation | Software update | Software Update
Z-0044-2012 | 60008 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA Toric (enfilcon A) TORIC soft (Hydrophilic) Contact Lenses, 6 SOFT CONTACT LENSES FOR ASTIGMATISM, CooperVision, Scottsville, NY 14546 USA; Juana Diaz, PR 00795 USA; Hamble SO31, 4RF, UK; Device Listing # D033736. |  |  | August 19, 2011 | 2011 | CooperVision Inc. | This recall was initiated because of the unintended presence of a residue (silicone oil) on lenses from certain lots of Avaira Toric contact lenses. The company received complaints of temporary hazy or blurry vision and discomfort, most likely caused by transient or reversible swelling of the surface layer of the cornea, which is known as corneal epithelial edema. Corneal epithelial edema results | Material/Component Contamination | The firm, CooperVision sent an "Urgent: Company Initiated Voluntary Medical Device Recall - Contact Lenses" notification letter to all affected consignees. The notice identified the product, the problem, and the actions to be taken. The letter includes an attached "Recall Response Form" that consignees are required to complete and fax to (585) 889-5688. Consignees were instructed to immediately examine their inventory, stop any further distribution, quarantine and return product subject to recall. In addition, consignees were told if they have further distributed the affected product to notify customers at once of this product recall. The letter states that accounts will either be credited or exchanged for an unaffected product. If you have any further questions, call 1-800-341-2020. |  | 778301 | At least ( 778,301 lenses: Domestic - 659,847 lenses; Foreign - 118,454 lenses; including EXPANSION: Domestic - 69,047 lenses) ( to be updated) | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Europe, Germany, Hong Kong, Italy, New Zealand, and Spain. | Open, Classified |  |  | 778301 | Not_Computer
Z-0049-2012 | 59978 | [6 empty] | ArthroCare brand MiniMagnum Knotless Fixation Implant Device with Inserter Handle Plus SpeedStitch Magnum Wire Suture Cartridge; Catalog Number: OM-6007; |  |  | May 11, 2011 | 2011 | ArthroCare Corporation | ArthroCare initiated the recall of one lot of the MiniMagnum Knotless Implant Fixation Device where the incorrect diver block was packed with the implant. | Packaging process control | The firm, ArthroCare Corporation, sent a letter dated May 10, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were asked to immediately return the product; check their inventory, refer to the copy of the Mini Magnum outer label to help identify the product, and if the affected product lot is in their inventory, notify ArthoCare Product Support at (800) 797-6520, press 2 and return the product immediately. The letter also stated, "In addition, if you do not have any product from this lot remaining in your inventory, we request that you contact ArthroCare at the number above, or by e-mail at ra@arthrocare.com, to inform us of the lack of remaining product. This will ensure that you will not be contacted again and that we have accounted for product shipments from this lot." If you have any questions, please call the Chief Regulatory Officer and Senior Vice President, Quality Systems and Assurance at (408) 735-6229. |  | 32 | 32 units | Nationwide distribution: USA including states of: CA, FL, NY, OH and OK. | Terminated | October 24, 2011 | 166 | 32 | Not_Computer
Z-0050-2012 | 53046 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Total ÂżhCG. |  |  | March 05, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that certain Access Total ÂżhCG reagent packs show decreased stability from the date of manufacture, which may lead to increased imprecision. | Nonconforming Material/Component | The recall communication was initiated on March 18, 2009 with Beckman Coulter forwarding an Urgent: Product Corrective Action (PCA) letter dated March 13, 2009 with attached PCA Response Form (via US Postal Service for US Customers) to all customers who purchased the Access Immunoassay Systems total ÂżhCG. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discontinue use of the reagent lot numbers 824542 and 825652, and discard any unused reagent packs.Customers were asked to review their historical records for low level quality control to ensure that the assay performance for the referenced lot numbers has been within their laboratory specifications. They were also asked to share the information with their laboratory staff and retain the letter for their Quality System documentation, and then complete and return the enclosed response form within 10 days so that they may be assured that consignees received the notification.Finally they were told to contact Customer Technical Support Center at 1-800-854-3633 in the US and Canada, or their local Beckman Coulter Representative with any questions. |  | 14764 | 14,764 | Worldwide Distribution -- US, Algeria, Angola, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, canada, Chile, China, Colombia, Czech Republic, Egypt, El Salvador, Ethiopia, Finland, France, French Polynesia, Germany, Ghana, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kenya, Korea, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Mexico, Monaco, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Reunion, Russian Federation, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Zambia. | Terminated | April 27, 2012 | 1149 | 14764 | Not_Computer
Z-0054-2012 | 59912 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | SARNS INTRACARDIAC SUCKER, BULK, NON-STERILE |  |  | September 13, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) recently discovered three instances of metal fragments inside the stainless steel tube of the Sarns" Adult Rigid Intracardiac Sucker. | Nonconforming Material/Component | On 9/13/2011 Terumo sent a letter entitled "URGENT MEDICAL DEVICE REMOVAL" to its customers that included the name and description of the device, lot numbers, the potential hazard, correction, the reason for the recall and instructions. Customers are to review the notice, ensure that all users are aware of the issue, and a recall response form to be completed and returned. Customers are to call Terumo CVS Customer Service to obtain a Return Goods Authorization number. The letter also included MedWatch reporting information and stated that Customers with questions should call 1-800-521-2818. |  | 5700 | 5700 | Worldwide Distribution -- USA, Belgium, Australia, Mexico, and Japan. | Terminated | December 10, 2012 | 454 | 15470 | Not_Computer
Z-0052-2012 | 59961 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Impac Medical Systems Sequencer |  |  | September 09, 2011 | 2011 | Impac Medical Systems Inc | Varian reported a problem related to certain collimator rotation values which are incorrect on TrueBeam compared to planned values in Mosaiq. | Nonconforming Material/Component | "Impac Medical Systems Inc., sent an USER NOTICE letter to all affected consignees on September 20, 2011. The letters identified the source characterization error, clinical impact, and final resolution. The letter states that Elekta Support will follow up with each consignee and remotely correct the source characterization. Customers were instructed to sign, date and return the attached Confirmation of Receipt Form by one of the three ways indicated. The letter instructs consignees to distribute this notice to any and all users of IMPAC software that may potentially be affected by the issue. For questions, concerns, and requests for upgrade, email support@impac.com or call. The letter provided consignees with a complete list of contact information.. |  | 11 | 11 devices | Nationwide Distribution (USA) - including the states of: AZ, CA, FL, NJ, NY, NV, TX, VA and WA. | Terminated | March 06, 2012 | 179 | 11 | Other | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0051-2012 | 59848 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES, 621-400ES, 621-400ESR. |  |  | September 07, 2011 | 2011 | B Braun Medical, Inc | B Braun Outlook ES Safety Infusion System "Run" Light Emitting Diodes (LEDs) pumps stop infusing and a backup alarm sounds, but the "Run" LEDs advance as if the pumps were infusing. | Software design | B. Braun Medical, Inc. sent an "URGENT: FIELD CORRECTION" letter dated September 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. A firm representative will contact all affected customers in order to schedule the field correction. The correction consists of revising the Door and Main Processor software, inspecting the main board, and replacing the board, if warranted. A Field Correction Information Form was included in the letter for customers to complete and return. Contact Technical Product Support at 1-800-627-7867 for questions regarding this notice. |  | 6321 | 6321 | Nationwide Distribution | Terminated | May 04, 2015 | 1335 | 6321 | Software | Alarm/Message | Revise software/insepct and replace board if needed | Software Update
Z-0057-2012 | 59442 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MicroMaxx ultrasound system and M-Turbo ultrasound system is Medical Diagnostic Ultrasound. |  |  | July 18, 2011 | 2011 | Sonosite, Inc. | This is a software anomaly. In some circumstances, if the Systolic/Diastolic (S/D) ratio value is interpreted directly from the display, the clinician could potentially use an incorrect value of the S/D ratio in his/her assessment. | Software design | The firm SonoSite, Inc., sent an "Urgent - Medical Device Correction" letter dated September 29, 2011 to its foreign and domestic customers. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the three (3) steps provided in the letter to verify if the extended OB/GYN Calculations (EOBC) is enabled on their system. Additionally, SonoSite will perform a software upgrade to M-Turbo and/or MicroMax ultrasound system in order resolve the issue. The customers were instructed to contact Sonosite Technical Support via website: www.sonosite.com/sdratio; email: regulatory@sonosite.com (include contact information and address), or call in U.S. and Canada: 1-877-590-4982. If you have any questions or concerns regarding this matter, do not hesitate to contact in US and Canada at 1-877-657-8118, Monday-Friday, 6 am-5 pm PST. Customer can also send an email to Technical Support at service@sonosite.com |  | 4236 | 4236 units (US) and 4089 units (internationally) | Worldwide distribution: USA (nationwide) including Puerto Rico; and countries of: Afghanistan, Algeria, American Samoa, Argentina, Australia, Bolivia, Botswana, Brazil, Canada, China, Costa Rica, Cyprus, Dominican Republic, Estonia, France, Gambia, Germany, Great Britain, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Tanzania, Uganda, United Arab Emirates, Viet Nam, and Zambia. | Terminated | September 14, 2012 | 424 | 4236 | Software | Output/Calculation | Software update | Software Update
Z-0060-2012 | 59998 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ANCHORLOK(R) LEADING EDGE(R) SOFT TISSUE ANCHOR SYSTEM, 2.5mm, 2.0 SUTURE, REF 5100-006, 1 each, STERILE R, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002 |  |  | September 20, 2011 | 2011 | Wright Medical Technology Inc | The product is labeled with an expiration date without adequate justification. | Incorrect or no expiration date | The firm, Wright Medical Technology, Inc., sent an "URGENT: PRODUCT RECALL NOTIFICATION" letter dated September 20, 2011 to all affected consignees. The letters identified the products, the problem, and the actions to be taken. The letter instructed consignees to identify affected inventory. If recalled products are located consignees are instructed to contact customer service at 1-800-238-7117, for specific instructions on how to return the recalled products. The letter asked that this communication be forward to any associated sites if the recalled products were further distributed. Consignees are to complete and return the enclosed notification form by fax to 901-867-7401 or by email to: cathy.park@wmt.com. For questions regarding this recall contact 901-867-4324. |  | 596 | 596 units | Worldwide Distribution - USA (nationwide) and countries of: Argentina, Canada, Denmark, France, Great Britain, Ireland, Japan, and Spain. | Terminated | July 17, 2012 | 301 | 11960 | Not_Computer
Z-0071-2012 | 58302 | No | LPA |  |  | Cardiovascular | PMA | ***REF #200***TAPSCOPE 550. Each pouch contains one (1) 18 French esophageal stethoscope with bipolar pacing and recording electrodes with a YSI 400 series male connector thermistor. LOT xxxxxxx 400095-0200-C 100902. www.cardiocommand.com. |  |  | October 25, 2010 | 2010 | Cardiocommand Inc. | CardioCommand, Inc. is recalling their TAPSCOPE 550 esophageal pacing/recording catheters because there maybe electrical discontinuity, which can result in the inability to reliably pace and/or record ECG. | Nonconforming Material/Component | On October 25, 2010, an Urgent Medical Device Recall Notice was sent by certified mail to all facilities advising them of the recall. The letter identified the affected product and reason for the recall. Users were asked to discontinue use and return the affected product to the firm. If all product was already used, then this should be indicated on the form provided. The firm will send replacement product in quantities equal to the returned product. The attached form should be completed and returned via fax. Customers were instructed to report any suspected incidents related to this issue to CardioCommand Customer Service at 1-800-231-6370. Also questions or comments should be directed to CardioCommand Customer Service. |  | 567 | 567 total | Nationwide Distribution -- GA, OR, TX, CO, KY, MT, MA, AL, CA, FL &amp; TN. | Terminated | August 28, 2012 | 673 | 567 | Hardware | Device Operation | Replace | Remove or Replace
Z-0068-2012 | 60000 | No | KGT | Drape adhesive. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STRIP T's STERILE ADHESIVE ORGANIZER, Model #SC5000-4R, Lot #090710, 510(k) exempt. One (1) adhesive organizer is packed in a Tyvek pouch and 60 pouches are placed into a 'Shelf Box'. The shelf box is placed into a shipping container. |  |  | March 05, 2008 | 2008 | Kapp Surgical Instrument Inc | A portion of the labeled expiration date is missing. The expiration date on approximately half of the pouch labels distributed are missing the last digit of the year of expiration. However, the shelf box containing the pouches is labeled correctly. | Incorrect or no expiration date | Surgical Concepts, Inc ( A subsidiary of Kapp Surgical Instrument Inc) sent a notification letter dated March 5, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm ceased distribution of the suspect product on December 5, 2007. The letter informs the firm's customers (both distributors and medical user/customers) of the expiration date labeling omission of the last digit of the year of expiration. It goes on to explain that the incomplete expiration date is present on the individual packages only and that the shelf boxes are labeled correctly with the complete/correct expiration date. The letter provides the customers with a contact person and a toll free telephone number 1-(800) 282-5277, to call in order to make arrangements for receive replacement labels for any misbranded packages received. |  | 59 | 59 boxes (3540 units) | Nationwide Distribution-- including the states of AK, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MO, MT, NC, NM, OH, PA, TN, TX, UT, and VA | Terminated | February 19, 2013 | 1812 | 59 | Not_Computer
Z-0070-2012 | 60007 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre 9F Pruitt F3 Outlying Carotid Shunt with T-port |  |  | October 04, 2011 | 2011 | Lemaitre Vascular, Inc. | Extended Expiration on the outer box label 2016-05, the correct expiration date is on the tray label: 2014-05 | Incorrect or no expiration date | The firm, LeMaitre Vascular, Inc., sent an "URGENT: 9F PRUITT F3 Carotid Shunt Device Field Safety Notice" dated October 4, 2011 to its customers. The notice describes the product, problem and action to be taken. The customers were instructed to identify their inventory; return the affected products for exchange with a properly labeled device; contact Le Matire customer service for replacement at 800-628-9470, and complete and return the RECALL NOTICE RECONCILIATION FORM via fax to 781-425-5049, even if no product is on hand. If you have any questions concerning this safety notice, please contact the Quality Assurance Engineer at 781-221-2266 ext. 183. |  | 144 | 144 units | Nationwide distribution: USA including states of: CA, FL, IA, IL, MA, MI, NE, NJ, NY, OH, SC, TX, and WI. | Terminated | March 14, 2012 | 162 | 144 | Not_Computer
Z-0069-2012 | 59889 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Inc. Dimension(R) EXL Cardiac Troponin I Flex(R) Reagent Cartridges |  |  | August 31, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed customer complaints for low recovery of QC and patient samples. | Process control | Siemens Healthcare Diagnostics, Inc. Urgent Field Safety Notice was sent to customers on August 31, 2011 regarding the issue. The letter instructs customers to immediately discontinue use of the affected lots and discard all remaining inventory. |  | 6072 | 6,072 cartons | Worldwide distribution: USA, Puerto Rico, Guam, Austria, Canada, Croatia, France, Greece, Germany, Italy, Japan, Mexico, Portugal, Spain and Switzerland. | Terminated | January 14, 2014 | 867 | 6072 | Not_Computer
Z-0076-2012 | 59911 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MICROPLASTY TAPERLOC 7.5 MM RESECTION GUIDE, STAINLESS STEEL, BIOMET ORTHOPEDICS 56 EAST BELL DRIVE PO BOX 587 WARSAW IN 46582 USA. |  |  | September 16, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the blueprint calls forengraved measurements on both sides on the Microplasty Taperloc 7.5MM Resection Guide. Theabove lots, as supplied to Biomet by a vendor, are engraved on only one side. To date, Biomet hasreceived 2 reports involving 2 units where the instrument was missing the engraving on one side. | Nonconforming Material/Component | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated September 16, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to locate and discontinue use of the affected product. Arrangements will be made for customers to return the affected product to Biomet. To confirm receipt of this notice call Biomet at 800-348-9500 ext. 3755 or 3756. Questions related to this notice should be directed to 574-371-3755 or 3756, Monday through Friday, 8 AM to 5 PM. |  | 22 | 22 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, IN, MI, OK, and VA and countries of Australia and Chile. | Terminated | March 07, 2013 | 538 | 22 | Not_Computer
Z-0075-2012 | 59968 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Total Hip Joint Replacement |  |  | July 14, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware that a size 3 Accolade stem was in a box labeled as a size 5 Accolade stem. | Mixed-up of materials/components | Stryker Orthopaedics sent an "URGENT: PRODUCT RECALL" letter dated July 14, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an attached Product Recall Acknowledgment Form via fax to 201-831-6069. Additionally, information is provided for customers to return the affected products. Please contact the Divisional Regulatory Compliance Manager at 201-972-2100 for questions regarding this notice. |  | 21 | 21 units | Nationwide Distribution | Terminated | November 15, 2013 | 855 | 21 | Not_Computer
Z-0079-2012 | 59948 | No | DXS | Cardiopulmonary bypass coronary pressure gauge. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack, which are designed as specified by the user. The following models contain a Pressure Display Set. Sterilized by Ethylene Oxide. Non-pyrogenic. Manufacturer: Medtronic, Inc. 710 Medtronic Parkway, Minneapolis, MN 55432-5604. |  |  | September 09, 2011 | 2011 | Medtronic Perfusion Systems | Pressure Display Set (models 61000, 61006, 62000, 62006) DFU will be updated to clarify that the intended use of the disposables is only for cardiopulmonary bypass surgery (procedures up to 6 hours). Use beyond 6 hours may result in air leakage in the system resulting in a pressure reading that is not representative of the actual line pressure. | Device Design | Medtronic Perfusion Systems, sent an "Urgent Medical Device Notification letter dated September 9, 2011, via UPS 2-day delivery to all affected customers. The letter identified the affected devices, described the issues, and the actions needed to be taken. The letter indicates when approved, the updated Directions for Use for the Pressure Display Box and Pressure Display Sets will be sent to customers. Customers were asked to complete the Customer Confirmation Certificate and return to Medtronic per the instructions on the form. The letter states that if the recalled products have been further distributed, forward this letter immediately to all customers who received the affected products, and gather proof of notification. For further questions concerning this notification, contact Medtronic Cardiovascular Technical Support at 877-526-7890 or Medtronic Sales Representative. |  | 74804 | 74,804 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and Virgin Islands (U.S.) and the countries of: Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, El Salvador, France, Georgia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Romania, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom, Venezuela, and Yemen. | Terminated | October 11, 2012 | 398 | 270088 | Software
Z-0083-2012 | 52105 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager -Software Versions 8.05, 8.06, 8.07, 8.08, and 8.09. |  |  | May 11, 2009 | 2009 | Data Innovations, Inc. | The "previous result" retrieved from the Specimen Management Database may not be the most recent "previous result". | Software design | Data Innovations sent an "IMPORTANT PRODUCT NOTICE" dated May 11, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter offers temporary workarounds for customers to use until corrected versions of the software is complete. Versions of the corrections will be provided on a staggered schedule. Customers will be notified by e-mail. Questions regarding this notice are directed to Support at 802-658-1955. |  | 188 | 188 | USA (nationwide) and the country of Canada. | Terminated | October 25, 2011 | 897 | 188 | Software | Output/Calculation | Software update | Software Update
Z-0085-2012 | 59892 | No | FCM | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | O'Neil Sterile Field Intermittent Urinary Catheter Kit labeled in mart:"***ORIGINAL CLOSED SYSTEM BY DR O'NEIL O'NEIL STERILE FIELD INTERMITTENT URINARY CATHETER KIT WITH IODINE SWABSTICKS (3)*** REF: ONSM 16KPES SIZE:16Fr Product No: AAM 8014KPES***SINGLE USE ONLY***Original design by: Dr. George O'Neil 200 Churchill Ave, Subiaco, WA 6008, Australia Tel.: +61 (0)8 9388 1700 Fax: +61 (0)8 9382 4952 European Representative: Advena Ltd Hereford, HR4 9DQ, UK DISTRIBUTED IN THE USA BY: American Australian Medical, Inc. 1-888-880-8871***" |  |  | September 03, 2011 | 2011 | Go Medical Industries Pty., Ltd. | O'Neil Sterile Field Intermittent Urinary Catheter Kits With Povidone Iodine Swabsticks are being recalled due to a recall on the Triad Povidone Iodine Swabsticks contained within the kit. H&P; Industries are recalling the Triad Povidone Iodine Swabsticks because it was determined that the swabsticks were manufactured without having a system for microbial testing in place at the time of release. | Nonconforming Material/Component | The firm, Go Medical Industries Pty Ltd, sent an e-mail on September 3, 2011 and follow-up with an "URGENT: MEDICAL DEVICE RECALL" letter dated September 07, 2011 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to quarantine and not sell the product in the market place; immediately perform a sub-recall and notify their customers and complete and return the RETURN RESPONSE FORM via fax at 716-883-8099 or mail to: American Australian Medical Inc, 1083 Delaware Avenue, Buffalo, NY 14209.For further details you can contact QA Manager at Go Medical Industries at +61 89 381 3226 or pmoretta@gomedical.com.au. |  | 14700 | 14,700 | Nationwide distribution. | Terminated | January 30, 2012 | 149 | 235243 | Not_Computer
Z-0104-2012 | 60004 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine Zero-P Instruments and Implants. |  |  | September 23, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION - Following review of labeling and promotional materials associated with the Zero-P Implant system, it was found that the materials reference a study where PEEK-CR is utilized as a cervical interbody spacer. PEEK-CR is not currently indicated for use as a cervical interbody spacer. | Labeling False and Misleading | Synthes (USA) sent an "URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated September 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to review the updated supplemental labeling page included with the letter and replace/attach the technique guide page in any Zero-P technique guide that customers may have. The Verification Section of the letter is to be completed and returned to Synthes via fax at 610-251-9005. For questions regarding this notice call 1-800-620-7025 ext. 5403. |  |  |  | USA (nationwide) including Puerto Rico, Washington, DC and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY. | Terminated | December 18, 2014 | 1182 |  | Not_Computer
Z-0090-2012 | 59837 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***REF QD11-6B***6 mm Fluted Ball. Use with QD11 Attachments. ***ANSPACH***Sterile, Sterilized using Irradiation***Rx Only***Made in the USA***Manufacturer: The Anspach Effort, Inc. 4500 Riverside Drive, Palm Beach Gardens, FL 33410 USA Tel: (800) 327-6887 / + 1 561 627 1080***Anspach Europe Ltd. Aston Court, Kingsmead Business Park Frederick Place High Wycombe, HP11 1LA United Kingdom Tel: +441494 616126***" |  |  | May 19, 2008 | 2008 | The Anspach Effort, Inc. | The Anspach Effort, Inc. initiated a recall on May 6, 2008 for their Sterile Bone Cutting Burrs. All affected customers were notified of the recall by letter dated 5/29/2008. | Employee error | The firm, Anspach sent an "URGENT PRODUCT REMOVAL" letter on May 29, 2008, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Screen their inventory immediately for all products listed on the attached "Product Replacement Form." Upon visual inspection if the pouch is sealed on both ends it is sterile. If one end of the pouch is not sealed it is non-sterile and should be returned for credit or replacement. 2) Complete the "Product Replacement Form" by indicating the quantity returned, and whether the customer would prefer to have their returned devices replaced, or their account credited. 3) Place products to be returned in a box and mark the box with the Return Authorization Number on the "Product Replacement Form." 4) Use the Anspach FedEx account number listed on the form to return the unused devices with the completed "Product Replacement Form." Following receipt of the form, Anspach will process the request immediately. For questions related to the product replacement, please contact Customer Service at 800-327-6887. For questions related to the Voluntary Removal, please contact Anspach Regulatory Affairs and Quality Assurance at 800-327-6887. |  | 7 | 7 | Worldwide Distribution - USA (nationwide) including the states of: CA, CO, FL, GA, IA, ID, SC, MA, MD, MI, MO, MN, MS, NY, OH, OR, PA, TN, TX, and WI, and the countries of: Africa, Netherlands, Panama, South Korea, Spain, and United Kingdom. | Terminated | October 24, 2011 | 1253 | 905 | Not_Computer
Z-0103-2012 | 60047 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hamilton brand Disposable Syringe Pack for use with Microlab F.A.M.E; Part Number: 147692 (case containing 96 syringes); Manufactured by: Treff AG, Degersheim, Switzerland for Hamilton Bonaduz AG, Bonaduz, Switzerland (Manufacturer of record); Product is distributed by Hamilton Company, Reno, NV |  |  | September 26, 2011 | 2011 | Hamilton Co | Certain lots of syringes have caused an increase in error message "DIS541", which indicates "improper dispensation". | Nonconforming Material/Component | The firm, Hamilton, sent an "Urgent: Product Recall Notification" letter dated September 26, 2011 to its customers. The letter describes the product, problem and action to be taken. The customers were instructed to return the syringes with a completed Hamilton Company Product Recall Questionnaire form via fax to: (775) 856-7259 or with an emailed signed pdf copy to wayne.mcauliffe@hamiltoncompany.com. The firm's Customer Service Department will contact customers with complete Returned Material Authorization information. If you have any questions, please do not hesitate to contact the Director of Quality at (775) 858-3000. |  | 1096 | 1096 units | Worldwide distribution: USA(nationwide) including states of: NJ, NV, TX and WA; and countries of: Brazil, P.R. China, France and Germany. | Terminated | March 06, 2012 | 162 | 1096 | Not_Computer
Z-0082-2012 | 60054 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew, HD1200 Autoclavable Camera Head |  |  | September 23, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | HD1200 camera heads may cause image issues | Process design | Smith Nephew sent an "URGENT  PRODUCT RECALL NOTIFICATION" letter dated sent September 23, 2011 to all affected consignees. The letter identified the product, problem and actions to be taken. Consignees were instructed to complete requested information, return affected product and order replacement product. To return product and order replacement product, consignees were instructed to contact the Smith &amp; Nephew Endoscopy Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. For questions call 405-917-8556. |  | 15 | 15 units | Nationwide Distribution (USA) including the states of: MI, NH, OK. | Terminated | February 16, 2016 | 1607 | 15 | I/O | Display/Image | Replace | Remove or Replace
Z-0081-2012 | 59981 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Medpro AccuFlo Elastomeric Infusion Device, Model #CT-0020-270C, packaged 12/box. The box label shows the product is Manufactured By: Medpro International (Thailand) Ltd. for Medpro Corporation Pte. Ltd., Singapore. |  |  | July 28, 2011 | 2011 | Progressive Medical Inc | The product may have a higher flow rate than specified resulting in over administration of drug solutions to the patients. | Nonconforming Material/Component | The firm, Progressive Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 28, 2011 via email and regular mail to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue their use; remove the products from inventory and quarantine them; inform their customers to return unused products to Progressive; return the product, and complete and return the RECALL RETURN RESPONSE FORM via fax to 314-961-5786. Progressive customer service will contact the customers to coordinate the return of the goods in question.Should you have any questions, please call 800-969-6331 or email to dsullivan@progressivemedinc.com. |  | 180 | 180 units | Nationwide distribution: USA including states of: FL, KY, PA, VA and WY. | Terminated | December 19, 2011 | 144 | 180 | Not_Computer
Z-0080-2012 | 59984 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | "***Vipat REF H0020-100***Lot 100533***2015-04***STERILE EO Elastomeric pump***2mL/Hr***100mL***Medpro International (Thailand) Ltd. 360, Moo 6, Chonburi Industrial Estate, Tambon Bowin, Amphur Sriracha, Chonburi 20230, Thailand***Victus Inc. 4918 SW 74 Court 33155 44400***" |  |  | September 20, 2011 | 2011 | Victus, Inc. | Victus, Inc. Miami, FL is recalling Vipat Brand Elastomeric Infusion Pump, Model H0020-100, Lot 100533. The product may have a higher flow rate than specified. The use of this product may lead to over-administration of drug solutions to patients. | Nonconforming Material/Component | Victus, Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 15, 2011 to affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use, remove the products from inventory and quarantine the products. Additionally, the firm indicated that their customer service department would contact customers to coordinate the return of affected products. Customers are directed to call 305-663-2129 ext. 110 for questions regarding this notice. |  | 63 | 63 units | USA (nationwide) including Puerto Rico and the states of FL and PA. | Terminated | January 12, 2012 | 114 | 63 | Not_Computer
Z-0105-2012 | 60017 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OPMI VARIO 700 Surgical Microscope. |  |  | September 26, 2011 | 2011 | Carl Zeiss Meditec, Inc. | An incorrect nut and bolt set was used in the axis three (3) spring assembly of the OPMI VARIO 700 surgical microscope distributed to three US locations. | Nonconforming Material/Component | Carl Zeiss Meditec had Field Service Engineers directly contact all affected customers. The Field Service Engineers will replace the defective component. Questions regarding this Field Corrective Action should be directed to the Director of Clinical/Regulatory Affairs at 925-557-4616. |  | 3 | 3 | USA (nationwide) including the states of FL, GA, and LA . | Terminated | November 04, 2011 | 39 | 3 | Not_Computer
Z-0106-2012 | 58984 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | "***Assembled for Invivo Corporation 12501 Research Parkway Orlando FL 32826***Made in U.S.A.***MRI Patient Monitoring System Model 865214***Use only with Model 989803169201 Power Adapter***REF 453564155341***SN US11401652***". |  |  | August 30, 2010 | 2010 | Invivo Corporation | Philips Invivo Corporation of Orlando, FL is recalling their Expression MRI Patient Monitoring System. During transport of the system's cart with a docked Display Controller Unit (DCU), a front caster (wheel) may become loose and fall off causing the cart to become unbalanced and fall. | Device Design | Philips Invivo Corporation issued a Field Safety Notice "URGENT - MEDICAL DEVICE RECALL" letter to all affected customers notifying them that the Expression MRI Patient Monitoring System could pose a risk for patients and users. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the user in order to prevent risks for patients or users and the actions planned by Philips Invivo to correct the problem. This notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. Customers were instructed to contact Philips Invivo Customer Support at (877) 468-4861 (option 1, then option 3) for questions or support concerning this issue. |  | 285 | 285 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI., SC, SD, TN, TX, VA, WA, and WI and the countries of: Australia, Austria, Canada, Denmark, Germany, Israel, Malaysia, Netherlands, Norway, Singapore, Spain, Switzerland and, United Kingdom. | Terminated | November 15, 2012 | 808 | 285 | Not_Computer
Z-0089-2012 | 59969 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | CORFLO 300 Enteral Pump; a volumetric rotary peristaltic intermittent enteral feeding pump; CORPAK MedSystems, Wheeling, IL 60090; catalog/reorder #20-0500. |  |  | September 19, 2011 | 2011 | CORPAK MedSystems | Corpak is removing the Corflo 300 Enteral Pump from the market due to the possibility that with fluid ingress or other damage, the keypad may not function, and the possibility of motor mount failure. | Process control | Corpak MedSystems sent Recall Notification letters dated September 19, 2011 to all CorFlo 300 Enteral Pump customers on the same date. The customers were informed that there is a possibility that with fluid ingress the keypad may become inoperable, and that there is a possibility of motor mount failure. Either event could possibly affect the rate at which formula is delivered to the patient, and fluid ingress could cause a potential of audible alarm failure. Corpak determined that the removal of the pumps from the market was the best option for their customers, and requested the return of all CorFlo 300 pumps to Corpak by no later than December 1, 2011. The customers were requested to contact Corpak at 1-800-323-6305 to arrange for the return of the pumps, and to complete and return the enclosed acknowledgement form, indicating the number of pumps on hand, via fax at 847-541-9526 or e-mail to questions@corpak.com. Dealers/distributors were requested to notify their customers of the recall. |  | 13000 | 13,000 pumps | Nationwide Distribution | Open, Classified |  |  | 13000 | Not_Computer
Z-0112-2012 | 58454 | No | BYF | Venturi mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mask adult, elastic strap style, Percent-O2 Lock air entrainment system, w/7' tube (Linde) - 50/cs; |  |  | March 25, 2011 | 2011 | Salter Laboratories, Division of Regulatory Affairs | A small number of oxygen mask components were found to have 100% obstructed airways rendering them incapable of delivering oxygen to patients undergoing respiration therapy. | Equipment maintenance | The firm, Salter Labs, called all their customers on March 25, 2011 and asked that they hold all affected product. In addition, the firm sent an "URGENT PRODUCT RECALL" letter dated April 4, 2011 and a follow-up letter dated April 7, 2011 via e-mail and certified mail to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory, return any affected lots listed, contact Salter Customer Service at 1-800-235-4203 to arrange for return of affected product and complete and return the Customer Reply Form via 661-854-6816 or toll free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 800 | 800 all varieties | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Colombia and Iceland. | Terminated | October 26, 2011 | 215 | 800 | Not_Computer
Z-0115-2012 | 59326 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE DS FULL FIELD DIGITAL ESSENTIAL MAMMOGRAPHY SYSTEM. |  |  | September 22, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurement on GE magnified images provided by Senographe DS and Essential Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | Software design | GE Healthcare sent an " URGENT MEDICAL DEVICE CORRECTION" letter dated September 21, 2011 to all affected customers. The letter describes the products, problem, and actions to be taken by the customers. The letter provides safety instructions and advises customers to complete and return a Customer Reply Form via fax to +33(0)1 30 70 39 30. A GE Healthcare service representative will contact customers if the updated software version is required. Contact the Call Center at 800-437-1171. |  | 1344 | 1344 | Worldwide Distribution-USA (nationwide) including Puerto Rico and Washington, D.C. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM BOTSWANA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA,GERMANY, GREECE, GUADELOUPE, GUYANA, HONDURAS, HONG KONG, HUNGARY, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF KOREA, RUSSIAN FEDERATION SAINT PIERRE AND MIQUELON, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE FORMER YUGOSLAV REPUBLIC OF MECEDONIA, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VENEZUELA. | Terminated | October 16, 2012 | 390 | 3457 | Software
Z-0107-2012 | 59522 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE HEALTHCARE, MSK 1.5T EXTREME MR SCANNER, MODEL AA5000. |  |  | August 29, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware that an exposed high voltage gradient wire internal to the MSK 1.5T Extreme system equipment may cause electrical shock, which may impact service personnel safety. This issue does not impact patient or operator safety. | Nonconforming Material/Component | GE Healthcare LLC sent a "Urgent Medical Device Correction letter" dated August 29, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to please ensure that all potential service personnel at your facility are made aware of this safety notification and the recommended actions.Contact your local service representative if you have any questions concerning this notification. For further questions please call ( 262 ) 513-4122. |  | 52 | 52 | Worldwide Distribution -- USA (nationwide) including the states of AL, CA, FL, ID, NY, NH, NJ, NV, OH, PA, TX, VA and WI., and countries of UNITED ARAB EMIRATES, UKRAINE, UNITED KINGDOM, SWITZERLAND | Terminated | May 29, 2012 | 274 | 52 | Not_Computer
Z-0074-2012 | 59933 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe Ventilator, Model 19250-001. Manufactured by: CareFusion, 22745 Savi Ranch, Yorba Linda, CA 92887 USA. EnVe is a portable critical care ventilator. It is suitable for a wide range of life support and critical care applications. The ventilator normally operates from external DC power. When the ventilator is portable, a Removable Battery Pack powers the unit. |  |  | September 12, 2011 | 2011 | CareFusion 203, Inc. | CareFusion has identified potential risks associated with the EnVe ventilator manufactured between December 2010 and May 2011. There are 3 issues that can result in ventilation delivery to the patient being interrupted. CareFusion will be contacting facilities to coordinate implementation of the necessary corrective actions. | Device Design | The firm, CareFusion, sent an "URGENT PRODUCT RECALL" letter dated September 12, 2011 to its customers.The letter identified affected product, described the 3 issues providing the description of the event, potential risk, and provided required action for Users. The letter also stated that CareFusion would be contacting the facility to coordinate the corrective action.If you have any questions and/or need support, please contact CareFusion Technical Support at 1-800-554-8933 from Mon. - Fri. 8:00am -5:00pm Pacific Time. |  | 128 | 128 units | Nationwide distribution: USA including states of: CA, CO, FL, IL, MI, MN, MT, ND, NJ, NM, NY, OK, PA, TX, UT, WI, and WV. | Terminated | November 22, 2012 | 437 | 128 | Not_Computer
Z-0131-2012 | 59147 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON CX7 Delta Clinical System Part #: 468205 |  |  | August 23, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed software limitations which affect only manually-loaded reagent cartridges on the CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. | Software design | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated September 20, 2010 with an attached Fax Back Response Form (via US Postal Service for US Customers) to all customers who purchased the SynchronÂż CX4/5/7 CE, CX4/5/7 Delta and CX9 ALX Clinical Systems. The letter provides the customers wtih an explanation of the problem identified and actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected product listed above to another laboratory, they were asked to provide a copy of this letter to them. Questions regarding this Product Corrective Action letter, were directed to Customer Support Center (Call Center) on the Beckman Coulter website athttp://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. Outside the United States and Canada, they were directed to their local Beckman Coulter Representative. |  | 1070 | 1070 systems total in US | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | June 28, 2012 | 675 | 1070 | Computer
Z-0122-2012 | 60041 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | Mountain First Aid Kit, Part Number 4118; UPC 0-91966-04118-3, contains 88 Pieces. |  |  | July 07, 2011 | 2011 | Lifeline First Aid LLC | The first aid kits are recalled because the kits contain Iodine Prep Pad single use wipes which were recalled by Triad. | Material/Component Contamination | Lifeline First Aid, LLC sent a " PRODUCT RECALL NOTICE-LIFELINE FIRST AID KITS WHICH INCLUDE H &amp; P [TRIAD] PVP IODINE WIPES" letter dated July 7, 2011 to all affected customers. The letter states that Lifeline First Aid, LLC has become aware of a recall of Povidone Iodine Wipes manufactured by H&P; Industries Inc. (TRIAD GROUP) and marketed under various brand names, including Triad and Swift. The recall of the Povidone Iodine Wipe is due to potential microbial contamination of the product. The letter instructs customers to collect and quarantine all affected products and to complete and return an attached Customer Response Form via fax to (503) 783-3721.Affected products will be replaced without charge and acceptable Lifeline first aid kits will be re-worked and returned using Lifeline's standard RGA process. Customers may call Lifeline customer service at 1 (877) 377-8243 for further instructions. |  | 498 | 498 kits | Nationwide Distribution | Terminated | March 29, 2012 | 266 | 26150 | Not_Computer
Z-0118-2012 | 60064 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Verardius Image Intensified X-ray |  |  | October 04, 2011 | 2011 | Philips Healthcare Inc. | Front plate broke in to two pieces causing the C-Arm to suspend on the single central bolt | Nonconforming Material/Component | Philips Healthcare Inc., sent an "URGENT - Field Safety Notice" dated September 30, 2011 to all affected customers. The notice identifies the following: what the problem is and under what circumstances it can occur; the actions that should be taken by the customer/user in order to prevent risks for patients or users and the actions planned by Philips Medical Systems to correct the problem. The notice contains important information for the continued safe and proper use of the equipment and customers were advised to retain a copy with the equipment Instruction for Use. The notice stated that Philips will replace the front plates for systems on the Unit Affected List by the end of June 2012. The notice instructed customers to contact Philips Healthcare representative for further information or support. |  | 120 | 120 units in the US | Worldwide Distribution - USA (nationwide) and the countries of: Australia, Austria, Belgium, Bharain, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Netherlands, Norway, Portugal, Russian Fed., Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Terminated | January 12, 2017 | 1927 | 266 | Not_Computer
Z-0113-2012 | 60025 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Radio used in the Precess MRI Patient Monitoring System. 3160 (Precess) and HE97B, 94044, 94045 (radio and radio kits). |  |  | September 23, 2011 | 2011 | Invivo Corporation | Philips Invivo has determined that it is possible for the system's alarm limits and other device settings to spontaneously change without user intervention. This problem is associated with intermittent and frequent loss of communication between the wireless Processing Unit and the Display Controller Unit. | Device Design | A Field Safety Notice entitled Urgent - Medical Device Correction letter, dated September 2011, with attached Field Safety Notice for customers were sent to distributors. The letter described the reason for the recall and the intention of the Field Safety Notice. Also, the hazards involved and instructions on how to identify units that are affected by the recall were provided . The actions to be taken by the user and the actions planned by the Philips Invivo were explained. Users were to follow the instructions provided on how to safely use the device until the unit can be serviced by Philips Invivo. For further information or support concerning this issue, customers are to contact Philips Invivo Customer Support at (877)-468-4861 or their local Philips Invivo representative. |  | 875 | 875 | Worldwide Distribution -- Australia, Austria, Brazil, Canada, China, France, Germany, Malaysia, Norway, Portugal, Spain, Sweden, The Netherlands, United Arab Emirates, United Kingdom, and US. | Terminated | August 07, 2013 | 684 | 875 | I/O | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0073-2012 | 59536 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho OSI Modular Table System; |  |  | July 29, 2011 | 2011 | Mizuho Orthopedic Systems Inc | Reports of injury related incidents while using the Mizuho OSI Modular Table System. | Use error | Mizuho Orthopedic Systems Inc sent a "Field Advisory Notice (Recall)" letters dated July 29, 2011 to all customers using the Modular Table System via Federal Express. The Field Advisory Notice reference the devices involved and explain that there is no need to return any product or to discontinue the use. The letter provides WARNINGS and RECOMMENDATIONS for all users to consider when using the devices and request that the response form be completed as instructed in the advisory communication. The letters also recommend that awareness, training, inspection prior to use, and proper maintenance be established and practiced. The letter recommend that the Field Advisory Notices be posted in prominent areas and communicated to all staff and potential users of this device. Additional support and training can be obtained by contacting Mizuho OSI's Customer Resource Group at 1-800-777-4674. For further questions concerning this notification, contact the Director of Customer Resource Group for assistance (1-800-777-4674, outside the USA 00+1-510-476-8199).Revised letters were approved by CDRH and sent out in January 2012. |  | 15011 | 15,011 units | Worldwide Distribution - Nationwide USA | Terminated | August 27, 2013 | 760 | 15011 | Not_Computer
Z-0138-2012 | 59866 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker Communications, 1410 Lakeside Parkway, # 100, Flower Mound, TX 75208 |  |  | August 23, 2011 | 2011 | Stryker Communications Corp | Installation records for the SwitchPoint Infinity, Booms, Lights, and Flat Panel systems are incomplete or inadequate. | Process control | Stryker sent an URGENT FIELD SAFETY NOTICE: RA2010-245 dated August 23, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached Customer Response Form within 5 business days to Stryker. A Stryker representative would contact them at their convenience to schedule a re-inspection for all of the units identified and still in use. In the event the inspected equipment does not meet Stryker's established specifications, their representative will visit their facility and hopefully resolve the issue. If a subsequent visit is required the date will be scheduled in full consideration of their services.For questions regarding this recall call 972-410-7310. |  |  |  | Worldwide Distribution including Puerto Rico, Brazil, Ecuador, Venezuela, Costa Rico, Bolivia, Peru, Mexico, China, India, Korea, Philippines, Singapore, Malaysia, Taiwan, Middle East, Russia, Africa, Poland, South Africa, Australia, New Zealand, Canada, Japan, Denmark, Finland, Norway, Sweden, Belgium, The Netherlands, France, Germany, Austria, Italy, Portugal, Spain, Romania, Switzerland, and United Kingdom. | Terminated | August 20, 2012 | 363 |  | Not_Computer
Z-0135-2012 | 60175 | No | FFX | Gastrointestinal motility monitoring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AR ManoShield, MSS-3599, 10 pack. |  |  | October 12, 2011 | 2011 | Sierra Scientific Instruments Inc | The recall was initiated after Given Imaging received a test report during a Canadian licensing process where the AR ManoShield exhibited pinholes and did not pass the test, the Field Safety Corrective Action (FSCA) Committee was convened. As a result of the data discussed during this meeting, a decision was made to voluntarily recall this product, and cease its further commercialization or distr | Nonconforming Material/Component | Given Imaging sent an "Urgent Product Recall 1st Notice" letter with an attached Device Recall Confirmation Form (via Fed Ex with signature for tracking purposes) to all customers who purchased the AR ManoShield, part numbers MSS-3599. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were instructed to locate the affected product and removed from use immediately. The letter stated that if the affected product have been further distributed, forward this letter immediately to all customers who received the affected product. Customers were asked to complete and return the enclosed confirmation form as per the instructions. Customers with any questions were instructed to contact Karla Guerrero at (310) 641-8492 ext. 2138 or email at karla.guerrero@givenimaging.com. |  | 292 | 292 units | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Canada, Chile, China, Columbia, Ecuador, France, Germany, Israel, Mexico, Poland, Portugal, Sweden, and Switzerland. | Terminated | August 30, 2012 | 323 | 292 | Not_Computer
Z-0133-2012 | 59521 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, InnovaÂż 3100IQ (Optima Edition) Fluoroscopic Vascular X-ray System. |  |  | August 24, 2011 | 2011 | GE Healthcare, LLC | A potential weakness on the monitor suspension end-stop device may impact user safety. The end-stops may break if the monitor-suspension is repeatedly pushed with excessive force on the end rails. | Nonconforming Material/Component | A GE Healthcare "Urgent Medical Device Correction" letter dated August 22, 2011 was sent out 8/24/11. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Users were told to operate the suspension without excessive force. If one of the end-stops breaks, then the system should not be used and their local GE Healthcare Service Representative should be contacted. GE will contact customers to replace all end-stops at no cost to the customer. Questions or concerns should be directed to 800-437-1171 in the US and 0120-055-919 in Japan. For other countries, customers can contact their local GE Healthcare Service Representative. |  | 15 | 15 | Worldwide Distribution -- USA, including the states of TX &amp; NC and the countries of TURKEY, POLAND, MEXICO, CHINA, CAMBODIA, BRAZIL, BOLIVIA, and SAUDI ARABIA. | Terminated | May 29, 2012 | 279 | 15 | Not_Computer
Z-0134-2012 | 60168 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6301 Dual Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product Code # 2M8064. |  |  | August 26, 2011 | 2011 | Baxter Healthcare Corp. | The Flo-Gard infusion Pump was released to the customer with OOS values for the Air Sensor Calibrations with the Air in Tubing test. | Employee error | A Baxter Healthcare technician initially notified the sole customer on August 26, 2011 of the defective product and requested that the product be quarantined. On September 14, 2011 a technician visited the site to correct the problem observed on the Flo-Gard Infusion Pump. All required tests were performed and the results were found acceptable. |  | 1 | 1 unit | Distributed only in Puerto Rico. | Terminated | October 31, 2011 | 66 | 1 | Not_Computer
Z-0141-2012 | 59909 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Presource Custom Sterile Surgical and Procedure Kits; surgical trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components for which Cardinal Health has been unable to confirm the manufacturer's compliance with section 502(o) of the FD&C; Act: |  |  | September 03, 2011 | 2011 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturer(s) does not have all necessary FDA registrations or filings needed for this product. | Component design/selection | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated September 2, 2011 to all affected consignees via UPS next day air, informing them that their Presource kits listed in the attached spreadsheet contain various components listed on the attachment, which are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels to affix to the front side of each affected Presource kit in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. If the consignee does not wish to accept overlabeled product or to overlabel the product in their possession, they were asked to contact Presource Sales Operations at 800-766-0706 or their Sales Representative for further instructions. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact the following: Hospitals - Presource Sales Operations at 800-766-0706, Distributors - Distributor Management Group at 800-635-6021, and Federal Government facilities - Government Service at 800-444-1186.For questions regarding this recall call 847-887-6412. |  |  |  | Nationwide Distribution | Open, Classified |  |  |  | Not_Computer
Z-0139-2012 | 60006 | [6 empty] | Picis Anesthesia Manager, PACU Manager,Critical Care Manager-8.2 Service Pack 4 or higher |  |  | August 12, 2011 | 2011 | Picis Inc. | Periop/Critical Care Manager EHR Software Application displayed one patient's demographic and health information in place of another patient that was in current view. | Software design | Picis Inc. sent a "Safety Notification letter" dated August 12, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed until this service pack is applied, we recommend: * For patients who have been transferred from other facilities within the same VISN, you should discontinue the use of these Picis applications. Manually check the patient's name with that displayed in the Picis patient blue banner to ensure you catch all of these cases. * For all other patients, you may safely continue the use of these Picis applications.For further questions please call (781) 557-3000. |  | 10 | 10 | Nationwide Distribution -- AZ, DE, FL, IL, NY, OH, PA, WI and WV. | Terminated | March 23, 2012 | 224 | 10 | Software | Output/Calculation | Software update | Software Update
Z-0142-2012 | 59582 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 including all Catalog Numbers to include: Catalog # 500AVHCT with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. |  |  | September 23, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo CVS identified a defective circuit board, which could potentially cause the CDI monitor to not function after a power loss. Once power is lost, the monitor will not power up again, even if AC or battery power are restored. | Component design/selection | Terumo Cardiovascular Systems Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" notice dated September 23, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides information on what to do in the event of a failure of the device. Additionally, a Customer Response Form was attached to the letter for customers to complete and return. Terumo CVS will schedule the return of affected products for inspection and replacement of defective parts. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice. |  | 432 | 432 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Chile, Japan, Philippines, and Thailand. | Terminated | April 11, 2013 | 566 | 432 | Hardware | Device Operation | Replace component | Remove or Replace
Z-0143-2012 | 55537 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Bioplate Titanium Fixation System, Sterile Kit. |  |  | March 31, 2010 | 2010 | Bioplate Inc | There is potential defect in seal of the sterile barrier containing the device. | Process design | BIOPLATE sent an URGENT - PRODUCT RECALL NOTIFICATION letter dated March 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to return any recalled product they have at their siteand if distributed further they should forward the information contained in the letter to all customers. If a device from one of the lots noted above has already been used, Bioplate would like to be advised of that use.They asked that consignees complete the enclosed Fax form and transmit the requested information to them as soon as possible at (310) 815-2126.Returned recalled product was to be sent to: Bioplate3643 Lenawee AvenueLos Angeles, CA 900163643 Lenawee Avenue, Los Angeles, CA 90016For any questions regarding this recall call (310) 815-2100. |  | 18246 | 16,020 Individual Kits; 2,226 Boxes. | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, Japan, Taiwan, Middle East and Australia. | Terminated | April 04, 2012 | 735 | 18246 | Not_Computer
Z-0144-2012 | 60057 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż Ventralex" ST Medium Circle with Strap Hernia Patch Diameter 6.4cm /2.5", Catalog Number: 5950008. |  |  | October 04, 2011 | 2011 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: BardÂż Ventralex" ST Medium Circle was packaged in a pre-printed carton identified for a BardÂżVentralex" ST Large Circle. | Packaging process control | Davol Inc. issued an Urgent; Medical Device Recall letter on 10/4/11 for the mislabeled Bard Ventralex ST Medium Circle. The letter identified the affected product and explained the reason for recall. Customers are to immediately examine their inventory and quarantine any affected product. In addition, if any product was further distributed then those customers should be notified of the recall as well. Accounts are to contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product to Davol. The attached recall effectiveness check form should be completed and returned. Questions should be directed towards Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Services &amp; Support Department at 1-800-562-0027. |  | 96 | 96 units | Nationwide Distribution | Terminated | September 20, 2013 | 717 | 96 | Not_Computer
Z-0145-2012 | 60046 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | MERZ AESTHETICS, RadiesseÂż Volume Advantage 1.5 CC, Injectable Implant, Part Number 8071M0H1 and 8071M0K1, Sterile, Rx ONLY. |  |  | August 18, 2011 | 2011 | Merz Aesthetics, Inc. | There is a potential leakage of product that may occur with Merz Aesthetics 1.5cc syringes of Radiesse Dermal Filler. When using the syringe, Health Care Professionals should be alerted to the potential for product that my leak pass the plunger at the proximal end of the syring. The defect may result in infection requiring medical intervention to prevent impairment or damage. To date Merz has n | Nonconforming Material/Component | On 8/18/11, consignees were sent a Merz Aesthetics "Urgent Medical Device Recall" letter dated August 15, 2011. The letter was addressed to "Dear Health Care Professional". The letter described the product and the problem and listed the details regarding this issue. |  | 311582 | 311,582 | Worldwide distribution including USA, Puerto Rico, Dominican Republic, Hong Kong, Israel, Netherlands, Germany, South Korea, and Canada. | Terminated | May 01, 2012 | 257 | 311582 | Not_Computer
Z-0146-2012 | 58162 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Labeling for RIO Robot Unit is comprised of three main components: 20399 RIO SURGICAL ARM, 201251 RIO GUIDANCE MODULE, 200294 RIO CAMERA STAND ASSEMBLY***PN 203999 SN ROB 125 2010-12 V 100/120/230 A 9.6/8.0/4.2 Hz 50/60 Class I Equipment. Conforms to IEC 60601-1/A2: 1995, EN 60601-1/A2: 1995 UL 60601-1: 2003, CAN/CSA-C22.2 No. 601.1-M90.***Manufactured in USA***MAKO SURGICAL CORP. 2555 DAVIE ROAD, FT. LAUDERDALE, FL 33317. |  |  | January 20, 2011 | 2011 | Mako Surgical Corporation | MAKO Surgical Corp. is recalling their RIO Robotic Arm Interactive Orthopedic System (RIO) due to software issue that exist that could potentially result in a bone resection. No adverse events reported. | Software change control | The firm, MAKO Surgical Corp., sent an "RECALL NOTICE 2.4. SITES" letter dated February 1, 2011 to its customers. The letter describes the product, problem, actions to be taken which includes a workaround and an updated user guide identifying the appropriate workflow. The customers were instructed to follow the user guide and workflow and to complete and return the enclosed confirmation memo to their MAKOplasty Specialist or return it in the pre-addressed, pre-paid envelope provided. MAKOplasty included with the letter new steps in the surgical workflow in the use of the RIO system to ensure error value accuracy; worked on a software update that will not allow the user to proceed with error values that do not support the required accuracy-this software was to be supplied to customers in February 2011, and trained MAKOplasty Specialist will make available written updates to the RIO use instructions as well.If you have any questions or concerns, please do not hesitate to speak with your MAKOplasty Specialist or contact the Manager of Customer Relations at 954.927.4022 x447. |  | 70 | 70 | Worldwide distribution: USA including states of: AZ, CA, CO, FL, GA, IA, IL, LA, MI, MO, MS, NC, NJ, NV, NY, OK, OH, OR, PA, RI, TX, UT, VA, WA, WI, and WV; and countries of: Scotland, Italy and Korea. | Terminated | December 13, 2011 | 327 | 70 | Software | Physical Safety Hazards | Software update | Software Update
Z-0014-2012 | 59949 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Dynamic Multileaf Collimator (DMLC) |  |  | May 25, 2011 | 2011 | Elekta, Inc. | It is possible to load different patients in the record and verify (R&V;) system and in the multilead control software (MCS). | Device Design | Elekta, Inc. sent an "IMPORTANT NOTICE" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several precautions customers should use until Elekta finds a solution. Customers will be contacted when a solution is found. |  | 27 | 27 | Nationwide Distribution-USA including the state of CT, FL, GA, IA, MN, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA , and WA. | Terminated | January 14, 2014 | 965 | 27 | Computer
Z-0147-2012 | 60240 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medrad(R) Sterile Disposable Stellant Syringe Kit, Catalog # SDS-CTP-SPK, Lot # 112865. |  |  | October 24, 2011 | 2011 | Medrad Inc | Packaging may exhibit open seals - sterility compromised | Packaging | Medrad sent a " Urgent Medical Device Recall letter" dated October 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed if they have any of the affected product lot, please contact MEDRAD customer support to make arrangements for the return of the product. Call 1 -800-633-7231, select option 5, then select option 1. Indicate that you are calling about the affected syringe kit recall.Complete and sign the attached response form and fax it back to MEDRAD at (412) 406-0942.Please complete and fax this form even if you do not have any of the referenced product so we can maintain proper accountability for all syring kits that have been shipped. |  | 17380 | 17,380 | Nationwide Distribution -- including the States of : AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, VS, WA, WV, and WI. | Terminated |  |  | 17380 | Not_Computer
Z-0149-2012 | 60034 | Yes | OAE |  |  | Cardiovascular | PMA | Medtronic Arctic Front Catheter, Sterile EO, Medtronic CryoCath LP, Kirkland, Quebec, H9H 5 H3, Canada, Model 2AF282 and 2AF231 ( not distributed in US) |  |  | August 29, 2011 | 2011 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic CryoCath has identified specific lots of Arctic Front CryoAblation Catheters that were distributedwith incorrect programming that may result in false expired catheter notifications. Impacted lots are 21571,25375, 25376, 25377, and 25381 of the Arctic Front Catheter Model Number 2AF282. Medtronic CryoCathis communicating this information to the FDA.On affected catheters, the incor | Employee error | Medtronic sent an "Important Device Recall" letter dated August 2011, hand delivered by Medtronic Field Representatives on August 29, 2011 to all affected customers. The letter described the product affected, the problem, and the actions needed to be taken. The customers were instructed to return all unused catheters from affected lots for credit and replacement. The letter states that Medtronic AF Solution representative will assist customers with the return of affected product. For further questions contact your local Medtronic AF Solutions representative. |  | 263 | 263 | Worldwide Distribution - USA (nationwide) including the states of: AK, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MN, MS, MO, NE, NY, NC, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI. and the countries of: Austria, United Kingdom. | Terminated | June 07, 2012 | 283 | 263 | Software | Alarm/Message | Return | Remove or Replace
Z-0150-2012 | 59991 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RSP Humeral Socket Reamer (Size 32 SML): |  |  | September 28, 2011 | 2011 | Encore Medical, Lp | DJO Surgical has determined that the alignment pins cannot be consistently cleaned using the technique provided in the cleaning/sterilization section of the instrument IFU. | Device Design | DJO Surgical sent an URGENT FIELD SAFETY NOTICE letter dated September 28, 2011, to all affected customers. The letter identified the product. the problem, and the action to be taken by the customer. Customers were instructed to return the RSP Humeral Socket Reamers and RSP Baseplate Rim Planers. The RSP Humeral Socket Reamers will be reworked and returned to the customer. The Baseplast Rim Planer will be replaced when a suitable design is available. During the interim until instruments are returned, the use of FMP acetabular reamers size 42mm, 43mm, 45mm (DJO PIN 803-05-226, 227, 200) for the humeral preparation was validated to be effective to prepare the humerus for the implant.Customers were instructed to immediately notify the hospitals and surgeons using these devices and document it on the Field Safety Notice Response Form. For questions regarding this recall call 512-832-9500. |  | 399 | 399 Reamers | Worldwide Distribution - USA including AR, AZ, CA, CO, FL,GA, ID, IN, IL, LA, MA, MD, ME, MI, MO, NJ, NM, NY, OH, OK, PA, PR, SC, TN, TX, UT, VA, WA, and WI and Internationally toGermany, South America, Italy, Saudi Arabia, United Kingdom, and Canada. | Terminated | March 29, 2013 | 548 | 532 | Not_Computer
Z-0158-2012 | 60254 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | "***BIOMET 3i***4555 Riverside Drive Palm Beach Gardens FL 33410 USA***REF AC4425***Conical 25 ANGLED ABUTMENT***4.1mm(D) X 4mm(H)***Abutment; Pilier: Abutment; Abutment; Pilastro; Pilar***RX Only***Sterile***BIOMET 3i Dental Iberica S.L. WTC Almeda Park. Ed. 1. Planta 1'***". |  |  | September 28, 2011 | 2011 | Biomet 3i, LLC | On 09/28/2011 Biomet 3i, Palm Beach Gardens, FL initiated a recall of their Conical Angled Abutment, Model # AC4425 Lot 887628-5 and their Standard Conical Abutment Cylinder Model # SWCA62 Lot 848478. The packaging for the referenced products may not have been completely sealed prior to shipment. | Employee error | Biomet 3i, LLC sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to inspect their inventory and return any affected product for replacement.A Product Recall Response Form was attached for customers to complete and return via fax to 561-514-6316. Customers are directed to call 1-800-342-5454 for questions regarding this recall. |  | 122 | 122 | Worldwide Distribution-USA (nationwide) and the countries of Chile, Italy, Japan, Korea, Spain, and Sweden. | Terminated | March 09, 2012 | 163 | 130 | Not_Computer
Z-0152-2012 | 60100 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ProScreen PSCup-5M Waived, Catalog Number: PSCup-5M Waived |  |  | June 07, 2011 | 2011 | Ameditech Inc | The recall was initiated because Ameditech, Inc. received a complaint from US Diagnostics (USD) that First Check Drug Screen Cups were packaged in the ProScreen Cup boxes. Some of US Diagnostics' customers received ProScreen 5-panel cups that contain First Check THC cups. The recalled products are mislabeled. If the USD customer receives the FirstCheck tes, there will be no instructions for use | Packaging process control | The recall communication was initiated on June 9, 2011 with Ameditech Inc. forwarding an Urgent Medical Device Recall notification with an attached Verification Form (via phone and written notification) to the affected distributor. The letter provides the distributor with an explanation of the problem identified and actions to be taken. The distributor was required to: cease distribution of the affected lot and provide remaining inventory count to Ameditech; destroy any remaining inventory of these lots and document destruction per site requirements; and ensure that all user of the device have received a copy of the Urgent Medical Device Recall. In addition, the distributor was instructed to complete and return the enclosed Verification Form with in 10 days.For any questions about the information contained in this notice, contact Ameditech at amd.complaint@alere.com or call 858-535-1968. |  | 2000 | 2000 test cups (800 kits) | Nationwide Distribution (USA) including the state of Alabama. | Terminated | January 18, 2012 | 225 | 2000 | Not_Computer
Z-0154-2012 | 60114 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | 37% Etching Gel-1ml syringe 20 pk, part # T06A |  |  | March 31, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | Pentron Clinical is voluntarily recalling all lots of products labeled as 37% Phosphoric Acid Etching Gel produced since April 2007 because the product continues to extrude out of the syringe even after pressure is no longer applied, and it contains a higher percentage of phosphoric acid than that indicated on the label. | Nonconforming Material/Component | The firm, Pentron Clinical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 31, 2011, with attached Return Form (via USPS 1st class mail) to all their consignees/customers. The letter describes the product, problem and action to be taken.All consignees/customers were instructed to complete the Recall Return Form via fax to: 1-877-677-8844 and to return any affected product in their inventory. Affected products will have expiration dates between April 2011 and April 2015.Consignees/customers with questions can contact Pentron Clinical Customer Service at (800) 551-0283 directly to handle the arrangements of a quick return and replacement. |  | 41744 | 41,744 units | Worldwide distribution: USA (nationwide) and countries of: Armenia, Austria, Australia, Antigua &amp; Barbuda, Barbados, Belarus, Bulgaria, Canada, Czech Republic, Croatia, Cyprus, Ecuador, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Kenya, Latvia, Lebanon, Malaysia, Mexico, Morocco, Netherlands, Poland, Romania, Russia, St. Vincent &amp; Grenadine, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, UAE, Ukraine, UK, and Vietnam. | Terminated | August 09, 2012 | 497 | 41744 | Not_Computer
Z-0160-2012 | 57234 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | Microseal Master Cone Part Number: 815-9033 |  |  | September 09, 2010 | 2010 | Ormco Corporation | The recall was initiated because SybronEndo has confirmed that the Microseal Master Cones has been mislabeled. | Error in labeling | Sybron Dental Specialties, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return the affected lots of product.. Products, would be replaced at no charge or credited to their account. Customers were instructed to contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. The affected product should be returned to SybronEndo at the following address. Please label returned product "RECALLED PRODUCT Attention: Customer Returns"SybronEndo 1332 South Lone Hik AvenueGlendora, CA 91740ADDITIONALLY, THEY KINDLY REQUEST CONSIGNEE COOPERATION IN COMPLETING AND FAXING BACK THE ENCLOSED ACKNOWLEDGEMENT/RETURN FORM. THIS FORM MUST BE COMPLETED AND RETURNED WHETHER OR NOT YOU HAVE ANY PRODUCT TO RETURN.Customers were also asked to identify and recover the affected product lots listed above that may have been shipped to their customers.For questions regarding this recall call 909-962-5600. |  | 200 | 200 units | Worldwide Distribution - USA including AR, CA, FL, PA, GA, IL, TX, WA, NY, IN, AL &amp; MT; Internationally to Canada, CIS (Commonwealth of Independent States, Israel, United Kingdom, Sweden, France, Germany, and Italy. | Terminated | November 21, 2011 | 438 | 200 | Not_Computer
Z-0161-2012 | 60003 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Spine T-PAL Spacer System. Transforaminal posterior atraumatic lumbar spacer system. |  |  | September 23, 2011 | 2011 | Synthes USA (HQ), Inc. | LABELING CORRECTION for Medical Device - previous labeling and promotional materials associated with the T-PAL Spacer system reference the use of T-PAL spacer with allograft material. T-PAL Spacer is currently indicated for use with autograft material. All accounts with record of receiving T-PAL Implant will receive notification of the labeling correction and a copy of the updated brochure. All | Error in labeling | SYNTHES sent an URGENT MEDICAL DEVICE LABELING CORRECTION letter dated September 23, 2011, to all affected customers. The labeling has been revised. Updated information packets have been sent to all accounts which included the final labeling and communications to accounts and sales force. Customers were instructed to review the updated supplemental labeling provided in the packageCustomers were instructed to attach the supplementary technique page to an T-PAL Spacer Brochure that may be in their facility. They can access an electronic version of the Brochure by accesssng Product Literature at the following link:http://us.synthes.com. The new version can be identified by code J10120 located at the bottom right corner of the back cover.To confirm receipt and understanding of the information, customers were requested to complete the Verification Section of the letter and return it to Synthes by either of the following methods:Mail provided postage paid envelopeFax 610-251-9005Scan/e-mail to FieldAction@synthes.comFor any questions call 1-800-420-7025 x5403 or contact their Synthes Spine Sales Consultant. |  |  | N/A - Labeling Correction | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, HI, KS, KY, LA, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. and Puerto Rico and Washington, D.C; Internationally to Canada and Japan | Terminated | December 18, 2014 | 1182 |  | Not_Computer
Z-0162-2012 | 60094 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | ARCHITECT Cyclosporine; an in-vitro diagnostic chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of cyclosporine in human whole blood Distributed by Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA; list 1L75-25. |  |  | October 07, 2011 | 2011 | Abbott Laboratories | The ARCHITECT Cyclosporine assay has the potential to generate falsely elevated or depressed results with patient samples due to precision exceeding the Package Insert claim of &lt;15% total CV. Calibrators and Controls also have the potential to be impacted by the assay imprecision. | Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated October 7, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to only run ARCHITECT Cyclosporine using Reaction Vessel (RV) lots manufactured with the same resin as follows: a) RV lots beginning with lot 06083P100 and higher can be used together or, b) RVs lots lower than 06083P100 can be used together. Any RV lot number that does not begin with a zero (e.g. 90032P100) should also be considered lower than 06083P100. Further instructions were included for the accounts to follow when running ARCHITECT Cyclosporine assays. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051. Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. A second letter was sent to the accounts on 11/23/11 to simplify the instructions for identifying the Architect Reaction Vessel lots that can be used together for use with the Architect Cyclosporine assay. |  | 8913 | 8,913 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bahrain, Bangladesh, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Brunei, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Terminated | November 07, 2012 | 397 | 8913 | Not_Computer
Z-0165-2012 | 60241 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Abbott brand IntraLase Femtosecond Laser System, Model 2; Catalog Numbers: 20003D/J/K, R20003D/J/K, 20005D/J/K; Product is manufactured and distributed by AMO Manufacturing USA, LLC, Milpitas, CA |  |  | October 12, 2011 | 2011 | AMO Manufacturing USA, LLC | The firm has identified a discrepancy in cut depth that may affect deep corneal incisions (e.g. 300 Âżm or deeper) created using the Laser System. | Process design | Abbott Medical Optics Inc. (AMO) issued an "ADVISORY NOTICE" letter dated October 12, 2011 to all affected customers. The letter identifies the affected products, the problem, and the actions needed to be taken. Customers were instructed to exercise caution when making cuts that approach the corneal endothelium. The letter includes a 1-page Operator's Manual addendum that clarifies the recommendations. The letter states that AMO will conduct a field correction to adjust the EOF on the affected FS Laser Systems. The letter includes a response form to be completed and returned as per the instructions. Customers were instructed to notify all appropriate staff/personnel of the Advisory Notice. For any further questions regarding the information or recommendations, contact an AMO representative at 1-800-266-3375 |  | 1226 | 1226 total laser systems | Worldwide Distribution - USA (nationwide) | Terminated | May 01, 2013 | 567 | 1226 | Not_Computer
Z-0164-2012 | 60143 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | BioPro Femoral Head 28 mm +18, Item 19010. |  |  | September 21, 2011 | 2011 | Biopro, Inc. | BioPro Femoral Head Outer Packaging was found lacking proper seal ,compromising sterility. | Packaging | Biopro, Inc notified all affected customers via telephone on September 21, 2011 with instructions to quarantine and return affected product to the manufacturer. The firm sent a follow up "RECALL NOTICE" dated September 29, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter gives instructions on how to return affected product. Contact Customer Service at 1-800-252-7707 for questions regarding this recall. |  | 8 | 8 units | USA (nationwide) including the states of FL, MI, and TX. | Terminated | February 15, 2013 | 513 | 12 | Not_Computer
Z-0174-2012 | 60255 | No | KZH | Syringe needle introducer. | General Hospital | General Hospital | 510(k) | 1.0mL 28 G x 1/2" blister packaged insulin syringe Catalog Number 329424; Lot Number 0039920 |  |  | June 15, 2011 | 2011 | Becton Dickinson &amp; Company | Experiment labels may be on shelf cartons of 1.0ml 28G x 1/2 blister packaged insulin syringes. | Labeling mix-ups | The firm, BD, sent an "PRODUCT RECALL NOTIFICATION" letter and response form dated June 15, 2011 to their consignee/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately review their inventory for the affected product, remove the product from inventory, quarantine the product and return it to BD; if the product was further distributed, they should identify their customers and notify them at once, and complete and return the attached Tracking/Verification Form via fax to: 201-847-4853.If you have any questions, contact Customer Service at 1-800-237-2762, Monday -Thursday, 8:00am-6:30pm and on Friday, 8:00-5:30pm Eastern Standard Time. |  | 10230 | 10,230 shelf cartons (1,0230,000 syringes) | Nationwide Distribution. | Terminated | July 27, 2015 | 1503 | 10230 | Not_Computer
Z-0173-2012 | 60183 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter System 1000 Tina Single Patient Hemodialysis Delivery System; Baxter Healthcare Corporation, Renal Division, McGaw Park, IL 60085; product code S1000L3P. |  |  | September 19, 2011 | 2011 | Baxter Healthcare Renal Div | The Tina System was released to the customer with OOS values for the Pressure Displays Accuracy and Leakage Rate Test. | Employee error | Baxter Healthcare Corporation sent a service technician to the lone customer site on September 19, 2011 to correct the affected product. The technician executed all tests as established in the Service Manual, with acceptable results. |  | 1 | 1 unit | Distributed in Puerto Rico only. | Terminated | November 07, 2011 | 49 | 1 | Not_Computer
Z-0169-2012 | 60069 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits; general surgery trays; Cardinal Health, McGaw Park, IL 60085; The kit catalog numbers listed in the code information section below contain one or more of the following listed components for which Cardinal Health has been unable to confirm the manufacturer's compliance with section 502(o) of the FD&C; Act: |  |  | September 30, 2011 | 2011 | Cardinal Health | Some of the components contained in the Presource kits are being recalled because Cardinal Health has determined that the component manufacturers do not have all necessary FDA registrations or filings needed for these components. | Component design/selection | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated September 30, 2011 to the affected wholesale consignees via UPS next day air, informing them that their Presource kits listed in the attached spreadsheet contain various components listed on the attachment, which are being recalled because the component manufacturer(s) do not have all the necessary FDA registrations or filings needed for these components. Enclosed with the letter were recall labels to affix to the front side of each affected Presource kit and shipping carton in their inventory. The labels have an ID code which coincides with the code on the component attachment to the letter. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of pack/modules labeled. For component credit, the consignees were instructed to contact Distributor Management Group at 800-635-6021. This wholesale recall was expanded to include additional kits via letter dated October 5, 2011, with the same instructions as listed above.For questions regarding this recall call 800-292-9332. |  |  |  | Nationwide Distribution including California, Florida, Illinois, Kentucky, Minnesota, Missouri, Ohio, Tennessee, Washington, West Virginia and Wisconsin | Open, Classified |  |  |  | Not_Computer
Z-0171-2012 | 58766 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | ***REF IOSS511***OSSEOTITE Certain Implant 5 x 11.5mm***LOT 910695***Sterile using Radiation***Rx Only***BIOMET 3i Dental Iberica S.L.WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n 08940 Cornella de Llobregat (Barcelona) Spain***Biomet 3i Dental Implants are intended for surgical placement in the upper or lower jaw to provide a means for prosthetic attachment. The Thread form titanium screws (cylindrical dental implants) designed to be placed into mandibular or maxillary jaws in order to affix a dental prosthesis (edentulous or partially edentulous segments) that replaces 1 or more missing teeth. |  |  | November 18, 2010 | 2010 | Biomet 3i, LLC | The product may contain a manufacturing defect, that if present, may prevent the driver; cover screw or abutment from fully engaging the implant. | Process control | The firm sent an Urgent: Medical Device Recall letter, dated November 2010, to notify customers that received the Osseotite Certain Implant (IOSS511) of the problem. The customers were instructed to return the unopened device(s) to Biomet 3i for replacement. If customers notice an issue, they were to contact Customer Service at 1-800-342-5454 for assistance. Customers were to respond via fax and return the implant(s) to BIOMET 3i using the address and incident number provided. Customers with questions or concerns were instructed to contact Mark Mashburn at 561-776-6906. |  | 241 | 241 | Worldwide Distribution -- USA , BR, JP, TW, CA, Germany, and Guam. | Terminated | November 09, 2011 | 356 | 241 | Not_Computer
Z-0176-2012 | 59684 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Microtainer MAP Microtube for Automated Process with K2EDTA. |  |  | February 11, 2011 | 2011 | Becton Dickinson &amp; Company | BD Microtainer MAP Microtube is being recalled due to the potential clotting of blood specimens resulting from a reduced amount of EDTA additive. | Device Design | A Becton Dickinson Sales Consultant conducted a "on-site" visit to the single Customer account, Florida Hospital, on February 11, 2011. All units were successfully recovered. All product recovered was quarantined and destroyed. Becton Dickinson believes that all reasonable efforts to communicate the issue to the affected consignee have been completed. If you have any questions or request additional information please do not hesitate to call (201) 847-5636. |  | 10800 | 10,800 units | Distributed only to Florida. | Terminated | November 10, 2011 | 272 | 10800 | Not_Computer
Z-0175-2012 | 59996 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter Assembly Manufactured by: Stryker Spine SAS Z. I. Marticot - 33610 - Cestas France; Distributed in the USA by: Stryker Spine 2 Pearl Court. Allendale, NJ 07401-1677 USA |  |  | June 24, 2011 | 2011 | Stryker Spine | There have been reports of difficulty in rod rotation using the newly designed Mantis Rod Inserter. | Device Design | Stryker notified branches/agencies, surgeons and hospitals beginning June 24, 2011 with an Urgent Product Recall letter, dated June 24, 2011, via Fed Ex. The letter identified the affected product, the issue, potential hazards, risk mitigations, and the return process. Customers were asked to examine their inventory and reconcile any inserters and shafts with the catalog numbers provided using the Customer Response Form. Customers are to fax the form to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should be retrieved and returned. Questions should be directed to 201-760-8298. |  | 1138 | 471 units US, 667 units international | Nationwide Distribution | Terminated | November 13, 2012 | 508 | 1138 | Not_Computer
Z-0177-2012 | 59425 | No | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Capintec CRC 25 PET, item number 5130-3217 with Software Revisions 2.01 and 2.02 |  |  | July 08, 2011 | 2011 | Capintec Inc | Software error in Rev. 2.01 and Rev 2.02 for Capintec CRC 25PET | Software design | Capintec, Inc. sent a Software Error Alert letter dated July 8, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contac their Capintec Customer Support Representative to receive a software upgrade. For questions call Capintec Customer Support at 1-800-631-3826. |  | 69 | 69 | Worldwide distribution - United States including CA, FL, IL, MA, NJ, OK, PA, and VA; and internationally to Brazil, China, India, Indonesia, Italy, Japan, Korea, South Korea, and Turkey. | Terminated | December 10, 2013 | 886 | 69 | Software |  | Software update | Software Update
Z-0181-2012 | 60022 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Clinac, Trilogy, Trilogy Tx and Novalis linear accelerators. |  |  | September 09, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Varian has received reports in which a user has remotely rotated the gantry into contact with the couch or with the patient, in both the manual mode and the automate mode. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated September 9, 2011 to all affected customers via Federal Express with attached instructions. The letter provides the description of problem; references, details, recommended user actions, and varian actions.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***In addition, the letter also states "THIS LETTER IS NOT A SUBSTITUTE FOR THE CAREFUL REVIEW OF ALL USER DOCUMENTATION ASSOCIATED WITH YOUR PARTICULAR EQUIPMENT AND SOFTWARE."If you have any questions, contact Varian Oncology Help Desk at USA and Canada: 1.888.VARIANS (888.827.4265); Europe +41 41 749 8844 and/or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 5823 | 5823 devices | Worldwide Distribution. | Terminated | July 24, 2012 | 319 | 5823 | Not Computer
Z-0179-2012 | 59918 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Quick Shield with Snappy Safety Holder, Greiner bio-one. |  |  | August 11, 2011 | 2011 | Greiner Bio-One North America, Inc. | Reports of the Vacuette QuickShield Safety Holders breaking during use due to a crack at the holder hub. | Nonconforming Material/Component | Greiner Bio-One sent a "QUICKSHIELD WITH SAFETY HOLDER" notification letter dated August 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to stop distributing and using the affected products. A Product Disposition Form was attached for customers to complete and return for replacements. Contact the firm at 888-266-3883 for questions concerning this recall. |  | 443950 | 443,950 pieces | Nationwide Distribution-USA (nationwide) including the states of CO, IL, KY, NY, PA and VA. | Terminated | March 20, 2012 | 222 | 443950 | Not_Computer
Z-0178-2012 | 60104 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LifeCell Strattice Reconstructive Tissue Matrix for Stoma Reinforcement |  |  | July 13, 2011 | 2011 | LifeCell Corporation | The use of Strattice for stoma reinforcement at the time of stoma creation is not within the product's cleared indications for use in the US market. | PMA | LifeCell began notifying their customers via phone on 7/13/2011. Recall Notice Letters and return response forms, dated July 12, 2011, were sent out on the same day. The letter identified the affected product and informed customers of the reason for the recall. Customers were asked to examine their inventory and quarantine the affected product. If customers have affected product, they were to contact Customer Solutions at 1-866-423-2433 to arrange for return of the product to LifeCell. Also, the attached Recall Notification form should have been completed and returned no later than August 17, 2011. Questions should be directed to Customer Solutions at 1-866-423-2433. |  | 111342 | 2651 sublots (of 42 lots) | Nationwide Distribution | Terminated | March 04, 2013 | 600 | 111342 | Not_Computer
Z-0180-2012 | 60048 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat cTnI cartridges |  |  | September 29, 2011 | 2011 | Abbott Point Of Care Inc. | i-Stat cTnI cartridges beginning with letter "T" may exhibit a higher than expected variability in reported results. This increased variability may result in falsely elevated or falsely depressed results. | Material/Component Contamination | APOC issued Urgent Recall Notices/Business Reply Forms dated September 2011 to consignees. The letter identified the affected product and explained the background of the problem. Customers are to discontinue use of the cartridges from the lots identified. All unused product will be credited and should be returned as instructed on the attached Business Reply Card. If the affected product was further distributed, a copy of the Recall Notice should be provided to those customers as well. APOC shipped replacement product in order to minimize any disruption. If replacement product has not been received then customers should contact Customer Service at 1-800-366-8020, option 5. Questions regarding this information should be directed to Abbott Point of Care Technical Support at 800-366-8020, Option 1 or customers' local APOC representative. |  | 1688755 | 1,688,755 cartridges | Worldwide Distribution | Terminated | March 11, 2013 | 529 | 1688755 | Not_Computer
Z-0172-2012 | 59829 | No | ISA | Therapeutic massager. | Physical Medicine | Physical Medicine | 510(K) Exempt | Shoulder Flex Massage Device |  |  | August 30, 2011 | 2011 | King International | The covered rotating mechanism can catch on user's clothing, jewelry (e.g. necklace), and hair. There have been reports of one death and a strangulation associated with using the device. | Component design/selection | An Urgent Product Recall letter was sent via regular mail to customers. The letter identified the affected product and described the reason for the recall. Customers were asked to stop using Shoulderflex massagers immediately and to safely dispose of them in the trash. The device components should be disposed of separately so that the device cannot be reassembled and used. This means that the power supply should be separated from the massager unit, and the massage fingers should also be removed and disposed of separately. Customers are to complete and return the enclosed the postage-paid card to King International in order to confirm receipt of the recall letter. If additional information is needed regarding this recall, customers are to contact King International LLC at 503-524-7046 or visit their website at www.shoulderflex.com or via mail to PO Box 2384, Beaverton, Oregon 97075.On 8/25/2011 FDA published a health alert on the FDA website. On 8/31/2011 the firm issued a nationwide press release. |  | 11934 | 11934 | Nationwide Distribution | Terminated | March 06, 2014 | 919 | 11934 | Not_Computer
Z-0182-2012 | 60036 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation-NM (EBW-NM) |  |  | September 15, 2011 | 2011 | Philips Medical Systems | A problem related to a software error has been detected in the Philips Multiple Gated Radionuclide Angiocardiography (MUGA) application. A false negative interpretation due to a falsely elevated ejection fraction or a false positive interpretation due to a false low ejection fraction could occur. | Software design | Philips sent all affected sites an Urgent Medical Device Correction Letter on September 27, 2011 for the MUGA issue. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy of the communication with the equipment instruction for use. Additionally, a field change order was released on September 26, 2011, customers will be receiving release notes as part of their correction informing them of the changes and corrections to the software. The new software will be installed on affected units by Philips Field Service Engineers.For customers in North America contact Customer Care Solutions Center at 1-800-722-9377. Select option 5. In all other countries contact the local Philips Healthcare office. |  | 1181 | 1181 systems | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | January 30, 2013 | 503 | 1181 | Software | Output/Calculation | Software update | Software Update
Z-0186-2012 | 60242 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Breeze Self-Adhesive Resin Cement, Part Number: N97A, |  |  | July 08, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | Pentron is voluntarily recalling the Breeze Self Adhesive Resin Cement because it was discovered that there may be a curing problem that may cause cemented crowns to fall off after use of the product. | Pending | Pentron Clinical sent an "Urgent: Medical Device Recall" letter dated July 18, 2011 with an attached Recall Return Form (via USPS First Clas mail) to all affected customers. The letter identified the affected product, explanation of the problem, and actions to be taken. Customers were instructed to discontinue use and return any affected product in their inventory. The letter asks that each customer complete and return the enclosed Recall Return Form via fax to 714-288-4609, in order to confirm receipt of the customer notification.Customers were instructed to contact Pentron Clinical Customer service at (800) 551-0283 (option 1) directly to handle the arrangements of a quick return and replacement. |  | 336 | 336 units | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Canada, Cypress, Greece, Germany, Italy, Jordon, Poland, Romania, Turkey and Vietnam. | Terminated | June 28, 2013 | 721 | 336 | Not_Computer
Z-0187-2012 | 59461 | No | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare 600 lb. Bariatric Bed, Head motor extension cable used on component numbers BAR5490 and BAR5490IVC. These components are part of the following Bed Model Number(s): BAR600, BARPKGCA, BAR600IVC, BARPKGIVC-1633, and BARPKG. |  |  | June 23, 2011 | 2011 | Invacare Corporation | Invacare Corporation decided to recall the products because risk analysis reveals that miskeying or disconnection of the head motor extension cable during bed set-up may result in product damage, or in rare instances, the potential for fire. | Labeling design | On 9/19/11, Invacare sent Recall Notificaiton Letters to their customers. They instructed their customers to quarantine the product, complete and return a response form. |  | 35747 | 35,747 | Distribution Nationwide and to United Arab Emirates, Dominican Republic, Egypt, Guam, Mexico, &amp; Virgin Islands. | Terminated | July 02, 2013 | 740 | 35747 | Not_Computer
Z-0184-2012 | 59919 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Greiner bio-one, 4ml K2E EDTA K2 Vacuette tube. Used for venous blood collection. Item number 454209. |  |  | August 12, 2011 | 2011 | Greiner Bio-One North America, Inc. | Vacuette 4ml EDTA DK2 tubes have been reported for short fills to no fills. | Process control | Consignees were notified by Greiner bio-one letter on August 11, 2011. They were instructed to: Stop distributing the material, complete the product Disposition form attachÂż and to fax it to Greiner bio-one and to destroy or return any of the affected products for replacement or credit. Distributors were asked to provide a list of their customers that have received the item, preferably in an Excel file, so that they could be notified of the recall. |  | 504000 | 504,000 pieces | Nationwide. | Terminated | March 12, 2012 | 213 | 504000 | Not_Computer
Z-0183-2012 | 60122 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | Pulstar(R) DVT PREVENTION SYSTEM, Anti-Embolism Wrap System, albahealth(R), A Subsidiary of Encompass Group, LLC. |  |  | October 12, 2011 | 2011 | Albahealth LLC | The wrap portion of the device was tearing at the seam. | Component change control | Albahealth, L.L.C. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 19, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory and to quarantine all affected product. A response form was enclosed for customers to complete and return. The firm provide information on how to return the affected product and receive replacement product. Questions concerning this recall are directed to call 865-354-5229 for assistance. |  | 2582 | 2,582 pairs | Nationwide Distribution- including the states of CA, CO, FL, GA, IN, KY, LA, NY, OH, PA, TX, and TN. | Terminated | March 30, 2012 | 170 | 2582 | Not_Computer
Z-0190-2012 | 60366 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Step 2, 3000-22000 GPD (Sodium Hydroxide 2.0N) cleaning solution. 4 oz plastic bottle. Lot # 11IC002, Serial #'s: 11IC00210 - 11IC002294. |  |  | October 24, 2011 | 2011 | Isopure Corp | The cleaning solution Step 2, 3000-22000 GPD labeled bottles (NaOH) were filled with Step 1, 3000-22000 GPD chemical (HCl). | Employee error | Isopure sent a Product Recall Notice dated October 24, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue the use of any Step 2 chemical from the affected product, place any remaining chemical into the original box, seal it with tape, and segregate the product to prevent it from being placed back into stock. The firm will pick up the product for return on a call tag. Customers would receive replacement product for all affected cases of the Step 2 chemical the firm received.For questions call 502-722-1000. |  | 48 | 91cs/12/4 oz plastic bottles | Nationwide Distribution including the following states: AL, AZ, CA, FL, GA, IN, KS, LA, MA, MO, MS, NC, NH, NJ, NM, NV, OH, OR, PA, SC, TN, TX, and VA | Terminated | February 24, 2012 | 123 | 48 | Not_Computer
Z-0188-2012 | 60370 | No | LKB |  |  | General Hospital | Enforcement Discretion | PDI Alcohol Prep Pads packaged in the INRatio kit. |  |  | October 11, 2011 | 2011 | Alere San Diego | Alere has initiated a recall of the Professional Disposables International (PDI) alcohol prep products due to the potential presence of low levels of the microorganism Bacillus cereus. | Material/Component Contamination | Alere sent a recall notification letter dated October 6, 2011, and Verification Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use an alternate prep pad that is not involved with the PDI recall, or use a sterile gauze pad with isopropyl alcohol. If customers purchase alcohol pads from a store or pharmacy, customers can confirm with their pharmacist that the pads they are purchasing are not associated with the PDI recall. INRatio/INRatio2 kits at Alere are currently being reworked to include a sterile PDI alcohol pad.Customers were instructed to complete and FAX the enclosed Verification Form within 10 days to confirm their receipt of the notice and to indicate the number of discarded/required replacement alcohol prep pads.Customers with questions about INRatioÂż/ INRatioÂż 2 were instructed to consult their healthcare provider, or call Alere Technical Service at 877-866-5313 (24 hours a day, 7 days a week). |  | 14248 | Estimated 14,248 Total units | Nationwide Distribution | Terminated | February 29, 2012 | 141 | 14248 | Not_Computer
Z-0189-2012 | 60059 | No | CDQ | Urea nitrogen test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Urea Nitrogen List Number 7D75-21 and 7D75-31 used for the Quantitation of urea nitrogen in human serum, plasma, and urine. |  |  | October 03, 2011 | 2011 | Abbott Laboratories, Inc | The reagents are being recalled due to particulate matter in some of the reagent cartridges. The particulate matter has been identified as Aureobasidium fungus species. | Process control | The firm decided to recall and they sent a Product Recall Letter to their customers on 10/03/2011. Questions regarding the letter should be directed to Customer Service at 1-877-4ABBOTT for US customers, and customers outside of the US should contact their local area Customer Service. Customers should determine if you have any remaining inventory. Please discontinue and destroy any remaining inventory of these lots. Order alternative lots of Clinical Chemistry Urea Nitrogen. Complete and return the enclosed Customer Reply form. |  | 5704 | 5704 units of Urea Nitrogen | Worldwide Distribution. | Terminated | May 11, 2012 | 221 | 5704 | Not_Computer
Z-0193-2012 | 60097 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Restoration ADM System X3 Acetabular Insert Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; Stryker France ZAC Satolos Green Pusignan Av de Satolas Green 69881 Mey Sieu Cedex France |  |  | June 03, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become ware that there is the potential for interpretation of the product labeling which may lead to an incorrect implant being used. | Labeling False and Misleading | The firm, Stryker, sent an "URGENT PRODUCT CORRECTION" letter and Product Correction Bulletin dated June 3, 2011 via Fed Ex to its customers. The letter and bulletin describes the issue, potential hazards, risk mitigation and product correction instructions. The customers were instructed to complete and return the enclosed PRODUCT CORRECTION ACKNOWLEDGMENT FORM within 5 days via fax to: 201-831-6069.If you have any questions, feel free to call (201) 972-2100 or Hip Marketing Team at 201-831-5280. |  | 1677 | 1677 units | Nationwide distribution. | Terminated | November 13, 2013 | 894 | 16891 | Not_Computer
Z-0198-2012 | 60145 | No | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Right IBO Blade , for use with Stryker TPS Reciprocating Saw, REF 5100-37-901S2, Stryker Instruments, Kalamazoo, MI. |  |  | October 18, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | There is a potential for these blades to break where the blade is welded on to the arbor. This hazard can result in the following harms: tissue damage due to the broken blade being aspirated or swallowed, additional surgery to remove a broken blade, infection due to an unretrieved device component, complications associated with increased time under anesthesia to retrieve a broken blade, osteomyeli | Device Design | The firm, Stryker, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated October 18, 2011 to Risk Managers, Office Managers, and Oral Surgeons. The letter described the products under recall, the problem and a series of actions to be taken. Customers were asked to immediately locate recalled products and quarantine. Products were to be shipped back to Stryker. The letter included a Business Reply Form to be faxed back to Stryker Instruments Regulatory Department, 866-521-2762. The letter instructs to forward this notification to all affected locations if the recalled products were further distributed. Customers with questions can call Stryker at 1-800-253-3210. |  | 113 | 113 | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Canada, China, Germany, Greece, Hong Kong, India, Jordan, Miramar, Saudi Arabia, Singapore, Spain, Switzerland, United Kingdom. | Terminated | July 12, 2013 | 633 | 12631 | Not_Computer
Z-0191-2012 | 60174 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia Apheresis System, REF: 61000, Caridian BCT, Lakewood, CO 80215. |  |  | September 21, 2011 | 2011 | Caridian BCT, Incorporated | Possible higher than expected platelet loss following therapeutic plasma exchange or plasmapheresis. | Software design | CaridianBCT sent a "Safety Alert" letter dated September 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm will be sending a recall letter to consignees in November 2011, stating they are voluntarily recalling the software and will provide an upgrade once the software has been approved.If you have any further questions please contact your CaridianBCT Representative, the CaridianBCT Support Center at 1-877-3-FYU-BCT ( US Toll free 1-877-299-4228 ) or (303) 231-HELP ) ( 1 (303) 231-4357, or your local CaridianBCT Customer Service office. |  | 852 | 852 units (US: 333; Foreign: 519) | Worldwide Distribution -- USA ( nationwide ) and the countries of Australia, Belgium, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Pakistan, Republic of Korea, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand and Vietnam. | Terminated | February 01, 2013 | 499 | 852 | Software |  | Software update | Software Update
Z-0185-2012 | 60202 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | The LIFEPAK 20 defibrillator/monitor is an AC powered external defibrillator with battery backup used by healthcare providers in hospital and clinic settings. |  |  | December 30, 2008 | 2008 | Physio Control, Inc. | When users do not disconnect the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform. | Employee error | Physio Control sent a " Urgent Medial Device Safety Alert" letter dated January 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The letter instructed the customers to follow the Operating Instructions for AED and Manual Defibrillation mode (Section 4-Therapy, pages 58 and 61) that instructs users to "Connect therapy electrodes to the therapy cable and confirm cable connection to the defibrillator" to prevent users from mistakenly leaving the test plug and reconnect the QUIK-COMBO electrodes to the therapy cable, the test plug signal can be misinterpreted as a patient waveform.Recommendations to users who use a Therapy Cable:-Disconnect the QUIK-COMBO test plug prior to patient use-Confirm therapy electrodes - to - therapy cable connection-Confirm therapy cable - to - defibrillator connection-Confirm therapy electrodes - to - patient connectionAdding 2005 American Heart Association software that includes audible and visual prompts to LIFEPAK 20 defibrillators manufactured prior to January 2007 is optional and not required. If the customers would like to have this optional kit installed in your defibrillators, please contact Physio Control, Inc.Customers can call Technical Support at 1-800-442-1142 for any questions about this recall or or visit our website at www.physio-controlnotices.com/LP20testplug. |  | 23467 | 23,467 units worldwide | Worldwide Distribution--USA ( nationwide ) and the countries of Bahrain, Lebanon , Cyprus, Netherlands, Egypt, Pakistan, Germany, Saudi Arabia, India, South Africa, Iran, Syria, Israel, Turkey, Jordan, UAE,Kuwait,Yemen, Asia Pacific, Australia, New Caledonia, Guam, New Zealand, Hong Kong, Latin America, Anguilla, Jamaica, Argentina, Mexico, Bahamas, Nicaragua, Brazil, Panama, Chile, Peru, Colombia, Trinidad, Costa Rica,Uruguay, El Salvador, Venezuela, Guatemala, Virgin Islands and Canada. | Terminated | November 21, 2011 | 1056 | 23467 | Other |  |  | Other
Z-0201-2012 | 60118 | No | DYW | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Spirit MB Brackets, part #494-1211. |  |  | May 18, 2011 | 2011 | Ormco Corporation | Ormco Corporation is voluntarily recalling three lots of Spirit MB Brackets due to a manufacturing error. | Device Design | Ormco Corporation sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 13, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product. If any of the affected product is found, customers are directed to contact Ormco Customer Care at 1-800-854-1741 to receive a RMA number. The RMA number will allow a quick return and replacement or credit. Additionally, a Recall Return/ Acknowledgement Form was attached for customers to complete and return via fax to 909-962-5605. |  | 135 | 135 packs | Worldwide Distribution-USA (nationwide) including the states of FL, MN, TX, and VA, and the countries of Australia, France, Germany, Italy, Japan, Mexico, Morocco, Netherlands, New Zealand, South Africa, Turkey, and Ukraine. | Terminated | April 24, 2012 | 342 | 145 | Not_Computer
Z-0209-2012 | 60420 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. |  |  | July 13, 2011 | 2011 | SCC Soft Computer | Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 3.17.5 to 3.17.9, 4.2.1 to 4.2.3, and 4.3.0 to 4.3.0.4. SoftPath architecture does not correctly handle AP macros and data fields inside report sections that are stored in the database. Improper information can be displayed in all places where the diagnosis from the database is read and displayed. | Software design | The firm, SCC Soft Computer, sent a "CORRECTION COMMUNICATION" notice on July 13, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to restrict security access to the template set-up files and to acknowledge receipt of the notice; grant SCC personnel permission to check their report templates and correct any that are affected; indicate if they wish to schedule the utility to be run and if they would like to have the SoftWeb parameter PathPDFLinkOnly set to "Y". SCC will review the templates on the customers systems; discuss any irregularities requiring change and move or remove AP macros and data elements as needed.If you have any questions, contact the Director of Quality Management at (727)789-0100. |  | 142 | 142 | Worldwide distribution: USA (nationwide) including states of: AR, AL, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MN, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI, WV including Puerto Rico; country of Canada. | Terminated | July 17, 2012 | 370 | 142 | Software |  |  | Other
Z-0207-2012 | 59990 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | LifeStar ACT III Platinum Sensor Ambulatory Cardiac Telemetry; an automatically activated system with a high performance triple channel EKG, office hookup and auto send/auto detect that requires no patient intervention to either capture or transmit an arrhythmia when it occurs; Manufactured by Card Guard, 2 Pekeris St, Rehovot, 76100, Israel; Model FG-00084 |  |  | September 26, 2011 | 2011 | Lifewatch Services Inc | LifeWatch received complaints of minor shocks and burns during use of the LifeStar ACT systems with shorts in the leads. | Nonconforming Material/Component | The firm, LifeWatch, sent an " Urgent Device Recall" letter dated September 27, 2011 to the affected customers. The letter describes the product, problem and action to be taken. The customers were instructed to immediately refrain from providing these devices to their customers; stop using the affected sensors, return their entire kit in the pre-paid mailer, and use the new device kit included with the recall letter. In addition, the letter informed the users that the replacement device has upgraded software and firmware that allows for detection of shorts in the cable, and if detected, cuts off current flow to the patient to prevent minor shocks and burns.If you have any questions, please call (847) 813-4258. |  | 442 | 442 sensors | Nationwide distribution. | Terminated | May 15, 2012 | 232 | 1094 | Computer
Z-0202-2012 | 59960 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed paradigm mio Infusion Sets MMT-921 |  |  | May 16, 2011 | 2011 | Medtronic Minimed | The consignee received four (4) boxes of mio infusion sets (MMT-921) that were shipped to them in error. These infusion sets were shipped in a package that included a MiniMed Paradigm REAL-TIME Insulin Pump among other items.The mio infusion sets were not authorized to be shipped to the consignee based on Medtronics internal product handling procedures and they may not function properly.T | Employee error | The firm, Medtronic, sent an "IMPORTANT RECALL INFORMATION" letter dated May 16, 2011 to consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue use right away; use the replacement mio infusion sets provided to them in this package; return all unused Lot 3700255 mio infusion sets using the emailed pre paid return label, and drop the package off at any UPS location or call UPS at 800.742.5877 to schedule a pick up."Please note that only the Lot 3700255 mio infusion sets were shipped in error. The other products included in the package were not shipped in error and do not need to be returned."If you have any additional questions relating to this notification or any questions in general call 800.999.9859 ext.63704. |  | 4 | 4 boxes | Nationwide distribution: AL only. | Terminated | December 14, 2011 | 212 | 4 | Not_Computer
Z-0205-2012 | 60392 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SUPER TORQUE MB Angiographic Catheter product is labeled in part: "***5F (1.65 mm)***110 cm .038" (0.97 mm)***MULTIPURPOSE A SUPER TORQUE MB***Max. 1200 psi (8274 kPa)***Markers***OPEN end***Sideholes***SPECIAL***Angiographic Catheter*** REF Cat No.***Lot***Use By***Sterile EO***Assembled in Mexico Cordis Corporation, 14201 NW 60th Ave. Miami Lakes, Florida 33014, USA***EC***REP*** Cordis Cashel, Cahir Road Cashel, Co Tipperary, Ireland***Cordis***a Johnson&Johnson; company***" |  |  | November 21, 2011 | 2011 | Cordis Corporation | Cordis Corporation is issuing a Correction for their SUPER TORQUE MB Angiographic Catheter after receiving a report of marker bands dislodged during a procedure. Model Numbers 532598A, 532598B, 532598C, 532598D, SRD5724MB and SRD5727MB are affected by the Cordis Correction. | Error in labeling | Cordis Corporation sent a Correction letter dated November 21, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers are instructed to read the "Description" and "Recommendations" sections, sign and return an enclosed Acknowledgement Form, and to maintain awareness of the communication until the information has been incorporated into the Cordis Super Torque MB Catheter labeling. Contact Cordis at 800-781-0282 for questions regarding this notice. |  | 310 | 310 units | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | July 23, 2013 | 610 | 119000 | Not_Computer
Z-0203-2012 | 60124 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | EHEC Kit Lot# 608096 Containing: Sample Diluent 18.8 ml Lot# 8602EXP.1-BUF WASH 20X II 100.0ml.Lot# 2607EXP.2-EHEC Detection Antibody 10.0ml Lot# 8603EXP.3-EHEC-Enzyme Conjugate 10.0ml Lot# 8604EXP.4-Substrate II 12.5ml Lot# 1601EXP.5-Stop-Solution II 10.0 ml Lot# 1601EXP.6-EHEC Positive Control 3.4 ml Lot# 8601EXP.7-EHEC Negative Control 3.4 ml Lot# 1603EXP.8-96 transfer pipettes-1 Microwell holder-3 Microwell strip sealers |  |  | October 10, 2011 | 2011 | Meridian Bioscience Inc | Internal testing has shown that the indicated kits produce false negative results when tested with some direct stool specimens. These false negative reactions are not observed for all devices tested, however, the rate of occurrence will likely impact clinical sensitivity as defined by product registration and associated instructions for use. | Nonconforming Material/Component | Meridian Bioscience, Inc. sent an Urgent Medical Device Recall letter dated October 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop use of the affected product, evaluate testing performed with the affected product., and segregate and destroy remaining inventory.Customers were asked to contact Meridian's Technical Service Department by phone or facsimile as necessary. For replacement product phone 800-343-3858, outside North America 513-271-3700, Fax 513-272-5432.Customers were also asked to complete and return the attached Customer Response Form. In the event that they have no remaining inventory, customers were instructed to indicate such on the Customer Response Form and return to Meridian Technical Service. For questions phone 800-343-3858, outside North America 513-271-3700. |  | 3236 | 3236 Kits | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, NC, NH, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI and WV and the countries of AUSTRALIA, CANADA, CHILE, COSTA RICA, ITALY, JAPAN and NEW ZEALAND. | Terminated | February 19, 2012 | 132 | 3236 | Not_Computer
Z-0229-2012 | 57585 | No | OJV | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Angiotech***Custom Bone Marrow Tray for University of CA, Davis*** 1 - 10cc Syringe L/L; 1 - 20cc Syringe L/L; 1 - 19ga x 1-1/2" Filter Needle; 1 - 21ga x 1-1/2" Needle; 1 - 25ga x 5/8" Needle; 1 - 8ga x 4" Bone Marrow Needed "J" Type; 1 15ga x 4" Bone Marrow Needle "I" Type; 1 - T- Lok Extraction Cannula; 1 - 10% Providone Iodine Swab Sticks (3/Pkg.); 1 - Probe Guide; 1 - Obturator; 2 - 1% Lidocaine (5mL); 1 - #11 Mini Scalpel; 2 - 4" x 4" Sponges in stacks of 5; 1 - 2" x 3" ADH Dressing; 1 - Towel; 1 - Fenestrated Drape; 1 - Male Luer Cap; 1 - CSR Wrap***Catalog Number: CB0140TL***Sterile - For single patient use. STERILE EO. Caution: Federal (USA) Law restricts this device to sale by or on the order of a physician |  |  | December 22, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc. dba Angiotech, is recalling HSG Procedure Trays and Bone Biopsy Trays distributed between December 2007 and December 14, 2010 for package integrity. | Packaging | The firm, Angiotech, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 22, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review their current inventory, segregate the affected lot numbers and complete and return the attached Customer Acknowledgement Form via fax to our Quality Assurance department at 1.352.338.0662 or 1.800.333.0440 to arrange for return of the identified product.If you have any questions concerning this notification, call (352) 338.0440 ext. 355 or ext 358. |  | 225 | 225 boxes | Worldwide distribution: US (nationwide) and countries of: Canada, Columbia, Costa Rica, Japan, Mexico, Peru and Saudi Arabia. | Terminated | November 30, 2011 | 343 | 32491 | Not_Computer
Z-0210-2012 | 59818 | No | LYJ |  |  | Neurology | PMA | VNS Therapy Aspire HC Generator and VNS, Model 105 |  |  | August 12, 2011 | 2011 | Cyberonics, Inc | The devices are being recalled because the output current delivered to the vagus nerve is less than the design intent and there is a potential charge imbalance at the lead cathode and generator-can during stimulation. | Nonconforming Material/Component | Cyberonics sent a Field Safety Alert on August 15, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm decided to recall and initiated contact with implanting facilities in the United States on 08/15/2011. A Field Safety Alert was delivered on 09/19/2011 to implanting facilities. Customers were instructed not to prescribe AspireHC pulse generators for implant until Cyberonics resumes shipment of AspireHC pulse generators that have the issue remedied. Cyberonics recommends that physicians continue their current practice of titrating programming parameters for patients implanted with an AspireHC pulse generator to achieve appropriate clinical outcomes, paying particular attention to report of:Lack or loss of efficacyPain, inflammation, swelling/edema in the neck area; and/orPatient perception of stimulation or muscle twitching in the area surrounding the implant. Cyberonics recommended routine systems diagnostic testing to aid in device and patient monitoring. Physicians may contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com.For questions regarding this recall call 281-228-7323. |  | 340 | 340 units | Worldwide Distribution - USA including TN, TX, IL, OH, FL, CA, CO, OR, AZ, NY, PA, MO, NJ, WI, OK, MI, ID, KY, IA, MA, NE, GA, LA, NY, and Norway. | Terminated | August 20, 2012 | 374 | 340 | Not_Computer
Z-0213-2012 | 60101 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | DICOM Option for the Norland Illuminatus Software Revs 4.2.0 though 4.3.1. Models Norland XR-600, XR-800, XR-46, EXCELL, XR-36 and Eclipse DXA Bone Densitometers. |  |  | September 23, 2011 | 2011 | Cooper Surgical, Inc. | The DICOM Software, Norland Illuminatus DXA Software with the DICOM Option (Part Number 434A133) version 4.2.0 through 4.3.1, which is installed on the XR-600, XR-800, XR-46 and EXCELL DXA Bone Densitometers, could lead to improper patient information getting into a previous patient's record. | Software design | The firm, CooperSurgical, sent a "Notice of RECALL-URGENT" letter dated September 23, 1011 to its customers. The letter described the product, "Problem/Situation", and "When can this problem occur?". CooperSurgical will send a new revision of Illuminatus software to the customers once available.If you have any concerns about your existing database of patients, please contact our Norland Technical Support team at (203) 601-5200 Extension 9888 or 9889. |  | 184 | 184 | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, FL, GA, IN, KS, KY, LA, MD, MA, MI, MO, NJ, NY, NC, OK, PA, TN, TX, UT, and WI ; and countries of: ARGENTINA, AUSTRALIA, BELGIUM SPAIN, CANADA, CHINA, DENMARK, EMIRATES, ENGLAND, FRANCE, GERMANY, GREECE, IRELAND, MALTA, POLAND, SAUDI ARABIA, SOUTH KOREA, and TAIWAN. | Terminated | April 09, 2013 | 564 | 184 | Software | Output/Calculation | Software update | Software Update
Z-0212-2012 | 60165 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) |  |  | June 06, 2011 | 2011 | AGFA Corp. | Incorrect entries in mapping files could lead to inaccurate measurement display in adult and pediatric echo reports when using the IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR). | Pending | The firm, AGFA HealthCare" sent an 'URGENT FIELD SAFETY NOTICE" dated May 27, 2011 via FED-EX by June 6, 2011 to all customers using the affected product. The letter describes the product, problem and actions to be taken. The customers were instructed to distribute this information within their facility to all those who need to be aware and to complete and return the attached FEEDBACK FORM via fax to 864-421-1668 or email to debbie.norris@agfa.com. AGFA will provide software corrections to all affected sites and will contact the customers to schedule the analysis and provide any corrections if required.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. |  | 280 | 280 | Worldwide distribution: AK, AZ, CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WV; and country of: Canada. | Terminated | May 08, 2014 | 1067 | 280 | Software | Output/Calculation | Software update | Software Update
Z-0211-2012 | 60154 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Results Management (RM) Software Versions RM 2.04.37.04 to RM 7.8 SU2 |  |  | June 06, 2011 | 2011 | AGFA Corp. | Improper merging of IMPAX Cardiovascular Results Management (RM) reports causes the signed report to become unavailable. | Software design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated May 27, 2011, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to carry out a merge post finalization of an RM reports. If it is necessary to modify a finalized report, AGFA recommends that they make the modification through one of the following actions:Amend the report through the RM interface, orCall AGFA Service to change the status of the report to "preliminary" and then proceed with the finalization of the report.AGFA issued an Acknowledgment, via FAX- Back that the information was received and understood to affected customers. For any questions customers were instructed to contact their AGFA Healthcare Service at 877-777-2432. |  | 323 | 323 | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 02, 2013 | 757 | 323 | Not_Computer
Z-0258-2012 | 60021 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Precedence System; |  |  | September 27, 2011 | 2011 | Philips Medical Systems | A mechanical malfunction of the button on the hand controller on the Philips Precedence and SKYLight SPECT systems may get stuck and cause unintended motion of a detector. | Device Design | Philips Medical Systems sent a "Customer Information" letter dated September 27, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter lists additional precautions the customers should use while operating the device. Philips will notify customers for replacing the affected product through an Action for Performance-Proactive Field Change Order to resolve the issue. Contact Philips Customer Care Solutions at 1-800-722-9377, option 5, for questions concerning this notice. |  | 658 | 658 - all units in recall | Worldwide Distribution-USA (nationwide) including Puerto Rico. | Terminated | December 21, 2012 | 451 | 658 | Not_Computer
Z-0260-2012 | 60173 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is a synthetic absorbable suture with needle and labeled as sterile. The product is sealed inside an inner foil pouch with an outside Tyvek pouch. The outer sterile barrier seal is intact. The product is boxed and shrink wrapped for protection against contamination. |  |  | January 02, 2009 | 2009 | C P Medical Inc. | The seal integrity of the outer product pouch may be compromised due to partial delamination of the two layers that make up part of the pouch. | Nonconforming Material/Component | CP Medical, Inc. sent a "RECALL NOTICE" letter dated January 5, 2009 to all affected customers.The letter describes the product, problem, and the actions to be taken by the customers. Customers are instructed to identify affected product and to return the product to the firm. Customers should call 1-503-232-1555 for a return goods authorization number and shipping account number. Contact your customer service representative at 1-800-950-2763 for questions concerning this recall. |  | 13632 | 13632 sutures | Nationwide Distribution | Terminated | December 13, 2011 | 1075 | 13632 | Not_Computer
Z-0261-2012 | 60405 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. |  |  | August 08, 2011 | 2011 | SCC Soft Computer | On 08/08/2011 SCC Soft Computer, Clearwater, FL initiated a correction on following SoftLab used with: SA BASE 3.1.4.x, 3.1.6.x and SA LabMic 4.0.0.x, 4.0.1.x, 4.0.2.x, 4.0.3.x, 4.0.4.x. A client reported a flagging error where the result was flagged as critical in the Instrument Menu but the printed report did not show the corresponding flag. All affected clients were notified of the issue on 0 | Software design | SCC Soft Computer notified all affected customers on August 8, 2011, of the recall via the SCC Soft Computer proprietary Task management system to manage client communication. Clients use this tool to log in requests for services and receive and provide updates to any issue. Each client user has a unique password and every action is date and time-stamped. Each Task has a unique identifier and a full audit trail.SCC recommended to the customer until Live on the correction avoid making edits in the instrument menu.For further questions please call (727) 789-0100. |  | 255 | 255 | Worldwide Distribution - USA ( nationwide) and Canada. | Terminated | November 22, 2013 | 837 | 255 | Not_Computer
Z-0265-2012 | 60301 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Endoscopy 3.55 MM Super 90 S-SERFAS Energy Probe; electrosurgical device. Model number: 279-351-300; |  |  | October 24, 2011 | 2011 | Stryker Endoscopy | An increased number of complaints for tip breaks were observed for the 3.5mm Super 90-S SERFAS Energy Suction Probes manufactured from June to August 2011. | Component design/selection | Stryker Endoscopy Urgent Medical Device Removal notifications datyed October 25, 2011, were sent to all consignees via traceable mail. Sales representatives and regional managers were notified by e-mail. Product is to be returned to Stryker Endoscopy using the pre-paid shipping label. Replacement product may be obtained by calling Stryker at 1-800-624-4422. The enclosed Acknowledgement receipt form is to be completed and returned. Questions or comments may be directed to Stryker at 1-800-624-4422 or e-mail at 90S@stryker.com. |  | 6307 | 6307 | Worldwide distribution including USA, Australia, Canada, China, Greece, Hong Kong, India, Italy, Japan, France, Netherlands, Poland, South Africa, Spain, Sweden, Switzerland and the United Kingdom | Terminated | February 29, 2012 | 128 | 6307 | Not_Computer
Z-0262-2012 | 60205 | No | KOY | Surgical skin degreaser or adhesive tape solvent. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PDI Electrode Prep Pads, Skin Prep B339 are non-sterile alcohol prep pads. They are included in the Spacelabs Healthcare Hook-Up kits. Each kit contains 2-4 pieces of PDI Electrode Prep Pads . |  |  | October 12, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | The non-sterile alcohol prep pads (B339) are included in Spacelabs Healthcare hook-up kits may have the presence of a bacterium, Bacillus cereus. These alcohol prep pads were manufactured and recalled by Professional Disposables International (PDI). | Material/Component Contamination | SpaceLabs Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated October 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using the affected product and dispose of them per their hospital or clinic protocol. For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 2for Technical Support. |  | 25000 | 25,000 pieces of PDI alcohol prep pads-15 different Spacelabs hook-up kits | Worldwide Distribution - USA (natiowide) and the countries of South Africa, Turkey, Bangladesh, Dominican Republic, Peru, Belgium, Ecuador, Romania, Austria, Honduras, Guatemala, Denmark, Republic of Korea, Argentina, Thailand, Italy, Austria, Chile, Bolivia, Mexico, Sweden, Netherlands, Switzerland, India, Serbia, Australia, Morocco, Turkey, ISpain, Czech Republic, Slovakia, Poland, China, Columbia, Portugal, Canada, Germany, United Kingdom, Singapore, France, Russian Federation, Bahrain, and Kuwait | Terminated | February 23, 2012 | 134 | 25000 | Not_Computer
Z-0264-2012 | 60365 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Super Sheath and Super Sheath R/O Introducer Sheaths, Sterile EO, Manufactured by Togo Medikit Co., Ltd. 17148-6 Aza Kamekawa, Oaza Hichiya, Hyuga City, Miyazaki Prefecture 883-0062, Japan, Distributed by Boston scientific Corporation One Boston Scientific Place, Natick, MA 01760-1537 USA. |  |  | November 03, 2011 | 2011 | Boston Scientific Corporation | Boston Scientific is initiating a medical device recall removal regarding nine lots/batches of Super Sheath and Super Sheath R/O Introducer Sheaths. Boston Scientific was informed through product complaints that the 0.035" ID dilator may be labeled incorrectly as a 0.038" ID dilator. | Packaging process control | Boston Scientific sent an URGENT MEDICAL DEVICE RECALL REMOVAL -IMMEDIATE ACTION REQUIRED letter dated November 21, 2011. The letter described the product and the problem. Customers were instructed to segregate the affected product immediately and return it to Boston Scientific. Distributors were advised that the recall removal depth is to the hospital level and the recall removal notification should be forwarded to their customers.For questions regarding this recall call 763-494-1133. |  | 2660 | 2660 | Worldwide Distribution - USA including AZ, AR, CA, CO, FL, GA, IL, KY, LA, MD, MA, MN, MO, NE, NJ, NM, NY, OH, OK PA, RI, TN, TX, UT, VA, WY, WI, and WV and the countries of BELGIUM, BRAZIL, CZECH REPUBLIC, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, ITALY, NETHERLANDS, RUSSIAN FEDERATION, SOUTH KOREA, SPAIN, SWEDEN, and UKRAINE. | Terminated | January 15, 2013 | 439 | 2660 | Not_Computer
Z-0270-2012 | 60278 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | "***UPLIFT TECHNOLOGIES INC.***Bringing quality to life.***Model No.***CA200***Description***Lifting commode chair***Weight Range***80-300+ lb-136+ kg***" |  |  | June 27, 2011 | 2011 | Uplift Technologies Inc. | On June 26, 2011 Uplift Technologies Inc. Nova Scotia, Canada initiated a recall of the Uplift Commode Assist Model No. CA200, Lot No. 1C0101. The plastic seat does not meet manufacturing specifications. The seat may not fit the unit properly and may interfere with the unit's normal locking mechanism. | Nonconforming Material/Component | Uplift Technologies, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 27, 2011 to all affected customers The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory and to quarantine the affected product. Customers are to call Uplift's Customer Service Department at 1-800-387-0896 to request a Return Materials Authorization number in order to return the product.. |  | 219 | 219 chairs | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 339 | 219 | Not_Computer
Z-0266-2012 | 60005 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | Rapid Strand Rx; the RAPID Strand Rx kit consists of custom loaded absorbable seeding spacers and I-125 radionuclide brachytherapy sources (seeds) spaced at prescribed distance and configuration within a sleeve (tube) made of absorbable suture material, stiffened, loaded into prostate seeding needles, packaged and then sterilized by Gamma sterilization method, accompanied by a calibration strand, a sealed source calibration certificate (100% assay report), a radiograph of the loaded needles in the needle-holding card, and a source certification technical data sheet.; Products Provided by: Oncura, 3350 North Ridge Avenue, Arlington Heights, IL 60004 USA, Prescription Filled by: AnazaoHealth, 5710 Hoover Blvd., Tampa, FL 33634 USA; product identifier: I125RSRX, model 7000. |  |  | September 20, 2011 | 2011 | Medi-Physics Inc. dba GE Healthcare | There is a potential brachytherapy over or under dose with the RAPID Strand Rx kit that may impact patient therapy. | Process control | GE Healthcare/Oncura sent a "Urgent Medical Device Correction" letter dated September 20, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were advised to assess the pre and post plan dosimetry for their patient, taking into account the possible variation of the seed strength. GE Healthcare/Oncura will contact the customers in the near future to confirm their receipt of the letter and to discuss any questions they may have.Please call if you have any further questions (847) 398-8400 and 011-44-1494-737251 (Amersham, United Kingdom, GMT) |  | 23 | 23 orders | Worldwide Distribution -- USA (nationwide) including the States of Arizona, Florida, Louisiana, Mississippi, New Jersey, New York, Pennsylvania, Puerto Rico, Virginia and Washington, and the countries of Canada and the United Kingdom | Terminated | January 20, 2012 | 122 | 23 | Not_Computer
Z-0271-2012 | 60374 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | TUM-E-VAC REF 2075, Rx ONLY, NON STERILE --- Ethox International, Buffalo, NY 14204, USA --- Intended use: gastric lavage |  |  | March 25, 2011 | 2011 | Ethox International, Inc. | Ethox Tum-E-Vac Gastric Lavage Kit #2075, Lot Number 031123255 was labeled with expiration date symbol and date of 2010/10 on the unit product label. The carton was labeled with a manufacturing date symbol and date of 2010/10. The Ethox Tum-E-Vac Gastric Lavage Kit #2075 does not have an expiration date. | Employee error | Ethox International, Inc. sent an "URGENT DEVICE RECALL" letter dated March 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine their stock immediately to determine if they have the affected product. Customers are to return the affected product to the firm by using UPS account 128959. Customers will be reimbursed for the returned goods. Additionally, a response form was enclosed for customers to complete and return. Contact the firm at (716) 842-4000 for questions regarding this recall. |  | 130 | 130 units | Nationwide Distribution-including the states of California, Connecticut, Illinois, Ohio, and Virginia. | Terminated | November 28, 2011 | 248 | 130 | Not_Computer
Z-0263-2012 | 60155 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) - Physiological Patient Monitoring Solution |  |  | October 18, 2011 | 2011 | Draeger Medical, Inc. | During product training on the Infinity(R) Acute Care System (IACS) Monitoring Solution (M540) it was discovered that the weight-based drug dosage calculation may indicate incorrect recommended values. In addition, a misalignment of ECG and ART waveforms delivered from an Infinity(RO Acute Care System (IACS) Monitoring Solution (M540) was observed on the Infinity Central Station. | Software design | The firm initiated their recall of this device on October 17, 2011 by issuing a letter to the consignee instructing customers to follow the Instructions for Use and use the Infinity Central Station only for remote assessment of a patient's status. They recommend that you do not consider the waveforms on Infinity Central Station for primary diagnosis purposes. In case of a high priority alarm situation, "Check the patient and treat if necessary".If the misalignment issue is observed, it can be cleared by undocking and redocking the M540 or by changing the order of waveforms displayed on the Cockpit screen (or any other change to the auto display screen layout). |  | 476 | 476 | Nationwide distribution - IL | Terminated | April 01, 2013 | 531 | 476 | Other | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0269-2012 | 60316 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Wolf Medical Supply Inc., WOLF-PAK REDI-FLO ELASTOMERIC INFUSION PUMP, FILL VOLUME 270 mL, FLOW RATE 2mL/Hr. |  |  | October 20, 2011 | 2011 | Wolf Medical Supply Inc. | Redi-Flo Elastomeric Infusion Pumps may have a higher flow rate than specified (more than +15%). | Process control | On 10/20/2011 Wolf Medical Supply Inc., customers were notified by telephone and email. Product was distributed to end users. |  | 24 | 24 cases (576 units) | Nationwide distribution, including Puerto Rico. | Terminated | April 25, 2012 | 188 | 40 | Not_Computer
Z-0272-2012 | 60164 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CardioVascular (CV) Admin Tool |  |  | June 06, 2011 | 2011 | AGFA Corp. | Potential for misattribution of data and for a report to become unavailable when the user requests for it to be moved to another study that already has a report associated with it. | Software design | An 'URGENT FIELD SAFETY NOTICE" was sent via FED-EX by June 6, 2011 to all the sites using the affected product. The letter described the issue and mitigation. The letter recommends that users not move reports to studies where a report already exists. If a report is being moved because the existing report is irrelevant to the study, then the existing report should be deleted prior to moving the new report to that study. If the report is being moved because the user wants to combine the content of the reports, then the "merge" function should be utilized. The letter also states that a software enhancement is currently available. Acknowledgment via FAX-Back was requested from the sites. Questions regarding the issue should be directed to customers' local Agfa HealthCare Service at 877-777-2432. The firm has identified 185 additional sites and an "URGENT FIELD SAFETY NOTICE" was sent on October 12, 2012. |  | 323 | 323 units | Worldwide Distribution -- USA, Virgin Islands and Canada | Terminated | June 26, 2012 | 386 | 323 | Software | Output/Calculation | Software update | Software Update
Z-0273-2012 | 60123 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | IBC INTERNATIONAL BIOPHYSICS CORP. 2101-2 EAST ST. ELMO, SUITE 275, AUSTIN, TX 78744 SUCTION HANDLE (STERILE) PART NUMBER: 1990S LOT NUMBER 041811-2065 EXPIRATION DATE: 18/APR/2015 |  |  | October 05, 2011 | 2011 | International Biophysics Corp | Potential for the sterility to be compromised. | Packaging | International Biophysics Corp. sent a" Medical Device Recall Notice" dated October 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Use of the product in the referenced lot number should cease immediately. Customers were asked to follow the instructions in the letter to return the affected product.If you have any further questions please call (512) 814-0070. |  | 1200 | 1200 units | Nationwide Distribution including the states of AZ, CA, CO, DE, MS, OR, PA, TN, and TX | Terminated | March 23, 2012 | 170 | 1200 | Not_Computer
Z-0278-2012 | 60430 | No | LNX |  |  | General Hospital | 510(k) | SCC Soft Computer, Softlab Laboratory information System with DORA versions: 2.0.0.4 with Hot Fix 1.2808, 1.2809, 1.2956, or 1.3476; 2.0.0.5; 2.0.0.6; 2.0.0.7 with Hot Fix 1.6595; and 2.1.0.3. |  |  | March 22, 2011 | 2011 | SCC Soft Computer | Potential for patient data to be sent to the wrong patient record. A software change affected how the Softlab handles recovery from certain types of table access conflicts. When such conflicts occur, this function has the potential to cause faulty writing to the database. | Software change control | Clients were notified through SCC Soft Computer's proprietary Task Management System On 03/22/201. All clients have received a hot fix that corrected the issue or made the hosparam change that prevents the issue. |  | 12 | 12 | Sold throughout US and Canada. | Terminated | October 29, 2012 | 587 | 12 | Software | Output/Calculation | Software update | Software Update
Z-0282-2012 | 59913 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Flexible Probe with Blocking Water, GM11002420, (Part of the Segmented Cylinder Set, GM11004150) |  |  | May 20, 2010 | 2010 | Varian Medical Systems, Inc. | The Instructions For Use ( IFU) distributed with product prior to 2004 lacks the correct limit to the sterilization cycles and checks regarding the firm seating of the blocking washer. | Labeling Change Control | Varian Medical Systems sent a Medical Device Recall (PNL-GM-42697-2) in May 2010 for the blocking washer on the flexible probe, Varian part number GM11002420, a component of the Segmented Cylinder Applicator Set. The letter was to inform users that a flexible probe with an unstable or shifted blocking washer must not be used due to the potential for treatment length error. The blocking washer may move from rough handling or sterilization. Users were told to discard any instances of the obsolete Instructions for Use (IFU) in use and replace it with the IFU (PGM11004150 4 March 24, 2010.pdf). A Receipt Verification Card was included requesting consignees to fill out, fax or e-mail immediately to Varian BrachyTherapy. |  | 81 | 81 parts | Worldwide Distribution: USA, Germany, Belgium, Canada, Mexico City, Spain, Brazil, Switzerland, Uzbekistan, India | Terminated | December 05, 2011 | 564 | 81 | Not_Computer
Z-0281-2012 | 60269 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity(R) Delta, Delta XL, Gamma XXL and Kappa Monitors are used to monitor, physiologically, the vital signs of patients (with Arrhythmia Detection or Alarms). |  |  | October 31, 2011 | 2011 | Draeger Medical, Inc. | Firm received 2 complaints reporting Ventricular Fibrillation (Vfib) arrhythmias were not recognized by the monitor and did not alarm when the Masimo SET pod is used for SpO2 measurements. | Software design | Firm initiated their recall of this device on October 25, 2011 by sending an "Urgent Medical Device Recall" letter to consignees. A Technical Service Bulletin will be generated and distributed globally that instructs the affected International Draeger subsidiaries to inform customers in their countries. The software anomaly has been corrected in software version 7.3. Draeger Service Representatives will schedule time with customers to upgrade the software in affected monitors. The software upgrade will be provided free of charge. |  | 14894 | 14,894 | Worldwide distribution- USA including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SD, TN, TX, UT, VA, WI and WV and International distribution including the countries of Algeria, Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Botswana, Brasilien, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Deutschland, Dominican Republic, Ecuador, Egypt, El Salvador, Equatorial Guinea, Estonia, Finland, France, Frankreich, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indien, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Jordanien, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Madagascar, Malaysia, Malta, Mauritius, Mayotte, Mexico, Moldavia, Moldova, Morocco, Netherlands, Norway, Norway, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Rumania, Russian Fed., Saudi Arabia, Serbia, Seychelles, Singapore, Slovenia, Slowenien, South Africa, South Korea, Spain, Sri Lanka, SÂżdafrika, Sweden, Switzerland, Syria, Taiwan, Thailand, Tschechische Republic, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and White Russia. | Terminated | July 01, 2013 | 609 | 14894 | Software | Alarm/Message | Software update | Software Update
Z-0279-2012 | 60305 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS 5085 and 5085SRT Surgical Tables, Catalog Numbers: 5085- ST01-410-1; 5085SRT- ST0-1420-6. |  |  | October 17, 2011 | 2011 | Steris Corporation | Some hydraulic column cylinders installed in certain 5085 and 5085SRT Surgical Tables were assembled by their supplier with incorrect snap rings and an anomaly was present in the tilt cylinder. This error may affect the user's ability to move the table top out of the full right tilt position. | Nonconforming Material/Component | On October 17, 2011, Steris sent 'Urgent Voluntary Field Correction Notices' to their customers via Fedex with tracking numbers for delivery confirmation. The letter identifies the affected product and notifies the user/customers of the problem with the surgical tables. The notification letter goes on to inform the customers that a STERIS Service Representative will be contacting them to arrange for the field correction of the affected unit(s). This involves the replacement of the hydraulic column cylinder of each surgical table(s). |  | 167 | 167 tables were distributed - 5085=108 and 5085SRT=59 | Nationwide distribution. | Terminated | March 19, 2013 | 519 | 167 | Not_Computer
Z-0276-2012 | 59980 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Sensar Foldable IOL with OptiEdge. Model number AR40e 21.5 Diopter. |  |  | September 20, 2011 | 2011 | Abbott Medical Optics Inc (AMO) | The firm is initiating this voluntary recall because of the possibility that certain Sensar IOLs are labeled with an incorrect diopter power. | Employee error | Abbott Medical Optics, Inc. sent an URGENT FIELD SAFETY NOTICE dated September 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using and remove the affected product from their inventory, complete and return the included facsimile form to Abbott Medical Optics and indicate the status of each affected serial number on the form. Customers were asked to return the form even if they have no inventory affected by the recall. For any questions customers were instructed to contact their Abbott Medical Optics Customer Service Representative.For any questions regarding this recall call 714-247-8656. |  | 96 | 96 | Worldwide Distribution including China, India, Indonesia, Korea, Malasyia, Taiwan, Thailand, Austria, Finland, France, Germany, Great Britain, Hungary, Italy, Netherlands, Portugal, and Saudy Arabia | Terminated | January 08, 2014 | 841 | 179 | Not_Computer
Z-0280-2012 | 60219 | No | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Inova Diagnostic Inc., NOVA Lite ANA KSL Kit |  |  | September 26, 2011 | 2011 | Inova Diagnostics Incorporated | The NOVA Lite ANA KSL kit label has the wrong fill volume for the IFA System Negative Control (P/N 508186) with 1mL. The component in the kit is correctly labeled and filled with the 0.5mL fill volume. | Process change control | On 9/26/11, Inova Diagnostics forwarded an Urgent Advisory Notice and Product Advisory Notice Fax Back Form to all their customers who purchased the Nova Lite Ana KSL, Product Code: 708390. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete the attached fax back form indicating how many of these kits they have in their inventory, quantity destroyed, method of destruction and the quantity of kits that need to be replaced. Customers were also instructed to notify their customers of the advisory notice. Customers were instructed to contact Tana Keivens with any questions or comments at (858) 586-9900 ext 351. |  | 500 | 500 kits | Worldwide distribution to: USA, Mexico, Kuwait, Uraguay, Romania, Italy, Philippines, India, Malaysia, Algeria, Czechoslovakia, Brazil, Colombia, Ireland, Korea, and Saudi Arabia. | Terminated | April 04, 2012 | 191 | 500 | Not_Computer
Z-0284-2012 | 60142 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System. |  |  | October 14, 2011 | 2011 | Siemens Medical Solutions USA, Inc. | Errors could potentially impact image interpretation. The acquisition software is exhibiting a defect in the energy window position when an isotope is peaked. There is a potential patient safety issue when running syngo MI Applications version VA60A with Foresight detectors firmware version v1.5.9.12 on the Symbia systems. | Component change control | Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated October 14, 2011 to all affected customers. The letters informed customers of the acquisition software defect in the energy window position when an isotope is peaked, which could potentially impact image interpretation, and that the AEB firmware version on their imaging system has been changed to a previously released version until acquisition software service pack is available. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at (800) 888-7436. |  | 3 | 3 units | Florida, Illinois and Texas | Terminated | October 16, 2012 | 368 | 5 | Software | Physical Safety Hazards | Software update | Software Update
Z-0274-2012 | 60328 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSDV), sterile EO, Rx Only, Synovis Surgical Innovations, A Division of Synovis Life Technologies, Inc, 2575 University Ave W., St Paul, MN 55114-1024. |  |  | October 26, 2011 | 2011 | Synovis Surgical Inovation Div. of | Synovis Surgical Innovations is voluntarily recalling one lot of Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) product because some of the devices were mislabeled. | Labeling mix-ups | Synovis Surgical Inovation sent a "Recall Notification" letter dated October 26, 2011. The letter described the problem and product affected. Customers were instructed to identify and quarantine the affected product and to complete and fax the Product Return/Replacement Form. |  | 132 | 132 units | Nationwide Distribution including CT, GA, NJ, NY, OH, TX, VA. | Terminated | April 24, 2012 | 181 | 132 | Not_Computer
Z-0277-2012 | 60436 | No | LNX |  |  | General Hospital | 510(k) | SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. |  |  | May 03, 2011 | 2011 | SCC Soft Computer | On 05/03/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the following SoftLab GUI versions 4.0.2.8-4.0.2.10; 4.0.3.12-4.0.3.13; and 4.0.4.1-4.0.4.4. A client reported the PT and INR results did not get flagged. The PROT should have been flagged as HIGH and the INR should have been flagged as PANIC. Neither test was flagged. All affected clients were notified of the issue on 05/03 | Software design | SCC Soft Computer initiated a Field correction on May 3, 2011. All customers with affected software have been notified of this issue. Clients were instructed to have a "valcheck" utility run to determine if there are any tests in their system affected by the issue. They were supposed to indicate whether they wanted to schedule to receive the patch or hot fix.If you have any further questions please call ( 727 ) 789-0100. |  | 69 | 69 | Worldwide Distribution-- USA ( nationwide) and Canada. | Terminated | August 07, 2013 | 827 | 69 | Software | Output/Calculation | Software update | Software Update
Z-0285-2012 | 60166 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Curix Opthos H, size 35x43 cm |  |  | October 04, 2011 | 2011 | AGFA Corp. | The film notch was located in the wrong position. | Process control | An "URGENT SAFETY NOTICE" letter was sent via email to the dealers on October 7, 2011. The letter describes the problem and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood has been requested from the dealer sites. Included in each dealer letter, Agfa also provided a customer notification letter to be sent to their end-use customers. Agfa has requested dealers and end-use customers to file a complaint at 1-877-777-2432 for any identified product still in their inventories. Agfa will replace product and arrange for return shipment. |  | 125 | 125 | Worldwide Distribution -- USA, including the states of IN, KS, MI, MN, MO, NC, NE, NY, OH, and TN and the country of Canada. | Terminated | December 22, 2011 | 79 | 125 | Not_Computer
Z-0292-2012 | 60413 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira LifeShield Latex-Free 150 mL Burette Set, with Calibrated Burette and 2 Clave Ports, Convertible Pin, 77 inch , Soluset Microdrip with Option-Lok; an Rx sterile medical device used to administer fluids; 20 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 19208-01 |  |  | November 10, 2011 | 2011 | Hospira Inc. | The float valve in the burette sticks to the burette wall and does not open or close properly. | Employee error | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated November 10, 2011 to the customers of record via UPS on the same date. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return Reply Form via fax to 1-866-913-0667 or email to: hospira2530@stericycle.com; return affected product to Stericycle using the label provided with this letter; and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability. |  | 12580 | 12,580 units | Worldwide distribution: USA (nationwide) including states of: Alabama, Alaska, California, Colorado, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kentucky, Massachusetts, Michigan, Minnesota, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oregon, Pennsylvania, Rhode Island, Tennessee, Texas, Vermont, Virginia, Washington and Wisconsin, and internationally to Canada | Open, Classified |  |  | 12580 | Not_Computer
Z-0288-2012 | 60355 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | ImmuneTech brand MyAllergyTest, Specimen Collection Kit, Part No. |  |  | November 04, 2011 | 2011 | Immunetech Inc | ImmuneTech is recalling certain Specimen Collection Kits that contain Triad brand non-sterile alcohol prep pads which have been recalled by Triad Group, Hartland, WI due to Bacillus cereus, microbial contamination. | Material/Component Contamination | The firm, IMMUNETECH, sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated November 4, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory and quarantine all affected kits; if they have further distributed the affected kits, contact Immunetech to accurately identify and immediately notify their customers and facilitate the return and exchange of all remaining kits. If you have any questions, call (650) 312-1066, Monday-Friday, 9:00am -5:00pm (pacific) or email: customersupport@immunetech.com. |  | 310 | 310 | Worldwide distribution. | Terminated | June 14, 2012 | 223 | 3340 | Not_Computer
Z-0294-2012 | 60031 | scale, stand-on, patient | FRI | Stand-on patient scale. | General Hospital | General Hospital | 510(K) Exempt | Health o Meter Professional Remote Display Digital Physician Scale, Model 498KL; digital patient weighing scale; Pelstar LLC, 11800 South Austin Avenue, Unit B, Alsip, IL 60803, U.S.A., Made n China; Model 498KL |  |  | October 07, 2011 | 2011 | Pelstar, Llc | There is a potential for an electrical short when installing or replacing the batteries in the remote display head of the scales. This short may cause the batteries to smoke, burn or burst. | Device Design | Pelstar LLC. sent a Urgent Device Recall letter dated October 7, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informed them of the potential of an electrical short caused when installing or replacing the batteries which may result in minor injuries to the person installing the batteries. The customers were requested to immediately examine their inventory and quarantine all product subject to the recall. In addition, if they distributed this product to end users, they were requested to identify all customers and notify them of this product recall at once in writing, using the enclosed recall notification letter. The customers were also requested to complete and return the enclosed mandatory response form as soon as possible. Any questions were directed to Health O Meter Professional at 800-815-6615 or 708-377-0600. |  | 2498 | 2,498 units | Worldwide Distribution -- USA (nationwide) including the countries of Canada, New Zealand, Israel and Lebanon. | Open, Classified |  |  | 2498 | Hardware | Physical Safety Hazards | Remove | Remove or Replace
Z-0293-2012 | 60304 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia 1.5T and 3.0T MRI Systems, Model Numbers: 781396 and 781377, 510(k) #K110151. System units equipped with software version: 4.1.1 are subject to recall/field correction. |  |  | August 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The firm determined that during scanning procedures, for some larger patients, some reddening of the skin of can occur in certain instances due to RF heating when the torso of the patient is positioned within the body coil. | Software design | Philips Medical Systems ( Cleveland) Inc. sent a Urgent Field Safety Notice dated August 3, 2011, to all affected customers. Philips Field Service Engineers were dispatched to the affected customer sites to deliver the recall notification and to perform the necessary software corrections. The letter lists the affected MRI systems; describes the problem; the actions to be taken by the customer/user; it also informs the customer of the upcoming software correction which will be applied to each unit by the Philips Field Service Engineers in order to address this issue. If any additional information is required, the customer is instructed to contact their local Philips Field Service Engineer. For further questions please call Customer Care Solutions Center (CCSC) at 800-722-9377. |  | 54 | 26 units - 12 units distributed in U.S. &amp; 16 distributed to foreign customers | Worldwide Distribution -- USA (nationwide) including the states of AL, AZ, CA, GA, OR, MI, NJ and TX., and the countries of AU, BE, CH, DE, FR, JP, NL and SE. | Terminated | February 22, 2013 | 569 | 54 | Software | Physical Safety Hazards | Software update | Software Update
Z-0297-2012 | 60233 | No | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Multigent Vancomycin Calibrators; |  |  | June 08, 2011 | 2011 | Microgenics Corp | Complaint from the sole distributor that the third party control material is recovering high, and outside the published range. | Nonconforming Material/Component | Microgenics sent a Product Removal notification letter dated June 15, 2011 to all affected customers. The letter identified the affected product, the problem, and the necessary actions to be taken. Customers were instructed to discontinue use and dispose of the affected product per local waste ordinances, review all patient results reported with the use of this lot, and retain a copy of this notification for their records. The letter states that if the affected product has been forwarded to another laboratory, a copy of the letter is to be provided to them. An alternative calibrator is available to customers impacted. It can be ordered by calling Thermo Fischer Scientific Customer Service at 1-800-232-3342.Questions regarding the information should be forwarded to Abbott Customer Service at 1-877-4ABBOTT. |  | 118839963 | 809 kits lot 59446762; 651 kits lot 59391741 | Nationwide Distribution (USA) in the state of ILLINOIS | Terminated | February 17, 2012 | 254 | 118839963 | Not_Computer
Z-0303-2012 | 59479 | Yes | DWB | Roller-type cardiopulmonary bypass blood pump. | Cardiovascular | Cardiovascular | 510(k) | Battery Module for Sarns Modular Perfusion System 8000 |  |  | October 21, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems is aware on one instance in which an incorrectly labeled printed circuit board for the Sarns Modular Perfusion System 8000 was shipped to a single customer. The affected board mislabeled two connection between the board and the transformer. | Employee error | Terumo conducted a Correction/Removal - Customer Communicated Script on November 7, 2011 to the affected customer. The phone script communicated the correction and reason for removal, potential hazards, affected population, and confirmation of possession of affected module. A replacement date was set for November 7, 2011. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818, hours are Monday to Friday. 8 AM -6 PM EST. |  | 10 | 10 units | Nationwide Distribution (USA) including the state of San Antonio, TX | Terminated | February 17, 2012 | 119 | 10 | Not_Computer
Z-0305-2012 | 60121 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5 |  |  | March 09, 2011 | 2011 | Elekta, Inc. | If "Confirm" or "Unconfirm" settings are clicked during the "Terminated Checking" phase of treatment delivery, Desktop Pro cancels the loaded beam and sets the Beam MUs to zero. Desktop Pro then sends the delivered MUs to MOSAIQ a a zero value. | Software design | Elekta sent an IMPORTANT NOTICE letter dated March 9, 2011 to affected customers. The notice identified the product, the problem, and the corrective action to be taken by the customer. The notice advise all customers to follow any instructions or recommendations covered in the Notice. For questions call +44(0)1293 654200. |  | 56 | 56 | Nationwide Distribution (USA) including the states of: AZ, CA, GA, IL, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, TN, VA, WI and Puerto Rico | Terminated | August 09, 2012 | 519 | 56 | Other |  |  | Other
Z-0306-2012 | 60192 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ON-Q Silver Dressing is a Antimicrobial wound dressing containing silver. It is individually wrapped in a foil pouch. Each pouch contains 1 single use 4" x 4.5" pad. There are 5 pouches per box. Ref: SD4X4.5; Part Number: IF-151023 and IF-151026. |  |  | June 30, 2008 | 2008 | Acrymed Incorporated | The Silver Dressing, lot H0705J and H0706J may contain dressings that are not properly sealed in the foil pouch, which could affect sterility of the dressing. | Process control | On June 30 2008, the firm e-mailed all affected customers with an Important Product Recall Notification letter.The customers were notified that two lots of SilverDressing may contain dressings that are not properly seated in the foil pouch.The customers were instructed to inspect the inventory, ceased distribution of the affected lots and quaratine the remaining inventory and notify AcryMed.Customers can call AcryMed at 503-624-9830 for any questions about this recall. |  | 41061 | 14,125 units from lot # H0706J and 25,525 units from lot # H0705J | Nationwide distribution -- DE, FL, KY, OR, IN, TX, NV, MT, CA and NY. | Terminated | December 20, 2011 | 1268 | 41061 | Not_Computer
Z-0299-2012 | 59581 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | The Terumo Advanced Perfusion System 1, Catalog Number 801764 System 1 Base, 220 V |  |  | November 08, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | 1 US customer was operating the Advanced Perfusion System 1 with an unapproved, modified cable, which was in product development, to connect the system an an occluder for the Sarns Modular Perfusion System 8000 . Five(5)of twelve(12) cables from Terumo Japan, which were in product development, were given to a US customer. | Component design/selection | Terumo sent an URGENT MEDICAL DEVICE CORRECTION: Recall Notification dated November 8, 2011 to all affected customers. The letter described the reason for the recall, known/potential hazards and request immediate location and correction of product. The letter instructs customers to stop using the modified cables and return to Terumo for replacement and destruction. Customers were asked to complete the attached response form and fax to the number indicated on the form. All affected customers will be receiving a safety advisory to advise them that Terumo do not recommend the alternative configuration of a System 1 Base with an 8K Occluder. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818 or 1-800-292-6551 (fax). |  | 12 | 12 units | Worldwide Distribution - (USA) Nationwide Distribution | Terminated | August 22, 2012 | 288 | 12 | Not_Computer
Z-0302-2012 | 60289 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1- 4" Roller Pump for Terumo System 1 |  |  | November 10, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo CVS discovered that our In house service center has incorrectly installed a felt washer into the main bearing of nine roller pumps that had been returned for repair or preventative maintenance unrelated to the felt washer. The felt washer was added as a validated running change to production to all roller pumps manufactured since February 2009. The felt washer should not have been installed | Nonconforming Material/Component | Terumo sent an URGENT MEDICAL DEVICE REMOVAL notification letter dated November 10, 2011 to all affected customers. The letter described the affected products, problem, potential hazards, and important clinical information. Replacement roller pumps will be configured at the customer site by a Terumo Field representative with the affected roller pumps being returned via normal returned goods process. Customers were asked to complete the attached response form and fax to the number indicated on the form. For questions or concerns contact Terumo CVS Customer Service at 1-800-521-2818. |  | 3 | 3 units | Worldwide Distribution - USA (nationwide) including the states of: CA, KY, MO, NE, PA and TX and the countries of Canada and Europe. | Terminated | September 06, 2012 | 301 | 9 | Not_Computer
Z-0304-2012 | 60204 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Ambulatory Blood Pressure (ABP) Monitor, Model 90217A. |  |  | October 12, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | If external pressure is applied to the cuff of the Ambulatory Blood Pressure, Model 90217A, during deflation, the monitor will stop working and display an EC75 error code. | Device Design | On October 21, 2011, a Medial Device Correction customer letter was mailed to all US customers. A return receipt was requested.The letter identified the affected devices and described the problem. If the device displays an EC75 error code, customers can cycle the power off and then on to restore normal functionality. Please advise patients of this mitigation to avoid a delay in obtaining results. The firm is in the process of upgrading their software to avoid the issue and will contact customers to arrange for an update to their monitors.Customers can call Spacelabs Healthcare at 1-800-522-7025, select 2 for Technical support for any question about this recall. |  | 785 | 785 units worldwide (83 in US and 702 outside US) | Worldwide Distribution | Terminated | May 30, 2012 | 231 | 785 | Software | Device Operation | Software update | Software Update
Z-0311-2012 | 60146 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Cobas c 111 Analyzer |  |  | October 18, 2011 | 2011 | Roche Diagnostics Operations, Inc. | Internal investigations have revealed a potential printer malfunction on the Cobas c 111 analyzer, where one or more numbers or characters can be missing on the hard copy result printout, or the printer may skip leaving blanks which may result in an incorrect print layout. The missing number of characters can be in any position, therefore it could be in the middle of a result, (e.g. 48.53 mmol in | Software design | Roche Diagnostics Operations, Inc sent a Urgent Medical Device Correction dated October 18, 2011, via UPS Ground (receipt required) to all affected customers and one distribution center, with instructions to disable the Auto Print Result feature until new software is available. All printed results will need to be confirmed against the results indicated on the display screen. Questions are to be directed to Roche Diagnostics Technical Support at 1-800 428-2336, |  | 135 | 135 | Worldwide Distribution - USA including West Virginia and Texas. | Terminated | February 21, 2013 | 492 | 135 | Software | Output/Calculation | Software update | Software Update
Z-0309-2012 | 60314 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Arrow International StimuCath(R) Continuous Nerve Block Kit, ASK-19608-SFH, Lot # MF0127039 |  |  | June 20, 2011 | 2011 | Arrow International Inc | Some of the outer trays used in the packaging of the ASK-19608-SFH Continuous Nerve Block Kits, lot # MF0127039 were found to have an insufficient seal. Therefore, the packaging may have been compromised, and sterility of the product cannot be guaranteed. | Packaging | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 20, 2011, to its customer. The letter included a stock response form to be completed and returned to the firm. The letter also described the product, problem and action to be taken. The customer was instructed to check their stock for the products included in the scope of the recall; cease use and distribution; quarantine all affected product immediately; complete and return the enclosed Recall Acknowledgement &amp; Stock Status Form to their sales representative, and return any affected product freight collect to: Arrow International, Inc., 312 Commerce Place, Asheboro, NC 27203, attn: Asheboro Quality Assurance Manager. Arrow International can provide replacement with a similar product or credit their account.If you have questions or concerns, please call this toll free number 800-233-3187. |  | 17 | 17 kits | Nationwide distribution: CT | Terminated | October 01, 2012 | 469 | 17 | Not_Computer
Z-0313-2012 | 60323 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Rel 3, a radiographic system, digital |  |  | November 01, 2011 | 2011 | Philips Healthcare Inc. | Missing a small, metal protective cover on the bottom side of the control handle for the ceiling suspension. exposes printed circuit | Process control | The firm, Philips Healthcare, sent a "URGENT - Field Safety Notice" letter dated October 25, 2011, to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed: Do not place your hand in the open slot below the Control Handle and ensure that no patient places his/her hands the opening; review the following information with all members of their staff who need to be aware of the contents of this communication, and retain a copy with the equipment instruction for use. Philips has issued a Field Change order instructing Service Engineers to check affected systems and install the cover where it is missing free of charge.If you need any further information or support concerning this issue call 978-659-4519. |  | 28 | 28 (7 US; 2 Canada; 19 OUS) | Worldwide distribution: USA (nationwide) including states of: NC, SC, TX, and WV; and countries of: Australia, Austria,Canada, Germany, Italy, Netherlands, Sweden, and Switzerland. | Open, Classified |  |  | 28 | Not_Computer
Z-0319-2012 | 60150 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular Collimator |  |  | March 25, 2011 | 2011 | Elekta, Inc. | Recent newspaper articles outlined improper use of SRS Cone Collimator accessories that injured patients on Brainlab and Varian systems. | Under Investigation by firm | Elekta sent an IMPORTANT NOTICE letter dated March 25, 2011 to all affected customers. The letter identified the affected product, the problem, key safety procedures and information to users for safe operation of the Stereotactic Circular Collimator. Elekta will issue a Mandatory Action Field Change Order which will include an updated Circular Collimator Instructions for Use - 102022201 and new labels. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. Also, Elekta is offering some free WebEx sessions which will include the procedures and recommended operation of these types of accessories. Customers are instructed to register at www.elekta.com/SafeAccessories. |  | 89 | 89 devices | Nationwide Distribution (USA) - including the states of: AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN, and WV | Terminated | October 02, 2012 | 557 | 89 | Not_Computer
Z-0310-2012 | 60282 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas c 311 Analyzer, Part Number: 04826876001 |  |  | October 26, 2011 | 2011 | Roche Diagnostics Operations, Inc. | The cobas c 311 analyzer software may incorrectly calculate the volume of the ISE Reference Electrolyte solution remaining on the analyzer. In some cases the incorrect results obtained will not be flagged with any data alarms. | Software design | October 26, 2011 Urgent Medical Device Corrections were sent to all direct accounts in the USA that have installed the cobas c 311 analyzer and included a faxback form to be completed and faxed to 1-877-766-7452. Consignees were instructed to temporarily replace the ISE Reference Electrolyte bottle based upon the frequency calculated from the maximum number of ISE samples run per day and activate the yellow threshold alarm on the system. Customers with question and concerns are directed to call Roche Diagnostics Technical Support at 1-800-428-2366, available 24-hours, 7 days a week. |  | 940 | 940 worldwide | Worldwide Distribution -- USA including the states of Maryland and Texas. | Terminated | July 09, 2012 | 257 | 940 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0314-2012 | 60287 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek Panalock Loop, Lupine Anchor |  |  | October 31, 2011 | 2011 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Incorrect suture configuration | Employee error | The firm, DePuy Mitek Inc., sent an "URGENT VOLUNTARY PRODUCT RECALL"letter dated October 31, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to not use or sell any of the identified problem products included in this recall; complete the enclosed business reply form and fax to: 1-508-828-3750 or email to: mitekcomplaints@its.jnj.com (note: It is important that we receive this business reply form, even if you have no product remaining in your inventory from the affected lots); isolate all inventory and call 1-877-379-4871 for return goods authorization and replacement product; and return all identified affected product to: (USA)- DePuy Mitek, ATTN: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324 or (foreign)-GMED Healthcare EDC Quality Dept Rue de Luxembourg 5, ZI Trazegnies, BE-6180 Courcelles, Belgium. If you have any questions or concerns with regard to this program, please contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-877-379-4871. |  | 298 | 298 units | Worldwide distribution: USA (nationwide) and countries of: Austria, Australia, Brazil, Canada, Estonia,Switzerland, China, Czechoslovakia, Denmark, Spain, Finland, France, Great Britain Israel, India, Japan, Korea, Mexico, Martinique, Norway, Portugal, Sweden, Singapore, Solvonia, and Slovakia. | Terminated | August 05, 2015 | 1374 | 1768 | Not_Computer
Z-0307-2012 | 60125 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma Plan |  |  | June 23, 2011 | 2011 | Elekta, Inc. | There is a risk that the user may interpret the instructions in the manual too literally, and adjust the TMR 10 dose description to obtain an ETCPD that matches previous clinical prescriptions, which in turn could lead to a prescription dose that deviates a few percent from the intended dose. | Labeling False and Misleading | Elekta sent an IMPORTANT NOTICE letter dated June 23, 2011 to affected customers. The notice identified the product, problem, and corrective actions to be taken by the user. The letter advise all customers to follow any instructions or recommendations covered in the Notice. The letter states that Elekta is working on a new software version and customers will receive the future version when released. For questions contact your local Elekta representative. |  | 2 | 2 | Nationwide Distribution (USA) - including the states of: NM and WA | Terminated | January 31, 2012 | 222 | 2 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0295-2012 | 59856 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Flexible Applicator Probe, Part GM11002560 - 3.2 mm diameter braced PVDF. Varian Medical Systems, Palo Alto, CA, Mfg. by Varian Medical Systems, Haan GmbH. |  |  | August 05, 2008 | 2008 | Varian Medical Systems, Inc. | GammaMed Flexible Applicator Probes may not have been heat annealed during manufacturing to prevent shrinkage after autoclave. | Process control | Varian Medical Systems, Inc. sent a Customer Technical Bulletin CTB-GM-603A on August 5, 2008 to all affected customers. The letter included description of device and problem. The bulletin also advised that if they find that the applicator and source guide tube combination is not within the indicated range on the length gauge, to call VBT technical support using the contact information on page 5 of the bulletin to obtain a replacement applicator probe. Customers are to complete and return the attached verification card. If support is needed the letter provides customer support contact information. |  | 653 | 653 parts | Worldwide Distribution -- USA (nationwide): Arkansas,California, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Kentucky, Maine, Maryland, Massachusetts, Minnesota, Montana, Michigan, Missouri, Nebraska, Mississippi, New Jersey, Nevada, New York, North Carolina, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington, West Virginia, and Wisconsin, and the countries of Austria, Azerbaijan, Brazil, British Columbia, Canada, Chile, Germany, Greece, India, India, Ireland, Israel, Mexico, Morocco, Pakistan, Poland, Republic of Macedonia, Russia, Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Ukraine, and Venezuela. | Terminated | December 07, 2011 | 1219 | 653 | Not_Computer
Z-0312-2012 | 60128 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Leksell Gamma Knife C 1.2, 4 and 4C |  |  | May 26, 2011 | 2011 | Elekta, Inc. | Several of the LMR03 actuators with bronze drive nut have failed unexpectedly, creating a potential safety hazard for the operator and patient. | Pending | Elekta sent an IMPORTANT NOTICE: Urgent Field Safety Notice letter dated May 12, 2011 to affected customers. The notice identified the affected product, problem, and actions to be taken by the user. The letter describes special care to be used in operating the helmet changer actuator LMR03. The letter states that Elekta will replace all actuators LMRO3. For more details contact your local Elekta representative.. |  | 44 | 44 | Nationwide Distribution (USA) - including the states of: CA, CO, DAL, FL, GA, IL, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OK, OR, PA, TN, TX, VA, WA, WI and Puerto Rico | Terminated | October 06, 2014 | 1229 | 44 | Not_Computer
Z-0326-2012 | 59779 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed Plastic Needle with Mandrin; 2.0 mm diameter, length 113 mm. Remote controlled radionuclide applicator system. Varian Medical Systems, Inc. Mfg by Varian Medical Systems Haan GmbH, Haan, Germany. |  |  | September 30, 2008 | 2008 | Varian Medical Systems, Inc. | Plastic needle tip could possibly crack or separate. | Nonconforming Material/Component | Varian Medical Systems sent a "CUSTOMER TECHNICAL BULLETIN" letter dated September 30, 2008 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructed users to inspect needle tips before and after every clinical application to prevent needle tip separation or cracking. |  | 1716 | 1716 parts | Worldwide Distribution - USA (nationwide) including the states of Arizonia, California, Florida, Illinois, Indiana, Michigan, North Carolina, and Vermont, and the countries of Argentina, Austria, Belgium, Canada, Estonia, France, Germany, India, Norway, Pakistan, Russia, Slovenia, So. Africia, Spain, Switzerland, United Kingdom, and Ukraine. | Terminated | December 07, 2011 | 1163 | 1716 | Not_Computer
Z-0331-2012 | 60229 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Model CAT-850B Catheterization Table, a component of the Toshiba INFX-8000 series X-Ray Interventional System. |  |  | October 03, 2011 | 2011 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that during patient transfer from a gurney to the catherization table, the tabletop may rotate unexpectedly when sufficient force is applied to the tabletop. The result of the tabletop rotating unexpectedly is the patient being tranferred may fall from the tabletop. | Device Design | Toshiba America Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 3, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers were instructed to follow the user instructions by positioning the tabletop at the foot-end position, and ensuringsufficient assistance is available, when transferring a patient from the gurney to the CAT-850B catheterization table. Customers were also instructed to follow the instructions for transferring a patient to the CAT-850B that are included in the operation manual. A Customer Reply Form was included with the letter for customers to complete and return via fax to (877) 349-3054. The customer reply form can also be sent via e-mail to raffairs@tams.com. Questions regarding the recall letter can contact the Director of Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 300 | 300 units | Nationwide Distribution. | Terminated | March 18, 2013 | 532 | 300 | Not_Computer
Z-0337-2012 | 59767 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Dual Incu i (Atom Infant Incubator Model 100) |  |  | August 23, 2011 | 2011 | Atom Medical Corporation | The caster with lock on the incubator or infant warmer may become loose or break. | Device Design | The firm, Atom Medical, sent an "Urgent: Field Safety Notice" letter dated August 18, 2011, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify all devices in their inventory and quarantine them until the rework is complete; identify any customers to whom they have further distributed the product; contact their customers at once and arrange a time in which the field service technician can perform the rework (the parts needed to complete the rework will be sent from Atom Medical's export department ); complete rework of all affected devices and document rework on the Distributor FSCA/Recall Response Form and return via Fax to: +81 3 3812-3199 or email to kira-utc@atomed.co.jp; t-sugino@atomed.co.jp or kenichi.shimada@atomed.co.jp (all forms should be completed and returned by January 31, 2012.) If you have any questions, contact the Manager Quality Control or Manager, Regulatory Affairs at email addresses above, phone: +81-3-3815-3632 or fax +81-3-3812-3199. |  | 105 | 105 units | Worldwide distribution: USA (nationwide) including state of: PA and countries of: Algeria, Austria, Australia, China, Egypt, Germany, Italy, Japan, Korea, Kuwait, Malaysia, Portugal, Russia, Saudi Arabia, Spain, Switzerland, Thailand, The Netherlands, and UAE. | Terminated | September 09, 2013 | 748 | 540 | Not_Computer
Z-0345-2012 | 59934 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Odyssey |  |  | November 04, 2010 | 2010 | PerMedics, Inc. | Anomaly 1 - Invalid Export of Beam Table and or Collimator Angle for DICOM RT Ion, DICOM Image and RTP Link export files:During PerMedics' internal code reviews for lEC 61217 compliance verification, an issue was identified by PerMedics engineers in three export files: D1COM RT Ion Plan, D1COM RT Image and RTP Link. In some cases, this issue can result in table and/or collimator angle values t | Device Design | PerMedics sent an Odyssey TPS URGENT NOTICE OF CORRECTION letter dated November 4, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Anomaly 1 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations were provided if they planned to use Odyssey RT Ion Plan, DICOM RT Image, and/or RTP Link export:Workaround for the invalid exported table and/or collimator angle errors: If, prior to the software correction, any of these affected file types were exported from Odyssey, manually verify that the table and collimator angles are correct. Additionally, apply appropriate changes to any angles deemed incorrect that were obtained from Odyssey's DICOM RT Ion Plan, DICOM RT Image, and/or RTP Link files.Only those machines with table and/or collimator angles that increase in the clockwise direction will have the incorrect angles.Anomaly 2 User Corrective Action:Install the mandatory new version correcting the software as soon as it is available (see PerMedics Corrective Action below). In the meantime, the following recommendations are provided if the user- Changes the patients scheduled start date prior to the software correction:Workaround for Incorrect Monitor Unit Values when Changing Treatment Start Date: If, prior to the software correction, the user changes the patient's treatment start date on any fully-computed IMRT plan, IMAT plan or plan containing a "Region DVH" prescription, the user should force a full recommendation oft hose patient plans prior to fixing and exporting the plan.The following procedure can be used to force a full re-computation of the plan:- Open the patient's plan worksheet- Change the algorithm type of the plan- Change the algorithm type back to its original value- Close t |  | 15 | 15 | Worldwide Distribution - USA including NY, GA, CA and the countries of Hungary, Brazil, China, India, Italy, and Mexico | Terminated | March 02, 2012 | 484 | 15 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0325-2012 | 59885 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VariSource iX, Software Version 1.1, Part Number AL20020000 VariSource iX Afterloader. |  |  | March 01, 2010 | 2010 | Varian Medical Systems, Inc. | VariSource iX console software may cause incorrect dwell times to be delivered to patient and may result in a mistreatment. | Software design | Varian Medical Systems sent a "MEDICAL DEVICE RECALL" letter dated April 1, 2010 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were asked to conduct a simple test to establish if their machine is affected by the problem. Device users were requested to complete an attached receipt verification form and return it to Varian BrachyTherapy via fax or e-mail. Call the Contact Service Dispatch number at 1-800-864-1672 for questions regarding this detail. |  | 3600 | 72 worldwide/50 in U.S. | Worldwide Distribution-USA (nationwide) | Terminated | May 30, 2013 | 1186 | 3600 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0336-2012 | 60286 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 10mm x 2.7mm RibLoc Screw, a component of the RibLoc Rib Fracture Plating System, Model RBP1202. |  |  | February 28, 2011 | 2011 | Acute Innovations Llc | RibLoc screw from lot L1008007 was out of specification. It would not engage with the hex tip properly. | Process control | The firm, ACUTE Innovation, sent an "URGENT NOTICE: DEVICE RECALL" letter dated February 21, 2011, to all consignees/customers via FED-EX on February 28, 2011. The letter describes the product, problem and actions to be taken. The customers were instructed to identify product having the lot code(s) and quarantine them; return all product from this lot to Acute Innovations immediately using shipping labels provided or return to: Acute Innovations, LLC, 21421 NW Jacobson Rd, Suite 700, Hillsboro, OR, 97124, and if they further distributed this product, notify those customers of this recall and request return of product(s).Customers can call ACUTE Innovation at 1-866-623-4137 for any questions about this recall. |  | 299 | 299 units (267 in US and 32 to South Africa) | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, GA, IL, MN, NC, NY, UT, and WA; and countries of: South Africa. | Terminated | December 06, 2011 | 281 | 299 | Not_Computer
Z-0330-2012 | 60231 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Varian brand GammaMed 12i Accessory: Clamping adapters for bronchial catheters, 1.8mm diameter, GammaMed 12i; Ref: BT-00264, Model 2-H63; Product is distributed by Varian Medical Systems Inc., Palo Alto, CA and Manufactured by Varian Medical Systems Haan GmbH, |  |  | October 06, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The Product Models 1 &amp; 2 have a design flaw problem where the clamping adapter canobstruct movement of the source cable in the event the bronchial catheter is dislodgedfrom the connector. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION LETTER/FIELD SAFETY NOTICE" letter dated October 6, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Contact the Varian Help Desk at 1-888-827-4265 for questions regarding this notice. |  |  |  | Worldwide Distribution--USA (nationwide). | Terminated | October 22, 2012 | 382 |  | Not_Computer
Z-0344-2012 | 60302 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips GEMINI TF 64 Diagnostic CT X-ray/PET Imaging System, Catalog/Model #882471, 510(k) #K052640. System units equipped with software version: 3.5.2 are subject to recall/field correction. |  |  | April 25, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | In response to complaints received by Philips Medical concerning certain identified software errors reported by their customers, the firm is conducting a recall/software correction to several of their GEMINI PET/CT X-ray system units in the field. The firm believes that if these software defects were to recur, it could impact the performance of reliability of the these PET/CT system units. | Software design | Philips Healthcare sent a "CPE CUSTOMER LETTER" dated October 28, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Additionally, device operation information is provided to customers. A Philips Field Service Engineer will be visiting each customer site to install the necessary software upgrades. Contact the Customer Care Solutions Center at 1-800-722-9377 for questions concerning this notice. |  | 10 | 10 system units were distributed. | Nationwide Distribution including the states of CO, LA, MD, MI, MN, PA, TX, VA, VT, and WI. | Terminated | February 22, 2013 | 669 | 38 | Software |  | Software update | Software Update
Z-0322-2012 | 59814 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZMR Porous Revision Hip Prosthesis Surgical Technique, utilized for the following ZMR devices: Porous Stem,Nitrited Porous Revision, Slotted Porous Bowed Stems with and without Compression Nut, Splined Straight Stem, Taper Stem and Bodies, Calcar Bodies, and Spout Bodies. |  |  | October 21, 2011 | 2011 | Zimmer Inc. | This recall was initiated because it may not be clear to the surgeon whether the device can be used based on the instructions for use. | Error in labeling | The firm Zimmer, sent an "URGENT DEVICE CORRECTION" letter dated August 26, 2011 to Marketing Personnel, Distributors, Sales Representatives, Distribution Operation Managers, and Surgeons. The letters identified the reason for the recall, the product, and required actions to be taken. The letter also included Medwatch reporting information and a Certificate of Acknowledgement form to fax back to the recalling firm to confirm receipt of the recall notice. For questions contact Zimmer at 1-877-946-2761. |  | 119284 | 119,284 total devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dubai/UAE, Eschbach, Finland, France, Germany, Greece, Hong Kong, Hungary, Italy, Japan, Jordan, Israel, Korea, Lebanon, Mexico, Netherlands, New Zealand, Peru, Poland, Portugal, Russia, S. Africa, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Taiwan, Thailand, Turkey, UK, Ukraine, and Venezuela, | Terminated | August 29, 2013 | 678 | 119284 | Not_Computer
Z-0320-2012 | 60250 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI R.4.5 and R4.6 |  |  | October 14, 2011 | 2011 | Elekta, Inc. | The software did not operate as expected. | Software design | Elekta sent an IMPORTANT NOTICE letter dated October 14, 2011 to all affected consignees. The letter identified the affected product, the problem, and the actions to be taken. This notice contained precautionary measures that are required for the operation of the equipment. The notice advise all customers to follow any instructions or recommendations covered in the Notice. The notice instructs Users to file this document in the Important Notice section of the appropriate User Manual. For questions contact your local Elekta representative. |  | 77 | 77 | Nationwide Distribution (USA) - including the states of: AZ, CA, CT, GA, ID, IL, LA, MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI and Puerto Rico | Terminated | November 03, 2014 | 1116 | 77 | Software |  | Instructions | Safety Notice/Insructions
Z-0333-2012 | 60244 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System, Models 6441*, 6442*, 6465*, 6488 (USA). |  |  | October 12, 2011 | 2011 | Boston Scientific CRM Corp | Current Boston Scientific defibrillator systems measure shock lead impedances up to 200 ohms, and export results to the programmer and LATITUDE monitoring system via the Daily Measurements feature. A measured shock lead impedance that is greater than 125 ohms or less than 20 ohms will generate an alert message on programmer screens and a Red Alert within LATITUDE (if activated). However, the mea | Software design | Boston Scientific sent an "IMPORTANT DIAGNOSTIC INFORMATION" letter dated October 2011 to all affected customers. The letter described the product, problem and actions to be taken by the customers. |  | 25300 | 25,300 (14,700 US; 10,600 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NORWAY, PORTUGAL, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, CANADA, AUSTRALIA, NEW ZEALAND, CYPRUS, CZECH REPUBLIC, GREECE, GUADELOUPE, HUNGARY, BANGLADESH, CHINA, COCOS ISLAND, HONG KONG, INDIA, INDONESIA, KOREA, MALAYSIA, PHILIPPINES, SINGAPORE, SOUTH KOREA, SRI LANKA, THAILAND, RUSSIA, SAUDI ARABIA, SLOVAKIA, SLOVENIA, SOUTH AFRICA, NEW CALEDONIA, POLAND, PORTUGAL, QATAR, REUNION,IRAN, ISRAEL, JORDAN, KUWAIT, LATVIA, LEBANON, LUXEMBOURG, MONACO, MOROCCO, JAPAN, ARGENTINA, ARUBA, BAHAMAS, BARBADOS, BELIZE, BERMUDA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EL SALVADOR, FALKLAND ISLANDS, HAITI, HONDURAS, JAMAICA, MEXICO, PANAMA, PERU, TRINIDAD &amp; Tobago, URUGUAY, VENEZUELA, AND NETHERLANDS ANTILLES. | Terminated | December 04, 2012 | 419 | 25300 | Not_Computer
Z-0332-2012 | 60238 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Head Ring Posts with part number 970.280 - re-usable components of the Frame Array Module (of the Optical Guidance Platform and Floorstand devices. |  |  | October 04, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the Head Ring posts used by both the Optical Guidance Platform FrameArray module and the Floorstand where the head ring posts may be damaged due to excessive mechanical stress resulting in possible failure during usage. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" letter dated October 4, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. Customers are advised to verify that they are using the current design of the head ring posts (part # 970.284) and to contact Varian if replacements are needed. Customers may contact the Varian Help Desk for questions regarding this notice at 1-888-827-4265. |  | 820 | 205 kits ( Each kit contains 4 Head Post Rings) 820 HPR's | Worldwide Distribution. | Terminated | July 25, 2012 | 295 | 820 | Not_Computer
Z-0335-2012 | 60228 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, 510(k) #K033326, Model Number: 728231, Serial Numbers: 9745, 90169, and 95414. The recalled device systems are equipped with Version 3.5 Beta 1 software. |  |  | September 22, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were found for the Brilliance Big Bore and Brilliance 64 Computed Tomography X-ray Systems equipped with software version; 3.5 Beta1. The non-conformances stem from software errors identified in the operating software. | Software design | Philips Healthcare released a 'Field Change Order' (FCO #72800531) on September 12, 2011. Philips sent an - URGENT Medical Device Correction letter dated October 17, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to this situation. The letter informed customers on how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update is installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. Customers were instructed to contact their Philips representative or local Philips Healthcare Office for further information or support for this issue. For North America and Canada customers were instructed to contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 3 | three (3) units were distributed. | Worldwide Distribution - USA including Texas and the countries of Italy, India and China. | Terminated | February 22, 2013 | 519 | 10 | Software |  | Softwarae update | Software Update
Z-0342-2012 | 60497 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI X-ray Volume Imaging System |  |  | October 20, 2011 | 2011 | Elekta, Inc. | A small number of sites have experienced situations with kV generator transformers overheating due to a diode failure on the interface board. | Nonconforming Material/Component | The firm, Elekta, issued a "FIELD CHANGE ORDER MANDATORY ACTION" notice dated October 20, 2011 to its customers to add fuse protection for the kV transformers. The notice describes the product, problem, provides instructions and procedures and actions to be taken. The customers were instructed to follow any instructions or recommendations covered in the notice and to file the document in the Important Notice section of the appropriate User Manual. Elekta will contact the customer if the FCO is applicable to their kV generator.If you have any questions, call 770-300-9725. |  | 123 | 123 | Nationwide distribution: USA including states of: AL, AK, AZ, CA, CT, FL, GA, IL, IA, KY, MD, MS, MI, MN, MS, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TX, VT, VA, WA, WV, WI, Puerto Rico, &amp; DC. | Terminated | August 09, 2012 | 294 | 123 | Hardware | Physical Safety Hazards | Instructions | Safety Notice/Insructions
Z-0341-2012 | 60153 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | TC pCO2 Electrodes, Model/Part #945-655-E5260, 510(k) #K043003. |  |  | September 27, 2011 | 2011 | Radiometer America Inc | Transcutaneous pCO2-electrodes were wrongly coded as: pO2 electrode on the hardware plug for the firm's TCM series monitors. This causes the monitor to interpret the electrode wrongly, causing it to fail all calibrations and never get ready for use. | Employee error | Radiometer contacted the affected U.S. customer and informed them of the recall via telephone on September 27, 2011. Radiometer requested the return of the affected product received from the firm. For questions call (440) 871-0463. |  | 4 | 4 units were distributed. | Worldwide Distribution - USA (nationwide) in the state of San Diego, CA. | Terminated | February 19, 2012 | 145 | 4 | Not_Computer
Z-0347-2012 | 60335 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part - Pouch label: "***NanoFUSE***DBM***Manufactured and Distributed By: Nanotherapeutics, Inc. 13859 Progress Blvd, Suite 300 Alachua, FL 32615 (386) 462-9663***Description: nanoFUSE DBM***Size***Order No.:***ID***Expiration***STERILE R***Donated Human Tissue Allograft***SINGLE PATIENT USE ONLY***Rx Only***nanotherapeutics***" |  |  | March 10, 2011 | 2011 | Nanotherapeutics, Inc | On 03/10/2011 Nanotherapeutics Inc, Alachua, FL initiated a recall of their Origen DBM with Bioactive Glass Catalog # 22-2002 (2cc), 22-2005 (5cc) and 22-2010 (10cc). The same product is also distributed as NanoFUSE DBM Catalog # NAN109-02 (2cc) NAN109-05 (5cc) and NAN109-10(10cc). The products were manufactured without an approved 510k. Direct accounts were notified of the recall on 04/08/2011 | No Marketing Application | Nanotherapeutics sent an Urgent: Medical Device Recall letter dated April 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove any unimplanted/unsold product shipped to them. Customers were asked to notify their customers and instruct them to make notification until the end user (hospital/physician) is reached. Recall Instructions:Immediately examine inventory and quarantine product subject to recall.Return all quarantined product subject to recall to:Nanotherapeutics, Inc.Attn: Receiving Department13859 Progress Blvd, Suite 300Alachua, Florida 32615For questions call 386-462-9663 |  | 131 | 131 units | Nationwide Distribuiton including TX, FL, AZ, CA, NY, and MA. | Terminated | April 03, 2012 | 390 | 12814 | Not_Computer
Z-0363-2012 | 60458 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. |  |  | July 20, 2011 | 2011 | SCC Soft Computer | On 07/20/2011 SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftPath GUI versions: 4.2.3.3-4.2.3.6 and 4.3.0.5-4.3.0.12. A client reported that the HPV results on a patient report were incorrect. The HPV results were from the previous order on the patient and not the current order. All affected clients were notified about the issue on 07/20/2011. | Software design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" dated July 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. SCC personnel will review the customized Lab Results AP macro and/or the Micro Results AP macro on customer systems and will either discuss with them the irregularities requiring change or reconfigure macros as necessary. |  | 97 | 97 | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | May 21, 2013 | 671 | 97 | Software | Output/Calculation | Upgrade | Repair
Z-0355-2012 | 60451 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. |  |  | June 13, 2011 | 2011 | SCC Soft Computer | On 06/13/2011, SCC Soft Computer initiated a correction on the following SoftPath GUI Versions: 4.3.0.10, 4.3.0.11, and 4.3.0.12. A client reported when they verified the contents of the fax reports, they found one of the specimen diagnoses to be missing. Parts A through D were present as well as Part F, however, Part E diagnosis was missing for one of the patient reports. All affected clients wer | Device Design | SCC Soft Computer initiated a Field Correction on June 13, 2011. All customers with affected software have been notified of this issue.SCC recommends until the corrected patch is loaded to temporarily use one of the following alternative working solutions: 1. Use MS Word Editor functionality within SoftPath application for creating patient reports. In order to ensure Text Control editor is not inadvertently activated, the templates for Text Control editor will be deleted and reinstalled at the clients request by their SCC representative after receipt of the correction. 2. In Report Template set-up, the bottom margin should be made slightly larger than the size of the footer plus its distance from the bottom page border. For further explanation, please consult your SCC representative for further details, or call (727) 789-0100. |  | 14 | 14 | Class II Recall - Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | May 31, 2012 | 353 | 14 | Software | Output/Calculation | Software update | Software Update
Z-0359-2012 | 60409 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15 Consult instruction for use. RaySearch Laboratories AB Sveavagen 25, SE-111 34 Stockholm Sweden. support@raysearchlabs.com |  |  | October 19, 2011 | 2011 | RAYSEARCH LABORATORIES AB | The firm has found an issue with DICOM RT Image representation of DRR images from RayStation 2.0 which needs to be clarified in the user manual. The symptoms involve the Pixel Spacing and Image Position attributes in the RT images being incorrect, thus enabling the user to make measurements that are erroneous and possibly misalign the patient. | Device Design | The firm, RaySearch Laboratories, sent an email with an "Urgent Field Safety Notice Medical Device Correction" notice dated October 19, 2011 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to pass on the information in the field safety notice to everyone involved and create routines to compensate for the incorrect DRR images; if the DICOM DRR images are not used to make any form of absolute measurements, no action needs to be taken. Raysearch laboratories will provide a script to update these DICOM files to have the correct values upon request.For further information please contact RaySearch Laboratories AB Quality and Regulatory Affairs Manager at +46 722 366 110 or david.hedfors@raysearchlabs.com; RaySearch Americas Director of Customer Support at +1 877 778 3489 or freddie.cardell@raysearchlabs.com. |  | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: MA, NJ, and NY; and country of: Netherlands | Terminated | May 11, 2012 | 205 | 19 | Software | Display/Image | Software update | Software Update
Z-0348-2012 | 60237 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Oscillating Saw Attachment/Large w/Key for Small Battery Drive. |  |  | October 17, 2011 | 2011 | Synthes USA (HQ), Inc. | The potential exists for the Large Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece, there is a potential for injury to the patient and/or user. | Device Design | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers return their affected device(s) for servicing. Instructions are provided for customers to follow. Customers having questions regarding this notice should contact the firm at 610-719-5450. |  | 5 | 5 | Nationwide Distribution including the states of CO, LA, NY, and OK. | Terminated | December 18, 2014 | 1158 | 5 | Not_Computer
Z-0349-2012 | 60245 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Tibial Broach P/N 801-01-013 |  |  | October 26, 2011 | 2011 | Encore Medical, Lp | The strike plate may dislodge from the broach stem handle during impaction of the device. | Nonconforming Material/Component | DJO Global sent an URGENT FIELD SAFETY NOTICE letter dated October 26, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information to all those who need to be aware within their organization or to any organization where the affected devices have been transferred. Customers should contact Customer Service to place a replacement order and receive an RMA number and return all affected devices using the RMA number. For any questions call (512) 834-6255. |  | 30 | 30 units | Worldwide Distribution - USA including AR, AZ, CA, LA, MD, NY, TX, UT and Puerto Rico and the countries of United Kingdom and Italy | Terminated | September 05, 2012 | 315 | 30 | Not_Computer
Z-0351-2012 | 60347 | No | FCK | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The EnCorÂż biopsy probe, designed for use with EnCorÂż and EnCorÂż Enspire Vacuum Assisted Biopsy (VAB) driver systems. |  |  | November 03, 2011 | 2011 | Bard Peripheral Vascular Inc | This recall was voluntarily initiated by the firm due to the product 's damaged packaging trays; cracked EnCor (Trademark) packaging trays, plastic PETG tray the probe is packaged in was cracked. | Packaging | The firm sent the customer notification letter on 11/07/11, via FedEx with proof of delivery notice. The letter notified customers of the reason for recall that units from the affected lot may have packaging damage resulting from transportation. The letter instructed customers to check all inventory locations , remove any identified product, to not use the affected product, and return the affected product to Bard Peripheral Vascular, Inc., 1415 W. 3rd Street, Tempe, Arizona 85281. The letter also instructed customers to fill out the recall and effectiveness check form and fax it to 1-800-894-6772. Customers were instructed to call 1-800-321-4254, Option #2 Ext 2727 for a return authorization number to facilitate the expedient return of the product.For questions regarding this recall call 480-894-9515 |  | 64570 | 64570 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Europe, LAPAC, and Australia | Terminated | March 29, 2012 | 147 | 64570 | Not_Computer
Z-0362-2012 | 60416 | No | LNX |  |  | General Hospital | 510(k) | SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. |  |  | October 25, 2011 | 2011 | SCC Soft Computer | On 10/25/2011 SCC Soft Computer, Clearwater, FL initiated a correction on the SoftReports version 1.1.8.2.4 with Hot Fixes 1.7199.7 or 1.7200.7. When SoftReports Designer opened a layout, the system did not load the layout with the previously saved value in the NewPage property. This caused multiple patient records to be printed on the same page. All affected clients were notified on the issue on | Software design | SCC Soft Computer sent a "Correction Notification" letter dated October 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The hot fixes that were previously loaded which introduced this issue were corrected and delivered to all affected customers as a mandatory correction in hot fixes 1.7199.8 and 1.7200.8 and higher. |  | 5 | 5 | Nationwide Distribution. | Terminated | March 16, 2012 | 143 | 5 | Software | Output/Calculation | Software update | Software Update
Z-0364-2012 | 60411 | [6 empty] | TrueBeam and TrueBeam STx, Model Number: H19, Ref/FSCA identifier: CP-06381 are intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors, and conditions anywhere in the body where radiation treatment is indicated. |  |  | October 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the TrueBeam and TrueBeam STx systems where, under certain tuning conditions, the electron beam emerging from the bend magnet may have an elongated spot shape. | Device Design | On 11/04/2011, an Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps. Varian is developing a corrective action for this issue. A customer service representative will contact you when the correction is available to schedule installation on your system. Questions regarding the letter should be directed to the Varian Oncology Help Desk Contact Information at USA and Canada: 1-888-827- 4265 or Europe +41-41-749-8844. |  |  |  | Worldwide distribution including the US and Internationally. | Terminated | July 25, 2012 | 268 |  | Other |  | Software update | Software Update
Z-0360-2012 | 59886 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography System, Model #728231, 510(k) #K 033326, Serial Numbers: 95504, 91003, 300030, 95130, and 95157. |  |  | September 06, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were logged during in-house testing and external validation testing of Brilliance 64 and Ingenuity CT software version 4.0 Beta 2. The non-conformances are due to software errors identified in these device systems. | Software design | Philips Healthcare released a 'Field Change Order' (FCO #72800542) on September 15, 2011. Philips sent an URGENT - Medical Device Correction letter dated October 10, 2011, to all affected customers. The letter listed the affected device systems; described the nature of the software problem; and any possible hazard inherent to the situation. The letter was mailed by Philips to their customers on October 27, 2011. The letter provided information telling the customer how to identify the affected products and what action to take. In this instance, the customer was asked to review the Customer Release Notes which were provided to each customer location once the software update was installed. The letter notified the customer that a Philips representative would be installing a software upgrade in order to address the referenced software issues free-of-charge, through a Field Change Order. The letter closes by providing information which all customers in Canada and North America can use to contact the firm through a toll-free telephone number 1-800-722-9377, option 5; enter side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 5 | 5 units | Worldwide Distribution - USA including Ohio and the countries of Australia, China, Belgium, Israel, and Sweden | Terminated | March 15, 2013 | 556 | 7 | Software |  | Software update | Software Update
Z-0353-2012 | 60338 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Servo-S ventilator system |  |  | September 19, 2011 | 2011 | Maquet Inc. | Servo-I and Servo--S ventilator systems are equipped with a specific revision of a printed circuit board (PC1772 revision 22) that includes a component that has generated a number of nuisance alarms due to technical error message 24 (TE 24). | Vendor change control | The firm, Maquet Getinge Group, sent an " PRODUCT RECALL URGENT - MEDICAL DEVICE CORRECTIVE ACTION" letter dated September 19, 2011, to its customers. The letter describes the product, problem and actions to be taken. Maquet service team will contact the customers within two weeks of receiving this letter to schedule an on-site service of the affected equipment. If you have any questions, call technical support at our toll free number at 1-888-627-8383. |  | 26 | 26 units | Nationwide distribution. | Terminated | July 22, 2015 | 1402 | 202 | Not_Computer
Z-0356-2012 | 60300 | No | FCW | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Integra Surgical Luxtec MLX Light Source |  |  | October 14, 2011 | 2011 | Integra LifeSciences Corp. | A loose screw may result in shifting of components and improper function of the turret or actuator assembly. | Process control | Integra sent an URGENT: Field Corrective Action letter dated October 14, 2011 to all affected customers via FedEx. The letter identified the affected product, the problem, and the actions to be taken with instructions to immediately inspect their inventory for potentially loose screws within the affected unit(s). If the unit(s) were observed with loose screws, the unit(s) were to be returned to Integra Surgical for repair and servicing. Customers were instructed to complete the Acknowledgement and Return Form per the instructions on the form. For questions call Integra Surgical Customer Service at 1-800-431-1123. |  | 66 | 66 units | Worldwide Distribution - USA (nationwide) including the states of: Arizona, California, Colorado, Florida, Indiana, Mississippi, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Montana, Washington, Wyoming, and Puerto Rico, and the countries of: Australia, Germany, Italy, Japan, Netherlands, New Zealand, Pakistan, Sweden, Switzerland, and Turkey. | Terminated | March 08, 2012 | 146 | 97 | Not_Computer
Z-0383-2012 | 34843 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Product is 40mm Flexible Tube. ***NON-STERILE PRODUCT***CAUTION: FEDERAL LAW RESTRICTS THIS DEVICE TO SALE BY OR ON THE ORDER OF A PHYSICIAN." |  |  | February 17, 2006 | 2006 | Walter Lorenz Surgical, Inc. | Lorenz Engineer discovered flex tube could break if bent too much. | Device Design | W. Lorenz Surgical notified consignees of recall by an Urgent Advisory Notice letter. Consignees were requested to acknowledge receipt of recall and provide product disposition. Initial letter was sent on March 2, 2006, March 15, 2006, follow-up letters sent on April 4, 2006 and May 15, 2006. For additional information call (904) 741-4400 or fax (904) 741-4500. |  | 76 | 76 pieces | Worldwide Distribution - USA (nationwide) including the states of: GA and TX and the country of Japan. | Terminated | December 12, 2011 | 2124 | 182 | Not_Computer
Z-0468-2012 | 51146 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilizer Hydrogen Peroxide Gas Plasma Sterilizer Product Code 10133. |  |  | December 10, 2008 | 2008 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has discovered a defect in the barcode labeling of the STERRAD NX System Cassettes 10133, Lot #08I032. This is the only lot number affected by this action. The barcodes for Lot #08I032 were printed improperly and are not capable of being read by the STERRAD unit. | Error in labeling | Advance Sterilization Products sent an "URGENT: PRODUCT RECALL" letter dated December 10, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to examine inventory and return all sealed cassettes, using an enclosed prepaid UPS return label to: Stericycle, 2670 Executive Drive Suite A, Indianapolis, IN 46241. Additionally, a Business Reply Card and packing slip were included with the letter for customers to complete and return. Questions or further assistance was directed to ASP Customer Care Center at 888-783-7723, option 2. |  | 776 | 776 cases | Nationwide Distribution. | Terminated | December 16, 2011 | 1101 | 776 | Not_Computer
Z-0487-2012 | 56540 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software in use with: |  |  | April 16, 2009 | 2009 | Stellate Systems | The following implied or stated indications in Stellate labeling with Harmoniae software have not yet been cleared by the Food and Drug Administration (FDA) for sale in the United States of America. | No Marketing Application | Stellate Systems sent a "DEVICE CORRECTION" letter dated April 6, 2009 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The firm provided information on how to verify if customers had the affected product. Software updates will be available to all customers in possession of the affected product. Customers were instructed to call the Customer Support line at 1-888-742-1306 with any questions or concerns they may have regarding the notification. |  | 1041 | 1,041 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Denmark, Greece, Hong Kong, Hungary, Italy, Japan, Mexico, Norway, Rep. DE Panama, Spain, Switzerland, Taiwan ROC, and Turkey. | Terminated | December 27, 2011 | 985 | 1041 | Computer
Z-0569-2012 | 57863 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Appraise/Heritage Combo A1C &amp; Blood Test Kit, Product K8300, each kit packed in a bag or envelope and contains 1 Triad Medium Alcohol Prep Pad, Cat. No. 10-3001, NDC 50730-3001-0, and blood specimen collection supplies, forms, and instructions. |  |  | January 09, 2011 | 2011 | Heritage Labs International LLC | Kits contain recalled Triad alcohol pads | Material/Component Contamination | Heritage Labs sent a Device Notification and Safety Alert letter dated January 12, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Heritage Labs sent replacement alcohol prep pads from a different supplier not affected by the recall to user of Heritage Labs kits. Heritage Labs recommended the replacements be on "as needed" basis. Upon opening kit, users should immediately dispose of sealed TRIAD brand alcohol pads. Users are to use the replacement wipes instead of the wipes contained in the kits. For any questions call (888)764-4120. |  | 4114 | 4,114 kits | Worldwide Distribution - USA (nationwide) and Germany | Terminated | December 26, 2012 | 717 | 683376 | Not_Computer
Z-0393-2012 | 60386 | No | HEB | Unscented menstrual tampon. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Kotex Natural Balance Security tampons, regular absorbency, unscented. packaged in 18 count or 36 count retail containers. Consumer Services, Kimberly-Clark Corp. Dept. KNBTR-18, P.O. Box 2020, Neenah, WI 54957-2020 USA. |  |  | November 03, 2011 | 2011 | Kimberly-Clark Corporation | A limited number of Kotex Natural Balance Security Unscented Tampons, regular absorbency, are being recalled after a raw material contaminated with a bacterium, Enterobacter sakazakii, was found when testing the plunger portion of the applicator. | Material/Component Contamination | Kimberly-Clark sent an Urgent Product Recall letter, dated 11/09/2011, via FedEx to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify, quarantine, and discontinue use immediately. Customers were asked to complete the response form and fax to +1 (920) 380-6467. A Kimberly-Clark representative would contact them regarding return of the product, replenishement and/or account credit. For questions regarding this recall call +1 (920) 216-8152. |  | 1400 | 1400 Cases (216 tampons per case) | Nationwide Distribution | Terminated | December 13, 2012 | 406 | 1400 | Not_Computer
Z-0388-2012 | 60418 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea (available in various configuration as Vitrea Enterprise Suite). Vitrea Enterprise Suite 1.2, Vitrea Enterprise Suite 1.3, Vitrea enterprise Suite 6.0 and Vitrea Enterprise Suite 6.1. |  |  | November 11, 2011 | 2011 | Vital Images, Inc. | Vital Images has found two issues that may display incorrect Standard Update Values (SUV) when viewing certain PET/CT images. This issue affects the VitreaCore PET/CT feature for Vitrea Enterprise Suite 1.2, 1.3, 6.0, 6.1, including all updates to those versions. | Software design | Vital Images, Inc. sent an URGENT DEVICE RECALL letter dated November 11, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to inform all potential users of the issues. Customers were to be contacted by their Service Account Manager to discuss and schedule resolutions.For questions customers were to contact Vital Images Customer Support with any questions toll free at 1-800-208-3005, in Europe +31 (0) 70 4135801, in Asia t +86 10 8588 1800, or online at support@vitalimages.com. |  | 338 | 338 (233 US) | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY and the countries of IRELAND, FRANCE, UNITED KINGDOM, SWITZERLAND, AUSTRALIA, SLOVENIA, GERMANY, NETHERLANDS, BELGIUM, PAKISTAN, BRAZIL, ARGENTINA, SWEDEN, QATAR, CHINA, FINLAND, ITALY, DENMARK, PORTUGAL, REPUBLIC OF KOREA, SPAIN, and CANADA | Terminated | December 03, 2013 | 753 | 338 | Computer
Z-0523-2012 | 60423 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MAMMOMAT Inspiration |  |  | September 15, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Lowering the swivel arm with "object table" while an operator's or patient's body parts are underneath the "object table" may cause the "object table" to collide with and injure the operator or patient. Therefore attention should be paid when adjusting "object table" height to avoid a possible risk of injury to the operator or the patient.The movement of the swivel arm may also cause an injur | Use error | Siemens Medical Solutions USA, Inc sent a Safety Advisory Notice dated January 8, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer.The notification included an Addendum to the User Manual.We appreciate your understanding and cooperation with this safety advisory and ask you to immediately instruct your operating personnel accordingly. Please ensure that this safety advisory is placed in the System's Operator Manual. Your operating personnel should maintain awareness over an appropriate defined period. Further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of FL, MD, MI, NE, NY, NC, ND, OH, and SC. | Terminated | March 05, 2014 | 902 | 12 | Not_Computer
Z-0522-2012 | 60426 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plaza |  |  | September 21, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Potential malfunction when using syngo.plaza. On syngo.plaza workplaces large studies can be loaded into the viewer if transferred from a DICOM node or imported even while image transfer is not yet completed. In these situations more images of that study may exist which are not shown in the syngo.plaza viewer. | Software design | Siemens Medical Solutions USA sent a Customer Safety Advisory Notice dated March 10, 2011 to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The letter informed the customer of the potential issues and provides additional information on the implemented resolution.For further questions please call (610) 219-4834. |  | 1 | 1 | Nationwide Distribution New York | Terminated | January 31, 2014 | 863 | 1 | Not_Computer
Z-0508-2012 | 60325 | No | HXK | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Offset Wire Fixation Bolt, catalog number 8180-50-006. The bolts are used with the DePuy Ace-Fischer External Fixation System and are packaged in a flexible polyethylene bag. |  |  | November 08, 2011 | 2011 | DePuy Orthopaedics, Inc. | DePuy has received three complaints concerning oversized bolt threads in the Offset Wire Fixation Bolts used in the Ace- Fischer External Fixation System. This could lead to the inability to properly thread the nut onto the bolt. | Component change control | On 11/9/2011 DePuy sent "Medical Device Recall Notice" letters to Distributors and to direct account hospitals. Recall notices included instructions to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN via distributors and sales representatives. The product is going to be discontinued and no reconditioning or reworking for the product is planned. For questions about recall information provided, please contact Katie Seppa, Manager of Customer Quality, 574-372-7333 (M-F; 8 am - pm EST.) |  | 890 | 890 | Nationwide Distribution | Terminated | March 07, 2013 | 485 | 890 | Not_Computer
Z-0350-2012 | 60434 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA AQUAFORM Comfort Science enfilcon A Daily Wear SPHERE Soft Contact Lenses (disposable), Rx only, STERILE --- CooperVision: Hamble, SO31 4RF, UK -- Scottsville, NY 14546 USA -- Juana Diaz, PR 00795 USA --- Device Listing # D033736. Packaged: 6 Soft Contact Lenses per unit carton or trial product 3 pack. |  |  | November 15, 2011 | 2011 | Coopervision Inc. | Presence of silicone oil residue on Avaira Sphere Soft Contact Lenses. | Process control | CooperVision Inc. issued a press release on November 15, 2011. An "URGENT: MEDICAL DEVICE RECALL- AVAIRA SPHERE CONTACT LENSES" letter dated November 16, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their inventory, stop further distribution of the affected product, contact their patients and to advise their patients to check the firm's website to confirm they have the affected product. Additionally, a response form was included with the letter for customers to complete and return via fax at 866-484-1048. Contact Coopervision at 800-341-2020 for questions regarding this recall |  | 4877897 | Domestic: approx. 3,155,781 lenses; Foreign: estimated 1,722,116 lenses | Worldwide Distribution--USA (nationwide) including Puerto Rico. | Open, Classified |  |  | 4877897 | Not_Computer
Z-0444-2012 | 60455 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab used with SA INST 4.0.2.0 to 4.0.2.51, 4.0.3.0 to 4.0.3.13, 4.5.0.0 to 4.5.0.9, 4.5.2.0 to 4.5.2.8 and SA INST 3.1.3.0 to 3.1.3.29, 3.1.6.0 to 3.1.6.17, 4.0.0.0, 4.0.1.0 to 4.0.1.32. |  |  | July 26, 2011 | 2011 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL, initiated a correction for the SoftLab used with SA INST. A test performed on the Roche Cobas Integra series of instruments when used with the GenInst/Dispatch listening program could cause incorrect flagging for a test result posted to the lab database. | Software design | The firm, SCC Soft Computer, sent a "Correction Communication" letter on July 26, 2011 to customers/consignees. The letter describes the product problem and actions to be taken. The firm will correct the software to recognize and support all known formats of the resulting date and time sent by an instrument. The software is available via a recommended patch or hot fix to the client's current release. The customers were asked to acknowledge receipt of the task and grant permission for SCC to correct the affected configuration files. If you have any questions, call (727) 789-0100. |  | 265 | 265 | Worldwide Distribution - USA (Nationwide) and Country of: Canada. | Terminated | May 21, 2013 | 665 | 265 | Computer
Z-0389-2012 | 60344 | [6 empty] | Magnus Hybrid operating table columns 1180.01A1 and 1180.01B1 |  |  | September 19, 2011 | 2011 | Maquet Inc. | A component of the Magnus hybrid OR system fails to respond to a position command. | Device Design | Maquet sent a PRODUCT RECALL URGENT MEDICAL DEVICE CORRECTIVE ACTION letter dated September 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. To avoid the potential of a malfunction occurring prior to their scheduled Service procedure, customers were instructed to ensure that the MAGNUS table-top is positioned at the "0" pivot position and that the MAGNUS column pivot brake is secured. If it is necessary to pivot the MAGNUS table-top prior to procedure start, ensure that the column pivot brake is completely released and then completely secured when returned to "0" pivot position. Customers were instructed to ensure that all user of the affected product and other necessary persons are informed of the product correction notice. For technical support or questions call 1-888-627-8383 |  | 45 | 45 UNITS | Worldwide Distribution - USA including VA, LA, MO, NC, MD, and AZ and the countries of AE, AU, CA, CH, CN, DE, DK, ES, IE, IL, SP, PL, RU, and TH. | Terminated | December 08, 2011 | 80 | 45 | Not_Computer
Z-0601-2012 | 60367 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter Âż HmX Hematology Analyzer with Autoloader, |  |  | October 13, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman has identified a problem with Integrated Circuit chips used in different boards throughout the analyzer. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated October 13, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to contact Beckman Coulter Service at (800) 526-7694 if they were experiencing failures. Customers not experiencing failures, but their system may have the suspect chips, were told they would be contacted by Beckman Coulter Service to correct their problem. Customers were also instructed to complete and return an enclosed Response Form. |  | 98 | 98 units total | Worldwide Distribution. | Terminated | June 01, 2012 | 232 | 98 | Not_Computer
Z-0394-2012 | 60440 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OR table TITAN, Surgical Table. |  |  | January 21, 2011 | 2011 | Trumpf Medical Systems, Inc. | Product is being recalled due to unintended movement on the Titan Tables. | Device Design | Trumpf Medical Systems sent an initial notification to customers dated January 24, 2011. A subsequent "URGENT SAFETY NOTICE-UPDATE" letter dated May 26, 2011 was sent to all affected customers. The letter described the product, problem, and actions to be taken by the customers. A Corrective Action Plan includes a software update to be installed by a Trumpf Technical Service Representative. Contact Technical Service at 888-474-9359 for questions regarding this notice. |  | 94 | 94 beds total for all types. | Worldwide Distribution. | Terminated | August 13, 2012 | 570 | 98 | Computer
Z-0417-2012 | 60447 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OSS INTERLOK BOWED IM STEM W/SCREW 14 MM X 150 MM TI 6AL 4V ALLOY. |  |  | November 01, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that Part 150367, (OSSCemented 1M Stem 13x150) was incorrectly labeled as part 150368 (OSS Cemented 1M Stem14x150). | Mixed-up of materials/components | Biomet, Inc. sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated November 1, 2011 to the one distributor with the affected product. The letter described the product, problem, and actions to be taken by the customer. The letter instructs the customer to immediately locate and remove the affected product from inventory. Additionally, the customer was asked to complete and return a Response Form via fax to 574-372-1683. Call 574-372-3983 for questions related to this notice. |  | 4 | 3 units &amp; 1 pending release | Distributed only in The Netherlands. | Terminated | September 10, 2012 | 314 | 4 | Not_Computer
Z-0454-2012 | 60359 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Dental cement under the brand name Cement-It and Natural Elegance (private label of Cement-It). |  |  | March 17, 2011 | 2011 | Kerr/pentron Dba Kerr Corporation And Pentron Clinical | The firm initiated a recall because the product sets faster than specified in the "Directions for Use" and in release specifications. | Other | The firm, Pentron Clinical, sent an " URGENT: MEDICAL DEVICE RECALL" letters dated April 4, 2011 and April 20, 2011 via USPS 1st class mail to its customers. The letter describes the product, problem and actions to be taken. The letter instructed customers to review the table with affected product lots and determine if the affected products was in their inventory; return the affected product for a replacement at no charge; contact Pentron Clinical Customer Service at (800) 551-0283; 1-800-243-3969 ext. 677 or email: aestridge@pentron.com directly to handle the arrangements of a quick return and replacement; and complete and fax back the enclosed Recall Return Form at 1-877-677-8844. The letter also request the customer to identify and recover the affected product lots that may have been shipped to thier customers.If you have any questions, contact Pentron Clinical Customer Service at (800) 551-0283. |  | 157 | 157 | Worldwide distribution: USA (nationwide) including states of: PA, IL, NV, WA, OH, TN, WI, NY, NJ, CT, IN FL, MI, OK, NE, HI, VA, CA, and TX; and countries of : Canada, Egypt, Isreal, Netherlands, Sweden and UK. | Terminated | January 22, 2014 | 1042 | 157 | Not_Computer
Z-0370-2012 | 60441 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Release 10 with RT Chart Release 10 (Client Build 10.0.28 or Build 10.0.34), An |  |  | October 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with Varian Eclipse and RT Chart where adding a reference pointin RT Chart and editing the field dose contributions may change the Prescribed Dose Per Fractionin Eclipse. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated October 31, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Varian has corrected the issue and will contact customers to schedule installation of the corrected version of the system. Contact the Varian Oncology Help Desk at 1-888-827-4265 for questions regarding this notice.. |  |  |  | Worldwide Distribution. | Terminated | June 06, 2014 | 949 |  | Not_Computer
Z-0596-2012 | 60504 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, pre-filled flush solution, for IV Flush only, sterile fluid path; 10 mL Single-Use Syringe with Male Luer Lock, 100 syringes per pack, 4 packs per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list no. 1978-20 |  |  | November 18, 2011 | 2011 | Hospira Inc. | There is particulate matter in the fluid pathway, which has been identified as the same material as the rubber tip of the syringe plunger. | Material/Component Contamination | Hospira sent a "URGENT DEVICE RECALL" letter dated November 18, 2011, to all affected customers via UPS. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to check their inventory and immediately quarantine any affected product complete and return Reply Form via fax to 1-866-912-2512. Return affected product to Stericycle using the label provided with the letter and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability.Hospira expanded the recall to include thirteen additional affected lots via letter dated January 31, 2012. The recall instructions remain the same as in the November 18,2011 letter. |  | 5817600 | 5,817,600 syringes | Nationwide distribution including states of: Alabama, Alaska, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Mississippi, Missouri, Nebraska, New Hampshire, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, West Virginia and Wisconsin. | Open, Classified |  |  | 5817600 | Not_Computer
Z-0488-2012 | 60452 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. |  |  | May 03, 2011 | 2011 | SCC Soft Computer | On 05/03/2011, SCC Soft Computer, Clearwater, FL, initiated a correction on the following SoftLab GUI versions: 3.18.1.0-3.18.1.21; 3.19.1.0-3.19.1.25; 3.19.2.0-3.19.2.21; 4.0.0.2; 4.0.1.0-4.0.1.16; 4.0.2.0-4.0.2.10; 4.0.3.0-4.0.3.16; 4.0.4.0-4.0.4.6; and 4.5.0.0-4.5.0.8. A client reported that results were not flagged correctly. This can occur when a test is ordered in Order Entry, then collected | Software design | SCC Soft Computer sent a "CORRECTION COMMUNICATION" notification letter to all affected customers. The letter describes the product, Risk-to-Health notification task, potential hazard, and corrective action. The customers were asked to acknowledge receipt of the task and grant permission for SCC to make the required and recommended changes. For any questions, call (727) 789-0100. |  | 220 | 220 | Worldwide Distribution - USA (nationwide) and the country of: Canada. | Terminated | August 07, 2013 | 827 | 220 | Computer
Z-0447-2012 | 60421 | [6 empty] | Box labeled in part, "REF: 72-1000***IQ SYSTEM IQ INTELLIGENT DRIVER 16.51CM (6.5")(L) RABALON, STAINLESS STEEL, GOLD PLATED***NON-STERILE PRODUCT*** MADE IN THE PHILIPPINES*** BIOMET MICROFIXATION 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA www.biometmicrofixation.com***" |  |  | July 08, 2011 | 2011 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc., Jacksonville, FL, recalled their IQ Intelligent Device, 72-1000 due to failure of the front bit sensor to detect the correct bit that is installed. When a screw is inserted with the device in drill mode, it is possible for the screw or bone to strip. | Component design/selection | Biomet Microfixation called, emailed, or sent an Urgent Medical Device Recall letter dated July 8, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the remaining device(s) at their facility and return immediately to PIVOT for the required software upgrade. Customers were also instructed to return the attached reconciliation form to:PIVOT International10916 Strang Line RoadLenexa, KS 66215Attn: Julie LawrenceFor any questions regarding this recall call 1-800-874-7711 or 904-741-4400, ext 9468. |  | 161 | 161 devices were distributed. | Worldwide Distribution - USA AZ, CA, CO, DC, FL, GA, HI, ID, IL, IN, KS, LA, NC, NH, NY, OK, OR, PA, TN, TX, and WI and the countries of Argentina, Brazil, France, Italy, Japan, South Africa, and Thailand. | Terminated | March 09, 2012 | 245 | 161 | Computer
Z-0462-2012 | 60461 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | syngo RT Therapist and syngo RT Therapist Connect/MOSAIQ OIS. |  |  | November 02, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The software release will bring current Installed Base system to the forward production version. It is a result of the final system testing addressing several safety issues that are not critical for the safe use of the device and multiple performance defects. | Software design | Siemens sent Customer Information letters and Update Instructions TH024/11/S for the installation of software release 4.2.108 beginning on November 2, 2011 to all affected consignees. Sites with active devices for the Oncology system were sent the current Update Instructions TH014/11/S for the installation of software release 4.2.108 and Customer Information letter beginning November 22, 2011.Consignees were asked to include the Customer Information in their syngo RT Oncologist System Owner Manual chapter "Safety Advisory Letters" where it should remain. For questions regarding this recall call 925-602-8157. |  | 147 | 85 affected customers RT Therapa[ist, 62 customers Oncology | Worlwide Distribution - USA (nationwide) including UT, OH, TN, FL, WI, IN, KY, CA, PA, SD, NY and the countries of India, Germany, United Kingdom, Australia, Poland, Japan, Croatia, India, Columbia, Belgium, Puerto Rico, China, Sweden, South Africa and Czech Republic. | Terminated | March 21, 2012 | 140 | 147 | Computer
Z-0446-2012 | 59231 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Osteopore PCL Scaffold Bone Filler: Osteopore Osteoplug-C, Osteoplug, and Osteomesh. |  |  | April 27, 2011 | 2011 | OSTEOPORE INTERNATIONAL PTE LTD | Product sterility may be compromised. | Nonconforming Material/Component | Osteopore sent an Urgent Voluntary Product Recall notice to their consignees by letter dated April 27, 2011. |  | 330 | 330 | Nationwide distribution - one distributor in TX. | Terminated | June 06, 2012 | 406 | 330 | Not_Computer
Z-0461-2012 | 60542 | No | NYQ |  |  | Immunology | PMA | SPOT-Light" HER2 CISHÂż Kit, catalog #84-0150, 20 tests/kit, labeled in part ***Invitrogen Corporation 7335 Executive Way Frederick MD 21704 www.invitrogen.com*** |  |  | November 14, 2011 | 2011 | Life Technologies Corporation | In vitro diagnostic reagent may be contaminated with a fungal contaminant. | Material/Component Contamination | Life Technologies sent an URGENT DEVICE RECALL letter dated November 14, 2011 by mail to all affected customers. The letter identified the product, description of problem, and actions to be taken by the user. Customers were instructed to examine inventory immediately, discontinue use, quarantine the material and return remaining inventory. Customers were advised to complete the Customer Response Sheet included for recovering inventory data and fax back to (+1) 716-774-6727 Technical Support or e-mail to techsupport@lifetech.com. Customers who had distributed the affected kits outside of their facility were instructed to immediately notify their customers of the recall. For questions call Technical Support, at (+1) 800-955-6288 opt. 2 then opt. 3. |  | 29 | 29 kits | Worldwide Distribution - USA (nationwide) in the states of CA and NV and the countries of: Canada, Venezuela, and the UK. | Terminated | February 10, 2012 | 88 | 29 | Not_Computer
Z-0466-2012 | 60459 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. |  |  | August 25, 2011 | 2011 | SCC Soft Computer | On 08/25/2011, SCC Soft Computer initiated a correction on the following SoftMic GUI versions: 4.0.3.0-4.0.3.15, 4.0.4.1-4.0.4.7. For clients using SoftMic Result Entry, when a certain series of keystrokes on the test grid are performed while entering results and statusing a test, the test status may not be maintained upon saving. The status may be removed or copied from another test on the order | Software design | SCC Soft Computer initiated a Field Correction dated August 25, 2011. to all affected customers.Customers with the affected software have been notified and are scheduling receipt of the hot fix..Further questions call (727) 789-0100. |  | 144 | 144 | Worldwide Distribution -- USA (nationwide) and Canada. | Terminated | October 03, 2014 | 1135 | 144 | Computer
Z-0445-2012 | 60565 | No | NAA | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | CrystalensÂż Accommodating Posterior Chamber Intraocular Lens, AT-50AO. |  |  | November 14, 2011 | 2011 | Bausch and Lomb, Incorporated | Bausch + Lomb voluntary initiated the recall due to a physician complaint of haze on the Crystalens AOwhich was visualized using an angled beam from a slit lamp during the post-operativeexamination. | Nonconforming Material/Component | Bausch + Lomb sent an "URGENT- MEDICAL DEVICE RECALL" letter dated November 14, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to remove the affected products from inventory and return the lenses to the firm. Additionally, a Recall Acknowledgement Form was attached to the letter for customers to complete and return with the affected products. Contact Bausch + Lomb at 949-521-7895 for questions regarding this recall.. |  | 62 | 62 units | Nationwide Distribution | Terminated | May 23, 2012 | 191 | 62 | Not_Computer
Z-0457-2012 | 60571 | Yes | LWS |  |  | Cardiovascular | PMA | St. Jude Medical Riata (8Fr), Models: 1560, 1561, 1562, 1570, 1571, 1572, 1580, 1581, 1582, 1590, 1591, 1592 |  |  | November 28, 2011 | 2011 | St Jude Medical CRMD | St. Jude Medical has confirmed the failures associated with all cause insulation failure on the St. Jude Medical Riata and Riata ST Silicone Endocardial Defibrillation Leads with specific emphasis on externalized conductors. | Device Design | St. Jude Medical sent a Medical Device Advisory Important Product Information Update letter dated November 28, 2011, Physician Communication dated December 15, 2010, Product Performance Report Extracts, and a Physician Device Advisory Notification to physicians by FedEx priority overnight service for receipt on November 29, 2011. The letters provides the customers with an explanation of the problem, the product, and the actions to be taken. The issuance of the recall letter is in conjunction with St. Jude's recent release November 2011 Product Performance Report (available online at http://sjmprofessional.com). The letter also explains about the Lead Peformance Summary, Root Cause, Clinical Implications, Rate of Occurrence from Complaints and Returns, New Peer Reviewed Publications, Data on Optim Insulated Leads (Durata and Riata ST Optim), and Recommendations and Mitigations. Physicians with any questions or concerns can contact their local St. Jude Medical representative or the Technical Services Department at (800) 722-3774. |  | 128000 | 128,000 units total | USA (Nationwide) and worldwide to: Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hungary, Iceland, Italy, Romania, Russia, Netherlands, Norway, Poland, Portugal, Russia, Serbia, Slovakia, Spain, Sweden, Switzerland, The Netherlands, UK, Australia, China, Hong Kong, Indonesia, Japan, Korea, Malaysia, Philippines, Singapore, Taiwan, Thailand, Vietnam, New Zealand, Argentina, Brazil, Colombia, Costa Rica, Ecuador, Mexico, Puerto Rico, Uruguay, Egypt, Greece, India, Iran, Israel, Joran, Kuwait, Lebanon, Oman, Qatar, Saudi Arabia, South Africa, Turkey, Malta, Cyprus, Iraq, Egypt, Lybia, Yemen, and U.A.E. | Terminated | May 09, 2012 | 163 | 128000 | Not_Computer
Z-0520-2012 | 60592 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Lifescreen Apnea is Lifescreen ECG Screening System, Version 3.31. |  |  | November 11, 2011 | 2011 | Del Mar Reynolds Medical, Ltd. | When Lifescreen is left open and a new recording is loaded from Sentinel Cardiology InformationManagement System, customers can get corruption in the data or a wrong ECG after the first32 hours. Both recordings have to be longer than 32 hours for this to happen. | Software design | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated November 11, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to close Lifescreen every time before loading another recording from the Sentinel database. Spacelabs is in the process of upgrading software and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 523 | total 523 units and 171 kits (34 units and 7 kits in US; 489 units and 164 kits outside US) | Worldwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, KY, LA, MA, MD, MI, MN, NC, NY, OH, TN, TX, UT, WV, and PUERTO RICO and the countries of AUSTRIA, BELGIUM, CHILE, CHINA, COLOMBIA, COSTA RICA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, KUWAIT, MALTA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, Slovakia, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM | Terminated | June 05, 2012 | 207 | 523 | Computer
Z-0489-2012 | 60502 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur CA19-9 Assay |  |  | November 11, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc | Positive bias of 14% and higher for some individual points for CA 19-9 results generated on the ADVIA Centaur CP system when they are compared to CA 19-9 assay results generated on either the ADVIA Centaur or ADVIA Centaur XP systems | Nonconforming Material/Component | Siemens Healthcare Diagnostics issued a Urgent Device Notification email dated November 11, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed the customers conduct a look back if they are using the ADVIA Centaur CA 19-9 interchangeably with the ADVIA Centaur or ADVIA Centaur XP systems and the ADVIA Centaur CP system and to review the contents of the recall notice with their Laboratory Director. A confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory actionFor further questions please call (508) 359.3825. |  | 21593 | 50 Test Kits =21,160; 250 Test Kits = 21,593 | Worldwide Distribution -- USA (nationwide) and the countries of Australia, Malyasia, Korea, Hong Kong, Singapore, Myanmar, Mexico, Egypt, Taiwan, South Africa, Brazil ,India, Japan, Argentina, Chile, Peru, Germany, Malaysia, Saudi Arabia, and Canada. | Terminated | April 12, 2013 | 518 | 21593 | Not_Computer
Z-0590-2012 | 59833 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DURAMAX 15.5 F x 36 cm Chronic Hemodialysis Catheter Set (With Cuff 31 cm from Tip), DuraMax Stacked Tip 36 cm Str. Basic Kit, Catalog No./REF 10302806, STERILE --- AngioDynamics, Inc. 603 Queensbury, NY 12804 USA |  |  | January 24, 2011 | 2011 | Angiodynamics, Inc. | AngioDynamics received customer complaints for breakage of the tunneler sleeve during product use. | Nonconforming Material/Component | The firm, AngioDynamics, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 28, 2011 and response form via Certified Mail, Return Receipt Requested, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to identify and segregate the recalled lots; complete and return the enclosed DuraMax HemoDialysis Catheter Recall Reply Form via fax to: Attn: DuraMax HemoDialysis Catheter Recall Coordinator at 518-798-1360 (even if they don't have any product); and do not use the affected product, return and ship the recalled product to ANGIODYNAMICS (replacement product will be shipped upon receipt and confirmation of the returned product). If the customers are in immediate need for replacement product, they should fax the completed Recall Reply Form, along with a copy of the product label, to Customer Service at (518) 798-1360 or call us at (800) 772-6446, x1358 or x1363.If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call ANGIODYNAMICS customer service at 1-800-772-6446 or email to: customerservice@angiodynamics.com. |  | 39 | 39 units | Nationwide distribution. | Terminated | February 02, 2016 | 1835 | 1124 | Not_Computer
Z-0519-2012 | 60518 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Smith &amp; Nephew, SPIDER2 Limb Positioner, including accessories Battery, foot pedal and charger |  |  | July 26, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | SPIDER2 may unlock (loss of traction) resulting inpotential movement of the patient when the device is within close proximity to an RF source. Lossof traction can potentially result in the unintended repositioning of the extremity and potential for tissue or cartilage damage | Device Design | Smith &amp; Nephew Advanced Surgical Devices Division sent an "URGENT- PRODUCT RECALL NOTIFICATION-URGENT" letter dated July 26, 2011 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return a form via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this recall. |  | 52 | 52 units | Worldwide Distribution-USA (nationwide) and the countries of Canada and France. | Completed |  |  | 52 | Not_Computer
Z-0459-2012 | 60525 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE and ONCOR linear accelerators with G31 mother board. |  |  | November 09, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Reports of gantry rotation without receiving a movement command caused by an overload within the drive signal power line. | Component design/selection | Siemens Medical Solutions USA, Inc. sent a Customer Information letter dated November 9, 2011, by federal express. The letter identified the product the problem and the action needed to be taken by the customer.The HW Update included information regarding the potential problem and provide instructions in the Safety update.Further questions call (925) 602-8157. |  | 13 | 13 active devices | Worldwide Distribution-- USA ( nationwide ) including the states of : MO, NC, OH, and OK and the countries of Turkey, Japan, Germany and Portugal | Terminated | April 15, 2013 | 523 | 13 | Not_Computer
Z-0606-2012 | 60626 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Eius Unicompartmental Knee System |  |  | August 17, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Two National Joint Registries suggest that the EIUS Unicompartmental System is associated with a higher revision rate then the unicompartmental devices. | Other | Stryker Orthopaedics sent "URGENT PRODUCT RECALL" letters dated August 17, 18 and 25, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an attached Product Recall Acknowledgement Form via fax to 201-831-6069. Contact the firm at 201-972-2100 for questions concerning this recall. |  | 26648 | 26,648 units | Nationwide Distribution. | Terminated | March 26, 2014 | 952 | 26648 | Not_Computer
Z-0595-2012 | 60628 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos Cellex Photopheresis System |  |  | September 01, 2011 | 2011 | Therakos, Inc. | There have been three confirmed reports of post reinfusion hemolysis or hemoglobinuria in patients treated on the Therakos Cellex Photopheresis System. | Other | Therakos sent an "URGENT PRODUCT CORRECTION NOTIFICATION" dated September 1, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Additionally, a Confirmation of Receipt was attached to the letter for customers to complete and return. Contact Therakos Technical Support at 1-877-865-6850 for questions regarding this notice. |  | 114 | 49 USA, 65 Foreign | Worldwide Distribution. | Terminated | March 27, 2014 | 938 | 114 | Not_Computer
Z-0456-2012 | 60636 | No | NZE | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ThermoSuit Manufactured for Life Recovery Systems. HD, LLC, 170 Kinnelon Rd. Kinnelon, NJ 07405 USA |  |  | October 21, 2011 | 2011 | Life Recovery Systems HD, LLC | ThermoSuit System Model TSP-50 Pump/Controller may have corrosion on the temperature circuit board which may result in incorrect patient temperature reading. | Process control | Life Recovery Systems sent recall letters/return response forms dated 10/24/2011 to all consignees. Life Recovery Systems has expanded their recall to include additional units and to provide additional service to products out in the field. A new recall letter was sent to consignees on 4/27/2012. |  | 28 | 28 units | Nationwide Distribution including NY, NC, LA, NJ, AL. PA TX, KY, MO, MA, and TN | Terminated | January 24, 2013 | 461 | 28 | Not_Computer
Z-0509-2012 | 60587 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Benephit XT Infusion System, Targeted Renal Therapy Infusion System, Catalog/REF 70035, STERILE -- AngioDynamics Inc. 603 Queensbury, NY 12804 USA --- COMMON/USUAL NAME - Continuous flush catheter; CLASSIFICATION NAME: Continuous flush catheter; Device Listing # D069932 |  |  | December 05, 2011 | 2011 | Angiodynamics Worldwide Headquarters | The Benephit XT Infusion System may contain hairline cracks in the sheath hub. | Nonconforming Material/Component | Angiodynamics Worldwide Headquarters sent a Urgent Medical Device Recall Letter and Response Forms dated December 2, 2011, to all affected customers via Certified Mail on December 5, 2011. The recall is being extended to the medical user level. Customers will be instructed not to use the product and to return it to AngioDynamics, Inc. in Queensbury, NY. A Return Authorization number (RA #) will be assigned to each customer in the recall notification letter and a UPS account will be provided for return freight. A reply form was provided with the recall notification letter to be completed by each consignee. ACTIONS TO BE TAKEN BY CUSTOMER / USER: (1.) Identify and segregate the recalled lots that are in your possession. (2.) Complete the enclosed Benephit XT Infusion System Recall Reply Form and fax it to the attention of the Benephit XT Infusion System Recall Coordinator at 518-798-1360. The form lists the catalog number, lot numbers, and quantity our records indicate your facility has received. It is important that even if you do not have any product remaining in your possession that you fill out the attached form noting zero quantity to be returned and fax the form to ANGIODYNAMICS. (3.) Ship the recalled product to ANGIODYNAMICS. Replacement product will be shipped upon receipt and confirmation of the returned product. Reference Return Authorization Number on the outside of the shipping box and include a copy of the Benephit Infusion System Recall Reply Form with your shipment. (4.) If you have distributed the affected product/lot, please inform your customers of this recall action immediately, and have them return the recalled units to you. --- Certified Mail will verify delivery to each domestic consignee. Non-respondents will be contacted by telephone or via e-mail. --- Corrective Action: Returned product will be quarantined and final disposition has not been determined at this time. A corrective action has been initiated by AngioDynam |  | 55 | 55 units domestically | Nationwide Distribution including the states of New York, Pennsylvania, Virginia, Kentucky, Iowa, Illinois, Missouri, Louisiana, Texas, and Utah. | Terminated | January 29, 2016 | 1516 | 55 | Not_Computer
Z-0597-2012 | 60456 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | Midmark M11 Ultra Steam Sterilizer, Distributed by: Midmark Corporation, 60 Vista Drive, Versailles, Ohio 45380. |  |  | October 24, 2011 | 2011 | Midmark Corportation | During and FDA inspection on 8/23/2011, it was discovered that the firm received numerous reports of the steam sterilizer malfunctioning. | Employee error | Midmark Corp. will begin sending customer communications out the week of January 9, 2012 to all affected customers. The letter will describe the product, problem, and actions to be taken by the customers. Additionally, a warning label and an illustration of where to place the label will be included with the letter. Customers may contact the firm at 1-800-643-6275. |  | 30255 | 30,255 units | Worldwide Distribution- USA (nationwide) including Puerto Rico and the countries of AUSTRALIA, BELGIUM, BRAZIL, CHILE, CHINA, EGYPT, FRANCE, GUAM, HONG KONG, HUNGARY, ICELAND, INDIA, KOREA, KUWAIT, MALAYSIA, MEXICO, NEW ZEALAND, POLAND, RUSSIA, SAIPAN MP, SAUDI ARABIA, SEOUL, SINGAPORE, TAIWAN, THAILAND, TURKEY, U.A.E., UNITED KINGDOM and WEST INDIES. | Terminated | June 20, 2014 | 970 | 30255 | Not_Computer
Z-0455-2012 | 60551 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Artiste Maverick Tint, Part Number: N36TB, Lot/Serial Numbers: 3416180, Product Description and Intended Use: The intended use of this device is for internal color characterization and includes 3 opaques for blocking out deep staining or metal substrates. |  |  | March 03, 2011 | 2011 | Kerr Corporation | The resin material separates when the product is extruded from the syringe. Although there are no safety issues related with this situation, please refrain from using the affected product. | Nonconforming Material/Component | Pentron Clinical sent an URGENT: Medical Device Recall letter dated April 20, 2011 to all affected customers. The letter described the product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine all product(s) subject to recall for return and exchange at no charge. In addition, the letter states that if the product was further distributed to contact their customers and notify them of the recall. Customers were also asked to complete and return the enclosed Recall Return Form. Contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangement of a quick return and replacement. |  | 100 | 100 Units | Worldwide Distribution - USA (nationwide) and the countries of: Armenia, Russia and Ukraine. | Terminated | January 09, 2012 | 312 | 100 | Not_Computer
Z-0308-2012 | 55294 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | ViroMax viral/bacterial filter. The filter reduces the risk of infection due to cross-contamination between patients |  |  | March 31, 2010 | 2010 | A M Systems Inc | Some ViroMax viral/bacterial filters are leaking at the point where the two halves of the filter are ultrasonically bonded together. | Nonconforming Material/Component | A - M Systems sent a Urgent Notification of Product Recall/Correction dated March 31, 2010, to all affected customers. The letter identificed the product the problem and the action needed to be taken by the customer.Customers were instructed to cease use and/or distribution of the affected product and to immediately notify any of their sub-accounts with the same information. Customers are asked to provide A - M Systems with the amount distributed and amount remaining in inventory and to then destroy or discard any remaining product in their possession.Please call 360.683.8300, if you have any questions or concerns. |  | 37800 | 37800 | Worldwide Distribution -- USA ( nationwide) including the states of AZ, CA, FL, KY, MA, MI, NC, OH, OR, and WI., and the countries Canada, Mexico, and Taiwan. | Terminated | January 05, 2012 | 645 | 37800 | Not_Computer
Z-0582-2012 | 48745 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Artificial Ligament Fixation Device 15MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 25MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 20MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 30MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device 35MM TI NS ToggleLoc (TM) W/Continuous Maxbraid Polyethylene Suture Loop TO Gal 4V Alloy/Polyethylene; Artificial Ligament Fixation Device #7 PE Ziploop (TM) ToggleLoc (TM). |  |  | April 24, 2008 | 2008 | Biomet Sports Medicine | Six sizes of ToggleLoc, Artificial Ligament Fixation Device, have been manufactured with the suture pull hole drilled in the wrong location on the device.The ToggleLoc implant is deployed with a suture that pulls the device through the tunnel. The hole being drilled in the wrong location will orient the device to pass through the tunnel with the angled end of the implant first. This may caus | Employee error | Biomet Sports Medicine sent an "URGENT MEDICAL DEVICE REMOVAL" notice dated May 15, 2008 to all affected customers. The letter described the product, problem, and the actions to be taken by the customers. The letter instructs customers to locate and discontinue use of the affected product. Customers should contact the firm at 800-535-8692 for questions related to this notice. |  | 203 | 203 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, GA, ID, IL, IN, KS, TX, and WA, and the countries of Australia, Chile, France, Poland, The Netherlands and Turkey. | Terminated | December 27, 2011 | 1342 | 203 | Not_Computer
Z-0604-2012 | 60479 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 101 Hematocrit/Oxygen Saturation Monitoring System. The CDI 101 Hematocrit/Oxygen Saturation Monitoring System is intended for use during cardiopulmonary bypass procedures where continuous monitoring of the blood's oxygen saturation, hematocrit and hemoglobin levels is needed. |  |  | December 07, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo is recalling the CDI 101 Hematocrit/Oxygen Saturation Monitoring System because they have not yet submitted a 510(k) to FDA to market the device. | No Marketing Application | 12/7/11 &amp; 12/8/11 All consignees received a "Transition and Notification Guide" dated 12/6/11 and a "URGENT MEDICAL DEVICE REMOVAL" letter dated 12/4/11, which will inform them of the removal of the CD l 101 and provide the user with various options. Terumo will remove CD l 101 units from the US market over the next 18 months. |  | 366 | 366 | Nationwide Distribution - Including the states of: NV, MO, SC, WA, MI, FL, GA, NY, CT, IL, VA, NY, TN, IN, HI, TX, AL, OH, MD, AZ, CA, KS, MD, PA, and NM. | Terminated | August 30, 2013 | 632 | 366 | Not_Computer
Z-0420-2012 | 60096 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Single Use PERPOS PLS System, Catalog #9045-01 |  |  | February 22, 2008 | 2008 | Interventional Spine Inc | The recall was intiated because Interventional Spine has confirmed that the PERPOS 2-in-1 Drill (6070) manufactured by a contract manufacturer for Interventional Spine, Inc., may break during use. | Pending | The recall communication was initiated on February 22, 2008 with Interventional Spine forwarding an Urgent Recall Notification to all their customers (via email). The notification informed the customers of the problem identified and instructions for the recall.If you have any questions, contact the Vice President, Quality, Regulatory and Clinical Affairs at (949) 525-1493 or email: jmetcalf@i-spineinc.com. |  | 65 | 65 | Worldwide distribution: USA (nationwide) including states of: CA, FL, MO, NY, OH, PA, TN, WA, and WI; and country of: Europe. | Terminated | February 29, 2012 | 1468 | 69 | Not_Computer
Z-0453-2012 | 60481 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Cepheid brand Xpert MRSA/SA-BC US-IVD, an In-vitro diagnostic test kit, Catalog Numbers: GXMRSA/SA-BC-10, Lots Affected: 2803 and 3001; |  |  | November 17, 2011 | 2011 | Cepheid | Cepheid has received reports from some of its customers regarding higher than desired (&gt;5%) rates of invalid results when using lot numbers 2803 and 3001 of the Xpert MRSA/SA BC (blood culture) product. | Nonconforming Material/Component | On 11/17/11, Cepheid issued letters via mail to all its direct consignees, informing them of the affected products and providing instructions on the recall. |  | 1289 | 1289 | Product was released for distribution in US. | Terminated | January 09, 2012 | 53 | 1289 | Not_Computer
Z-0515-2012 | 60424 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, versions VA70A and VA70B Client/Server Installations |  |  | September 19, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, versions VA70A and VA70B Client/Serve installations. Loading additional images/series from Patient Browser using the thumbnails can cause an unintended deletion of images/series. | Software design | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory letter dated August 24, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistenceis in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 12 | 12 | Nationwide Distribution including the states of CA, IN, IA, NV, NE, NY, OH, TN and TX. | Terminated | January 24, 2014 | 858 | 12 | Computer
Z-0390-2012 | 60407 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Maxcem Elite(TM) self-etch/self-adhesive resin cement, Refill Kit, net wt. 10 g, REF/Part # 33872, CLEAR --- Refill Kit Contents: 2 Syringes (5 g ea), 10 Automix tips (regular) 4:1, 10 Automix tips (wide) 4:1, 10 Intra-oral tips |  |  | November 10, 2011 | 2011 | Henry Schein, Inc. | The specific lots of Maxcem Elite Refill Kit appear not to be authentic. | Pending | The firm, Henry Schein, Inc., sent an "URGENT: Product Recall (Dental Practitioners/User Level) letter and response form dated November 10, 2011, via first class mail to its customers. The letter instructed the customers to examine their inventory to verify if they have any of the specified lots on hand. If so, remove it from their shelves. The product may be returned for credit within 30 days, to the following address: Henry Schein, Inc., 41 Weaver Road, Denver, PA 17517. Additionally, the letter instructed the customers to complete, sign, date, and return the enclosed Henry Schein Response Form via fax to: (631) 843-5557 or mail to: Regulatory Affairs Department (E-355) 135 Duryea Road, Melville, NY 11747, Attn: Regulatory-RR. The letter stated "Only returns of the above noted recalled item/lot numbers purchased from Henry Schein, Inc. will be credited to your account. In order to expedite your credit, please include a copy of your response form and invoice, if available, with your return".If you have any questions regarding this matter, please contact Recall Coordinator at 631-390-8000 ext. 8476. |  | 18855 | 18,855 units total: Domestic - 18,820 units; Foreign - 35 units | Worldwide distribution: USA (nationwide) and countries of: Bermuda, Martinique FWI, Nassau Bahamas, St Maarten NA, Australia, and South Africa. | Open, Classified |  |  | 18855 | Not_Computer
Z-0507-2012 | 56545 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga Unified Intelligence System (UIS), This is a software-only device, WHIE Implementation v1.1i. A calculator/data processing module, for clinical use. |  |  | July 27, 2010 | 2010 | Microsoft Corporation Health Solutions Group | The database may incorrectly match patient data resulting in the incorrect patient data being displayed. This could potentially present a risk to patient safety. | Software design | Microsoft notified the sole consignee by FAX on 7/22/2010 with an Urgent Medical Device Safety Information. The letter identified the affected product, the description of the problem, and actions to be taken to correct the problem, along with some additional information for the consignee. In the interim, the firm proposes to block access to patient records identified as mismatched. The firm also stated that the issue has been addressed in version 2.0. The customer should report any incidents to Michael Stokes, Director of Policy Compliance at 208-938-5397 or via e-mail at Michael.Stokes@microsoft.com. |  | 1 | 1 | Nationwide Distribution -- Wisconsin. | Terminated | December 20, 2011 | 511 | 1 | Not_Computer
Z-0505-2012 | 60412 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | REF UT8150 Urethal Insertion Tray 14Fr Red Rubber Catheter with PVP Swabs Qty: 20 For Rx Only Sterile unless package is damaged or opened. Latex Free Manufactured for: Coloplast A/S 3050 Denmark STERILE EO Distributed by: Coloplast Corp. Minneapolis, MN 55411 USA 1-800-328-3863 Distribue par: Coloplast Canada Corporation 3300 Ridgeway Drive Unit 12 Mississauga, Ontario Canada L5L 5Z9 1-905-820-7588 |  |  | September 02, 2011 | 2011 | Nurse Assist, Inc | Nurse Assist, Inc. received a recall notice from the Triad Group. This recall has been initiated due to a recall of the PVP Solution Packs manufactured by the Triad Group that are contained in certain procedure trays. | Component design/selection | Nurse Assist sent a recall notice dated September 2, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately and quarantine the affected product. Customers were also instructed to identify their customers and notify them at once of the product recall. Customers were instructed to complete and return the enclosed response form as soon as possible. For those distributors that choose to destroy the affected product they must use the attached Notice of Destruction to report the specifics of the destroyed product. and contact Nurse Assist Customer Service for a RGA # prior to destruction. For any questions call Nurse Assist Customer Service at 1-800-649-6800. |  | 968 | 968 cases | Worldwide Distribution - USA (nationwide) and the countries of Canada and Saudi Arabia. | Terminated | October 26, 2012 | 420 | 55118 | Not_Computer
Z-0296-2012 | 60099 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100, 200 and 202 Ventilators |  |  | October 11, 2011 | 2011 | Respironics, Inc. | Trilogy 100, 200 and 202 units are being recalled due to low impeller press-force used during the manufacturing process. If the press force were low enough, theoretically the impeller could move up the shaft and come into contact with the inside of the blower housing/impeller enclosure. This may impact therapy. | Process design | Philips Healthcare, began notifying customers via telephone on October 11, 2011. The telephone notification discussed the product, problem and actions to be taken. The customers were instructed to not remove a ventilator from a patient without providing a replacement unit; provide a fax number or email address Philips can use to send the Business Reply Form to complete and return via fax to: 800-733-9962 or email: report.request@philips.com; locate their affected devices (if the device is with a patient, replace it with a non-affected unit-Philips will provide a loaner overnight if needed)-(if the device is in their inventory, quarantine it and return it using the RA number provided in the email/fax that Philips sends them); and If loaner units are needed, they are to call 1-877-387-3311 to arrange return of your affected device and to obtain a loaner unit.If you have any questions or concerns, contact Philips Respironics at 877-387-3311. |  | 127 | 127 | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, GA, ID, IL, KS, MI, MO, NE, NV, NJ, OH, PA, TN, and TX; and countries of: Australia, Austria, Canada, China, Denmark, Germany, Great Britain, India, Italy, Japan, Korea (Republic of [South] Korea), Netherlands, Philippines, Saudi Arabia, Spain (Espana), and Turkey. | Terminated | November 06, 2012 | 392 | 127 | Not_Computer
Z-0518-2012 | 60381 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | The IS Can Disposable CO2 absorber x 6 --- |  |  | February 09, 2011 | 2011 | Intersurgical Inc | Intersurgical Incorporated has issued a voluntary recall of the IS Cans Carbon Dioxide Absorber due to a possibility that loose granules may enter the inspiratory path of the breathing system. | Device Design | Intersurgical Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 8, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected product found. A Recall Response Form was attached to the letter for customers to complete and return via fax at 315-451-3696, Attn: Customer Service. Contact Customer Service at 1-800-828-9633 if you have questions concerning this recall. |  | 110 | 110 units | Worldwide Distribution- USA (nationwide) including the states of Florida, North Carolina, New York, Ohio, and the country of Canada. | Terminated | January 15, 2013 | 706 | 110 | Not_Computer
Z-0428-2012 | 60516 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Level 1Âż Normothermic I.V. Fluid Administration Sets, |  |  | November 21, 2011 | 2011 | Smiths Medical ASD, Inc. | Disconnected/ loose Luer connectors at the patient end of the triple lumen tubing of the HOTLlNE Disposable Administration Sets and Level 1 D/DI-60HL Normothermic IV Administration Sets | Device Design | Smiths Medical notified consignees by letter dated November 28, 2011 titled: Urgent Field Safety Notice. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to inspect inventory and segregate unused affected product. Distributors were instructed to immediately stop distributing and quarantine all inventory. The letters included a Confirmation Form to be complete and return, by fax to 781-610-9859 or by email to hotlineconnectorrecall@smiths-medical.com. For questions consignees were instructed to contact Smiths Medical at 1781-763-9330. |  | 7631 | 7631 units | Worldwide Distribution  (USA) (nationwide) - Argentina, Austria, Australia, Belgium, Bahrain, Bermuda, Canada, Chile, China, Columbia, Czech Republic, Germany, Denmark, Estonia, Egypt, Finland, France, Great Britain, Guernsey, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel India, Iceland, Italy, Japan, South Korea, Kazakhstan, Lebanon, Luxembourg, Mexico, Malaysia, New Caledonia, Netherlands, Norway, New Zealand, Panama, Poland, Portugal, Qatar, Singapore, Sweden, Slovak Republic, South Africa, Spain, Sri Lanka, Switzerland, Taiwan, and Vatican City State. | Terminated | August 07, 2014 | 990 | 805506 | Not_Computer
Z-0168-2012 | 60236 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | SafeSheath(R) CSG(R) Coronary Sinus guide Braided Core Hemostatic Introducer, SafeSheath(R) Worley Telescopic Braided Series, and Situs LDS2 Lead Delivery System - Catalog numbers: CSG/WORLEY/BCor-1-09; CSG/WORLEY/BCor-2-09; CSG/WROLEY/L/BCor-1-09; CSG/Worley/R-1-09; CSGWORBC10M; CSGWORBC29M; CSGWORLBC19M; LVI/66-5.5-CR; LVI/66-5.5-RE; LVI/75-5-62-07-HO; LVI/75-5-62-07-HS; LVI/75-5-62-07-MP; LVI/75-5-62-07-RE; LVI/75-5-66-5.5-RE; LVI/Tele/B-60-07-HO; LVI/Tele/B-60-07-HS; LVI/Tele/B-60-07-MP; LVI/Tele/B-60-07-RE; RL065; RL066; RL067; and RL068 |  |  | October 07, 2011 | 2011 | Thomas Medical Products Inc | Radiopaque (RO)/soft tip of the braided core component in the CSG kits may detach from the braided core tube. | Nonconforming Material/Component | The firm,GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 7, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately segregate all affected product in a manner that ensures it will not be used; return any product with reference to the return authorization (RA) number or contact customer service representative to facilitate return of product; if any affected product has been forwarded to another facility, contact that facility to arrange return, and complete and return the attached product reconciliation form.If you have any questions, you may contact the Customer Service number at 1-866-446-3003 Monday through Friday from 8:30am to 5:00pm EDT. |  | 33134 | 33,134 | Nationwide distribution: CA only. | Terminated | August 29, 2012 | 327 | 33134 | Not_Computer
Z-0381-2012 | 60214 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | UltraLite Headlight module-grey |  |  | October 10, 2011 | 2011 | Integra LifeSciences Corp. | Headlight assembly (module) may become loose and separate from the headband assembly | Device Design | Integra Life Sciences notified consignees by URGENT DEVICE RECALL letter dated October 10, 2011, via FedEx, email or telephone conference calls. The letter identified the affected product, the problem, and the actions to be taken. Customers were instructed to identify and return any affected product. Customers were requested to complete and return the attached Field Corrective Action Acknowledgement Form to Integra Surgical indicating receipt and review of this notification. For assistance and additional information contact your local and Integra Sales Representative or call Integra Surgical Customer Service at 1-800-431-1123. |  | 60 | 60 units | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of: Africa, Algeria, Australia, Austria, Brazil, Canada, Check Republic, China, Chile, Costa Rica, Denmark, Dominican Republic, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece Hungary, Guatemala, India, Ireland, Israel, Italy, Iraq, Japan, Jordan, Kenya, Kuwait, Lebanon, Lithuania, Macedonia, Mexico, Netherlands, New Zealand! Norway, Oman, Pakistan, Peru Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, Qatar, United Kingdom, Yemeni United Arab Emirates, Uruguay, and Venezuela. | Terminated | October 01, 2012 | 357 | 7216 | Not_Computer
Z-0486-2012 | 60357 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare AW Server version 2.0 with release 2.0 - 4.0 |  |  | November 17, 2011 | 2011 | GE Healthcare, LLC | 2 Unrelated Software Issues: An incorrect patient name maybe displayed when switching between two Volume Viewer tabs with different patients. Also, if a second patient exam is selected while the application is starting, the second patient's data will be displayed with the first patient's name display on the applications tab. | Software design | Consignees were sent on 11/17/11 a GE "Urgent Medical Device Correction" letters dated November, 16 2011 The letter was addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letter described the Safety Issue, Safety Instructions Affected Product Details, Product Correction and Contact Information. For the first issue, customers were asked to contact their field engineer to install the correct configuration file. Also, customers should not apply the recommended changes from the Customer Release Note as those changes cause the issue to occur. With the second issue, customers are to wait until the application has fully initialized before making another patient exam selection. GE Healthcare will correct all affected AW server installations via a software upgrade at no cost to the customer. A representative will contact customers to arrange for the correction. Questions should be directed to customers' local GE Healthcare Service Representative or at the numbers provided in the letter. |  | 308 | 308 | Worldwide Distribution -- USA, including states of AZ, CA, CO, CT, DE, DC, FL, HI, IL, IN, LA, MD, MI, MN, MS,MO, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, and WI and countries of UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, REUNION, QATAR, POLAND, NORWAY, NEW ZEALAND, LEBANON, SOUTH KOREA, JAPAN, ITALY, ISRAEL, HONG KONG, GUADELOUPE, GREAT BRITAIN, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CANADA, BRAZIL, BOSNIA AND HERZEGOVINA, BELGIUM, and AUSTRALIA. | Terminated | July 03, 2012 | 229 | 308 | Computer
Z-0365-2012 | 60117 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14 |  |  | October 11, 2011 | 2011 | Abbott Molecular | Abbott Molecular has received reports in which the liquid waste sensor and components in the m2000sp lower instrument cabinet overheated and/or caught fire. | Component design/selection | Abbott Molecular sent a Field Correction Recall letter dated October 10, 2011, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until further notice, and to complete the enclosed reply sheet and fax it to (847) 775-6728 at their earliest convenience to acknowledge receipt and understanding of the letter.For questions regarding this recall call 224-361-7133. |  | 432 | 432 units | Worldwide Distribution -- USA (nationwide) including the countries of Algeria, Australia, Austria, Azerbaijan, Belgium, Brazil, Burundi, Cameroon, China, Croatia, Estonia, Ethiopia, Finland, France, Gaza &amp; Jericho, Germany, India, Ireland, Israel, Italy, Kazakhstan, Kenya, Libya, Malawi, Malaysia, Mali, Mozambique, Netherlands, New Zealand, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, United Arab Emirates, Uganda, Ukraine, United Kingdom, Vietnam and Zambia | Terminated | February 23, 2016 | 1596 | 432 | Not_Computer
Z-0452-2012 | 59562 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC and Synchron LX Clinical Systems 100 ÂżL Sample Syringe, 100UL-PKGD, Part Number 474171 |  |  | September 02, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Modular Chemistry (MC) and Cartridge Chemistry (CC) 100 uL sample syringes may experience premature wear as indicated by brown discoloration on the plunger and/or inside the glass syringe barrel. Excessive wear of the 100 uL sample syringe barrel and plunger could lead to leaking and may affect results. | Component design/selection | Beckman Coulter sent an URGENT: Product Corrective Action letter to all affected customers. The letter provided customers with an explanation of the problem identified and actions to be taken. The letter instructs customers to share this information with their laboratory staff and retain a copy as part of their laboratory Quality System documentation. In addition, the letter states that if the affected product was further distributed to provide a copy of the letter to them. Customers were also asked to complete and return the enclosed Response Form. For any questions contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the United States and Canada. |  | 3278 | 3278 total (3083 in US, 195 in Canada) | Worldwide Distribution - USA (nationwide) and the country of Canada. | Terminated | June 30, 2014 | 1397 | 3278 | Not_Computer
Z-0418-2012 | 59739 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FOUNDATION HIP SYSTEM Series 460 Hip Stem STERILE R CEMENTED Sz 16.5mm |  |  | August 17, 2011 | 2011 | Encore Medical, Lp | The sterility has the potential to be compromised. | Packaging | Encore Medical Lp sent a Urgent Field Safety Notice dated August 17, 2011. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately quarantine these devices if you have not already done so. Contact Customer Service to obtain replacement devices as well as an RMA. Return all affected devices by October 31, 2011, using the RMA Customer Service provided in step #3.If you have any further questions you may call (512 ) 834-6330. For International Customer Service please call (512) 834-6275. |  | 169 | 169 units | Worldwide Distribution -- USA ( nationwide ) including the states of WI, NY, MO, TX, FL, CA, UT, AZ, Puerto Rico, and the country Japan. | Terminated | May 11, 2012 | 268 | 169 | Not_Computer
Z-0372-2012 | 60307 | No | BZT | Esophageal stethoscope with electrical conductors. | Anesthesiology | Anesthesiology | 510(k) | Cincinnati Sub Zero, Steri-Probe Temperature Probe, Esophageal/Rectal Temperature Probe, Part Number: 39622, Catalog #491B. TA disposable temperature probe used in conjunction with a measuring/monitoring device to continuously measure/monitor a patient's core body temperature. |  |  | September 26, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | Reports that the Rectal/ Esophageal Temperature Probes are not reading correctly. This situation causes the Blanketrol device, which the probes connect to, to shut down operation and display a 'check probe' error. | Nonconforming Material/Component | Cincinnati Sub-Zero Products, Inc., issued an "Urgent Medical Device Field Action" letter on November 1, 2011 to all affected customers. The letter includes; description of product and problem, and a list of the affected lot numbers. The letter instructs the customers to discontinue use of the device. All unused product is to be returned to the Biomedical/Clinical Engineering and/or Service Department for an replacement and all used probes are to be disposed of in accordance with hospital policy. The customers are also instructed to complete and return an attached Recall Response Form ASAP. For additional information, customers may contact the Beatrice Washington at 1-800-989-7373 or (513) 772-8810. |  | 369680 | 369,680 | Nationwide distribution and Canada. | Terminated | August 13, 2013 | 687 | 369680 | Not_Computer
Z-0506-2012 | 60581 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Walgreens Quick Read Digital Thermometer; an electronic digital thermometer; Model WIC551226 |  |  | November 23, 2011 | 2011 | Walgreen Company | The thermometers have been reported to have inaccurate readings | Labeling design | On November 23, 2011, Walgreens issued a recall notice electronically to their stores, instructing them to pull and quarantine the Thermometers and placing a Stop Scan on the products in the computer cash register system on the same date. Walgreens also removed the thermometer from their internet sales site. |  | 92000 | estimated 92,000 units | Nationwide through Walgreens retail stores and their internet site. | Terminated | August 14, 2017 | 2091 | 92000 | Not_Computer
Z-0391-2012 | 60371 | stocking, medical support (for general medical purposes) | FQL | Medical support stocking. | General Hospital | General Hospital | 510(K) Exempt | Truform Therapeutic Classic Medical Style Compression Stockings 20-30 mmHG, Truform Therapeutic Classic Medical Style Compression Stockings 30-40 mmHG, A Div of SAI, Cincinnati, OH 45209 |  |  | October 07, 2011 | 2011 | Surgical Appliance Industries | During an audit with the FDA, it was discovered that a new compression testing device used for testing hosiery had been implemented and that the operators were not properly trained. Additionally, the acceptance criteria for the new device were not changed from the old testing device to reflect additional data (compression profile percentages). This resulted in the device history files containing s | Employee error | The firm, Surgical Appliance Industries sent an URGENT: MEDICAL DEVICE RECALL letters dated November 7, 2011 to all affected customers. The letter described the products, problem and actions to be taken. Customers were instructed to immediately examine their inventory and quarantine all product(s) subject to recall for return and exchange. In addition, the letter states that if the product was further distributed to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Recall Response Form. For any questions contact Quality Systems Manager, at 513-271-4594 ext. 282. |  | 351896 | 351896 pairs | Worldwide Distribution - USA (nationwide) and the countries of: Australia, Canada, Colombia, Hong Kong, Peru, Singapore, Spain, and Venezuela. | Terminated | March 05, 2013 | 515 | 363715 | Not_Computer
Z-0463-2012 | 60375 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew 5.5mm DYONICS BONECUTTER ELECTROBLADE Resector |  |  | November 07, 2011 | 2011 | Smith &amp; Nephew, Inc. Endoscopy Division | Manufacturing issue: The distal tip of the outer sheath (return electrode) may detach and become loose in the joint during use. | Process control | Smith &amp; Nephew Inc. Endoscopy Division notified facilities on 11/7/11 via an "Urgent - Product Recall 1st Notification - Urgent" letter and/or phone. The letter identified the affected product and the reason for the recall. It also discussed the potential risks, customer actions, and instructions for the return of the product. Customers were to check their inventory and locate any of the affected product. They were to also complete the requested contact information and return the notification letter. In addition, customers were to contact Smith &amp; Nephew Endoscopy Division Returns Group at endo.andreturns@smith-nephew.com for information on how to return the affected product and receive replacement product. Questions, contact Hoangthi Le directly by phone at 508-337-3731. |  | 1743 | 1743 units | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Dubai, Finland, France, Germany, Norway, South Africa, Switzerland, United Kingdom, and Venezuela. | Terminated | February 22, 2016 | 1568 | 1743 | Not_Computer
Z-0603-2012 | 60390 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The GSP 2021-0010 (Genetic Screening Processor) instrument is a fully automated, high through put batch analyzer for time- resolved and prompt fluorescence analysis of samples in microtitration plates. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. The chemistries running on the GSP instrument are used to test blood specimens dried on filter paper as an aid in screening newborns for congenital disorders. |  |  | October 24, 2011 | 2011 | PerkinElmer LAS, Inc | Malfunction of the GSP disk remover unit may cause a punched disc to remain in the well during the measurement of the plate. This may cause a false negative screening result being reported for 17-OHP, resulting in failure to identify cases of congenital adrenal hyperplasia (CAH). | Device Design | Recall (correction) letters were sent to all customers on October 24, 2011. An improved disk remover design is under development to eliminate the reported problem. Until the final field correction is available, customers should follow the precautions provided. The needed protocol update will be performed by the local service engineer. If customers observe an increase in the occurrence of the problem, the laboratory should contact their local PerkinElmer representative. |  | 52 | 52 | Worldwide Distribution - USA, including the states of: Missouri, Florida, Texas, Oregon, Alabama, Louisiana, New Jersey, Georgia, Virginia, Kentucky and the countries of: Australia, Belgium, Brazil, Egypt, France, Germany, Italy, Netherlands, Norway, Qatar, Saudi MOH, Switzerland, and the UK. | Terminated | July 20, 2012 | 270 | 52 | Not_Computer
Z-0591-2012 | 60625 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa Computed Radiography Systems with DX-G Digitizer. |  |  | December 06, 2011 | 2011 | AGFA Corp. | Occasionally images can be lost or sent to an incorrect folder or the digitizer may stop either during startup or during a cassette scan cycle. | Software design | Agfa HealthCare sent a 'URGENT SAFETY NOTICE" letter dated December 6, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Agfa Service will contact your facility and schedule a software upgrade if this has not already been performed. Please complete the attached feedback form as soon as possible and fax it to 864-421-1662.If you have any questions about this matter, please contact your local Agfa HealthCare Service at 877-777-2432. |  | 125 | 125 | Worldwide Distribution-- USA (nationwide) including the states of AZ, CA, CO, FL, GA, HI, IL, IN, KS, LA, ME, MI, MN, MO, NE, NH, NJ, NY, OH, OK, PA, SD, TX, VA, WA, WI and WV, and Canada. | Terminated | April 11, 2012 | 127 | 125 | Computer
Z-0382-2012 | 60437 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System: Lung Optimized Treatment option. |  |  | November 07, 2011 | 2011 | Accuray Incorporated | User Facility reported an anomaly where inhale and exhale CT pairs used for treatment planning did not represent the same magnitude of respiration that was displayed during treatment delivery. | Labeling design | Accuray sent a Urgent Medical Device Correction letter dated November 7, 2011, via Federal Express to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The information provided in this letter will be incorporated in user manuals shipped with the CyberKnife system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at 1. (877) 668-8667 ( USA) 1. (408) 716-4700 ( non USA ) |  | 30 | 30 | Worldwide Distribution -- USA ( nationwide ) including the states of FL, TN, IL, NJ, CO, PA, MD, CA, MS, MN, SC, WA, OK, GA, AZ, and TX., and the countries of France and United Kingdom. | Terminated | February 27, 2012 | 112 | 30 | Not_Computer
Z-0387-2012 | 59812 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | ActiveCare DVT |  |  | August 28, 2011 | 2011 | Medical Compression Systems | This voluntary field action is being undertaken because MCS has received a limited number of reports involving the use of a non-MCS AC/DC adaptor with the ActiveCare DVT and ActiveCare+S.F.T. devices. The firm found that an ActiveCare DVT device may work improperly if connected to a non-MCS AC/DC adaptor with a higher nominal voltage. As specified in each product's User Manual, the ActiveCare DVT | Device Design | Medical Compression Systems (DBN) Ltd, sent out "VOLUNTARY FIELD ACTION" letters dated September 15, 2011 to all affected customers. The letter included description of the device and problem. Customers were instructed to check their inventory and make a list of serial numbers of all applicable ActiveCare+S.F.T. and ActiveCare DVT devices in their inventory, include the list on the attached excel file (Field Action Verification Form) and email to maya@mcsmed.com; even if they do not have any such devices. The letter also states that within 15 days from the email, they will receive instructions for the next steps. In addition, the letter stated -YOU MAY CONTINUE TO USE THE ActiveCare+S.F.T.Âż DEVICES WHILE AWAITINGTHE INSTALLATION OF THE OVP.In case you have questions or you need any assistance, you can contact the firm at (317) 485-7887 or e-mail to: maya@mcsmed.com or mahlet@mcsmed.com. |  | 2217 | 2217 units | Nationwide distribution. | Terminated | July 11, 2012 | 318 | 2768 | Not_Computer
Z-0513-2012 | 60643 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Dimension Vista(R) Cardiac Troponin I Calibrator (KC678) lot 1DD085, in vitro diagnostic |  |  | November 17, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed positive bias on QC and Patients. | Nonconforming Material/Component | Siemens healthcare Diagnostics initiated their recall of lot 1DD085 on November 17, 2011 by sending an Urgent Field Safety Notice letter to all affected customers. Customers were requested to immediately discontinue use and discard any remaining inventory of the CTNI CAL lot 1DD085. |  | 2281 | 2281 cartons | Worldwide distribution, including USA, Austria, Belgium, Canada, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, New Zealand, Portugal, Puerto Rico, Slovenia, South Korea, Spain and Switzerland. | Terminated | August 26, 2014 | 1013 | 2281 | Not_Computer
Z-0473-2012 | 60608 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Y-Type Solution Set, 78" (2.0 m), 2 Injection Sites, Male Luer Lock Adapter; an Rx sterile, nonpyrogenic fluid pathway for the administration of sterile I.V. fluids to the patient; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 1C8581 |  |  | December 06, 2011 | 2011 | Baxter Healthcare Corp. | During production of the IV sets, solvent was inadvertently applied to the locking collar of the male luer connector, which may have damaged the component, preventing proper engagement of the locking collar to secure a connection. | Employee error | The firm, Baxter, sent an "Urgent Product Recall" letter dated December 6, 2011, via first class mail to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to inspect all locations where their facility's inventory is located for products with the affected lot numbers; return the affected product; request a credit or replacement by calling Baxter Healthcare Center for Service at 1-888-229-0001 between 7:00am to 6:00pm Central time, Monday through Friday; complete and return the attached Customer Reply Form via fax to: 1-847-270-5457, or scan and email to fca@baxter.com; and if they are a dealer, wholesaler, or distributor/reseller that distributed any affected product to other facilities, please notify their customers of this recall.If you have any questions regarding the communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 2352 | 2,352 units | Worldwide distribution: USA (nationwide) including states of: Alaska, California, Colorado, Florida, Georgia, Hawaii, Illinois, Kansas, Kentucky, Massachusetts, Maine, Michigan, Minnesota, Missouri, North Carolina, New Hampshire, New Jersey, New York, Ohio, Pennsylvania, Rhode Island, South Carolina, Tennessee, Virginia, Washington, Wisconsin and West Virginia; and country of: New Zealand. | Terminated | November 01, 2012 | 331 | 48480 | Not_Computer
Z-0464-2012 | 60568 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | CD HORIZON (R) Spinal System, MULTI-AXIAL SCREW FOR 5.5mm ROD W/DUAL LEAD THREAD, Size 6.5mm X 40mm, Mat'l TITANIUM ALLOY, REF X0809439, Rx only, Medtronic Sofamor Danek USA, 1800 Pyramid Place, Memphis, TN 38132 |  |  | November 30, 2011 | 2011 | Medtronic Sofamor Danek USA Inc | One screw may be longer then size etched on the screw. | Mixed-up of materials/components | Medtronic Sofamor Danek USA, Inc sent a Voluntary Field Correction letter dated November 30, 2011, via FedEx to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The firm will visit each consignee to perform a visual check of inventory.Please utilize the attached questionnare to account for the affected product and to record any patient injuries that may have occurred specifically related to the use of these items. If you need assistance, or if you have any related questions or concerns please call (800) 876-3133 extension 3173 |  | 252 | 252 units | Nationwide Distribution including the states of AL, PA, IN, CO, MI and IA. | Terminated | April 05, 2012 | 127 | 282 | Not_Computer
Z-0583-2012 | 60258 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | INOmax DS (Delivery System) Model 10003 &amp; 10007. |  |  | November 30, 2010 | 2010 | Ino Therapeutics | Through the review and monitoring of its complaint data, Ikaria identified a product anomaly with the INOmax DS Model 1003. Ikaria concluded the anomaly risk to the patient safety was low and non reportable (4/23/10). During a subsequent Level 3 inspection by FDA, it was concluded that the action should have been reported under 21 CFR &amp; 806.10. | Nonconforming Material/Component | Ikaria sent a "RECALL- CUSTOMER PRODUCT REMEDIATION" notice dated December 22, 2011 to all affected customers. The letter describes the product, problem, actions taken by the firm. No action was required by customers. The notice served as correspondence to inform customers of the issue. Customers may direct their questions to Ikaria Customer Care at 1-877-566-9466. |  | 2446 | 2446 | Worldwide Distribution-USA (nationwide) and the country of MALAYSIA. | Terminated | May 29, 2012 | 546 | 2446 | Not_Computer
Z-0460-2012 | 60051 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | SRS FAST SET PUTTY 3 CC BONE VOID FILLER-STERILE, Part Number SRS-005-FSP, Lot number N001101. |  |  | September 29, 2011 | 2011 | Synthes USA | A recent shipment carrying Norian SRS was lost by the carrier. Since the product was lost before final product release testing, Synthes is unable to confirm that Part Number SRS-005-FSP- Lot Number N001101 meets final product release requirements. The product contains all final labeling and cannot be visually distinguished from the product that has met final product release requirements. | Labeling Change Control | All Synthes accounts with a history of purchasing the product in the past two years will be notified by letter of the incident. All firm sales consultants will also receive written notification of the incident. |  | 131 | 131 units | Synthes USA sent a "NOTICE: MISSING SHIPMENT-NORIAN SRS" letter dated September 29, 2011 to all accounts with a history of purchasing the product in the past two years. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory to verify they have not received the affected product. Call 1-800-620-7025 ext. 5450 for questions concerning this notice. | Terminated | December 18, 2014 | 1176 | 131 | Not_Computer
Z-0598-2012 | 60521 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Infusion Pump, Single Channel Volumetric Infusion Pumps. Product Code: 2M8161. |  |  | May 08, 2009 | 2009 | Baxter Healthcare Corp. | There was no documentation to show the required tests were performed following the replacement of the battery and battery harness. | Process control | Baxter sent a service technician to the consignee location on 05/08/2009 to perform the battery check and discharge test. The pump was found to be within specification. |  | 1 | one pump | Puerto Rico | Terminated | January 03, 2012 | 970 | 1 | Not_Computer
Z-0421-2012 | 60454 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | MAXLOCK EXTREME EXTREMITY PLATING SYSTEM WITH VARIABLE ANGLE TECHNOLOGY. |  |  | August 22, 2011 | 2011 | Orthohelix Surgical Designs Inc | The firm discovered a testing protocol error, in both dynamic and static cantilever comparative testing, during the initial comparative testing. | Nonconforming Material/Component | Orthohelix Surgical Designs Inc. sent an URGENT RECALL NOTIFICATION letter dated August 22, 2011 to all affected customers. The letter identified the affected products, the problem, and the actions to be taken. Customers were instructed to immediately inventory, remove and return the recalled product to the firm via FedEx or any shipping method with tracking capability. Customers were also requested to notify any affected user facilities of the recall. For any questions call 330-247-1444 or 330-247-1445. |  | 2722 | 2,722 units | Nationwide Distribution (USA) - including the states of: CA, CO, CT, FL, GA, IA, IL, KY, LA, MD, MI, MO, NC, NJ, NM, NY, OH, OK, PA, SD, TN, TX, UT, VA, WA, WI and WY. | Terminated | February 19, 2012 | 181 | 2722 | Not_Computer
Z-0517-2012 | 60583 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Rusch Flexi-slip Endotracheal Tube Stylet with Soft Distal Tip, 6FR, Catalog numbers: 502501, Teleflex Medical Sdn Bhd, P.O. Box, Industrial Estate, 24600 Kamunting, Perak, Malaysia. |  |  | December 02, 2011 | 2011 | Teleflex Medical | Complaints of difficulty removing stylet from the endotracheal tube and possible PVC sheath separation were received. | Pending | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall Notification" letter dated December 2, 2011 to its Consignees/Customers. The letter described the product, problem and actions to be take. The customers were instructed to immediately discontinue use and quarantine any products with the listed catalog and lot numbers; return the product; complete and return the enclosed Immediate Attention Requested Recall Acknowledgement Form via fax to 1-866-804-9881, Attn: Customer Service (a customer service representative will contact the customer with a returns good authorization number and provide instructions for the return of product to Teleflex medical), even if they have no affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 73281 | 73,281 units | Worldwide distribution: USA (nationwide) and countries of: Canada and Europe. | Terminated | April 29, 2014 | 879 | 73281 | Not_Computer
Z-0485-2012 | 60528 | No | RER |  |  |  |  | Raymond Laser alignment guide - Model #1096597/002 Red Laser. |  |  | November 28, 2011 | 2011 | The Raymond Corporation | Raymond shipped a laser alignment guide as an option on its powered industrial trucks from February 2, 2010 through April 21, 2011, but was not aware of its obligation under 21 C.F.R. 1010.2 and 1010.3 to certify and label. | Radiation Control for Health and Safety Act | Raymond will be initiating a Field Campaign to ensure that all customers with Alignment Guides have the confirmatory laser caution decals and manufacturer label, an updated installation and service manual, the operator instructions, and informational brochure. This campaign will be performed at no cost to Raymond's dealers and customers, and will include documenting the laser serial number for each truck configured with an Alignment Guide for product traceability. |  | 493 | 493 total | Nationwide Distribution | Open, Classified |  |  | 493 | Not_Computer
Z-0521-2012 | 60519 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Triage Drugs of Abuse Panel Plus TCA 25 Test Kit, PN (Model Number) 92000 |  |  | July 01, 2011 | 2011 | Alere San Diego | The recall was initiated because Alere San Diego has confirmed false positive THC results obtained on Triage Drugs of Abuse Plus TCA 25 devices from lots 235991 and 238928. | Nonconforming Material/Component | A recall communication was initiated on July 1, 2011 with Alere San Diego forwarding a Product Notification and a Verification Form (via fax, e-mail or direct mail) to all their customers who purchased the Triage Drugs of Abuse Plus TCA 25 Test Kit, PN 9200, Lots 235991 and 238928. The letter provides the customers with an explanation of the problem identified and an action to be taken. The firm reminds customers that clinical consideration and professional judgment must be applied to any test result. Customers were instructed to retain the Product Notification letter for their records as documentation that they have received the information and to ensure that all users have a copy. Alere will replace the lots at the customer's request. As the customers complete the fax back form, they were instructed to indicate whether they wish to replace the affected product. Customers were instructed to complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice. If customers have any further questions or need to report any concerns they have about the Triage Drugs of Abuse Plus TCA performance, they were instructed to contact Technical Services at (877) 308-8287 or by e-mail at techservice@alere.com. |  | 997 | 997 total kits | Worldwide Distribution -- US and to Canada. | Terminated | April 04, 2012 | 278 | 997 | Not_Computer
Z-0594-2012 | 60549 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning System, software version 9.0. |  |  | November 18, 2011 | 2011 | Philips Medical Systems (Cleveland), Inc. | Phillips Medical Systems have recently determined that a software (Pinnacle3 Software version 9.0) nonconformance can cause incorrect beam geometry. | Software design | Consignees were sent on 11/18/11 a Philips "Urgent Medical Device Correction" letter dated 2011 Nov 16. The letter identified the affected product and described the Problem, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer/User, Actions Planned by Philips and Further information and Support. If customers observed the issue, they were to manually make the suggested change. In addition, customers were to place the letter in the Instructions for Use until otherwise notified. The firm will deliver 9.2 software release free of charge. Customers in North America are to contact Customer Care Solutions Center at 1-800-722-9377, if they have questions or need support concerning the recall. |  | 715 | 715 (463 US) | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and countries of AUSTRALIA, AUSTRIA, BELGIUM, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, MALAYSIA, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, and CANADA. | Terminated | February 28, 2014 | 833 | 715 | Computer
Z-0511-2012 | 60476 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Urine Reagent Strips (URS) 10 Parameters |  |  | October 14, 2011 | 2011 | Teco Diagnostics | This recall was initiated in response to findings during a FDA inspection of the firm. The firm's recall has been initiated due to improper manufacture of sub-component nitrite paper Lot 35134, Expiration Date: 01/2014. Discrepancies were found on the information recorded on formulation sheet and manufacturing log sheets (Refer to next paragraph). Sampling size submitted for QC testing was insuff | Nonconforming Material/Component | Teco Diagnostics sent an "URGENT: MEDICAL DEVICE RECALL" letter dated October 14, 2011 to affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructed customers to examine their inventory for the affected product and to quarantine any product found. A Recall Return Response Form was included with the letter for customers to complete and return. Contact the firm's Technical Support Department at (800) 222-9880 for questions regarding this recall. |  | 550 | 550 kits | Distributed in CA and Poland. | Terminated | April 09, 2013 | 543 | 550 | Not_Computer
Z-0416-2012 | 60263 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Grafton DBM A-Flex; Unit Size 60 mm diam; 1ea |  |  | August 08, 2011 | 2011 | Osteotech Inc | Aseptic conditions during the production of these products may have been compromised. | Material/Component Contamination | Osteotech sent a recall letter dated August 8, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all units of the affected product. Customers were asked to contact Osteotech Customer Service at 732-578-6699 within two business days of receiving the letter to acknowledge receipt of the letter and to arrange for a return and exchange or credit for the product.For questions regarding this recall call 732-542-2800. |  | 18 | 18 unts | Worldwide Distribution | Terminated | December 08, 2011 | 122 | 33 | Not_Computer
Z-0516-2012 | 60684 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Portal Vision, R-Arm with 4D Integrated Treatment Console; |  |  | November 23, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | An anomaly was discovered when using the 4DITC in combination with an R-Arm to acquire double exposure MV images. | Software design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION/URGENT FIELD SAFETY NOTICE" dated December 1, 2011 to all affected sites via Certified mail, return receipt. Canadian accounts were sent by Fed Ex and internationally to Varian International Customer Support Managers for translation and notification of customers. The notice described the product, problem and actions to be taken. The customers were instructed to always acquire both the treatment port and the open port images BEFORE displaying the digital graticule, when using a standard double exposure template; and to advise the appropriate personnel working in their radiotherapy department of the content of this letter.Special instructions for customers outside the USA and Canada: these customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems, once they have read this document. Varian is developing a correction for this issue and a Customer Service rep will contact the customers when the correction is available to schedule its installation on their systems.If you have any questions, contact USA and Canada at 1-888-827-4265; Europe: +41 41 749 8844; Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: China.apps.helpdesk@varian.com; Japan: Japan-Apps.Helpdesk@varian.com and Latin America: soporte.al@varian.com or Internet: Oncology systems customer site: www.myvarian.com and Varian Medical Systems public site: www.varian.com. |  | 1486 | 1486 devices | Worldwide distribution: USA (nationwide) and countries of: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France,, Georgia, Germany,Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Lativia, Macedonia, Malaysia, Martinique, Mexico, Montenegro, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | August 06, 2014 | 987 | 1486 | Not_Computer
Z-0510-2012 | 60148 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | AVEA Ventilator (Catalog Code 17310, 17311, 17312). |  |  | September 08, 2011 | 2011 | Carefusion 211 Inc | The recall was initiated because Carefusion has identified a potential risk associated with certain AVEA ventilators and affected replacement parts manufactured between March 1, 2009 and June 30, 2011. Carefusion is voluntarily initiating a field correction of the affected devices to preclude the possibility of this risk. | Nonconforming Material/Component | CareFusion sent an "URGENT PRODUCT RECALL" letter dated September 25, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers were instructed to contact the firm by telephone to coordinate implementation of the corrective action at their site. Customers were instructed to return an enclosed response card to expedite the correction process and acknowledge receipt of the recall notification. Carefusion does not require that the customers return their devices. Customers will be contacted by a member of the CareFusion Technical Support Department to arrange for onsite remediation of the affected devices. In the interim, if any affected AVEA ventilator units in their facility exhibits a sustained Extended High Ppeak alarm followed by the opening of the Safety Valve, customers were instructed to remove the ventilator from service, provide alternate ventilation and contact Carefusion Technical Support (800) 213-2466 to report the issue. All ventilator dependent patients should be constantly monitored by qualified personnel to ensure that if a malfunction were to occur, alternate ventilation can be provided. |  | 6797 | 6797 total, 1891 in US, 4906 worldwide | Worldwide Distribution-USA (nationwide) and the countries of Cyprus, Estonia, Brazil, Kuwait, United Arab Emirates, Slovenia, Lithuania, Latvia, Belgium, Bermuda, Ecuador, Guatemala, Italy, Canada, Australia, Netherlands, Germany, United Kingdom, Spain, Czech (Rep), Nicaragua, China, Bangladesh, Morocco, Bolivia, Vietnam, Slovakia, Paraguay, Argentina, Turkey, Chile, Singapore, Japan, Colombia, Egypt, Hong Kong, Honduras, India, Croatia, Venezuela, Saudi Arabia, Greece, Hungary, Mexico, Austria, Ireland,Indonesia, Dominican Republic, Portugal, South Africa, Panama, Philippines, Libyan Jamahiriya, Jordan, Malaysia, France, Thailand, Peru, Costa Rica, Korea, and Taiwan. | Terminated | October 29, 2013 | 782 | 6797 | Not_Computer
Z-0592-2012 | 60524 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Orbit and Orbit Micro |  |  | March 01, 2011 | 2011 | ICU Medical, Inc. | The recall was initiated because ICU Medical Inc. has become aware of the market recall of the Triad Group's alcohol prep pads, alcohol swabs, and alcohol swabsticks manufactured by Triad in the United States and marketed under various brand names. The Triad alcohol prep pads have the potential contamination of the bacteria, Bacillus cereus. The Triad Group alcohol prep pads have been copackaged | Pending | The firm, ICU Medical, Inc., sent an "IMPORTANT-Important Information About Triad Group's Alcohol Prep Pads" letter dated March 2, 2011 (via email) to their distributors. Hard copy notices were shipped (via next day express delivery with receipt confirmation) to all their customers. The Important letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were informed that the Orbit and Orbit Micro subcutaneous infusion sets as well as the IV start kits are not contaminated and may continue to be used according to the directions for use. The issue is confined to the actual Triad alcohol prep pad products. Patients and healthcare providers should not use the alcohol prep pads packaged with these devices and should instead use an alternate alcohol prep pad that is not involved with the Triad Group recall. ICU Medical has halted all shipments of product that contain the Triad alcohol prep products to its distribution network. An alternate wipe has been qualified and is now being shipped in new orders. Customers were instructed to forward the Important letter to their customers and on to any end users of the devices. For Orbit and Orbit Micro devices, patients are instructed to immediately discontinue using the Triad alcohol prep pads included in the Orbit packaging and dispose of those pads in the trash. Patients should use an alternative prep pad that is not subject to the Triad recall. For IV Start Kits, distributors and users are instructed to place the enclosed label on the device prior to use. Additional labels can be obtained by contacting ICU Customer Service Monday through Friday between the hours of 8:30AM and 4:00PM Pacific Standard time: (800) 866-9025 and select option 8, or email the following address: customerservice@icumed.com. Customers may also contact regulatory@icumed.com for MSWord template file for Avery 5160 labels and print additional labels. Customers were instructed to complete |  | 376290 | 376,290 units | Worldwide distribution: USA (Nationwide) and countries of: Austria, Belgium, Chile, Cyprus, CZ, France, Germany, Great Britain, Greece, Israel, Italy, Lithuania, Netherland, Poland, Russia, Saudi Arabia, Sweden, andSwitzerland. | Terminated | February 22, 2013 | 724 | 957990 | Not_Computer
Z-0585-2012 | 60132 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery XR650, Digital Radiographic Systems. |  |  | November 22, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the patient barrier rotating arm that is a part of the Definium 8000 and Discovery XR650 image pasting systems that may impact patient and user safety. The patient barrier rotating arms that are used to position and stabilize the patient during image pasting procedures may fail to retain residual friction at the arm pivot. Thi | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 17, 2011 to its customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to ensure that all potential users in their facility are made aware of this safety notification and the recommended actions; discontinue use of the patient barrier and remove from service (away from patients) if either one the rotating arm locking levers is damaged (does not hold the rotating arm in place when engaged); take extra precautions to hold the patients barrier rotating arm during release of the locking lever and repositioning of the arm and store the rotating arms in the lowest pivot position (down) when not in use. The letter also stated that the patient barrier may continue to be used if the rotating arm locking lever securely holds the rotating arm in place when it is engaged.Additionally, GE Healthcare will modify the rotating arms to restore the residual friction so that the arm will retain position (rather than swing freely) while the locking levers are disengaged. Also, GE will update the product labeling to improve awareness and identify the hazard relative to proper storage.If you have any questions, call Center in US: 800-437-1171; Japan: 0120-055-919 and for other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 229 | 229 | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID,IL, IN, IA, KS, KY, LA,ME, MD, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY,NC, OH, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, PORTUGAL, NORWAY, NEW ZEALAND, NETHERLANDS, MARTINIQUE, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, HONG KONG, GREECE, GERMANY, FRENCH POLYNESIA, FRANCE, FINLAND, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BOSNIA AND Herzegovina, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | October 16, 2012 | 329 | 860 | Not_Computer
Z-0430-2012 | 60134 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Anthem CRT-P models PM3110 and PM3210 |  |  | September 23, 2011 | 2011 | St Jude Medical CRMD | St. Jude Medical has received reports of out of range pacing lead impedance (PLI) measurements. These measurements are low and out of range (&lt;200 ohms) but measure in range during visits to the clinic. The out of range readings are concluded as false. | Pending | The firm, St. Jude Medical, sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated September 22, 2011 to its customers by FedEx priority overnight service for planned receipt on September 23, 2011. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed that in order to prevent a false reading, a new programmer software version (12.1.0.1) is available. This software when used to interrogate an Accent DR or Anthem CRT-P pacemaker will eliminate the potential for this anomaly to occur. The onetime upgrade is performed automatically on affected devices and will not change the operation of the implanted device. A St. Jude Medical Sales Representative will assist the customers in loading the new programmer software onto their Merlin programmer. Customers with questions or concerns, can contact their local St. Jude Medical representative or their Technical services department at (800) 722-3774. |  | 14000 | 14,000 units | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Australia, Asia Pacific, Belgium, Bulgaria, Canada, China, Colombia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Hungary, India, Iran, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latin America, Lebanon, Malaysia, Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, UK, and U.A.E. | Terminated | April 10, 2012 | 200 | 166000 | Computer
Z-0607-2012 | 60660 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Imaging Systems, Model 882050; |  |  | December 08, 2011 | 2011 | Philips Medical Systems | During preventative maintenance, a crack was discovered in the lower portion of the old style, non-common arm for the detector assembly, which could result in the detector dropping in an uncontrolled manner. | Device Design | Urgent Medical Device Correction/Field Safety Notices, dated Dec 08, 2011, were mailed to affected customer sites on December 9th and 12th, 2011 by Return Receipt certified mail. The letter identified the affected product, described the issue and when it could occur, along with the hazards involved. Customers were advised to stop using the system and contact a Philips service representative for an immediate inspection. Customers were also to follow the Instructions For Use to avoid the issue. Philips service will contact customers to arrange a time for Philips to inspect the system. The letter should be placed in the Instructions for Use. The Field Change Order will be released on December 16, 2011. If further information or support is needed, customers should contact Customer Care Solutions Center at 1-800-722-9377. |  | 129 | 129 | Worldwide Distribution -- USA, Canada, Italy, Spain, France, United Kingdom, Sweden, Switzerland, Germany, Japan, Denmark, and the Netherlands. | Terminated | August 13, 2012 | 249 | 129 | Not_Computer
Z-0399-2012 | 60230 | No | KRZ | Calibrator for hemoglobin or hematocrit measurement. | Hematology | Hematology | 510(K) Exempt | VARIANT II Link Hemoglobin A1c Program Reorder Pack, |  |  | October 25, 2011 | 2011 | Bio-Rad Laboratories, Inc. | It was discovered that the Variant II Turbo Link Calibrator Values card contains incorrect values for IFCC units. The numbers are missing a decimal point. | Process control | Bio-Rad Laboratories, Inc. sent a Medical Device Correction Notification letter dated October 25, 2011, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The correction was initiated by the firm, customers were instructed to contact their regional Bio-Rad office for any questions.For questions regarding this recall call 510-741-6114. |  | 63 | 63 | Nationwide Distribution including MI, SC, OH, IL, MN, MD, VA, NY and WI. | Terminated | February 02, 2012 | 100 | 63 | Not_Computer
Z-0613-2012 | 60049 | [6 empty] | OPUS SmartStitch M-Connector Product Number OM-8007 is a suture passer |  |  | August 08, 2011 | 2011 | Innovatech Medical Resources L.P. | Product is being removed from the field due to an over-label issue. | Reprocessing Controls | Innovatech Medical Resources L.P. decided to recall all products resterilized and overlabeled at their facility. A firm representative would visit each customer and request that they remove the overlabeled product from their shelves and secure it. Customers could call the MSI office to get a RMA (Return Materials Authorization) for a refund on this product.For questions regarding this recall call 817-714-9066. |  | 113086 | Lot Number (QTY): 113086 (1 unit), 107304 (1 unit), 110511 (1 unit), and 112473 (2 units) | Nationwide Distribution including Texas | Terminated | September 27, 2012 | 416 | 99641488 | Not_Computer
Z-0467-2012 | 60590 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | MORPHEUS CT PICC with STYLET, 5F Dual Lumen 65 cm Insertion Kit, Catalog No./REF 1200923, STERILE, AngioDynamics Queensbury, NY 12804 USA --- COMMON/USUAL NAME: Percutaneous, implanted, long-term intravascular catheter; CLASSIFICATION NAME: Percutaneous, implanted, long-term intravascular catheter; Device Listing # D095274. |  |  | December 05, 2011 | 2011 | Angiodynamics Worldwide Headquarters | AngioDynamics distributed the affected PICC catheters that were beyond their expiration date. | Incorrect or no expiration date | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated December 2, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. A Recall Reply Form was attached to the letter for customers to complete and return to the firm via fax at 518-798-1360. Contact your local sales representative or AngioDynamics Customer Service at 800-772-6446 for questions regarding this notice. |  | 16 | 16 units domestically | Nationwide Distribution-including the states of Florida, Georgia, Montana, and Tennessee. | Terminated | September 27, 2012 | 297 | 16 | Not_Computer
Z-0438-2012 | 60138 | No | NRA |  |  | Orthopedic | PMA | "***BUECHEL-PAPPAS CRUCIATE SACRIFICING BEARING ***MATERIAL UHMWPe***STERILE EO***ENDOTEC***Endotec, Inc. 300 Sunport Lane, Suite 500 Orlando, FL 32809 USA***CAT.NO.***SIZE***" |  |  | September 21, 2011 | 2011 | Endotec, Inc. | On 09/21/2011 Endotec, Orlando, FL initiated a recall for the Beuchel-Pappas Mobile Bearing Knee System (BP Knee). The BP Mobile Bearing Knee System consists of unapproved mobile (rotating) bearing tibial and patellar components. | PMA | Endotec sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 21, 2011 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. The letter instructs customers to examine and quarantine product subject to the recall. Additionally, a Response Form was included with the letter for customers to complete and return to the firm. Call 862-703-6730 for questions regarding this recall. |  | 60 | 60 | USA (nationwide) including the states of AZ, CA, CT, FL, NE, NJ, OH, and OK. | Terminated | September 27, 2013 | 737 | 2400 | Not_Computer
Z-0514-2012 | 60425 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS, all versions |  |  | September 20, 2011 | 2011 | Siemens Medical Solutions USA, Inc | Unintended behavior when using syngo Imaging XS, all versions. The "Rename" functionality can cause an unintended patient merge if the renamed patient is sent to another system. | Labeling False and Misleading | Siemens Medical Solutions USA, Inc sent a Customer Safety Advisory Notice letter dated August 3, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you detect patient data inconsistencies in your system caused by this problem, please contact Siemens Service for support.For further questions please call (610) 219-4834. |  | 38 | 38 | Nationwide Distribution including the states of CA, FL, GA, IN, IA, KS, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | March 05, 2014 | 897 | 38 | Computer
Z-0448-2012 | 60387 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Microcyn Solution with Preservatives (Rx), 16.9 fl oz (500 mL) and 33.5 fl oz. (990 mL) |  |  | November 09, 2011 | 2011 | Oculus Innovative Sciences Inc | Unapproved drug claims and those not approved under 510K clearance. | Process control | Oculus Innovative Sciences sent an"URGENT MEDICAL DEVICE RECALL" letter dated November 15, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter instructs customers to identify and remove all affected product from inventory. A Business Reply Card was attached to the letter for customers to complete and return via fax to 1-415-462-5181. Call Oculus Innovative Sciences Customer Service at 1-800-759-9305 for questions regarding this recall. |  | 4890 | 1363 x 500 mL; 2037 x 990 mL. | Nationwide Distribution. | Terminated | December 07, 2012 | 394 | 19918 | Not_Computer
Z-0414-2012 | 60467 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VenaCure EVLTÂż NeverTouch-FRS PROCEDURE KIT Featuring Gold-Tipped Fiber, 65 cm, VC NT FRS 65cm Clear Non-Echo Needle, Catalog No./REF 51403005, STERILE -- AngioDynamics, Inc. 603 Queensbury, NY USA 12804. This is a Special Order. |  |  | November 10, 2011 | 2011 | Angiodynamics Worldwide Headquarters | During an in-process inspection the hemostasis sidearm assemblies were identified to have the potential for flash to exist on the housing at the base of the threads. This flash has the potential to break off and relocate within the fluid path. | Process control | AngioDynamics sent an URGENT: Medical Device Recall letters and response form (dated November 10, 2011) to all affected customers. The letter identified the affected products, the problem, and the actions to be taken by customer/user. The customers were instructed to do the following: 1) Identify, segregate and return the recalled products to ANGIODYNAMICS and 2) Complete the enclosed Recall Reply Form even if they do not have any products remaining in their possession, noting zero quantity to be returned and fax to 518-798-1360. The letter states that a returned authorization number (RA#) will be assigned to each customer in the recall notification letter, an UPS account will be provided for return freight and a replacement product will be shipped upon receipt and confirmation of the returned product. If you have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please do not hesitate to contact your local sales representative or call the Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. You may also e-mail your questions to customerservice@angiodynamics.com. |  | 105 | 105 units | Worldwide Distribution - USA (nationwide) and the countries of: Canada, Great Britain, France, India, New Zealand, Netherlands, Saudi Arabia, Turkey, and Singapore | Terminated | February 01, 2016 | 1544 | 12205 | Not_Computer
Z-0605-2012 | 60658 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | ETCO2 Divided Sampling Cannula with Simultaneous Oxygen Delivery, Model number 4707 7-7-24. |  |  | December 07, 2011 | 2011 | SALTER LABS | Mis-labeling: report of cases labeled as 4707 (cannulas with male luer lock connector) actually contained 4707F (with female luer lock connectors). | Process control | The firm notified the three affected customers directly by phone, followed by a recall form for each direct customer indicating the product, part number, lot number and shipment date and customer order specific for each firm. Customers were to complete and return the form to the fax number listed or call a Customer Service Representative at 1-800-628-4690. |  | 240 | 240 cases | Nationwide Distribution -- TX and VA. | Terminated | January 09, 2012 | 33 | 240 | Not_Computer
Z-0366-2012 | 60419 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer ZMR Femoral Spout &amp; Cone Bodies and Trabecular Metal Primary Hip Prosthesis Femoral Stem |  |  | November 16, 2011 | 2011 | Zimmer Inc. | Lack of assurance of proper metal fatigue strength due to metal grain structure anomaly. In 2006 a recall was initiated as a result of the grain anomaly; however, these 4 lots had been split off during normal production and had not been included in the 2006 recall. | Nonconforming Material/Component | On 11/16/11, Zimmer sent letters titled "URGENT MEDICAL DEVICE RECALL" to distributors via electronic mail and to risk managers and surgeons via FedEx indicating that the product must be immediately removed from use and quarantined." Stop using the affected lots and place affected product in quarantine." Review available product on the shelf" Identify those part/lot combinations listed in the recall communication" Immediately remove any implants found that match part/lot combinations listed in the recall communication and quarantine" Return affected lots and signature of acknowledgement to Zimmer |  | 2 | 2 (1 device is domestic) | Worldwide distribution: USA (NY), AUSTRALIA, SWITZERLAND, and DENMARK | Terminated | August 28, 2013 | 651 | 4 | Not_Computer
t | o | o |  | b | i | g | [18 empty] | Not_Computer
Z-0832-2012 | 60615 | No | NZG | Pediatric medical crib. | General Hospital | General Hospital | 510(K) Exempt | Bassinet Models: BSNT/01; BSNT/02, BSNT/03, BSNT/04 |  |  | October 14, 2011 | 2011 | Nemschoff Chairs, Inc. | To correct issues with casters, door hinges, drawer slides, and the bassinet surround. Also, to add labeling to the product regarding specific instructions for cleaning the bassinet tub and drawer weight limits. | Device Design | The firm, Nemschoff, sent an "Important Medical Device Correction" letter and email dated October 14, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to check inventory, quarantine and hold product until they receive instructions for servicing the units; once corrections are complete on the units at their locations, document all corrections on the Correction Acknowledgment Letter and return to Nemschoff, Inc. as instructed. The customers were also instructed to complete and return the enclosed Receipt Acknowledgment Form by October 21, 2011.Nemschoff will provide instructions and assistance with locating and repairing the units.If you have any questions, call Nemschoff, Inc at (800) 203-8916 and ask for our Quality Representative or your Customer Service Representative. |  | 1214 | 1214 units. | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IA, IL, IN, KS, LA,MA, MD, MI, MO, MS, NC, ND, NJ, NY, OH, OK, SC, TX, VA, WA, WI, and WY; and country of: Canada. | Terminated | November 27, 2012 | 410 | 1214 | Not_Computer
Z-0935-2012 | 60842 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | FloWire Doppler Guide wires. |  |  | October 06, 2009 | 2009 | Volcano Corporation | Reports of no signal displayed on the monitor when FloWire Doppler Guide wires were connected to the Combomap. | Employee error | Volcano Corporation sent recall letters dated October 6, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and fax the attached Customer Reply Letter so that an RMA could be issued for the return of the unused affected product, and any replacement or credit could be arranged. Customers were instructed to return the Customer Reply letter even if they no longer have the affected product in their inventory. For questions regarding the use of the product. customers were instructed to contact their Sales Representative.For questions regarding this recall call 916-281-2790. |  | 71 | 14 in US; 42 to EU and 15 to Japan | Worldwide Distribution - USA including MI, AZ and FL and the countries of Japan and Europe | Terminated | February 01, 2012 | 848 | 71 | Not_Computer
Z-0934-2012 | 60911 | No | NQQ | Ultrasonic pulsed echo imaging system. | Radiology | General &amp; Plastic Surgery | 510(k) | Imalux Niris OCT Imaging System Probe, Model Number 1300, Serial Number F464019, IMALUX Corporation, Cleveland, OH |  |  | December 22, 2011 | 2011 | Imalux Corporation | During an FDA inspection, the firm was notified that the wrong tubing was used during the manufacture of the probe for the unit. | Process control | Imalux Corp contacted their one customer via telephone on December 14, 2011. They informed the customer of the problem and requested the probe returned. The probe involved was returned and received by Imalux on December 22, 2011, and will be stored offsite in quarantine.Any further questions please call (216) 502-0755 ext. 301. |  | 1 | One Unit | Nationwide Distribution only to Texas. | Terminated | January 09, 2014 | 749 | 1 | Not_Computer
Z-0942-2012 | 52807 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Vigilance II Monitor Model: VIG2/VIG2E Software: versions 00.51, 00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, |  |  | September 12, 2008 | 2008 | Edwards Lifesciences, Llc | Edwards Lifesciences has made the decision to initiate a field corrective action for the Edwards Vigilance II Model: Software: versions 00.51, 00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, 00.71, 00.73 due to the potential of inaccurate Injectate Cardiac Output (ICO) measurements. | Device Design | Edwards initiated two Urgent Product Notification - Action Required letters to all affected customers. The letters identified the affected product, problem, and actions to be taken. The letters instructed users to immediately discontinue use of in-line temperature probes for Intermittent Cardiac Output (ICO) measurements on the Vigilance II instrument and recommends use of the bath probe only (model 9850A) to measure the temperature of the injectate solution for ICO calculations, until software update have been completed. Customers were instructed to complete and return a software upgrade form to Edwards as record of completion. For questions call Edwards Technical Support department at 800-822-9837. |  | 6934 | 6934 instruments | Worldwide Distribution - Nationwide Distribution (USA) - and the countries of: Argentina, Australia/ New Zealand, Brazil, Canada, China, Columbia, Ecuador, Europe, Indonesia, Japan, Korea, Macao, Malaysia, Mexico, Middle East, Peru, Puerto Rico, Singapore, Thailand, Taiwan and Venezuela | Terminated | February 08, 2012 | 1244 | 6934 | Computer
Z-0936-2012 | 60936 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva FD Software release: PBL 6.1.1 |  |  | January 09, 2012 | 2012 | Philips Healthcare Inc. | The skin dose and skin dose rate values are calculated using an incorrect point of reference. Therefore, the AKR as displayed and stored in patient records is too low, although the dose as received by the patient is not affected. | Software design | Philips Healthcare issued on January 9, 2012 a mandatory Field Safety Notice to all affected customers. The letter identified the affected product, problem description, hazard involved and actions that should be taken by the customer/user. The letter contains important information for the continued safe and proper use of the equipment. Customers were instructed to review the following information with all members of there staff and retain a copy with the equipment "Instruction for Use". For information or support concerning this notice contact your local Philips representative 866-767-2822. |  | 28 | 28 units (US) | Worldwide Distribution - Nationwide Distribution (USA) including the states of AR, AZ, CA, CO, FL, IA, KS, LA, MI, MO, NC, NY, OH, PA, TX, VA, VT, WA and WV | Open, Classified |  |  | 28 | Not_Computer
Z-0941-2012 | 61010 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DRAD-3000E: Radrex-i X-ray is a general purpose X-ray System. Labeling reads in part "TOSHIBA***DIGITAL RADIOGRAPHY SYSTEM RADREX-i DRAD-3000E Systems Data No. MSDXR0144EA***" |  |  | December 29, 2011 | 2011 | Toshiba American Medical Systems Inc | It has been found that some of the starter cables used for x-ray tube high speed rotation are not sufficiently durable. | Nonconforming Material/Component | The recall notification titled "URGENT: MEDICAL DEVICE CORRECTION", was sent on 12/29/11 to customers via certified mail. The letter informed the customer of an issue with the starter cable used for high speed rotation on the Radrex X-Ray System. The letter also informs customers of the corrective measures: a replacement cable will be installed on your system to prevent this issue from occurring. An 800 number was provided for any questions regarding the letter. |  | 39 | 39 | Class II Recall - Nationwide Distribution (USA) - Including the states of VA, IL, CA, PA, TN, OH, LA, TX, OK, AR, AZ, MS, WV, TN, WI, NJ, MN, NY, IA, FL, and Washington D.C. | Terminated | March 24, 2014 | 816 | 39 | Not_Computer
Z-0938-2012 | 60900 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | KIMBERLY CLARK MIC GASTRO-ENTERIC FEEDING TUBE, 22F. |  |  | December 21, 2011 | 2011 | Kimberly-Clark Corporation | The product is misbranded, i.e. the Gastric (center) port may be mislabeled as the Jejunal port, and the Jejunal (side) port mislabeled as the Gastric port. | Labeling False and Misleading | Kimberly-Clark Health Care sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter and e-mail to all affected customers. The communications described the product, problem, and actions to be taken by the customers. Customers were instructed to locate and quarantine any affected product. A Product Recall Response Sheet was enclosed for customers to complete and return to the firm via fax at 920-380-6682. Contact the firm at 770-587-7400 for questions regarding this recall. |  | 50 | 50 devices | Worldwide Distribution-USA (nationwide) including the states of CA, CO, ID, IL, MD, MO, NC, NJ, NY, PA, SC, UT, VA, WA, WI, and WV and the country of Canada. | Terminated | July 02, 2012 | 194 | 50 | Not_Computer
Z-0939-2012 | 60016 | No | OGL | Portable liquid oxygen unit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Broselow/Hinkle Pediatric Emergency System |  |  | October 17, 2011 | 2011 | Vital Signs Devices, a GE Healthcare Company | Povidone Iodine Pads included in the Vital Signs Broselow Pediatric Emergency Kit may be contaminated with Elizabethkinga meningoseptica. | Material/Component Contamination | GE Healthcare sent a recall letter/return response form dated October 17, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to not use the affected product, isolate and discard all affected product.Fill out the attached Confirmation form and fax back to 800-535-7923 per the instructions to obtain replacement Busse IV Start Kits. Also, consignees were instructed to return the attached response form even if no recalled product was in their inventory.If they forwarded any affected product to any other healthcare institutions, they should forward a copy of the letter to those institutions.For any questions regarding this recall call 1-800-932-0760. |  | 6675 | 6675 kits | Nationwide and Foreign | Terminated | January 15, 2013 | 456 | 6675 | Not_Computer
Z-0940-2012 | 61030 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Aera/Skyra, model numbers 10432914 and 10432915. |  |  | January 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During product monitoring firm has discovered that on a few systems the gradient cable connections did not meet their specifications after installation. Locally, the cable connections overheated, making servicing necessary. The potential existed for a cable connection to become overheated, leading to the emission of smoke and possible material damage. Firm preventively checked the gradient cabl | Nonconforming Material/Component | Siemens Medical Solutions sent a "CUSTOMER SAFETY ADVISORY NOTICE" letter dated November 24, 2010 to all affected customers. The letter described the product, problem, and actions to be taken. Contact the firm at 610-219-4834 for questions regarding this recall. |  | 128 | 128 | Nationwide Distribution-including the states of AK, AZ, CA, CT, DE, FL, GA, IL, IN, KS, MD, MA, MI, MN, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, and WI. | Terminated | January 31, 2014 | 736 | 128 | Not_Computer
Z-0937-2012 | 60969 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF 16 PET/CT System &amp; GEMINI TF 64 PET/CT System, Philips Medical Systems, Cleveland, OH |  |  | October 24, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips became aware of an issue relating to a UPS cabinet that contained a melted battery which caused damage to other batteries lower in the cabinet, leading to battery acid leaking on the floor. | Material/Component Contamination | Philips Medical Systems (Cleveland) Inc communitcated the problem to the facilities via telephone. No customer letter was sent since the correction was already performed by the supplier. |  | 2 | 2 | Nationwide Distribution including Maryland and Colorado | Terminated | June 18, 2013 | 603 | 2 | Not_Computer
Z-0947-2012 | 58937 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***CAT 7205305***DYONICS***Arthroscopic Blade***3.5mm FULL RADIUS, 13cm long***Fits Power/EP-1 Motor Drives***Beige***SAMPLE LATEX-FREE***Made in the USA***Manufacturer: ASCENT HEALTHCARE SOLUTIONS***Lakeland, FL*** Lot number: 199468***Expiration Date: 12/2011. |  |  | January 26, 2009 | 2009 | Ascent Healthcare Solutions | On 01/26/2009, Ascent Healthcare Solutions, Lakeland, FL initiated a recall of various models of Dyonics, Stryker, and Linvatech Reprocessed Arthroscopic Shaver Blades. The sterility of the devices may be compromised due to a potential breach in the seal of the Tyvek cover to the plastic tray. | Packaging process control | Ascent Healthcare Solutions sent an "URGENT MEDICAL DEVICE VOLUNTARY RECALL" letter dated January 26, 2009 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Customers were instructed to examine their inventory, remove and quarantine all affected devices. A Recall Effectiveness Check Form was attached for customers to complete and return via fax to 1-866-325-4760. Contact the firm at (888) 888-3433 ext. 5519 for questions regarding this recall. |  | 330 | 330 devices | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NH, NV, NY, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WI, and WY, and the country of Canada. | Terminated | March 09, 2012 | 1138 | 3702 | Not_Computer
Z-0944-2012 | 60836 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra PAC1 Mod 2 Cables used with the Integra MPM, CAM01 Monitors |  |  | October 20, 2011 | 2011 | Integra LifeSciences Corp | Some of the connector cables associated with Integra's intracerebral pressure (ICP) monitors may cause temporary interference with the ICP waveform on the monitor display. | Under Investigation by firm | Integra sent a Urgent Product Recall Notification letter dated October 20, 2011, via FedEx, email, and fax to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed to examine their inventory at your earliest convenience. If you have an effected PAC1 cable, please complete the attached Recall Acknowledgement Form and return it to Integra Customer Service per the instructions on the form. An Integra Sales Representative will contact you to assist in the return and replacement of new cables at no cost. In the meanwhile, you may continue to use the cables until you have been contacted and replacement product has been provided. Distortion of the ICP wave form can be stopped or prevented by keeping the PACI catheter receptacle out of human contact.Please contact Integra Customer Service at 1-800-654-2873. |  | 221 | 221 Units | Worldwide Distribution-- USA (nationwide ) . | Terminated | June 18, 2013 | 607 | 221 | Not_Computer
Z-0946-2012 | 60977 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance Big Bore Model #728243 and #728244, Philips Medical Systems. |  |  | December 13, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The Block and Multi Leaf Collimator (MLC) shape is not propagated to the beam that is modified by the Oppose operation. The contours can appear with the incorrect color in the Beam's Eye View Viewer, resulting in organs being incorrectly placed or incorrectly shaped. | Software design | On 12/13/2011, the firm sent Field Safety Notices, entitled URGENT - Medical Device Correction, to their customers. The letter identified the affected product, the problem, the hazard involved, and the actions to be taken by customers and the firm. Customers were to follow the instructions provided in order to address or avoid the issues. Philips will be installing a software update to address the identified issues free of charge. If customers need further information or support in regards to these issues, they are to contact their local Philips representative or local Philips Healthcare office. In North America and Canada, customers can contact the Customer Care Solutions Center at 1-800-722-9377. |  | 142 | 142 Units | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, FL, IA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NH, NM, NY, OH, OR, PA, PR, TN, TX, VA, &amp; WA and countries of AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, DENMARK, DOMINICAN REPUBLIC, FRANCE, GABON, GERMANY, HUNGARY, ISRAEL, ITALY, JORDAN, LUXEMBOURG, MALI, MEXICO, NETHERLANDS, OMAN, SLOVAKIA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UKRAINE, UNITED ARAB EMIRATES, &amp; UNITED KINGDOM. | Terminated | February 22, 2013 | 437 | 142 | Computer
Z-0945-2012 | 61012 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian Clinac Accelerators with Mark, Millennium MLC 120 (Multileaf Collimator) and BrainLAB m3 MCL, Reference/FSCA CP-05591, Model Number: H53, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | November 10, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Failure to retract or park the Varian MLC leaves prior to use of the demountable BrainLAB m3 could result in the Varian MLC leaves interfering with the treatment field defined by the BrainLAB m3 without an interlock being asserted. | Device Design | The firm, Varian Medical Systems, Inc., sent an "URGENT MEDICAL DEVICE CORRECTION-URGENT FIELD SAFETY NOTICE" letter dated November 10, 2011 to all affected users. The letter describes the product, description of problem, details, recommended actions and Varian's actions. Varian Service will contact customers to arrange installation, once modifications are available.***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter also included special instructions for customers outside the USA and Canada. The customers were instructed to complete and return the attached Proof of Notification or Receipt Verification Card to Varian Medical Systems once they have read the document.If you require further clarification, feel fee to contact your local Varian Customer Support District or Regional Manager at: USA and Canada: 1-888-VARIANS (1-888-827-4265); Europe: +41 41 749 8844; or Email: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com; Latin America: soporte.al@varian.com or Internet: Oncology Systems customer site-www.myvarian.com and Varian Medical Systems public site-www.varian.com. |  | 418 | 205 units***12-12-12 AMENDED TO: 213 units*** | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Algeria, Austria, Australia, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Slovenia, South Korea, Spain, Sri Lanka, Sweden, Taiwan, Turkey, United Kingdom and Venezuela. | Terminated | February 19, 2015 | 1197 | 418 | Not_Computer
Z-0963-2012 | 60852 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sterile surgical kits: |  |  | June 14, 2011 | 2011 | DeRoyal Industries Inc | Custom surgical kits contain recalled Kimberly Clark surgical drapes which were found to contain metal shavings in the folds of the drapes. | Nonconforming Material/Component | DeRoyal Industries Inc. sent an "URGENT VOLUNTARY PRODUCT RECALL OF KIMBERLY-CLARK THREE-QUARTER DRAPE 79147" letter dated June 14, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify and quarantine all affected product. Further instructions by the firm asks customers to complete and return a Recall Response Spreadsheet via fax to 865-362-3744. Customers should contact their Deroyal Sales Representative for questions or assistance regarding this recall. |  | 3881 | 3881 kits | Nationwide Distribution. | Terminated | June 27, 2013 | 744 | 3881 | Not_Computer
Z-0964-2012 | 60391 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Analyzer, Model Number: 10-004, 14-874, 14-875, in vitro diagnostic. |  |  | November 16, 2010 | 2010 | Alere San Diego | Alere has identified that a bias on results for certain analytes tested on the Alere Cholestech LDX System may exist due to humidity variability. | Device Design | Alere sent an Updated Urgent Recall Notification Letter in March 2012 to all their customers who are in possession of one or more Cholestech LDX Systems. The updated letter provides the customers with important information about the problem identified and the actions to be taken. Customers are instructed to operate their device in a controlled environment that maintains humidity within 40-60% relative humidity if the device has not been upgraded with ROM pack v3.40. If humidity cannot be monitored, the device is no longer CLIA waived. Customers are instructed to complete the customer verification form. A previous notice from Alere sent in November 2010 instructs users to contact Alere to receive the ROM pack upgrade. Customers may email LDXROM@alere.com to request an electronic form to complete and return via email. Customers can expect their ROM packs approximately 10 to 14 business days after Alere receives their request. If customers have transferred ownership to another party, they can email Alere at LDXROM@alere.com with the updated contact information. If customers are not currently using their Cholestech LDX System, it is recommended that the ROM Pack upgrade be installed prior to restarting use of any Cholestech LDX analyzer. Customers with questions regarding this communication can call Alere at (877) 441-7440, option 1 or email at LDXROM@alere.com. |  | 22000 | 22,000 meters | Worldwide distribution including USA | Terminated | August 22, 2013 | 1010 | 22000 | Not_Computer
Z-0943-2012 | 60422 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Foot Rest Kit (pN 6252700-003) for Models 6251 and 6252 Stryker (EMS Stair -PRO), OE 2065939, 1739631, 2066066, 6066803 and 20666941 with optional , Stryker Medical, Portage, MI 49002. |  |  | October 31, 2011 | 2011 | Stryker Medical Division of Stryker Corporation | The Foot Rest Kit may have the wrong pivot spacers making the joint more susceptible to relaxation and prone to fatigue. There is a moderate risk of injury to the patient should the handles become detached during use. | Mixed-up of materials/components | Stryker Medical spoke with all customers who received these shipments prior to 10/31/2011 requesting information on kits affected and recalled foot rest units installed. December 5, 2011 Urgent Medical Device Recall Letters, with instructions on how to to either install a new Foot Rest Kit ( enclosed ) or dispose of any kits that were not installed were sent to both customers who could have installed recalled foot rest units. Customers needing help or assistance with installation are asked to call Stryker Technical Support a 1-800-327-0770. Customers were to complete and return a postage paid postcard confirming receipt. Customers with questions or concerns are asked to contact Kristin Dudek at 269-389-6927 M-F m 8 a.m. - 5 p.m. (EST) |  | 3 | 3 units | Nationwide Distribution -- MI &amp; MA. | Terminated | August 28, 2013 | 667 | 3 | Not_Computer
Z-0966-2012 | 61018 | No | LLT |  |  | Obstetrics/Gynecology | Contact ODE | Bravo pH Monitoring capsule, UPC 7290101361695 (FGS# 0312) and UPC 7290101361688 (FGS# 0313). |  |  | January 11, 2012 | 2012 | Given Imaging Ltd. | Failure of the Bravo capsule to attach to the esophagus or alternatively, failure of the capsule to detach from the placement device. | Nonconforming Material/Component | Given Imaging sent a Urgent Product Recall letter to all affected consignees. The letter identified the product, description of problem, actions to be taken and advised customers to read this important information packet in its entirety. Customers were advised to complete the Customer Response Sheet included and return per the instructions provided. Also attached to the letter was a Customer Bulletin: Update to Vacuum Settings. For questions call 1-800-448-3644 option 3. |  | 21885 | 21,885 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Belgium, Brazil, Denmark, Canada, Cyprus, Finland, Germany, Hong Kong, India, Indonesia, Israel, Italy, Japan, Mexico, Netherlands, Norway, New Zealand, Panama, Portugal, Singapore, Sri Lanka, South Africa, Spain Thailand, Turkey, United Kingdom and Venezuela. | Terminated | September 14, 2012 | 247 | 21885 | Not_Computer
Z-0967-2012 | 60987 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | ACROBAT Stabilizers- |  |  | November 16, 2011 | 2011 | Maquet Cardiovascular, LLC | There may be a small perforation in the sterile pouch of select lots of the ACROBAT V, ACROBAT SUV, and ACROBAT Mechanical Stabilizers and the XPOSE 3 and XPOSE 4 Positioners. | Employee error | Maquet sent an Urgent Device Removal Immediate Action Required letter dated November 16, 2011 and a Field Action Response forms via Federal Express to to affected customers. The letter identifies the problem, product. risk factors and actions to be taken. Customers were instructed to examine their stocks immediately to determine if they have any of the recalled product. If so, discontinue dispensing (distributing) the lot and complete the Field Action Response form attached to complete and fax the attached Field Action Response form. In addition, it provides instructions for how to return the recalled product to MAQUET. Distributors are instructed to provide the same instructions to their customers supplied with the recalled products. If you have any questions contact the Recall Hotline Center at 1-888-613-0271. If you have any additional questions, contact your local MAQUET Cardiovascul'ar sales representative or Customer Service at 1-888-880-2874 (6AM -5PM PST). |  | 9436 | 9436 unit total (US AND WORLDWIDE) | Worldwide Distribution - USA Nationwide | Terminated | January 10, 2014 | 786 | 9436 | Not_Computer
Z-0965-2012 | 53127 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Total Knee Junior Model Number 1100 |  |  | August 10, 2009 | 2009 | Ossur H / F | Some units of the Total Knee Junior produced between November 25th, 2008 and through April, 2009 contain balancing units that may fail under extreme use over a long period. The fault is noncritical and no incidents or injuries resulting from this situation have been reported to Ossur at this stage. Over long term use (estimated to be in the four to six year range) the prosthetic device may fail a | Nonconforming Material/Component | Ossur sent a Safety Notice dated August 10, 2009 by mailed to all affected consignees. Ossur informed consignees of the above mentioned public reason for recall and asked consignees to take the following action: Identify the users of each of the Total Knee Junior model devices listed above and check inventory and isolate any recalled devices on hand but have not fitted on users. The letter states that Sue Fitzgerald from customer service will be contacting consignees in the next few days to coordinate replacement of the affected devices. For question call 1-(800) 219-2150, Ext: 4109. In addition, Ossur website: www.ossur.com we will be posting updates as appropriate. |  | 141 | 141 units | Worldwide Distribution - USA (nationwide) in the state of: AL, AR, AZ, CA, CO, DE, FL, GA, IL, IN, KS, KY, LA, MA, MI, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, SC, TX, and WV and the countries of: Australia, Canada, China, France, Germany, Hong Kong, Israel, Japan, Malaysia, Netherlands, New Zealand, Portugal, Puerto Rico, Saudi Arabia, South Africa, Spain, Sweden, Taiwan, United Kingdom and United States. | Terminated | February 06, 2012 | 910 | 141 | Not_Computer
Z-0973-2012 | 60867 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano pcFM Software Kit ; |  |  | June 11, 2010 | 2010 | Volcano Corporation | Internal review of validation of pcMF Software kit found that instructions directs user to remove the rear panel, causing the potential for electric shock. | Software design (manufacturing process) | The firm, Volcano Corporation, sent an "IMPORTANT PRODUCT INFORMATION" letter dated June 17, 2010 to all customers. The letter describes the product, problem and actions to be taken. The customers were instructed to return the pcFM Software kit to: Volcano Corporation at 2870 Kilgore Road, Rancho Cordova, CA 95670 and contact (800) 228-4728 for instructions on returning the product. The customers were also instructed to complete and return the attached Customer Reply Letter form to Volcano even if they no longer use the pcFM Software Kit via fax to Volcano Customer Support at (916) 638-8812.If you have any questions, contact your Sales Representative or Customer Service at (800) 228-4728. |  | 8 | 8 units | Worldwide distribution: USA (nationwide) including states of: CA, MI, TX, and WI; and countries of: The Netherlands, Sweden and United Kingdom. | Terminated | February 07, 2012 | 606 | 8 | Computer
Z-0971-2012 | 60794 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż Chemistry Products Calibrator Kit 24, Catalog No. 680 1708. Common/Usual Name: VITROS Calibrator Kit 24. For in vitro diagnostics use only. VITROS Chemistry Products Calibrator Kit 24 is used to calibrate the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System for the quantitative measurement of albumin in urine. Ortho-Clinical Diagnostics Inc. |  |  | October 27, 2011 | 2011 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received numerous complaints of lower than expected results for control fluids and patient samples upon initial calibration or following a recalibration using VITROSÂż Chemistry Products Calibrator Kit 24 and VITROSÂż Chemistry Products mALB Reagent. | Nonconforming Material/Component | All consignees were sent "Important Product Correction Notification" communications (dated 10/27/11) beginning on 10/27/11 informing them of the issue, and instructing them to discontinue use of their remaining inventory of VITROSÂż Chemistry Products Calibrator Kit 24, VITROSÂż Chemistry Products mALB reagent, and VITROSÂż Chemistry Products mALB Performance Verifiers. For questions, please contact OCD's Customer Technical Services at 1-800-421-3311. |  | 4436 | 4436 units worldwide (domestic - 2201; foreign - 2235) | Worldwide Distribution - USA, including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | February 12, 2016 | 1569 | 4436 | Not_Computer
Z-0979-2012 | 60577 | Yes | OAE |  |  | Cardiovascular | PMA | EZ Steer ThermoCool Nav Part Number D-1292-01-S Diagnostic/Ablation Catheter instructions for use. |  |  | November 22, 2010 | 2010 | Biosense Webster, Inc. | Biosense Webster, Inc. has recently become aware of the omission of the following two Warning and Precaution statements on the Instructions For Use (IFU) for the above catheter:1. If phrenic nerve location is a concern, precautionary measures are recommended to evaluate the proximity of the nerve to the ablation electrode, such as pacing maneuvers.2. To minimize risk of atrio-esophageal fist | Labeling design | Biosense Webster, Inc. sent an Urgent Field Safety Notice letter dated Novembe 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to continue use of the affected product and that no further action is required on their part. Consignees were advised that the updated IFU can be found at www.jnjgatewayifu.com. Consignees were instructed to contact their local Biosense Webster representative if they have any questions or require additional information.For questions regarding this recall call 909-839-8604. |  | 24713 | 24,713 units in total | Nationwide Distribution | Terminated | February 08, 2012 | 443 | 24713 | Not_Computer
Z-0970-2012 | 60721 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Hemotherm 48ÂżC Dual Reservoir Heater/Cooler System, Model Number: 400 and Model Number: 400MR. The 48 o ( Degree) C Hemotherm Dual Reservoir Heater/Cooler I provides a flow of temperature-controlled water to a heat exchanger. The device, in conjunction with a heat exchanger, is used to warm/cool the fluid that enters and exits a heat exchanger. The conductive heat/cool transfer is accomplished through any type of blood oxygenator/heat exchanger unit. The operator of the unit regulates the desired fluid temperature. Frequent observation of the actual blood/fluid temperature and body temperature is required. The Hemotherm does not contact the patient of the fluids; the water being circulated by the Hemotherm is separated from the fluid by the heat exchanger. To active temperatures above 42 Degrees C, the user must follow explicit instructions in to the owner's manual in order to activate the higher setting. Without following the explicit instructions, the device can only operate within the normal maximum temperature range of: 42 degrees C. The Hemotherm can be used with adult and pediatric patients for extracorporeal circulation of the patient's limb, entire body, or a chemo medication lavage of the abdominal area. |  |  | October 27, 2011 | 2011 | Cincinnati Sub-Zero Products Inc | The suspect hyper-hypothermia device systems sold to physicians for research and/or the treatment of AIDS and Cancer patients, were not appropriately cleared by the FDA prior to distribution. | Software design | CSZ sent an "URGENT: FIELD ACTION NOTICE" letter dated January 20, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to become familiar with the amended manual page and to get approvals from the Institutional Review Board before further use of the device. Additionally, a response form was enclosed with the letter for customers to complete and return via fax to (513)772-9119. Contact CSZ Service Department at 1-800-989-7373 for questions regarding this notice. |  | 60 | A total of 20 devices were distributed in the U.S. &amp; 3 units were distributed internationally. | Nationwide Distribution-including the District of Columbia and states of CA, GA, MA, MD, MI, NJ, NY, OH, and UT. | Terminated | February 26, 2013 | 488 | 76 | Not_Computer
Z-0975-2012 | 60919 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU Clinical Chemistry Analyzers: Model Number: |  |  | July 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman has confirmed that the AU Clinical Chemistry Analyzers may experience cuvette overflow. | Under Investigation by firm | Beckman Coulter sent an "UPDATE TO PCA -14563: URGENT PRODUCT CORRECTION" letter dated April 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter is a communication which describes an alternative approach to detecting cuvette overflow. Additionally, a response form was enclosed for customers to complete and return to the firm. Contact the Customer Support Center at (800) 854-3633 or on the Beckman Coulter website at www.beckmancoulter.com/customersupport/support/default.asp for questions regarding this notification. |  | 8225 | 8225 units | Worldwide Distribution-including USA (nationwide). | Terminated | August 04, 2014 | 1484 | 8225 | Not_Computer
Z-0978-2012 | 60011 | No | DWH | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Terumo, Sarns Sternal Saw I System, Catalog 6090. The Sarns Sternal Saw I is indicated for use in splitting the sternum during thoracic surgery. |  |  | November 17, 2011 | 2011 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received one report of minute metal shavings leaking from the Sarns Sternal Saw I System at the point where the saw blade is inserted into the saw handpiece. | Equipment maintenance | On 11/16/2011, 11/23/11, 02/3/12, 03/22/12, and 4/24/12 the firm sent a letter entitled "URGENT MEDICAL DEVICE RECALL" to it's domestic direct accounts that included the name of the device, catalog and lot numbers, dates of distribution, and instructions to discontinue use and return the device to the firm for destruction, with a Fax-back response form to be submitted. Medwatch reporting information was also included in the letter. Questions and comments are to be addressed by Customer Service @ 1-800-521-2818, M-F, 8-6 EST |  | 10186 | approx. 10186 | Worldwide Distribution - USA, and the countries of Colombia, Argentina, Lebanon, Belgium, Chile, India, Sri Lanka, Jordan, Egypt, Mexico, Brazil, Taiwan, Honduras, Canada, Germany, Singapore, Republic of Georgia, Hong Kong, Japan, UAE, Philippines, Thailand and South Africa. | Terminated | December 06, 2012 | 385 | 10186 | Not_Computer
Z-0972-2012 | 60662 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens CLINITEST hCG -Pregnancy Test (qualitative), Radio immunoassay,Human Chorionic Gonadotropin (hCG) test system. Package of 25. Product Code 1760/ Catalog# 06484105). |  |  | December 15, 2011 | 2011 | Siemens Healthcare Diagnostics | A potential for false positive or increase in borderline results with patients and commercially available negative control material. | Process control | Siemens issued Overnight Mailings of Urgent: Recall Notices on December 15, 2011 for delivery beginning December 16, 2011 to Distributors and end users. Customer bulletin has been issued instructing customers to discontinue use of product and discard impacted inventory. Customers are to contact their distributor regarding credit for unused product. For technical questions contact Siemens Technical Solutions Center at 877-229-3711. |  | 48578 | 48,578 Kits (25 tests / kit) | Worldwide Distribution -- US, Austria, Germany, Great Britain, Denmark, Norway, Sweden, Poland, Netherlands, Luxembourg, Belgium, Finland, France, Greece, Ireland, Italy, Portugal, Qatar, Saudi Arabia, Slovakia, Slovenia, and Switzerland. | Terminated | March 30, 2015 | 1201 | 48578 | Not_Computer
Z-0984-2012 | 60800 | No | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | VITROSÂż Valproic Acid (VALP) Reagent, REF/Catalog No. 6801710 --- Ortho- Clinical Diagnostics Rochester, NY USA -- Ortho- Clinical Diagnostics High Wycombe, UK --- An in vitro diagnostic. |  |  | May 10, 2011 | 2011 | Ortho-Clinical Diagnostics | Results using VITROSÂż Chemistry Products VALP Reagent were negatively biased compared to other manufacturers. OCD confirmed that patient results were also affected by this issue. | Manufacturing material removal | All domestic end-user consignees were initially sent "Important Product Correction Notification" recall communications (dated 5/10/11) by Federal Express overnight mail on 5/10/11. The letters informed users of the issue, and advised them to discontinue using their VITROS VALP reagent until adjusted calibrator values are provided for GEN 14 of VITROS VALP Reagent. The adjusted calibrator values for VITROS VALP Reagent were sent to customers via Assay Data Disk (ADD) Data Release Version (DRV) 5632. Customers were to refer to the Technical Support Communication for details. Users were to confirm receipt of the notice by completing and returning the Fax Reply Form - Confirmation; confirm that all customers or distributors were informed of the issue and instructed to discontinue use of affected product; and lastly confirm that all customers or distributors were sent adjusted calibrator values. |  | 5800 | Approximately 5800 units worldwide (domestic - 4549, foreign - 1287) | Worldwide Distribution -- US, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | February 12, 2016 | 1739 | 5800 | Not_Computer
Z-0985-2012 | 60898 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva IITV (MD-EII) with Software Release: PBL 6.1.1 |  |  | January 09, 2012 | 2012 | Philips Healthcare Inc. | Incorrect Skin Dose calculation. AKR (Air Kerma Rate) displayed and stored in patient records is too low | Software design | Philips Healthcare issued a mandatory Urgent Field Safety Noticedated October 3, 2011 to all affected customers. The letter identifies the affected product, problem description, hazard involved, action to be taken by customer/user and actions planned by Philips to correct the problem. The letter provide important information for the continued safe and proper use of the equipment. The letter instructs customers to review the following information with all members of their staff who need to be aware of the contents of this communication and retain a copy with the equipment Instruction for Use. For further information or support concerning this issue contact your local Philips representative at 866-767-2822. |  | 10 | 10 units (US) | Worldwide Distribution - USA (nationwide) including the states of CA, IL, OH, SD, TX, WY and the countries of: Australia, Austria, Canada, Egypt, Germany, mexico, Turkey and UK | Open, Classified |  |  | 10 | Not_Computer
Z-0148-2012 | 59982 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Signa, Discovery and Optima. |  |  | September 15, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of missing labels on the alignment lasers. | Labeling design | GE Healthcare will execute a field action in which GE Healthcare Field Engineers will perform an inspection of all suspect MR systems. Any systems lacking any of the 3 laser warning labels will be remediated by applying the missing labels). We expect this field action will be completed by June 30, 2012.1. Notification of all distributors and purchasers be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. The FDA district office coordinator noted below is to be included in the notifications.2. Corrections be made at no cost to the purchasers. |  | 4614 | 4614 | Nationwide Distribution | Terminated | June 29, 2012 | 288 | 4614 | Not_Computer
Z-0986-2012 | 60938 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra OS Osteoconductive Scaffold Putty (2.5 cc and 5 cc). Integra OS Putty, combined with bone marrow aspirate, is intended for use as a bone void filler to fill voids or gaps of the skeletal system. Following placement in the bony void or gap (defect), Integra OS Putty is resorbed and replaced with bone during the healing process. |  |  | October 07, 2011 | 2011 | Integra LifeSciences Corp | Integra OS Osteoconductive Putty is labeled with incorrect storage conditions. | Error in labeling | Integra issued an "Urgent: Product Recall Notification" to customers and sales specialists on 10/7/2011. They were advised of the problem and provided instructions on how to handle the affected product. |  | 45 | 45 units | Nationwide Distribution - Including the states of CA, OK, MD and VA. | Terminated | May 21, 2013 | 592 | 45 | Not_Computer
Z-0987-2012 | 59092 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ***REF C9944A***"FULL RADIUS RESECTOR"***2.9mm***Qty 6***STERLING***STERILE***Rx Only.*** Made in the USA***Manufacturer: CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***ConMed Linvatec Biomaterials, Ltd.Hermiankatu 6-8L 33720 Tampere, FINLAND.***Lot number # 95987. |  |  | November 12, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL recalled catalog number, C9944A, lot # 95987, 2.9MM Full Radius Resector Blade in November 2009 because there was a possibility that the device was assembled with an incorrect outer tube. | Employee error | A Linvatec Corporation dba ConMed Linvatec sales representative contacted the affected customers by telephone on November 12, 2009. The customers were also contacted by the Recall Coordinator. The product, problem, and action to be taken by the customers were discussed. Customers with the product in their inventory were instructed to return the product to the firm. Contact the firm at 727-399-5205 for questions regarding this recall. |  | 30 | 30 Units | Product Distributed in the states of Georgia and New York. | Terminated | February 13, 2012 | 823 | 30 | Not_Computer
Z-0988-2012 | 60902 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Gamma camera systems: AXIS SPECT Systems &amp; IRIX SPECT Systems, Model #: AXIS- 882130, IRIX- 882140. Philips Healthcare. |  |  | November 18, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Philips was informed that the detector buckets may experience weakness over time. As a result, this could cause the detector bucket to separate from the mounting plate of the gantry and eventually result in the detector falling. | Nonconforming Material/Component | On 11/15/2011 Philips sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected product, the problem, the hazard involved, and the actions to be taken by users and the Philips. Customers are to place the letter in their User Documentation. Philips service will arrange a time to inspect customers' systems and conduct the appropriate field safety correction. Customers are to contact their local Philips representative if they need further support or information. In the US and Canada, customers can contact Philips Healthcare Customer Care Solutions Center at 1-800-722-9377. |  | 795 | 795 units | Worldwide Distribution -- USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PRO, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY and the territory of Puerto Rico; and countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Egypt, Finland, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Japan, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, Peru, Philippines, Poland, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, &amp; United Kingdom. | Terminated | February 22, 2013 | 462 | 795 | Not_Computer
Z-0990-2012 | 60693 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Flow Tec Nano-Hybrid Flowable Affected Product |  |  | March 03, 2011 | 2011 | Sybron Dental Specialties | The recall was initiated because Pentron Clinical has confirmed that the Flow Tec Nano-Hybrid Flowable Composite material has been found to be difficult to extrude. | Nonconforming Material/Component | Pentron Clinical sent an initial "URGENT: MEDICAL DEVICE RECALL" letter dated April 20, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and fax back an enclosed recall return form in order to confirm their receipt of the recall notification. A second "URGENT: MEDICAL DEVICE RECALL" letter dated November 10, 2011 was sent to all affected customers because Pentron Clinical expanded the recall. On December 8, 2011 another recall communication was forwarded to all affected customers to change the wording on the letter to make it more clear to the customers about the recall expansion. Contact Pentron Clinical Customer Service at 1-800-551-0283 for questions regarding this recall. |  | 77289 | 77,289 Units total | Worldwide Distribution: USA (nationwide) and the countries of Canada, Egypt, France, Germany, Italy, and Spain. | Terminated | October 05, 2012 | 582 | 77289 | Not_Computer
Z-1008-2012 | 60929 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | S1827 Crea Calibration solution 1, 944-135, Lot DV-02 Radiometer Medical, Bronshoj Denmark |  |  | December 16, 2011 | 2011 | Radiometer America Inc | When the user activates the dosicap (containing the crea additive) and mounts it onto the bottle, the threads will not seal correctly causing approximately 10% of the solution to leak out when the bottle is placed horizontally for dissolving the additive. This will cause the concentration of creatinin and creatin in the calibration solution to be wrong, and results in a bias on the Crea results. | Device Design | Radiometer sent a Medical Device Recall letter dated December 28, 2011 to their customers. The letter provided details on the affected devices, description of the problem and actions to be taken. The firm instructed their customers to immediately examine their stock and return any remaining product to Raiometer America Inc. for disposal. The letter states that Radiometer will replace all bottles of the affected lots, at no charge. Customers were asked to complete the attached Recall Response Fax Form and fax to 800-736-0601. For questions regarding the Recall contact Radiometer America Technical support at 1-800-736-0600 opt. 4. |  | 42 | 7 boxes/6/175mL bottles | Nationwide Distribution (USA) including the states of: NY, PA and TN. | Terminated | January 17, 2014 | 763 | 42 | Not_Computer
Z-1000-2012 | 59936 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Artiste Flowable Value Pack Part Number N28VA, Syringe Lot Number 165741 |  |  | October 22, 2010 | 2010 | Sybron Dental Specialties | The recall was initiated because Pentron Clinical has confirmed the Flow Tec Nano Hybrid Flowable Composite material has been found to be difficult to extrude. | Nonconforming Material/Component | Sybron Dental sent a Recall Letter dated November 15, 2010 to consignees informing them of the above mentioned public reason for recall and asking them to determine if they have any affected products in their inventory and requesting their cooperation in returning any affected product which they still have in their stock. Product was to be returned and replaced at no charge.They also requested that consignees complete and fax back an enclosed return from , whether or not they had any product to return.Consignees were told to call Pentron Customer Service at (800) 551-0283 directly to handle the arrangements of a quick return and replacement.Authorized Petron Clinical distributors were asked to identify and recover the affected product lots that may have been shipped to their customers. |  | 771 | 771 units in total | Worldwide Distribution - USA (nationwide) and the countries of Cyprus, Germany, Italy, Poland, Canada, Ecuador, and Kuwait | Terminated | February 13, 2012 | 479 | 771 | Not_Computer
Z-1011-2012 | 60835 | [6 empty] | Zoll pedi.padz Reduced Energy Multi-Function Electrodes (MFE) |  |  | December 30, 2011 | 2011 | Bio-Detek, Inc. | Lack of an adequate silicone seal around the periphery of the attenuator enclosure may allow fluid to enter. | Process control | Zoll Medical issued Urgent Device Correction notices dated 12/30/11 via UPS/Registered Mail. The letter described the issue and identified the affected product. ZOLL recommended that customers alert their staff of the issue, that the electrodes should be kept away from a fluid environment, to use caution in handling the electrodes if they are subjected to a fluid environment and maintain safe defibrillation practice, and seek a replacement if available. Also customers should contact Zoll 1-800-348-9011 for product replacement. The firm plans to replace the affected lots free of charge. |  | 3080 | 3,080 units | Worldwide Distribution -- US and Canada. | Terminated | December 11, 2012 | 347 | 3080 | Not_Computer
Z-1003-2012 | 61038 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Custom surgical kits: |  |  | January 10, 2012 | 2012 | DeRoyal Processing Center | The firm distributed surgical kits which contained DermaHook Neuro Elastic Retractors which were subsequently recalled by Teleflex. | Nonconforming Material/Component | DeRoyal Industries sent a Customer letter dated January 6, 2012, to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer. An email was also sent on January 10, 2012, to the affected customer. The customer was requested to follow these instructions. 1) Identify any affected inventory using the product and lot numbers on Attachment 1.2) Attach the provided warning label to each affected kit and tray.3) Forward this recall to any of your end users that may have affected product.4) Complete and return Attachment 1 by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter, even if you no longer have affected inventory.If you have questions or need assistance with the recall, including additional labels, please call (865) 362-1037 . We apologize for any inconvenience this may caused you. |  | 214 | 214 units | (USA) including Indiana | Terminated | November 05, 2012 | 300 | 214 | Not_Computer
Z-1012-2012 | 61044 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Model #03001-02: Periophil Biphasic, Synthetic Biphasic Dental Bone Graft Granules, 1.0 cc. |  |  | January 11, 2012 | 2012 | Cytophil, Inc | Cytophil Corporation is recalling Periophil Biphasic Periophil Biphasic, Periobase Biphasic, and Periobone Biphasic because devices were inadvertently shipped after receiving an inadequate dose of radiation and this product may not be sterile. | Process control | Consignees were sent an "Urgent Medical Device Recall" letter dated January 11, 2012 by Cytophil. The letter described the problem and the product involved in the recall. The firm advised consignees to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, they were requested to identify their customers and notify them at once of this product recall. For questions, call 262-642-2765. |  | 547 | 547 | Nationwide Distribution - Including the states of WI, GA, IL, SC, MO, and TX. | Terminated | May 01, 2012 | 111 | 547 | Not_Computer
Z-1010-2012 | 61004 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical lighting systems iLED and TruLight in use with the ALC+ function. |  |  | January 09, 2012 | 2012 | Trumpf Medical Systems, Inc. | The ALC+ function may not be properly locked during use resulting in the possibility of the handle falling off into the surgical field causing injury to the patient. | Device Design | The firm, TRUMPF Medical Systems, Inc., sent an "Urgent Safety Information" notice/package dated January 9, 2012 to its customers. (Note: Some were delivered by TRUMPF representatives in person and some via mail with Urgent Safety identification on the envelope). The notice describes the product, problem and actions to be taken. TRUMPF will replace all sterilizable handles of this type in the field with new handles, in addition, TRUMPF will send new handles and verify disposal of the old version at the customers location. Distribution of new handles will start January 2012. The customers were instructed to ensure that all users and other relevent persons are informed of the Urgent Safety Information; store the information until the measure has been completed and complete and return the Confirmation of User Receipt form via mail to TRUMPF Medical Systems, Inc. 1046 LeGrand Blvd., Charleston, SC 29063 or email: Lindsey.ronnenberg@us.trumpf-med.com.If you have any questions, call 843-822-6939. |  | 8733 | 8,733 units | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, CT, DE, FL, GA, IL, IN, MA, MD, MO, NC, NM, NJ, NY, OH, PA, TX, VA, VT, WA, and WI; and countries of: Australia; Austria, Belguim, China, Cyprus, Denmark, Estonia, Finland, France, Germany, Great Britain, Hong Kong, Indonesia, India, Ireland, Italy, Japan, Kazakhstan, Kenya, Korea, Latvia, Luxembourg, Mexico, Netherlands, Norway, Pakistan, Peru, Philippines, Poland, Romania, Russian Federation, Saudi Arabia, Seychelles, Singapore, South Africa (Zuid Afirka), Spain, Sweden, Switzerland (Confederation of Helvetia), Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Terminated | October 09, 2012 | 274 | 8733 | Not_Computer
Z-1009-2012 | 60988 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD 30G1/2" Precision Glide" Needle |  |  | October 08, 2011 | 2011 | Becton Dickinson &amp; Company | Becton Dickinson has identified clogged needles. | Device Design | Becton Dickinson sent Urgent Product Recall letters dated October 18, 2011 and October 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory to deterine if they have any of the affected product. If they have affected product customers should return the product to Beckton Dickinson. For any questions regarding this recall call 201-847-4344. |  | 480000 | 480,000 units | Nationwide Distribution including IL and NJ | Terminated | December 11, 2012 | 430 | 480000 | Not_Computer
Z-1004-2012 | 60943 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CellTracks Analyzer II. |  |  | August 17, 2011 | 2011 | Veridex, LLC | Control cell images may i intermix with Patient cell images while changing between the Control review screen and Patient review screen on the Cell tracks Analyzer II. | Software design | Veridex, LLC sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated August 17, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter provides procedure sheets with step by step directions for customers to follow. Additionally, a Confirmation of Receipt form was enclosed for customers to complete and return to the firm. Contact Customer Technical Services at 1-877-837-4339 for questions regarding this notice. |  | 37 | 37 Units | Worldwide Distribution. | Terminated | June 17, 2015 | 1400 | 37 | Not_Computer
Z-0997-2012 | 60901 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DUET TRS 60 4.8MM STRAIGHT SULU |  |  | January 12, 2012 | 2012 | Covidien LP | Duet TRS Straight and Articulating Single Use (SULU)Loading Staplers used in thoracic surgery may result in serious injury or death | Under Investigation by firm | Covidien issued an Urgent Medical Device Recall letter dated January 12, 2012 to all affected customers. The letter identified the affected products, problem, and actions to be taken. Customers were instructed to immediately discontinue use of the device in thoracic surgery, Immediately advise all surgeons/surgery personnel of this recall and return products intended for thoracic use to Covidien by contacting Customer Service at SDFeedback@Covidien.com or 1-800-722-8772, option 1, to obtain a Return Goods Authorization prior to returning the affected units. Customers were advised to complete the Recall product return form in its entirety and return it with the returned units. For questions contact your Covidien representative or Quality Assurance at 203-492-5232. |  | 9687 | 9,687 units | Worldwide Distribution - USA (nationwide) and the countries of: ÂżeskÂż republika, ASSAGAY, Australia, Austria, Belgium, Bosnia, Herzegovina, Bulgaria Cairo,Egypt, Canada, Canary Islands, China, Croatia,Cyprus, Czech Republic, Denmark, Egypt, England, Finland, France,Germany, Guadeloupe, Hong Kong, Hungary, Iceland, India, Iran, Ireland,Israel, Italy, Japan, Jordan, Kingdom of Bahrain, Korea, Republic of Kuwait,Latvia, Lebanon, Libya, Luxembourg, Nederland, Nelspruit, New Redruth,New Zealand, Norway, Nouvelle Caledonie, Philippines, Poland, Portugal, Qatar, Reunion, Russia, Romania, Rustenburg, Saudi Arabia, Serbia,Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Turkey, Taiwan, Thailand, Tunisia, Uae, Uae - Dubai, Umhlanga Rocks,United Kingdom, Vietnam and Wales. | Open, Classified |  |  | 666880 | Not_Computer
Z-1014-2012 | 50701 | No | NOP | West Nile virus serological reagents. | Microbiology | Microbiology | 510(k) | West Nile Virus IgM Capture ELISA |  |  | October 08, 2008 | 2008 | Inverness Medical | It has come to our attention that users of these product lots may experience an elevation in the rate of false positivity when compared with previous product lots. Internal quality assurance testing has shown that the products are performing within the release specifications, although there has been a slight shift in the cut-off which may lead to an increase in the number of positive results obta | Nonconforming Material/Component | Inverness Medical sent an Urgent In Vitro Diagnostic Recall letter dated October 6, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use, destroy any unused or partially used test kits of the affected lot(s), and complete the table and certification of compliance. Customers were asked to fax back the completed form to inverness Medical at 1-609-527-8054. Replacement stock would be forwarded to them at no cost when the new lot becomes available.For questions regarding this recall call 609-627-8038. |  | 289 | 289 Kits | Worldwide Distribution including ND, PA, TX, MN; and the country of Canada | Terminated | February 15, 2012 | 1225 | 289 | Not_Computer
Z-1013-2012 | 59094 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | CI healthcare***Model: CI8092***S/N: 5-60-85-044 MFD: JUN, 2010 MADE IN CHINA***cihealthcare.com |  |  | February 17, 2011 | 2011 | Cardiopulmonary Instrumentation, Inc | Cardiopulmonary Instrumentation, Inc. dba CI Healthcare, Miami, FL, is recalling their Stainless Steel Patient Transport Stretcher, model # CI-8092, Lot 002, because of a malfunction that occurs when weight is applied to the backrest. This malfunction causes the backrest support bar, when secured in the first hook, to fall to a flat position. | Device Design | CI Healthcare sent an "URGENT RECALL NOTIFICATION" letter dated February 17, 2011, to its customer. The letter describes the product, problem and actions to be taken. The customer were instructed to do the following: 1) Examine their inventory and isolate the product(s). 2) If they have distributed any of the products to other services or facilities, please forward this information as appropriate. 3) Discontinue use of the product until a correction has been properly executed and 4) Complete and return the "Recall Response Form" as soon as possible. Should you have any questions, please call Robert Nunez at (305) 592-8196. |  | 72 | 72 Units | Nationwide Distribution - to one distributor in Puerto Rico | Terminated | April 19, 2012 | 427 | 72 | Not_Computer
Z-1032-2012 | 60996 | Yes | OAD |  |  | Cardiovascular | PMA | The 1500T9 Cardiac Ablation Generator, RO#lBl-89000 is part of the Therapy Cardiac Ablation System, along with the Cool Point irrigation Pump, the Therapy 1300 Series Steerable Ablation Catheters, the Therapy Cool Path Ablation Catheters and accessories. The generator is a microprocessor-controlled device that produces a continuous unmodulated radiofrequency (RF) output delivered from the generator between its tip electrode and a dispersive pad for the treatment of atrial flutter. |  |  | January 05, 2012 | 2012 | St Jude Medical Inc | Two St Jude Medical 1500T9 RF Generator units were distributed to customers while still under FDA review as part of a PMA supplement to the approved PMA applications. | Vendor change control | The notification letter was dated January 5th, 2012 and titled "Product Recall RF Generator 1500T9-CP". The letter informed customers that a St Jude Medical 1500T9-CP RF Generator unit had been distributed to them while currently under FDA review as part of a PMA supplement application. The letter requested that the use of the impacted generator be discontinued. Customers were requested to return the field action form and return the unit identified in the letter. The contact number provided in the letter is: 651-756-2000. |  | 2 | 2 units | Nationwide Distribution, including the states of Kansas and Maryland. | Terminated | September 04, 2014 | 973 | 2 | Not_Computer
Z-1016-2012 | 60947 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | DYNASTY(R) BIOFOAM(R) SHELL, SIZE: 46mm GROUP B, REF DSBFGB46, Rx ONLY, STERILE R, Wright Medical Technology, Inc., hip prosthesis acetabular component. |  |  | January 03, 2012 | 2012 | Wright Medical Technology Inc | The possibility exists that particulate debris from the packaging spacer could be unintentionally implanted into a patient and could potentially lead to complications that might necessitate a revision surgery. | Packaging | Wright Medical sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated December 27, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product and to return any affected product to the firm. A Product Recall Verification Form was enclosed for customers to complete and return via fax to 901-867-7401. Contact Customer Service at 1-800-238-7117 for questions regarding this recall. |  | 160 | 160 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, Denmark, and Italy. | Terminated | March 13, 2013 | 435 | 5342 | Not_Computer
Z-1036-2012 | 60946 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion 100-BGM TeleHealth Monitor, Model No. 09377512, Product Code/Part No. 81842478, Bayer HealthCare LLC 555 White Plains Road, Tarrytown, NY 10591-5097 USA --- Manufactured in Japan for Bayer HealthCare LLC -- Common Name: Viterion100 TeleHealth Monitor -- Classification Name: Monitor, Physiological, Patient (Without Arrhythmia Detection or Alarms) [ref. 21CFR 870.2300] |  |  | May 25, 2011 | 2011 | Viterion TeleHealthcare Llc | Software defect; Specifically, blood glucose (BG) results from the Roche Aviva meter that are tagged by the user with either the meter hypoglycemic or general "asterisk" flags (*) will not be automatically transmitted through the Viterion 100-BGM monitor containing software version 3.2.0 or displayed to the health care provider through the data server. | Software design | On May 12, 2011, Bayer Healthcare (aka Viterion) telephoned their consignees regarding the issue with the Viterion 100-BGM and Roche ACCU-CHEK Aviva Flagged Data and followed-up with a letter (dated May 27, 2011) via e-mail. The letter notified customers of the issues and provided information on how to work around the problems. The letter provided the following recommendations: (1) The information concerning the general asterisk (*) flag be communicated within their organization as an amendment to the current patient Safety Alert Notice that was distributed and posted in April 2011. (2) For patients who need to use the Aviva Hypo and/or the asterisk flag, please assign Manual Data Entry for blood glucose on the V100-BGM. The letter informs customers that Bayer has completed the software development to rectify the aforementioned data transmission issue and will upgrade all affected products. Bayer request that customers confirm their receipt of this memo in writing and for any questions concerning this situation, contact Thuy Nguyen, Manager Quality Assurance and Regulatory Affairs at 914-333-6821. |  | 3905 | 3905 units domestically | (USA) Nationwide Distribution | Open, Classified |  |  | 3905 | Computer
Z-1040-2012 | 61098 | Yes | NIO |  |  | Cardiovascular | PMA | S.M.A.R.T. CONTROL NITINOL STENT SYSTEM ILIAC: Self-Expanding Vascular Stents delivered by catheter for the indication of improving luminal diameter in patients with symptomatic atherosclerotic disease of the iliac arteries. catalog numbers are CxxyyyMB, CxxyyySB, CxxyyyML, and CxxyyySL, where xx is the stent diameter in mm and yyy is the stent length in mm; but not including catalogs where yyy is 120 or 150 . The Removal only includes Catalog Numbers in the above sequence with S.M.A.R.T.Âż CONTROLÂż in the logo/description. All unexpired lots from 15112447 through 15525516. |  |  | January 31, 2012 | 2012 | Cordis Corporation | Cordis Corporation initiated a recall of S.M.A.R.T. CONTROL Nitinol Iliac Stent System due to a potential sterility breach in the pouches. | Packaging | All affected consignees were notified by an "Urgent Medical Device Recall (Removal) Letter" dated January 31, 2012 and mailed by overnight delivery on February 1st and 2nd, 2012. Consignees were instructed to identify and quarantine affected product; review, complete, sign, and return an Acknowledgement Form; return the affected product per instructions. Questions should be directed to Customer Service at (800) 551-7683. |  | 132000 | 132,000 units | Worldwide Distribution - USA, and the countries of Japan, Canada, Mexico, South Korea, and Taiwan. | Terminated | August 08, 2014 | 920 | 150000 | Not_Computer
Z-1033-2012 | 61129 | No | OHP | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | Dental | Dental | 510(k) | Provent -SR - Sleep Apnea Therapy-SR (standard resistance). |  |  | February 03, 2012 | 2012 | Ventus Medical, Inc. | A discrepancy was found during inspection of a returned unit that the lot number and SKU on cartoon did not match. | Labeling mix-ups | Ventus Medical sent a Urgent Medical Device Correction Letter dated February 3, 2012, to all affected customers. The letter identified the problem, the product and the action needed to be taken by the customer.There is no health hazard associated with the affected devices or with the continued use of the affected devices.Ventus Medical wanted Customers to be aware of the problem so they can respond to patients if inquiries are received. A sample letter to patients was also provided, should physicians choose to provide it to patients. Questions may be directed to Ventus Medical Territory Business Manager at 650-632-4199. |  | 288 | 288 units | Nationwide Distribution including the states of OH, TX, IL, OR, CA, FL and NY. | Terminated | February 15, 2012 | 12 | 288 | Not_Computer
Z-1038-2012 | 60942 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Temporary Pacing Wire, BIPOLAR, Models BM, V, VE and VF, Rx Only, STERILE/EO, Single use only -- Mce (MEDICAL CONCEPTS EUROPE) --- Manufacturer ECM B.V. Oost-Om 54, P.O. Box 53, NL 5420 AB GEMERT --- Shipping or Unit Package: 20 individual units packaged per box with 6 boxes per case. --- Device Classification Name: Electrode, pacemaker, temporary. ---The device is supplied sterile and intended for single use. |  |  | April 18, 2011 | 2011 | European Custom Manufacturing B.V. | The crimping operation during product production was not performed properly. As such, the pacing wire may separate from the connection pin, which causes non-functionality. | Process design | ECM sent an "ECMbv RECALL OF MCE TEMPORARY PACING WIRES" letter dated May 2, 2011 to all affected customers. Customers were also informed of a potential problem via e-mail and/or phone between April 18 and 21, 2011. The letter describes the product, problem, and actions to be taken by the customers. Customers were instructed to place the affected product on hold and to return the product to the firm. For questions on this matter, contact ECM or MCE Customer Service at 1-877-886-1985. |  | 2120 | 2,120 individual units (106 boxes) | Nationwide Distribution-including the states of AL, AR, CA, FL, GA, KY, NC, NV, NY, TX, VA, and WI. | Terminated | February 15, 2012 | 303 | 2120 | Not_Computer
Z-1037-2012 | 61117 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Orthofix rotating cutter used in spinal fusion procedures with the FORZA Spacer System. Catalog Number 89-0452 (8mm Rotating Cutter), 89-0454 (10mm Rotating Cutter), and 89-0030 (FORZA Discectomy Case). Manual surgical instrument comprised of stainless steel and titanium. |  |  | February 06, 2012 | 2012 | Orthofix, Inc. | The 8mm and 10mm Rotating Cutters may fracture during use, specifically during disc space preparation. If an 8mm or 10mm Rotating Cutter does Fracture during use, it may result in a surgical delay estimated to be 30 minutes or greater due to preparing the disc space using other instruments or methods and/or retrieving the instrument fragment. | Device Design | Orthofix mailed recall letters to the distributors on 02/06/2012. Orthofix does not sell or consign these instruments to hospitals. |  | 52 | 52 | Nationwide and Puerto Rico | Terminated | June 11, 2012 | 126 | 52 | Not_Computer
Z-1034-2012 | 61111 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A: AQUILION ONE; System, X-Ray, Tomography, Computed |  |  | January 03, 2012 | 2012 | Toshiba American Medical Systems Inc | The recall was initiated because Toshiba has confirmed that their is a potential software issue for the TSX-301A Aquilon One and the TSX-301B Aquilon One. | Software design | Toshiba sent an Urgent Medical Device Correction letter to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete and fax the attached form to (877) 349-3054. The form can also be sent via email to raffairs@tams.com.Customers with any questions regarding this letter were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968. |  | 82 | 82 units total | Nationwide Distribution - USA (nationwide) and Puerto Rico | Terminated | April 08, 2013 | 461 | 82 | Computer
Z-1042-2012 | 60866 | No | ETP | Air or water caloric stimulator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Leica M822 Surgical Microscope Device Controller Software (MDC version 3.6, pack F, affecting the control unit component, Article #10448446, in combination with a Leica M822 Optics Carrier-XY-Focus-Tilt-Unit, article #10448587; Leica Microsystems (Schweiz) AG, Max Schmidheiny-Strasse 201, CH-9425, Heerbrugg, Switzerland |  |  | December 21, 2011 | 2011 | Leica Microsystems, Inc. | Leica Microsystems received complaints stating that the zoom function of the Leica M822 surgical microscope locked up during use. | Device Design | The firm, Leica Microsystems, Inc., sent a "Medical Device Correction" notice dated December 21, 2011 to all affected customers by e-mail. The product, problem, and actions to be taken by the customers were addressed. Customers were instructed- DO NOT USE use the affected product until the software could be upgraded and complete and return the Acknowledgement Form via fax to: 1-847-607-3147 or Scan and send to elizabeth.culotta@leica-microsystems.com. Leica Microsystems representatives will contact the customers to make necessary arrangements to complete the software upgrade.If you have any questions, contact Director, Global Product Quality Management at 1-847-317-7209. |  | 29 | 29 units (10 units-USA) | Worldwide distribution: USA (nationwide) including states of: California, Colorado, Illinois, New York, Oregon, South Carolina; and countries of: Australia, Japan, Korea, Pakistan, Romania, South Africa, Switzerland and Turkey. | Terminated | February 16, 2012 | 57 | 29 | Computer
Z-1045-2012 | 60990 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | SLIDEX Staph kit (250 Tests). |  |  | December 21, 2011 | 2011 | bioMerieux, Inc. | Reports of false negative results in organism identification. | Nonconforming Material/Component | Biomerieux sent an "URGENT PRODUCT REMOVAL NOTICE" dated December 21, 2011 to the affected customer. The letter describes the product, problem, and actions to be taken by the customer. The letter instructs the customer to stop using and to destroy the affected products. The firm recommends steps to follow when interpreting results. An acknowledgement and Product Replacement Form was enclosed for the customer to complete and return to the firm. Contact Biomerieux Customer Service representative at (800) 682-2666 for information regarding this notice. |  | 100 | 100 kits | Distributed to one customer in California. | Terminated | September 17, 2013 | 636 | 100 | Not_Computer
Z-1050-2012 | 60599 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | "***REF NL82059-03***INTEGRA NEUROSCIENCES***Touhy Needle***Non Sterile for Single Use Only***14g, 6"***Plainsboro, NJ***" |  |  | November 28, 2011 | 2011 | Integra Neurosciences PR, Inc. | The current packaging configuration for the non-sterile shunt accessories could lead to the potential for confusion regarding product sterility. | Labeling mix-ups | Integra sent an Urgent Product Recall Notification letter dated November 28, 2011. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out and sign the attached Recall Acknowledgement Form noting the part numbers, lot numbers and quantity of the affected product in their inventory. Customers were asked to contact Integra LifeSciences Corporation International Customers Service at 609-936-5400 to arrange for Returns. If customers do not have any of the affected product in their inventory they are required to complete and sign the form and indicate they have NO affected product in their inventory. Customers should fax or email their Recall Acknowledgement Form to:609-750-4259 or LAPAC_CS@Integralife.comFor questions call 609-936-2495. |  | 357 | 357 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Canada, Italy, Australia | Terminated | October 29, 2013 | 701 | 1303 | Not_Computer
Z-1044-2012 | 61019 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | STA - System Control N + P |  |  | August 23, 2011 | 2011 | Diagnostica Stago, Inc. | Factor VIII quality control results are being recovered below the assigned ranges for STA System Control N+P. | Nonconforming Material/Component | Stago sent a recall letters via email dated August 23, 2011 to all affected customers . The letter identified the product, description of problem, and corrective actions to be taken. Customers were asked to immediately examine their inventory and follow the corrective actions provided. Customers were instructed to complete the response form attached to the letter. Because of incorrect manufacturer's instructions included in the initial letter another letter was emailed by the firm on August 30, 2011. For questions or technical assistance for updating the parameters please call the Diagnostica Stago 24 hour hotline at 1-800-725-0607. |  | 1259 | 1259 units | Worldwide Distribution and USA (Nationwide) and including Canada | Terminated | September 04, 2013 | 743 | 1259 | Not_Computer
Z-1043-2012 | 60982 | No | GGW | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade(R) Actin(R) FSL Activated PTT Reagent |  |  | December 19, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | The listed lots of Dade(R) Actin(R) FSL Reagent show an increase in Heparin sensitivity over the shelf life which is demonstrated by testing with Dade(R) Citrol(R) Heparin Control, low and high. | Unknown/Undetermined by firm | Siemens sent an Urgent Field Safety Notice letter dated December 2011 to all customers of the affected lots. The letter identified the product, description of problem, and actions to be taken by the user. DadeÂż ActinÂż FSL customers were instructed to exercise one of two options: 1) If customers would like to continue working with their Dade(R) Actin FSL reagent lot (if one of the impacted lots) they must verify on a regular basis (monthly) that their therapeutic range is still valid, e.g. by using a Heparin containing control. In the case that it is not valid, they are to adjust their existing therapeutic range or establish a new one. 2) If customers decide not to continue working with the Dade(R) Actin(R) reagent lot (if one of the impacted lots) they are to contact their Siemens representative at 1-800-242-3233, option 5, for a replacement lot. Once they have established a new normal, therapeutic, and quality control ranges, they are to discard the remainder of the old lot. Customers were advised to complete the response form included to confirm receipt of Safety Notice and fax back to (302) 631-8467. If you have any questions or concerns, please contact the Technical Solutions Center at 1-800-242-3233, option 1. |  | 24672 | 24,672 Kits | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Mongolia, Netherlands, Oman, Philippines, Poland, Portugal, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom and Vietnam. | Terminated | April 03, 2014 | 836 | 24672 | Not_Computer
Z-1046-2012 | 61102 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity CT (Computed Tomography) System, 510 (k) #K033326, Model #728326, Serial Numbers: 300003, 300004, 300005, and 300010. |  |  | December 08, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Design defect. Philips learned of this device issue following their receipt of a customer complaint reporting that a 'gap' is created in the side covers when the gantry of the device is tilted to +30 degrees (forward tilt). A hazard to the operator/user presents itself if the operator puts his or her fingers or hand inside the gap and the gantry tilt is returned in position while the fingers or | Device Design | The firm, Philips, sent an 'URGENT - Medical Device Correction Ingenuity CT***' letter dated December 08, 2011. The letter describes the affected product, the nature of the problem along witth the circumstances under which this problem can occur; the hazard involved and actions to be taken by the customer/user. The customers (operator/users) were instructed not to insert hands in the gap which will be located between the gantry and stationary support column located on either side of the CT System (illustrations are provided). The letter notifies the customers that a Philips representative will be visiting each customer site and installing a hardware upgrade through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, option 5; enter site 10 or follow prompts). |  | 8 | Four (4) device system units were distributed. Two (2) in U.S. and two (2) overseas | Worldwide distribution: USA (nationwide) including states of: IN and OR; and countries of: Australia and Germany. | Terminated | March 01, 2013 | 449 | 8 | Not_Computer
Z-0833-2012 | 57842 | No | RCE |  |  |  |  | GE eXplore Locus Micro CT cabinet x-ray system model RS9-80. |  |  | January 31, 2011 | 2011 | GE Healthcare, LLC | Product fails to comply with the federal performance standard for cabinet x-ray systems (Title 21 CFR 1020.40(c)(1)(i)). Specifically, the x-ray emission limit of 0.5 milliroentgen in one hour could be exceed. Highest emission measured was 3 mR/hr. | Radiation Control for Health and Safety Act | GE Healthcare will notify purchasers of the problem, the affected product, instructions for users, and specify the maximum emission rate measured. GE Healthcare will bring defective systems into compliance free of charge. This will be done via Field Modification Instructions IFMI) 11101 which is scheduled for deployment on March 1, 2011 with an estimated completion date of July 1, 2011. |  | 21 | 21 | Worldwide Distribution | Terminated | September 06, 2012 | 584 | 21 | Not_Computer
Z-1059-2012 | 60972 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ENDURAMESH CORPECTOMY SPACER SYSTEM 013MM, 017MM, 023MM by varying heights. Dist. By: Lucero Medical, PO Box 67, Richfield, Ohio 44826 |  |  | January 11, 2012 | 2012 | Lucero Medical LLC | During an FDA inspection the firm was informed that the device marketed is not the device cleared for the 510K associated with the device. The product is sold unlabeled. | Device Design | Lucero Medical LLC made a personal visit to facilities in regards to an Urgent Medical Device Recall as of January 12, 2012. The firm instructed the customer that they were removing and quarantining all product from the facility. If you have any further questions, please contact Lucero Medical, PO Box 69, Richfield, Ohio 44826, Director of Quality Assurance/Natalie Williams/ (440) 670-9583 |  | 190 | 190 units | USA Nationwide Distribution - in the state of OH. | Terminated | April 11, 2014 | 821 | 190 | Not_Computer
Z-1060-2012 | 60855 | No | KXE | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Ascension Orthopedics First Choice Partial Ulnar Head System ("DRUJ System"). Head sizes 14.5 mm, 16.0 mm, 17.5 mm, and 19.0 mm each with stem sizes 4.5 mm, 5.5 mm, or 6.5 mm. |  |  | December 03, 2011 | 2011 | Ascension Orthopedics, Inc | A recent review conducted by Ascension Orthopedics revealed that certain modifications were made to the First Choice System and should have been preceded by a submission of a filing with the FDA. | Device Design | The Firm, Ascension Orthopedics, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated December 3, 2011 via FedEx, email or fax to all customers. The letter described the product, problem, and actions to be taken. The customers were instructed to immediately cease any further use of any affected product; remove from their stockrooms until it is shipped back to Ascension Orthopedics; promptly return any affected product to: Ascension Orthopedics, 8700 Cameron Road, Austin, TX 78754, Attn: First Choice DRUJ Recall Administrator, and complete and return the reply form even if you do not have affected product in stock, to Regulatory Affairs Dept. via fax to 512-836-6933.If you have any questions, please do not hesitate to contact customer service at 1-512-836-5001. |  | 1303 | 1303 units released to customers | Worldwide distribution. | Terminated | January 14, 2013 | 408 | 1303 | Not_Computer
Z-1056-2012 | 61110 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | TruFlight Select PET/CT System equipped with software version 3.5.3.10, Model #459800111661. |  |  | December 13, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances have been noted in the TruFlight Select PET/CT System. A number of these non-conformances have been corrected to improve the overall quality and functionality of the suspect system units. | Software design | Philips sent an 'URGENT - Medical Device Correction TruFlight Select PET/CT Serial Numbers: 17002, 17003 letters. The letter was dated December 13, 2011. The letter describes the affected product, the nature of the problem along with the circumstances under which this problem can occur; the hazard involved and actions to be taken by the customer/user. In this case, the customers (operator/users) are provided with separate recommended actions. The letter goes on to notify the customers that a Philips representative will be visiting each customer site and installing a software update (3.5.3) in order to address several of the non-conformance issues through a Field Change Order. For further information or support concerning this issue contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center at (1-800-722-9377, Customers in all other countries are advised to dial their local Philips Healthcare office. |  | 2 | two (2) units were distributed | Nationwide Distribution -- OH and MD. | Terminated | February 22, 2013 | 437 | 2 | Computer
Z-1055-2012 | 60513 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | ImmunoCAP Allergen Ri206, Cockroach American, Periplaneta americana |  |  | January 03, 2012 | 2012 | Phadia US Inc | The firm found this lot of ImmunoCAP Allergen Ri206, American Cockroach does not contain sufficient levels of some cockroach allergens e.g. tropomyosin. Approximately 2% of the positive samples have a risk to fall below the Limit of Quantization (o .1 kUA/L). | Release of Material/Component prior to receiving test results | ThermoFisher sent an URGENT RECALL NOTIFICATION letter dated December 22, 2011, to all affected consignees. The letter identifed the product, the problem, and the action to be taken by the conignee. Consignees were recommended to cease testing on the affected lot of product and, if so desired, substitute with ImmunoCAP Allergen i6, German Cockroach, until ImmunoCAP Allergen Ri206 was in stock. For any questions regarding this notification, contact Technical Support at 800.346.4364, option 2. |  | 10 | 780; all sent to US, with 10 assays per sleeve | Nationwide to include: MN, VA, NC, MO, SC, TX, MA, MI, FL, IL, PA, MD, UT, CO, GA, CA, | Terminated | February 04, 2013 | 398 | 10 | Not_Computer
Z-1054-2012 | 60974 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Concentrated Mouse Monoclonal Antibody CD7 (LP15) Reagent; an in vitro diagnostic primary antibody for use as part of an antibody panel for the classification of tumors of T-cell origin; Catalog No. NCL-L-CD7-580; packaged in 1 mL vials and 0.1 mL vials; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. |  |  | January 10, 2012 | 2012 | Leica Microsystems, Inc. | The primary antibodies to CD2 (clone 11F11) and CD7 (clone LP15) have been adulterated. | Labeling mix-ups | Leica Microsystems, Inc. sent a Medical Device Field Notification letter dated January 4, 2012, via first class mail on January 10, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the product and to return the product to the firm. Customers who have received the affected lot numbers of the listed products were requested to return any unused/part used reagents to the Buffalo Grove, IL, location for replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the account's Leica Representative. For further questions please call (847) 405-0123 |  | 72.1 | 60 - 1 mL and 11 - 0.1 mL vials | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | April 18, 2012 | 99 | 339 | Not_Computer
Z-1062-2012 | 61052 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Agfa DX-D Imaging Package - (The detector power cables described in the Report of Correction and Removal are connected to portable detector panels used in the DX-D systems) |  |  | January 16, 2012 | 2012 | AGFA Corp. | The firm received a complaint that when using the DX-D Mobile DR X-Ray unit they noticed the detector cable was hot. | Device Design | AGFA Corp. sent an "URGENT FIELD SAFETY NOTICE" letter dated January 16, 2012 to affected consignees. The letter described the affected product, concerns, issues and actions to be taken. Customers have been requested to place the Varian warning label included in a highly visible and suitable position on the cable of the portable detector. Customers were instructed to complete and return the feedback form by fax to 864-421-1664, as soon as possible. For questions about this matter contact your local AGFA HealthCare Technical Support at 877-777-2432. |  | 42 | 42 units | Nationwide Distribution (USA) including the states of: FL, IL, IN, KS, MI, MN, MS, NE,NY, OH, SC and WV. | Terminated | May 10, 2012 | 115 | 42 | Not_Computer
Z-1061-2012 | 60863 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | ASSISTO Arm System Table-mounted self retaining instrument holder system consisting of stainless steel tubular, articulated arms. |  |  | December 16, 2011 | 2011 | Geomed Medizin-Technik Gmbh &amp; Co. | The product may fracture during surgical procedures. | Nonconforming Material/Component | Geomed Medizin-Technik Product sent an Urgent Voluntary Medical Device Recall notification on December 9, 2012 and December 16, 2012 to affected customers. The letter identified the affected products, reason for recall and actions to be taken. Customers were instructed to (1) fill out the enclosed return form and fax back immediately at +49 (0)7461-935593, (2) return the affected products to the address provided, including the completed return form, (3) use the UPS account number 03710R provided to return the products free of charge and (4) immediately forward a copy of this recall notification to third parties to which you have distributed affected products. For questions and support, please contact Mr. Hanno Haug, VP -AMB at +49 (0)7461 9355-30 or by email: hhaug@geomed.de. |  | 1205 | 1,205 units | Worldwide Distribution - USA (nationwide) including the countries of: Australia, Belgium, China, Denmark, Finland, France, Germany, Great Britain, India, Italy, Japan, Lebanon, Libya, Lithuania, Luxemburg, Netherlands, Russia, Sweden, Spain and Turkey. | Terminated | April 15, 2014 | 851 | 1205 | Not_Computer
Z-1063-2012 | 60058 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE SENOGRAPHE 2000D FULL FIELD DIGITAL MAMMOGRAPHY SYSTEM ( SENO 2000D). |  |  | October 17, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate measurements on GE magnified images provided by Senographe 2000D Mammography systems when reviewed on IHE compliant workstations. This issue may impact patient safety. | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 17, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. Contact the Call Center at 800-437-1171 for questions regarding this notice. |  | 1060 | 1060 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA,CT, DC,FL, GA, HI, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, VA, WV,and WI and countries of VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIA, QATAR, PORTUGAL, POLAND, NETHERLANDS, MOROCCO, MEXICO, MALAYSIA, LEBANON, KUWAIT, KOREA, JAPAN, JAM,AICA, ITALY, IRELAND, HUNGARY, HONG KONG, HONDURAS, GREECE, GERMANY, GEORGIA, FRANCE, EGYPT, DOMINICAN REPUBLIC, COLUMBIA, CHINA, CHILE, CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | March 12, 2014 | 877 | 1060 | Not_Computer
Z-1064-2012 | 61071 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Imaging XS. |  |  | January 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | In September 2011 firm issued two Customer Safety Advisory Notices for syngo Imaging XS. Firm became aware of an unintended behavior when using syngo Imaging XS, version VA70A and VA70B with Client/Server Installations. Loading additional images/series from the Patient Browser using the thumbnails may cause unintended deletion of images/series. In response, firm issued a Customer Safety Advisor | Software design | Siemens Medical issued an "UPDATE INSTRUCTIONS SY055/11/S/UPDATE PACKAGE VA70B_0211" to all affected customers. The package described the product, problem, and instructions for customers to follow. |  | 128 | 128 | Nationwide Distribution-including the states of CA, CO, CT, FL, GA, IL, IN, IA, KS, MD, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | September 29, 2012 | 247 | 128 | Computer
Z-1065-2012 | 60126 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | REVOLUTION XR/D DIGITAL RADIOGRAPHIC IMAGING SYSTEM WITH TOMOSYNTHESIS |  |  | October 26, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue associated with the hand grip that is clamped to the table top of your Revolution XRD X-ray systems that may impact patient safety. The hand grip accessory that is clamped to the table top may not be properly adjusted and can dislodge during use. Should this happen during patient loading or unloading, patient injury could occur. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to remove the hand grip accessory from use until the product is corrected and to take extra precautions when loading or unloading patients from the table while the hand grips are not in use. A GE Healthcare service representative would contact the consignee to arrange for the correction. For questions regarding this recall call 262-513-4122. |  | 772 | 772 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WY, WI and the countries of VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TAIWAN, SWITZERLAND, SPAIN, SLOVENIA, SAUDI ARABIA, REPUBLIC OF KOREA, NORWAY, NEW ZEALAND, NETHERLANDS, MALTA, KUWAIT, JORDAN, JAPAN, ITALY, ISRAEL, HONG KONG, GERMANY, FRANCE, ECUADOR, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, and AUSTRALIA. | Terminated | August 02, 2012 | 281 | 772 | Not_Computer
Z-1067-2012 | 60751 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE SIGNA EXCITE MR SURGICAL OPTION (MR450w Surgical Suite). |  |  | December 15, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the failure of the transfer board release lever associated with the OR Compatible patient transfer table of your Optima MR450w MR Surgical Suite. | Nonconforming Material/Component | GE Healthcare, sent a Urgent Medical Device Correction letter dated December 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed discontinue use of OR Compatible table until issue is corrected. GE Healthcare will correct all affected systems by providing a replacement table cradle incorporating design changes to the cradle release lever mechanism at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification. For further questions please call (262) 513-4122. |  | 2 | 2 | Worldwide Distribution -- USA including Wisconsin and the country of Austria. | Terminated | May 29, 2012 | 166 | 2 | Not_Computer
Z-1070-2012 | 60828 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Children's Medical Ventures 1 ml Amber Oral Syringe. The product is a Liquid Medication Dispenser. |  |  | December 12, 2011 | 2011 | Benlan, Inc. | The tip caps may come loose from the syringe, potentially resulting in leakage of medication from the syringe. | Process control | The firm initiated their recall of the specified lots of product in November, 2011 by sending a "Customer/Distributor Medical Device Correction" letter to their consignee. The letter requested they examine their inventory, quarantine any affected product and complete a verification and response form to be returned to the firm. Questions were directed to (905) 829-5004 ext 233. |  | 143250 | 143,250 | Nationwide Distribution, USA - including only the state of Pennsylvania | Terminated | March 19, 2013 | 463 | 143250 | Not_Computer
Z-1041-2012 | 59827 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | V24-00400-29 cable, an accessory to the Newport HT50 ventilator. |  |  | August 10, 2011 | 2011 | Flight Medical Ltd. | The recall was initiated because Flight Medical Innovations Ltd. has identified a problem may occur when contact between an exposed unused wire and the grounding wire occurs. If this contact occurs the ventilator may reset and switch off. | Device Design | Newport Medical Instruments, Inc., the U.S. distributor for Flight Medical Ltd., sent a letter dated August 23, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue use of the affected product. An Acknowledgement Form was attached to the letter for customers to complete and return via fax at (714) 427-0489 or e-mail to Regulatory@NewportMedical.com. Replacement product will be provided once the form is received. Contact Newport Medical's Technical Service Group at 800-451-3111 ext. 500 for questions regarding this notice.Updated: The firm posted a press release on 02/12/12 Isreal time, notifying customers of update regarding its voluntary recall of remote alarm cable, part number V24-00400-29, used to connect a Newport HT50 mechanical ventilator to a nurse station. The press release notified: The recalled cable may negatively affect Newport HT50 ventilators manufactured prior to 2007. Affected units may shutdown unexpectedly. The FDA has classified this action as a Class I recall.On August 23, 2011, Flight Medical sent an urgent Recall Notice to affected customers through its distributor, Newport Medical Instruments, Inc. (Newport Medical), to stop using the alarm cable immediately and to contact Newport Medical in order to receive an upgraded version of the cable. In the notification letter, the customers were notified of the affected cables and that an upgraded version of the cable would be provided to them free of charge. Customers, including Newport Medical are requested to dispose of the affected cables properly.Instructions to customers: Customer inquiries related to this action should be addressed to Newports Technical Service Group at (800) 4513111, extension 500. |  | 363 | 363 total units, all in the US | Nationwide Distribution. | Terminated | May 21, 2012 | 285 | 363 | Not_Computer
Z-1072-2012 | 60881 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż 3600 Immunodiagnostics Systems, Catalog # 6802783; COMMON/USUAL NAME: VITROSÂż 3600 --- NOTE: VITROSÂż 3600 that are e-Connected are not affected by this issue. |  |  | April 13, 2011 | 2011 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints regarding the loss of calibration data following the installation of a development Assay Data Disk (ADD) that was available in the on-screen mailbox on the Main Menu System Status screen of their VITROSÂż 5600 Integrated System. OCD's investigation determined that ADD, Data Release Version (DRV) 9241 or 9251 was electronically transferr | Software design | Ortho-Clinical Diagnostics (OCD) sent an Important Product Correction Notification letters dated April 13, 2011 via FedEx overnight to customers with an affected VITROSÂż 3600 or 5600 System (those not e-Connected). Foreign affiliate consignees were informed of this action via e-mail on April 13, 2011. The letters described the affected products, problem and actions to be followed. The letters states that an OCD Field Engineer will contact customers who has the affected products and perform a Modification (MOD) procedure that will remove ADD DRV 9241 or 9251 from the on-screen mailbox of their system(s). Customers were instructed to complete and return the accompanying Confirmation of Receipt Form and fax to1-888-557-3759 or 1-585-453-4110. The letter asks that distributors forward a copy of the enclosed customer communication to all customers who were shipped an affected VITROSÂż 3600 or 5600 System. For questions, please contact OCD's Customer Technical Services at 1-800-421-3311. |  | 96 | Domestic - 15 units, foreign - 81 units | Worldwide Distribution - USA (nationwide) including the countries of: Brazil, Canada, Chile, China, England, France, Germany, India, Italy, Japan, Mexico, Singapore, Spain and Venezuela. | Open, Classified |  |  | 228 | Computer
Z-1071-2012 | 60971 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Discovery Elbow System Humeral Condyle Set- Hexalobular, REF 114700, Lot 797550 Alloy, Sterile, Biomet Orthopedics Warsaw IN 46581. |  |  | December 14, 2011 | 2011 | Biomet, Inc. | A Discovery Condyle Kit w/ Hexalobula was opened during an elbow joint replacement surgery, and was found to contain two male condyles. | Process change control | December 14, 2011 Urgent Medical Device Recall Notices were sent to Distributors and hospital OR managers describing the product, problem, risks associated with the recall, and actions to be taken. The notice included Instructions to locate and remove the product. Customers were asked to complete a response form and fax to 574-372-1683, prior to returning product by priority mail to the attention of Angi Dickson, Biomet , 56 East Bell Drive Warsaw, IN, USA 46580 with a copy of the response form. Receipt of the notice should also be confirmed by calling 800-348-9500, extension 3983 or 3009. Questions related to this notice should be directed to 574-371-3983 or 574-371-3009 Monday through Friday 8 AM to 5 PM. |  | 10 | 10 units | International Only Distribution -- Netherlands &amp; UK. | Terminated | December 12, 2012 | 364 | 10 | Not_Computer
Z-1066-2012 | 61050 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI 4.5 and R4.6 |  |  | January 26, 2012 | 2012 | Elekta, Inc. | If a user, using an on-line correction protocol, clicks the accept correction button, an approval dialog box appears. If cancel is selected, the table correction values return to their initial values, but the image panes continue to display the changed position. | Software design | ELEKTA sent an Important Field Safety Notice (20001507042) dated January 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.For questions regarding this recall call 770-670-2548 |  | 96 | 96 units | Nationwide Distribution including AR, AZ, CA, CT, FL, GA, ID, IL, LA MD, MA, MI, MN, MO, NJ, NY, NC, ND, OH, OK, OR, PA, RI, TN, VA, WA, WI, and Puerto Rico | Terminated | April 17, 2012 | 82 | 96 | Not_Computer
Z-1069-2012 | 61124 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | AngioDynamics SPECIAL ORDER Morpheus Smart PICC CT --- 4F Single Lumen 65cm --- Max CT Flow Rate: 4mL/sec --- Kit Contains: --- 1 each Nitinol Guidewire, 0.46mm (0.018in.) x 130cm --- Catalog No. [REF] 52110902 --- [Mfg symbol] AngioDynamics, Inc., 603 Queensbury, NY USA 12804. |  |  | February 06, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Morpheus Smart PICC label states the kit contains a Nitinol guidewire; however the kit contains a Platinum guidewire. | Labeling mix-ups | AngioDynamics sent an "URGENT-MEDICAL DEVICE RECALL" letter dated February 2, 2012 to the single customer. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs the customer to identify and segregate the affected product. A Morpheus Smart PICC Recall Reply Form was enclosed for the customer to complete and return via fax to 518-798-1360. Contact AngioDynamics Customer Service at 1-800-772-6446 for questions regarding this notice. |  | 25 | 25 units | Distributed to one customer in Wisconsin. | Terminated | February 10, 2016 | 1465 | 25 | Not_Computer
Z-1081-2012 | 60748 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | LH Lithium Heparin tubes, Greiner bio-one, North America, 4238 Capital Drive, Monroe, NC 28110. |  |  | December 14, 2011 | 2011 | Greiner Bio-One North America, Inc. | The lithium heparin tubes may have a contamination with EDTA potassium. | Material/Component Contamination | Greiner Bio-One sent a "RECALL OF LITHIUM HEPARIN TUBES WITH EDTA CONTAMINATION" letter dated December 13, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete an attached Product Disposition Site Confirmation Form and to fax the form to 800-726-0052. This form will confirm that customers have discarded/destroyed all affected product. The firm will replace the product after receipt of the form. Contact the firm at 800-515-8112 for questions regarding this recall. |  | 2082000 | 2,082,000 pieces | Nationwide Distribution. | Terminated | June 21, 2012 | 190 | 2082000 | Not_Computer
Z-1080-2012 | 61036 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CORFLO-Ultra Jejunal Tube with CORLOCK-Twoomey Y-Adaptor for use with the CORFLO-Max PEG; enteral feeding tube,10 Fr, 36" (91 cm), 3 Gram weighted bolus; individually packaged, 5 tubes per case; CORPAK MedSystems, Wheeling, IL 60090; catalog/reorder #30-7361; For delivery of enteral nutrition |  |  | January 26, 2012 | 2012 | Corpak Med Systems | The wrong size Twoomey adapter was packaged with the Jejunal tube. | Mixed-up of materials/components | Corpak MedSystems telephoned customers on January 26, 2012 and sent a follow-up Recall Notification letter dated January 27, 2012 via e-mail reiterating the telephone conversations. Customers were informed of the urgent product recall, the problem and the actions to be taken. Customers were instructed to inspect inventory, segregate any affected product and contact Corpak at 1-800-323-6305 to arrange for the return and replacement of the units found. Customers were advised to complete and return the enclosed acknowledgment form, indicating the number of affected units found in their inventory. Distributors were requested to notify their customers. For questions contact Corpak MedSytems at 1-800-403-3400. |  | 145 | 145 units | Worldwide Distribution - USA (nationwide) including the states of Georgia, Indiana, Kentucky, Maryland, Michigan, Minnesota, Tennessee, Utah and the countries of to Sweden and New Zealand. | Terminated | March 21, 2012 | 55 | 145 | Not_Computer
Z-1082-2012 | 60890 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga - previously known as Microsoft Amalga Unified Intelligence System (UIS) 2009 |  |  | January 04, 2012 | 2012 | Microsoft Corporation Health Solutions Group | Healthcare providers reported that when sharing a single browser session on Amalga Web, patient A's data can be displayed under patient B's name. | Software design | Microsoft Corporation sent an "Urgent Field Safety Notice" letter dated January 5, 2012 by email to all affected customers. The letter described the product, problem, and actions to be taken by the customers. If you have any questions about this corrective action program, please contact Microsoft Corporation at 425-882-8080. |  | 9 | 9 units | Worldwide Distribution - USA (nationwide) including the states of California, China, Kentucky, North Carolina, Wisconsin and country of Switzerland. | Terminated | December 12, 2012 | 343 | 9 | Not_Computer
Z-1074-2012 | 60959 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Strep A Rapid Test Cassette (K031784), IST-502 Cassettes. |  |  | September 01, 2011 | 2011 | Alere San Diego | Alere San Diego is expanding the scope of the voluntary recall of the Strep A Rapid Test Cassette and Twist devices from the single lot that was recalled in January 12, 2011 to all lots of product that were on the market within the expiry period. | Process change control | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 1, 2011 to its customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to retain the recall letter for their records as documentation that they have received the recall information. Customers were instructed to follow the directions and return the attached Verification Form. Customers were instructed to use the table in the Verification Form to document all inventory at any of their inventory locations. Customers may use multiple forms to document all sites, products, and lots. Alere will issue credit against PO#ASD811STA for any unused portion of the product based on completion and return of the attached Verification Form. Customers with any technical questions about the information contained in the notification were instructed to contact Alere by phone (877) 866-5309, fax (877) 866-9304 or email: ca2services@Alere.com. |  | 1859430 | 1,859,430 devices (228 lots) | Worldwide distribution: USA (nationwide) and countries of : Austria, Canada, Cyprus, Denmark, Finland, Germany, Greece, Hungary, Iceland, Italy, Poland, Portugal, Romania, Slovenia, Spain, Switzerland, and United Kingdom. | Terminated | June 28, 2012 | 301 | 4657030 | Not_Computer
Z-1077-2012 | 60962 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Precision Aspheric Lens (PAL) Intraocular Lens |  |  | August 02, 2011 | 2011 | Aaren Scientific, Incorporated | The recall was initiated because Aaren Scientific has confirmed that a specific group of lenses were shipped without a complete expiry date showing on the label. | Incorrect or no expiration date | Aaren Scientific sent a recall communication letter dated August 2, 2011, August 3, 2011, and August 30, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Aaren Scientific is requesting that you return the lenses and we will rework the product and return them back to you. Please let me know and I will issue an RGA number and inform you of our FedEx account number to reference.In the meantime, we are conducting an internal investigation regarding this issue. Once this has been closed out then we will notify you of our findings.For further questions please call (909) 937-1033 ext. 108. |  | 3 | 3 units | Worldwide Distribution - USA including CA and TN, and the countries of India and Lebanon. | Terminated | August 10, 2012 | 374 | 153 | Not_Computer
Z-1083-2012 | 60967 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | UltraVision Detection System; |  |  | January 11, 2012 | 2012 | Lab Vision Corporation | Firm was notified by vendor that a component in kits was not meeting internal specifications. Declination in product performance confirmed by Thermo Fisher. | Nonconforming Material/Component | Lab Vision Corporation of Thermo Fisher Scientific sent a Product Recall Notification letters dated January 11, 2012 to all affected customers. The letter identified the affected products, reason for recall and necessary actions to be taken. Customers were instructed to examine inventory, discontinue use, destroy any remaining inventory, follow normal procedure to order an alternate lot of the affected products, retain a copy of this letter for their laboratory records, if the affected products were further distributed provide a copy of this letter to them and complete the attached Customer Recall Response Form and return within 5 days. For questions regarding this information, U.S. customers should call Technical Support at 510-771-1595. Customers outside the U.S. should contact their local area Customer Service. |  | 31 | 31 kits in total | Worldwide Distribution - USA (nationwide) FL, IL, MA, TX and countries of Denmark, Turkey, Italy, Poland, Canada, Japan, Germany, Egypt and Greece. | Terminated | July 20, 2012 | 191 | 31 | Not_Computer
Z-1093-2012 | 60932 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 DIAGNOSTIC ULTRASOUND SYSTEM with R1.1.0 Software. |  |  | February 07, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue where an image from a different patient, or an image from the current patient with title information from a different patient maybe noticed. | Software design | Consignees were sent on 2/7/12 a GE Healthcare "Urgent Medical Device Correction" letter dated February 6, 2012. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Head of General Imagining Ultrasound Department. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Users were instructed not to press the PRINT key during the specified view in order to avoid seeing the issue. GE will provide a software update at not cost to customers in order to remove the malfunction. Customers are to contact their local GEHC service representative or call 1-800-321-7937 if they have questions or concerns. |  | 161 | 161 | Worldwide Distribution -- US, including the states of NJ, PA, and WI and countries of UNITED KINGDOM, THAILAND, SOUTH KOREA, SLOVENIA, SINGAPORE, ROMANIA, REPUBLIC OF MOLDOVA, POLAND, MALAYSIA, JAPAN, ITALY, INDIA, GERMANY, FRANCE, EUROPE, CZECH REPUBLIC, AUSTRIA, AUSTRALIA, and ARAB EMIRATES. | Terminated | July 06, 2012 | 150 | 161 | Computer
Z-1098-2012 | 61045 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System with 4D Planning procedure option of MultiPlan Treatment Planning System, versions 3.0, 3.1, 3.5.1, 3.5.2 and 3.5.3 with Ray-Tracing dose calculation. |  |  | January 16, 2012 | 2012 | Accuray Incorporated | An anomaly was discovered during internal regression testing, where the dose information is displayed incorrectly during treatment planning during a specific workflow using the optional 4D Ray Tracing dose calculation algorithm. As a result, it is possible that the dose calculation will display a lower dose than the intended dose prescribed for treatment delivery. | Software design | Accuray sent an "URGENT DEVICE CORRECTION" letter dated January 16, 2012 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter provides a work around to prevent the issue from happening. Contact Accuray Customer Support at 1-877-668-8667for questions regarding this notice. |  | 207 | 207 | Worldwide Distribution-USA (nationwide) and the countries of Korea, Hong Kong, China, Canada, Turkey, India, Russia, Italy, France, Belgium, Poland, Germany, Netherlands, United Kingdom, Spain, Greece, Saudi Arabia, Switzerland, Ukraine, Czech Republic, and Japan. | Terminated | September 28, 2012 | 256 | 207 | Not_Computer
Z-1094-2012 | 60824 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | ORIGIO MIDATLANTIC DEVICES CRYOPETTE (r) Rx only Single use only CRY-PETTE-CL LOT 123456 2013-10-10 Exp. date *** MADE IN THE USA STERILE R ***Origio, Inc. 3400 Hunter's Way, Charlottesville, VA 22911 www.origio.com |  |  | November 28, 2011 | 2011 | ORIGIO HUMAGEN PIPETS INC. | Cryopettes were rupturing during the warming process following vitrification. | Device Design | ORIGIO sent a Field Safety Corrective Action letter dated November 21, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ORIGIO feels it is important to eliminate this risk completely. As soon as we are confident that this danger of rupture has been eliminated, we will make Cryopettes available again. Please use the attached UPS call tag to return any unused Cryopettes to ORIGIO Inc. Upon receipt, we will issue a credit to your account.We sincerely apologize for the problems this issue has caused in your laboratories. For further questions please call (434) 979-4000. |  | 73850 | 73850 Cryopettes | Worldwide Distribution - U.S. included states of AL, AR, DE, PR, NY, MA, NH, VT, IL, CT, NJ, PA, VA, WV, NV, NC, SC, GA, FL, OH, MD, MS, IN, MI, PR, WI, MN, SD, MT, MO, TX, CO, UT, AZ, CA, HI, OR, AK, WA and in countries of Turkey, Saudi Arabia, Kuwait, Jordon, So. Korea, Taiwan, Brazil, So. Africa, Oman, Kazakhstan, Honduras, Slovenia, Serbia, Poland, Czech Republic, Romania, Croatia, Ukraine, Netherlands, Bulgaria, Hungary, Russia, Lebanon, Singapore, Malaysia, Vietnam, China, Thailand, Nigeria, Japan, Spain, Greece, Canada, Denmark, France, United Kingdom, Italy, Germany, and Australia, | Terminated | February 04, 2013 | 434 | 73850 | Not_Computer
Z-1097-2012 | 61206 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | BD Oral Dispensing Syringe Oral Syringe Tip Caps |  |  | October 05, 2011 | 2011 | Becton Dickinson &amp; Company | There are complaint reports for loose tip caps in various types of oral syringes | Device Design | Becton Dickinson (BD) sent an Urgent: Product Recall letters dated October 5, 2011 to all affected consignees. The letter identifies the affected product, problem and actions to be taken. Customers are instructed to follow the instructions on how to handle the affected product: 1) Check their inventory to determine if they have any Oral Syringe Tip Caps from the impacted lot numbers. 2) If they have any of the impacted product return the Oral Syringe Tip Caps to BD for replacements as directed on the attached recall packing list. 3) If they use the Oral Syringe Tip Caps and require replacement, follow the instructions on the attached form. 4) If they have no inventory of the affected lots, indicate this on the website. |  | 22827500 | 22,827,500 units | (USA) Nationwide Distribution | Terminated | February 01, 2013 | 485 | 22827500 | Not_Computer
Z-1092-2012 | 61121 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE, ONCOR and PRIMUS Linear Accelerators with RT Therapist Workspace, version 4.1. Manufactured by Siemens AG, Kemnath, Germany. The indication is to deliver X-ray, photon and electron radiation for the therapeutic treatment of cancer. |  |  | January 22, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Software upgrade to installed base to the latest release of the RT Therapist version 4.1.144 which corrects multiple safety issues that could lead to patient mistreatment. | Software design | Siemens customers with active devices will receive the software update through the distribution of the "Update Instructions" with a release date of January 22, 2012. The software update will be delivered and installed by a qualified Field Service Engineer utilizing a SW kit and instruction distributed by the Update instruction. Consignee notification of the Update Instructions will be delivered by Customer Service or by certified mail. |  | 182 | 182 active devices | Worldwide Distribution - USA, including the states of WI, UT, PA, TN, OH, LA, DE, MI, NJ, FL, NE, WV, MO and the countries of Germany, Spain, Poland, Australia, Malaysia, New Zealand, India, Republic of Korea, Canada, Italy, Egypt, Philippines, France, Belgium, Saudi Arabia, South Africa, Norway, Turkey , Hungary, Russian Federation, P.R. China, Lebanon, Iran, and the United Kingdom. | Terminated | November 07, 2012 | 290 | 182 | Computer
Z-1090-2012 | 61074 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian 52-leaf Standard MLC multileaf collimator; |  |  | January 16, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | When using the Standard and Millennium 52-leaf MLC multi-leaf collimators for beam shaping, under specific treatment conditions, dose may be delivered outside of the intended treatment field that exceed the expected transmission dose. | Device Design | Varian Medical Systems sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated January 16, 2012 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter provides recommended actions when planning and treating dynamic fields on the 52 leaf MLCs. Contact the Help Desk at 1-888-827-4265 for questions regarding this recall. |  |  |  | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Canada, China, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, United Kingdom, Singapore, Bahrain, Egypt, Hong Kong, India, Indonesia, Israel, Malaysia, Morocco, Nepal, South Africa, South Korea, Taiwan, Turkey, Ukraine, United Arab Emirates, Japan, Brazil, Chile, Colombia, Mexico, Panama, Venezuela, Saudi Arabia, and Australia. | Terminated | July 26, 2012 | 192 |  | Not_Computer
Z-1096-2012 | 60945 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDE Patient management System, Model 6488, software version 7.2. It is indicated for use to remotely communicate with a compatible Boston Scientific pulse generator and transfer data to a central database. |  |  | December 19, 2011 | 2011 | Boston Scientific CRM Corp | Boston Scientific is conducting a recall involving the LATITUDE Patient Management System - Model 6488, software version 7.2 because patients are not able to perform setup. There is no risk to health associated with this issue. | Software design | Beginning on December 19, 2011, for the patients that have attempted to set-up their communicators, Boston Scientific called the patients and their associated LATITUDE clinics to let them know that they can now complete the set-up. |  | 1 | 1 | Nationwide Distribution, USA - including the states CA, MA, MI, NC, OK, PA, RI, TX, and WA. | Terminated | October 25, 2012 | 311 | 1 | Computer
Z-1104-2012 | 61048 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Presource Custom Sterile Surgical and Procedure Kits containing Ethicon Ethilon Black Nylon Nonabsorbable Surgical Sutures in new Tyvek packaging; surgical trays; Cardinal Health. The kits are custom packed with all of the specific disposable class 1 and 2 devices required by the customer for specific types of surgery and medical procedures. These packs are then ethylene oxide sterilized and are ready to be used by the customer for the specific procedure. |  |  | January 27, 2012 | 2012 | Cardinal Health | The Presource kits were assembled with Ethicon Sutures in new Tyvek packaging which could no longer be placed in kits that will undergo ethylene oxide sterilization. | Packaging change control | The firm, Cardinal Health, sent an "URGENT: PRODUCT RECALL" letter dated January 26, 2012 to the affected consignees via UPS next day air. The letter identified the product, problem, and actions to be taken. Enclosed with the letter were recall labels, instructing the customer not to use the affected sutures and to remove and destroy the sutures upon opening the kit, to affix recall label to the front side of each affected Presource kit and shipping carton currently in their inventory. The customers were also requested to complete and fax back to Cardinal at 847-689-9101 or 614-652-9648 regardless of whether or not you have inventory. For component credit/replacement the consignees were instructed to contact the following: Presource Sales Operations at 800-766-0706; Government Customer Service at 800-444-1166; and Distributor Management Group at 800-635-6021. For any questions, require additional labels or desire special assistance relating to this recall contact Cardinal Health Quality Systems, 800-292-9332. |  | 18300 | 18,300 kits | Worldwide distribution: USA (Nationwide) including the states of: Alabama, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Illinois, Indiana, Kentucky, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Jersey, New Mexico, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, West Virginia, Wisconsin and Washington; and country of Canada. | Open, Classified |  |  | 18300 | Not_Computer
Z-1099-2012 | 61180 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Retcam3 &amp; Retcam Shuttle with version 6.0 software. |  |  | February 02, 2012 | 2012 | Clarity Medical Systems Inc | Still images extracted from video recordings may be mislabeled as to right and left eye. | Software design | Clarity Medical Systems, Inc. sent a Safety Alert letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customer was instructed the following :1. Do not use the video mode until the software has been updated.2. Install the software update using the enclosed CD. The instructions for installation are also enclosed. This software update will correct the anomaly.3. If you have extracted still images from video recordings, review all potentiallyaffected patient files to determine if any right eye images are mislabeled as left eye images and/or left eye images are mislabeled as right eye images. If any images are mislabeled, make the appropriate corrections in the annotation section of the affected file. Consider whether there is any diagnostic impact or need to re-irnage the patient.We apologize for the inconvenience. If you have questions or concerns, please contact Clarity at 800-215-6005. |  | 33 | 33 RetCam Shuttles and 20 RetCam 3 - total 53. | Worldwide Distribution -- USA (nationwide) including the states of TN, MO, TX, FL, NJ and Puerto Rico and countries of Italy, Portugal, Czech Republic, Thailand, Chile, Oman, Hong Kong, Russia, Philippines, Sweden, Finland, Ecuador, Australia and Canada. | Terminated | June 07, 2012 | 126 | 33 | Computer
Z-1103-2012 | 61082 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Medical Zippie TS wheelchair |  |  | January 13, 2012 | 2012 | Sunrise Medical (US) LLC | One Zippie TS Wheelchair was made per customer order was shipped with transit brackets for use a seat in a motor vehicle, however the chair was not configured with a back tested seat for motor vehicle use. | Employee error | Sunrise Medical contacted customer by telephone on January 17, 2012 to arrange for removal of transit brackets. The Sunrise Sales rep then visited the customers on January 18, 2012 and removed the bracket from the chair making it appropriate for intended use.If you need any additional information or have any questions, call 559-348-2572 or email: laurie.roberts@sunmed.com. |  | 1 | 1 chair | Nationwide distribution: OK only. | Terminated | February 27, 2012 | 45 | 1 | Not_Computer
Z-1101-2012 | 60981 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Oral-Eze (R) Oral Fluid Collection System, manufactured by Capitol Vial, Inc., Auburn, AL 36832 |  |  | October 07, 2011 | 2011 | Capitol Vial Inc | Low buffer fill volumes on Oral-Eze, an oral fluid collection system. | Device Design | Thermo-Fisher Scientific/Capital Vial sent a Recall Notification letter dated October 7, 2011 to the sole consignee. The letter identified the affected product, problem, result impact and necessary actions to be taken. The consignee was instructed to discontinue use of the affected product and follow the necessary actions provided. If you have any questions regarding this information, call Michael Hect, Manufacturing Manager through Customer Service at (800) 772-8871. |  | 135000 | 135,000 units | (USA) Nationwide Distribution - in the state of KS | Terminated | August 24, 2012 | 322 | 135000 | Not_Computer
Z-1102-2012 | 60976 | No | DRT | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | V Series Monitor |  |  | October 14, 2011 | 2011 | Mindray DS USA, Inc. d.b.a. Mindray North America | A system database corruption may occur that causes the V Series Monitor to reset. | Process control | The firm, Mindray, sent an "URGENT PRODUCT CORRECTIVE ACTION" letter dated October 14, 2011 on October 17, 2011 via certified mail, return receipt requested, signature required to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to contact their Mindray Service Representative at 1-800-288-2121, to arrange for software upgrade of the V Series Monitor(s) in their facility.If you have any questions, contact Manager, Regulatory Affairs and Clinical Affairs at 201-995-8407 or email: d.arpino@mindray.com. |  | 695 | 542 units US ; 153 units International | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, IN, MO, NE, PR, and SD; and countries of: Australia, Chile, Columbia, United Kingdom, and Venezula. | Terminated | June 11, 2013 | 606 | 695 | Computer
Z-1100-2012 | 61235 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Breeze Self-Adhesive Resin Cement, part number: N97F. The intended use is for cementation of porcelain, resin, and metal-based inlays, onlays, crowns, bridges, and posts. |  |  | December 02, 2010 | 2010 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that the Breeze Self-Adhesive Resin Cement exhibits faster than expected gel and set times. If the work time is exceeded when using this material, it could possibly result in inadequate bond strength. | Nonconforming Material/Component | An "Urgent Medical Device Recall" letter was sent on 12/17/2010 via USPS 1st class mail to all their customers. The letter provides the customers with an explanation of the problem identified and the action to be taken. Customers were instructed to complete and fax back the enclosed recall return form. Customers with questions were instructed to call (800) 551-0283. |  | 622 | 622 units total (12 units in US) | Worldwide Distribution - Nationwide (USA) and the countries of Vietnam, Jordan,Turkey, and Italy. | Terminated | February 24, 2012 | 449 | 622 | Not_Computer
Z-1106-2012 | 61049 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Premier Care Recliner Model 5570. |  |  | October 17, 2011 | 2011 | Winco Mfg., LLC | On 10/17/2011 Winco Mfg. LLC, Ocala, FL initiated a recall of The Premier LifeCare Recliner and Premier Care Recliner Models 5400 and 5570 distributed between January 4, 2011 and October 12, 2011. The Position Lock Pin(s) could potentially fail if exposed to excessive force when the chair is in the Locked Recline position. This failure will result in an inoperable Position Lock Bar feature. | Device Design | Winco Mfg. sent a "POSITION LOCK PIN RECALL NOTIFICATION" letter dated October 24 , 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers. The letter instructs customers to complete and return an enclosed response form and provide product disposition. Winco will ship retro-fit parts to install on the affected product. Instructions will explain step by step, how to implement the fix. Contact Winco at 800-237-3377 for questions regarding this recall. |  | 79 | 79 chairs | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, KS, MA, MI, MO, MS, NC, NM, NY, OK, PA, SC, SD, TN, TX, VA, WI, and the country of Canada. | Terminated | May 04, 2012 | 200 | 133 | Not_Computer
Z-1111-2012 | 60790 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Discovery NM 630 |  |  | January 11, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with the collimator cart and lead cover of your Discovery NM/CT 670 and/or Discovery NM 630 system that may impact operator safety. 1.The cart locking pin release and cart disengagement could result in the collimator falling down and could lead to operator or service personnel injury.2.Another issue is that the | Nonconforming Material/Component | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated January 13, 2012 to its Consignees/Customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The customers were instructed to continue to use product, ensure that the collimator exchange monitor the collimator cart are properly locked in place; pay attention to any unexpected cart motion during the exchange procedure; cease the process immediately if the cart has moved from the docking spot, dock cart in place again, and complete the exchange; and if a 3rd party service engineer is servicing rather than the firm, GE, advise according to the instructions in letter.A GE Healthcare Service representative will perform the required hardware update on each affected system.If you have any questions or concerns regarding this notification, contact your local GE Healthcare Service Representative. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 10 | 10 | Worldwide distribution: USA (nationwide) including states of: CA, FL, GA, IL, NJ, NY, NC, TX, and WI; and countries of: AUSTRALIA, BELGIUM, CANADA, CHINA, DENMARK, FRANCE, GERMANY, INDIA, ISRAEL, ITALY, KOREA, JAPAN, NETHERLANDS, NORWAY, POLAND, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, UNITED KINGDOM. | Terminated | December 06, 2012 | 330 | 89 | Not_Computer
Z-1109-2012 | 61006 | No | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | BBL[TM] Lysine Iron Agar, catalog #211363,500 g bottle 2014-03-31. Lysine Iron Agar was developed for use in the differentiation of enteric organisms based on their ability to decarboxylate or dominate lysine and to form hydrogen sulfide. Becton, Dickinson and Company, Sparks, MD 21152 USA 38800 Le Pont de Claix, France. |  |  | January 11, 2012 | 2012 | Becton Dickinson &amp; Co. | Lysine Iron Agar may not produce the correct biochemical reactions with Quality Control organisms over the shelf life of the product.. | Process design | BD Diagnostics Systems sent an "URGENT PRODUCT RECALL" letter dated January 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard any remaining packages. A form was attached for customers to complete and return via fax to: Attention: Regulatory Compliance at 410-316-4258. Contact the BD Technical Services Department at 1-800-638-8663 for questions regarding this recall. |  | 33 | 33 bottles | Worldwide Distribution-USA (nationwide) including the states of AL, AR, CA, CO, CT, DC, FL, GA, IL, IN, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OR, PA, SD, TX, VA, and WI, and the countries of Belgium, Canada, Chile, China, Columbia, Costa Rica, Mexico, El Salvador, Ecuador, Japan, Manila, Singapore, Peru, Taiwan, and Venezuela. | Terminated | March 30, 2012 | 79 | 297 | Not_Computer
Z-1057-2012 | 60992 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart G3 Automated External Defibrillators (Non-Wearable) and include the following models: Powerheart 9300A, 9300E, 9300P, 9390A, 9390E; CardioVive 92532, 92533; Nihon-Kohden 9200G, 9231; and GE Responder 2019198, 2023440 AEDs. |  |  | January 13, 2012 | 2012 | Cardiac Science Corporation | The affected Automated External Defibrillators (AEDs) contains a circuit board manufactured with a component that may fail unexpectedly due to a supplier manufacturing defect. If the component were to fail during a rescue attempt, the AED may not deliver defibrillation therapy. | Process control | Cardiac Science sent an "URGENT-Voluntary Medical Device Recall" letter dated January 13, 2012 to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to immediately return the device to the manufacturing location for repair. The affected devices will receive a hardware correction, and the same serial number device returned to the consignee in most cases.Customers in the US can call 1-888-402-2484 and customers outside the US can call 1-425-402-2482 or email at aed210@cardiacscience.com to arrange delivery of the shipping materials.Customers can also visit www.cardiacscience.com/aed210 for more information about AEDs affected by this recall. |  | 11019 | 11,019 units (6142 units in the US and 4877 units international) | Worldwide distribution: USA (nationwide) and Puerto Rico; and to the following countries: Australia, Belgium, Brazil, Canada, China, Colombia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Malaysia, Maldives, Malta, Netherlands, New Zealand, Norway, Panama, Peru, Poland, Portugal, Qatar, Romania, Russia, Saint Lucia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, Ukraine, and United Kingdom. | Terminated | September 27, 2013 | 623 | 11019 | Not_Computer
Z-1117-2012 | 60831 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | VITROSÂż Chemistry Products PHBR Slides, REF 822 1384 |  |  | June 23, 2011 | 2011 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints of intermittent imprecision, outliers, or shifts in control fluid values when using various lots of VITROSÂż Chemistry Products PHBR Slides (Phenobarbital) from Coating 0053. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated June 23, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue using and discard all remaining lots of the affected product. A Confirmation of Receipt Form was attached so customers may complete and return to the firm. The firm recommends that customers consult with their Laboratory Medical Director and requesting physician to resolve any concerns they may have regarding previously reported results. Contact Customer Technical Services at 1-800-421-3311 for questions regarding this notification. |  | 2977 | Domestic - 2345 sale units, Foreign - 632 sales units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 2977 | Not_Computer
Z-1114-2012 | 61123 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics BENEPHIT CV Infusion System Targeted Renal Therapy Infusion System --- Catalog No. [REF] 60040 --- Manufactured in USA --- [Mfg symbol] AngioDyanamics, 603 Queensbury Avenue, Queensbury, NY USA 12804. |  |  | February 06, 2012 | 2012 | Angiodynamics Worldwide Headquarters | The Benephit CV Infusion System contains a Sheath Y Connector that has the potential to develop cracks. | Nonconforming Material/Component | Field Safety Notices (FSNs), entitled Urgent - Medical Device Recall, were sent to consignees on 02/06/2012 via USPS certified mail, return receipt requested. The letters identified the affected devices, the issue, hazard involved, and actions to be taken by the customer and AngioDynamics. The FSN directs consignees to identify and segregate the recalled lots, returning any affected product to AngioDynamics, Inc., Queensbury, NY 12804. The Recall Reply Form should be completed and faxed to 518-798-1360. --- If customers have distributed the affected product/lot, they are instructed to inform their customers of this recall action immediately, and then return the recalled units. --- If customers have any additional questions or concerns or need more detailed instruction on how to comply with this notice, please contact the local sales representative or call Julie Blair, Customer Service Manager, at ANGIODYNAMICS Customer Service at 1-800-772-6446. Customers may also e-mail questions to customerservice@angiodynamics.com |  | 4 | 4 units | Nationwide Distribution -- Florida, Iowa, Maryland, New York, and Texas. | Terminated | February 10, 2016 | 1465 | 15 | Not_Computer
Z-1116-2012 | 61031 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer Genetic Screening Processor (GSP) Product code: 2021-0010. It is intended for in vitro quantitative / qualitative determination of analytes in body fluids. |  |  | January 09, 2012 | 2012 | PerkinElmer Life and Analytical Sciences, Wallac, OY | GSP instruments with defective washer modules may cause a false negative screening result being reported for T4, TSH and IRT. For 17-OHP, a false positive result may be generated if the failure occurs for a normal sample. | Nonconforming Material/Component | Perkin Elmer issued an "Urgent Medical Device Recall" letter dated 1/25/12 to all customers and informed them about the possible issue and requested permission for service personnel to check their instrument. The customers were instructed that the instrument could not be used and the washer modulerequired replacement. Customers were told to expect to be contacted by the firm's local representative for further details. |  | 10 | 10 units | Worldwide Distribution - USA, including the states of Florida, Ohio, Oregon, Texas and the countries of Brazil, Belgium, France, Germany, and the UK. | Terminated | March 06, 2013 | 422 | 10 | Not_Computer
Z-1115-2012 | 61239 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Brand Name: Fusio Liquid Dentin, part number N21SB, N21SC. |  |  | August 19, 2010 | 2010 | Kerr Corporation | A recall was initiated because Pentron Clinical has confirmed that the Fusio Liquid Dentin is not as flowable as ideally expected. | Nonconforming Material/Component | Pentron Clinical sent a Medical Device Recall letter dated September 8, 2010 (via USPS 1st class mail) to all affected customers. The letter identified the affected product, the reason for the recall and the actions to be taken. Customers were instructed to examine their inventory for the affected product, return any remaining in stock and complete and fax back the enclosed recall return form. Customers were advised to contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 382 | 382 units total (207 units in US) | Worldwide Distribution - USA (nationwide) in the countries of Austria, Bulgaria, France, Great Britain, Netherlands, Japan, Lebanon, Malaysia and South Africa. | Terminated | February 28, 2012 | 558 | 382 | Not_Computer
Z-1126-2012 | 61184 | No | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems (a Nxstage Company) Hemodialysis Fistula Needle Set with MasterGuard Ant-Stick Needle Protector 15 GA x 1" |  |  | January 31, 2012 | 2012 | Medisystems Corp | Mislabeled outer carton: Outside shipping cartons is stamped 15Ga x 1", product contained within properly labeled as 17Ga x 1" needle | Employee error | Medisystems Corporation issued notice of recall to distributors on January 31, 2012. Distributors to notify affected end users by letter. Each letter was accompanied by a response form and instructions on how to return the response form. End users were given the option of returning involved product for credit or replacement. Contact Medisystems Quality Assurance Manager at (978) 332-5973 for information. |  | 20750 | 20,750 units (83 cases) | USA (nationwide) including the states of AZ, CA, CO, OR, NV, NY, TN, and WA. | Terminated | September 12, 2012 | 225 | 20750 | Not_Computer
Z-1015-2012 | 60995 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | B. Braun Infusomat Space Pump - US Version Infusion pump. Model number 8713050U. |  |  | July 12, 2011 | 2011 | B. Braun Medical, Inc. | If the IV set anti free flow clip is incorrectly inserted into the pump and the pump door forced closed, the clip catch on the inside of the pump door may break. If the clip catch is broken and the door opened, free flow protection is ensured. If IV set is then removed from pump without closing the roller clamp, the clip may not re-engage the tubing, leaving an open fluid path with free flow pote | Device Design | B Braun sent a letter on August 16, 2011, to their customers, that detailed the proposed G03 implementation in response to reported customer complaints. The software upgrade with the door jam was agreed upon and the eight hospitals were upgraded in November 2011. |  | 5769 | 5769 | Product was distributed to hospitals in CA and MO. | Terminated | November 03, 2014 | 1210 | 5769 | Computer
Z-1127-2012 | 61141 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza with software version VA20B_HF04 |  |  | January 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction when using syngo.plaza VA20B_HF04 with USB dongle and archive license. In syngo.plaza VA20B_HF04, there is a scenario where data in Short-Term Storage (STS) could reaming unprotected and not archived. this can happen when an archive server with a USB dongle-based license is enabled with auto archiving. If the disk fill level is reached which triggers auto deleti | Software design | Siemens sent a Customer Safety Advisory Notice letter dated January 20, 2012 to all affected consignees. The letter identified the affected product, problem, steps to take to avoid this issue, and actions to be taken. Customers were instructed to contact their local Siemens Service to plan for their software update and ensure that this safety advisory is placed in the system's instruction for use. |  | 7 | 7 | Nationwide Distribution (USA) - CA, FL, GA, MO, NY and WA. | Terminated | September 29, 2012 | 253 | 7 | Computer
Z-1129-2012 | 59285 | No | REK |  |  |  |  | DirectFRAP (Direct Flourescence Recovery After Photobleaching) Imaging System is an extension unit for Axio Observer that is used for microscopic analyses of dynamic processes within living cells. The DirectFRAP module is a component used in conjunction with Laser Scanning Microscopes, Non-Medical Devices for Research Use. The DirectFRAP Imaging System is a Class 3B laser system. |  |  | April 28, 2011 | 2011 | Zeiss, Carl Inc | The manufacturer confirmed that the opto-semiconductor sensors can be triggered by exposure to incandescent lighting. The safety interlock is in place to prevent unintended exposure to Class 3B laser radiation. | Nonconforming Material/Component | Carl Zeiss MicroImaging LLC, through trained personnel, first visited each user site, provided additional user instructions, re-labeled the components to provide adequate instruction and warning and further placed a label over the access to the locking screw which secures the component to the microscope. Carl Zeiss Inc. issued "IMPORTANT LASER SAFETY NOTICE: URGENT: FIELD CORRECTIVE ACTION" notices to each user by express courier in late April 2011, Carl Zeiss personnel will confirm that the relabeling has been performed and updated instructions for use have been provided. For questions, please contact the company.The firm will implement 2 actions. Service Representatives will visit each user site, provide additional user instructions, relabel the components to provide adequate instruction and warning, and add a label over the access to the locking screw which secures the component to the microscope. Secondly, when new replacement sensors are available, Carl Zeiss field service technicians will perform the component replacements. Modified componentsare anticipated to be available within 6 months and replacement completed by April 2012. |  | 11 | 11 units | Worldwide Distribution -- USA (nationwide) | Terminated | December 11, 2012 | 593 | 111 | Not_Computer
Z-1125-2012 | 61043 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips Intellispace Portal, Release 4.0.0. |  |  | January 27, 2012 | 2012 | Philips Medical Systems Nederlands | There maybe incorrect values in the diffusion/perfusion maps. | Software design | Philips issued URGENT - Medical Device Correction Notifications, dated January 20, 2012. The letter identified the affected product, all Intellispace Portal with software release 4.0.0 Installed, and gave a description of the problem. Users were provided with the actions to be taken prior to the software upgrade( to version 4.0.1). Users are to follow the instructions provided in order to generate maps with the correct values and to make sure that the setting is correct before using the generated data. Philips will contact customers in order to install the update free of charge. Customers should contact their local Philips representative at 866-767-2822 |  | 25 | 25 units | Worldwide Distribution -- US, including states of IN and TX, and countries of Australia, Austria, Czech, Denmark, France, Germany, Indonesia, Netherlands, Poland, Spain, and Switzerland. | Open, Classified |  |  | 25 | Computer
Z-1118-2012 | 61187 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Liebel-Flarsheim Optistar MR 60 mL Syringe w/lateral flow needle, Mallinckrodt, Cincinnati, OH 45237. Made in Mexico. Devices packaged in Plastic and Tyvec pouch, with 50 pouches per case. |  |  | February 10, 2012 | 2012 | Mallinckrodt Inc | The machine used to package the product was missing a tool, which could prevent the packaging from sealing properly, causing the sterility of the product to be compromised. | Packaging process control | Covidien sent an Urgent Device Recall letter dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product immediately. Customers were to use the attached business response form to report the quantity of each lot that is currently in their stock. If customers did not have any of these lots in stock, they should indicate by checking the appropriate box on the business report form. When completed, customers should email the completed form to covidien2600@stericycle.com or fax to 888-257-7910. For any questions customers should call 866-761-9571. |  | 665 | 665 Cases | Worldwide Distribution - USA including Distributed worldwide: USA including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, Puerto Rico and the countries of Canada, Australia, Belgium, Brazil, Chile, Ecuador, Guatemala, Mexico, Panama, Singapore, and Uruguay. | Terminated | December 24, 2013 | 683 | 665 | Not_Computer
Z-1133-2012 | 61185 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | BardPort M.R.I. Implanted Port with Attachable 6 Fr. ChronoFlex Open-Ended Single-Lumen Venous Catheter and Peel-Apart Introducer Kit. Subcutaneous implanted intravascular infusion port and catheter is used to provide repeated access to the vascular system for the delivery of medications, intravenous fluids, parenteral nutrition solutions, and blood products. They are also indicated for the withdrawal of blood samples. |  |  | February 16, 2012 | 2012 | Bard Access Systems | The end flap label on the kit incorrectly describes the product as "Bard Port" MRI DUAL Implanted Port instead of a "Bard Port MRI Implanted Port" as correctly indicated on the main label of the kit. | Labeling Change Control | A "Labeling Discrepancy Notification" letter was sent to US consignees via FED EX on February 16, 2012 and on February 20, 2012 for the foreign consignee explaining the labeling discrepancy and providing a visual example of the incorrect label. The customers were given the option of continuing to use the product with the understanding the end flap label was incorrect or they could return the affected products for replacement (Call 1-800-290-1689 to arrange a replacement). |  | 545 | 545 units (Lot RESF0856 - 150 units; Lot RESH0067 - 295 units; Lot RESH1150 - 100 units) | Worldwide Distribution -- USA, including the states of CA, CT, FL, KS, IL, MI, MN, OH, OK, SD, TN, TX, VA, WA and the country of Belgium. | Terminated | March 30, 2012 | 43 | 545 | Not_Computer
Z-1132-2012 | 60888 | No | INX | Air-fluidized bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | Clinitron Rite Hite Bed headboard and foot side rails. |  |  | February 03, 2012 | 2012 | Hill-Rom, Inc. | The firm identified in October 2011 that the Clinitron Rite Hite Bed and Clinitron At Home Bed headboards failed to meet the safety standard force requirements. The firm identified in October 2011 that the Clinitron Rite Hite Bed and Clinitron At Home Bed outer rail, which is extruded Aluminum tubing, is intended to be one piece joined together by a weld at the lower center. The parts receiv | Nonconforming Material/Component | Hill-Rom, Inc. sent a Modification Notice Urgent Medical Device Correction Notification dated February 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The modification notice instructed all customers to find and replace all headboards and food side rails that do not meet design intent. All technicians will inspect all affected product in the service centers and all at home and rite hite beds for an "x" or "L" stamp. The technicians will inspect the beds that are still in use during scheduled maintenance or before each rental. All will be completed before 01-JUL-12.For support call 1-800-445-3720 |  | 103 | Foot Siderail = 70; Headboards = 103 | Worldwide Distribution -- USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV, and WY. and the country of Canada. | Terminated | December 05, 2012 | 306 | 103 | Not_Computer
Z-1130-2012 | 60980 | No | DIS | Barbiturate test system. | Toxicology | Toxicology | 510(k) | Product Brand Names: TriageÂż Drugs of Abuse Panel 8 Test Kit (K924476), TriageÂż Drugs of Abuse |  |  | January 04, 2012 | 2012 | Alere San Diego | Alere San Diego received a rise in customer complaints regarding false positive Barbiturates (BAR) results on the Triage Drugs of Abuse Panel Plus TCA. | Nonconforming Material/Component | Alere sent a "PRODUCT NOTIFICATION" letter dated January 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Verification Form was enclosed for customers to complete and return to the firm. Contact Technical Services at 1-877-308-8287. |  | 693054 | 693054 kits | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, EU, Guatemala, Japan, Lebanon, and Saudi Arabia. | Terminated | July 05, 2013 | 548 | 693054 | Not_Computer
Z-1095-2012 | 61070 | No | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Inspira AIR Balloon Dilation System, |  |  | January 30, 2012 | 2012 | Acclarent, Inc. | Acclarent received reports of difficulty deflating the balloon during the procedure, which could potentially result in airway obstruction. | Component design/selection | Acclarent sent an "Urgent Voluntary Product Recall" letter dated January 31, 2012 by US mail, return receipt requested to all affected customers. Additionally, customers known to have scheduled procedures using the affected product were notified by phone to advise them of the recall. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, discontinue use and return any remaining in stock per the instructions provided. The letter provides additional information and photographs illustrating how to identify the recalled product. Customers were advised to complete the enclosed Business Reply Card regardless of whether they have or do not have product subject to this recall in their possession and mail to Stericycle. Customers were instructed to distribute this information to all staff within their department who uses the affected product. For further questions call 1-866-781-1173 or contact your Acclarent Sales Representative. |  | 633 | 633 units | Nationwide Distribution (USA) only sixe 18 x 40Worldwide distribution other sizes.Distirbution to France, Canada, Saudi Arabia | Terminated | August 29, 2013 | 577 | 633 | Not_Computer
Z-1137-2012 | 61073 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | Perfector |  |  | January 06, 2012 | 2012 | Ion Genius, Inc. | Wires do not meet performance standards under 21 CFR 898 and have unprotected electrode configurations. Additionally, devices were sold/distributed without 510(k) or PMA | Nonconforming Material/Component | Ion Genius sent a Medical Device Product Recall letters to all affected customers on January 6, 2012. The letter identified the affected products, problem and actions to be taken. The letter instructed customers to check their inventory for the affected products and return immediately to Ion Genius for replacements. Customers returning affected products should call the firm so that a shipping label can be provided.Customers were instructed to be completed the Response Form and returned to Ion Genius by fax at 808-395-0787. For questions and concerns contact Xanya Sofra-Weiss at 808-222-2664 or email x@arasysperfector.com. |  | 72 | 72 cables in total, amounts of other devices not reported | Worldwide Distribution - (USA) Nationwide Distribution | Terminated | May 21, 2015 | 1231 | 72 | Not_Computer
Z-1135-2012 | 61107 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, LATITUDEÂż Patient Management System Communicator, Models 6476. |  |  | January 19, 2012 | 2012 | Boston Scientific CRM Corp | One (1) Model 6476 LATITUDE Communicator is unable to interrogate the patient's pulse generator. Boston Scientific has determined that this is due to a patient receiving an incorrect communicator. Therefore the communicator did not find and was unable to interrogate the assigned patients PG. Additionally when the patient attempted to set-up this incorrect communicator, the system reflected | Packaging process control | On January 19, 2012 Boston Scientific telephoned the affected patient and their clinic. The patient was notified that the communicator they received will not be able to activate successfully and that their communicator and sensors would be replaced. The clinician was informed of the patient's communicator was sent incorrectly to the other patient who attempted to activate it, the monitored status of the website was incorrect and that a new communicator would be shipped to the patient that will work with their device. |  | 1 | 1 | Nationwide distribution: UT only. | Terminated | July 05, 2012 | 168 | 1 | Not_Computer
Z-1136-2012 | 61221 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Brand name: Orthos CM Bracket, Orthodontic Metal Bracket, |  |  | February 13, 2009 | 2009 | Ormco Corporation | The recall was initiated because Ormco has confirmed that manufacturing error resulted in the Orthos CM Brackets being incorrectly manufactured with a torque of -9Âż. The packaging was labeled with +9Âż torque; however the brackets contained inside the packaging are actually Orthos CM Brackets with a torque of -9Âż. | Process control | Ormco Corporation sent an Urgent Field Safety Notice dated February 13, 2009, to all affected customers. The recall communication for consignees in the United States, Canada, Colombia, HongKong, Kuwait, United Arab Emirates, Mexico, Russia, South Africa and Croatia was sent on February 13, 2009 via USPS First Class Mail. The recall communication for consignees in Japan was sent on February 17, 2009. European consignees were sent recall notification on February 23, 2009. The Australian consignees and New Zealand consignees were the sentrecall communication on February 25, 2009. Customers were instructed to contact Ormco Customer Care at (800) 854-1741 to receive an RMA number. The affected product should be returned to Ormco Corporation at the following address: Ormco Corporation, 1332 South Lone Hill Avenue, Glendora, CA 91740.Customers were instructed to label their return product "RECALLED PRODUCT Attention: Customer Returns". Customers were instructed to complete the enclosed Return Form and return it by fax to (909) 962-5605.For questions regarding this recall call 909-962-5600. |  | 304 | 304 units | Worldwide Distribution - USA (nationwide) and worldwide to Canada, Colombia, Hong Kong, Kuwait, United Arab Emirates, Mexico, Russia, South Africa, Croatia, Japan, France, Bulgaria, Germany, Estonia, United Kingdom, Hungary, Ireland, Italy, Netherlands, Australia, and New Zealand. | Terminated | March 02, 2012 | 1113 | 304 | Not_Computer
Z-1134-2012 | 61127 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Patient Monitors: |  |  | February 09, 2012 | 2012 | Philips Healthcare Inc. | Central station monitors may not alarm (either visual or audible) when an alarm is announced at the affected patient monitor | Software change control | Philips Healthcare sent a Urgent Medical Device Correction letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips has recently received reports that, under certain circumstances, alarms announced at the patient monitor are not announced (either visual or audible) at the central station. This issue affects the following IntelliVue patient monitor models with SW rev. H.xx.xxMP20 (M8001A), MP30 (M8002A), MP40 (M8003A), MP50 (M8004A), MP60 (M8005A), MP70 (M8007A), MP80 (M8008A), MP90 (M8010A), Intelligent Display D80 (M8016A), MX600 (865242), MX700 (865241) and MX800 (865240)If this issue occurs, the primary alarm function at the bedside monitor is not affected.Philips is conducting this voluntary correction to upgrade software on affected devices. Please refer to the following pages, which provide instructions for actions to be taken. Follow the Action to be taken by Customer/User section of the instructions. This issue has been reported to the appropriate regulatory agencies.Ensuring that you have the highest quality medical devices, accessories and supporting documentation is our top priority. Your satisfaction with Philips products is very important to us.Review this information with all staff members who interface with the central station and need to be aware of the contents of this communication.The Product Number and Serial Number is contained on the devices product label, located on the front of the device. The SW revision can be accessed via the Revision Screen at the bedside monitor. charge.A Philips Healthcare representative will contact customers with affected devices contact Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this correction to arrange an upgrade of the Intellivue software. |  | 83534 | 83,534 units | Worldwide Distribution -- USA (nationwide) and the countries of ALBANIA, ANGOLA, ANTILLES, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BYELORUSSIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, GHANA, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LEBANON, LIBYAN, ARAB, JAMAHIRI, LITHUANIA, LUXEMBOURG,MACAU, MALAYSIA, MOROCCO, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NEPAL, NETHERLANDS, ANTILLES, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PALESTINE, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIET NAM, YEMEN, ARAB REP and ZAMBIA | Completed |  |  | 83534 | Computer
Z-1143-2012 | 61100 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OSS AVL Poly Tibial Bushing Set |  |  | February 07, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after complaints were received where OOS AVL Bushing Sets were opened during OSS knee replacement procedures, and were found to contain two tabbed bushings. | Employee error | BIOMET Orthopedics sent an URGENT MEDICAL DEVICE recall notice dated February 7, 2012, to all affected consignees. The recall notice instructed consignees to immediate locate and remove the devices from circulation and to fax back the response form to 574-372-1683. If product was further distributed, consignees must notify hospital personnel of this action via the enclosed notice. Consignees were charged with the location and return of the affected product. For questions regarding this recall call 574-372-3983 or 574-371-3009. |  | 27 | 27 units distributed US and foreign | Worldwide Distribution - USA including IN. KS, NJ, CA, and OHInternationally: The Netherlands | Terminated | February 06, 2013 | 365 | 27 | Not_Computer
Z-1147-2012 | 61125 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Titanium SURFIX ALPHA SCREWS, a component of the Newdeal Hallu Lock Plate System. Surfix Alpha Variable angle locking system exists in different diameters and lengths. The product can be sold as a replenishment part individually packaged sterile, or as one (1) unit which is kitted within an instrumentation tray at which point, the product is non-sterile until the tray in which it is kitted is sterilized a the designated facility. The product is indicated for use in fixation of fractures, osteotomies or arthrodesis of the first metatarso-phalangeal joint. |  |  | December 21, 2011 | 2011 | Integra LifeSciences Corp. | Manufacturing error. The part number which corresponds to the length dimension of : 14 mm. was incorrectly etched on to screws with the actual measurement of: 12 mm. | Error in labeling | Integra contacted the Sales Representatives and All USERS on December 21, 2011, with a 'SUBJECT: MEDICAL DEVICE RECALL' letter. The letter was sent via FedEx and/or E-mail. The letter describes the product, problem, and the actions to be taken by the customers and the firm. The customers are instructed to immediately examine their inventories for the presence of the recalled product, and complete and return the attached, MEDICAL DEVICE RECALL ACKNOWLEDGMENT AND RETURN FORM via E-mail or FAX #1-609-0750-7999. Any questions concerning the recall document are to be addressed to the customers' regional managers or to the E-mail address and/or telephone number provided on the recall notification. |  | 105 | 105 units | Nationwide Distribution, USA - including the states of OH, FL, WI, CT, OK, TN, OR, ID, CA, PA, MD, VA, FL, MS, NJ, LA, IA, GA, AL, MA, NY, MN, TN, IL, AR, NE, AZ, MI, WA, OR, NC, KY, MO, UT, NV, IN, SC, CO and the countries of the United Kingdom, France, Switzerland, and Austria. | Terminated | February 19, 2013 | 426 | 105 | Not_Computer
Z-1146-2012 | 59959 | No | ELW | Impression material. | Dental | Dental | 510(k) | Correct VPS Vinyl Polysilocane Impression Material, Product Number Q01B. The intended use of this device is as a dental impression material. |  |  | November 04, 2010 | 2010 | Sybron Dental Specialties | Pentron Clinical is voluntarily recalling the following lot of mislabeled Correct VPS Medium Body Tubes. It was discovered that the individual tubes inside the Correct VPS Medium Body Tube boxes are missing the lot/expiration date sticker. | Packaging process control | Pentron Clinical sent "Urgent Medical Device Recall" letters dated November 17, 2010 to consignees explaining the above mentioned public reason for recalls and asking them to determine if they have any of the affected products in their in stock. Product that is returned will be replaced at no charge. They also asked that they complete and fax back an enclosed recall return form.Questions can be directed to Customer Service at (800) 551-0283. |  | 176 | 176 units | Worldwide Distribution, USA including the states of FL, NY, WA, CA, IL, MA, NC, NJ, AZ, FL, KY, CT, PA, IN, NJ, PA, TN, NV, TX and the countries of Canada and Pakistan. | Terminated | March 14, 2012 | 496 | 176 | Not_Computer
Z-1142-2012 | 61265 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Brand name: Bond-1 Primer/Adhesive, a resin tooth bonding |  |  | May 27, 2010 | 2010 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that various lots of Bond-1 Primer/Adhesive have been found to have gelled and therefore becomes unusable. | Nonconforming Material/Component | Pentron Clinical sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 7, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to complete and return an enclosed Recall Return Form via fax at 1-877-677-8844. Contact Pentron Clinical Customer Service at (800) 551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 30276 | 30,276 units total | Worldwide Distribution-USA (nationwide) and the countries of Germany, Canada, Ecuador, Egypt, Hong Kong, India, Israel, Japan, Kenya, Lebanon, Morocco, Saudi Arabia, Singapore, El Salvador, Thailand, Turkey, Taiwan, UAE, Vietnam, Armenia, Austria, Belgium, Bulgaria, Byelorussia, Cyprus, Germany, Spain, Russia, Finland, England, Greece, Hungary, Ireland, Italy, Latvia, Netherlands, Poland, Romania, Sweden, and Serbia. | Terminated | March 02, 2012 | 645 | 30276 | Not_Computer
Z-1144-2012 | 61041 | No | OYK | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TibiaxysÂż Ankle Arthrodesis Plate Compression Guide, Model #159635, part # 159 635, Batch #: EAHZ, EAJ1, ELGV, ENQE &amp; ENQH. |  |  | November 17, 2011 | 2011 | Integra LifeSciences Corp. | During 6/30/2010 to 8/17/2011 the firm received five (5) incident complaints of breakage of the threaded part of the compression guide which potentially occurred when: 1) a compression is applied between talus and tibia. (Step 3 of the product description) or when 2) the surgeon screws the compression guide into the plate which most likely was attributed to improper surgical techniques used durin | Nonconforming Material/Component | Integra sent a "Field Safety Notice" letter dated November 17, 2011 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review the Surgical Techinique Manual and follow the recommended guiidelines to avoid breakage. ensure that this information is distributed to their surgical teams, and fill out the enclosed "Field Safety Notice Acknowledgement and Return Form" and return to bonnie.kuzbyt@integralife.com or fax to 1-609-750-7999. For questions call 609-936-6932 or email at sean.luland@integralife.com. |  | 168 | 168 units | Worldwide Distribution - USA (nationwide) including the states of; AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KY, MA, MD, MI, MO, MN, NC, NJ, NM, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA &amp; WI. Product was also shipped to the following countries: Austria, Belgium, Canada, Czech Republic, France, Germany, Ireland, Israel, Italy, Lebanon, Saudi Arabia, South Africa, Spain, Sweden, Switzerland &amp; United Kingdom. | Terminated | February 19, 2013 | 460 | 168 | Not_Computer
Z-1145-2012 | 59740 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 42-422561 SIGNATURE QTY, 1 TKA GDE/MDL (Prosthesis, knee) SET 04-05 DGHUMPHR-NOV52LKA-J_K-00432022 SLIDEX DISTAL-MP/PREMIER TIBIAL CUSTOM-MADE DEVICE |  |  | August 04, 2011 | 2011 | Biomet, Inc. | Biomet has initiated this action following investigation that identified two (2) specific lots of part number 42-422561 were incorrectly labeled. One (1) unit (lot 025640) was received back from distributor's office as a product complaint. | Labeling mix-ups | The firm, BIOMET Orthopedics, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated January 23, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:-Immediately locate and remove the above identified Signature " Guide and Model Set.-Carefully follow the instructions on the enclosed FAX Back Response Form .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.- Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action, via the enclosed recall letter. This letter MUST be given to hospital personnel responsible for receiving recall notices. Confirm receipt of this notice by calling, 800-348-9500, extension 3983. Questions related to this notice should be directed to 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 2 | 2 units | Nationwide distribution:Virginia and Ohio only. | Terminated | September 06, 2012 | 399 | 2 | Not_Computer
Z-1148-2012 | 60657 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AdvanSync is an orthodontic device intended for the orthodontic treatment of Class II malocclusions. The AdvanSync part numbers and lot numbers affected by this recall are as follows: |  |  | October 02, 2009 | 2009 | Ormco Corporation | Manufacturing Error: There is a potential for a braze joint failure in the upper telescopic rod assembly. Although an incident could occur in which the eyelet loses connection from the rod, there is no risk of swallowing or aspirating the appliance because it will still be attached to the patient's crown. | Process control | Ormco sent a Urgent Medical Device Recall letter dated October 9, 2009,( via USPS 1st class mail) to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are instructed to return the affected lot numbers of the 5-Pack Left and Right Telescopic Rod Assemblies (600-4040 and 600-4041) and the Left and Right Telescopic Rod Assemblies contained within the AdvanSync Kit (600-4000). The affected parts are identified by the items contained inside the wells titled, "Left Rods" and "Right Rods." The remaining items within the kit are not affected by this recall notification; therefore, there is no need to return the entire kit.Customers were instructed that any product which they have in stock will be replaced at no charge or credited to their account. Customers were instructed to contact Ormco Customer Care at (800) 854-1741 (ext. 1) directly to handle the arrangements of a quick return and replacement. The affected product should be returned to Ormco Corporation at the following address: Ormco Corporation, 1332 South Lone Hill Ave., Glendora, CA 91740. Customers were instructed to label their returned product "RECALLED PRODUCT Attention: Customer Returns". In addition, customers were asked to complete the enclosed Return Form and return it by fax at (909) 962-5605. |  | 3635 | 3,635 Units (237 kits, 315-pack left assemblies and 305-pack right assemblies) | Worldwide- USA (nationwide) and the countries of Europe, Australia, Canada and New Zealand. | Terminated | March 05, 2012 | 885 | 3635 | Not_Computer
Z-1153-2012 | 61208 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5/C5 Control panel for 2 mast roller pumps 85, Part No.: 28-95-85, Sorin Group Deutschland GMBH, Lindberghstrasse 25, 80939 Munchen, Germany, Sorin Group USA, Inc., 14401 W. 65th Way, Arvada, CO 80004, USA |  |  | February 16, 2012 | 2012 | Sorin Group USA, Inc. | Touch screen may be unresponsive, inhibiting user input. | Mixed-up of materials/components | All affected US customers and distributors were notified of the Sorin S5 HLM Touch Screen Field Correction by certified mail on 02/15/12. The letter informed users of the potential for unresponsive S5 HLM touch screens and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to replace affected touch screens on site.Customers and distributors were also provided with a Response Form to confirm they had received, read and understood the Field Correction Notice. They were instructed to fax the completed form to (303) 467-6502.For questions regarding this recall call 303-467-6306. |  | 2 | 2 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Belgium, Bulgaria, China, Denmark, Germany, Ecuador, Egypt, Finland, France, Georgia, Greece, United Kingdom, Hong Kong, Indonesia, Iraq, Iran, Israel, Italy, Japan, Jordan, Canada, British Columbia, South Korea, Kuwait, La Reunion, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Morocco, Mexico, Monaco, New Zealand, Netherlands, Norway, Oman, Austria, Phillippines, Poland, Portugal, Puerto Rico, Qatar, Serbia, Russia, Saudi Arabia, Sweden, Switzerland, Singapore, Spain, South Africa, Syria, Thailand, Czechoslavakia, Tunisia, Turkey, Hungary, Vietnam, and the United Arab Ermirates | Terminated | November 09, 2012 | 267 | 187 | Not_Computer
Z-1159-2012 | 61251 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio, AXIOM Luminos dRF, AXIOM Luminos Agile/TF, Uroskop Omnia |  |  | February 15, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Firm has become aware of unintended behavior when using Ysio, AXIOM dRF, AXIOM Luminos Agile/TF and Uroskop Omnia with RAD Fluoro Uro function using software version VC10A. Due to an error in the image pipeline, images may display dark and not suitable for diagnostic. Due to this software error customers have to repeat x-ray exposure to receive diagnostic images. | Software design | Siemens sent Update Instruction XP057/11/S by mail on February 15, 2012, which releases software update package VC10D to remedy the described issues.For questions regarding this recall call 610-219-4834. |  | 58 | 58 | Nationwide Distribution including AK, CA, CO, FL, GA, IL, IN, IA, KY, MD, MA, MO, NE, NY, NC, OH, OR, PA, SC, TX, and UT. | Terminated | January 14, 2014 | 699 | 58 | Computer
Z-1160-2012 | 60285 | No | FEM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Mobile Workstations (Endoscopy Carts), Models WM-DP1, WM-NP1 and WM-WP1 |  |  | October 28, 2011 | 2011 | Olympus America Inc. | Replacement wheels may break off from the base of the cart | Equipment maintenance | OLYMPUS sent an Urgent - Device Correction letter dated October 28, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. OLYMPUS required customers indicate on the attached questionnaire the contact information for their facility so that they can schedule a convenient time for an Olympus Field Service Representative to visit. Service personnel would inspect, and replace, as necessary any user installed casters on the WM-P1 cart. For questions regarding this recall call (483) 896-5688. |  | 5 | 5 | Nationwide Distribution including AL, CA, NY and WV. | Terminated | January 14, 2014 | 809 | 5 | Not_Computer
Z-1176-2012 | 61055 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Cell Marque CD7 (MRQ-12); |  |  | January 23, 2012 | 2012 | Cell Marque Corporation | Letter received from supplier that material received in October 2009 was labeled as CD7, but was actually CD 2. | Mixed-up of materials/components | Cell Marque sent a Field Notification letter dated January 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and return the attached form indicating they have received the notification. Customers were also asked to destroy any existing inventory of the affected product. If the product was shipped to their customers, customers should confirm that they have contacted those customers. Customers were asked to complete, sign and date and return the attached Acknowledgement Form by faxing it to 1-916-746-8989 or by scanning it and emailing it to bdevos@cellmarque.com.For questions regarding this recall call 1-800-665-7284 or 1-916-746-8900. . |  | 110 | 110 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Greece, Turkey, Italy, The Netherlands, Spain, Portugal, France, Israel, S. Korea, Pakistan, Morocco, Russia and Brazil. | Terminated | October 19, 2012 | 270 | 110 | Not_Computer
Z-1174-2012 | 60937 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Synchron Phenobarbital (PHE) Reagent, P/N 469785, |  |  | August 11, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed reports of variability in results with the lots of Phenobarbital (PHE) reagent and calibrator. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 26, 2009 to all customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were instructed to share the recall information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Customers were asked to complete and return a response form. Contact Beckman Coulter Customer Support Center (Hotline) at (800) 854-3633 for questions regarding this recall. |  | 31559 | 31,559 units total (30,617 units in the US) | Worldwide Distribution--USA (nationwide) and the country of Canada. | Terminated | December 05, 2013 | 1577 | 31559 | Not_Computer
Z-1165-2012 | 61088 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Depuy Mitek Lupine Loop Plus Anchor w/Orthocord |  |  | January 27, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Lupine BR Anchor w/ Orthocord may be assembled with an incorrect suture configuration and result in a loose repair | Process control | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated January 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to immediately check all inventories to determine if they have any affected product. Customers were asked to refer to the instructions to report their inventory status, return and obtain replacement or credit for the product. For questions regarding this recall call 1-877-379-4871. |  | 32408 | 32408 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, AE, Argentina, Australia, Austria,Belgium,Brazil, Bahrain, Chile, China, Columbia,Cyprus, Check Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Finland, Great Britain, Guadeloupe, Hong Kong, Hungary,Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Luxembourg,Latvia, Martinique, Montenegro, Mexico, Maylasia, Netherlands,Norway, New Zealand, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Tunisia, Turkey,Uruguay, Venezuela, and South Africa | Terminated | August 05, 2015 | 1286 | 190853 | Not_Computer
Z-1175-2012 | 61196 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and True Beam STx versions 1.0 through 1.5, Model number H19; Varian Medical Systems, Palo, Alto, CA 94304. TrueBeam Radiotherapy Delivery System is intended to provide stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation treatment is indicated. |  |  | January 30, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the respiratory gating software of TrueBeam. When importing breath-hold gating protocols, the gating thresholds can be reset to default values, rather than retaining the thresholds established during planning. | Device Design | Urgent Medical Device Correction/Urgent Field Safety Notices were sent on January 30, 2012 by certified mail to US and Canadian consignees.***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***Varian Sales, Marketing and Service Organizations also were informed via letter the same day. The Field Safety Notice with customer distribution list was e-mailed to Varian International Customer Support Managers for translation (if necessary) to be sent to affected sites with return proof of notification to be sent to Varian. The notice informed users of the recommended actions. Questions were directed to the Varian Help Desk 1-888-827-4265 (US and Canada). |  | 382 | 182 units ***12-12-12 AMENDED TO: 200 units*** | Nationwide Distribution (USA) and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Netherlands, New Zealand, Norway, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Thailand, Turkey and the United Kingdom. | Terminated | April 16, 2014 | 807 | 382 | Computer
Z-1172-2012 | 61224 | No | EKS | Dental hand instrument. | Dental | Dental | 510(K) Exempt | K3 Engine File G-Pack, |  |  | August 01, 2008 | 2008 | Ormco Corporation | A recall was initiated because SybronEndo has confirmed that the K3 Engine File G-Pack has label packaging and file mis-marking error. | Process control | A recall communication was initiated on 07/31/2008 with SybronEndo forwarding an Urgent Medical Device Recall letter (via USPS 1st class mail) to all their customers who purchased the K3 Engine File G-Pack (Part No. 830-9925). The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to contact SybronEndo Customer Care at (800) 346-3636 directly to handle the arrangements of a quick return and replacement. Customers were instructed to complete the Return Form and return any affected product in their inventory. |  | 199 | 199 units | Worldwide Distribution -- US, including the territory of St. Thomas, and countries of Australia, Canada, New Zealand, China, Singapore, India, and Chile. | Terminated | March 08, 2012 | 1315 | 199 | Not_Computer
Z-1177-2012 | 61262 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Pelton &amp; Crane Spirit 3000 dental operative unit with the TLC SAFEdrive option, 1727 Fuehauf Dr. Charlotte, NC 28273 |  |  | February 20, 2012 | 2012 | Pelton &amp; Crane Company | Modification of the Pelton &amp; Crane Spirit 3000 unit disabled the TLC SAFEdrive feature rendering it ineffective in detecting whether the handpiece is potentially damaged or heating up. | Employee error | Pelton &amp; Crane sent an Urgent Medical Device Recall letter dated February 20, 2012 to all affected customers. The letter identifies the affected product, problem and actions to be taken. The letter states that a Pelton &amp; Crane representative will be contacting the customers to schedule a service appointment with the end user of the device so that the affected TLC SAFEdrive circuit board can be replaced and returned to Pelton &amp; Crane. Customers were instructed to complete the attached Recall Acknowledgement/Return Form and fax it to Pelton &amp; Crane, attention Frank Ray, at (704) 587-7250. The letter also requested customers to mail the original Recall Acknowledgement/Return Form in the enclosed self addressed stamped envelope. |  | 2 | 2 units | (USA) Nationwide Distribution | Terminated | March 28, 2012 | 37 | 2 | Not_Computer
Z-1161-2012 | 61198 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse Treatment Planning System, versions 7.3, 8.0, 8.1, 8.2, 8.6, 8.9 and 10.0; |  |  | January 30, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has been identified an anomaly with the Eclipse Treatment Planning System where the calculated dose distribution is not updated or invalidated when a recalculation or reload is performed while the compensator or fluence editor is open. | Device Design | Varian sent an URGENT MEDICAL DEVICE CORRECTION - URGENT FIELD SAFETY NOTICE letter dated January 30, 2012 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the affected product, description of problem, provide recommended user actions, and special instructions for customers outside the USA and Canada. The letter stated that this document will be posted on the customer support website: http://www.MyVarian.com. Customers were instructed to complete the attached Proof of Notification or Receipt Verification Card and return to Varian Medical Systems. For further clarification and questions contact your Varian Customer Support District or Regional Manager. |  | 25604 | 12,077 units ***AMENDED TO 13,527 units*** | Worldwide Distribution - (USA) Nationwide. | Terminated | August 15, 2014 | 928 | 25604 | Not_Computer
Z-1186-2012 | 60321 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH (iPTH) Assay -assay (lot 138) |  |  | October 03, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc | Poor sensitivity and imprecision on ADVIA CentaurÂż systems; Serum samples low bias; EDTA plasma samples low bias | Process control | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated October 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructed customer to immediately discontinue the use of ADVIA Centaur Intact PTH (iPTH) assay kits that contain intact PTH (iPTH) reagent lot 138. It also recommended that for serum samples analyzed on the ADVIA Centaur CP system, the content of the Urgent Field Safety Notice be discussed with their laboratory director to determine the need to review previous test results, conduct patient follow up, and repeat testing for iPTH results that were less than the upper limit of the reference range, and had a companion calcium result that was elevated or at the upper limit of normal. Additionally, a Completion Notification form was attached for customers to complete and return. Customers should contact their local technical support provider for questions regarding this recall. |  | 500 | 100 Test =5768; 500 Test =1778 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Argentina, Chile, Columbia, Egypt, Finland, France, Hong Kong, India, Israel, Japan, Malaysia, Mexico, Netherlands, New Zealand, Poland, South Korea, Saudi Arabia, Singapore, South Africa, Spain, Taiwan, Thailand, Turkey, Uruguay, and Venezuela. | Terminated | March 08, 2012 | 1252 | 500 | Not_Computer
Z-1179-2012 | 61114 | No | LXB | Hearing Aid. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | SoundBite Hearing System; Microphone Tube Size 0. For use in patients with moderately severe, severe or profound sensorineural hearing loss in one ear and normal hearing in the other or patients with conductive hearing loss where the pure tone average bond-conduction hearing threshold is better than or equal to 25 dB HL. Sonitus Medical, San Mateo, CA. |  |  | February 02, 2012 | 2012 | Sonitus Medical Inc. | Sonitas received reports of stress cracking on the SoundBite Microphone Tube. | Nonconforming Material/Component | Letters were sent to all direct SoundBite customers on February 9, 2012 via certified mail. Customers were asked to remove the old microphone tubes and replace with the new enclosed microphone tubes. Return mail instructions were provided for the old microphone tubes. Customer questions were directed to 650-251-4185. |  | 676 | 676 microphones, all sizes. | Class III Recall - Nationwide Distribution, including the states of CA, AZ, PA, MI, UT, NC, NY, WI, TX, FL and DC. | Terminated | March 23, 2012 | 50 | 676 | Not_Computer
Z-1184-2012 | 61237 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Siemens mMLC (ModuLeaf) safety bar and trolley label. |  |  | February 07, 2012 | 2012 | Siemens Medical Solutions USA, Inc | This correction is being performed to update the safety knob of the Siemens brand mMLC (ModuLeaf) to conform to tie IEC 60601-1 safety standard. | Nonconforming Material/Component | The firm, SIEMENS, issued a "Customer Information" letter to its customers. The letter described the product, problem and actions to be taken. Siemens developed a new steel knob that complies with the latest standards; Siemens checked all ModuLeaf mMLC trolleys and if necessary refitted with the safety label and also arranged for an inspection and modification of the customers system- a Release of the safety update UI TH 026/11/S to distribute and hardware update commenced on February 7, 2007 by certified mail.If you have any questions, contact the Regulatory Specialist II at 925-602-8175. |  | 46 | 46 active devices | Worldwide distribution: USA (nationwide) including states of: CO, IA, KY, NJ, OH and TX; and countries of: Algeria, Austria, Belgium, Brazil, China, Egypt, France, Germany, Hungary, India, Italy, Japan, Kuwait, Malaysia, Mexico, P.R., Russian Fed., Saudi Arabia, South Africa, Spain, Turkey, and United Kingdom. | Terminated | August 14, 2012 | 189 | 46 | Not_Computer
Z-1185-2012 | 61238 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON G5 ventilators with software version V2.0x. |  |  | February 06, 2012 | 2012 | Hamilton Medical Inc | Hamilton-G5 ventilators with software version V2.0x may experience an unintentional change of the displayed ventilation mode under certain conditions. | Software design | Hamilton Medical sent a Medical Device Safety Alert and Correction Action letter dated February 17, 2011 to all affected customers via a traceable method. The letter identified the affected product, problem and immediate actions required by operators, distributor and manufacturer. The letter states that a Hamilton Medical Account manager will contact customers to schedule a convenient time for the Technical Support Department to install the software update. Questions or concerns may be directed to Technical Support Manager at 800-426-6331, Ext 125. |  | 456 | 456 devices in US. | (USA) Nationwide Distribution including the states of: CA, CO, FL, GA, MA, MI, MO, NC, NV, NY, PA, TN, TX, VA and WV. | Terminated | October 23, 2013 | 625 | 456 | Computer
Z-1194-2012 | 61182 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON Rheumatoid Factor (RF) Reagent, Part Number: 475136 |  |  | June 20, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman has confirmed that the all in date lots of Synchron Rheumatoid Factor (RF) Reagent might generate false positive results or calibration failures. | Nonconforming Material/Component | Beckman Coulter sent an Urgent Field Safety Notice/Product Correction letter dated June 20, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that Beckman Coulter is developing a new RF assay. In the meantime due to current issues related to shelf life and low end sensitivity, Beckman has decided to take the product off the market effective immediately.Beckman will keep the customers updated on their progress and advise when the new assay becomes available.Customers were instructed to complete and return the enclosed Response Form with in 10 days. Customers with questions regarding this notice can contact their local Beckman Sales representative or contact Customer Support Center at www.BeckmanCoulter.com/CustomerSupport/Support or call (800) 854-3633 in the US and Canada. |  | 22081 | 22,081 units total (12,462 units in US) | Worldwide Distribution - USA (nationwide) and worldwide to foreign countries:Australia, Belgium, China, Costa Rica, Czech Republic, Finland, France,French Polynesia, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy,Japan, Kuwait, Lebanon, Lithuania, Macao, Mexico, Netherlands, NewZealand, Poland, Russian Federation, Serbia, Slovakia, South Africa, Spain,Switzerland, Taiwan, Turkey, United Arab Emirates, and the United Kingdom. | Terminated | September 02, 2014 | 1170 | 22081 | Not_Computer
Z-1193-2012 | 61092 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | PointCare NOW CD4NOW Gold Reagent Pack, or Gold Pack, |  |  | November 02, 2011 | 2011 | Pointcare Technologies, Inc. | Stability not assured at labeled temperature 2-42ÂżC | Storage | The firm, PointCare, issued a "Customer Service Advisory (CSA-00016)" dated November 2, 2011 to its customers to inform them of the new storage requirements of 2-30ÂżC as opposed to the previously claimed 2-42ÂżC. The letter describes the product, problem and action to be taken. The customers were instructed to correct the labeling on any Gold Reagent Packs and Gold Reagent and ensure that the Gold Reagent Packs and Gold Reagent they currently have in stock is stored within the 2-30ÂżC storage range. Future lots will be labeled with the updated storage temperatures.A member of our customer support team will provide additional information relative to this advisory.Should you have any questions or require further assistance, you may email us at customersupport@pointcare.net or call our offices directly at 1-508-281-6925. |  | 463 | 463 packs | International only: including countries of : Abuja, Haiti, Indonesia, Natal, Nigeria, Senegal, South Africa, Uganda, and United Kingdom. | Terminated | April 26, 2012 | 176 | 463 | Not_Computer
Z-1191-2012 | 61021 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga UIS 2009 is Microsoft Amalga Unified Intelligence System (UIS), Calculator/data processing module, for clinical use. |  |  | January 11, 2012 | 2012 | Microsoft Corporation Health Solutions Group | The problem occurs when a user filters the data in Amalga datagrid to exclude specific values. If the field specified in the filter criteria contains NULL values, the filtering will exclude those records as well. | Software design | Microsoft Corporation Health Solutions Group sent a "Urgent Field Safety Notice" email dated January 19, 2012, to all affected customers. Notification of customers in China is still pending due to the Chinese New Year Holiday. Customers were instructed to :1. If the user notices varying row counts or inconsistent values in a given column and they are using the exclude filter, they may have to look for alternative ways to filter the data that do not use the above mentioned filter conditions.2. To avoid the indicated issue in your implementation there are two steps to be taken:a. Ensure that all parsers are inserting empty strings () rather than NULLwhen ingesting data into the system andb. Identify all Baseview definitions (i.e. the sql view that a baseview isaccessing) that include an outer join and update those definitions toinclude the ISNULL replacement function [e.g. ISNULL( column_name, )]on any columns which may return NULL due to a non-matching row in theouter join tabAs of January 23, 2012 The firm's members of Services Team have spoken by phone with representative of every customer.Customers should call 425-882-8080 for any questions about this recall. |  | 28 | 28 units | Worldwide Distribution -- USA (nationwide) including the states of AR, CA, KY, HI, FL, MA, MD, MN, NC, NJ, NY, WA, WI and countries China, Switzerland, and United Kingdom. | Terminated | December 12, 2012 | 336 | 28 | Not_Computer
Z-1192-2012 | 61279 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | Visionsense Stereoscopic Vision System, Brand name VSII, Model/Part Numbers 172-0030 and 172-0040 --- Classification Name: Neurological Endoscope --- NOTE: Model Numbers 172-0032, 172-0044, 172-0070, 172-0072, 172-0080 were also included in the firm's recall action but these never entered US commerce. These were distributed directly from Israel to the foreign consignees. --- The VSII is intended for viewing internal surgical sites during general surgical procedures and for use in visualization of ventricles and structures within the brain during neurological surgical procedures. These devices do not carry an expiration date. However, they are automatically taken out of service and replaced after 40 uses. |  |  | March 21, 2011 | 2011 | Visionsense, Ltd. | A complaint was received and a unit returned that exhibited the metal and glass shaft window becoming dislodged during scope cleaning. Additionally, during in-process visual inspection of the metal and glass shaft window sub-assembly from a specific batch that was received from the supplier it was noted that the window at the tip of the scope was not seated properly. It was determined that the wi | Nonconforming Material/Component | Visionsense Ltd. sent URGENT DEVICE RECALL and URGENT DEVICE RECALL/FIELD SAFETY NOTICE letter dated March 21, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer and included a return "fax" form in the recall communication package. The letters were sent via UPS express service. All direct consignees were contacted, and follow up phone calls conducted if needed. The recall required the return of all affected units for inspection and/or repair. Visionsense provided a return authorization number and packing instructions to the consignees. |  | 21 | Domestic - 17 units; Foreign - 4 units | Worldwide Distribution - USA including Arizona, Georgia, Pennsylvania, California, and the country of Turkey | Terminated | March 08, 2012 | 353 | 21 | Not_Computer
Z-1190-2012 | 61173 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Right Long Tactile Probe, Part #962011S, Medtronic Navigation, Louisville, CO 80027. |  |  | February 06, 2012 | 2012 | Medtronic Navigation, Inc. | When subjected to bending, twisting, and malleting, the probe tips break off when used in the sacrum for which they are not indicated for use. These probes are approved for use in the lumbar vertebra. | Device Design | An Urgent - Product Safety Notice customer letter was sent via FED EX on 2/6/12 and 2/6/12 to direct consignees in the US identifying the affected product and the issue with it. The letter states that Medtronic Navigation is removing the product from clinical use and will provide a sales credit to customers. In the letters to Medtronic field personnel, it states that local sales and service personnel will be identified to remove the affected product from customer sites and return them for scrapping. Customers are to contact a sales representative if they have any questions regarding the letter. |  | 5 | 5 units | Worldwide Distribution -- US and Republic of South Korea. | Terminated | March 26, 2012 | 49 | 11 | Not_Computer
Z-1187-2012 | 61131 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra InterFix CT Adapter (CTA) |  |  | January 23, 2012 | 2012 | Integra LifeSciences Corp | Material used in the superior region of the InterFix CT Adapter (CTA) has a different density than that of the Treatment Table Couch Adapter (TTA). Resulting dosimetry can be affected if posterior fields are used. Integra estimates that a target dose decrease of 1.2% would result from a typical 360 degree TomoTherapy arc. | Device Design | Integra sent a Product Technical Bulletin dated January 23, 2012, to all affected consignees via FedEx, email or telephone conference calls. The revised Operator Manual was provided with a Technical Bulletin to all consignees that have been shipped this system. Consignees were asked to distribute the information in the bulletin to members of the Radiotherapy Planning Team within their institution. Customers were also asked to sign and return the enclosed Product Technical Bulletin Acknowledgement Form. For questions regarding this recall call 888-772-7378. |  | 33 | 33 units | Worldwide Distribution - USA including DC, IN, FL, HI, LA, NY, MD, MI, OH,OR, WI, WY,Foreign: Great Britain, Italy, and France | Terminated | October 01, 2012 | 252 | 33 | Not_Computer
Z-1195-2012 | 61145 | No | CIN | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCESS Ostase |  |  | June 22, 2011 | 2011 | Beckman Coulter Inc. | The Access Ostase reagent lot 021281 shows instability resulting in falsely high patient results. | Process design | A recall communication was initiated on 6/15/2011 with Beckman forwarding an "Urgent Product Correction" (PCA) letter to all their customers who purchased the Access Immunoassay Systems Ostase Reagent Kits. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to discontinue use of the Access Ostase reagent lot identified on the PCA letter and discard all remaining inventory. As stated in the Instructions for Use for Access Ostase, results should be interpreted in light of the total clinical presentation of the patient.All results generated with reagent lot number 021281 where patient results are higher than expected should be evaluated. The affected lot has been removed from Beckman's inventory, and alternate lots will be provided for all current and future replenishment orders or requests.Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions can contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call(800) 854-3633 in the US and Canada. |  | 318 | 318 units total (44 units in US) | Worldwide Distribution -- US, including states of FL, IN, VA, CA, and KY and countries of China, Japan, and Panama. | Terminated | March 09, 2012 | 261 | 318 | Not_Computer
Z-1197-2012 | 61267 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva FD, Software R5.1.x |  |  | February 24, 2012 | 2012 | Philips Medical Systems Nederlands | Displayed skin dose rate is calculated using an incorrect point of reference, resulting in too low displayed values. | Software design | Philips Healthcare issued Urgent Field Safety Notices dated February 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. No action is necessary by the customer/user. A mandatory Field Safety Correction (FCO) Action will be issued to solve the problem. The FCO will consists of a Software update. For further support customers were instructed to contact their Phillips representative.For questions regarding this recall call 978-687-1501. |  | 20 | 45 US; 20 Canada;142 International=207 units | Worldwide Distribution - USA (nationwide) and the following countries:Canada, AustraliaAustriaBelgiumBahrainCzech RepublicDenmarkEstoniaFranceFinlandGermanyHong KongIsraelItalyLatviaLuxemburgMalaysiaMexicoNetherlandsNew ZealandNorwayRussian FedSAUDI ARABIASerbiaSlovakiaSolveniaSouth AfricaSpainSwitzerlandSwedenSyriaThailandTurkeyUnited Kingdom | Terminated | March 26, 2013 | 396 | 59 | Computer
Z-1198-2012 | 60892 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Flat Proximal Pressurizer. |  |  | January 13, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | The firm identified two complaints have been reported indicating that the product inside the package did not match with the label information. The product was labeled as Flat Proximal Pressurizer P/N 0206-556-000 while the product inside the package was Femoral Pressurizer P/N 0206-566-000. | Mixed-up of materials/components | Stryker Instruments sent an "Urgent Medical Device Recall notification" dated January 13, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to : Immediately check all stock areas or operating room storage and quarantine any affected product found. Mark product as RECALLED PRODUCT . Please indicate on the BRF the quantity of affected Pressurizers you are returning and fill in the form completely. Sign the BRF (even if you do not have any affected product).Note: Your signature on the BRF indicates that you received and understand this notification and are returning all affected product from your facility.Fax the Business Reply Form to Stryker Instruments Regulatory Department,866-521-2762.Upon receipt of the BRF, Stryker will contact you to coordinate the return of allaffected product that you have on hand. Send back all affected product using the pre-paid shipper provided to you by Stryker.Upon receipt of the product, a replacement will be issued to your account |  | 10 | 10 boxes (60 units) | Worldwide Distribution -- Nationwide, including the states of AL and OK and the country of Canada. | Terminated | July 24, 2013 | 558 | 10 | Not_Computer
Z-1199-2012 | 61179 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Nucletron Adapter and Transfer Guide Tube, part number AL 13005000, a component used with the VariSource" IX HDR Afterloader. |  |  | January 05, 2012 | 2012 | Varian Medical Systems, Inc. | Components of medical device used in brachytherapy may be defective and cause a shifted dose distribution to patients. | Device Design | Varian Medical Systems sent an Urgent Medical Device Correction - Urgent Field Safety Notice letter dated January 5, 2012 to all affected consignees. The letter identified the affected products, description of problem and recommended user actions to be taken. Consignees were requested to cease use of the affected products and contact Varian BrachyTherapy Helpdesk immediately to arrange return of their affected parts. Consignees were requested to return the enclosed proof of notification form. Consignees were instructed to advise the appropriate personnel working in their radiotherapy department of the content of this letter. For future reference, this document will be posted to the Varian Customer Support website: http://www.MyVarian.com. For further clarification contact your local Varian District or Regional Manager. |  | 122 | 122 total units | Worldwide Distribution - USA (nationwide) Including the states of CA, FL, IN, KY, MI, MN, MO, OH, TX, VA and the countries of Canada, Japan and Sweden. | Terminated | November 21, 2012 | 321 | 122 | Not_Computer
Z-1201-2012 | 61160 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX SP Power Wheelchair, Model #'s: TDXSP-CG, TDXSP-CG-GT, TDXSP-MCG, TDXSP-MCG-GT, Distributed By: |  |  | February 10, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the product because use of this actuator could result in reduced cross sectional current carrying area due to wear/chafing of the wires. | Nonconforming Material/Component | The firm, Invacare, sent an "URGENT: RECALL INFORMATION" letter dated February 10, 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. Invacare has developed retrofit kits intended to address this issue. The consignees were instructed to contact the customers that have potentially affected product and perform the necessary updates. The consignees were also instructed to complete and return the Provider Response Form to order the retrofit kit.If you have any questions, please contact Customer Service at 1-800-333-6900. |  | 102 | 102 wheelechairs | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CT, FL, GA, IA, IL, IN, KY, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, RI, SD, TX, VA, WA, and WI; and countries of: Australia, Canada and New Zeland. | Terminated | August 25, 2014 | 927 | 102 | Not_Computer
Z-1202-2012 | 61156 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S. Surgical System, da Vinci Si and da Vinci Si-e Surgical Systems. |  |  | February 06, 2012 | 2012 | Intuitive Surgical, Inc. | The holding brake, a component of the da Vinci S Surgical System, may allow passive uncontrolled motion due to gravity during specific power-off conditions. | Device Design | Intuitive Surgical sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated February 3, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provided new instructions for use. Additionally, a Device Correction Response Form was enclosed for customers to complete and return via fax to 408-523-1390. Contact Intuitive Surgical Customer Service at 1-800-876-1310 for questions regarding this notice. |  | 1966 | 1966 systems are in commerce at hospital locations. | Worldwide Distribution. | Terminated | May 09, 2012 | 93 | 1966 | Not_Computer
Z-1204-2012 | 61296 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Reprocessed Mitek Arthroscopic Wand is considered a soft tissue ablators. Soft tissue ablators are radiofrequency (RF) surgical tools designed for removal and dissection of tissue in arthroscopic surgeries. A generator or controller serves as the power unit. Soft tissue ablators are available in a wide range of sizes and angles; some models include thermal or suction functions. |  |  | June 01, 2009 | 2009 | Stryker Sustainability Solutions, Inc. | On 06/09/2009, Stryker Sustain ability Solutions, formerly Ascent Healthcare Solutions, recalled lots 674201 &amp; 674290 of their Reprocessed Mitek Arthroscopic Wand due to the lots not being tested in accordance with the functional test requirement during the time frame of 04/14/09 to 05/25/09. | Process control | Ascent Healthcare Solutions notified their customers by presenting a letter dated June 1, 2009 from an Ascent Representative. The representative was to recover any products that may have been affected. Questions were directed to 1-888-888-3433. |  | 48 | 48 | Nationwide Distribution (USA), including the states of: TX, NY, MN, CA, and CT. | Terminated | May 07, 2012 | 1071 | 48 | Not_Computer
Z-1203-2012 | 61275 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Coupler System II |  |  | February 27, 2012 | 2012 | Mizuho Orthopedic Systems Inc | Mizuho determined that current labeling of The Coupler II system required revisions after investigation of a complaint. | Labeling design | Mizuhosi OSI sent an Urgent Medical Device Recall Communication letter dated February 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Mizuho OSI provided their customers of the affected product an updated Owner's manual. Customers were instructed to ensure they remove their existing Owner's Manual from use and repalce with the current Owner's Manual. If customers were unsure of the device's use, they should refer to the included Owner's Manual and ensure all users are properly trained, able to assess the device's status prior to and after each use. If customers were unsure of the status of their device, or of its use, they should call 1-800-777-4674 or outside of the USA 00+1=510-476-8199.Customers were asked to respond that the recall has been received, read, understood and communicated to those affected by this activity. A response should be sent via e-mail to couplerII@mizuhois.com.For any questions regarding this recall call 1-800-777-4674 or outside of the USA 00+1=510-476-8199.was sent out on February 27, 2012 to all affected sites. This was after the original was sent to CDRH for review on January 31, 2012, but approval of revisions was not received until 2/27/12. |  | 1077 | 1077 devices | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Canada, France, Germany, Hong Kong, Israel, Japan, South Korea, Lebanon, Malaysia, Netherlands, New Zealand, South Arabia, Singapore, Germany, South Africa, Spain, Sweden, UA. E., United Kingdom, and Australia. | Terminated | September 25, 2012 | 211 | 1077 | Not_Computer
Z-1207-2012 | 61103 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Wet PVP-I Small Procedure Preoperative Skin Prep Tray containing 1 - 59 mL bottle Povidone Iodine Cleansing Solution (Scrub) USP 7.5% and 1 - 59 mL bottle Povidone Iodine Topical Solution (Paint) USP 10% with Powder-Free Esteem Gloves; 20 trays per case; a sterile, latex-free, single use only tray, solution not exposed to sterilant; U.S. components packaged in Mexico except as indicated: Gloves and Cotton-Tipped Applicators made in China, Sponge Stick made in Mexico; CareFusion, McGaw Park, IL 60085 USA; tray catalog number 4420 |  |  | February 07, 2012 | 2012 | Carefusion 2200 Inc | The povidone-iodine (PVP-I) solution bottles in the trays may have loose caps, allowing leakage and also elevated levels of ethylene oxide (EO) sterilant and/or by-products within the PVP-I solution, which may represent a patient safety hazard. | Packaging process control | CareFusion sent e-mails and follow-up Urgent Recall Notices dated February 7, 2012 to all of the affected customers on the same date, requesting that the customers follow the steps outlined below as mitigation until product can be replaced.1. Please immediately destroy/discard or return any affected product you may have in-stock at your facility. Do NOT further distribute these products.2. Forward this RECALL NOTICE to any customers to whom these products were sold notifying them of the potential risk (reference attachment for affected production lot numbers). End users must be advised to discard or return any CareFusion Skin Prep Tray with the noted product/lot number; distributors should manage this return process per their internal procedures. End users MUST NOT use these products. 3. Complete and return the attached Acknowledgement Form indicating that they have received this RECALL NOTICE and that they will communicate the Recall Notice to all customers who have purchased CareFusion Preoperative Skin Prep Tray catalog numbers 4420, 4468, and 4469A. 4. Contact CareFusion Customer Service at (800) 523-0502 ext. 53576 for details on return and/or replacement of stock. |  | 38440 | 38,440 trays | Nationwide Distribution including Arizona, California, Colorado, Connecticut, Florida, Illinois, Missouri, Nebraska, Nevada, New York, Ohio, Puerto Rico, Rhode Island, South Carolina, South Dakota, Texas, Utah and Virginia. | Terminated | June 05, 2015 | 1214 | 1644150 | Not_Computer
Z-1205-2012 | 61270 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Brand Name: TempSpan CMT Temporary Cement, a temporary cement. |  |  | March 03, 2011 | 2011 | Kerr Corporation | The recall was initiated because Pentron Clinical has confirmed that the TempSpan Temporary Cement may have black plastic material on the surface of the paste, and it may set faster than specified. | Nonconforming Material/Component | Pentron Clinical sent an Urgent Medical Device Recall letter dated April 20, 2011, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to complete and fax back the enclosed recall form to 1-877-677-8844 in order to confirm receipt of the customer notification. Customers were instructed to contact Pentron Clinical Customer Service at (800) 551-0283, option 1, directly to handle the arrangements of a quick return and replacement.For questions regarding this recall call 714-516-7400. |  | 627 | 627 units total (601 units in the US) | Worldwide Distribution - USA (nationwide) and the countries of Canada, Philippines, Italy, Netherlands, and the UK. | Terminated | March 14, 2012 | 377 | 627 | Not_Computer
Z-1206-2012 | 61273 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, MedfusionÂż 3500 Syringe Pumps, Software version 5.0.0, Rx Only. |  |  | January 02, 2012 | 2012 | Smiths Medical ASD, Inc. | Pumps rendered inoperable at power up (displaying a blank screen and alarm). | Vendor change control | Smiths Medical ASD, Inc. contacted two consignees by telephone in November 2010 and January 2011. Smiths Medical worked with one of the two customers to correct 540 of the 595 devices (91%) while on-site in January 2011. The remaining devices were corrected by hospital personnel.The other customer was also made aware of the issue, root cause and corrective action. This customer reported to Smiths Medical that they understood the issue, but did not want Smiths Medical to perform the correction on their pumps at the time (January 2011). They preferred to send their pumps to Smiths Medical for repair, as and when needed. As of December 2011, The hospital has returned 67 devices to Smiths Medical.For questions regarding this recall call 651-628-7477. |  | 1459 | 595 (WI )+ 864 (DC) = 1459 | Nationwide Distribution including WI and DC | Terminated | July 25, 2013 | 570 | 1459 | Not_Computer
Z-1213-2012 | 61090 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System. |  |  | January 31, 2012 | 2012 | Idev Technologies, Inc. | The firm is taking this action because the device was one of eight that may have been mislabeled with the incorrect size (length). | Labeling mix-ups | IDEV Technologies sent an "URGENT PRODUCT RECALL" letter dated February 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to sign the letter and return a copy to the firm via fax to 281-525-2001 to indicate that they have received the notification. Contact the firm at 281-525-2052 for questions regarding the notice. |  | 7 | Seven devices | Nationwide Distribution-USA (nationwide) including the states of AL, IL, NJ, NV, and PA. | Terminated | September 09, 2013 | 587 | 7 | Not_Computer
Z-1215-2012 | 61093 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA EXcel Sterile Tips Rx Only |  |  | October 14, 2011 | 2011 | Integra LifeSciences Corp | A small number of non-conforming CUSA Excel 23 KHz Standard Tip and flue (item Number C4601S) were released for distribution that did not meet Integra's in-process requirements for release. It was discovered during in-process inspection that the max power reading exceeded the specification of 80 watts, causing the stroke range of the tip to exceed it required range by up a value up to 0.002 inches | Release of Material/Component prior to receiving test results | Integra sent an Urgent Product Recall Notification letter dated October 14, 2011 to all affected customers via US Mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory, discontinue use, quarantine and return any available affected lot numbers. Customers were instructed to distribute this communication to any customers to whom they have shipped the affected product. Customers were advise to complete the enclosed Recall Acknowledgement and Return Form and return it as soon as possible to Integra Customer Service. For questions contact Integra Customer Service at 609-936-2659. |  | 24 | 24 | Worldwide Distribution | Terminated | October 19, 2012 | 371 | 24 | Not_Computer
Z-1211-2012 | 61348 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | MPI Rapid Response HUT Table, Model #RRHUT 1017. The firm name on the label is Medical Positioning Inc., Kansas City, MO. |  |  | March 05, 2012 | 2012 | Medical Positioning Inc | The warning label regarding electrical shock was not placed on the device warning to have the power switch in the "OFF" position prior to connecting the unit to electric power. | Error in labeling | Medical Positioning Inc. sent a "RAPID RESPONSE HUT TABLE LABELING CORRECTION NOTIFICATION" letter dated March 5, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included an Owner's Manual insert page which clearly states the warning and how to avoid the electric shock hazard and a warning tag to attach to the product's power cord near the plug end using the provided wire tie. Contact the firm at 816-595-2356 for questions regarding this notice. |  | 5 | 5 tables | Worldwide Distribution-USA (nationwide) including the states of IN, KS, WI, NY, IL, NC, OH, and MS and the country of Thailand. | Terminated | July 11, 2013 | 493 | 9 | Not_Computer
Z-1216-2012 | 61204 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips EasyDiagnost Eleva DRF, Fluoroscopic x-ray system |  |  | February 14, 2012 | 2012 | Philips Healthcare Inc. | Misinterpretation of Mirror Icon-Mirror icon on the bottom of each image could be misinterpreted as aRIGHT lead marker for side indication if the image is mirrored within a PACS system | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 8, 2012 to all affected customers. The letter identified the affected product, problem description, hazard involved, actions to be taken by customer/user and actions planned by Philips. The letter contains important information for continue safe and proper use of the equipment. Philips will contact customers to schedule the installation of the updated software. For further information or support contact your local Philips representative at the Technical Support Line 866-767-2822. |  | 60 | 60 units | Worldwide Distribution - USA Nationwide | Open, Classified |  |  | 60 | Computer
Z-1218-2012 | 60934 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Elecsys 2010 Analyzer Software version 06-07 |  |  | January 31, 2012 | 2012 | Roche Diagnostics Operations, Inc. | Elecsys 2010 customers have reported getting no value and flag 48 results below expected value range after performing automatic dilution for a sample that previously recovered above the measuring range. Incorrect but believable results may occur for the CEA, Rubella IgG or SHBG assays. The other affected assays will generate either no value or a negative result, both of which are not bel | Software design | Roche sent an Urgent Medical Device Correction letter dated January 31, 2012, via UPS Ground (receipt required) to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: use one of the workarounds outlined in this Urgent Medical Device Correction (UMDC) if you are running an automatic dilution on an Elecsys 2010 analyzer with software version 06-07; complete the attached faxback form (4600-00-0112) and fax it to 1-888-912-8460; and file this UMDC for future reference.For questions regarding this recall call 1-800-428-2336. |  | 607 | Software version 06-07 was installed on Elecsys 2010 analyzers at 13 accounts. | Worldwide Distribution - USA including NY, GA, DE, SD, WI, ME, OR, IA, AND CA and the countries of Austria, Cyprus, Czech Republic, Estonia, France, Germany,Greece, Hungary, Iceland, Italy, Latvia, Norway, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, and Switzerland | Terminated | September 06, 2012 | 219 | 607 | Computer
Z-1220-2012 | 61007 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | DuoDiagnost, model 707015, Universal RF system |  |  | December 26, 2011 | 2011 | Philips Healthcare Inc. | The device is currently not compliant with FDA's amended design criteria for this system, the product is not commercially available in the US. However, it was inadvertently shipped to Puerto Rico due to a logistical error. | No Marketing Application | Philips sent an Urgent - Field Safety Notice dated December 26, 2011 to the two affected customers. The letter identified the affected product, problem description, hazard involved, actions to be taken by customer/user and actions planned by Philips. The letter states that customers may continue to use the DuoDiagnost system until the new systems are replaced by Philips. A Philips representative will contact customers to arrange for the replacement of current system. For further information or support contact your local Philips representative. |  | 2 | 2 units | USA Nationwide Distributed in Puerto Rico | Terminated | November 01, 2013 | 676 | 2 | Not_Computer
Z-1217-2012 | 61193 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FastPlan Treatment Planning System, Model HZ3, Varian Medical Systems, Palo Alto, CA. Varian FastPlan Treatment Planning System is a radiation treatment planning system composed of a high resolution UNIX graphics computer, a printer, a digitizing tablet, and radiation treatment planning application software. |  |  | January 30, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the CT Localization module of the FastPlan treatment planning system. If the module reports a significant difference in tilt or spin on a slice and the use elects to "Discard" that slice, an incorrect calculation of the isocenter location may result. | Device Design | "Urgent Medical Device Correction/Urgent Field Safety Notification" letters were sent on January 30, 2012 by certified mail to consignees in the US and by Federal Express for Canada. International sites will be notified by Varian International Customer Support Managers, who will translate letters if necessary. The notice provided a description of the issue, the actions that can be taken to avoid or mitigate the issue, and the steps Varian is taking to address the issue. Customer service number for US and Canada is: (888) 827-4265. |  | 201 | 201 | Worldwide Distribution - Nationwide (USA) and the countries of Canada, China, Germany, Ireland, Spain, Singapore, Georgia, Hong Kong, India, Macau, Malaysia, Russia, Sri Lanka, Taiwan, Japan, Chile, Columbia, Venezuela and Australia. | Terminated | October 21, 2013 | 630 | 201 | Not_Computer
Z-1219-2012 | 61250 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Brilliance Workspace Portal is a software component of the Computed Tomography X-Ray System. |  |  | February 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Incorrect Standardized Uptake Value ("SUV") values are displayed in two instances:1. After loading PET data (with a Slope value different than 1) into the Quick Review (QR) application from a remote device; and2. During the loading the above mentioned type of PET data into the CT viewer application from a remote device. | Software design | The Philips Healthcare, sent an "URGENT-Medical Device Correction" Field Safety Notice dated February 08, 2012 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to review the information with all members of their staff and to retain a copy with the equipment Instruction for Use. The letter notifies the customers that a Field Service Engineer will be contacting them to schedule a visit to implement the corrective action by installing the software update version 2.6.1.If you need further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 822 | 822 | Worldwide distributions: USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, VA, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, VT, WI, and WV; and countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech, Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Poland, Portugal, Reunion, Russia, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uzbekistan, and Venezuela. | Terminated | February 22, 2013 | 365 | 822 | Computer
Z-1221-2012 | 59992 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Crystal Metal Disposable Blades size Mac 4 (part number 50674, 50674/F, 50674/BMI) and Mac 3 (part numbers 50675, 50675/F, 50675/BMI) |  |  | September 21, 2011 | 2011 | Penlon, Ltd. | Disposable laryngoscope blades incorrectly labeled with wrong size. | Error in labeling | InterMed Penlon sent and Urgent Field Safety Notice letter dated September 21, 2011 to affected customers. The letter identified the affected product, problem and advice on actions to be taken by user. Customers were instructed to complete the Reply Form and send to the address provided. |  | 24940 | 24,940 cartons | Worldwide Distribution - USA (nationwide) in the countries of Belgium, France, Greece, Ireland, Italy, Lithuania, Netherlands, Norway, Portugal, Switzerland, Canada, China, Hong Kong and Malaysia. | Terminated | June 19, 2012 | 272 | 24940 | Not_Computer
Z-0910-04 | 35768 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Quick-set Plus Infusion Sets. Models MMT-359S6; MMT-359S9; MMT359L6; &amp; MMT-359L9. (differences in models is variation in length, 6 versus 9mm) |  |  | March 02, 2004 | 2004 | Medtronic MiniMed | Possibility of interrupted insulin flow. This posting is for a recall that occurred in March 2004, also posted on the FDA website in 2004. | Component design/selection | Firm will ship all users (customers) one box at a time every three weeks for the next several months. A recall notification letter was issued to users and to healthcare professionals, and distributors. Letters were mailed from March 2-5, 2004. |  |  | Undetermined | Nationwide | Terminated | March 19, 2012 | 2939 |  | Not_Computer
Z-1124-2012 | 59944 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova X-ray Imaging Systems. |  |  | May 04, 2011 | 2011 | GE Healthcare, LLC | It was discovered that when the Diamentor M4-KDK is installed on an Innova system the dose display is disabled on the certified Innova monitor and enabled on the Diamentor M4-KDK which is not certified as required by FDA regulation (Subpart J). | Radiation Control for Health and Safety Act | GE Healthcare will alert affected Innova users with a Diamentor M4-KDK unit of this issue and the workaround to mitigate this issue. GE Healthcare will bring affected Innova units in the field into compliance by enabling the Innova certified display and providing affected customers with an Operators Manual Addendum and Service Manual Addendum which will address this topic.GE Healthcare will correct all affected systems by updating the system's Operator and Service Correction manuals and enabling the innova dose display on the Live and Reference monitors and the DLconsole. The Diamentor M4-KDK can still be used to display additional informative dose values. GE Healthcare will, without charge, remedy the defect or bring the product into compliance; the details of which will be included in a subsequent communication to you or through a GE field engineer site visit. GE Healthcare updated the Innova User Instructions, as part of the recall, to indicate to meet regulatory requirements based on the international safety standards IEC 60601-2-43:2000 and/or 21 CFR 1020.30 (k), use as reference the exposure dose displayed on the DL console, live or reference monitor. If you have any further questions or concerns please call 800-437-1171. |  | 19 | 19 | USA Nationwide Distribution | Terminated | March 29, 2012 | 330 | 19 | Not_Computer
Z-1226-2012 | 61183 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Integra Endoscopic Gastrocnemius Release (EGR) System. The device consists of one (1) unit which is kitted within an instrumentation tray. The Integra EGR System includes a range of instruments for endoscopic approach to the Gastrocnemius Aponeurosis Recession. The cutting instrument features a retracting blade for cutting the gastrocnemius aponeurosis. Associated instruments include an elevator, a cannula, and a cannula obturator, which are all consumable instruments provided sterile. The EGR System is indicated for the treatment of posterior heel cord contracture (Equinus) in those patients who fail to respond to conservative management. |  |  | January 06, 2012 | 2012 | Integra LifeSciences Corp. | The firm received several complaints from customers reporting that the blade of the EGR system may not retract back into the device following surgery. The firm attributes this problem to user error or improper surgical technique on the part of the physician. | Labeling mix-ups | The firm sent a 'A Medical Device Correction Notice' with an attached Acknowledgement and Return Form to all consignees on 1/6/2012 via FedEx and E-mail. The Correction Notice informs customers of the firm's receipt of several complaints regarding the Endoscopic Gastrocnemius Release (EGR) System and the problem with non-retraction of the blade back into the device during surgery. Although no adverse health consequences or patient injuries have been reported to date, the notice is being issued by the firm in order to reinforce the importance of following the additional clarification to the Surgical Technique Manual. According to the firm, this action should serve to reduce the possibility in which the blade may fail to retract into the device. Questions were directed to (609) 936-6832. |  | 850 | 850 units | Nationwide Distribution, including the states of: MI, CA, TX, VT, OK, OH, NC, MN, CT, CO, TN, NY, NJ, MO, FL, WA, AR, MT, VA, IL, PA, NE, GA, AZ, WI, OR, MD, LA, IA, AL, MA, ID, KY, UT, NV, IN and SC. | Terminated | January 23, 2014 | 748 | 850 | Not_Computer
Z-1225-2012 | 61138 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew BIOSURE HA Interference Screw, 7 mm x 25 mm |  |  | February 06, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | Mislabeled: The device was labeled as a standard thread screw (P/N 72201772) on pouch and box,however the device contained in the packaging is a reverse thread screw (P/N 72201773). | Packaging process control | Smith &amp; Nephew Endoscopy sent an "URGENT-PRODUCT RECALL 1st NOTIFICATION" dated February 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to inspect their inventory and quarantine any affected product. For instructions on returning product, contact the Smith &amp; Nephew Endoscopy Division Returns Group at 800-343-5717 ( 0ption 3) or email to endo.andreturns@smith-nephew.com. Additionally, customers were asked to complete the bottom portion of the letter and to fax it to 1-508-261-3620. Contact the Smith &amp; Nephew Endoscopy Division for questions regarding this notice at 1-508-261-3731. |  | 194 | 194 units | Worldwide Distribution-USA (nationwide) and the states of AK, AZ, CA, CO, CT, IA , MD, MI, MN, MO, MS, NE, OH, OK, PA, SC, and VA and the countries of Australia, Canada, Germany, South Korea, Spain, Switzerland, Thailand, and United Kingdom. | Terminated | February 29, 2016 | 1484 | 194 | Not_Computer
Z-1227-2012 | 61243 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Duet TRS Articulating 60- 3.5MM, Single Use Loading Unit |  |  | February 16, 2012 | 2012 | Covidien LP | Missing component results in the stapler not firing. | Process control | The firm, Covidien, sent a "Medical Device Recall" letter dated February 16, 2012 to customers via federal express. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of Lot# N1H003LX of Duet TRS Articulating Loading Unit and segregate affected product; return inventory to Field Returns Department, 195 McDermott Road, North Haven, CT 06473 and contact customer service at SDFeedback@Covidien.com or (800) 722-8772, option #1 for RGA # prior to returning; and complete and return the Recall Product Return Form via fax to (203) 822-6009 (all affected units must be returned with completed form through the Distributor). If you have any questions or concerns, please do not hesitate to contact your Covidien representative or Quality Assurance at 203-492-5232, or Customer Service at 800-722-877 option 1 customer service. |  | 1158 | 1158 units | Worldwide distribution: USA (nationwide) and country of: Venezuela. | Open, Classified |  |  | 1158 | Not_Computer
Z-1223-2012 | 61274 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Vanguard Knee Instrumentation Tibial Resection Tower Universal with alignment tower geometry, REF 32487557, LOT 2561334, Biomet UK Ltd, South Wales, UK |  |  | February 14, 2012 | 2012 | Biomet U.K., Ltd. | This recall was initiated following a report that a bag labeled as REF. 32-487557/Lot no 2561334 (Vanguard Tibial Resection head universal w/ alignment tower) actually contained a Ref. 32-487002/Lot no. 2561329 (Vanguard Distal cutting Block w/out handle attachment). Investigation found packaging of the affected product may contain the incorrect item. | Packaging change control | Biomet U.K. sent a Field Safety Corrective Action notice dated February 14, 2012 to all affected customers. The notice identified the affected products, problem and actions to be taken. Customers were instructed to locate, discontinue use and quarantine the recalled products pending return to Biomet. The letter recommended customers provide the notice to any other organization who may have received the affected instruments. Customers were instructed to complete and return without delay the attached "Fax-Back Response Form" acknowledging receipt of notice and indicating the quantity of products to be returned. For questions call +44 (0) 1656 761658, fax: +44 (0) 1656 645454 and e-mail uk.complaints@biomet.com |  | 4 | 4 units | Worldwide Distribution - in the countries of France, Japan, Italy, Netherlands, UK , no devices distributed in the USA | Terminated | November 06, 2012 | 266 | 4 | Not_Computer
Z-1229-2012 | 61122 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron CX Immuno-Protein Cal, part number: 442840, |  |  | August 26, 2009 | 2009 | Beckman Coulter Inc. | A recall was initiated because Beckman has confirmed that the standardization of the Immuno Protein Calibrator Set has been re-established to IFCC Reference Preparation for Plasma Proteins, BCR-470 beginning with lot number M901330. | Nonconforming Material/Component | Beckman Coulter sent an Urgent Field Safety Notice letter dated August 24, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to share the information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers with questions regarding the product correction action were instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-993-8916. |  | 1923 | 1,923 units total (1,882 units in US) | Worldwide Distribution - USA (nationwide) and Canada | Terminated | April 17, 2012 | 965 | 1923 | Not_Computer
Z-1230-2012 | 61197 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 Systems, Product Code M-4800-01 |  |  | February 09, 2012 | 2012 | Biosense Webster, Inc. | Some CARTO 3 Systems may get radio frequency interference through certain interconnected cable configurations resulting in jumping or disappearing catheter icon and/or error messages during ablation. | Nonconforming Material/Component | Three Urgent Customer Notification recall communications were initiated 2/9/12 with Biosense forwarding the notification to customers with affected backplane cards (electrophysiology lab managers, a notification for customers with affected back plane cards (risk managers), and a notification for customers with affected backplane cards and smarttouch module along with customer acknowledgement forms. The letters provides the customers with an explanation of the problem identified and steps to take in the event that the issue is observed. If the steps provided do not resolve the issue, customers are to contact their Biosense Webster representative. A Biosense Webster representative will contact customers to schedule an appointment for the replacement of the Backplane card.Customers were instructed to complete and return the attached Acknowledgement Form in accordance with the instructions listed on the form. If customers have questions, they can contact their BWI representative. |  | 134 | 134 units total (32 units in US) | Worldwide Distribution -- US, Japan, China, Ireland, Germany, Italy, Slovenia, Indonesia, Poland, Spain, United Kingdom, Switzerland, Belgium, and Canada. | Terminated | September 06, 2012 | 210 | 134 | Not_Computer
Z-1271-2012 | 61029 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITEÂż/IMMULITEÂż 1000 CK-MB LKMB --- Catalog number: LKMB1 (100 tests), LKMB5 (500 tests) --- Test code: CMB Color: Dark Blue. |  |  | January 25, 2012 | 2012 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics has observed that the median for the reference range with current kit lots of the CK-MB assay differs from the guidelines published in the Instructions for Use (IFU). This difference applies to all IMMULITE platforms. | Nonconforming Material/Component | Siemens sent an "URGENT DEVICE RECALL" letter dated January 2012 to all affected customers. The letter identifies the product , problem, and actions to be taken by the customers. Siemens recommends that customers reestablish their own medians and reference ranges, in line with the recommendations made in the IFU for these products. A Field Correction Effectiveness Check form was included for customers to complete and return to the firm. Customers with questions or who need additional information should contact their Siemens Technical Solutions Center. |  |  |  | Worldwide Distribution. | Open, Classified |  |  |  | Not_Computer
Z-1231-2012 | 61104 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | REF 905463, 2.9 mm Micromax Anchor Long Shaft, Size 2 Blue/White Maxbraid (Lactosorb L-15 resorbable copolymer ) Sterile, EO; Biomet Sports Medicine, Warsaw, IN |  |  | January 25, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action following an internal audit which discovered the component level package insert was included in the product packaging in place of the end level product package insert (01-501072).This meant that the Indications, Contraindications, Possible Adverse Effects, and some Warnings are not present with the packaged device. | Packaging | Biomed Inc. sent an Urgent Medical Device Recall Notice dated January 25, 2012 to distributors, hospitals, and OR managers via US mail. The notice identified the product, description of problem, possible adverse health consequences, and actions to be taken, including directions to immediately locate devices and stop device use ,with directions for return. A copy of the correct package insert was included with directions to report any procedures performed outside of indications for use. Receipt confirmation was requested by calling 800-348-9500, extension 3983 or 3009. The notice included a list of units invoiced to the account with a FaxBack Response form to be filled out and sent to the Attention of Angie Dickson, Biomet, 56 East Bell Drive Warsaw, In 46580. Questions concerning the recall should be directed to 574-371-3983 or 574-371-3009, M-F 8-5PM(EST) |  | 787 | 787 | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Brazil, Chile, Colombia, Israel, Mexico and Netherlands. | Terminated | April 12, 2013 | 443 | 4212 | Not_Computer
Z-1241-2012 | 61255 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 780 Analyzer, |  |  | December 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Auto Stop and Auto Transmit features for Quality Control become disabled after scanning assay values from Coulter 5C and Retic-C Cell Controls assay sheets using the 2D Barcode scanner. | Software design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 30, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers who use the Auto-Stop or Auto-Transmit to ensure the settings are enabled in Control Setup after scanning the assay sheets. Additionally, a Response Form was included for customers to complete and return. Contact Beckman Coulter Customer Service at (800) 526-7694 for questions regarding this notice. |  | 8922 | 8,922 units total (4,073 units in US) | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the countries of Algeria Andorra Angola Antigua and Barbuda Argentina Australia Bahrain Bangladesh Belgium Bermuda Brazil Bulgaria Burundi Cameroon Canada Chile China Colombia Croatia Czech Republic Denmark Ecuador Egypt Ethiopia France Germany Greece Guadeloupe Guam Guatemala Honduras Hong Kong Hungary Iceland India Indonesia Iraq Ireland Israel Italy Japan Jordan Korea, Republic of Kuwait Latvia Lebanon Libyan Arab Jamahiriya Lithuania Macao Malaysia Mayotte Mexico Monaco Morocco Myanmar Namibia Netherlands New Zealand Norway Oman Pakistan Panama Philippines Poland Puerto Rico Qatar Reunion Romania Russian Federation Saudi Arabia Serbia Singapore Slovakia Slovenia South Africa Spain Sri Lanka Sweden Switzerland Taiwan Thailand Trinidad and Tobago Tunisia Turkey United Arab Emirates United Kingdom Uruguay Venezuela Viet Nam Yemen Zimbabwe | Terminated | September 02, 2014 | 1006 | 8922 | Not_Computer
Z-1239-2012 | 61214 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Ostase QC, Part Number: 37309. |  |  | November 10, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman confirmed that the Access Ostase Calibrator and QC lots do not meet expiration date claims. The Access Ostase Reagent Kit lots identified were quality control tested and released using Access Ostase Calibrator and QC lots that contain the implicated raw material. | Process design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 9, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard all remaining inventory. In addition, a Response Form was enclosed for customers to complete and return. Contact Beckman's Customer Support Center at (800) 854-3633 for questions regarding this notice. |  | 4961 | 4961 units total (1321 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Belgium, Canada, China, Colombia, Czech Republic, Denmark, France, Georgia, Germany, Greece, Hong Kong, Hungary, India,Ireland, Italy, Japan, Jordan, Lebanon, Liechtenstein, Mexico, Morocco, Netherlands, New Zealand, Panama, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | June 26, 2012 | 229 | 4961 | Not_Computer
Z-1214-2012 | 61298 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | D-SPECT Cardiac Scanner System -- COMMON/USUAL NAME: Emission Computed Tomography System --- SPECTRUM DYNAMICS --- Spectrum Dynamics (Israel) Ltd. 22 Bareket St. North Industrial Park, P.O. Box 3033, Caesarea, 30889 Israel |  |  | July 15, 2011 | 2011 | Spectrum Dynamics | The firm received a report related to the detector housing on the D-SPECT Cardiac Scanner. Specifically, it was reported that if the detector housing securing mechanism is compromised, the detector housing may pull away from its secured position on the gantry and potentially come intocontact with a patient. The reported event was discovered during set up and thus did not involve a patient or p | Nonconforming Material/Component | Spectrum Dynamics sent a "Voluntary Field Correction" letter dated July 15, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken.To implement this fix, a BC Technical Field Service Engineer contacted users to schedule a service visit. In the meantime, Spectrum Dynamics did not believe that this issue warranted the discontinuation of use of the D-SPECT Cardiac Scanner. However, in an abundance of caution, if the detector housing appeared to be loose, or appeared to be disconnecting from the gantry, do not use the system. Please contact your local BC Technical Field Service Engineer immediately.Enclosed is a verification form. Please fax the completed form at your earliest convenience to 415.663.4225. For further information or support concerning this issue, please call 925.922.2615. |  | 42 | Domestic- 35 units, Foreign - 7 units | Worldwide Distribution -- USA (nationwide) and the countries of France and United Kingdom | Terminated | April 01, 2015 | 1356 | 42 | Not_Computer
Z-1270-2012 | 61308 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ System; |  |  | February 14, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Images used in treatment planning may belong to another patient. Risk of treatment to wrong area. Software update to upgrade Siemens installed base to the latest release of MOSAIQ version 2.00X3 SP5, for RT Therapist version 2.1A | Device Design | Siemens sent a Customer Information letters to all affected customers. The Customer information letters inform customers of the release of MOSAIQ 2.00X3 SP5 which specifically addresses a safety issue that is only found whenMOSAIQ is used with COHERENCE Therapist 2.1 MQ (also called RTT2.1A ). This letter is to be kept with Customer information documentation. Update Instruction TH004/12S are being issued with the update. The Software update is delivered and installed by qualified Field Service Engineers. For questions, contact your responsible SIEMENS Service Engineer or Application Specialist. |  | 44 | 44 affected Siemens Customers | Worldwide Distribution - USA (nationwide) including the states of KY, PA, TN, WI, and P.R. and the countries of Australia, Belgium, Bosnia Herzeg. China, Colombia, Croatia, Czech Republic, Germany, India, Ireland, Japan, Lebanon, Malaysia, Mexico, New Zealand, Philippines, Poland, Russian Federation, South Africa, Saudi Arabia, Thailand, Turkey, and United Kingdom. | Terminated | September 21, 2012 | 220 | 44 | Computer
Z-1275-2012 | 57924 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CRYO-206V and CRYO-206F Procedure Kit, which includes: Cryo Probes, Urethral Catheter, and temperature probes. |  |  | February 27, 2012 | 2012 | Healthtronics, Inc. | HealthTronics has initiated this recall of these products because there have been instances where the temperature probes (aka TempProbes) in the kit may be mislabeled. In some instances, the label near the prove does not match the label at the connector. If the labeling error went undetected, the TempProbes would not be monitoring temperature in the area believed to be under surveillance. | Labeling mix-ups | Endocare sent an Urgent: Device Recall letter dated February 27, 2012 to all affected consignees. The letter identified the affected product, problem and actions to be taken. The letter instructed consignees to immediately cease all distribution, examine their stock and the stock of accounts that have been supplied and remove any affected devices from all inventories. If the product have been further distributed contact Endocare for instruction. If consignees have any of the affected products: 1) Contact Customer Service at 877-367-4873 to obtain a RMA number, 2) Complete the Verification Section on page two, and 3) Return this letter with the affected product to the address provided. For questions about this recall contact Endocare, Inc at 877-367-4873. |  | 9274 | 9274 units | Worldwide Distribution--USA (nationwide) and the countries of Barbados, Brazil, Canada, China, Italy, Jakarta, Malaysia, Madrid Spain, Taiwan, and United Kingdom. | Terminated | March 19, 2013 | 386 | 9274 | Not_Computer
Z-1269-2012 | 61294 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Touch Panel/Control Unit Carrier of the Swingarm M820/M844 C19, article no. 10448194MO; a component of the Leica M820 C19 Surgical Microscope, article 10448196, and the Leica M844 C19 Surgical Microscope, article 10448197; ceiling mounted surgical microscopes; Leica Microsystems (Schweiz) AG, Max Schmidheiny-Strasse 201, CH-9425, Heerbrugg, Switzerland;; The intended use of the microscope is for improving the visibility of objects through magnification and illumination. It can be applied for observation and documentation and for human and veterinary medical treatment. |  |  | February 27, 2012 | 2012 | Leica Microsystems, Inc. | The screws, which lock the bottom plate to the spindle, pull through allowing the carrier arm with touch panel/control unit of the surgical microscope to disconnect, which potentially could result in them falling. | Device Design | The firm, Leica Microsystems, sent a "Medical Device Correction" letter dated February 27, 2012 to the affected customers via e-mail on the same date. The letter described the product, problem and action to be taken. The customers were informed that the Leica M820 and M844 C19 surgical microscopes require upgrading. An upgrade kit will be provide for replacement of the flange and axis of the touch panel/control unit carrier. The Customers were instructed to NOT USE the affected Leica surgical microscope until the unit has been upgraded by their Leica Microsystems representative and to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the serial numbers of the unit at their facility and fax it back to Leica at 1-847-236-3747. If they have any questions, contact Director, Global Product Quality Management at 1-847-317-7209. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including states of: California, Connecticut, Kansas, Missouri, New York, North Dakota, Ohio, Pennsylvania, Texas and Utah; and countries of: Japan and South Africa, | Terminated | May 18, 2012 | 81 | 33 | Not_Computer
Z-1246-2012 | 61314 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | ENSEALÂż 5 mm Diameter Tissue Sealer G2 14 cm Length Straight Jaw, Model # NSLG2S14 and ENSEALÂż 5 mm Diameter Tissue Sealer G2 25 cm Length Straight Jaw, Model # NSLG2S25. Ethicon Endo-Surgery, LLC Guaynabo, Puerto Rico 00969 |  |  | January 26, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary global recall for specific production lots of ENSEALÂż G2 Curved and Straight Tissue Sealers due to two potential issues that may occur related to the activation button: (i) Continuous Activation: The ENSEALÂż device/system may continue to activate when the activation button is released prior to reaching the end of cycle (tone 3). (ii) No Activation | Nonconforming Material/Component | Ethicon Endo-Surgery sent an Urgent Medical Device Recall letter dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop use of the affected product, fill out the attached Business Reply Form (BRF) and return to and return the affected product to:Ethicon Endo-SurgeryATT: ENSEAL G2 Recall 4545 Greek RoadCincinnati, Ohio 45242Customers were instructed to choose one of the following response options:Return the Business Reply Form to their sales representativeCall 1-800-873-3636, Option 6Fax the BRF to 1-513-337-4138For any questions regarding this recall call 513-337-3419. |  | 84 | 84 pieces | Worldwide Distribution - USA including AZ, CA, CT, FL, IL, IN, KS, KY, MA, MI, MO, NC, NY, OH, PA, TN &amp; TX and the countries of Arab Emirates, Austria, Belgium, France, Germany, Great Britain, Italy, Jordan, Kingdom of Saudi Arabia, Lebanon, Portugal, Slovenia, Sweden and Turkey. | Terminated | February 15, 2013 | 386 | 1035 | Not_Computer
Z-1263-2012 | 60753 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Wound Management Tray, catalog# 900-540 |  |  | December 08, 2011 | 2011 | Customed, Inc | The firm has determined that there is a possibility that packaging seals may be compromised. The issue could result in an injury to the patient due to product contamination or loss of sterility condition. | Packaging | Customed Inc sent an Urgent Medical Device Recall letter dated December 8, 2011 to distributors via e-mail; then by certified (return and receipt) mail. The letter identified the product, description of problem, and actions to be taken. The letter included the following instructions:1. Product should be removed from sale (Quarantine them) and cease distribution. Products affected shall NOT be used.2. Return the enclosed acknowledgement form to Customed, Inc. Recall Coordinator whether or not the consignee has affected products. 3. If products has been futher distributed, sub-recall products - It is extremely important to notify your customers of this action by including a copy of this leller along with your formal notification. Please, instruct your customer to follow your directions for credit/replacemenl if they purchased the product from you. For qucstions. contact Cuslomed, Inc. Recall Coordinator at 787-622-5151 Ext. 7510. |  | 400 | 400 units | USA Nationwide Distribution - including Florida and Puerto Rico. | Terminated | April 12, 2017 | 1952 | 184527 | Not_Computer
Z-1268-2012 | 61236 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Vitamin D Diluent 2-pack, |  |  | February 16, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Automatic and system-performed dilution calculation factor for Vitamin D dilutions is incorrect and causesdiluted patient samples to under recover by approximately 50%. | Nonconforming Material/Component | The firm, Siemens Healthcare Diagnostics, sent an "Urgent Device Recall Notice" on February 16, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to discontinue using the automatic and system-performed dilution features for Vitamin D; to discontinue the use of ADVIA Centaur VitD Diluent, 2-pack (REF 10494100); to manually dilute samples with values greater than the assay range of 150 ng/mL (375 nmoI/L), using the ADVIA Centaur VitD Diluent, 1 bottle (REF 10632114) as stated in the Vitamin D assay Instructions For Use (IFU); to review the contents of the recall notice with their Laboratory Director and retest the previous sample results generated using system- performed dilutions, and complete and return the confirmation fax-back form via fax to TECHNICAL SOLUTIONS CENTER at (302) 631-7597.If you have any questions or need additional information, please contact your local Technical Support Provider or Distributor or call (508) 668-5000. |  | 442 | 442 units | Worldwide distribution: USA (nationwide) and countries of: Columbia, China, Denmark, Egypt, and Hong Kong. | Terminated | October 17, 2013 | 609 | 442 | Not_Computer
Z-1278-2012 | 61351 | No | KNN | Wheelchair component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Permolock C3, Docking system for Power wheelchair in vehicle. |  |  | February 21, 2012 | 2012 | Permobil Inc | The fasteners used between June 3, 2011 and January 31, 2012 to attach the Permolock C3 pins to wheelchairs manufactured in the United States do not contain the required nylon ring, which allows for a more secure attachment. | Mixed-up of materials/components | The firm, Permobil, sent an "URGENT DEVICE RECALL" letter dated February 27, 2012 via fax to all customers. For those customers where Permobil did not have a valid fax number, the notices were mailed via first class United States mail. The letter describes the product, problem and actions to be taken. Permobil will designate Permobil personnel to complete an update to the Permolock C3 device.If a Permobil representative has not contacted you already, please call 800-736-0925 to schedule the service. |  | 147 | 147 units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, MD ,MI, MN, MO, NC, NE, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, VT, WV, and WY; and countries of: Australia, Canada, France, Netherlands, Sweden, and United Kingdom. | Terminated | August 13, 2013 | 539 | 147 | Not_Computer
Z-1277-2012 | 61175 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products --- ALT Slides --- [REF] 165 5281 --- [Mfg Symbol] Ortho-Clinical Diagnostics, Inc., Rochester, NY USA. The recalled device is the calibration disk contained within the VITROS Calibrator Kit 3. The is can be identified with the label "VITROS Chemist y Systems --- Universal Calibration Diskette --- [REF] 871 6607 --- Ortho-Clinical Diagnostics" |  |  | February 02, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics received complaints of lower than expected results (&lt; 6 U/L) for samples that were known to be outside of the measuring (reportable) range (&gt; 1000 U/L) when using Calibration Diskette/Assay Data Diskette Data Release Version (DRV) 5653 through 5666 and VITROSÂż Chemistry Products ALT Slides, Generation (GEN) 26 (Product Code 1655281). | Software design | Ortho Clinical Diagnostic sent a Important Product Correction letter dated February 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed customers to: Install Calibration Diskette/Assay Data Diskette DRV 5667 on your VITROS System. Following a successful recalibration, you may resume using VITROSÂż ALT Slides, GEN 26 remaining in your inventory. To please complete and return the Confirmation of Receipt form upon receipt of this notification. Please return this form no later than February 17, 2012. Post this notification by each VITROSÂż System in your facility that utilizes the VITROSÂż ALT Slides, or with the VITROSÂż user documentation. Forward this notification if you have distributed this product outside of your facility.We apologize for the inconvenience this may have caused your facility. If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 4306 | 4306 | Worldwide Distribution - USA (nationwide) | Open, Classified |  |  | 4306 | Computer
Z-1276-2012 | 61312 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Osteonics Shoulder Glenoid Spanner |  |  | November 17, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Two product complaints reported that the Glenoid Reamer Spanner (Lot F3W7017) would not properly connect to the Glenoid Spherical Reamer. | Nonconforming Material/Component | Stryker sent an "URGENT PRODUCT RECALL" letter dated December 21, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Product Recall Acknowledgment Form was attached for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-972-2100 for questions regarding this notice. |  | 5 | five units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | December 19, 2012 | 398 | 5 | Not_Computer
Z-1248-2012 | 61356 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic EnTrust, Dual Chamber Implantable, Cardioverter Defibrillator with Atrial and Ventricular Therapies, models D153ATG and D154ATG. The device is indicated for use in ICD patients with atrial tachyarrhythmias, or who are at significant risk of developing atrial tachyarrhythmias. The device is intended to provide ventricular antitachycardia pacing and ventricular defibrillation for automated treatment of life-threatening ventricular arrhythmias. In addition, the device is intended to provide pacing, cardioversion, and defibrillation for treatment of patients with symptomatic, drug-refractory atrial fibrillation and/or life-threatening ventricular tachyarrhythmias. Sterile. Medtronic, Inc., Minneapolis, MN 55432 USA. |  |  | March 06, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | A small percentage of EnTrust ICDs may not meet expected longevity or provide at least three months of device operation between the Elective Replacement Indicator (ERI) and End of Life (EOL) due to a more-rapid-than-expected drop in battery voltage. | Device Design | An "Urgent Medical Device Correction" letter was sent to physicians beginning 3/6/2012. The letter described the issue and provided patient management recommendations. Additionally the letter stated that ongoing updates to the product performance report will be posted on a Medtronic website. |  | 68804 | 68,804 devices worldwide | Worldwide Distribution -- USA, and the countries of: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Bolivia, Brazil, Canada, Cayman Islands, Chile, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jamaica, Jordan, Kuwait, Lebanon, Macedonia, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Reunion, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | July 27, 2012 | 143 | 68804 | Not_Computer
Z-1281-2012 | 61420 | No | KZA | Liquid crystal vein locator. | General Hospital | General Hospital | 510(K) Exempt | VeinViewer (R) by Luminetx, Model Number VV1.1 GS, Part No. P00800-L, Luminetx Corporation, Memphis, TN 38104 |  |  | March 14, 2012 | 2012 | Christie Medical Holding Inc | Instrument may tip over if not moved correctly. | Employee error | Christie Medical Holdings, Inc. sent an Urgent Medical Device Correction by letter on March 14, 2012 to all affected customers. The letter identified the product, problem and actions to be taken. The letter includes "Correct Transport" and "Do Not Pull labels to be applied to the unit. Customers were instructed to immediately examine their inventory, quarantine all product subject to correction and apply the labels prior to use. A Service Bulletin is enclosed that should be kept and placed with the User Guide for future reference and a Response Form to be completed and return as soon as possible. For questions contact Technical Service at 901-721-0300. |  | 600 | 600 units | Worldwide Distribution - USA (nationwide) and the countries of: Argentina, Australia, China, Brazil, Czech Republic, Canada, Germany, Denmark, Spain, Great Britain, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, South Korea, Kuwait, Mexico, Malaysia, Netherlands, Norway, Philippines, Russian Federation, Saudi Arabia, Singapore, Thailand, Turkey, Taiwan, and United Arab Emirates | Terminated | October 23, 2012 | 223 | 1050 | Not_Computer
Z-1279-2012 | 61205 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Codman ICP Skull Bolt Kit, Product Code: 82-6638. Single Use, disposable drill kit for gaining access to the cranium to allow for insertion and retention of intracranial cranial pressure sensor, included in the kit. |  |  | February 09, 2012 | 2012 | Codman &amp; Shurtleff, Inc. | Codman ICP Skull Bolt wrench is out of specification and cannot be used to turn the screw on the drill | Nonconforming Material/Component | Codman issued "Urgent - Voluntary Product Recall Notification" letters on 2/9/12 via Fedx in the US and requested accounts examine inventory and return product. An Acknowledgement Form was requested to be returned. Outside the US affiliates were notified to recall to their accounts. Questions were directed to Customer Service (800) 225-0460. |  | 4056 | 4056 kits | Worldwide Distribution - USA, and the countries of Canada, Austria, Brazil, Belgium, Columbia, China, Denmark, Estonia, Spain, Finland, France, Great Britain, Germany, India, Italy, Japan, Lithuania, Netherlands, New Zealand, Oman, Mexico, Portugal, Russia, Singapore, SIovenia, and Switzerland. | Terminated | September 11, 2015 | 1310 | 4056 | Not_Computer
Z-1287-2012 | 60718 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Soft-Flow Aortic Cannula; Aortic cannula, 7.0 mm (21 Fr) OD with 3/8" connector, 12" (31 em) long. |  |  | February 23, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During in-house inspection, Terumo Cardiovascular Systems discovered one Soft-Flow Aortic Cannula with a hole in the tip, just above the suture ball. | Nonconforming Material/Component | Terumo sent an URGENT MEDICAL DEVICE RECALL letter dated February 23, 2012 via Federal Express to all US customers. Terumo Foreign Affiliates received the recall notification via email on February 24th, 2012. The letters identified the affected products, reson for removal, potential hazard, affected population and actions to be taken. Customers were instructed to review the Medical Device Removal notice, assure that all users are aware of the notice, confirm receipt of this communication by faxing the attached Customer Response Form to the fax number provided or email to tcvs.recall@terumomedical.com and call Terumo CVS Customer Service for a Returned Goods Authorization (RGA) number. For questions call Terumo CVS Customer Service at 1-800-521-2818. |  | 188530 | 188,282 Cannulae and 248 Custom Tubing Packs | USA Nationwide Distribution including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of: CANADA, CHILE, India, Japan, Singapore, AUSTRALIA, Malaysia, BELGIUM, BELGIUM, BRAZIL, CHILE, BRAZIL, BRAZIL, and Philippines. | Terminated | April 03, 2013 | 405 | 4718744 | Not_Computer
Z-1284-2012 | 61278 | No | MIH |  |  | Cardiovascular | PMA | Talent Captivia Stent Graft System - FreeFlow configurations only. |  |  | February 21, 2012 | 2012 | Medtronic Vascular, Inc. | Complaints received related to the inability/difficulty to release the proximal (FreeFlow) bare stent using the tip capture mechanism. | Process design | Medtronic, Inc. sent an Urgent Medical Device Safety Information letter dated February 2012, to all affected customers. The letter identified the product, the problem, and the corrective action taken by Medtronic. Medtronic updated the Instructions for Use (IFU) manual.For questions regarding this recall call 707-566-1110. |  | 1523 | 1523 units | Nationwide Distribution | Terminated | July 12, 2012 | 142 | 1523 | Not_Computer
Z-1228-2012 | 61268 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II Left Ventricular Assist System (HM II LVAS) Sealed Outflow Graft Bend Relief. |  |  | February 23, 2012 | 2012 | Thoratec Corporation | A trend of disconnected bend reliefs on the sealed outflow graft used with the HeartMate II LVAS has been identified | Device Design | Thoractec Corporation sent an Urgent Medical Device Correction letter on February 23, 2012 via overnight courier and e-mail to all affected consignees. The Urgent Medical Device Correction notification letters identify the affected products, description of the problem with illustrations, symptoms of problem and immediate actions to taken. All users are asked to review the attached labeling, including new caution statements and share with all surgical personnel. Consignees are instructed to completed and returned the Acknowledgement Form indicating that the information was received and understood. Additionally, an update concerning the Worldwide Medical Device Correction Notification is enclosed with each letter.Questions should be directed to Thoratec Regulatory Affairs at 925-847-8571. |  | 3852 | 3852 units in commerce | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Belgium, Denmark, Germany, Greece, France, Hong Kong, Kuwait, Ireland, Israel, Italy, Lebanon, Liechtenstein, Lithuania, Luxembourg, Malaysia, Netherlands, Turkey, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand and United Kingdom. | Terminated | November 08, 2012 | 259 | 3852 | Not_Computer
Z-1283-2012 | 61365 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Diagnost Eleva DRF, Software release 4.0 |  |  | March 09, 2012 | 2012 | Philips Medical Systems Gmbh, DMC | Underexposed image due to lag in switching betweenauxiliaries | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated February 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The firm provided a work around for customers to use until a software update can be installed. Customers should contact their local Philips representative for questions regarding this notice. |  | 88 | 88 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Germany, Indonesia, Malaysia, Saudi-Arabia, Singapore, and Thailand. | Terminated | May 23, 2013 | 440 | 88 | Computer
Z-1282-2012 | 61305 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | Fildes Enrichment, BBL brand, catalog #211866, in 5 ml tubes,10 tubes/package, manufactured by BD Diagnostic Systems, 250 Schilling Circle, Cockeysville, MD 21030. |  |  | February 20, 2012 | 2012 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent is defective and may fail to enhance the growth of microbiological cultures as intended. | Nonconforming Material/Component | Becton Dickinson (BD) sent an "URGENT PRODUCT RECALL" letter dated February 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard any remaining affected product. BD will issue replacements for the discarded product. Additionally, an Acknowledgment Form was attached for customers to complete and return to the firm via fax to 410-316-4258. Contact BD Technical Services Department for questions regarding this notice. |  | 6930 | 6930 tubes | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, China, Japan, Mexico, and Taiwan. | Terminated | June 18, 2012 | 119 | 6930 | Not_Computer
Z-1315-2012 | 61282 | [6 empty] | QRS Diagnostic Universal ECG 12 Channel ECG (IEC), Rx only, REF Z-7000-0400, Manufactured for QRS Diagnostic, 14755 27th Ave North, Plymouth, MN 55447 USA. |  |  | February 16, 2012 | 2012 | Qrs Diagnostic | QRS Diagnostic has initiated a recall involving Universal ECG 12 Lead IEC ECG Devices. This recall has been initiated because the chest lead wires on these devices may have been improperly labeled. Incorrect labeling of the lead wires could result in improper lead wire placement on the patient which may result in inaccurate waveforms being displayed to the user and an inaccurate interpretatio | Process control | QRS Diagnostic sent an Urgent Medical Device Recall letter dated February 9, 2012 via e-mail and mail to all affected consignees. The letter identified the affected product, description of the problem, and actions to be taken. If the affected product have been further distributed, consignees are instructed to notify them at once of the product recall. Consignees were instructed to contact Uni-com immediately for a replacement device. Uni-com will provide this replacement at no charge and will arrange to have the device returned to QRS Diagnostic for correction. Additionally, consignees were instructed to complete the attached Medical Device Correction Receipt Confirmation Form and return to QRS Diagnostic via fax at country code =1-763-559-2961 or email support@QRSDiagnostic.com. For questions contact Technical Support at support@QRSDiagnostic.com or by phone 800-465-8408. |  | 13 | 13 devices | Worldwide Distribution and the countries of: Belgium, France, Italy, Mali and Senegal. | Terminated | August 02, 2012 | 168 | 13 | Not_Computer
Z-1285-2012 | 61363 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Whole body Computed Tomography X-Ray System (Brilliance 64 and Ingenuity CT), Brilliance 64 M/N 728231, Ingenuity CT M/N 728326, Dist. By Philips Medical Systems, Cleveland, OH |  |  | February 17, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were identified during internal and external validation testing of the Brilliance 64 and Ingenuity CT systems with software version Beta 3 (version 4.0.0.xx368). | Software change control | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 17, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of the contents in the communication. Customers were asked to retain a copy of the communication with the equipment instruction for Use. For further information or support concerning this issue, customers should contact their local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 440-483-7000. |  | 7 | 7 units | Worldwide Distribution - USA including OH and the following countries: Australia, China, Belgium, Israel &amp; Sweden. | Terminated | February 22, 2013 | 371 | 7 | Computer
Z-1286-2012 | 61207 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Product is labeled in part: "***REF: 01-7036, LOT 999999, "1.5MM" SYSTEM PLATE L-SHAPE, 3 X 3 HOLE, 100 DEGREE, QTY 00001, CE 0086, 2010-04 *** NON-STERILE PRODUCT ***, LEFT, X-LONG, 0.6 MM TITANIUM, BIOMET MICROFIXATION, 1540 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA, 01-736, LOT 999999, QTY 00001, L-SHAPE, 3 X 3 HOLE, 100 DEGREE, LEFT, X-LONG, 0.6 MM, CE 0086.***" |  |  | April 26, 2010 | 2010 | Biomet Microfixation, LLC | On 4/26/10 Biomet Microfixation, Jacksonville, FL initiated a recall of their 1.5 3/3 hole 100 degree left X-long L plate, Model #01-7036, Lot #193690. Product is labeled incorrectly. | Error in labeling | Biomet Microfixation, LLC sent an Urgent Medical Device Safety Alert letter dated April 21, 2010 to all consignees affected. Each consignee was notified by E-mail, letter or telephone call the week of April 26, 2010. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to fill out the attached Inventory Reconciliation sheet and FAX it to the attention of the Global Complaint Coordinator at 904-741-9425. If the affected product have been further distribute, consignees are instructed to notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field action notices. For questions contact Christy Cain at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 23 | 23 | USA Nationwide Distribution including the states of: California and Texas. | Terminated | April 18, 2012 | 723 | 23 | Not_Computer
Z-1318-2012 | 60647 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Anesthesia Machines (digital display assembly); |  |  | November 29, 2011 | 2011 | Penlon, Ltd. | Penlon, Ltd has identified the potential for the premature failure of a component in some Electronic Flowmeter sensor display unit PCB assemblies. | Component design/selection | Penlon Limited sent a Technical Bulletin dated November 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. As a precaution a field upgrade part was released that required all existing PCB assemblies to be removed and the updated item installed. The old PCB must be returned to Penlon with the serial number of the affected machine. The required replacement parts were included with the technical bulletin.For further information customers should contact the Technical Support Team at techsupport@penlon.com. |  | 160 | 160 units | Worldwide Distribution - USA (nationwide) and the countries of Peru, Uganda, Kenya, Malaysia, Poland, Saudi Arabia, United Arab Emirates, Philippines, Equador, and Thailand | Terminated | July 03, 2012 | 217 | 160 | Not_Computer
Z-1317-2012 | 61269 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Version 2.0.0.15, Radiation Therapy Treatment Planning System. |  |  | December 06, 2011 | 2011 | RAYSEARCH LABORATORIES AB | An issue was found with the dose calculation in the RayStation 2.0, that could result in a potential overestimation of the dose. This could lead to the potential under-dosage of a patient. | Device Design | The firm decided to recall and sent Urgent Field Safety Notice Medical Device Correction letters to consignees on 12/06/2011. The letters described the issue, identified the affected product, actions to be taken by the user, a solution, and contact information. Users are to follow the two part workaround, and educate their staff and all users on the workarounds. The letter states that the issue will be resolved in the next version of RayStation. If further information is needed, users are to contact either Freddie Cardel, Director of Customer Support, RaySearch Americas, at 1-877-778-3849 or freddie.cardel@raysearchlabs.com; or Niclas Borglund, Director of Service, RaySearch Laboratories AB, at 46-8-5450-6130 or niclas.borglund@raysearchlabs.com |  | 21 | 21 units (16 domestically and 5 internationally) | Worldwide Distribution -- US, including the states of FL, MA, NJ, and NY and the country of the Netherlands. | Terminated | May 11, 2012 | 157 | 21 | Not_Computer
Z-1316-2012 | 61222 | No | LBS | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products --- dHDL Slides --- Ortho Clinical Diagnostics, a Johnson&Johnson; company --- [Mfg Symbol] Ortho-Clinical Diagnostics, Inc., Rochester, NY USA --- VITROS dHDL Slides are used to quantitatively measure HDL cholesterol (HDLC) concentration in serum and plasma using VITROS Chemistry and Integrated Systems. --- Product codes: 6801895 (60 slide count); 6802469 (18 slide count) --- indicated to quantitatively measure HDL cholesterol (HDLC) concentration in serum and plasma using VITROS Chemistry and Integrated Systems. |  |  | February 02, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) has revised the VITROS dHDL Slides Instructions for Use (IFU) and removed EDTA plasma as a recommended sample type. | Device Design | Ortho Clinical Diagnostics sent an Important Product Information letter dated February 2, 2012, to all affected customers via Federal Express overnight mail. The letter informed customers of the issue, notifying them of the removal of EDTA plasma as a recommended sample type and providing the revised VITROSÂż dHDL Slides Instructions for Use. J&J; Foreign affiliate consignees were notified by email informing them of the issue on 02 February 2012. Customers were instructed to replace the current pages in their VITROS Chemistry Products Instructions for Use Manual with the updated instructions for Use and Table of Contents.Customers were asked to return the enclosed Confirmation of Receipt form regardless of whether their laboratory uses the affected product. For any questions regarding this recall call 1-800-421-3311. |  | 1366236.9 | 120,225 of 60-slide count, 11.364 of 18-slide count | Worldwide Distribution | Open, Classified |  |  | 1366236.9 | Not_Computer
Z-1322-2012 | 61170 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BBL (tm) Lowenstein-Jensen Medium Deeps, Catalog #221256, packaged in boxes of ten/600 mL tubes. Product Usage: Lowenstein-Jensen Medium is used for the isolation and cultivation of mycobacteria. The medium tubed as deeps is used for the semi-quantitative catalase test as an aid to the classification of mycobacteria. |  |  | January 30, 2012 | 2012 | Becton Dickinson &amp; Co. | Media fails to perform as intended with quality control organism Mycobacterium kansasii. | Nonconforming Material/Component | The firm, BD Diagnostic Systems, sent an "URGENT PRODUCT RECALL" letter dated January 2012 to Ex-US BD sites via email and its customers via UPS ground shipment on 1/30/12. The letter identified the product, the problem and the action to be taken. The customers were instructed to discontinue use of the recalled lot numbers and discard any remaining packages. Customers were also instructed to complete and return the attached form whether or not they had any inventory remaining via fax to: BD Regulatory Compliance at 410-316-4285. The letter also informed the customers that BD has decided to discontinue this product.If further assistance is needed regarding credit, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 3900 | 3900 tubes | Worldwide Distribution - USA (nationwide) including states of: AL, CA, CT, IA, IL, KY, LA, MD, MI, MN, NC, NJ, NY, OH, PA, TN, VA, WA, and WV; and countries of : Belgium and Singapore. | Terminated | April 02, 2012 | 63 | 3900 | Not_Computer
Z-1323-2012 | 61271 | No | IRQ | Powered heating unit. | Physical Medicine | Physical Medicine | 510(K) Exempt | The LactoSorb Heat/Contouring Pen is a disposable alkaline-battery powered device. |  |  | November 15, 2011 | 2011 | Biomet Microfixation, LLC | On 11/15/2011 Biomet Microfixation, Jacksonville, FL initiated a recall of their Acticort Heat/Contouring Pen, Model #915-3100, Lot #861200. Product is incorrectly labeled | Error in labeling | Biomet Microfixation, LLC sent an Urgent Medical Device Safety Alert letter dated November 15, 2011 via E-mail, letter or telephone call to all consignees affected. The letter identified the affected product, problem and actions to be taken. Each consignee was Instructioned to take the following actions: 1) If you have any unopened boxes of the product in question, please place this notice with it; 2) Please fill out the attached Inventory Reconciliation sheet and FAX to the attention of the Global Complaint Coordinator; and 3) If you have further distributed this product , you must notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field action notices. For questions please contact Christy Cain at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 42 | 42 | Worldwide Distribution - USA (nationwide) including the state of Texas and the country of Japan. | Terminated | August 04, 2014 | 993 | 42 | Not_Computer
Z-1321-2012 | 59573 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Base, 220/240V, Catalog 801764, 6 " Roller Pump Occlusion: Reference AA-2011-015 C. indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures. |  |  | March 16, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo CVS has received reports of 27 instances in which the 6 inch Roller Pump jammed during cardioplegia delivery. There has been one report of the pump jamming during a case that concluded with an adverse patient outcome, although there was no indication that the pump jam contributed to the adverse patient outcome. | Device Design | "URGENT MEDICAL DEVICE RECALL" letters were sent on 3/16/12 to all consignees detailing the correction as well as also adding an Addendum to the Operator's manual for the Terumo System 1 to incorporate additional information and warnings for use when operating a 6" roller pump with dual-sized tubing. For questions call 1-800-521-2818. |  | 648 | 648 units | Worldwide Distribution--USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, Puerto Rico and the countries of INDIA, CANADA, CHILE, HONG KONG, EGYPT, PAKISTAN, KUWAIT, SAUDI ARABIA, JORDAN, UNITED ARAB EMIRATES, SOUTH AFRICA, SOUTH KOREA, PHILIPPINES, JAPAN, PHILIPPINES, SINGAPORE, AUSTRALIA, MALAYSIA, TAIWAN, RUSSIA, TURKEY, BELGIUM, GERMANY, COLUMBIA, COSTA RICA, ARGENTINA, HONDURAS, COLUMBIA, DOMINICAN REPUBLIC, GUATEMALA, MEXICO, VIETNAM, and THAILAND. | Terminated | May 02, 2013 | 412 | 1642 | Not_Computer
Z-1178-2012 | 61311 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare EasyDiagnost Eleva DRF 4.0 |  |  | April 01, 2011 | 2011 | Philips Healthcare Inc. | It was discovered in some instances the detector did not reset as quickly as the user prompted a change form one image receptor (auxiliary) to the next, and then back again. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter, dated February 28, 2012, was sent to customers. The letter identified the affected product, a description of the issue, the hazard involved, along with the actions to be taken and planned. Customers were to use one of the three work-a-rounds provided. The issue will be corrected with Field Change Order 70600064, which consists of a Software Update 4.0.2SP Service Pack installed on the system. If further information or support is needed, customers can contact Dominic Siewko at 978-659-7936. |  | 96 | 96 | Worldwide Distribution - Including USA. | Completed |  |  | 96 | Computer
Z-1324-2012 | 61177 | No | MQL | Reproductive media and supplements. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sydney IVF Culture Oil, REF, K- SICO- 50, 50 mL. and K-SICO-200, 200 mL,Store at 2-8 C, Sterile, Cook Medical, William A Cook Australia LTD, |  |  | February 15, 2012 | 2012 | Cook, Inc. | Cook Medical is recalling Sydney IVF Culture Oil due to a number of recent customer complaints that have identified it as a possible cause of embryo degeneration at day 5/6. | Process control | Cook Medical sent an Urgent Recall Notices letter dated February 15, 2012 to all affected customers. The letter identified the recall product, reason for the recall and actions to be taken. The notice included directions to pull the recalled product from stock, isolate it and forward the notice to any sub-accounts. Customers were requested to complete and return the recall acknowlegement and product tracking form by to Fax: 812-829-1801 or Email: julie.cooper@cookmedical.com. For questions call 800-457-4448 x7241. |  | 54 | 54 total from both lots | Worldwide Distribution - USA (nationwide) including the states of IL, KS, KY, MA, NJ. TN, TX, UT, VA, WI, WV, and the country of Barbados. | Terminated | July 03, 2012 | 139 | 54 | Not_Computer
Z-1319-2012 | 61377 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific Soloist Single Needle Electrode: 18 cm long, 16 gauge, Sterile, packaged with the Soloist Electrode Power Cord. Boston Scientific, Nalick, MA. |  |  | March 02, 2012 | 2012 | Boston Scientific Corporation | Device found labeled with the incorrect expiration dates. The Soloist Single Needle Electrode has been qualified for a 3 year expiration date. Product manufactured between the dates of May 24, 2011 and November 8, 2011 has an incorrect expiration date of 20 years labeled on the pouched needle electrode and 5 years labeled on the carton. | Incorrect or no expiration date | Boston Scientific sent a Urgent Medical Device Recall letter dated March 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed to cease use, quarantine and return unused product with directions to forward the notification to the hospital level. Customers were directed to fill out the Reply Verification Tracking Form and fax to 508Âż683Âż5578, for reimbursement. Returned goods are to be sent to:Boston Scientific CorporationDistribution CenterAttn: QA Returns500 Commander Shea BlvdQuincy, MA 02171We regret any inconvenience that this action may cause and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. For further questions please call (508) 683-4678. |  | 682 | 682 (1555 used) | USA (nationwide) and the countries of Canada, Columbia, Viet Nam, Germany, France, Great Britain, Italy, Jordan, Kuwait, Netherlands Pakistan, Portugal, S. Arabia, Slovakia, Spain, Sweden and Tunisia. | Terminated | November 14, 2012 | 257 | 682 | Not_Computer
Z-1224-2012 | 60991 | No | RBS |  |  |  |  | Hamamatsu Corporation, Bridgewater NJ 08807 |  |  | March 30, 2011 | 2011 | Hamamatsu Corporation | It was discovered that the Micro Focus X-ray Source could be experiencing software related problems that would allow the unit to exceed the maximum power setting that was set by the operator. | Software design | Hamamtsu Corporation sent a Warning letter dated December 22, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. CDRH approves your CAP under the following conditions:1.Hamamatsu is able to work with your manufacturer customers to assure the revised manual and new label are provided to your customers' end users.2.Hamamatsu informs CDRH if you receive information which indicates this control defect has occurred. Relevant information includes direct communication from customers or through observation of excessive wear on returned x-ray tubes.For questions regarding this recall call 908-231-0960, ext 2300. |  | 100 | 100 | Worldwide Distribution-- USA (nationwide) and the country of Canada. | Open, Classified |  |  | 100 | Computer
Z-1212-2012 | 61028 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare MD Eleva system. |  |  | November 19, 2011 | 2011 | Philips Healthcare Inc. | It was discovered there was a software error where the air karma rate is not displayed correctly on the user interface. | Software design | A field correction action will be undertaken to correct the software error in affected systems worldwide. Your proposed action to bring the product into compliance with regulatory requirements is to institute a worldwide field correction to replace the software error in affected systems by updating the current software to version R6.1.2.The timeframe for completion of this field action will be 6 months post field action publication. Preventive actions: The software has been corrected moving forward in subsequent releases for these products.For further questions please call (978) 787-1501. |  | 92 | 92 ( 24 Eleva 11s, and 68 Eleva FDs ) | Worldwide Distribution - USA (nationwide) | Open, Classified |  |  | 92 | Computer
Z-1334-2012 | 61329 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion CO2 sampling lines and water traps under the label: |  |  | February 16, 2012 | 2012 | Oridion Medical 1987 Ltd. | CO2 sampling lines may have paper and plastic particulate present in the packaged units at levels that may exceed visual specification. | Packaging process control | Oridion notified Hospitals, Clinics, EMS with an "Important Communication" letter to clinicians dated February 12, 2012. The letter was intended to make US clinicians aware of an important issue relating to certain Oridion CO2 sampling lines, which may not fully meet quality requirements. In a few cases, paper and plastic particulate may be present in the packaged units at levels that exceed the visual specification. Recommended actions were included in the letter. Questions were directed to 1(781)-972-1252. |  | 2000000 | 2 million | Nationwide Distribution | Terminated | September 13, 2013 | 575 | 2000000 | Not_Computer
Z-1331-2012 | 61276 | preservative, cytological | LEA | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Surgipath Cyto Jar Fixative; a 120 mL clear, leakproof container prefilled with 45 mL of preservative containing polyethylene glycol for the convenient collection, fixation and preservation of all cytologic specimens, 50 jars per case; for in vitro diagnostic use only; Leica Biosystems, 5205 Rt. 12, Richmond, IL 60071; Made in USA |  |  | February 20, 2012 | 2012 | Leica Biosystems Richmond Inc. | Product was distributed with the incorrect lot number and expiration date. The lot number and expiration date should have been listed at 073013 and instead was noted as 073014 which is beyond the maximum expiration dating of 18 months for this product. | Process design | The firm, Leica Microsystems, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 17, 2012, via FedEx on February 20, 2012 to all affected customers. The letter describes the product, problem, and actions to be taken. A second letter was sent to the customers on February 29, 2012, clarifying that the incorrect lot number and expiration date was 073014, and should have been 073013. The customers were instructed to discontinue use of the product; return the product and complete and return the enclosed Acknowledgement Form via fax to 815-678-2216 within one week of receiving the notification.If you have any questions, contact the Regulatory Affairs Manager at toll free (800) 225-8867 or email: richmond@leica-microsystems.com. |  | 11 | 11 cases | Nationwide distribution: USA including states of: Michigan, Missouri and New York. | Terminated | September 20, 2012 | 213 | 11 | Not_Computer
Z-1330-2012 | 61375 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sorin C5 (perfusion system) |  |  | February 23, 2012 | 2012 | Sorin Group Deutschland GmbH | The pump speed control knob is attached to the shaft of an encoder that is used to adjust the speed. Insufficient lubrication within the encoder increases friction over time making it more difficult to turn the pump speed control knob. Excessive friction causes premature wear of the encoder shaft which can result in a stuck shaft. | Nonconforming Material/Component | Sorin Group sent an Urgent Field Safety Notice dated February 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:Before each use of the device during the set-up procedure, rotate all pump speed control knobs back and forth to check for smooth operation. If abnormal friction or any binding is detected, please do not use the system until it is repaired.If an affected product is currently in use with a patient and an increase of friction is detected please prepare for a replacement pump.Please complete and return the attached Confirmation Form by fax to 303-467-6502.Customers were instructed to ensure that the Urgent Field Safety Notice is distributed to all personnel in their organization who need to be aware of the notice. If the product was transferred to a third party, customers were asked to share the information with them as well as Sorin Group USA Customer Service at 1-800-850-2623.For questions regarding this Urgent Field Safety Notice, contact Sorin Group USA Customers Service at 1-800-650-2623.All affected customers and distributors were notified of the Field Correction by certified mail beginning on February 23, 2012. They will be contacted by Sorin Service Group to arrange an appointment to replace affected shaft encoders on site. |  | 39 | 39 | Worldwide Distribution | Terminated | November 27, 2012 | 278 | 1147 | Not_Computer
Z-1188-2012 | 61194 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aerolase Corporation. The products involved are the FriendlyLight YAG lasers, a Model Family first identified and reported. |  |  | June 07, 2011 | 2011 | Aerolase | It was discovered that customers were not supplied with a procedure and schedule for calibrating the FriendlyLight YAG lasers. | Radiation Control for Health and Safety Act | Aerolase issued a letter to customers, dated 11/30/2011, identifying the affected product and the reason for the recall. Aerolase is providing, at no charge, the Appendix to the Operating Manual for the LightPod lasers. This Appendix describes the procedure and frequency for checking whether the laser is within calibration, but it is not a procedure for customers to actually calibrate the laser themselves; any attempt to do so by any parties other than a company technician employed by Aerolase would void the warranty. Please contact us any time at 914- 345-8300 with any questions. |  | 176 | 176 | Worldwide Distribution | Open, Classified |  |  | 176 | Not_Computer
Z-1355-2012 | 61054 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Critical Care Bed Model 2031, 230 V , 50/60 Hz, Max Load = 225 Kg,, Stryker, Portage, MI |  |  | March 12, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Stryker Medical has determined that units manufactured between September 1, 2004 through November30, 2007, did not receive updated preventive maintenance information. (Failure to maintain the bed brakes including the brake plate kits, may result in reduced brake holding force. Reduced braking force may result in a patient fall or caregiver muscle strain, either of which may require medical int | Component change control | Stryker Medical sent an Urgent-Medical Device Correction letter dated March 12, 2012 to all consignees. The letter identified the affected products, problem, updated preventive maintenance information and requested the following immediate actions: 1) Locate the beds listed in this notice, 2) Ensure that the brakes are holding properly, 3) If the brakes are not holding, contact Stryker (1-800-327-0770, option 4) to order the maintenance parts, 3) Continue to periodically inspect brake functionality, as noted on enclosed preventive maintenance schedule/checklist, 4) Order replacement brake components, as needed, 5) Return the enclosed post card to confirm receipt of this notification, 6) If you have loaned or sold any of the beds listed in this letter, please, forward a copy of this notice to the new users and advise us of their new location, 6) If you have disposed of any of the beds and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers. For questions or concerns, please, contact Sandy Cliche at 269-389-6025. |  | 507 | 507 | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: DEDEDO GU, HAMILTON ON, SANTIAGO, SHATIN, NT , CAPITOL FEDERAL, ITAPEVI SP, ATHENS, OSAKA, SEOUL, MEXICO DF MX, MONTREUX, BALERNA, NORTH POINT, WARSZAWA, CANOVANAS, KOWLOON BAY, SINGAPORE, NEWBURY, and JOHANNESBURG. | Terminated | August 29, 2013 | 535 | 121982 | Not_Computer
Z-1349-2012 | 61143 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion Surestream CO2 Sampling Lines- |  |  | December 01, 2011 | 2011 | Oridion Medical 1987 Ltd. | Surestream CO2 Sampling Lines-Connector may cause breakage of the microswitch inside the CO2 inlet | Device Design | Oridion issued a Technical Note communication in March 2011 to field staff and end users globally. The Technical Note identified the affected product and provided instructions for user on the proper insertion of Surestream products. For questions contact your local distributor or Oridion at technicalsupport@oridion.com |  | 481 | 481 boxes (12025 untis) | USA Nationwide Distribution including the states of: AZ, CA, FL, HI, LA, MA, MO, NJ, NY, OH, OR, TX and WI. | Terminated | June 04, 2012 | 186 | 841 | Not_Computer
Z-1336-2012 | 61362 | No | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BacT/ALERT Blood Collection Adapter |  |  | March 05, 2012 | 2012 | bioMerieux, Inc. | The BacT/ALERT Blood Collection Adapter Cap is no longer compatible with Becton Dickinson (BK) Vacutainer Safety Lok Blood Collection Set. The incompatibility is the blood collection set needle backing out of the adapter cap that results in an exposed sheathed needle that may lead to an accidental needlestick. | Nonconforming Material/Component | bioMerieux sent an Urgent Product Correction Notice letter dated March 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions: 1. Ensure the letter and guidance on how to maintain control of the luer connector by securing it between the thumb and forefinger are distributed to all appropriate personnel within their organization. 2. Fill out and return the Acknowledgement Form by Fax to confirm receipt of the recall notice. For additional assistance or questions customers were instructed to contact their local bioMerieux Customer Service Representative at (800) 682-2666, option 3, then option 3 again. Industry customers should call 800- 634-7656, option 3. For questions regarding this recall call 314-731-8537. |  | 274800 | 274,800 boxes (32,862,600 caps) | Worldwide Distribution - USA (nationwide), Australia, Brazil, Canada, Chile, Columbia, Costa Rica, France, China, Indonesia, Korea, Mexico, Peru, Singapore, Ecuador, Dominican Republic, Philippines, Guatemala, Curacao, Thailand, United Arab Emirates, Andorra, Austria, Burkina Faso, Congo, Switzerland, Cameroon, Germany, Algeria, Estonia, Spain, France, Great Britain, French Guiana, Croatia, Israel, India, Italy, Kazakhstan, Lithuania, Latvia, Macedonia, Malta, New Caledonia, Netherlands, French Polynesia, Poland, Portugal, Reunion, Russia, Sweden, Slovenia, Chad, Turkey, South Africa, Saudi Arabia, Congo, CoteDIvoire, Denmark, Djibouti, Finland, Greece, Guadeloupe, Libyan Arab Jamahiriya, Jordan, Morocco, Martinique, Norway, Senegal, Togo, Yemen, Bangladesh, Georgia, Ghana, Uganda, and Qatar. | Terminated | September 04, 2014 | 913 | 274800 | Not_Computer
Z-1337-2012 | 61367 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | IntelliSpace Portal, M/N 881001, IntelliSpace Portal (ISP) is a multimodality (CT, NM, and MR) thin- client applications server that delivers full diagnostic viewing and clinical applications to the enterprise. IntelliSpace Portal (ISP) is a medical software system that allows multiple users to remotely access IntelliSpace Workspace Portal (ISP) from compatible computers on a network. |  |  | February 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A number of non-conformances were logged for IntelliSpace Portal running software version 4.0 during in-house random verification testing. When loading a study to Review, after failing to load it through Add to Running Application the Analysis application is not being closed and different (1st loaded study) patient appears in analysis (CCA). | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. Customers were asked to retain a copy with the equipment instruction for use. For any further information or support customers should contact their local Philips representative or local Philips Healthcare office. Customers in North America and Canada should contact the Customer Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 440-483-7000. |  | 90 | 90 units | Worldwide Distribution - USA including AL, AR, FL, GA, IO, IN, MI, NJ, NY, OH, OK, OR, PA, SC, TN, TX &amp; WA and the countries of Australia, Canada, China, Czech Republic, France, Germany, India, Indonesia, Israel, Italy, Poland, Spain, Sweden, Switzerland &amp; United Kingdom | Terminated | February 22, 2013 | 366 | 90 | Computer
Z-1340-2012 | 61373 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | RTVue CAM with Corneal Power Measurement, Cornea Power Kit (software) |  |  | February 21, 2012 | 2012 | Optovue Inc. | A few doctors in the US were inadvertently sent the Cornea Power software version kit for international users instead of the domestic version. | Employee error | Field Correction letters were sent to all doctors affected by the issue on February 23, 2012. The letter identified the affected product and the problem. The letter also stated that customers may continue to use the RTVue until the correction was made. Once the correction has been made, customers are to sign and return the attached acknowledgement form. If customers had questions, they are to contact Optovue at 1-866-344-8948 or via mail at 45531 Northport Loop W, Fremont, CA 94538. |  | 8 | 8 | Nationwide Distribution -- CA and WA only | Terminated | April 02, 2012 | 41 | 8 | Computer
Z-1338-2012 | 61391 | No | LKK |  |  | General Hospital | PMA | Medtronic, Model 8870, Application Software Card is part of the SynchroMedÂż II Drug Infusion System. The System includes the SynchroMedÂż II implantable drug infusion pump 8637-20, 8637-40, N'Vision clinical programmer 8840 and the Software application card 8870. |  |  | March 12, 2012 | 2012 | Medtronic Neuromodulation | Medtronic has confirmed that an algorithm used in the Model 8870 application card software has resultedin nine (9) occurrences of an incorrectly displayed Schedule to replace the pump by date. It is estimated that there are more than 140,000 SynchroMed II pump implants worldwide. A patient with a pump reaching EOS prior to replacement may experience the return of underlying symptoms and/orwit | Software design | Medtronic mailed an "Urgent Medical Device Correction " letter addressed to Dear Healthcare Professional on March 12, 2012. The letter described the problem, Background, Clinical Manifestations, Recommendations, and Additional Information.To ensure there is no interruption in therapy, pump replacements should be scheduled before the ERI alarm sounds and the pump reaches its 7-year end of service.Customers were advised to forward the Urgent Medical Device Correction letter to patients whose name appears on their Pump Replacement List.Reimbursement SupportFor assistance with prior authorization, contact Medtronic Reimbursement at the numbers below:Therapy for Severe Spasticity: (800) 292-2903Pain Therapies: (866) 962-9909For questions regarding the letter call 800-328-0810. |  | 6338 | 6338 | Nationwide Distribution | Open, Classified |  |  | 6338 | Computer
Z-1339-2012 | 61293 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Optiflux F160NRe Hemodialyzer |  |  | March 06, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Risk of internal blood leaks affecting the hemodialyzers. | Process control | Fresenius North America notified consignees on 3/6/12 by recall notification letter titled: Urgent FMCNA Optiflux F160NRe Hemodialyzer Recall, by certified mail with return receipt. The letter identified the affected product and described the reason for the recall. Customers were instructed to examine their stock immediately to determine whether they have any of the affected lot on hand. If customers have the affected lot, they have been instructed to immediately discontinue use and place all unused units in a secure area for return to Fresenius Medical Care North America. A Reply Response from is to be completed and returned. Customers have been instructed to contact their Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. Contact 1-800-323-5188. |  | 30576 | 30,576 units | Nationwide Distribution | Completed |  |  | 30576 | Not_Computer
Z-1351-2012 | 61440 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Standard Plus Implant 04.8mm Wide Neck |  |  | March 12, 2012 | 2012 | Straumann USA, LLC | Mislabeled: Vial lid label on some of the dental implant vials from this lot incorrectly state 04.8 SP 10 RN instead of 04.8 SP 10 WN | Packaging process control | Straumann sent a "FIELD CORRECTIVE ACTION NOTICE" dated March 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product and to return any affected product to the firm. Additionally, a Customer Confirmation Form was enclosed for customers to complete and return. Contact the firm at 978-747-2514 for questions regarding this notice. |  | 346 | 346 units | Nationwide Distribution. | Terminated | May 31, 2012 | 80 | 346 | Not_Computer
Z-1358-2012 | 61295 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex eXpert DC Intraoral Dental X-Ray System |  |  | December 29, 2011 | 2011 | Gendex Dental Systems | Improper hardware was used during installation of the eXpert DC Intraoral X-Ray system units causing drifting of the tubehead assembly. | Nonconforming Material/Component | Gendex Dental Systems sent a Urgent Medical Device Correction letter dated December 29, 2011, to both affected customers. The Customers were provided background information and details on how the nonconforming installations of the devices would be corrected. Gendex Dental Systems apologizes for any inconvenience caused by this issue, and request your corporation as we correct this problem. The customers were instructed to complete and return the Acknowledgement Form upon receipt of the recall notification letter and fax to (215) 997-5665.For further questions please call (215) 954-0365. |  | 52 | 52 | USA including the states of Iowa and Illinois. | Terminated | June 01, 2012 | 155 | 52 | Not_Computer
Z-1359-2012 | 61272 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch Verio IQ Blood Glucose Monitoring System Kit; |  |  | February 21, 2012 | 2012 | Lifescan Inc | An error was discovered in the OneTouch IQ meter software so that the meter turns itself off when a user attempts to view results in the "Results Log" when the log has 256 or a multiple of 256 items to display. | Software design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated February 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to 1. Identify and sequester all lots of OneTouchÂż Verio"IQ Meters they have in inventory.UPC No. 353885007702 LifeScan Part No. 022267 NDC No. 53885026701Note: Only OneTouchÂż Verio"IQ Meters are subject to this notice. Pleasecontinue to sell all other OneTouchÂż Brand products including OneTouchÂż Verio" TestStrips and all OneTouchÂż UltraÂż Brand Meters, Test Strips and Ancillary Products.2. Communicate this replacement program to their distribution centers and retail stores that have received OneTouchÂż Verio"IQ Meters. Request that they return only OneTouchÂż Verio"IQ Meters per their normal return procedures.3. Once you received all of the OneTouchÂż Verio"IQ Meters to be returned, call Inmar Corp. at 336-631-7744 for a Returned Goods Authorization (RGA) number and product return instructions.When product has been returned, a credit memo for an amount equivalent to the invoiced price of the returned product. No deductions will be allowed.If you have any questions about this action, please contact your LifeScan Sales Representative or call the LifeScan Customer Experience Line at: 1-800-600-7226, Option 1. Customers were advised that they can continue use of their affected product until an updated meter is provided free of charge.For questions regarding this recall call 408-942-5937. |  | 101882 | 101,841 units. 2-27-12 - this figure is an estimate and will be updated when more accurate number is received. | Worldwide Distribution - USA (nationwide) and Canada | Terminated | February 05, 2013 | 350 | 101882 | Computer
Z-1360-2012 | 61234 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | Removal of Literature/Brochures for Antalgic-Trak. "***What is Antalgic-Trak?***Articulating Non-Surgical Spinal Decompression****Antalgic-Trak By Spinetronics***Note: "Decompression" as used herein refers to unloading due to distraction and positioning.***Antalgic-Trak is an improved form of non-surgical spinal decompression*. |  |  | August 11, 2011 | 2011 | Spinetronics, LLC | On 08/11/2011 Spinetronics, LLC initiated a recall of Antalgic-Trak Product Brochures which contained language and/or claims that were not cleared under the original 510k, and represented a major change or modification in the intended use of the device. | Labeling False and Misleading | Customer notification letters were mailed by certified mail via USPS. The medical user/customer was instructed to either return or destroy any remaining recalled product literature. Customers were asked to complete and return a response form noting the number of available brochures. |  | 1600 | 1600 brochures | Worldwide Distribution, including Nationwide (USA) and the country of Canada. | Terminated | April 20, 2012 | 253 | 1600 | Not_Computer
Z-1364-2012 | 61488 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Verio |  |  | March 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During product monitoring, Siemens has discovered that the gradient cable connections on a few systems did not meet their specifications after installation. The cable connections overheated, leading to the emission of smoke and possible material damage. | Nonconforming Material/Component | The firm, Siemens Medical Solutions USA, Inc., sent a "Customer Safety Advisory Notice" including Update Instruction MR028/10/S dated March 2012 to affected customers via hand delivery by Siemens Service Engineer . The notice described the product, problem and actions to be taken. The customers were instructed to create a COBR image and to make a copy of the attached "Completion Protocol" fill it out and file in their corresponding binder.Siemens Service Engineer performed an all gradient cable connection visual inspection as well as torque check on the connecting screws at the sites and the factory prior to their release to the field. Siemens also issued a current Update Instruction MR024/11/S. If you have any questions regarding this report, please contact the Quality Engineer, Product Complaints &amp; Field Support at 610-219-4834; email: anastasia.mason@siemens.com or Director, Product Complaints &amp; Field Support at 610-448-4634 or email: josephine.mcbride@siemens.com. |  | 245 | 245 | Nationwide distribution to Washington, DC and the following States: AL, AK, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WI. Government Accounts include: BROOKE ARMY MEDICAL CENTER, 3851 ROGER BROOKE DR, FORT SAM HOUSTON, TX 78234; NATIONAL INSTITUTE OF HEALTH, 10 CENTER DRIVE, BETHESDA, MD 20892; NAVAL MEDICAL CENTER, 34800 BOB WILSON DR, SAN DIEGO, CA 92134; US AIR FORCE REGIONAL HOSPITAL, 5955 ZEAMER AVE, ANCHORAGE, AK 99506; VA MEDICAL CENTER, 3350 LA JOLLA VILLAGE DR, SAN DIEGO, CA 92161; VA MEDICAL CENTER, 830 CHALKSTONE AVE, PROVIDENCE, RI 02908; VA MEDICAL CENTER, 109 BEE ST, CHARLESTON, SC 29403; VA MEDICAL CENTER, 10000 BAY PINES BLVD, BAY PINES, FL 33744; VA MEDICAL CENTER, 800 BAILIE DR, AUGUSTA, GA 30901; VA MARTINEZ 612, 150 MUIR RD, MARTINEZ, CA 94553; VA MEDICAL CENTER, 1501 SAN PEDRO SE, ALBUQUERQUE, NM 87108; VA MEDICAL CENTER, 1670 CLAIRMONT RD, DECATUR, GA 30033; VA MEDICAL CENTER - ST LOUIS, 915 N GRAND AVE, ST LOUIS, MO 63106; and VA MEDICAL CENTER HARLINGEN, 2106 TREASURE HILLS BLVD, HARLINGEN, TX 78550. "Other" is National Institute of Health. | Terminated | April 03, 2014 | 738 | 245 | Not_Computer
Z-1357-2012 | 61333 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | The Polarus Plus Humeral Rod (part number HR-0824-S) is intended as an intramedullary fixation device for humeral fractures. It is sealed in an internal sterile blister pack within a box that is labeled in part: "Polarus Humeral Fixation System Size: 8mm X 240mm Polarus Plus Humeral Rod***Titanium***REF HR-0824-S***LOT 266220***RX only***Manufactured or distributed by: ACUMED 5885 NW Cornelius Pass Road Hillsboro, OR 97124 USA***" |  |  | March 07, 2012 | 2012 | Acumed LLC | The packaging contains the wrong product. It does not contain model HR-0824-S as the package indicates. | Labeling mix-ups | Acumed sent an Urgent Notice Device Recall letter dated March 5, 2012, to all affected consignees. A follow-up notification letter was sent to the domestic customers on March 9, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were instructed to further notify their consignees about the recall and to cease use of and return the affected product. Consignees were requested to respond using the enclosed response form. The form should be faxed to 503-520-9618. For questions regarding this recall call 503-627-9957, ext. 1302. |  | 12 | 12 rods | Worldwide Distribution - USA (nationwide) and the countries of Italy, Ireland, and Great Britain. | Terminated | July 25, 2012 | 140 | 12 | Not_Computer
Z-1361-2012 | 61445 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Systems. Brilliance 64 Model# 728231, Ingenuity CT Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH. |  |  | March 12, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare is performing an upgrade to the field for issues related to Philips Brilliance 64 and Ingenuity CT scanner having Software Version 3.5.1. When the user plans a scan with tilt applied, the scan increment changes from the user defined values (noncontiguous) to a contiguous scan. | Software design | . |  | 440 | 440 units | Worldwide Distribution-USA (nationwide) including the states of AK, AR, AZ, CA, CO, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MT, ND, NE, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA and WV, and the countries of Australia, Bahamas, Bangladesh, Belgium, Bolivia, Brazil, Chile, China, Colombia, Croatia, Denmark, Finland, France, Georgia, Germany, Great Britain, India, Indonesia, Iraq, Israel, Italy, Japan, Kazakhstan, Korea, Laos, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Poland, Russia, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Terminated | June 18, 2013 | 463 | 440 | Computer
Z-1365-2012 | 61392 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System, Version 3.10.00. |  |  | March 05, 2012 | 2012 | Computerized Medical Systems Inc | An unnecessary dose refresh can occur in certain situations which causes inconsistencies in the Beam Doses and the Total Dose. | Software design | An "IMPORTANT SAFETY NOTICE" was issued between 3/5-9/12 via e-mail, FAX, or traditional mail. The Notice described the correct functioning of the system, the problem, when and why the issue occurs, the clinical impact, and the workaround. Patch 3.10.01 is now available. Customers can download their patch at the link provided in the Notice. The Notice also provided a list of offices where customers can contact if they have questions or further support is needed. |  | 44 | 44 | Worldwide Distribution -- USA, including the states of OH, IN, WI, CA, FL, IL, LA, NJ, MA, NC, and MO and the countries of Canada, Australia, Belgium, China, France, Germany, Japan, The Netherlands, New Zealand, Portugal, Switzerland, Taiwan, and the United Kingdom. | Terminated | April 01, 2014 | 757 | 44 | Not_Computer
Z-1362-2012 | 61327 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo-i ventilator system |  |  | December 22, 2011 | 2011 | Maquet Inc. | The Servo-i and Servo-s ventilator systems equipped with specific revision os of the AC/DC convertor can result in electric shock during servicing. | Device Design | Maquet sent a Product Recall Urgent-Medical Device Corrective Action letter dated December 22, 2011 to all affected customers. The letter identified the affected products, patient risk, product issue and actions to be taken. Customers were instructed to complete and return the attached Customer Fax Response form as soon as possible. Maquet Field Service Technicians will update all affected products at all the concerned customer sites. Customers are advised untill Maquet has updated all the affected products (all of which are still under warranty), inform all the concerned biomedical staff about the content of this Device Correction Letter. For technical support call 1-888-627-8383. |  | 3136 | 647 US; 2462 Foreign; 27 Spare Parts Foreign | USA Nationwide Distribution | Terminated | October 22, 2013 | 670 | 3186 | Not_Computer
Z-1371-2012 | 52174 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | The EM 2400 is a Pharmacy Compounding System consisting of three sub-systems for compounding solutions for intravascular administration. |  |  | June 03, 2008 | 2008 | Baxa Corporation | Remote possibility of an unintended ingredient delivery as a result of a damaged valve actuator on the EM 2400. | Device Design | The firm, Baxa, sent a "Exacta-Mix 2400 Compounder Safety Alert" dated June 3, 2008 to all customers. The alert describes the product, problem and actions to be taken. The customers were instructed to not use the main module if any of the 24 valve actuators are damaged; do not infuse any bags if damage is noticed after they have been pumped; do not try to repair or adjust a damaged or bent valve actuator; immediately alert all compounder users of this issue to ensure that when the device is used, the valve actuators are always inspected for damage;" inspect all valve acutators before setup and after tear-down to ensure they are not damaged or bent. Do not use the compounder if the valve actuators are damaged or bent."; and if a damaged or bent valve is noticed after using the EM2400, immediately hold all pumped bags and contact Baxa Technical Support. If you need additional assistance, please contact Baxa Technical Support (800-678-2292). Technical Support is available 24 hours a day, 7 days a week. |  | 1000 | 1000 units (approximately) | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, China, New Zealand, and United Kingdom. | Terminated | April 10, 2012 | 1407 | 1000 | Not_Computer
Z-1367-2012 | 61359 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Extended Brilliance Workstation Computed Tomography X-Ray Systems, M/N 728260, Distributed by: Philips Medical System, Cleveland, OH |  |  | January 26, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | TumorLOC will save the incorrect contour when the user saves the contour while in the edit mode, and the software does not force the user to exit the edit mode before saving. In TumorLOC, when saving while a contour is in edit mode, that contour will shift to the slice corresponding to the most recently drawn contour of the last organ in the Organ list. | Software design (manufacturing process) | Philips sent an Urgent - Medical Device Correction letters to all affected customers. The letter identified the product, what the problem is and under what circumstances it can occur, important information for the continued safe and proper use of the affected equipment, and actions to be taken to minimize the effect of the problem. Philips instructed their customers to put a copy of the letter with their "Instructions for Use" manual. For question contact your local Philips representative or local philips healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site Id or follow the prompts). |  | 49 | 49 units | Worldwide Distribution - USA (nationwide) including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY. Product was also shipped to the following countries: Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Bulgaria, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakstan, Kenya, Kuwait, Libya, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Russian Fed, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom &amp; Venezuela. | Terminated | February 22, 2013 | 393 | 1110 | Computer
Z-1335-2012 | 61499 | No | RFK |  |  |  |  | U-SPECT-II/CT and VECTor/CT cabinet x-ray products. |  |  | April 15, 2012 | 2012 | Milabs Bv | These products have doors which are not equipped with at least one safety interlock which physically disconnects power to the high voltage power supply for the x-ray source. Additionally, these products do not have a certification label and the identification labels do not include the month and year of manufacture. | Radiation Control for Health and Safety Act | A notification will be sent to all affected purchasers by April 15, 2012. The manufacturer will bring the products into compliance with the applicable performance standards within 60 days. Specifically at least one door safety interlock will be replaced per door and revised labels will be produced and applied to the system. These corrections will be made in accordance with MILabs instructions by a MILabs installation and service engineer. |  | 2 | 2 | Nationwide Distribution -- Indiana and North Carolina. | Terminated | August 29, 2012 | 136 | 2 | Not_Computer
Z-1370-2012 | 61316 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Hologics SecurView DX workstation, software version 7.3.0 |  |  | February 28, 2012 | 2012 | Hologic, Inc. | SecurView DX with software version 7.3.0 defects when used to view mammography tomosynthesis images | Software design | Hologic Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides instructions on how to re-review the stored images and correct any annotation defect. A SecurView DX 7.3.0 Confirmation of Notification form was enclosed for customers to complete and return to the firm via fax at 866-652-8674. Contact Hologic's Help Desk at 1-877-371-4372 for questions regarding this notice. |  | 7 | 7 units | Worldwide Distribution-Distributed in the state of Texas and the countries Belgium, Brazil, and Canada. | Terminated | May 07, 2012 | 69 | 7 | Computer
Z-1382-2012 | 60152 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500, 500V. |  |  | February 03, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Replacement of the SBC batteries (system batteries) should be managed as "routine maintenance" instead of as an in-house Service Procedure because some systems may be relying on batteries beyond their useful life. | Labeling design | On 2/3/12 all consignees received an URGENT MEDICAL DEVICE RECALL letter via Federal Express. The letter described the reason for the recall, identified the affected product population, what to do in the event of a failure, the potential hazard, and correction. Terumo CVS will update the Operator's Manual for the CDI 500 with instructions to replace the battery within its useful life. Users are to review the notice and ensure that all users are made aware of the issue. In addition, the Customer Response Form should be completed and returned. Users can contact their local Terumo Cardiovascular Systems representative or call 1-800-521-2818 with questions or concerns, to check the last time system batteries were replaced, or to schedule maintenance for systems approaching 10 years. |  | 3829 | 3,829 total units distributed | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, WV, WA, and WI and the countries of Japan, Germany, Barbados, Saudi Arabia, Columbia, Uruguay, Ecuador, Trinidad, Costa Rica, Paraguay, Panama, Belgium, Jordan, Libya, Mexico, Brazil, Honduras, Canada, Korea, Dominica Republic, El Salvador, Guatemala, Vietnam, Chile, Hong Kong, United Arab Emirates, Singapore, Australia, Malaysia, Taiwan, Philippines, Mexico, Thailand, Vietnam, and South Africa. | Terminated | August 29, 2013 | 573 | 3829 | Not_Computer
Z-1377-2012 | 61084 | No | DFC | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | Wako Autokit Lp(a) Calibrator, model code# 998-41491, packaged in 1 ml bottles, 2 bottles/kit, and labeled in part ***Manufactured by: Wako Pure Chemical Industries, Ltd. 1-2, Doshomachi 3-Chome, Chuo-Ku Osaka 540-8605, Japan***Distributed by: Wako Chemicals USA, Inc. /Wako Diagnostics |  |  | January 30, 2012 | 2012 | Wako Chemicals USA, Inc. | In-vitro diagnostic kit calibrator is defective and could impact patient profile for cardiovascular disease. | Device Design | Wako Diagnostics notified its customers by telephone on January 26, 2012 and sent an Urgent Medical Device Recall letter dated January 30, 2012 to inform them of the product recall. The notifications identified the affected product, problem and actions to be taken. Consignees were requested to: 1) Complete and sign the attached Return Materials Authorization (RMA) Form to receive replacement material and 2) Package the recalled Product to be returned and send to the address listed on the RMA form. For questions call Technical Support Team at 1-877-714-1924. |  | 10 | 10 kits | The product had been sold to 2 customers in the US. | Terminated | May 02, 2012 | 93 | 10 | Not_Computer
Z-1375-2012 | 61118 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Assay Protocol File (APF) Software versions for UniCel DxC 600i Systems. |  |  | September 11, 2008 | 2008 | Beckman Coulter Inc. | The recall was initiated because Beckman confirmed a customer report that a timing conflict can occur while aspirating samples for the Access Total BhCG assay. | Software design | Beckman Coulter sent a "PRODUCT CORRECTIVE ACTION" letter dated September 24, 2008 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers were instructed to complete and return the enclosed response form. Customers with questions regarding this notification were instructed to contact Technical Support at (800) 854-3633 or their local Beckman Coulter representative. |  | 2948 | 2948 units total (2762 units in the US) | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | June 26, 2012 | 1384 | 2948 | Not_Computer
Z-1376-2012 | 61436 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OneTouch(R) Ping(R) Insulin Pump glucose management system. The device is intended to deliver insulin therapy. |  |  | March 01, 2012 | 2012 | Animas Corporation | There is an error in the pump software where the OneTouch(R) Ping(R) insulin pump did not allow the user to save changes to the time and date on the pump if those changes were made on February 29, 2012 (Leap Day). Users who did not attempt to change the time or date on their pump on February 29, 2012 are not affected by this issue. | Software design | The firm initiated their recall on March 1, 2012 and sent Urgent Notification letters dated March 6, 2012 to all consignees by UPS with delivery confirmation. Customers were notified of the issue and the appropriate steps to take in order to mitigate the issue. A Web Letter for users was also posted to the Animas web site. Questions were directed to (855) 230-7577. |  | 91239 | 91,239 | Worldwide Distribution, (USA) including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and the country of Canada and the territory of Puerto Rico. | Terminated | December 24, 2013 | 663 | 91239 | Computer
Z-1389-2012 | 52024 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Pharmacy compounding system . It is a twenty-four source compounder for multi-source fluid mixing of both macro and micro ingredient down to 0.2 ml quantities. |  |  | March 23, 2009 | 2009 | Baxa Corporation | Software issue with compounding pump resulting in over delivery of component ingredients during routine operations. | Software design | Baxa sent a Safety Alert letter dated March 23, 2009 to all affected consignees. The letter identified the affected product, problem and recommended actions with specific instructions related to the issue. Additionally, the firm will be providing a warning label for the affected product as soon as possible and is currently working on a product upgrade to correct the issue. For assistance contact Baxa Technical Support at 800-678-2292. The firm is also providing a 24 hour a day, 7 days a week Technical Support for the affected product. |  | 1322 | 1322 units | Worldwide Distribution - USA Nationwide including Canada. | Terminated | April 10, 2012 | 1114 | 1322 | Computer
Z-1386-2012 | 61405 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Nano Blood Glucose Meter |  |  | March 12, 2012 | 2012 | Roche Diagnostics Operations, Inc. | Accu-Chek Aviva Nano Blood Glucose meters that were manufactured for the purpose of validating a manufacturing process for product to be distributed outside the US were distributed in the US through eBay by an unauthorized third party. These meters were not intended for distribution, were not verified to meet final product specifications and were not distributed to customers with the necessary pr | Employee error | Roche sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to immediately return their ACCU-CHEK Aviva Nano Meter to the firm in an enclosed pre-paid envelope via USPS. Contact Roche Product Removal Services at 1-888-345-5352 between 8:00 am and 5:00 pm for questions regarding this notice. |  | 82 | 82 meter distributed on Ebay | Nationwide Distribution-USA (nationwide) including the states of AZ, NY, RI, TN, NJ, IN, TX, AL, NV, CA, OR, KY, IL, AR, NM, NE, MD, MO, GA, and SD. | Terminated | November 06, 2012 | 239 | 82 | Not_Computer
Z-1390-2012 | 61393 | No | JFT | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Cortisol Assay |  |  | March 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Cortisol Assay reagent kit lots do not meet onboard stability and calibration interval claims | Component design/selection | The firm, Siemens Healthcare, sent an "Urgent Field Safety Notice" via e-mail to all affected customers on March 6, 2012 in the United States and International. The Urgent Field Safety Notice describe the product, problem, and actions to be taken. The notice instructed the customers to perform daily calibration on all Cortisol reagent kits to maintain consistently accurate quality control and patient values; to complete and return the confirmation fax-back form 'Field Correction Effectiveness Check' via fax to the Technical Solutions Center at (302) 631-7597 and forward this notification to whomever they many have distributed this product.If you have any questions, contact Quality Engineer 5- Quality Systems and Compliance at 508-660-8540 or www.siemens.com/diagnostics. |  | 2410 | Total: 40023. 50 Test Kits = 21734;250 Test Kits= 15879;250 REF Test Kits= 2410 | Worldwide distributed: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, AUSTRALIA, BANGLADESH, BD, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, EGYPT, GERMANY, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, JAPAN, KOREA, MALAYSIA, MEXICO, NEW ZEALAND, PAKISTAN, PHILIPPINES, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, TAIWAN, THAILAND, URUGUAY, VENEZUELA, and VIETNAM. | Terminated | June 26, 2015 | 1207 | 2410 | Not_Computer
Z-1384-2012 | 61178 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SenDx Medical Inc |  |  | January 30, 2012 | 2012 | Sendx Medical Inc | The product can cause the analyzer to report values outside the published performance specifications for pH, pCO2, cNa+, cK+, cCa2+ and cCl-. | Nonconforming Material/Component | SenDx Medical sent Customers notification letters dated January 30th, 2012, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. We have recently become aware of a possible condition in a limited number of ABLBO FLEX analyzers configured with a specific Control PCe lerinted Circuit Board), which can cause the analyzer to report values outside the published performance specifications for pH, pCOz, cNa*, cK*, cCa2* and cCl-.Customers were instructed to immediately implement Alternative 1 or 2 until the signal data PCB is replaced. Alterative I:Stop the use of all affected analyzers until corrective actions can be performed on these analyzers.Alternative 2:Verify proper function of the affected analyzers immediately following everypatient sample measurement by measuring a manual QC sample.Customers were asked to complete the fax form acknowledging receipt of the notice and provided a contact for Technical Support at 1-800-736-0600 opt.4. |  | 863 | 819 analyzers and 44 signal data PCBs | One US consignee in state of OH. Other consignees are international such as Australia, France, Netherlands, Denmark, China, United Kingdom, Turkey, Switzerland, Japan, India, Poland, Spain, and Germany. Distribution list submitted with supporting doc.s | Terminated | July 13, 2012 | 165 | 863 | Not_Computer
Z-1383-2012 | 61483 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Luminos dRF solid state x-ray imager |  |  | February 16, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, Siemens became aware of a potential issue with the AXIOM Luminos dRF. Firm became aware of a potential issue with torque wrenches - during a regular scheduled calibration of the wrench, it was discovered that the wrench was out of the specification. The wrench was over-torquing approximately 21% for low-end readings (50 ftlbs) and 9% for high-end readings (250 | Nonconforming Material/Component | Siemens initiated a recall/correction on February 16, 2012 by issuing Update Instruction XP056/11/S to the affected customers. All screws on the unit support on AXIOM Luminos dRF units will be replaced by a Siemens service engineer.For information regarding this recall call 610-219-4834. |  | 2 | 2 | Nationwide Distribution including California | Terminated | November 06, 2014 | 994 | 2 | Not_Computer
Z-1385-2012 | 61364 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | VariSeed 7.1, software application, part number H62, manufactured by Varian Medical Systems. VariSeed 7.1 is a computer based software application for planning and evaluating prostate brachytherapy procedures. |  |  | January 25, 2012 | 2012 | Varian Medical Systems, Inc. | Brachytherapy device requires the user to enter radioactive source activity, but software may specify unintended units for activity and lead to overdose in patient treatment. | Labeling design | Varian Medical Systems notified direct account healthcare practitioners/ medical physicists with "Urgent Medical Device Correction/Urgent Field Safety Notice" letters by mail on 01/19/2012. Varian's advisory described how to mitigate user error and the firm's intentions to revise Instructions for Use. Technical Support phone numbers were provided for US and international customers. |  | 1377 | 1377 units | Worldwide Distribution | Terminated | November 07, 2012 | 287 | 1377 | Computer
Z-1392-2012 | 61395 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | BioCare Medical Vina Green Chromogen Kit; 25 mL, 100ml IVD. Kits include: Vina Green Chromogen 1. 0 mL; Vina Green Buffer, 25 mL and Vina Green Dropper Bottle. BioCare Medical, Concord, CA 94520. |  |  | March 13, 2012 | 2012 | Biocare Medical Llc | Reports of sporadic failure of specific lots of Vina Green Chromogen Kit, either through crystal formation that obscures positive staining, color not right, or a complete fading of staining by the chromogen. | Nonconforming Material/Component | "Medical Device Recall Notification" letters were sent on March 13, 2012 to all direct account. Letters informed customers of the reason for recall and the specific lots and catalog numbers involved, along with the reason for recall, how to identify the recalled product, and the actions to be taken. Customers were to immediately discontinue use of the affected product lot(s). They should dispose of the product according to national, state, and local laws and regulations. The enclosed Return Response Form should be completed and returned. Customers will be credited upon receipt of the Response Form. |  | 74 | 74 kits | Worldwide Distribution -- USA, including the states of WI, KY, CA, GA, IN, MA, IL, MD, NC, LA. AZ, TX, NC, NJ, MO, NY, PA, OH, DC, MI, WA, and CO, and the countries of Canada, Sweden, Germany, Italy, and Switzerland. | Terminated | September 20, 2012 | 191 | 74 | Not_Computer
Z-1391-2012 | 61347 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink(TM) Informatics Systems Software Version 5.0 and Software Version 5.0 Service Packs 1 through 4 used with the Dimension Vista(R) System. |  |  | December 02, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | QC results may be unexpectedly associated with a Pending Control. When this situation occurs, Pending QC Controls maybe be created or reused. | Device Design | Siemens Healthcare Diagnostic Inc, sent a Urgent Field Safety Notice dated December 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were asked to follow these instructions. The first option is Extending the "Minimum Inactive Days" to 999 days. The second option is to proactively monitor the system for unexpected Pending QC Controls.Long Term Action Plan: A corrective and preventive action (CAPA) investigation has been issued. Currently firm is actively investigating.Further questions please call (302) 631-0516. |  | 787 | 787 | Worldwide Distribution -- USA (nationwide) and the countries of CANADA, FRANCE, GERMANY, SPAIN, PORTUGAL, ITALY, JAPAN, NORWAY, AUSTRALIA, AUSTRIA, NEW ZEALAND, BELGIUM, SWITZERLAND, NETHERLANDS, MALAYASIA, SLOVENIA, SOUTH KOREA, DENMARK, SAUDI ARABIA, SLOVAKIA, PUERTO RICO, INDONESIA and TAIWAN. | Terminated | March 21, 2014 | 840 | 787 | Not_Computer
Z-1387-2012 | 61249 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Elecsys Troponin I Immunoassay |  |  | March 12, 2012 | 2012 | Roche Diagnostics Operations, Inc. | An investigation has confirmed a considerably reduced recovery of Li-heparin plasma samples for Troponin I lot 163176 and Troponin I STAT lot 163177. In the worst case, Li-Heparin sample recovery may be as low as 50% of the serum recovery. Correct serum recovery has been confirmed for the affected lots.The issue is caused by the raw material lot (Poly-L-Lysin) instability. Poly-L-Lysin is the | Nonconforming Material/Component | The firm, Roche Diagnostics Corporation, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 12, 2012 to all customers that received the affected lots of Elecsys Troponin I or Elecsys Troponin I STAT. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the affected product; discard the affected product from their inventory according to their site's local regulations; if their facility has distributed the affected product to other sites, ensure that this letter is provided to those sites; complete and return the attached fax form via fax to 1-888-912-8457 and file this letter for future references.If you have any questions about the information contained in this letter, please contact Roche Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-428-2336. |  | 739 | 739 kits | Worldwide distribution.Argentina, Kenya, Austria, South Korea, Belgium, Kuwait, Bahrain, Morocco, Brazil, Mexico, Switzerland, Nigeria, Chile, New Zealand, Columbia, Philippines, Germany, Pakistan, Algeria, Poland, Egypt, Qatar, Spain, Tunisia, France, Turkey, United Kingdom, Taiwan, India, Ukraine, Iran, Venezuela, Italy, South Africa, Jordan, and Armenia. | Terminated | February 06, 2013 | 331 | 2027 | Not_Computer
Z-1394-2012 | 61505 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, IMAGER II Angiographic Catheter, Sterilized with ethylene oxide gas. Made in Ireland, Manufactured for Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537. |  |  | March 15, 2012 | 2012 | Boston Scientific Corporation | Boston Scientific Corporation is initiating a Medical Device Recall of certain lots/batches of the Imager" II Catheter (Urology) and the Imager" II Angiographic Catheter . An investigationhas found that, during shipping, some units in the 5-pack box may experience a physical interaction between the carton and the area onthe pouch where the hub is located. This interaction was found to cause pi | Package design/selection | Boston Scientific sent an " Urgent Medical Device Recall" letter dated March 14. 2012. the letter was addressed to "Dear Risk/Materials Manager. The letter described the problem and the product involved in the recall. Advised consignees to cease the distribution of any remaining unused product affected. Advised distributors that the recall notification should be forwarded to their customers. Requested consignees to complete and return the Reply Verification tracking Form enclosed by fax to 508-683-5578.For questions regarding this recall 763-494-1700.Recall Expanded 4/24/12:Imager II Angiographic customers:Eight total number of customers not notified initially (3 USA and 5 APA customers were sent the same initial letter with a a revised send date of April 18, 2012, 329 total number of US Customers notified initially with an incomplete product list were sent a revised notification letter. Urology: Imager II Torque customers:1 US Customer was not notified initially. This customer was sent the same initial letter with a revised send date of April 18, 2012.Fifteen US Customers notified with incomplete product list. These customers were sent a revised notification letter. |  | 406450 | 406,450 (416,555 USA, 389,895 OUS) | Worldwide Distribution - USA (nationwide) and the countries of Arab Emir, Azerbaijan, Austria, Australia, Azerbaijan, Belgium, Bulgaria, Bahrain, Bermuda, Brazil, White Russia, Canada, Switzerland,Chile, China, Colombia, Cyprus, Czech Republic, Germany, Denmark, Ecuador, Estonia, Egypt, Spain, Finland,France, Great Britain, Greece, Guatemala, Hong Kong, Honduras, Croatia, Hungary, Ireland, Israel, Italy,Jordan, South Korea, Kuwait, Lebanon, Lithuania, Morocco, Malta, Mexico, Malaysia, Netherlands, Norway,New Zealand, Oman, Panama, Peru, Pakistan, Poland, Portugal, Qatar, Romania, Serbia, Russian Fed., SaudiArabia, Sweden, Singapore, Slovakia, El Salvador, Syria, Thailand, Tunisia, Turkey, Trinidad, Tobago, Taiwan,Ukraine, Vietnam, South Africa, Dutch Antilles, Sri Lanka, Libya, and Yemen | Terminated | April 14, 2014 | 760 | 417790 | Not_Computer
Z-1395-2012 | 61408 | No | LPC |  |  | Cardiovascular | PMA | 12Fr SLS II Laser Sheath. The laser sheath is intended for use as an adjunct to conventional lead extraction tools in patients suitable for transvenous removal of chronically implanted pacing or defibrillator leads constructed with silicone or polyurethane outer insulation. |  |  | March 13, 2012 | 2012 | Spectranetics Corp. | Isolated lots of the 12Fr SLS II Laser Sheath can have a split in the outer jacket. | Nonconforming Material/Component | Spectranetics sent Recall Notices to all customers on March 14, 2012. The product issue was described and requested actions were included. Questions were directed to Customer Service at 1-800-231-0978 (US). |  | 615 | 615 | Worldwide Distribution -- USA , and the countries of Australia, Belgium, Canada, Czech Republic, France, Germany, Italy, Ireland, Japan, Netherlands, Russia, Sweden, Switzerland, United Kingdom and Venezuela. | Terminated | May 02, 2012 | 50 | 615 | Not_Computer
Z-1396-2012 | 61434 | No | KSR | Empty container for the collection and processing of blood and blood components. | Hematology | Hematology | 510(k) | DeRoyal (R) Umbilicup, REF 72-8000, Rx Only, STERILE EO, Manufactured by DeRoyal, Powell, TN 37849 |  |  | January 20, 2012 | 2012 | DeRoyal Industries Inc | Device's needle may become dislodged from the cup during shipment or during use. | Nonconforming Material/Component | DeRoyal Industries sent Recall communication by letter to all affected consignees of the Non-Sterile Bulk and Single Sterile Units on January 20, 2012 and of the Kits and Trays on Febraury 6, 2012. The letters identified the affected products, problem and actions to be followed. Customers were instrructed to complete the enclosed Notice of Destruction form and return this form by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter. DeRoyal will use your response on the Notice of Destruction form to send replacements as soon as they become available. Distributors were instructed to forward this recall to any of their end users that may have affected product.If you have questions or need assistance with the recall, please call 865.362.1020 or Robby Lockett at 865.362.1034. |  | 2472 | 2472 units | Worldwide Distribution - USA Nationwide and the countries of: Canada, Dominican Republic, South Africa, and the United Arab Emirates. | Terminated | July 11, 2013 | 538 | 9434 | Not_Computer
Z-1420-2012 | 52120 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa EVA TPN bags for administering parenteral nutrition solutions. Sold under three brand names: Exacta Mix EVA Bags, Exacta Mix Dual-Chamber EVA Halobag, and Baxa Calibration Bags. |  |  | May 08, 2009 | 2009 | Baxa Corporation | Spike port flor EVA TPN Bags contains low levels of DEHP which could leach out. | Labeling False and Misleading | Baxa sent a Customer Notice letter dated May 8, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Current Baxa product is mislabeled as non-DEHP but this does not present a risk to patients. Baxa is committed to providing products that are non-DEHP. We are working with our supplier to replace the spike port with a non-DEHP as soon as possible. In the meantime, product labeling will be updated to reflect this finding.For further questions please call (303) 690-4204. |  | 17374700 | 347,494 cases/ 50 bags each | Worldwide Distribution -- USA (nationwide) | Terminated | April 13, 2012 | 1071 | 17374700 | Not_Computer
Z-1401-2012 | 61446 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Non-DEHP Microbore Administration Set with Non-Vented Spike Packaged Assembly |  |  | March 20, 2012 | 2012 | MOOG Medical Devices Group | Ambulatory administration sets used with the Curlin Ambulatory Pumps including the 4000 CMS Ambulatory Pump, 6000 CMS Ambulatory Pump and the Pain Smart IOD Ambulatory Pump may leak during use. | Nonconforming Material/Component | Moog Medical Device Group sent a Medical Device Recall Notification letter dated March 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product from their inventory, contact the distributor or Moog Customer Service at 1-800-970-2337, prompt #7, then return product to distributor or Moog.For additional questions regarding this recall call Moog's Customer Service Department at (800) 970-2337. |  | 11628 | 11,628 | Worldwide Distribution - USA including AZ,,, CA,,, FL,,, GA,,, KS,,, IL,,, MA,,, MD,,, MI,,, OH,,, PA,,, TX,,, VA,,, WI,,, and the countries of Germany,,, Italy,,, Morocco,,, New Zealand,,, and the United Arab Emirates,,, | Terminated | July 26, 2012 | 128 | 43176 | Not_Computer
Z-1400-2012 | 61370 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | TeleSentry, Multi-parameter Ambulatory Telemetry, Model #TS01. Dist. By: ScottCare Corporation, Cleveland, OH. |  |  | January 19, 2012 | 2012 | Scottcare Corporation | Certain TeleSentry Devices contain a Bad SD Card. When the SD Card goes bad, the TeleSentry Device cannot read or write data to the SD Card. The TeleSentry Device then stops functioning. | Component design/selection | ScottCare sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 19, 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to immediately discontinue use and quarantine any affected products. A Recall Acknowledgement Form was enclosed for customers to complete and return via fax to (216) 257-6129. Contact the firm at (216) 362-0550 ext. 113 for questions regarding this recall. |  | 513 | 513 units | Worldwide Distribution-USA (nationwide) including the states of CA, DE, FL, MI, NJ, NV, NY, OH, PA and TX, and the country of Canada. | Terminated | September 27, 2013 | 617 | 513 | Not_Computer
Z-1432-2012 | 53049 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Ammonia (AMM) Reagent |  |  | March 26, 2009 | 2009 | Beckman Coulter Inc. | The Ammonia Reagent lot M808278 may fail calibration or cause a low shift in Quality Control recovery on some SYNCHRON Systems. | Unknown/Undetermined by firm | Beckman Coulter mailed a Product Corrective Action (PCA) Letter on April 13, 2009 dated April 10, 2009 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers to discontinue use and discard appropriately the lot number(s) of the affected product and contact their Customer Service Representative at (800) 526-3821 (US), or (800)463-7828 for reagent replacement. Customers were instructed to complete and return the enclosed response form within 10 days so that Beckman Coulter may maintain records.Customers were told to call the Customer Support Center (Hotline) at (800) 854-3633 with any questions regarding the PCA letter. |  | 3388 | 3,388 kits | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 18, 2012 | 1119 | 3388 | Not_Computer
Z-1430-2012 | 60661 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Double foot switch, type KF 2 1S/1 S-MED-AP, manufactured by the company Seute Schaltgerate. GmbH &amp; Co. KG for Ziehm Imaging GmbH Germany, and used with Ziehm Imaging GmbH Mobile C-arm as a component thereof. Two versions were made: 2 pedal version KF 2 1S/1 S-MED-AP foot switch and 4 pedal foot switch version KF 2 1S/1 S-MED-AP This type of foot switch includes a pair of the dual foot switches assembled on a single metal bracket. The following labels may appear on the foot switch: |  |  | April 29, 2011 | 2011 | Ziehm Imaging Inc | On 04/29/2011, Ziehm Imaging, Inc., Orlando, FL recalled the double foot switch, type KF 2 1S/1S-MED-AP, a component used with Ziehm Imaging GmbH Mobile C-arm due to the potential of unintentional continued activation of fluoroscopic X-ray as the foot switch may not become inactive when the pedal is released. | Component design/selection | Ziehm Imaging, Inc., Orlando, FL sent an Urgent Recall Notice letter dated April 11, 2011, that included a return receipt request, envelope, and a reply questionnaire. to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. They also called consignees. After contact, Ziehm Imaging, Inc. shipped a replacement foot switch with instructions for the customer/end user on how to install the replacement foot switch. If no signed confirmation was received, Ziehm Imaging, Inc. visited the consignee and replaced the foot switch. Customers were asked to fax the completed response form to 866-839-7893. |  | 49 | 49 units | Nationwide Distribution including Florida, South Carolina, Georgia, Alabama, Virginia, Ohio, Massachusetts, New York, Pennsylvania, Kentucky, Connecticut, Rhode Island, Indiana, Illinois, Nebraska, Minnesota, Arizona, California, Oregon, and Washington. | Terminated | April 18, 2012 | 355 | 49 | Not_Computer
Z-1429-2012 | 61455 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Norco Folding Shower Bench; |  |  | March 20, 2012 | 2012 | North Coast Medical Inc | Due to a consumer complaint, North Coast Medical determined that due to a design flaw, there was not enough plastic support material to retain bracket bolts when subjected to a force greater than 150 lbs. | Device Design | North Coast Medical, Inc. sent a Recall Notification Letter dated March 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and discard the affected product. Complete the enclosed Form and return the form even if they no longer have the bath bench. Return the form by fax to 408-775-1430. Customers should contact the Quality Department if they have any questions, concerns or need any assistance following the instructions.For questions regarding this recall call 800-821-9319, ext 121. |  | 370 | 370 benches | Nationwide Distribution | Terminated | July 20, 2012 | 122 | 370 | Not_Computer
Z-1428-2012 | 61352 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | PantaÂż Arthrodesis Nail Support device, 510 (k) #K050882. One (1) unit of product is kitted within an instrumentation tray. The product is available in six (6) sizes. All nails are color-coded for easy size identification. Bony fixation is achieved using two (2) tibial screws, two (2) calcaneal screws and one (optional) talar screw. The nail, cross-locking screws and end cap are manufactured from a titanium alloy: TI-6AI-4V ELI ISO 5832-3, ASTM F136. The PANTA Arthrodesis Nail is intended for use in tibiotalocalcaneal arthrodesis and treatment of trauma to the hindfoot and distal tibia. |  |  | May 19, 2011 | 2011 | Integra LifeSciences Corp. | The firm received multiple complaints from their user customers (surgeons) who reported encountering difficulty inserting a calcaneal or tibial screw through the holes of the PANTA nail. | Use error | Integra issued a Field Safety Notice with an attached Acknowledgement and return Form to all Integra Sales Representatives and Customers via Fed Ex and E-mail on May 19, 2011. The Field Safety Notice informs the customers of the firm's receipt of information concerning several reported incidents that lead to the misalignment of the calcaneal of tibial when used with the Panta Arthrodesis Nail Support Device. The notice goes stated that Integra determined that the misalignment of these screws may be attributed to improper surgical technique and to that end, the Field Safety Notice is being issued in order to reinforce the importance of following the recommended guidelines our lined in the product's Surgical Technique Manual (a copy of the referenced guidelines are provided with the Field Safety Notice (FSN). The FSN goes on to direct the reader/customer to applicable section sections of the manual to which particular attention should be directed in order to minimize the potential for misalignment of the screws. The customers are asked to distribute the information provided to the surgical teams at their various institutions. Moreover, they are also asked to sign and return the enclosed Field Safety Notice Acknowledgement and Return Form so as to document their receipt of the notification and the material provided. Should the customer(s) have questions, they are directed to contact Sean Luland, Director of Marketing-Lower Extremity Reconstruction at Integra LifeScience at: 513-533-7929, for additional information.For questions regarding this recall call 609-275-0500. |  | 64 | 64 units were distributed | Nationwide Distribution including MD, IL, NU, CA, GA, CO, OH, SD, AR, FL, TX, MA, MO, NC, OK, KY, NJ, UT, VA, MN, DC, PA, WA, CT, AZ, IL, MI, AL, NE, TN, NM, NV, LA, OR, DE, WI, ND, NH, MS, SC, IN and Puerto Rico. | Terminated | February 19, 2013 | 642 | 64 | Not_Computer
Z-1424-2012 | 61304 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Attune Impaction Handle Warsaw, IN 46582 |  |  | March 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a Recall Notice for nine lots of the Attune" Impaction Handle due to the potential for the handle trigger to break during surgery. The Attune Impactor is provided as a part of a set of tools and is designed specifically for the installation of the Attune Knee | Device Design | DePuy Orthopaedics, Inc. sent an URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE letter dated March 16, 2012 to all consignees affected which extended to the Depuy Distributor level, surgeon and hospital level. The affected DePuy Distributors were notified via email on March 16, 2012. The sales representatives notify hospitals and surgeons by mail or in person with written communication. The letter identified the affected product, reason for recall notice, clinical implications and actions to be taken. Consignees were instructed to complete and return the Response cards to DePuy by fax to 574-372-7567 or email to:kseppa@its.jnj.com. When the new impaction handles are available, the DePuy sales representative will be responsible for providing the new handles, removing and returning the affected handles. For product-related questions, please contact your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact DePuy's Scientific Information Office at 1-888-554-2482 (M-F; 8 am - 5 pm EST.) For questions about recall information provided, please contact Katie Seppa, Manager of Customer Quality, 574-372-7333 (M-F; 8 am - 5 pm EST.) |  | 136 | 136 impaction handles | USA Nationwide Distribution including the states of: MA, MD, IL, IA, MN, WA, CA, FL, NC, VA, and IN. | Terminated | January 09, 2014 | 664 | 136 | Not_Computer
Z-1427-2012 | 61538 | [6 empty] | Hudson RCI, BiteGard Oral Bite Block, Single Patient Use, Rx Only, Teleflex Medical, Research Triangle Park, NC. The Teleflex Medical BiteGard Molar Bite Black is a disposable single patient use device that fits between the patients molars to prevent the patient from closing their mouth entirely and from obstructing the Laryngeal Mask Airway (LMA) tube, in the case where one is used. |  |  | April 03, 2012 | 2012 | Teleflex Medical | Teleflex Medical received complaints that the bite block can become separated from the handle. If separation occurs, medical intervention may be required to remove the bite block from the patient's trachea or esophagus/GI tract. | Process control | Consignees received "Urgent Medical Device Recall Notification" letters on or about 04/03/2011. The letters described the issue and provided recommended actions regarding remaining product and returns. Questions were directed to Customer Service at (866) 804-9881 (US). |  | 333724 | 333,724 units | Worldwide Distribution, including Nationwide (USA) and the countries of Canada, Singapore, China, Denmark, Netherlands, India, Australia, Japan and the United Kingdom. | Terminated | May 20, 2014 | 777 | 333724 | Not_Computer
Z-1425-2012 | 52748 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Total Bilirubin, 6L45-20 and 6L45-40. |  |  | July 09, 2009 | 2009 | Abbott Laboratories Inc. | Abbott Laboratories has issued a recall on the Clinical Chemistry Total Bilirubin due to an under-recovery on one clinical neonatal specimen resulting in a value less than the assay linear limit. | Material/Component Contamination | Abbott Laboratories plans to release a Product Correction letter to customers instructing customers to follow the steps provided in order to revise their Total Bilirubin assay parameters. The letter explained the reason for the recall, along with the patient impact. Customers were to keep the letter as part of their laboratory records and also forward a copy to any other laboratories that they may have forwarded Total Bilirubin kits. U.S. Customers should contact Customer Service at 1-877-4ABBOTT, if they have questions. |  | 11282 | 11,282 | Worldwide Distribution -- USA, Guatemala, Chile, Brazil, Canada, Germany, Hong Kong, Thailand, Puerto Rico, Australia, New Zealand, Trinidad &amp; Tobago, Bermuda, Curacao, Cayman Island, and Jamaica. | Terminated | March 07, 2014 | 1702 | 11282 | Not_Computer
Z-1431-2012 | 61547 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Zuma Mobility System Models: Zuma Mobility Assist 412045-01 and Zuma Mobility Trainer 412046-01. |  |  | April 06, 2012 | 2012 | KCI USA, Inc. | The Zuma Mobility Assist and Trainer have the potential to bind during lowering operation. | Device Design | KCI sent an Urgent - Voluntary Medical Device Correction notification letters dated April 6, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. KCI will contact affected customers to schedule a time to upgrade the device. For questions contact KCL's Customer Technical Service Center at 1-800-275-4524. Select option 3 followed by option 2. |  | 4 | 4 units | USA Nationwide Distribution including the states of ID and FL | Terminated | June 19, 2012 | 74 | 4 | Not_Computer
Z-1426-2012 | 61342 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PFC Steinman Pin/Drill Pack Sterile (WPC 667-2012) Drill Bit and Steinman Pin Packet (1 unit) |  |  | March 12, 2012 | 2012 | DePuy Orthopaedics, Inc. | Sterility of the pins is at risk because product packaging, which is designed to include 2 pouches (inner pouch and outer pouch) around the pin mounting card only, contains a single pouch, which could compromise product sterility. | Packaging process control | The firm, DePuy Orthopedics Inc., sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" dated March 12, 2012 to its customers via mail or in person with written communication. The notice described the product, problem and actions to be taken.The customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN. The sales representatives are expected to aid customer in returning of recalled products. Distributors and sales representatives are to return the recalled products through the normal DePuy Returns process, to attention of Returns and marking H12-01 on the outside of the box.The customers were also instructed to complete and return the Reconciliation Form to their Depuy Sales Representative or fax to 574-372-7567. Follow-ups for the Reconciliation Form for non-responders will be conducted with the Distributors until the form is returned. For Hospitals that are now out of business or refuse to sign the response card, the Distributor will indicate this on the Reconciliation form and return.For questions about the recall information provided, please contact Manager of Customer Quality, at 574-372-7333 (M-F; 8am - 5pm EST). For product-related questions, please contact your local DePuy Orthopaedics sales representative. For clinical-related questions from surgeons, please contact Depuy's Scientific Information Office at 1-888-554-2482 (M-F; 8am - 5pm EST). |  | 4800 | 40 of each lot, 120 total | Nationwide distribution: USA including states of: FL, LA, TX, OR, CA, AZ, KY, GA, MS, and PA. | Terminated | March 06, 2013 | 359 | 4800 | Not_Computer
Z-1421-2012 | 61403 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Sterile wound closure strips, packaged as follows: |  |  | February 07, 2012 | 2012 | Deroyal Industries, Inc. Lafollette | Packages of Episeal wound closure strips and surgical kits containing them were recalled because the adhesive backing on Episeal wound closure strips may not release from the backing rendering the product unusable. | Nonconforming Material/Component | DeRoyal Industries sent a Recall of Episeal Wound Closure Strip letter dated February 7, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following action:1. Please identify and quarantine any affected inventory using the product and lot numbers listed above.2. Render affected product unusable, destroy and discard according to your facilitys guidelines.3. If you are a distributor, forward this recall to any of your end users that may have affected product or provide us with the list and we will contact them.4. Please complete the enclosed Notice of Destruction form, indicating the affected products you found in your inventory and destroyed. PLEASE COMPLETE AND RETURN THIS FORM EVEN IF YOU NO LONGER HAVE THE PRODUCT IN INVENTORY.Return this form by fax (865-362-3716) or email (recalls@deroyal.com) within two weeks of the date of this letter. DeRoyal will use your response on the Notice of Destruction form to send replacements as soon as they become available.If you have questions or need assistance with the recall, please contact DeRoyal Customer Support at 1-800-251-9864.We apologize for inconvenience this may cause you |  | 126720 | 126,720 units | Worldwide Distribution -- USA (nationwide) | Terminated | July 10, 2013 | 519 | 205661 | Not_Computer
Z-1433-2012 | 61465 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista(R) MG Flex(R) Reagent Cartridge |  |  | February 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed a negative shift of approximately 0.2mg/dL [0.08 mmol/L] in patient and QC samples when using Vista Magnesium | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notification letter dated February 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter explained the correction needed and instructed customers to forward the letter to anyone they may have distributed the affected product.For questions or concerns call 800-441--9250.For questions regarding this recall call 301-631-6299. |  | 9178 | 9,178 | Nationwide Distribution including Puerto Rico and within the US to Washington, DC and the following States: AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Puerto Rico. | Terminated | March 10, 2015 | 1117 | 9178 | Not_Computer
Z-1437-2012 | 61622 | No | KDL | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ViaSpan Cold Storage Solution 1000mL Bag distributed under the Barr/Duramed Pharmaceuticals label (10 x 1000mL bags). |  |  | March 30, 2012 | 2012 | Teva Pharmaceuticals USA, Inc. | Lack of assurance of sterility | Material/Component Contamination | Teva Pharmaceuticals USA sent an "URGENT MEDICAL DEVICE RECALL-RETAIL LEVEL" letter dated March 30, 2012. The letter identifies the problem, product, and actions to be taken by the customers. The letter instructs customers to check their inventory and to discontinue distribution of the affected product. A Stock Response Form was enclosed for customers to complete and return via fax to 817-868-5362. Contact Teva Customer Services at 800-545-8800 for questions regarding this notice. |  | 4070 | 4,070 | Nationwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CT, DE, FL, GA, IA, IL, KS, KY, LA, MA, MD, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV and Washington, D.C. | Terminated | October 25, 2012 | 209 | 4070 | Not_Computer
Z-1434-2012 | 55440 | No | CIX | Albumin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry Albumin BCG (AlbG) List Number 7D53-21 |  |  | March 10, 2010 | 2010 | Abbott Laboratories Inc. | Observed particulate matter in some reagent cartridges. | Material/Component Contamination | Abbott Laboratories sent a Product Recall notification letter dated March 10, 2010, to all affected customers. The notification letter contained an explanation, patient impact, neccassary actions, and contact information.Customers were instructed to determine if they currently have any of the affected product in their inventory. Discontinue use and destroy any remaining inventory of the affected product in accordance with their institution's policies and procedures. Completed the enclosed Customer Reply Form. If customers forwarded any of the affected product they should provide a copy of the notification letter to their customers.For any questions customers should call Customer Service at 1-877-422-2688.For questions regarding this recall call 626-440-0705. |  | 2718 | 2718 (U.S. 652) | US nationwide and internationally to Germany, Hong Kong, Australia, New Zealand, Trinidad/Tobago, and Jamaica. No dsitribution of Canada or Mexica. | Terminated | April 18, 2012 | 770 | 2718 | Not_Computer
Z-1435-2012 | 61620 | No | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | MEDPOR PLUS SST EZ 22mm Sphere Implant, Howmedica Osteonics Corp, 15 Dart Road, Newnan, GA 30265. MEDPOR PLUS Orbital Implants are intended to restore the volume of an enucleated or eviscerated eye globe. |  |  | March 20, 2012 | 2012 | Howmedica Osteonics Corp dba Stryker Craniomaxillofacial | The MEDPOR PLUS SST EZ 22mm sphere implants are not within specification of 22mm +/- 8, therefore implant is not easily released from the syringe. | Process control | The sole distributor was notified by "Urgent Product Recall" letter on March 20, 2012. "Urgent Medical Device Recall" letters were sent to Stryker Customers on March 26, 2012. The letters described the product issue, hazards, and actions recommended to mitigate the risk. Customer number provided: (800) 962-6558. |  | 31 | 31 units | US Distribution, the state of Michigan. | Terminated | August 10, 2012 | 143 | 31 | Not_Computer
Z-1441-2012 | 49059 | No | LGW |  |  | Neurology | PMA | OMG (Observational Mechanical Gateway) Connector - A Model Number SC-9305 |  |  | April 22, 2008 | 2008 | Advanced Bionics Corporation | The recall was initiated because there is a possibility that the OMG Connector-A is incorrectly packaged with an OMG Connector-M Instruction manual. | Labeling mix-ups | Algotec Limited initiated a Field Safety Notice Urgent Medical Device Correction letter dated April 22, 2008, with Boston Scientific forwarding a recall letter to one customer who received the recalled product. The letter instructed the consignee to replace the previously packaged manual with the OMG Connector-A manual that was supplied with the notification. The replacement of the manuals is to be confirmed via the Verification form.Customers were instructed to complete and fax the Verification Form to their local Boston Scientific Customer Service to Fax 00 44 1442 411732 by 06 May 2008.For questions regarding this recall call 661-949-4481. |  | 6 | 6 Units | Worldwide Distribution including Great Britian | Terminated | April 19, 2012 | 1458 | 6 | Not_Computer
Z-1440-2012 | 49410 | No | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Synaptic 3600 and 3800 with Cranial Electrical Stimulation manufactured by Synaptic USA, Aurora, CO. |  |  | July 18, 2008 | 2008 | Synaptic Corp | Device was distributed without marketing approval. | PMA | Synaptic Corp, sent a Urgent Medical Device Recall letter dated July 14th, 2008, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.The customer was instructed to return your 3600 and or 3800 devices so that the cranial stimulator can be removed and destroyed. Please call 1-800-685-7246 to schedule pick up of your device..Complete the enclosed Confirmation Form and return it ( postage paid) to Synaptic Corp. It is important that you return the attached response confirmation form even if you do not have the (CES) module in your model.Should you have any questions about this recall please call 1-800-685-7246 ext. 101. |  | 59 | 59 units | Worldwide Distribution -- USA (nationwide) and the country UK. | Terminated | April 19, 2012 | 1371 | 59 | Not_Computer
Z-1438-2012 | 61089 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF. CP113456, MAK OSS No Head Lock Pin, NO HEAD / WITH PLUG CO-CR-MO ALLOY/ARCOM UHMWPE |  |  | March 20, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after receiving complaints of the MAK OSS No Head Lock Pin dislodgingfrom its intended position (post-initial surgery). An investigation was initiated and the available lots ofthe Lock Pin were evaluated and measured against print specifications. | Component design/selection | Biomet sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter dated March 20, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to continue monitoring implanted patients for disassociation of the pin component and report to Biomet. The Alert included a complete listing of device invoice history and patient information (if supplied). Contact the firm for confirmation of receipt of the notice and for questions at 574-372-3983. |  | 64 | 64 units distributed | Distributed in California and Pennsylvania. | Terminated | February 07, 2013 | 324 | 64 | Not_Computer
Z-1443-2012 | 61453 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PREMIER(R) Anterior Cervical Plate System Adjustable Drill Stop, REF 6905712, QTY: 1 EA, Rx only, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132 |  |  | March 13, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | Drill stops may be bypassed when the device is used with powered drilling instruments. | Device Design | Medtronic notified their US Sales force on March 5, 2012, of the upcoming field action. The written "Voluntary Recall" communications were disseminated March 13 and March 14 to the US consignees. The letter identified the affected products, problem and actions to be taken. Customers are instructed to 1) disseminate this information to personnel within their facility as appropriate; 2) review work area, district sales office, and sample inventory for affected product; 3) complete the attached Risk Manager Questionnaire and return any recalled product; and 4) Sign and acknowledge the attached Sales Rep Questionnaire indicating receipt of this letter. For questions or concerns, contact your Medtronic Sales Representative or Global Quality Department at 800-876-3133, extension 3197. |  | 1411 | 1411 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Barbados, Brazil, Netherlands, Hong Kong, Japan, Korea, Columbia, Chile, Canada, India, Jamaica, Switzerland, Costa Rica, Ecuador, and Mexico. | Terminated | April 30, 2013 | 413 | 4198 | Not_Computer
Z-1439-2012 | 61188 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Optilock, Locking Screw 3.5 mm x 28 mm, non sterile, REF 37728, Biomet Trauma Parsippany, NJ 97054. |  |  | March 15, 2012 | 2012 | Biomet, Inc. | Biomet initiated this action after a complaint was received that item 36728, Non-Locking Screw 3.5mm x 28mm was in a package labeled as 37728, Locking Screw 3.5mm x 28mm. | Labeling mix-ups | A March 15 2012 Urgent Medical Device Recall Notice was sent to Distributors, Operating Room Managers, and Physicians with directions to locate and remove recalled items with instructions to provide the recall notice to any sub-accounts. The letter also identified the affected product and explained the reason for the recall. OR Managers and Physicians are to confirm recall notice receipt by calling 800-348-9500, extention 3983 or 3009 and make arrangements for product return. Distributors are to complete and FAX back the response tracking form to the attention of Angi Dickson, Biomet @ 574-372-1638. Questions should be directed to 574-372-3983 or 574-371-3009 M- F, 8 AM- 5 PM, (ET). |  | 45 | 45 | Worldwide Distribution -- USA, including NJ, TX, VA, MI IL, KT, &amp; Puerto Rico, and country of the Netherlands. | Terminated | August 29, 2013 | 532 | 45 | Not_Computer
Z-1445-2012 | 61390 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific LATITUDEÂż Patient Management System Communicator, Model 6476, Distributed by Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798, Made in Malaysia. The LATITUDE Patient Management system is intended for use to remotely communicate with a compatible Guidant or Boston Scientific pulse generator and transfer data to a central database. |  |  | February 24, 2012 | 2012 | Boston Scientific CRM Corp | Boston Scientific LATITUDEÂż Patient Management System Communicators were shipped to the incorrect patients. | Process control | The patients were notified via phone starting on 2/24/12 that their communicator will not be able to activate successfully and that their communicator and sensors will be replaced. If the patient could not be reached via phone or voice mail, a Boston Scientific letter dated 2/29/12 was sent to the patient. The letter described the problem, explained What is a LATITUDE Communicator and What should you do. Questions were directed to Patient Services at (866) 484-3268. |  | 8 | 8 | Nationwide Distribution, including the states of: AL, CA, FL, IL, IN, MO, NJ, NY, OH, OR, TX VA, and WA. | Terminated | August 02, 2012 | 160 | 8 | Not_Computer
Z-1446-2012 | 61637 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Triage TOX Drug Screen (MTD) Panel, Model # 94400. |  |  | March 12, 2010 | 2010 | Alere San Diego dba Biosite Innovacon Hemosense Inc | The recall was initiated because Biosite Inc. received an increase in customer calls regarding unexpected positive THC results when using the Triage TOX Drug Screen. | Nonconforming Material/Component | Biosite sent a "PRODUCT ADVISORY NOTICE" letter dated March 12, 2010 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter states that customers may continue to use the product. Customers were instructed to complete and fax an enclosed verification form to Technical Services at 1-858-695-7100. Customers with any questions about the information contained in the notification were instructed to contact Biosite Inc. at (888) 246-7483. |  | 1258002 | 1,258,002 kits | Worldwide Distribution-USA (nationwide) and the countries of Canada and European Union. | Terminated | April 23, 2012 | 773 | 1258002 | Not_Computer
Z-1447-2012 | 50072 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT/AEROSET Activated Aspartate Aminotransferase |  |  | June 30, 2008 | 2008 | Abbott Laboratories Inc. | This letter is to inform you that some samples may generate A-AST results of less than (&lt;) 5 U/L or results within the reference range when actual concentration exceeds the non-Flex linearity of 1985 U/L. In other cases the following errors may be generated.-ARCHITECT cSystems: Error Code 1053 (Unable to calculate result, rate reaction linearity failure)-AEROSET: RL% Error CodeRevised P | Nonconforming Material/Component | Abbott Laboratories sent a Product Correction letter dated August 5, 2008, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:Update your assay parameters for A-AST1. Follow the steps in the appropriate attachment: - Use Attachment A for ARCHITECT cSystems - Use Attachment B for AEROSET.2. All A-AST results of less than (&lt;) 5 U/L should be evaluated by comparing to other clinically relevant assaysRetain a copy of this letter for your laboratory records.For questions regarding this recall call 626-440-0705. |  | 1646 | 1,646 | Worldwide Distribution including Germany | Terminated | April 23, 2012 | 1393 | 1646 | Not_Computer
Z-1452-2012 | 60774 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sarns Modular Perfusion System 8000 Catalog number 148689 or 78-8067-3878-3 Temperature-Pressure Board, Cardioplegia Monitor (service part only) |  |  | March 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo received multiple reports of malfunctions of the Cardioplegia and arterial monitors for the Sarns Modular Perfusion System 8000. Upon review of complaint investigation, the malfunctions are caused by or could have been caused by soldering anomalies on the circuit board of the monitor. The compromised solder joints could cause intermittent or loss of display, or could cause monitor reset, po | Device Design | TERUMO sent an URGENT MEDICAL DEVICE RECALL- SAFETY ADVISORY to all consignees on March 21, 2012 via Federal Express. Terumo issued the safety advisory to remind users how to intervene if they experienced intermittent or complete loss of pressure display on the arterial monitor or cartiplegia monitor, or a related pump stop. Customers were asked to review the Medical Device Correction notice, assure that all users are aware of the notice, confirm receipt of the communication by faxing the attached Customer Response Form to 1-800-292-6551.For questions regarding this recall call 1-800-521-2818. |  | 3022 | 3,022 | Worldwide Distribution - USA including AK, AL , AR, AZ , CA, CO , CT, DC, DE , FL, GA, HI, IA , ID , IL, IN , KS, KY , LA, MA , MD, MI, MN , MO, MS , MT, NC , ND, NE, NH, NJ, NM , NY, OH , OR, PA , RI, SC , TN, TX , VA, WA , WI, WV, WY, and Puerto Rico and the countries of PERU, Philippines, REPUBLICA, SANTIAGO, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tujunga, United Arab Emirates, Uruguay, Venezuela, Vietnam, Wooddale, and Yemen. | Terminated | July 11, 2013 | 477 | 3022 | Not_Computer
Z-1449-2012 | 60595 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ref, S331140 Selex M2A Magnum Modular Head, 40 mm head diameter plus 3 mm neck,1 Taper, Sterile, Biomet Orthopedics, Warsaw, IN |  |  | September 13, 2011 | 2011 | Biomet, Inc. | Devices with a specification for a +6mm neck length were incorrectly labeled for a +3mm neck length specification. | Nonconforming Material/Component | Biomet contacted consignees with device inventory by telephone on September 13, 2011 requesting all associated product be quarantined immediately. An Urgent Medical Device Recall Notice was sent via U.S. mail March 23, 2012 to consignees with directions to immediately locate, discontinue use of the affected product and return to Biomet. The letter instructed consignees to carefully follow the instructions on the enclosed "FAX Back Response Form" and fax a copy of the Response Form to 574-372-1683 prior to return of product. The notification letter to consignees contained directions to confirm receipt of notice by calling 800-348-9500, extension 3983 or 3009. Questions concerning recall are directed to 574-372-3983 or 574371-3009 M-F 8 A.M. to 5 P.M. (ET). |  | 5 | 5 devices | USA Nationwide Distribution including the states of: MI, NC., OK and WI. | Terminated | February 21, 2013 | 527 | 5 | Not_Computer
Z-1448-2012 | 60865 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Software, User Manuals, and Handi-Guides for acute and chronic peritoneal dialysis. Software versions earlier than 2.8. |  |  | March 30, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Increased risk of intraperitoneal volume (IIPV), referred to as Overfill | Software design | Fresenius Medical Care sent a Urgent Product Information letter dated March 30, 2012, and April 3, 2012, to all affected customers. The letter was sent via Certified Mail, Return Receipt Requested. The letter identified the product the problem and the action needed to be taken by the customer.Recall to occur in 2 phases:Phase I: Fresenius Medical Care notified all clinics with patients using Liberty Cyclers with software versions earlier than 2.8 that these patients are receiving updated User Manuals and Handi-Guides. These new materials will include updated descriptions of the safety warnings included with the materials provided with the release of software version 2.8. Clinics will also be informed that their patients will be receiving replacement cyclers with updated software within 12 months. These notifications via telephone and certified letter (return receipt requested). Following clinic notification, Fresenius Medical Care will contact patients operating Liberty Cyclers with software versions earlier than 2.8 by phone, informing them that they will be receiving updated User Manuals and Handi-Guides. In addition, these patients will be informed that they will be receiving replacement cyclers with updated software at some point in the future and that Fresenius Medical Care anticipates that the full timeline of cycler replacements will occur over several months. A certified letter return receipt requested will be sent to all patients after the initial phone contact containing the updated User Manuals and Handi-Guides. Phase II: Liberty Cyclers with updated software will be issued to customers using Liberty Cyclers with software versions earlier than 2.8 as inventory is available.Please complete and return the enclosed Reply Form, indicating receipt and understanding of this communicationFurther questions please call (1-800) 227-2572. |  | 10146 | 10,146 units | Worldwide Distribution -- USA (nationwide) | Completed |  |  | 10146 | Computer
Z-1399-2012 | 61553 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Automatic Mobile X- Ray (AMX) series: Brivo XR285 15KW, Brivo XR285 30KW, Optima XR200 15KW, Optima XR200 30KW, Optima XR220 15KW, and Optima XR220 30KW. |  |  | December 23, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a discrepancy on the x-ray source assembly of the mobile radiographic product, related to the light field to x-ray field edge alignment. | Radiation Control for Health and Safety Act | GE Healthcare will issue an Important Electronic Product Radiation Warning letter to customers. The letter will describe the defect and related hazards, along with the safety instructions, affected product, and product correction GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with an approved plan, the details of which will be included in a subsequent communication to customers or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the external tube port on the source assembly to address the issue.Questions or concerns can be directed to the Call Center at 800-437-1171 or your local GE Healthcare Service Representative. |  | 130 | 130 installed in US | USA (nationwide) | Terminated | January 16, 2014 | 755 | 130 | Not_Computer
Z-1458-2012 | 61457 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | CARE CLINER 6550 (STD)/6570 (X-LG) with Drop Arm products are labeled in part: "***CARE CLINER 6550/6570 (w/ NYLON CASTERS)***Winco ISO 9001-2008 Certified***All of our best features, plus the added versatility of drop arms***Leave it to Winco to combine the best clinical chair features in the industry with the convenience of dual drop-arms! Our drop-arm care cliners make assisting your patients easier. Our side-panels lift, making clean-ups quick and easy.***KEY FEATURES***Standard head-rest cover***Padded drop-down armrests***Dual fold down side-tables***Easy-to-clean side panels***Trendelenburg Positioning***1.800.237.3377***5516 SW 1st Lane***Ocala, Florida 334474***www.wincomfg.com***". |  |  | February 06, 2012 | 2012 | Winco Mfg., LLC | Winco Manufacturing LLC is recalling the Care Cliner, XL Care Cliner, Drop Arm Care Cliner, Elite Care Cliner, and XL Elite Care Cliner chairs with nylon casters distributed between 04/01/2009 and 08/31/2009. Certain 5" nylon casters shipped with the chairs may be defective. There is a potential of the caster's axle working itself out of the assembly under unconventional conditions. Should this | Device Design | Winco Mfg. sent an "URGENT DEVICE RECALL" letter dated February 1, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete the Caster Recall Replacement Form and return it to the firm's Customer Service Department via fax to 352-854-9544. Contact the Customer Service Department at 800 237-3377 for questions regarding this notice. |  | 1311 | 1,311 Casters (437 Chairs) | Worldwide Distribution-USA (nationwide) and the countries of Canada and Singapore. | Terminated | July 27, 2012 | 172 | 5132 | Not_Computer
Z-1454-2012 | 61482 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Dimension Vista(R) software version 3.4 |  |  | January 16, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints on Vista(R) software version 3.4 for Below Manufacturer Assay Range error flags associated with QC and patient test results that are within the assay range as defined in the Instructions For Use for the respective method. This issue is isolated to Vista(R) software version 3.4. Customers who have not made any changes to the default method configurations | Software change control | Siemens sent an Urgent Field Safety Notice dated December 19, 2011 to all affected customers alerting them to Invalid Below Manufacturer Assay Range Flags on Dimension Vista(R) Systems with Software Version 3.4. The UFSN included instructions on how to correct the situation and also included an effectiveness check form to be faxed to the firm at 302-631-8467. On January 16, 2012, the firm issued an additional UFSN alerting consignees to Invalid Above and Below Manufacturer Assay Range Flags. This notice also included directions and a Field Correction Effectiveness Check form to be completed and faxed to 302-631-8467.For technical questions customers were instruced to call 800-441-9250.For questions regarding this recall call 302-631-6299. |  | 1683 | 1,683 | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland &amp; Baltics, France, Germany, Italy, Japan, South Korea, Malaysia, Netherlands, New Zealand, Norway, Portugal, Puerto Rico, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland and United Kingdom. | Terminated | March 05, 2014 | 779 | 1683 | Computer
Z-1455-2012 | 61346 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Quintex Cervical Plating System - Screws. The Quintex Cervical Plating System is intended for the treatment of cervical spinal instability. The Quintex Cervical Screws are available in three different versions: semi-constrained, constrained, and dynamic. The screws all have a head diameter of 4.0mm and vary in length from 10mm to 22mm. Rescue screws are also available with a 4.5mm head diameter and lengths of 11-19mm. |  |  | January 24, 2012 | 2012 | Aesculap Implant Systems | The firm has received complaints that the locking ring that sits in the head of the screw can come out of the screw head when engaged with the screwdriver if the screwdriver is improperly angled. | Nonconforming Material/Component | The firm initiated this recall by sending an "Important Recall Notification" letter to consignees on January 20, 2012. This letter instructed consignees not to use the product and requested that all stock be returned to the recalling firm promptly. Customers may contact 1-(866) 229-3002 for questions. |  | 891 | 891 | Nationwide Distribution, including the states of CA, CO, FL, MO, and TX. | Terminated | December 13, 2012 | 324 | 891 | Not_Computer
Z-1462-2012 | 61514 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura X-ray System. |  |  | April 02, 2012 | 2012 | Philips Healthcare Inc. | The monitor Ceiling suspension may fail, which may cause the monitor to drop. | Process control | Philips sent an "URGENT-FIELD SAFETY NOTICE" dated March 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides information for continued safe and proper use of the equipment. Contact your local Philips representative at 866-767-2822 for further support regarding this issue. Recall expanded 5-2013 to additional 460 units distributed worldwide. |  | 145 | 145 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Belgium, Brazil, Canada, Czech Rep, Chili, China, Columbia, Denmark, Egypt, France, Germany, India, Indonesia, Italy, Japan, Malaysia, Mexico, Morocco, Netherlands, Oman, Russia, Saudi Arabia, Singapore, South Korea, Spain, Thailand, United Kingdom, and Vietnam. | Terminated | May 09, 2014 | 767 | 145 | Not_Computer
Z-1456-2012 | 61473 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Certain Low-Profile, Angled Abutment, Model Number ILPAC3217. Product is labeled in part: "***Biomet 3i***4555 Riverside Drive , Palm Beach Gardens, FL 33410 USA***REF ILPAC3217***LOT 2011091179***Certain Low Profile 17 Abutment***3.4mm(D) X 2mm(H)***Abutment:; Pillar; Abutment; Abutment; Pilastro; Pillar***CE dose***Use by 2016-09***Sterile using Radiation***RX Only***Do not reuse***Consult accompanying documents***Do not re-sterilize***Do not use if package is damaged***ILPAC3217***2011091179***"P-LBLNY3P***Rev. D***". |  |  | January 10, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i initiated a recall for a small percentage of the 33mm (ILPAC3217) abutment packages that may contain the 5mm (ILPAC5217) abutment. | Material/Component Contamination | Biomet 3i , LLC sent Urgent: Medical Device Recall notices on 1/10/2012 to each customer in via e-mail, facsimile, telephone or postal mail. The letters identified the affected product and the reason for the recall. Each customer was instructed to check their respective inventory for the affected product and return any unused product to Biomet 3i at the address provided. Biomet 3i will issue replacement product at no charge. Questions or concerns should be directed to 1-800-342-5454. |  | 400 | 400 | Worldwide Distribution -- Canada, Chile, Columbia, France, Germany, Hong Kong, India, Israel, Italy, Japan, Mexico, Netherlands, Norway, Poland, Singapore, South Africa, Spain, Sweden, Switzerland, UK, US, and Vietnam. | Terminated | April 30, 2012 | 111 | 400 | Not_Computer
Z-1470-2012 | 61409 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS&amp;amp;amp;amp;amp;reg; 5600 Chemistry System, Product Code: 6802413 |  |  | March 15, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho-Clinical Diagnostics is recalling VITROS Chemistry Systems due to possible internal hard drive failure. | Nonconforming Material/Component | Ortho Clinical Diagnostics (OCD) sent an Important Product Correction Notification letter dated March 15, 2012 via FedEx overnight courier to all affected customers. Foreign affiliates were informed by email on 15 March 2012 of the issue and instructed to notify their consignees of the issue and actions. The letter identified the affected products, problem and actions to be taken. The letter instructed customers to: 1) Consider performing the backup procedure more frequently to help prevent the loss of data on your system; 2) Contact OCD Customer Technical Service representatives if a hard drive failure occurs on your system; A Field Engineer will be dispatched to replace the hard drive on your system; 3) Complete and return the Confirmation of Receipt form upon receipt of this notification; and 4) Post this notification by each VITROSÂż System in your facility or with the user documentation. For questions, refer to the Questions and Answers section. For additional questions, contact Customer Technical Services at 1-800-421-3311. |  | 417 | 417 | Worldwide Distribution | Open, Classified |  |  | 785 | Not_Computer
Z-1463-2012 | 61472 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Bolus Cord; a remote bolus cord and switch for use with GemStar 7 Therapy and Pain Management Pumps, list 13000, 13150, 13087 and 13088; Hospira, Inc., Lake Forest, IL 60045; List Number 13027; |  |  | March 28, 2012 | 2012 | Hospira Inc. | Hospira has received customer reports of bolus delivery failures when using the GemStar bolus cord. Possible failures may be a result of cracked or damaged bolus cords which may result in an open or short circuit of the Bolus Cord. | Nonconforming Material/Component | Hospira, sent an "URGENT DEVICE FIELD CORRECTION RECALL" letter dated March 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to take care when removing the bolus cord from the base of the GemStar unit, inspect the cord for damage, discontinue using any damaged cords, and to call Hospira at 1-800-241-4002 for replacement. A Reply Form was enclosed for customers to complete and return via fax to 1-866-912-7085 or e-mail to: Hospira2627@stericycle.com. Contact Hospira Customer Care at 1-877-946-7747 for technical assistance. |  | 87095 | 47,127 cords - U.S. and 39,968 cords - ROW | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Malta, Mexico, Netherlands, New Zealand, Nola, Norway, Oman, Philippines, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay and Venezuela. | Open, Classified |  |  | 87095 | Not_Computer
Z-1465-2012 | 61510 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Anatomical Pathology Software Client Releases- |  |  | March 30, 2012 | 2012 | Medical Information Technology, Inc. | Meditech Anatomical Pathology software is being recalled because data on a pathology specimen may be inadvertently deleted. | Software design | MEDITECH sent a notification beginning on January 19, 2012 to customers who are LIVE with MEDITECH Client Server Anatomical Pathology 5.6 and higher. This notification was made via e-mailed task updates that informed customers of the software issues and the code corrections that are available. |  | 280 | 280 | Worldwide Distribution - USA Nationwide and the countries of Bahamas, Canada and United Kingdom. | Terminated | August 13, 2014 | 866 | 280 | Computer
Z-1466-2012 | 61516 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage CardioProfiler Panel, PN 97100CP. The test is used as an aid in the diagnosis of myocardial infarction (injury), an aid in the diagnosis and assessment of severity of congestive heart failure, an aid in the risk stratification of patients with heart failure, and as an aid in the risk stratification of patients with acute coronary syndromes. |  |  | March 05, 2012 | 2012 | Alere San Diego | Alere San Diego, Inc. is initiating a voluntary recall of the Alere Triage CardioProfiler Panel PN 9700CP, Lot W49569 and the Alere Triage Profiler SOB Panel PN 97300, Lot W48990. Alere San Diego, Inc. has determined that these lots have an increased frequency of Troponin I results &gt;0.05ng/mL for samples which are found to be below 0.05ng/mL upon additional testing. | Process control | A recall communication was initiated on 3/15/12 with Alere San Diego, Inc. forwarding a "Urgent Medical Device Recall" letter to all their customers who purchased the Alere Triage CardioProfiler Panel PN 97100CP, Lot W49569 and Alere Triage Profiler SOB Panel PN 97300, Lot W48990. The letter informed the customers about the problems identified and the actions to be taken. Customers with any questions about the information contained in this notification should contact Alere San Diego, Inc. at (877) 308-8287. |  | 1080 | 1080 kits | Worldwide Distribution, including Nationwide (US) and the countries of Brazil, China, India and Kuwait. | Terminated | December 04, 2012 | 274 | 2396 | Not_Computer
Z-1464-2012 | 61528 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Systems with Eleva software version 2.x and stitching option |  |  | April 03, 2012 | 2012 | Philips Healthcare Inc. | Potential misdiagnosis due to improperly automatically stitched images, improper image ruler handling. | Labeling design | The firm, Philips Healthcare Imaging, issued an "URGENT-FIELD SAFETY NOTICE" dated March 30, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions carefully and to not use the Philips Image Ruler for measurements. The firm plans to supply an Addendum to the Instructions For Use and a Philips Service Engineer will contact the customer as soon as the Field Action Kit is ready to be implemented.If you need any further information or support concerning this issue, please contact your local Philips representative or call 1-866-767-2822. |  | 606 | 606 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech. Rep, Denmark, Finland, France, Germany, Guadeloupe, India, Indonesia, Iran, Iraq, Italy, Japan, Malta, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Salkowski, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Ver.Arab.Emir. | Terminated | March 13, 2014 | 709 | 606 | Not_Computer
Z-1475-2012 | 60610 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | UMP Personal Sentry Deluxe Pull String Monitor, Model #91650. |  |  | December 06, 2011 | 2011 | Stanley Security Solutions, Inc. | Monitor gives false alarms | Nonconforming Material/Component | Stanley Healthcare Solutions sent an "IMPORTANT RECALL NOTICE" dated December 8, 2011. The letter identifies the product, problem, and the actions to be taken by the customers. The letter instructs customers to return affected product through Customer Service and to have the appropriate authority sign the Acknowledgement Form. Customers with questions can call UMP Customer Service at 1-800-206-1045. |  | 592 | 592 units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | September 24, 2013 | 658 | 592 | Not_Computer
Z-1473-2012 | 61535 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | ASP 6025 Vacuum Tissue Processor; an automated modular tissue processor; Leica Biosystems Nussloch GmbH, Heidelberger Str. 17-19, D-69226 Nussloch, Germany; The intended use of this device is as an automated system used to process tissue specimens for examination through fixation, dehydration, and infiltration. |  |  | March 30, 2012 | 2012 | Leica Microsystems, Inc. | There is a potential malfunction with the device in autorotation mode, which can lead to incorrect reagent handling of the device. The end result may lead to damage or to loss of tissue specimens. | Software design | Leica Microsystems sent a Urgent Medical device Recall letter dated March 30, 2012, to all affected customers in the U.S. on the same date. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed of the potential malfunction of the device in the autorotation mode, which could lead to damage or loss of tissue specimens. The letter provided instructions on what the customer should do with the recalled product in the interim before the software fix can be installed. They were instructed to follow the steps below: 1. Remove and dispose all reagents including Paraffin and clean all bottles and Paraffin bathes.2. Conduct a "Smart Clean" (IFU, chapter 6.1.4).3. In Submenu "System Setup", here at "Program options", respectively "Station sequence", customer shall switch from "Concentration" to "Auto by RMS".4. Prestored Xylene protocols shall be used instead of autorotation protocols (IFU, chapter 5.6.2 "Time - optimized Xylene infiltration programs"). Here you find the five validated programs which need to be copied to Favorite programs (IFU, chapter 5.5.3).5. Use the Reagents, bottle set up from "Time optimized Xylene programs" (IFU, chapter 5.6.2., page 134) and implement them into submenu "Reagents" (IFU, chapter 5.2. following pages).6. In submenu "The RMS system adjusting warning thresholds rules" (IFU, chapter 5.2.2) the validated reagent thresholds need to be implemented into Submenu "Reagents" (IFU, chapter 5.2).7. Current Reagent management values are visible within Submenu "Reagent Status" (IFU, chapter 5.2.3)The customers were also informed that their Leica Microsystems representative will them to make the necessary arrangements to complete a software upgrade at no charge. The customers were also requested to sign and date the enclosed acknowledgement form and fax it back to Leica at 847-236-3747 or e-mail it to LMGRA@leica-microsystems.com within 10 working da |  | 67 | 27 units - U.S., 40 units - ROW | Worldwide Distribution -- USA (nationwide)including states of: California, Florida, Georgia, Kentucky, Massachusetts, Minnesota, Mississippi, Nebraska, New York, North Carolina, the District of Columbia; and countries of: Australia, Belgium, France, Germany, China/Hong Kong, Italy, Japan, the Netherlands, Singapore, Switzerland, Turkey and United Kingdom. | Terminated | October 10, 2012 | 194 | 67 | Computer
Z-1474-2012 | 61328 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips eCareManager System, Model 865325, software versions 3.7, 3.7.1 and 3.8. |  |  | February 22, 2012 | 2012 | Visicu, Inc. | Three versions of software for the eCareManager failed to display anticoagulants on the medications screen. The problem arose due to an unexpected change in the National Drug Data File. | Software change control | Philips notified all affected customers via a preliminary email message advising them of the issue. An URGENT - Field Safety Notice follow-up letter was mailed to all customers on February 22, 2012. This letter identified the affected product, described the problem and circumstances in which it can occur; recommend actions the customer should take to prevent risks for patients, and list the actions planned by Philips to correct the problem. For questions contact Philips eICU Customer Support 877-374-2872. |  | 12 | 12 | Nationwide Distribution (USA) including states of MA, MO, WI, AR, AZ, PA, SD, VA, and IN. | Terminated | July 06, 2012 | 135 | 12 | Computer
Z-1476-2012 | 60698 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Micro-Tech Informer Plus Silver Monitor, Model 81830. The firm name on the label is Stanley Security Solutions, Inc., Lincoln, NE. |  |  | December 13, 2011 | 2011 | Stanley Security Solutions, Inc. | Connection between the call cord and the receptacle may not make a correct connection resulting in a continuous or intermittent (false) signal to the nurse call system | Component design/selection | Stanley Healthcare Solutions, Inc. contacted affected customers via phone and email on December 13, 2011. A follow-up certified letter was sent on December 14, 2011. Customers were additionally sent and "Important Recall Notice letter on December 16, 2011 after errors in the original letters were detected. The letter identified the affected product, problem, actions to be taken, and included a reply form for the customer to complete and return. The recalling firm requested their customers to stop using the monitor and contact them for return of the unit. Contact information (phone, fax, &amp; email) was included. For question contact Micro-Tech Customer Service at 1-800-206-1045. |  | 824 | 824 units | Worldwide Distribution - USA Nationwide including the states of CT, FL, KY, MN, NC, NE, NH, NJ, NY, PA, TX and the country of Canada (5 units). | Terminated | September 24, 2013 | 651 | 824 | Not_Computer
Z-1478-2012 | 61658 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | REF 28234 Troch Nail, Long 11MM, Right 34 CM |  |  | March 30, 2012 | 2012 | Biomet, Inc. | The Trochanteric Nail, lot 486130, wasmislabeled as a right nail orientation but the contents are actually a left nail orientation | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice dated April 18, 2012 to all affected consignees. The letter identified the affected product, reason for the recall including risks involved and actions to be taken. The notice provided directions to immediately locate and return the affected product. Receipt Confirmation was requested by calling 800-348-9500 Ext 3009 or 3983. FAX Back Response Form should be completed and returned to 574-372-1683. Questions should be direct to 574-371-3009 or 574-372-3983, M-F, 8 AM to 5 PM. |  | 30 | 30 units in total | Worldwide Distribution - Nationwide (USA) | Terminated | March 08, 2013 | 343 | 30 | Not_Computer
Z-1477-2012 | 61635 | No | KPJ | Topical oxygen chamber for extremities. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EPIFLO; Transdermal continuous oxygen therapy device; single use; disposable medical device. Model Number EPIFLO-15. Mfr.By: Neogenix, LLC, dba Ogenix, 3401 Enterprise Pkwy, Suite 340, Beachwood, OH 44122. |  |  | April 16, 2012 | 2012 | Ogenix Corporation | Ogenix Initiated a voluntary recall for the EPIFLO-15 devices due to a mislabeling of the expiry date for this lot. Instead of March 2013, this lot carried May 2013 as the expiry date. | Labeling mix-ups | Ogenix sent a letter dated March 16, 2012 to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to confirm the status of the devices and ship product with the recalled lot back to Ogenix and complete the form provided with the letter. |  | 38 | 38 units | Worldwide Distribution-USA (nationwide) including the states of FL, MN, OH and the county of Turkey. | Terminated | February 19, 2013 | 309 | 38 | Not_Computer
Z-1479-2012 | 61604 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew 560P High Definition Camera Control Unit |  |  | March 29, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | May experience intermittent loss of videodue to a defective board component | Nonconforming Material/Component | Smith &amp; Nephew Endoscopy Division sent an "URGENT PRODUCT RECALL 1st NOTIFICATION" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory and to complete and return the bottom portion of the letter via fax at 405-917-8650. Contact the firm at 405-917-8556 for questions regarding this notice. |  | 239 | 239 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Canada, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, and UK Asia-pacific: Australia, China, Hong Kong, India, Japan, Singapore, South Korea, and Thailand, Mexico, Puerto Rico, South Africa, Turkey, and Venezuela. | Terminated | January 13, 2016 | 1385 | 239 | Not_Computer
Z-1480-2012 | 61653 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems Winged Infusion Set, 22Ga, 1", |  |  | April 16, 2012 | 2012 | Churchill Medical Systems, Inc. | Mislabeled: MBS-2210 code on case label -22Ga Winged Infusion Set, ; however, the product label states 20Ga MBS-2010 code listed on devicelabel. | Labeling False and Misleading | The firm, VYGON, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 16, 2012 to its customers via overnight, courier service. The letter described the product, problem and actions to be taken. The customers were instructed to check their stock for the products included within the scope of the recall; count and document their affected inventory; contact Vygon's Customer Service Department at 1-800-473-5414 for return merchandise authorization number; return the affected product; complete and immediately return the Acknowledgement and Inventory Return Form via fax to 215-672-6740 and send the original form with their shipment . Distributors were instructed to communicate this recall to their customers and provide them with the Acknowledgement and Inventory Return Form for completion and return.If you have any questions, please contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by e-mail at customerservice@yygonus.com. |  | 552 | 552 units | Worldwide distribution: USA (nationwide) including states of: IN, KS, KY, NY,MA, MI , and TX; and country of: Canada. | Terminated | March 28, 2013 | 346 | 552 | Not_Computer
Z-1485-2012 | 61602 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | VALOR(TM) TTC FUSION, INSTRUMENT KIT, REF 4150-KIT1, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | April 05, 2012 | 2012 | Wright Medical Technology Inc | The possibility exists that the handle of the targeting guide may rotate and/or break under normal usage. | Device Design | Wright Medical Technology, Inc. sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated April 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their stock and to return the affected product to the firm. Additionally, a Product Recall Verification Form was enclosed for customers to complete and return via fax to 901-867-7401. Contact the firm at 901-867-4324 for questions regarding this notice. |  | 133 | 133 units | Worldwide Distribution-USA (nationwide) and the country of Australia. | Terminated | January 28, 2013 | 298 | 133 | Not_Computer
Z-1481-2012 | 61411 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coloplast, Restorelle DirectFix, Posterior Polypropylene Mesh, Catalogue Number 501460, Rx Only, Manufacturer, Coloplast A/S 3050 Denmark |  |  | March 02, 2012 | 2012 | Coloplast Manufacturing US, LLC | Coloplast is conducting a field correction based on the conclusion that two lots of Restorelle DirectFix Posterior mesh were mislabeled. The incorrect size was printed on the patient label. Specifically the patient label depicts the size as 16x11 cm, however, the correct size is actually 14x15 cm. It is important to note that both the product and box labels were correct. | Error in labeling | Coloplast sent a notification letter dated March 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed their coloplast representative is providing replacement device (s) and new patient labels. For further questions please call 1 (800) 258-3476. |  | 134 | 134 | USA ( nationwide) including the states of AL, AZ, AR, CA, CO, FL, GA, ID, IL, KY, MD, MS, MO, NY, PA, TN, VA,WA and WI. | Terminated | October 11, 2012 | 223 | 134 | Not_Computer
Z-1488-2012 | 61464 | No | EQN | Nasopharyngoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | BOMImed Reusable Integrated Fiber Optic Laryngoscope Blades, MacIntosh and Miller styles - all sizes. Instrument is used to aid with intubation. |  |  | December 08, 2010 | 2010 | Bomimed Inc | On 9/1/2010 BOMImed, Inc. Winnipeg, MB, Canada initiated a recall of their BOMImed Integrated Reusable Laryngoscope Blades (Macintosh &amp; Miller styles - all sizes), Model #OL-32D0 to OL-32D4 &amp; OL-32E00, OL-32E0, OL-32E1, OL-32E2, OL-32E3, OL-32E4. | Process design | BOMImed sent a "Medical Device Safety Alert &amp; Correction Action" letter by Fax and Mail on 11/25/2010 to the consignee. The product issue was described and recommended actions to mitigate risk were provided. Contact number is 1-800-667-6276 ext. 234. |  | 30 | 30 units | US Distribution to the state of California. | Terminated | May 08, 2012 | 517 | 30 | Not_Computer
Z-1486-2012 | 61545 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens-branded linear accelerator with a 550 TxT Treatment Table with a serial number &lt;_ 1894. Siemens brand ARTISTE, ONCOR and PRIMUS Linear Accelerator systems include the 550 TxT" Treatment Table. |  |  | April 03, 2012 | 2012 | Siemens Medical Solutions USA, Inc | SIEMENS has received reports where a dropping down of the 550 TxT" Treatment Tablewas observed while the table was being moved downwards during operation by means ofthe hand control. | Software design | Siemens sent Customer Information letters to all affected customers beginning April 3, 2012 via Siemens Field Service Engineers. The letter identified the affected product, what is the issue and what has Siemens done to address this issue. The letter included an Update Instructions TH006/12/S - TxT Table Software Update VD00A. SIEMENS has arranged for a modification of the affected system. A SIEMENS Service Representative will ensure that the new software version VD00A is installed on affected system. The letter instructs customer to please include this Customer Information in their Digital Linear Accelerator System Owner Manual chapter Safety Advisory Letters where it should remain. |  | 624 | 624 active devices | Worldwide Distribution - USA (nationwide) including P.R and the countries of: Africa, Angola, Argentina, Australia, Austria, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Russian Federation, Saudi Arabia, Serbia, South Thailand, Sri Lanka, Spain, Switzerland, Trinidad, Tobago, Turkey, Ukraine and United Kingdom, . | Terminated | March 14, 2013 | 345 | 624 | Computer
Z-1489-2012 | 61613 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | EG-3870UTK, Pentax Ultrasound Video Gastroscope. |  |  | March 14, 2012 | 2012 | Pentax Medical Company | Endoscope ultrasound balloon evacuation tube may become clogged with patient body fluid if users fail to properly follow instructions for pre-procedure check, operating the suction valve during procedure and pre-cleaning procedure. | Equipment maintenance | Pentax sent an "URGENT MEDICAL DEVICE RECALL" letter on March 14, 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. A response form was included for customers to complete and return to the firm. Contact the firm at 800-431-5880 ext. 2064 for questions regarding this notice. |  | 194 | 194 units | Nationwide Distribution. | Terminated | July 21, 2015 | 1224 | 194 | Not_Computer
Z-1483-2012 | 61656 | Yes | NIK |  |  | Cardiovascular | PMA | Cardiac Resynchronization Therapy (CRT) Leads |  |  | April 03, 2012 | 2012 | St Jude Medical CRMD | Visual observations of externalized conductors due to abrasion of the silicone insulation in the distal portion of these QuickSite and QuickFlex leads. | Device Design | A recall communication was initiated on 4/3/12 with St. Jude Medical forwarding a Medical Device Advisory to all doctors who have received the St. Jude QuickSite and QuickFlex Left Ventricular CRT leads models 1056T, 1058T, 1156T, and 1158T. The letter provides the doctors with the problems identified and actions to be taken. Doctors with any questions or concerns are instructed to contact their local St. Jude Medical representative or their Technical Services Department at (800) 722-3774.A Physician Device Advisory Notice for the QuickSite and QuickFlex Left Ventricular CRT Leads was issued on 4/3/12 to physicians who have received the recalled products. The notice provides the physicians with the problem identified and actions to be taken. Physicians are instructed to return any packaged product in hospital inventory to SJM Sales Representatives. For implanted product, due to the nature of the issue, externalized conductors may be observed with no electrical or performance anomalies. Therefore, prophylactic explant is not recommended. It is recommended that physicians continue to monitor their patients' implanted system at regularly scheduled intervals. A Medical Advisory and Inventory Management Information letter was sent on 4/3/12 to hospital administrators to notify them of the problems identified and the actions to be taken. |  | 101000 | 101,000 units (59,000 units in US) | Worldwide Distribution -- Algeria, Argentina, Australia, Austria, Bahrain, Balkan, Baltic, Belarus, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Mauritius, Montenegro, Morocco, Netherlands, New Caledonia, New Zealand, Norway, Oman, Paraguay, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Senegal, Serbia, Serbia Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, UK, United Arab Emirates, US, Venezuela, and Yemen. | Terminated | July 17, 2012 | 105 | 101000 | Not_Computer
Z-1492-2012 | 61413 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Optima MR450W with GEM. |  |  | April 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare is conducting a recall on Optima MR450W with GEM. During scanning of a patient, if an improperly torqued gradient connection exists, it may cause the gradient lead to produce excessive heat. If a second and compounding issue of improperly routed Rear Pedestal Air Supply Hose is also present, routed in such a way to cause the hose to come in contact with the overheated gradient lead | Nonconforming Material/Component | GE Healthcare is not sending a Communication Letter to customers. There is no plan to contact customers with an affected Optima MR450W with GEM system. A GE Healthcare Field Service Engineer will be dispatched to sites to inspect and correct all potentially affected systems. As an effectiveness check, GE Healthcare will track all service requests for the field action to completion. For any questions or additional information, call 262-548-2608. |  | 41 | 41 | Worldwide Distribution - USA (Nationwide) including the states of :CO, FL, IA, MA ,MN, NJ, NY, NC, OH, TX. and the countries of BRAZIL., BRITAIN, FINLAND, FRANCE, GREAT GERMANY, ITALY, SAUDI ARABIA, SWEDEN, SPAIN, KUWAIT, and SWIITZERLAND. | Terminated | August 09, 2012 | 119 | 41 | Not_Computer
Z-1491-2012 | 61680 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeShield Primary Plumset, CLAVE port, CLAVE Y-Site, 104 inch, 15 drops/mL, 48 sets per case; a sterile Rx device; Hospira Inc, Lake Forest, IL USA. Made in Costa Rica; product 120300412, list number 12030-12. |  |  | April 25, 2012 | 2012 | Hospira Inc. | The Lifeshield Primary Plumsets may cause a cassette test failure alarm when the set is loaded onto the infusion pump. A set which results in a cassette test failure alarm may result in delay in therapy, requiring medical intervention. | Process control | Hospira, sent an "URGENT DEVICE RECALL" letter dated April 25, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to check their inventory and immediately quarantine any affected product to return for credit. A Reply Form was enclosed for customers to complete and return via fax to 1-888-386-2073 or e-mail to: Hospira8503@stericycle.com. Contact Hospira Customer Care at 1-877-946-7747 for information regarding product availability. |  | 145875 | 145,875 sets | Worldwide Distribution - USA (nationwide) including the states of: Alabama, Colorado, Georgia, Illinois, Indiana, Iowa, Kentucky, Maryland, Massachusetts, Michigan, Minnesota, Missouri, North Carolina, North Dakota, Oklahoma, Pennsylvania, Wisconsin and the country of Australia. | Open, Classified |  |  | 145875 | Not_Computer
Z-1487-2012 | 53340 | No | FFX | Gastrointestinal motility monitoring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Insight rev H software. |  |  | February 23, 2009 | 2009 | Sandhill Scientific, Inc | Software error may produce incorrect esophageal contraction readings. | Software change control | Sandhill Scientific sent a Urgent Product Recall letter dated February 23, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were informed All Suite revision H, BioVIEW Analysis Software rev 5.5.1 are affected. Replacement of the Suite revision H, BioVIEW Analysis Software rev 5.5.1, will be provided at no cost to you. Sandhill Scientific will pay for shipping/handling charges associated with replacement of this software.We anticipate that by the first week of March, we will be able to provide you with the Software that resolves this problem.For further questions please call toll free 800-468-4556. |  | 60 | 60 programs | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, FL, IA, IL, MA, MD, MI, MN, NC, ND, NE, NM, NY, OH, OR, PA, SC, SD, TX, UT, VA and WA and the countries of Belgium, Canada, Malaysia, Mexico, New Zealand, South-Korea, Switzerland and United Kingdom. | Terminated | May 07, 2012 | 1169 | 60 | Computer
Z-1493-2012 | 61580 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Linear Accelerators in combination with 160MLC (Multi Leaf Collimator). Linear Accelerators that may be affected are: ARTISTE MV System; ONCOR Avant Garde ONCOR Impression ONCOR Impression Plus ONCOR Expression. Manufactured by Siemens AG, Medical Solutions, Kemnath, Germany; Registered by Siemens Medical Solutions, Concord, CA. ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | April 03, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety risk related to the behavior of the 160 MLC when using the "Motion Stop" button followed by "Reset" several times during the treatment of one beam or one segment (issue 1). Additionally, there may be a potential for collision if the operator does not ensure clearance of the extended flat panel positioner for all gantry rotation angles (issue 2). | Labeling design | Siemens sent a Safety Advisory Notice with Update to all affected customers. The notice informed the customers of the risk of pressing the Motion Stop button repeatedly during treatment. Work Steps were provided for customers to follow. |  |  |  | Worldwide Distribution - USA including the states of CA, DE, FL, IL, KY, LA, MA, MI, NE, NJ, NY, OH, PA, SD, TN, UT, WI, and WV | Terminated | July 25, 2012 | 113 |  | Not_Computer
Z-1496-2012 | 61709 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Doctors Research Group Kryptonite Bone Cement packaged in 5cc and 10 cc Kits: |  |  | April 06, 2012 | 2012 | Doctor's Research Group, Inc. | Bone cement strength and stiffness fall below specification at body temperature | Component design/selection | Doctors Research Group (DRG) issued a Voluntary Product Recall Notification letter dated April 2012 to all affected customers via email or Federal Express. The letter identified the affected product, problem and actions to be taken. Customers were instructed to identify, segregate and return any affected products in their possession. Customers were asked to fill out and return the enclosed verification form Distributors were requested to notify their customers. For questions contact DRG at 203-262-9335. |  | 3100 | 3100 kits | USA Nationwide Distribution | Terminated | July 14, 2015 | 1194 | 3100 | Not_Computer
Z-1499-2012 | 61610 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Dried Overnight Gram Negative Panels; |  |  | June 16, 2011 | 2011 | Siemens Healthcare Diagnostics, Inc. | Three lots were confirmed to not meet performance claims due to random non-biological contamination which is only apparent (precipated) after inoculation of the panels | Process control | Customer Notification letters were sent to all customers of the three MicroScan panels by traceable mail. The letter identified the affected product and asked customers to discard any remaining inventory in addition to contacting customer service for replacement product. The letter also recommended that customers consider the need to review previous test results, conduct patient follow-up, and or repeat tests as needed. Customers were to complete and return the customer reply card to confirm that they received and understood the recall communication. |  | 683 | 683 boxes (20 panels per box) | Worldwide Distribution -- US, Turkey, Poland, and Great Britain. | Terminated | May 08, 2012 | 327 | 2050 | Not_Computer
Z-1500-2012 | 61646 | No | LJY | Mumps virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | AtheNA Multi-Lyte MMV Test System- A93111G |  |  | March 08, 2012 | 2012 | Zeus Scientific, Inc. | Select lots of AtheNA Multi-Lyte MMV Test System have lost reactivity and may produce false positive results in seronegative patients who have not been previously affected with or vaccinated against Mumps. | Nonconforming Material/Component | The firm, Zeus Scientific, sent an "URGENT NOTICE; Voluntary Product Recall" letter dated March 8, 2012 their customers via Fed Ex. The letter described the product, problem and actions to be taken. The customers were instructed to remove and discard any remaining inventory of the listed product; if you are not the end-user and/or you have distributed these products to customers, notify them accordingly; and complete, sign and return the verification notice via fax to: 908-526-2058 or email a PDF to support@zeusscientific.com.If you have any questions, call 908-526-3744 or 800-286-2111. |  | 180 | 180 kits | Nationwide distribution. | Terminated | July 03, 2013 | 482 | 180 | Not_Computer
Z-1495-2012 | 61693 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | True Beam; |  |  | April 09, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified in TrueBeam when using respiratory gating with patients in the prone position. The gating curve is inverted when a prone patient plan is loaded into TrueBeam. | Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice dated April 13, 2012 to all affected customers. ***12-12-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.***The letter identified the affected product, description of problem, details and recommended user actions. Varian will provide a software version that will mitigate the issue described in the letter and Varian Customer Support will notify each affected facility when the software is available. The letter instructs customers to retained a copy of this notice along with the most current product labeling. For future reference, this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. For questions contact Varian Oncology Help Desk at 1-888-827-4265 US and Canada. |  | 423 | 207 devices ***12-12-12 AMENDED TO: 216 UNITS*** | Worldwide Distribution - USA Nationwide and the countries ofAustralia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom | Terminated | April 16, 2014 | 737 | 423 | Computer
Z-1494-2012 | 61578 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens-branded ONCOR" or PRIMUS" linear accelerator with an OPTIVUE" or MVision" imaging system; |  |  | April 04, 2012 | 2012 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Siemens-branded ONCOR or PRIMUS linear accelerator with an OPTIVUE or MVision imaging system that has the potential collision problem of the imaging panel if the panel does not retract before the gantry rotates; this may result in patient injury. | Process control | Siemens sent an Update Instruction containing the Customer Information letters beginning April 4, 2012 delivered by Siemens customer Service Group or by certified mail to all affected customers. The letter identified the affected product, what is the issue and what has Siemens done to address this issue. The letter instructs customer to include this Customer Information in their Digital Linear Accelerator System Owner Manual chapter Safety Advisory Letters where it should remain. |  | 73 | 73 units | Worldwide Distribution - USA Nationwide | Terminated | September 17, 2012 | 166 | 73 | Not_Computer
Z-1504-2012 | 53828 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Caridian BCT COBE Spectra System Blood Warmer, Catalog number 950000000. |  |  | October 26, 2011 | 2011 | CaridianBCT, Inc. | Potential for air entry into the fluid pathway of the COBE Spectra due to a leak at the return luer connection to an elevated bloodwarmer when directions for use are not followed. | Device Design | Caridian BCT sent a letter to customers dated June 2011. It was a Field labeling addendum to the COBE Spectra Apheresis System Operator's Manual stating: "When connecting a blood warmer tubing set to the return line, ensure that the tubing connection is tight. Put the luer connection no higher than 20 in (50 cm) above the return access to prevent the possibility of air entering the tubing." This correction was distributed to all consignees and is included with the current version of the owner's manual. |  | 2459 | 2459 machines | Worldwide distribution including USA, Canada, Algeria, Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Netherlands, UK, Hong Kong, Hungary, India, Italy, Iran, Ireland, Israel, Jordan, Korea, Kuwait, Lebanon, Libya, Luxembourg, Egypt, Finland, France, Germany, Japan, Malaysia, Mexico, Norway, Oman, Pakistan, Philippines, Poland, Portugal, Romania, Russia, Spain, Saudi Arabia, South Africa, Sweden, Switzerland, Taiwan, Thailand, Ukraine, Uruguay, Vatican, Vietnam, and Greece. | Terminated | May 11, 2012 | 198 | 2459 | Not_Computer
Z-1503-2012 | 61384 | No | GHM | Sickle cell test. | Hematology | Hematology | 510(k) | Sickle- Scan Hemoglobin S Cell Reagent Set, Catalog No. 13-2101 100 Tests For the Determination of Hemoglobin S Two bottles Sickle Cell Solubility Buffer, 100 ml , 2 vials Sickle Cell Solubility Powder, Size 100 tests, Store 20-30 C, Distributed by True-Medix Diagnostic Brooklyn, NY 11228, 800-952-3349. |  |  | February 23, 2012 | 2012 | MichClone Associates, Inc. | Two complaints concerning the powder vial included in the 100 test kit has been observed producing pressure while sealed. When the powders are opened, the pressure from within the vial releases a small amount of powder into the air.This powder. if released near the face and inhaled can cause minor irritation resulting in sneezing and nasalcongestion. Investigation found no affect on the perf | Nonconforming Material/Component | Michclone Associates, Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify affected lots, notify sub accounts, and complete and return the response form to arrange return of unused product and replacement . Forms can be returned by FAX 248(307)-9498, mail: Michclone Associates Inc. 680 Ajax Drive, Madison Heights, MI 48118 or Email: toben@greatlakesdiag.com. Questions and concerns should be addressed to Technical Support Department at 248-307-9494. |  | 400 | 400 kits | Nationwide Distribution. | Terminated | December 13, 2012 | 294 | 770 | Not_Computer
Z-1505-2012 | 61675 | No | EBZ | Ultraviolet activator for polymerization. | Dental | Dental | 510(k) | DISPOSA-SHIELD. The Disposable Curing Shield For Your Curing Light. The disposable Disposa-Shield are sold in different sizes from 6 mm, 8 mm, 9 mm, 11 mm and 13 mm. |  |  | April 23, 2012 | 2012 | O'Ryan Industries Inc | Disposa-Shield Disposable Curing Shield is recalled due to lacking 510K, PMA or IDE, and inadequate quality system. | No Marketing Application | The firm, O'Ryan Industries Inc., began to contact consignees by telephone and sent email on April 25, 2012. O'Ryan sent a recall notification letter dated May 7, 2012 to their customers through the Purchasing Agents and Order Department Personnel. The letter described the product, problem, and actions to be taken by the customers. Customers were advised to return the Disposa-Shield back to O'Ryan Industries.Customers can call O'Ryan Industries at 800-426-4311 or fax to 360-892-6742. |  | 11376 | 11,376 units | Nationwide distribution: USA including states of NC, MN, OR and WI. | Terminated | August 10, 2012 | 109 | 11376 | Not_Computer
Z-1520-2012 | 61577 | No | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Product is labeled in part: "***is immunosimplicity***HSV 1&amp;2 IgM***Enzyme Immunoassay Test Kit***For the Detection of HSV 1&amp;2 IgM antibodies***Catalog No: 720-350***Exp. Date:***Contains Sufficient for 96 Tests***C - Corrosive Harmful***Lot No: ***For In Vitro Diagnostic Use***Caution, consult accompanying documents.***Manufacturer***Diamedix***A Subsidiary of IVAX Diagnostics, Inc. MIAMI, FL 33127 USA-1-800-327-4565***Biohazard***" |  |  | January 16, 2012 | 2012 | Diamedix Corporation | On 01/16/20112 Diamedix initiated a recall of IS-HSV 1&amp;2 IgM Test Kit, Catalog # 720-350, Lot 60701, Exp. June 30, 2012. The Positive Control for this lot is Out of Specification. | Nonconforming Material/Component | Diamedix sent Recall notification letters dated January 16, 2012, to all affected customers by Certified Mail. Consignees were asked to review their inventory and then contact Diamedix to request replacement product. Customers were instructed to call 305-324-2314 with questions. For questions regarding this recall call 305-324-2560. |  | 442 | 442 kits | Nationwide Distribution and Puerto Rico. | Terminated | May 18, 2012 | 123 | 442 | Not_Computer
Z-1508-2012 | 61713 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Product is labeled in part: Inner Label "***SMART Control Transhepatic Biliary Stent***REF Cat. No.***Lot No.***Use By***Cordis S.M.A.R.T. CONTROL Nitinol Stent Transhepatic Biliary System***Assembled in Mexico***Cordis Corporation 14201 North West 60th Avenue Miami Lakes, Florida 33014, USA***Cordis S.M.A.R.T. CONTROL BILIARY***Cordis a Johnson &amp; Johnson company***Endovascular***" |  |  | April 09, 2012 | 2012 | Cordis Corporation | Cordis Corporation is recalling 32 lots of Cordis S.M.A.R.T. Nitinol Stent Transhepatic Biliary System and S.M.A.R.T. Control Nitinol Stent Transhepatic Biliary System. There is a potential for a sterility breach due to holes or tears in the Tyvek packaging. | Packaging process control | Cordis Corp sent a " Urgent Medical Device Recall (Removal) letter dated April 9, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A Response Form was provided with each notification. Customers were asked to complete and return Response Form. The customers were instructed to :" Immediately identify and set aside all product listed above in a manner thatensures the affected product will not be used." Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form." Either return any affected product with reference to the RA# on the form, orcontact your local sales representative, or contact Customer Service (see Available Assistance section) to facilitate return of the affected product.Replacement product will be provided in a short time frame. Credit will beprovided for customer-owned inventory upon request." Pass on this notice to anyone in your facility that needs to be informed." If any product listed below has been forwarded to another facility, contact thatfacility to arrange return." Maintain awareness of this notice until all product listed above has beenreturned to Cordis.For questions related to the issue or product replacement, please contactyour local sales representative, or contact the special Customer Servicenumber (800) 551-7683, Monday through Friday from 8:30 AM to 8:00 PMEDT. |  | 278 | 278 units | Worldwide Distribution-- USA (nationwide) and the country Canada. | Terminated | December 24, 2013 | 624 | 1262 | Not_Computer
Z-1521-2012 | 61685 | No | JXE | Nerve conduction velocity measurement device. | Neurology | Neurology | 510(k) | Ultra-DriveTool, 6 MM Plug Puller, REF 423868. Recommended Power Setting 70 For Single Use Only, Sterile, Biomet Orthopedics, Warsaw, IN 46581 |  |  | April 18, 2012 | 2012 | Biomet, Inc. | Incorrect product picture label was on the product box. | Process design | Biomet Orthopedics sent an Urgent Medical Device Recall Notice dated April 18,2012, to all affected customers. The letter identified the reason for the recall including directions to immediately locate, discontinue use of the device, and return product to Biomet, 56 East Bell Drive, Warsaw, IN 46580. Consignees were asked to confirm receipt of the notice by calling 1-800-348-9500 extension 3983 or 3009. A FAX Back Response Form was to be filled out and sent within 3 days of receipt to 574-372-1683. Questions related to the notice should be directed to 474-372-3983 or 574-371-3009, Monday- Friday 8 a.m. to 5 p.m.For questions regarding this recall call 574-371-1021. |  | 30 | 30 units | Worldwide Distribution - USA (nationwide) and Europe, Austrailia | Terminated | April 18, 2013 | 365 | 30 | Not_Computer
Z-1509-2012 | 61630 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ARCHITECT STAT Troponin-I Reagent Kit; list numbers 2K41-28 (500 test kit); an in-vitro diagnostic test kit; each kit contains bottles of microparticles, conjugate and assay diluent; Abbott Laboratories, Abbott Park, IL 60064 |  |  | April 11, 2012 | 2012 | Abbott Laboratories | The lot of Troponin-I is demonstrating a shift in expected results in some cases, and varies from kit to kit. A concentration change of up to +/-55% could occur when switching kits without recalibrating. | Material/Component Contamination | The firm, Abbott Laboratories, sent a "Product Recall Immediate Action Required"letter dated April 11, 2012 to all Troponin-I customers who received the affected reagent lot internationally. The customers were directed to determine if they are currently using and/or have inventory of lot 74264UN11; to discontinue use and destroy any remaining inventory of the lot; to follow their facility's policies and procedures for communicating with the heath care providers they serve; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; to retain a copy of the letter for their records and to complete and return the Customer Reply Form (sent to Abbott international affiliates via e-mail on 4/11/2012) via fax to: 1-800-777-0051 or e-mail: QAGCO@abbott.com. Any questions were directed to their local Customer Service representative or call the field action coordinator at 847-938-1923 or email: albert.chianello@abbott.com. |  | 8723 | 8,723 kits | International distribution only: including countries of: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bosnia &amp; Herze, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Mali, Malaysia, Mauritania, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela and Vietnam. | Completed |  |  | 8723 | Not_Computer
Z-1516-2012 | 61661 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Acute care(TM) NT - proBNP TestPak reagent (CPBNPM) |  |  | April 05, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed a low frequency of calibration failures due to slope errors that are not resolved with routine maintenance with the specific SCS Pro-BNP lots listed. This issue only causes the slope error failures but does not affect patient results when the calibration is within specifications. | Component design/selection | The firm, SIEMENS, sent a "Urgent Field Safety Notice" dated March, 2012 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to discontinue use of the lots once revised expiration date occurs; contact their local Siemens Customers Service Center at 888-588-3916 for replacement; complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the TECHNICAL SOLUTIONS CENTER at (302) 631-8467 as soon as possible, and forward this notification to anyone to whom they may have distributed this product. The remaining inventory at the firm was destroyed.If you have any technical questions regarding this information, please contact the Siemens Technical Solutions Center at 800-405-6473. |  | 1899 | 1,899 cartons | Nationwide distribution: USA including states of: AK, AR, CA, CO, CT, DC, FL, HI, ID, IN, IL, IA, KS, KY, LA, ME, MI, MN, MS, MO, MT, NE, NH, NM, ND, NY, OH, OK, OR, PA, SD, TX, UT, VT, WA, WV, WI, and WY. | Terminated | March 19, 2013 | 348 | 1899 | Not_Computer
Z-1514-2012 | 61481 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental Trabecular Metal Implant, TMM4B11, 4.1mm x11.5 MM; 3.5 mm implant TM, MTX Rx, Sterile R, Zimmer Dental, Carlsbad, CA 92008, USA |  |  | March 22, 2012 | 2012 | Zimmer Inc. | This Zimmer Dental voluntary device recall resulted from an investigation into the February 2012 complaint involving an apical tip of a 4.1mm D Trabecular Metal Implant which separated from the implant assembly during surgery on a patient with a dense (Type D1), thick, inferior border. The recall correction includes adding precautions to the Instructions for Use regarding the placement of a 4.1mm | Labeling Change Control | Zimmer sent an Urgent Medical Device Correction notification letter dated April 24, 2012, with an attached Technical Bulletin listing precautions present in the new Informaiton for Use sent to distributors and customers via E-mail and FED EX. Notifications identified the issue and risks found with the device listing responsibilties, precautions, and procedural precautions. Notifications contained directions to forward information to any sub accounts and clinicians. Zimmer had telephone follow-up with each distributor and customer to assure information was received and is understood. Questions should be directed to Technical Service (800)8511-7019 orCustomer Service(760) 929-1300 or Regulatory Affairs at 1-800-854-7019 .For questions regarding this recall call 574-372-4807. |  | 233 | 233 | Worldwide Distribution - USA (nationwide) and the countries of Germany, Italy, France Egypt, Spain, Russia, Bulgaria, Turkey, Israel, UAE and CANADA | Terminated | December 03, 2012 | 256 | 1209 | Not_Computer
Z-1506-2012 | 61517 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Siemens Axiom Sensis XP diagnostic programmable computer |  |  | February 22, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction in the connection of a Sensis EP system with Carto(R) 3 dedicated cable set (part number 14411343) with the Carto 3 system from Bionsense Webster. | Nonconforming Material/Component | Siemens sent a recall/correction letter dated February 22, 2012, via Update Instruction AX049/11/S of this product to all affected customers. The letter informed customers of the potential issues and provides additional information on the implemented resolution. An Update Instuction AX048/11/S will be available immediately. The update instruction provides a new version of the Carto 3 addendum to the Operator Manual and the exchange of the previous stim cables with the latest cable.For questions call 610-219-4834. |  | 3 | 3 | Nationwide Distribution including FL, NC, and SC. | Terminated | January 24, 2014 | 702 | 3 | Not_Computer
Z-1519-2012 | 61522 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Instrumentation Laboratory (IL)HemosIL von Willebrand Factor Activity Assay Kit |  |  | April 11, 2012 | 2012 | Instrumentation Laboratory Co. | Falsely elevated levels may be reported with the HemosIL von Willebrand Factor Activity assay. | Nonconforming Material/Component | Instrumentation Laboratories mailed Recall Letter via USPS to US and Canadian customers on April 11, 2012, through Beckman Coulter. The customer mailing and tracking of the Urgent Recall Notification and Recovery Procedure in the United States and Canada will be handled through the distributor, Beckman Coulter. Accounts are requested to check inventory for the product lots of Hemosll von Willebrand Factor Activity, PN 0020004700, listed in the table and document their destruction on the providedtracking form and return the completed form. Questions or concerns regarding this notification, please call Beckman Coulter Technical Service at (800) 526-7694.On March 29, an e-mail communication was circulated to the IL Affiliates in Europe and Japan requesting them to translate the Urgent Field Safety Notification [title of document outside US and Canada] and product destruction tracking form, mail the notification to their customer base and notify theirCompetent Authorities..Fpr questions regarding this recall call 781-861-4467. |  | 9320 | 9320 Kits | Worldwide Distribution - USA (nationwide) and the countries of Canada, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELARUS BELGIUM, BRAZIL, CHILE, CHINA, CZECH REPUBLIC, DENMARK FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAN (ISLAMIC REPUBLIC OF), ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF KUWAIT LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS NORWAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TURKEY, and the UNITED ARAB EMIRATES | Terminated | August 24, 2012 | 135 | 9320 | Not_Computer
Z-1517-2012 | 61681 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 8" x 4" 16-ply X-Ray Gauze, hospital fold, banded in 10's; 200 pieces per bag, 10 bags per case; Distributed By: International Medsurg Connection, Schaumburg, IL USA 1-847-619-9926; Made in China; item IMSGZX8416-10 |  |  | April 09, 2012 | 2012 | International Medsurg Connection, Inc. | The lots of X-Ray Gauze are contaminated with foreign material, including remnants of cotton seeds and lint. | Process control | International Medsurg Connection notified their direct accounts via telephone, e-mail and/or fax on April 9, 2012, informing them of the foreign material contamination of the gauze, and requesting them to quarantine all affected lots in inventory, cease distribution of the affected product, and notify any customers at the retail level to remove and discard these products. Any questions were directed to (847) 619-9926. |  | 276 | 276 cases | Nationwide Distribution including Illinois, South Carolina, Texas and Virginia | Open, Classified |  |  | 291 | Not_Computer
Z-1541-2012 | 53196 | No | KZE | Nonelectrically powered fluid injector. | General Hospital | General Hospital | 510(k) | Trima Accel Automated Blood Collection System, catalog number 917000000, Caridian BCT, Lakewood, CO 80215 USA. |  |  | June 01, 2010 | 2010 | CaridianBCT, Inc. | Labeling clarification and software updates to include fixes for end of run summaries, increasing boot reliability, corrections to the double red blood cell count collection TBV limit, and to be in compliance with the most current FDA guidance for post donation platelet count targets. | Software design | CaridianBCT sent a Letter notification dated March 2011 to consignees The letter identified the upcoming software updates to version 6.0.1 as well as operations manual changes. For questions contact your CaridianBCT Representative, the CaridianBCT Support Center at +1-877-339-4228 or 1-303-231-4357 or your local CaridianBCT Customer Service office. |  | 2278 | 2278 units | USA Nationwide Distribution | Terminated | July 25, 2012 | 785 | 2278 | Computer
Z-1539-2012 | 52454 | No | GWG | Neurological endoscope. | Neurology | Neurology | 510(k) | Mitaka Point Setter, holding system for surgical instruments. |  |  | February 27, 2009 | 2009 | Mitaka USA, Inc. | Use of a potentially weaker spring in the device than required to meet specifications to hold the Point Setter ARM in place during surgery. | Component design/selection | Mitaka USA sent a letter dated to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter included instructions for returning the product. Contact the Customer Support Department at 800-421-0837, ext. 7577 for questions regarding this notice. |  | 18 | 18 units (3 internationally and 15 within the U.S.) | Worldwide Distribution-USA (nationwide) including the states of AZ, FL, CA, SC, OH, WA, TX, VA, MI, MO, GA, and NC and the country of Canada. | Terminated | May 22, 2012 | 1180 | 18 | Not_Computer
Z-1530-2012 | 61644 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Terumo Custom Cardiovascular Procedure Kit-PEDIATRIC PACK |  |  | April 12, 2012 | 2012 | Terumo Cardiovascular Systems Corp | Loose connections (i.e., connections between PVC circuit tubing and Oxygenator ports) may result in leaks and patient blood loss | Nonconforming Material/Component | Terumo Cardiovascular Systems issued an Urgent Medical Device Recall notification on 4/12/12 to all user facilities/consignees who have received affected devices. The letter identified the affected product, stated the reason for the recall, the potential hazard involved, and the correction. Users are to review the letter, ensure that all users are made aware of the issue, discontinue use of the affected product, and return it to Terumo CVS. In addition, users are to return the response form to the manufacturer. Affected units will be replaced with new units. Contact Terumo CVS Customer Service @ 1-800-521-2818 or tcvs.recall@teruomedical.com |  | 37 | 37 units | Worldwide Distribution -- US, including the states of AL, AZ, CA, MO, and NY and the country of Canada. | Terminated | January 23, 2013 | 286 | 843 | Not_Computer
Z-1522-2012 | 61786 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | The Evolution steam sterilizers, manufactured by Steris Corporation, utilize a loading system to assist the operator in transporting, loading and unloading items to be sterilized. It consists of two major components: |  |  | March 23, 2012 | 2012 | Steris Corporation | Through Customer feedback and field service experience, STERIS learned that operators may experience difficulty when inserting or removing equipment from the Evolution sterilizer using the transfer carriage due to the following issues with the transfer carriage. 1) Deformation of the latch pin or cracked ball bearings preventing the roller from turning.2) The rollers on the ends of the pivot a | Device Design | Steris Corporation sent an Urgent Voluntary Field Correction Notice dated March 23, 2012 via FedEx with tracking numbers for delivery confirmation to each affected customer. The letter identified the affected product, problem and actions to be taken. The letter stated a STERIS Field Service Representatives will contact each affected customer to schedule a service visit to upgrade the transfer carriages. For technical questions or questions regarding STERIS visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 523 | 523 | Worldwide Distribution - USA Nationwide including AL, AZ, CA, CO, CN, DE, FL, GA, IL, IN, KA, KY, LA, MD, MA, MI, ME, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, &amp; WY and the countries of Brazil and Canada. | Terminated | June 11, 2013 | 445 | 523 | Not_Computer
Z-1536-2012 | 61527 | No | LLI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | LUMEX A Graham-Field Brand, Bed End Sert. Low/Full-electric, Plastic Panels, Item 690-7604-000 Manufactured for : GF Heath Products, Inc., Atlanta, Georgia 30360 Made in China. |  |  | April 17, 2012 | 2012 | GF Health Products, Inc./d.b.a.Basic American Metal Products | GF Health Products, Inc., is conducting a recall on the Lumex Patriot Full Electric Homecare "Bed Ends" manufactured 3/27/09, 5/9/09, 5/30/09 and 6/20/09. The supplier deviated from the print specifications of a machined pin to a sheared pin. This deviation has the potential to provide room for the pin to move during use. If failure were to occur, the bed could descend rapidly causing non-typic | Component change control | GF Health Products, Inc. sent an "URGENT CORRECTION AND REMOVAL" letter dated April 17, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately cease use of the affected product, secure the product, and to prevent the product from further distribution to customers and end users. Contact the firm at 1-800-556-6508 for questions regarding this notice. |  | 3095 | 3095 pairs of bed ends | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WI and the countries of EL SALVADOR, CHILE, GUATEMALA, TRINIDAD, TOBAGO, and BAHAMAS. | Terminated | July 15, 2013 | 454 | 3095 | Not_Computer
Z-1537-2012 | 61636 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | HemaSyst, Dual Cannula Tip with Adapter 16 ga. x 4" (10.2 cm). Product is labeled in part:"***Pouch label***HAEMACURE***HemaSyst***tm***Contents: One ***DUAL CANNULA TIP***WITH ADAPTER***16 ga. x 4" (10.2 cm)***REF H-1604*** CE 0297***LOT 111144***2014/10***40680b/20004-02***CAUTION: U.S. Federal Law restricts this device to sale by or on the order of a licensed physician (or properly licensed practitioner).***Contents of unopened undamaged package are sterile***CE 0297***Manufactured by: ***Micromedics*** St. Paul, MN 55121-1385 USA***Distributed by:***Haemacure Corporation***Sarasota, FL 34236 USA***Toll Free 1-877-872-4583***www.haemacure.com***Box label***Contents: 5 packages of 1 each***DUAL CANNULA TIP***WITH ADAPTER***16 ga. x 4" (10.2 cm)***REF H-1604***CE 0297***STERILE EO***LOT 11144***2014/10***40681B/20544***". |  |  | November 16, 2009 | 2009 | Haemacure Corporation | This product has been recalled due to the lot number being different on the box and pouch. The lot number on the box label was incorrectly printed as 11144 instead of 111144. The pouch label referenced the correct lot number, 111144. | Labeling mix-ups | Haemacure Corporation, Sarasota, FL initiated letters via e-mail on 11/16/2009 with delivery and read receipt confirmation to their two consignees. The letters identified the affected product and stated the reason for the recall. The letter also states that the firm will replace each order upon receipt of the mislabeled boxes at no charge to the customer. Customers were asked to return their unused devices. Questions should be directed to customer service at 1-877-872-4583. |  | 7 | 7 boxes | Nationwide Distribution -- Illinois and California. | Terminated | May 14, 2012 | 910 | 7 | Not_Computer
Z-1540-2012 | 61629 | No | MTF |  |  | Immunology | PMA | Dimension Vista(R) Total Prostate Specific Antigen Flex(R) reagent cartridge |  |  | April 02, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed a report of leaking Flex(R) reagent cartridge lot 11300BB from Dimension Vista(R) TPSA; when a leaking Flex(R) is used, patient test results may be falsely depressed. | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notification letter dated April 14, 2012 to all Dimension Vista(R) System customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to complete the attached form and fax it to 302-631-8467 to indicate that you have received the notification. For technical questions or concerns, contact the Siemens Technical Solutions Center at 800-441-9250. |  | 1103 | 1,103 cartons | USA Nationwide Distribution including the states of: AL, AZ, CA, CO, CT, FL IL, LA, MD, MI, MN, MS, MT, NC, NH, NJ, NV, NY, OH, PA, SD, TN, TX, UT, WI ,WV including Puerto Rico. | Terminated | April 02, 2013 | 365 | 1103 | Not_Computer
Z-1577-2012 | 61468 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00-5976-020-17, NexGen CR Complete Knee Solution, Cruciate Retaining (CR) Articular Surface Anterior Constrained AC, Size Purple, AE Micro 17mm Height , Sterile R, Rx ,Zimmer, Inc. Warsaw, IN |  |  | February 24, 2012 | 2012 | Zimmer Inc. | Zimmer has received 5 complaints where a NexGen CR micro articular surface was used with a standard CR femur, even though the compatibility chart and product labeling indicates that these components are NOT compatible. Investigation identified additional instances where a CR micro articular surface and standard CR femur were sold together on the same purchase order. CR micro articular surfaces are | Labeling design | Zimmer sent an Urgent Correction Notice dated March 5, 2012, to all affected customers via telephone, E-mail and letter notifying users of the appropriate compatibility of micro sizes of NexGenÂż CR and CR-Flex Articular Surfaces and informing them of known cases of incompatible devices being implanted, including associated risks. Customers and surgeons were asked to acknowledge receipt by calling 1-888-912-7349. Questions and concerns or help in notiying accounts about this correction were to be addressed to 1-888-548-8514.For questions regarding this recall call 574-372-4807. |  | 83322 | 83,322 all sizes | Worldwide Distribution - USA (nationwide) and the countries of Poland, Taiwan, Denmark, Korea, Singapore, China, Canada, Australia, Japan, France, Italy, UK, Germany, Spain, Belgium, Portugal, Switzerland, Sweden, India, Thailand, Peru, The Netherlands, Malaysia, and Austria. | Terminated | April 18, 2014 | 784 | 83322 | Not_Computer
Z-1542-2012 | 61475 | exerciser, non-measuring | ION | Nonmeasuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Norco Shoulder Pulley |  |  | July 11, 2011 | 2011 | North Coast Medical Inc | Potential for plastic D-ring to break during use and cause injury. | Device Design | North Coast Medical began contacting customers of the recall on July 13, 2011 by telephone. A phone script was used to inform customers of the affected product, problem, and actions to be taken by the customers. Information was provided for customers to return affected product. Affected devices were re-worked and inspected. Contact the firm by e-mail at www.ncmedical.com for questions regarding this recall. |  | 2364 | 2364 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Switzerland, Germany, Great Britain and the Netherlands. | Terminated | September 06, 2012 | 423 | 2364 | Not_Computer
Z-1543-2012 | 61699 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PATIENT SPECIFIC INSTRUMENTS |  |  | March 23, 2012 | 2012 | Materialise USA LLC | March 23, 2012 a field representative reported patient specific -0242-L tibia guide case packaging contained a different patient specific -0864-R tibia guide. The two patient specific products had been switched in packaging. Neither products had been delivered to the end user. | Labeling mix-ups | Materialise telephoned both field representatives on March 23, 2012 and confirmed recalled product mix-up. Materialise requested both cases be returned and replacements were provided. No affected product was received by the end user. |  | 1 | 1 unit | Nationwide Distribution including the states of Ohio and Minnesota. | Terminated | January 23, 2013 | 306 | 1 | Not_Computer
Z-1546-2012 | 61500 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | RF IgA Calibrator, a component of the EL-RF/3 Rheumatoid Factor Test Kit; an in vitro diagnostic; TheraTest Laboratories, Lombard, IL 60148; component catalog #12833, kit catalog #303-305. The TheraTest EL-RF/3 is intended as an aid to the diagnosis of Rheumatoid Arthritis (RA). |  |  | March 29, 2012 | 2012 | Thera Test Laboratories, Inc. | The RF IgA Calibrator component of the EL-RF/3 Rheumatoid Factor test system may be unstable, resulting in a drop in the RF IgA calibrator OD value which could invalidate an RF IgA test run. | Nonconforming Material/Component | TheraTest sent "Urgent Device Recall" letters dated March 28, 2012 to customers that received the affected kit, requesting them to return or destroy the remainder of the product. TheraTest issued a replacement kit to all customers that had unused product as well as any material they consumed which resulted in a failed run. Customers were also instructed to complete the enclosed confirmation form and fax the completed form back to TheraTest. Any questions were directed to TheraTest at 800-441-0771. |  | 207 | 207 kits | Nationwide Distribution (US) | Open, Classified |  |  | 207 | Not_Computer
Z-1545-2012 | 61804 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Instructions For Use: "***SmartLOX***Instructions for Use Cervical Plate System***STERILIZATION***The SmartLOX Cervical Plate System components are supplied clean and not sterile. All implants and instruments should be cleaned and sterilized prior to surgery. AORN recommended practices for in hospital sterilization should be followed. Sterilization testing of components has shown the following recommendations for sterilization are effective Method: Steam***Cycle: Pre vacuum***Temperature 270 F (132 C)***Exposure Time: 3 minutes***Contact Information: Captiva Spine Inc. 967 Alternate AA #1 Jupiter, FL 33477. Tel: 877-772-5571***CAPTIVA SPINE***". |  |  | March 22, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine has initiated a correction of the SmartLOX Instructions for Use (IFU) for SmartLOX Cervical Plate System. The pre vacuum sterilization cycle [270 (132 C) for 3 minutes] has not been validated. | Labeling False and Misleading | Captiva Spine sent a "VOLUNTARY CORRECTION NOTIFICATION" letter dated March 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides steps for the correction and a verification form for customers to complete. Customers should contact the firm at 877-772-5571 for questions concerning this notice. |  | 14 | 14 units | Nationwide Distribution including the states of FL, IN, UT, GA, PA, TX, AL, NC, MI, and MN. | Terminated | November 18, 2013 | 606 | 14 | Not_Computer
Z-1612-2012 | 61418 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical Naturalyte Acid Concentrate for bicarbonate Dialysis 45x , 55 gallons (208.2 Liters) |  |  | March 16, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Sodium content does not meet product requirements | Process control | Fresenius Medical Care North America sent an "URGENT PRODUCT RECALL" letter dated March 28, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to examine their stock immediately to determine whether they have any affected product. A fax back form was included for customers to complete and return. Contact FMCNA Customer Service Team at 1-800-323-5188 for questions regarding this notice. |  | 154 | 154 units | Nationwide Distribution. | Terminated | January 28, 2014 | 683 | 154 | Not_Computer
Z-1614-2012 | 59861 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Coseal Surgical Sealant. |  |  | July 28, 2011 | 2011 | Baxter Healthcare Services L.P. | The recall is being conducted as a precautionary measure due to out of specification results observed at the 24 months time point (out of 24-month licensed shelf life to expiration) during a Coseal stability study. The parameter that is out specification is an indicator of possible failure for the product to gel appropriately . Coseal's failure to gel does not represent risk for the patient's life | Device Design | Baxter Healthcare sent an"URGENT PRODUCT RECALL" letter dated July 27, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately stop the usage/distribution of the affected product and to quarantine any units that are in inventory. Instructions for obtaining replacement product were attached to the letter. Contact the firm at 1-800-423-2090 for questions concerning this recall. |  | 11160 | 11160 | Nationwide Distribution. | Terminated | May 31, 2012 | 308 | 11160 | Not_Computer
Z-1615-2012 | 61719 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare, Model Numbers QC200403, v.02; QC200403, v.03; QC200403, v.04; QC200405 (USA) and QCE200403, v.02; QCE200403, v.03 (UK). Manufacturer Living Independently Group, Inc., New York, NY. QuietCare is intended for use in monitoring the environmental conditions and activity (motion) of an individual living in a senior housing community. |  |  | July 30, 2011 | 2011 | Intel-GE Care Innovations LLC | The QuietCare production system experienced a partial outage on 7/29/11. As a result of the partial outage, some monitoring data about residents' movements was not processed in the production system, and alerts were therefore not sent to caregivers. The problem was identified and corrected and the unprocessed data was recovered and processed. However, because of the time lag in processing, a fe | Employee error | Initial notification to the affected facilities was done by telephone or email on 7/30/11. This was followed up with Urgent Field Notification letters to the customers on 8/12/11 via e-mail. For questions, customers were instructed to contact Customer Support at 1 (866) 772-8243. |  | 164 | 82 units affected by the outage: Domestic - 80 units, Foreign - 2 units | Worldwide distribution - USA and the United Kingdom. | Terminated | May 21, 2012 | 296 | 164 | Not_Computer
Z-1616-2012 | 61803 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S1000 ultrasound systems running software version 1.0, Siemens Medical Solutions, USA, Inc. Mountain View, CA. The product is indicated for use as an ultrasound imaging system. |  |  | April 25, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | Software design | A field corrective action was initiated on 04/25/2012 under Siemens update program US019/12/S in which a "Customer Safety Advisory Notice" was sent to all customers notifying them of the problem and the recommended actions to avoid potential risk. A follow-up field corrective action will be initiated in which the affected software will be updated in the field. |  | 45 | 45 | Worldwide Distribution -- USA, including the states of TX, NJ, GA and the countries of Germany, United Kingdom, Republic of Korea, Australia, Japan, Poland, Israel, Singapore, France and Saudi Arabia. | Terminated | October 19, 2012 | 177 | 45 | Computer
Z-1617-2012 | 61694 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TrueBeam and TrueBeam STx V1.0, 1.5, 1.6.95 and below. TrueBeam Radiotherapy Delivery System is intended to provide stereotactic radiosurgery and precision radiotherapy anywhere in the body where radiation treatment is indicated. Varian Medical Systems, Palo Alto, CA. |  |  | April 09, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian received a report involving a Gantry collision and is sending a notification to remind users of the collision protection tools available with TrueBeam and TrueBeam STx, including actions the operator should be aware of. | Labeling design | An "Urgent Medical Device Correction / Field Safety Letter" was distributed to all affected users on April 9, 2012, with a description of the problem and user corrective action steps. The Letter was also distributed to the Varian Sales, Marketing, and Service organizations. Customers in the US and Canada may contact (888) 827-4265 for questions.*** An additional letter was sent on 11/5/12 to additional consignees identified by Varian as being affected the the recall but not previously notified. |  | 2484 | 207 devices 12/7/12 - amended to 216 devices. | Worldwide Distribution -- USA, and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Netherlands, New Zealand, Norway, Portugal, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, and the United Kingdom. | Terminated | April 16, 2014 | 737 | 2484 | Not_Computer
Z-1613-2012 | 61657 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Galil Medical, 17G Cryoablation Needle Holder FPRPR 4014, Rx only, Sterile EO. |  |  | April 12, 2012 | 2012 | Galil Medical, Inc. | Galil Medical is recalling 17G Cryoablation Needle Holder, FPRPR 4014. This voluntary recall has been initiated due to the possibility of excessive bioburden on the Needle Holders. | Material/Component Contamination | Galil Medical sent an "Urgent Medical Device Recall" letter dated April 12, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken. The letter instructs customers to return all unused affected products, use a revised needle preparation procedure and forward and communicate this Recall Notice and instructions. Customers were instructed to complete and return the attached Response Form with returned product or fax to 1+877-510-7757. For any questions contact Galil Medical Customer Service. |  | 26 | 26 kits | Worldwide Distribution - USA (nationwide) including the states of CA, NY, NC, WA and the countries of CANADA, DENMARK, FRANCE, NETHERLANDS, TAIWAN and UNITED KINGDOM. | Terminated | August 16, 2012 | 126 | 26 | Not_Computer
Z-1618-2012 | 61715 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cutting Edge Acetabular Spherical Reamer 44 M, Stryker Howmedica Osteonics Howmedica Osteonics Corp. 325 Corporate Drive Mahwah, NJ 07430 USA. |  |  | December 23, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | A 46mm acetabular reamer was lasermarked as a 44m reamer and distributed to the field. | Device Design | Stryker sent an "URGENT PRODUCT RECALL" letter dated December 29, 2011 to all affected customers. Notificationletters and Product Accountability Forms were sent to Stryker accounts via FedEx on January 3, 2012 with return receipt.Hospital Risk Management and Chief of Orthopaedics Notification letters and Product Recall Acknowledgement Forms were sent on January 10, 2012.Surgeon Notification letters and Product Recall Acknowledgement Forms were sent on January 10, 2012.Notification sent via e-mail toCanada.EMEA, LatinAmerica.Asia Pacific on December 23, 2011.The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to return the affected product. A Product Recall Acknowledgment Form was attached to the letter for customers to complete and return. Contact the firm at 201-972-2100 for questions regarding this recall. |  | 134 | 134 units | Worldwide Distribution. | Terminated | July 17, 2013 | 572 | 134 | Not_Computer
Z-1623-2012 | 61819 | No | NPL | Bone grafting material. | Dental | Dental | 510(k) | a translucent, rectangular, resorbable, collagen membrane sheet derived from bovine tissue. |  |  | February 28, 2012 | 2012 | Kensey Nash Corp | Recall has been initiated due to concerns regarding the sterility of the product. The firm initiated the recall because of use of inadequately sterilized product might result in patient infection, | Employee error | Reimser, Inc. sent an "URGENT PRODUCT RECALL-COLLAGUIDE COLLAGEN DENTAL MEMBRANE" letter dated March 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to cease use, secure and quarantine affect product. Additionally, a Recall Acknowledgement Form was attached for customers to complete and return to the firm via fax to 1-800-505-4982. Contact the firm at 484-713-2197 for questions regarding this recall. |  | 13189174161 | total all sizes: 6387 units; p/n 20650-03 - 765 units | Worldwide Distribution. | Terminated | April 11, 2013 | 408 | 13189174161 | Not_Computer
Z-1624-2012 | 61801 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Pelleve Non-Ablative Wrinkle Treatment Handpiece, Product Codes D110FSHP, D115FSHP (reusable) --- Pelleve --- Elegance through science --- Contents: One (1) Pelleve Handpiece (Reusable) --- Non-ablative treatment of mild to moderate facial wrinkles and rhytids for skin phototypes I-IV. |  |  | May 04, 2012 | 2012 | Ellman International, Inc. | Ellman International is recalling its Pelleve reusable handpiece because of possible cable failure and sparking. | Equipment maintenance | Ellman will notify all consignees initially via UPS shipment with an Urgent: Medical Device Recall letter containing notification of the recall and instructions. The letter identifies the affected product, problem, and instructs customers to identify and return the handpieces in the enclosed prepaid mailer. Customers are to complete and return the Product Recall Verification Form and fax it to the number provided. In addition, customers are provided with a toll-free number, 1-877-522-7141, to call in the event that they have questions or require additional instruction. Customers will initially be provided with a fixed-life disposable GlideSafe handpiece and may choose to receive a replacement reusable handpiece with modified (more robust) cable, or additional GlideSafe handpieces. |  | 1088 | 1,088 units | Worldwide Distribution | Terminated | February 12, 2016 | 1379 | 1088 | Not_Computer
Z-1619-2012 | 61343 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Alignment Handle |  |  | October 10, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker received a Product Experience Report (PER) that a swivel block was improperly oriented on a Triathlon Tibial Alignment Handle, catalog no. 6541-2-807. | Device Design | Stryker sent an Urgent Product Field Action letter dated October 10, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Stryker Sales Representative to arrange for return of the affected product . Customers were asked to fax back the attached Poroduct Field Action Acknowledgement Form as soon as possible to 201-831-6069. For any questions customers were instructed to call 201-972-2100.For questions regarding this recall call 201-831-3523. |  | 6126 | 6,126 Units | Worldwide Distribution - USA (nationwide) and the countries of Asia Pacific, Canada, Europe, Middle East and Africa (EMEA), and Japan | Terminated | September 10, 2014 | 1066 | 6126 | Not_Computer
Z-1625-2012 | 61859 | No | KDH | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 50 Portal Vein Cannula REF 61.110.10.030 Size: Fr 10 Length: 12 Inches LL Connector LOT 120493 NON STERILE MEDDEV USA, INC. 100 Highland Park Village/Suite 200 DALLAS, TEXAS 75205 |  |  | May 10, 2012 | 2012 | MEDDEV USA, INC. | The device was marketed without a 510(k) and is not sterile. | No Marketing Application | MEDDEV USA, Inc decided to recall and notified their customer by phone on 05/10/12. Customers were instructed to return the entire shipment to MEDDEV .For questions regarding this recall call 214-692-3559. |  | 20 | 20 units | Nationwide Distribution including Illinois | Terminated | August 15, 2013 | 462 | 20 | Not_Computer
Z-1626-2012 | 61860 | [6 empty] | 8F Plastic Dignity Mid-Sized CT Port w/Silicone Filled Suture Holes and w/Attachable ChronoFlex Polyurethane Catheter. Catalog # MRCTI8000S. The Power Injectable, Implantable Infusion Port is an implantable single fluid reservoir port with a polyurethane (8F) catheter attachable for application by the inserting physician. |  |  | April 12, 2012 | 2012 | Medical Components, Inc dba MedComp | 50 trays of Dignity Low Profile CT Port were labeled as Dignity Mid-Sized CT Port. | Employee error | On April 12, 2012, MedComp sales representatives and the foreign distributor were notified via e-mail with a "Product Alert" notification. The issue was described and requested actions were provided to recipients. |  | 50 | 50 trays | Worldwide Distribution - - USA, including the states of CA, NJ, AZ, PA, TN and the country of Spain. | Terminated | July 18, 2013 | 462 | 50 | Not_Computer
Z-1627-2012 | 61750 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The ECHELON 60 Endoscopic Linear Cutter - Straight, compact ECHELON 60 Endoscopic Linear Cutter - Straight, and the Long ECHELON 60 Endoscopic Linear Cutter - Straight are sterile, single patient use instruments that deliver staples while simultaneously dividing tissue between rows. ECHELON 60 instruments with white, blue, gold, and green reloads deliver two triple-staggered rows of staples. The instruments have a staple line that is approximately 60 mm long and a cut line that is approximately 57 mm long. A staple retaining cap on the reload protects the staple leg points during |  |  | March 02, 2012 | 2012 | Ethicon Endo-Surgery Inc | There is a possibility of a component breakage, which may cause the device to stay locked on tissue after firing. If the component exhibits this issue, the device does not open either with the anvil release button or with the manual firing release lever. | Vendor change control | Ethicon Endo-Surgery sent an Urgent Device Recall letter dated March 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have the affected product and to remove the affected product. Fill out the Business Reply Form and fax it to 1-513-337-4138 within 3 business days, even if they do not have the affected product in their inventory. Ethicon will automatically replace any affected product returned by June 29, 2012. Customers were asked to share this information with appropriate staff.If customers should need clinical or product support call 1-800-873-3636, option 6, or contact their local sales representative.For questions regarding this recall call 513-337-7000. |  | 3120 | 3120 | Worldwide Distribution - USA including AL, AR, CA, CO, FL, IA, IL, KS, LA, MD, MO, MS, NC, NM, NY, OH, SC, TX &amp; VA; and the countries of Argentina, Ecuador, Mexico, Australia, Egypt, Singapore, Belgium, Greece, Slovenia, Brazil, India, South Africa, China, Israel, United Arab Emirates, Colombia, Japan &amp; Venezuela. | Terminated | March 22, 2013 | 385 | 3120 | Not_Computer
Z-1630-2012 | 61444 | No | MKQ |  |  | Pathology | PMA | BD Focal Point"100V (Refurbished), Catalog number: 490651 |  |  | October 06, 2011 | 2011 | Tripath Imaging, Inc. | The cable tie mounts that adhere to the inside of the FocalPoint do not adhere as intended and can therefore become detached. | Component change control | BD Diagnostics first notified consignees by letter on October 6, 2011. BD Diagnostics also sent a letter dated February 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Field Representatives were advised to inspect all cables, adhesive cable ties and connections to ensure they are secure. Inspections and repairs, if necessary, would be carried out at the next service call for each instrument, or no later than the next regularly scheduled preventative maintenance visit. For questions consignees were instructed to contact BD Diagnostics Technical Support at 1-877-822-7771 option 2 in the US, their local distributor, or their local BD office for any questions.For questions regarding this recall call 918-563-1624. |  | 1 | see line 1. | Worldwide Distribution - USA including AZ, CT, IL, IN, FL, MN, NC, OK, OR, and TX and the countries of Australia, Belgium, Brazil, Canada, Japan and Thailand | Terminated | December 03, 2012 | 424 | 73 | Not_Computer
Z-1628-2012 | 61663 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Silicone Cuffed Wire Reinforced Extended Connect Pediatric Perfect Fit Neck Flange Tracheostomy Tube, Size 4.0 mm I.D. Catalog Number 353-040. |  |  | April 18, 2012 | 2012 | Arcadia Medical Corporation | A defect in 3 tracheostomy tubes. The defect observed was a separation in the bond between the proximal junction of the neck flange and the cannula. | Process control | Arcadia notified the 3 consignees with a MEDICAL DEVICE VOLUNTARY RECALL notification dated 4/18/12. The notification identified the affected product, along with the reason for the recall. Consignees were instructed to cease distribution/use of any unopened product that they have remaining and return to firm following the steps provided. Arcadia will replace the product at no charge to the customer. Consignees were to confirm receipt of the recall notification by emailing Trena Depel, trena@arcadiamedical.com.Phone: 925-918-0595Fax: (415) 762-5489 |  | 1103307 | 01103307 (50 units) and 01107055 (73 units) | Worldwide Distribution - US, including Mariemount, OH and countries of UK and France. | Terminated | August 13, 2012 | 117 | 1103307 | Not_Computer
Z-1620-2012 | 61568 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Product is labeled in part: "***is immunosimplicity***ENA-6 Screen***Enzyme Immunoassay Test Kit***Catalog No: 720-230***Exp. Date: 2012-05-31***Contains Sufficient for 96 Tests***C - Corrosive Harmful***Lot No: S0000***For In Vitro Diagnostic Use***Caution, consult accompanying documents***Manufacturer DIAMEDIX***a Subsidiary of IVAX Diagnostics, Inc.***Miami, FL 33127 USA-1-800-327-4565***Biohazard***" |  |  | June 08, 2011 | 2011 | Diamedix Corporation | On 06/08/2011 Diamedix Corporation initiated a recall of Is-ENA-6 Screen Test Kit, Catalog # 720-320 Master Lot S0000, Exp. Date May 31, 2012. The labeling on the kit box has an expiration date of 2012-05-31 while the components have an expiration date of 2012-04-30 for an incorrect expiration date. | Error in labeling | Diamedix Corporation sent a Customer notification letter dated June 8, 2011, to all affected customers via Certified Mail. Customers were instructed to search their inventory and to return affected product to Diamedix.Diamedix apologizes for any inconvenience this has caused our valued customers. Please contact the Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any kits you might still have of the lot in question. Please rest assured that Diamedix values you as a customer and is in the process of taking the proper corrective and preventive actions. |  | 312 | 312 Test Kits | Worldwide Distribution -- (USA) nationwide including the states of FL, NY, NC, MS, KY, MA and Puerto Rico and the countries of Italy, Chili and Venezuela. | Terminated | November 15, 2013 | 891 | 312 | Not_Computer
Z-1647-2012 | 48179 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | EZ STEER Catheter D-1267-04-S / 8mm Non-Navigation Bi-Directional Catheter |  |  | December 15, 2007 | 2007 | Biosense Webster, Inc. | The recall was initiated because the catheter unit was shipped with a label expiration date of 9/30/2010, which is 11 months beyond the correct expiration date of 10/31/2009. | Incorrect or no expiration date | The firm, Biosense Webster, sent a "Medical Device-Voluntary Field Removal" letter dated December 17, 2007 to its customer. The letter identified the product, problem and an action to be taken. The customer was instructed to complete and return the Product Voluntary Field Removal Certification Form via fax to Quality Compliance at (909) 839-7207or mail using the enclosed self-addressed stamped envelope. If you have additional questions, please follow up with your local Biosense Webster representative.or contact Customer Service at (909) 839-8500. |  | 1 | 1 unit | Nationwide distribution: CA only. | Terminated | May 23, 2012 | 1621 | 1 | Not_Computer
Z-1645-2012 | 61770 | No | MTG |  |  | Immunology | PMA | Access Hybritech Free PSA, Part Number: 37210 |  |  | October 11, 2011 | 2011 | Beckman Coulter Inc. | A recall was initiated because Beckman Coulter has confirmed a change to the recommended diluents for the Access Hybritech free PSA reagent. | Process design | Beckman Coulter sent a Product Correction letter dated October 11, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed regarding a change to the recommended diluents for the Access Hybritech free PSA reagent. Patient safety is not affected by this change." Do not dilute elevated free PSA patient samples with Wash Buffer II." Do not dilute elevated free PSA patient samples using the UniCel DxI onboard dilution feature." Dilute elevated free PSA samples with Access Hybritech free PSA Calibrator S0 only." A review of historical patient results is not required because there is no effect on clinical interpretation.* The UniCel DxI onboard dilution feature is only available in limited geographies." An updated version of the UniCel DxI Assay Protocol File (APF) will remove the Access Hybritech free PSA (dfPSA) onboard dilution assay from the list of available assays." The Access Hybritech free PSA IFU will be updated to remove Wash Buffer II as an acceptable diluent for elevated patient samples. Correct any procedure manuals or other documents created in your laboratory that list Wash Buffer II as an acceptable diluent for the Access Hybritech free PSA assay.Please share this information with your laboratory staff. If you have forwarded any of the affected product listed above to another laboratory, please provide a copy of the letter to them.Please complete and return the enclosed response form within 10 days so we are certain you received this important communication.If you have any questions regarding this notification, please contact our Customer Support Center" Via our website, http://www.beckmancoulter.com/customersupport/support" Via phone, call 1-800-854-3633 in the United States and CanadaOutside the United States and Canada, contact your local Beckman Coulter Representative. We apologize for the inconvenie |  | 90965 | 90,965 units total (24,951 units in US) | Worldwide Distribution -- USA (nationwide ) and the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia, Herzegovina, Brazil, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, French, Polynesia, Georgia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Libyan, Arab, Jamahiriya, Macao, Mexico, Moldova, Republic of Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Panama, Philippines, Poland, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab, Emirates, United Kingdom, Uruguay ,Venezuela, Viet Nam, Yemen and Zimbabwe. | Terminated | March 18, 2013 | 524 | 90965 | Not_Computer
Z-1646-2012 | 61809 | No | OUO | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/MRI System v.3.7.1, 510(k) #K103483, Model #882380. |  |  | April 13, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software errors. A number of non-conformances were identified and corrected by Philips Medical for their Ingenuity TF PET/MRI Systems running software version 3.7.1. The firm states that these non-conformances stemmed from software and/or documentation errors, and that software corrections are being performed in order to 'improve the overall quality and functionality of the Ingenuity TF CT/MRI S | Software design | The firm, Philips Medical, sent an 'URGENT - Medical Device Correction' notification letter dated April 13, 2012 to all affected customers via Certified mail. The letter informs the customers of the firm's upcoming software update. It tells the customers that the software for the Ingenuity TF PET/MR system unit will be upgraded from version 3.7.1 to version 3.7.2 and that the update/upgrade will provide a number of improvements to the overall quality and functionality of the Ingenuity TF PET/MR System. The notification document states that Philips will be installing the system update to version 3.7.2 software through a Field Change Order. This will be implemented free-of-charge and will include complementary PET/MR application training. For further information, call toll-free telephone number which they can call 1-800-722-9377 and follow the menu options to reach a Customer Solutions Engineer. |  | 6 | 8units (2 units in U.S. &amp; 6 units in other countries) | Worldwide distribution: USA (nationwide) including states of: OH and NY; and countries of: China, Finland, Germany, South Korea, Spain, and Switzerland. | Terminated | August 05, 2013 | 479 | 6 | Computer
Z-1644-2012 | 61746 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, Versions 4.30.00-4.64.00. |  |  | April 30, 2012 | 2012 | Computerized Medical Systems Inc | XiO will underestimate the dose delivered when the closed leaf gap is set to 0mm on XiO and a value other than 0mm is set on the accelerator. | Software design | Computerized Medical Systems, Inc. sent an Important Safety Notice the week of April 30, 2012, by regional support managers via e-mail, fax, or regular mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. North American customers were instructed to go to www.elekta.com, select the SupportPlus Login at the top of the page, enter their portal, select Downloads/Updates to download the latest patches.For Customers outside North America, go to www.elekta.com,select the SOFTWARE button on the left, Treatment PlanningSoftware, then Contact TPS Support tab, Download TPSSoftware Patches..............................................................Customers were also instructed to distribute this notice to any and all users of CMS software at their organization who are potentially affected by this issue.For questions regarding this recall call 408-380-8023. |  | 1652 | 1,652 sites received the versions | Worldwide Distribution - USA (nationwide) and the country of Canada | Terminated | April 01, 2014 | 701 | 1652 | Computer
Z-1651-2012 | 61812 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MLCi/MCLi2. |  |  | April 19, 2012 | 2012 | Elekta, Inc. | The Y2 diaphragm block was found loose. | Nonconforming Material/Component | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated April 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to carry out the recommended weekly physics check of the X-ray and light field size calibration. If an error is found, customers are instructed to contact their local Elekta service engineer. |  | 386 | 386 | Nationwide Distribution including Puerto Rico. | Terminated | October 06, 2014 | 900 | 386 | Not_Computer
Z-1648-2012 | 61811 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Accessories for the Captiva Spine SmartLOX Cervical Plate System. Products are labeled in part: |  |  | November 22, 2011 | 2011 | Captiva Spine, Inc | Captiva Spine is recalling the SmartLOX Cervical Plate components 12 &amp; 14 mm Drill Bits, Single Barrel Variable Drill Guides, and Double Barrel Fixed Drill Guides. The components are part of the instrumentation kit used to complete an anterior screw fixation of the cervical spine. There is a potential risk for the drills to not properly fit through the drill guides due to an overlap in the dime | Nonconforming Material/Component | Captiva Spine notified customers of the recall by letter on November 23, 2011, and advised that the affected products would be replaced. Customers were asked to return the drill guides and unused drills with the provided return labels.For questions customers should call 877-772-5571.For questions regarding this recall call 561-277-9480. |  | 212 | 212 units | Nationwide Distribution including TX, IN, UT, AZ, GA, CA, AL, MI, and MN. | Terminated | November 18, 2013 | 727 | 317 | Not_Computer
Z-1650-2012 | 61701 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Celsus prosthetic foot, Product Number 103658-00. The Celsus is intended for low impact level lower limb prosthesis users weighing up to 300 lbs. |  |  | April 05, 2012 | 2012 | College Park Industries, Inc. | An incorrect shorter part ( pyramid) was installed in the foot. A foot of this size and configuration requires a longer pyramid. Customers requiring the extra firm configuration may have a minimal risk of component failure, which could be implicated in a potential fall. | Mixed-up of materials/components | All customers who received affected product were initially contacted via phone on April 5, 2012 and made aware of the recall, including instructions to notify customers immediately and arrange replacements. A "Recall Notice" letter dated April 24, 2012 was sent to customers describing the problem and the recommended actions. Contact reference person is Kimberly Light, 1 800-728-7950, info@college-park.com. |  | 11 | 11 | Worldwide Distribution - USA, including the states of CA, MA, ME, WA, TX, VA, AZ and the country of England. | Terminated | July 31, 2012 | 117 | 11 | Not_Computer
Z-1653-2012 | 61732 | No | MUI | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Other-Sonic Generic Ultrasound Transmission Gel 5 Liters Hypoallergenic |  |  | March 02, 2012 | 2012 | Pharmaceutical Innovations, Inc. | Other-Sonic GENERIC ULTRASOUND TRANSMISSION GEL is being recalled due to possible microbial contamination. An FDA Safety Communication was issued for this product on April 18, 2012. | Employee error | Phamaceutical Innovations Inc. sent an Urgent Message Notice letter dated March 02, 2012 with a followup letter dated March 9, 2012 and an Inventory Withdrawal letter dated March 23, 2012 via email and fax to all affected customers. Customers were instructed to examinne inventory, refrain from distribution and return remaining inventory for a replacement. For questions call 973-242-2900. |  | 40164 | 40,164 units in distribution worldwide - unable to obtain units in distribution Nationwide | Worldwide Distribution - United States (nationwide) including the states of: FL, MA, MS, NJ, OH, TN and VA, and the countries of: Canada and Australia. | Open, Classified |  |  | 40164 | Not_Computer
Z-1656-2012 | 61754 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | ECMO Heaters Model 333W, 115V - Part #86135, and 230V &amp; 240V - Part #86136, Cardiopulmonary by-pass temperature controllers. The ECMO Heaters are used in conjunction with a heat exchanger to warm the fluid temperature. The device is composed of a heater, a circulating pump, fan, safety high limit, and a microprocessor board. Water is heated and pumped from the unit to the heat exchanger. The device can be set so that the operator can monitor the actual water temperature or, with the use of a 400 series probe, the fluid (external) temperature. |  |  | February 03, 2012 | 2012 | Cincinnati Sub-Zero Products Inc | Design error. The probe jack contacts (terminals on certain of the firm's ECMO Heater devices are located too close to the solid state relay contacts (terminals) which can lead to the terminals coming into contact with one another. If this situation occurs, a Service Technician, Operator/User, and/or Patient could be exposed to electric shock during device servicing or use. This situation was ori | Device Design | Cincinnati Sub-Zero (CSZ) issued an 'URGENT Medical Device Field Action Letter to their customers dated January 31, 2012. The letter was mailed out starting on February 3, 2012. The letter includes: a description of the product and problem, describes the affected Serial Number Range subject to recall/correction and instructs the customers to discontinue using the affected serial-numbered units.The customers are instructed to return the product to their Service Department or Biomedical Engineering Dept. for correction and to complete and mail to CSZ the attached 'Field Action Form.' For additional information, the firm can be reached on a toll-free # 1-800-989-7373. |  | 298 | 298 units (115V=295 and 230V/240V=3) | AL, AR. AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL,K IN,, KS, LA, MA, MD, MI, MN, MO, NC, NE, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WI | Terminated | February 26, 2013 | 389 | 298 | Not_Computer
Z-1654-2012 | 61621 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PlasmaPrO" |  |  | November 29, 2010 | 2010 | Apheresis Technologies, Inc. | Apheresis Technologies, Inc., Palm Harbor, FL, initiated a voluntarily withdrawal on November 29, 2010 of their five (5) units of the Plasma Pro Plasma Pump and later classified as a recall (removal) rather than a withdrawal, Model # PP-05, Serial Numbers: E08003041, E08003042, E08003043, E08003044 and E08003045. | PMA | Apheresis Technologies sent a Letter by FedEx to their two consignees. The letter notified customers of the voluntary recall initiated, affected product and description of the problem. For questions call 727-784-2353 or 1-800-382-5556. |  | 5 | 5 units | US Nationwide Distribution including the states of Colorado and Utah. | Terminated | May 25, 2012 | 543 | 5 | Not_Computer
Z-1655-2012 | 61825 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, GuideLiner V2 Catheter, Rx only, Sterile EO, Catalog Number 5571, 6F, Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | January 17, 2012 | 2012 | Vascular Solutions, Inc. | Vascular Solutions has became aware of a potential defect with their 6F GuideLiner V2 catheters (Model 5571) with lot number 555038. Use of this product could result in the catheter shaft separating from the pushwire on the device. This may cause a procedural delay and/or injury possibly requiring medical intervention. | Material/Component Contamination | Vascular Solutions sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to check their inventory for the affected product , immediately remove the product and place the product in a secure area. Additionally, customers were asked to complete and return a Field Action Customer Inventory Form to the Customer Service Department via fax at 763-656-4251 or e-mail at customerservice@vasc.com. Customers may contact their local sales representative for questions concerning this recall. |  | 330 | 330 | Nationwide Distribution including the states of AL, AZ, AR, CO, CT, DC, FL, GA, HI, IL, IN, IA, KS, KY, LAM ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, WA, WV, and WI. | Terminated | July 19, 2012 | 184 | 330 | Not_Computer
Z-1657-2012 | 61692 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Becton Dickinson Phoenix System Software, catalog # 448035/448046, component of BD automated microbiology system, and used with BD Gram Positive MIC or MIC/ID panels. |  |  | April 10, 2012 | 2012 | Becton Dickinson &amp; Co. | BD identified an increase in vancomycin-resistant strains that may exhibit erroneous results on all PMIC or PMIC/ID panels and could lead to ineffective treatment regimen in affected patients. | Nonconforming Material/Component | Becton Dickinson (BD) sent an Urgent Field Correction Action letter dated April 2012 and an Update letter dated April 18, 2012 to all affected customers via e-mail and UPS overnight delivery. The letter identified the affected products, problem and actions to be taken. The letter provided specific workaround instructions for device users pending software upgrades. A BD Service Engineer will contact affected customers to arrange for the installation of the updated software when it becomes available. The letter included an Acknowledgement Form to be completed. For questions contact Technical Services at 800-638-8663. |  | 3454993 | 3,454,993 total products (U.S.) | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT ,DC ,FL , GA, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, AND WY. and the countries of Canada, Australia, Philippines, Thailand, Malaysia, India, New Zealand, Belgium, Singapore, Japan, Taiwan, China, Chile, Colombia, Korea, Brazil and Venezuela. | Terminated | August 08, 2013 | 485 | 3454993 | Computer
Z-1673-2012 | 61540 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | HPBLUE: HARMONICÂż Blue Hand Piece; HP054: HARMONICÂż Hand Piece; Non-Sterile products are individually packed in a poly bag and then placed into an individual carton. Mfr. By: Ethicon Endo-Surgery, LLc, 475 Calle C, Guaynbo, PR 00969. |  |  | March 16, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated a voluntary recall due to some lot numbers of the HARMONICÂż Hand Piece (HP054) and the HARMONICÂż Blue Hand Piece (HPBLUE), which impacts their compatibility with the HARMONICÂż Generator 300 (GEN04). The problem occurs after successful set up and initial use of HARMONICÂż Hand Pieces (HP054 and HPBLUE) with the Generator 300 (GEN04). The procedure is interrupted when | Nonconforming Material/Component | Ethicon Endo-Surgery Inc sent an URGENT: Product Correction Notification letter dated March 2012 to all affected customers. The letter identified the product, the problem, the action to be taken by the customer. Labels were enclosed that are used to identify the recalled product. Customers were instructed to fill out the Business Reply Form and fax it to Stericycle at 1-888-912-7091 within three business days for a credit.For questions regarding this recall call 513-337-3419. |  | 1896 | 1896 instruments | Worldwide Distribution-USA including: AK,AL,AR,AZ,CA,CO,CT,DE,FL,GA,HI,IA,ID,IL,IN,KS,KY,LA,MA,MD,ME,MI,MN,MO,MS,MT,NCND,NE,NH,NJ,NM,NV,NY,OH,OK,OR,PA,RI,SC,SD,TN,TX,UT,VA,VT,WA,WI,WV, and WY.and the countries of: Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Ecuador, Egypt, Greece, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Korea (South),Kuwait, Lebanon, Malaysia, Mexico, Netherlands, Antilles, Pakistan, Panama, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Taiwan, Thailand, Turkey, United Arab Emirates Uruguay, Venezuela &amp; Yemen. | Terminated | March 05, 2013 | 354 | 1896 | Not_Computer
Z-1675-2012 | 61760 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens MAGNETOM Essenza, Nuclear Magnetic Resonance Imaging System |  |  | March 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is an error in the compatibility data sheet which is part of the System Owner Manual | Error in labeling | Siemens issued a Customer Safety Advisory Notice letter to affected customers on March 30, 2012. The letter identified the affected product, problem, potential risk and what steps can the user take to avoid the potntial risk of this issue. Siemens is releasing corrected pages of the Operator Manual in order to deliver clear guidance and recommendation for the users. Customers are instructed to ensure that the safety advisory and updated pages are replaced in the System Owner Manual. |  | 97 | 97 | US Nationwide Distribution - including the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, MS, MO, NE, NJ, NY, NC, OH, OR, PA, TN, TX, WA, WI, WY and Puerto Rico. | Terminated | April 03, 2014 | 734 | 97 | Not_Computer
Z-1436-2012 | 52013 | No | LNX |  |  | General Hospital | 510(k) | Abacus TPN Calculating Software |  |  | January 19, 2009 | 2009 | Baxa Corporation | Ordering salt-based ingredients on an ion-based template resulted in medication errors when calculating amounts of ingredients required. | Labeling design | Baxa Corporation sent a "Safety Alert" letter dated January 19, 2009 via FedEx to all affected customers. The letter described the affected product, problem, and actions to be taken. Customers were instructed to ensure that users create either a salt-based template OR an ion-based template for ordering, and that both methods of ordering are not used in one template. Customers were provided a list of safe practices to follow on the enclosed ISMP Safety Alert. Customers were advised upon receipt of the Safety Alert that Baxa Technical Support will review and address any issues with the customer on their configuration setup for calcium and magnesium. For questions contact Baxa Technical Support at 1-800-678-2292. |  | 667 | 667 units | Worldwide Distribution - US (Nationwide) and the countries of: Canada, China, Australia, Panama and Columbia. | Terminated | September 12, 2012 | 1332 | 667 | Not_Computer
Z-1674-2012 | 61532 | No | KMA | Powered nasal irrigator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Water Pik Sinusense battery operated device including a blue semi-transparent water reservoir, a white hand held button initiated injector device, and a soft seal nozzle affixed to the white injector. The product is sold individually in white, blue red and orange cartons with a photo of the device as well as a stream of water flowing through it. The contents of the container include the device as well as a package insert with utilization instructions. |  |  | April 03, 2012 | 2012 | Water Pik, Inc. | Customer complaints related to cross contamination of sinus rinse solution between the battery compartment and solution reservoir used for rinsing nasal passages. | Device Design | Water Pik has generated a Press Release letter dated May 2012. The letter included safety information and instructions for the consumer to contact Waterpik for a replacement model if they experience water intrusion into the battery compartment of their device. For additional information go to www.waterpiksinusense.com/exchange. |  | 105257 | 105257 individual units | US Nationwide Distribution. | Terminated | August 07, 2012 | 126 | 105257 | Not_Computer
Z-1677-2012 | 61338 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration HA Reduced Neck 255 mm Long Stem |  |  | October 27, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | An incorrect laser mark was noted on specific lots of product, incorrectly implying a stem size of 9, when the product is actually size 8. | Error in labeling | Stryker sent an "URGENT PRODUCT RECALL" letter dated October 27, 2011 and October 28, 2011 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs the customers to return the product to the firm. An Acknowledgement Form was included for customers to complete and return. Contact the firm at 201-972-2100 for questions regarding this recall. |  | 31 | 31 Units | Nationwide Distribution and the country of Canada. | Terminated | August 22, 2013 | 665 | 31 | Not_Computer
Z-1678-2012 | 61814 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Energy Source REF 60-9500-120 Tissue Fusion Energy Source ConMed Corporation 525 French Rd, Utica, NY 13502, MDSS GmbH, Germany. |  |  | May 07, 2012 | 2012 | ConMed Electrosurgery | An investigation of the ConMed Electrosurgery Altrus Energy Source revealed that the device could have been calibrated incorrectly resulting in error messages or delays in surgery. | Equipment maintenance | ConMed Electrosurgery sent an "URGENT: MEDICAL DEVICE" letter dated May 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to review their inventory for the affected product and to complete and return a Business Reply Form to the firm via fax to 303-699-9854. |  | 122 | 122 units | Nationwide Distribution and the country of Canada. | Terminated | June 19, 2012 | 43 | 122 | Not_Computer
Z-1676-2012 | 61756 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SOLUS Anterior Lumbar Interbody Fusion Spinal |  |  | March 19, 2012 | 2012 | Alphatec Spine, Inc. | Alphatec Spine has received several complaints (n+4) on the Alphatec Solus Distal Deployment Wrench for twisting or bending, which has been discovered either during or after instrumentation use. A fifth complaint received in which it was reported that the spline tip of the Alphatec Solus Proximal Deployment Wrench had sheared from the instrument upon use. | Device Design | Alphatec Spine sent an Urgent Medical Device Recall Notification dated March 30, 2012, to all affected customers. The notification letter informed customers of the removal/recall and included instructions on what to do with the recalled products. Initial letter was delivered via FedEx, and followed by telephone call with a total of 4 attempts. Customers were instructed to return all of the affected product. Upon receipt of of the Medical Device Recall Notification, customers were instructed to contact Alphatec Spine's Customer Service Department immediately for instructions on how to return the instruments and fill out the lower portion of this letter to confirm that they have read the notification and have taken all necessary removal actions: Customers should return a signed copy of the form to:Alphatec Spine, Inc.5818 El Camino Real,Carlsbad, CA 92008ATTN: Rand JadanForms may also be returned via email to rjadan@alphatecspine.com or fax to 760-431-0289.For questions regarding this recall call 760-431-9286. |  | 18 | 18 devices | Worldwide Distribution - USA (nationwide) and Germany | Terminated | June 20, 2012 | 93 | 18 | Not_Computer
Z-1679-2012 | 61821 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion, Tissue Fusion Handpiece 5 mm O.D., 36 cm Shaft, Catalog Number 60-9522-001, Sterile/EO. |  |  | May 07, 2012 | 2012 | ConMed Electrosurgery | The devices were distributed prior to the firm completing performance qualifications for their package sealing and adhesive bonding. | Process control | ConMed Electrosurgery sent an "URGENT: DEVICE RECALL" letter dated May 2, 2012 to all affected customers. The letter identifies the product, problem,and actions to be taken by the customers. The letter instructs customers to review their inventory and to complete and return a Business Reply Form via fax to 303-699-9854. |  | 153 | 153 units | Class II Recall-Nationwide Distribution-including the states of OH, GA, IL, PA, AZ, KY, FL, NY, MO, TN, TX, AL, IN, SC, MD, CT, NJ, and CA. | Terminated | June 05, 2012 | 29 | 237 | Not_Computer
Z-1684-2012 | 61808 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON S2000 Ultrasound Systems, and the Advanced SieClearTM spatial |  |  | April 23, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | When using the Advanced SieClear spatial compounding feature a synchronization error may occur which can cause a mismatch between the image and the depth scale. | Software change control | Siemens initiated a Customer Safety Advisory Notification letter Field correction on April 23, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers' local Customer Service Engineer will install an update to their system as soon as it becomes available. For questions customers were instructed to contact their local service support person. Customers were also asked to pass the notification leter to all those within their organization who need to be aware of the issue until the corrective action is completed.For questions regarding this recall call 650-694-5747. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Austria, Australia, Belgium, Canada, Switzerland, Chile, China, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Israel, India, Italy, Jordan, Japan, South Korea, Kuwait, Lebanon, Myanmar, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Sweden, Slovenia, El Salvador, and Thailand. | Terminated | July 16, 2013 | 449 |  | Not_Computer
Z-1661-2012 | 61895 | No | FPK | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Epidural Admin Set with Non-Vented Bag Spike and 0.2 Micron Filter Package Assembly, product code: 340-4126, Method of Sterilization using Ethylene Oxide. |  |  | May 18, 2012 | 2012 | MOOG Medical Devices Group | Moog Medical Device Group became aware on March 23, 2012 via customer complaint, of a reversed pump segment on an Administration Set. This malfunction could reverse the flow of fluid or medication backwards from what was intended. | Process design | The firm, Moog Medical Device Group, sent an "Urgent Field Safety Notice" dated May 21, 2012 via certified letter and a press release direct to consignees/customers. The notice describes the product, problem and actions to be taken. The customers were instructed to do the following: If purchased by distributor -remove suspect product from inventory; contact your distributor to arrange replacement product and request a shipping return label as per the distributor's process and ship product to their location.If purchased by Moog-remove suspect product from inventory, contact Moog Customer Service at 1-800-970-2337, prompt #7 and Moog customer service will provide a call tag for convenient product return.For additional information or clarification, please contact the Post Market Surveillance Manager at (801) 264-1001, ext. 224 or via email at: lbrewer@moog.com. |  | 65160 | 65,160 | Nationwide distribution: USA including states of: AZ, CA, FL, IL, KS, MA, MI, MD, NC, ND, NY, OK, OH, and TX. | Terminated | February 22, 2013 | 280 | 544900 | Not_Computer
Z-1683-2012 | 57371 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Terumo BCT Lakewood, CO 80215, Catalog No 61000 Spectra Optia Apheresis System, Country of Origin: US. |  |  | May 04, 2012 | 2012 | CaridianBCT, Inc. | Defective product may result in decreased efficacy. | Component design/selection | Terumo BCT sent a Voluntary Medcal Device Safety Alert letter dated May 18, 2012 to each of their consignees via certified letter. The letter identified the affected product, problem, actions to be taken and instructions for healthcare providers until a corrective action is identified and implemented. Customers were instructed to share this safety alert with all those who need to be aware and complete the attached Acknowledgement Form and fax to +1.303.876.9277 or email to Terumo BCT. For questions contact your Terumo BCT Representative, the Terumo BCT Support Center at (U.S. Toll Free +1.877.339.4228) or +1.303.231.4357, or your local Terum BCT Customer Service Office. |  | 372 | 372 units | US Nationwide Distribution and the country of Canada. | Terminated | November 27, 2012 | 207 | 372 | Not_Computer
Z-1690-2012 | 61762 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 16 - Model #728246, CT Scanning System is subject to recall/correction |  |  | March 26, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Component defect. Following their receipt of a customer complaint, Philips Medical determined that the metal blocks which support the scissor slide rails underneath the scanning table (couch) of several of the firm's CT scanning systems were not manufactured to specification and may crack. | Nonconforming Material/Component | Philips issued an 'URGENT Medical Device Correction' letter dated March 27, 2012 to all affected customers. The letter identified the affected products, problem, hazard involved, how to identify affected products, and actions to be taken by customer/user. The letter informed customers that Philips intends to dispatching service personnel to each customer site in order to correct this issue. In the meantime, the letter recommend customer to contact their Philips representative if they hear unusual noises when moving the couch vertically, or if there is unintended vertical motion of the couch. For questions contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center 1-800-722-9377, option 5: Enter Site ID or follow the prompts. |  | 10 | 10 system units were distributed in the U.S. | Worldwide Distribution - U.S. (Nationwide) including the states of: CA, FL,, HI, IA, IN, KY, LA, MA,, MI, MO, MS. NM, NV, OH, PA, PA, SC, TN, TX, WA, and Puerto Rico | Terminated | February 22, 2013 | 333 | 35 | Not_Computer
Z-1689-2012 | 61779 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Practix Convenio Mobile X-ray System |  |  | May 02, 2012 | 2012 | Philips Healthcare Inc. | Unexpected Low-level X-ray pulse | Device Design | Philips sent a Urgent Field Safety Notice dated April 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of the communication. Please retain a copy with the equipment instruction for Use. A Philips Service Engineer will contact account as soon as the Field Action Kit is ready to be implemented. Communicate with Philips with regard to this program, please reference Field Change Order 70400049.Further questions please contact your local Philips representative at 1-800-722-9377.On 2/6/15, Philips reported that the recall was extended to an additional 88 units that were unit stock that did receive the correction. Philips issued recall notifcation letter (FCO70400049) on 1/20/2015 to users. Control boards will replaced by a Philips Service Engineer. |  | 76 | 10 US; 66 internattional | Worldwide Distribution-- USA ( nationwide ) including the states of IA, LA, LA, ND, NY, OH and TX and the countries of Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Germany, Hong Kong, Indonesia, Italy, Malaysia, Mexico, Netherlands, Philippines, Qatar, Russia, Saudi Arabia, Singapore, South Africa,Thailand and Ukraine. | Open, Classified |  |  | 76 | Not_Computer
Z-1688-2012 | 61858 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance iCT and iCT SP are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | May 07, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | On 3/5/2012, Philips Healthcare was notified of one incident involving a Brilliance iCT scanner that resulted in damage to the gantry and slight injury to the user when the heat exchanger detached from the cooling unit due to some missing welds. | Nonconforming Material/Component | Philips Healthcare sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated May 8, 2012 to all affected customers. The letter identifies the product and problem. No action is required by the customers. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. Philips Healthcare will inspect the cooling units on all iCT and iCT SP systems for this issue and if the welds on a cooling unit do not pass the inspection, Philips will replace the cooling unit free of charge. |  | 372 | 372 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MS, NC, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, VT, WA and WV and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Great Britain, Greece, Hungary, Iceland, India, Indonesia, Iraq, ISRAEL, Italy, Japan, Jordan, Korea, Lebanon, Mexico, Netherlands, New Zealand , Norway, Philippines, Poland, Romania, Russian Federation, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Arab Emirates and Uzbekistan. | Terminated | February 26, 2013 | 295 | 372 | Not_Computer
Z-1686-2012 | 61672 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E Series SPECT System; Single-Photon Emission Computed Tomography. The Siemens Symbia E system is a high performance dual detector variable angle system with flexible patient positioning for planar imaging, whole-body scans, and SPECT studies. |  |  | May 15, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | Siemens has identified a potential patient safety issue for e.cam and Symbia E camera systems that have had the radial drive motor assembly replaced during a Service action. There is a remote chance that during the motor replacement the radial drive gear may have been replaced incorrectly. If the gear is installed incorrectly the mechanical brake may not engage creating a patient hazard if the sys | Employee error | On May 15, 2012, Siemens Medical Solutions USA, Inc. sent an "Urgent Field Correction Recall" letter dated April 11, 2012 to all affected customers. The letters informed customers of the potential patient safety issue for e.cam and Symbia E camera systems that have had the radial drive motor assembly replaced during a Service action. Customers were informed that a Siemens Service Engineer will be inspecting the radial drive gear on their camera system to ensure it is functioning correctly. For questions or assistance they were instructed to contact Siemens Medical Solutions USA, Inc. at 800-888-7436. |  | 9 | 9 units | Worldwide Distribution, US (Nationwide) and the countries of Australia, Belgium, Canada, China, France, Germany, Italy, Japan, Republic of Korea, Mexico, Netherlands, Portugal, Spain, Taiwan and the United Kingdom. | Open, Classified |  |  | 146 | Not_Computer
Z-1695-2012 | 61908 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Monoject Insulin Safety Syringe 3/10mL 30 X 5/16" |  |  | May 21, 2012 | 2012 | Covidien LLC | Safety shield may separate from the syringe, exposing the needle, which could potentially result in a needle stick | Equipment maintenance | Covidien sent an Urgent Product Recall letter dated May 21, 2012, via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately stop using any affected product. Distributors were requested to notify their customers. Customers were instructed to complete the Monoject Insulin Safety Syringe recalled product return form even if they have no product to return. The form should be faxed to 1-800-895-6140 or e-mailed to sdfeedback@covidien.com. Customer Service questions regarding the recall please use one of the following methods.Telephone:1-800-722-8772 option I (Weekdays, 8am to 6:30pm Eastern)e-mail:sdfeedback@covidien.com |  | 77500 | 77,500 syringes | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia, Belgium, and Panama. | Terminated | June 10, 2013 | 385 | 28069500 | Not_Computer
Z-1699-2012 | 61790 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System Software Version 1.6 &amp; Below |  |  | April 26, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling VITROS 5600 Integrated System Software Version 1.6 and below because of an anomaly that may inadvertently add a surfactant to a CurveTip position. | Software design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letter dated April 26, 2012, via FedEx overnight courier to US consignees with VITROSÂż 5600 Systems. The letter identified the affected product, problem and actions to be taken. Foreign affiliates were informed by email on April 26, 2012 of the issue and instructed to notify their consignees of the issue and actions. OCD recommend customers to group all DAT assays into a single run as described in the letter, until the next Software Version 2.0 is installed on their VITROSÂż 5600 System. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311 |  | 1357 | 1357 units | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 1357 | Computer
Z-1698-2012 | 61623 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Transforminal Lumbar Interbody Fusion (TLIF) PEEK Inserter used in the Genesys Spine Apache" Interbody Fusion System. |  |  | August 24, 2011 | 2011 | Genesys Orthopedic Systems, LLC | The inserters have the potential to break. | Reprocessing Controls | Genesys Orthopedic Systems, LLC decided to recall the affected product and called their customers on August 24, 2011. The firm removed the inserters from distribution from 8/24-31/2011. |  | 4 | 4 inserters | Nationwide Distribution including Florida and Texas. | Terminated | December 04, 2012 | 468 | 4 | Not_Computer
Z-1700-2012 | 61755 | No | CFP | Immunoreactive insulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE --- Insulin --- Catalog Number: LKIN1 (100 test), LKIN5 (500 tests) --- Test Code: INS, Color: Orange |  |  | April 30, 2012 | 2012 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics has observed a negative bias in the Immulite Insulin assay. | Component change control | Siemens sent an Urgent Device Recall letter dated April 2012 via fax to affected U.S. customers and an Urgent Field Safety Notice to customers outside the U.S. The letter identified the affected kit lots, reason for voluntary recall, risk to health and actions to be taken by the customer/user. Customers were instructed to discard any inventory of the affected kit lots and complete and return the attached Field Correction Effectiveness Check questionnaire within thirty (30) days to request a no-charge replacement kits. Customers were requested to post this field correction by each IMMULITE system to alert users to the situation. For questions contact your Siemens Technical Solutions Center. |  |  |  | Worldwide Distribution - US (nationwide) | Terminated | March 16, 2017 | 1781 |  | Not_Computer
Z-1702-2012 | 61673 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microsoft Amalga UIS 2009 (previously known as Microsoft Amalga Unified Intelligence System), Calculator/data processing module, for clinical use. |  |  | April 12, 2012 | 2012 | Microsoft Corporation Health Solutions Group | Microsoft Amalga UIS 2009 has an issue with the display of lab results in the Lab Details window of the lab result viewer. The Lab Details view does not freeze/pin the row header when the user scroll horizontally. | Software design | Microsoft Corporation Health Solutions Group sent an "URGENT FIELD SAFETY NOTICE" dated April 12, 2012 to all affected customers via e-mail. The letter identifies the product, problem, and actions to be taken by the customers. The firm plans to revise its labeling for Amalga to clarify that the row headers will not be visible as the user scrolls horizontally. Customers were asked to complete and return to the firm a response form via fax at 425-706-7329 or e-mail to vlpigin@microsoft.com. |  | 13 | 13 units | Worldwide Distribution-USA (nationwide) including the states of AR, MA, MD, NC, NY, and WA and the country of China. | Terminated | December 12, 2012 | 244 | 13 | Not_Computer
Z-1707-2012 | 61823 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Heparin Lock Flush Solution 100 Unit/mL, USP |  |  | May 07, 2012 | 2012 | Excelsior Medical Corp | Some lots of Heparin Lock Flush, USP Syringes (3 mL fill in 10 ml syringe, 4 mL fill in 10 ml syringe, 5 mL fill in 10 ml and 6 ml in 10 ml syringe), failed Anti-Factor IIa assay for potency. | Nonconforming Material/Component | Excelsior Medical sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue using the affected product and to complete and return an attached Customer Reply Form via fax to (866) 807-3714. Contact Excelsior Medical Customer Service at (800) 487-4276 for questions regarding this recall. A second letter which included additional lots was sent out on 5/25/2012 with the date of 5/24/2012. A return response form was also included. |  | 34230500 | 34,230,500 units (596 lots) | Worldwide Distribution-USA (nationwide) and the countries of Brazil, Canada, and Colombia. | Terminated | September 12, 2013 | 493 | 34230500 | Not_Computer
Z-1706-2012 | 61766 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Xia 3 Anti-Torque Key Manual Surgical Instrument; Catalog Number 48237026. |  |  | January 12, 2012 | 2012 | Stryker Spine | Stryker received four reports regarding a handle which separates at the point where it is welded to the shaft on the Xia 3 Anti-Torque Key during surgery. | Mixed-up of materials/components | Stryker Spine sent an Urgent Product Recall letter dated January 12, 2012 with an attached Customer Response forms via FEDEX to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory to identify the affected product and return using the prepaid mailing label provided. The letter asked customers to complete the Customer Response Form and fax to 201.760.8370. For questions call 201.760.8298. |  | 95 | 95 units | US Nationwide Distribution | Terminated | March 21, 2013 | 434 | 95 | Not_Computer
Z-1703-2012 | 61834 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Brand Name: QuietCare, Model Numbers QC200403, v.02; QC200403, v.03; QC200403, v.04 -- Classification Name: System, Environmental Control, Powered. -- Firm on label: Manufacturer Living Independently Group, Inc., A wholly owned subsidiary of General Electric Company, PO Box 1466, New York, NY 10116 USA, 1-866-772-8243 --- The system is a passive behavior-monitoring device. QuietCare uses strategically placed, unobtrusive, passive wireless infrared pyroelectric sensors in the senior's residence to detect their movement. Data from these sensors are sent to a base station that periodically transmits the information to a centralized secure file server through standard phone lines over a toll-free number. Sophisticated algorithms analyze the data, learning each individual's normal patterns of behavior. When the system detects significant deviations from these baselines it automatically alerts designated caregivers. QuietCare also generates alerts of late wake-ups, potential bathroom falls and if household temperatures become dangerously low or high. The behavioral changes that the system tracks can often serve as subtle indicators of the potential emerging health problems. These changes include: nighttime bathroom usage, entering or exiting the bedroom at a different time than usual, change in the number of times the refrigerator door is opened/closed change in the number of times the medicine cabinet door is opened/closed. The intended users of the QuietCare device are the caregivers at the Assisted Living Facilities and Independent Living Facilities. The monitored individuals (seniors) do not actively operate the device and the device function is not dependent upon the senior resident's ability to use the device. --- Class 2 device, 510(k) exempt, Device Listing Number D117156. |  |  | August 01, 2011 | 2011 | Intel-GE Care Innovations LLC | On 7/29/11, the QuietCare production system experienced a partial outage from about 15:00 hours until 22:30. During the engineering investigation into the cause of the service disruption, it was discovered that some residents did not have the wander and night motion features properly activated, and alerts were not being transmitted. Failure to generate alerts in a timely manner could result in a | Employee error | Intel-GE Care Innovations sent an "URGENT FIELD NOTIFICATION" e-mail dated August 12, 2011 to all affected customers. The notification identifies the product and problem. No action is required by the customers. Contact Customer Service at (866) 772-8243 for questions regarding this notice. |  | 61 | 61 units were identified as affected | Nationwide Distribution-USA (nationwide) including the states of Arizona, California, Connecticut, Idaho, Illinois, Minnesota, New Jersey, New York, Ohio, Pennsylvania, Texas, and Utah. | Terminated | June 05, 2012 | 309 | 61 | Not_Computer
Z-1705-2012 | 61919 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Product is the Spacelabs Medical Xprezzon Bedside Monitor, Model 91393. |  |  | May 22, 2012 | 2012 | Spacelabs Healthcare, Llc | The screw used to connect the external power supply to the Spacelabs Medical Xprezzon Bedside Monitor, Model 91393, may damage an internal capacitor. The damage may result in an inability of the monitor to power on, power off, or involuntarily shut down. | Device Design | The firm, Spacelabs Healthcare, sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 22, 2012 via return receipt requested, to US customers. On May 29, 2012 a translated customer letter was emailed to all international subsidiaries and distributors of record. The letter describes the product, problem and actions to be taken. The customers were instructed to notify staff of the potential failure and to only tighten the power supply screw with their fingers. In addition, the firm will schedule a time to upgrade the monitors at the customers site.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 1898 | 1898 monitors (Domestic: 1188) (International: 710) | Worldwide distribution: USA (nationwide) and countries including: Australia, Canada, Chile, China, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, India, Italy, Kuwait, Malaysia, Netherlands, Panama, Poland, Portugal, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | May 07, 2013 | 350 | 1898 | Not_Computer
Z-1704-2012 | 61550 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Knee walker, one unit per box. Manufacturer Kenstone Metal (Kunshan) Co., LTD. The intended use is to provide additional mobility to individuals who have difficulty walking. |  |  | April 06, 2012 | 2012 | Invacare Corporation | Adverse event data prompted testing of device, which revealed that the rod supporting the knee pad was not made to specifications. Upon placing a load on the knee pad, the rod may break or bend causing the walker to collapse. As a result, the user may fall. | Nonconforming Material/Component | Invacare notified all consignees and distributors with a "Urgent: Recall Information" letter dated May 16, 2012. The product issue was described and recommended actions were provided to customers. Questions were directed to 1-888-257-7912. |  | 1640 | 1640 | Worldwide Distribution--USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the country of China. | Terminated | March 01, 2013 | 329 | 1640 | Not_Computer
Z-1772-2012 | 61767 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.0 MM DIAMETER 50 MM THREAD LENGTH EXTERNAL FIXATION HALF PIN HA TI-6AL-4V, REF 71932962, QTY: (1), STERILE R, Smith &amp; Nephew, Inc., Memphis, TN 38116 |  |  | April 13, 2012 | 2012 | Smith &amp; Nephew Inc | Plastic trays which encase these products have burrs on the hinged packaging which can puncture the single barrier sterile pouch, voiding sterility. | Packaging | Smith &amp; Nephew Sales Representatives and Smith &amp; Nephew International Distributors were notified of the problem with the affected product and the recall via email on April 13, 2012. All affected direct consignees were notified via Fed Ex on April 13, 2012. The Urgent Medical Device Recall letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease use and distribution of the product and quarantine them immediately for return. Customers were instructed to complete the Inventory Return Certification Form included with the letter. Customers should fax or email the form to 901-566-7975 or fieldactions@smith-nephew.com. The affected product should be returned to:Smith &amp; NephewAttn: Global Field ActionsSafety Affairs1450 Brooks RoadMemphis, Tennessee 38116If the affected product was distributed further customers were instructed to provide a copy of the recall letter to these customers. |  | 4 | 4 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Malaysia, Germany, Japan, Portugal, United Arab Emirates, Spain, France, South Africa, Norway, Sweden, Denmark, Switzerland, Belgium, China, The Netherlands, Finland, Mexico, Italy, UK, Venezuela, Colombia, Thailand, Chile, South Korea, India, and Turkey | Terminated | September 22, 2015 | 1257 | 112205 | Not_Computer
Z-1773-2012 | 61954 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Sterrad Cyclesure 24 Biological Indicator, P/N: 14324 |  |  | February 17, 2012 | 2012 | Advanced Sterilization Products | The recall was initiated by Advanced Sterilization Products because they recently determined that certain lots of raw materials that were rejected during incoming inspection acceptance test were used for the manufacture of a specific lot range of Sterrad Cyclesure 24 Biological Indicator. | Process control | Advanced Sterilization Products a Johnson &amp; Johnson company sent an Urgent Product Recall letter dated March 16, 2012, to all affected customers. The letter provided the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to read the "Details on the Affected Devices/Description of the problem" section in the letter. Customers were instructed to immediately identify and set aside all product listed above in a manner that ensures product will not be used. The lot number and the part number are printed on the label placed around the ampoules. Customers were instructed to return any affected product in accordance with the "Product Return Instructions" section of the letter. Customers were instructed to complete the requested information on the enclosed postage paid business reply card and return it to Stericycle even if they do not have the affected products listed on the letter in stock. Customers with questions were instructed to contact Stericycle at (877) 492-4795.For questions regarding this recall call 949-453-6400. |  | 2284 | 2,284 cases (137,040 Biological Indicators) | Worldwide Distribution - USA (nationwide) and the countries of AR, AU, BE, BR, CA, EC, HK, JP, MX, and PR | Terminated | October 17, 2013 | 608 | 2284 | Not_Computer
Z-1783-2012 | 61925 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Accessories for the Captiva Spine FuseLOX Lumbar Cage. Products are labeled in part: "FuseLOX Lumbar Cage Lumbar Interbody Fusion Device***Description:***Part Number:***Lot Number:***Quantity: 1***Material: PEEK-OPTIMA LT1, Tantalum ASTM F560***Non-Sterile-Single Use Only (See package insert for sterilization parameters and labeling.)***CAUTION: Federal (United States) law restricts these devices to sale by or on the order of a physician.***Manufactured and distributed by: CAPTIVA SPINE***captivaspine.com***967 Alternate A1A #1 Jupiter, FL 33477 Tel: 877-772-5571 Fax: 866-318-3224 Email: info@captivaspine.com***" |  |  | April 27, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine, Inc. initiated a recall of FuseLOX Lumbar Implants, Trials, and Inserters after FDA advised the firm that there is inadequate support for implemented design changes to the cleared device Pivotec Lumbar Interbody Fusion Device. | Finished device change control | Captiva Spine sent a Voluntary Correction Notification letter dated May 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and return affected devices. Customers were also instructed to complete and return a Tracking/Verification Form.For questions customers were instructed to call 877-772-5571.For questions regarding this recall call 561-277-9480. |  | 136 | 136 units | USA (nationwide) including FL,TX, UT, and NJ. | Terminated | December 24, 2013 | 606 | 254 | Not_Computer
Z-1788-2012 | 37654 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Revolution XR/d; GE Healthcare, 3000 North Grandview , Waukesha, WI, 53188 |  |  | May 31, 2006 | 2006 | General Electric Med Systems LLC | An incident was reported that an operator's finger was pinched between the fixed roller and the gliding tabletop of a Proteus XR/a Table while moving the tabletop with the thumb placed over its edge and fingers underneath | Device Design | GE Healthcare sent an "URGENT SAFETY INFORMATION" letter dated May 31, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included new operator instructions and warning labels. |  | 3796 | 3,796 units | Worldwide Distribution. | Terminated | July 31, 2012 | 2253 | 3796 | Not_Computer
Z-1791-2012 | 61866 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance iCT and iCT SP are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | May 07, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issuing a software update (version 3.2.4) to the Brilliance iCT and iCT SP systems currently running software version 3.2.0, 3.2.1 or 3.2.3, addressing a number of non-conformances. A review was conducted of all nonconformances that are fixed by the software upgrade. None of these nonconformances result in unacceptable hazards.The nonconformance documented below had the highest sc | Software design | Philips sent an Urgent - Medical Device Correction letter dated May 8, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. Customers were instructed to review the Release Document provided them with information about the software update, new feature and enhancements. For North America and Canada customers should contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Side ID or follow the prompts.For questions regarding this recall call 440-483-4918. |  | 241 | 241 | Worldwide Distribution - USA including AL, CA, DC, DE, FL, GA, IA, IN, MA, MD, MN, MS, NC, NY, OH, OR, PA, TN, TX, VT, WA &amp; WV and the countries of Argentina, Australia, Belgium, Brazil, Canada, China, Czech Republic, Denmark, EGYPT, France, Germany, Greece, Hungary, ICELAND, India, Indonesia, Israel, ISRAEL, Italy, Japan, Jordan, Lebanon, Netherlands, New Zealand, PHILIPPINES, POLAND, Portugal, Qatar, Russia, SINGAPORE, South Korea, SPAIN, SWEDEN, Switzerland, Taiwan, THAILAND, Turkey, UNITED KINGDOM &amp; Uzbekistan. | Terminated | July 25, 2013 | 444 | 241 | Computer
Z-1792-2012 | 61853 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Lipofilter 3000, model ASP-CAN-2, labeled in part ***MicroAire, 1641 Edlich Drive, Charlottesville, VA 22911, USA*** |  |  | April 26, 2012 | 2012 | MicroAire Surgical Instruments, LLC | Medical device was not manufactured under Good Manufacturing Procedures and a user may be unable to sterilize the device according to procedures. | Component change control | MicroAire Surgical Instrument sent a Device Removal Notice letter dated April 26, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were advised to recover the affected product from inventory and quarantine. Customers were instructed to contact MicroAire Customer Service Department at 800-975-8000 or 434-975-8000 to arrange for return and replacement of the affected product. For questions call 434-975-8336 or email toddm@microaire.com. |  | 15 | 15 devices | US Nationwide Distribution including the states of CA, OH and NY. | Terminated | November 21, 2012 | 209 | 15 | Not_Computer
Z-1794-2012 | 61951 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PlasmaJet Open Surgery Handpiece (Part of the PlasmaJet Neutral Plasma Surgery System) |  |  | April 27, 2012 | 2012 | Plasma Surgical Inc. | Potential failure in the glued seal between the stem and the tip of the handpiece leading to the possible leakage of a very small drop of distilled water coolant from the handpiece in use. | Nonconforming Material/Component | The firm, Plasma Surgical, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 27, 2012 to its Consignees/Customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their physical inventory; locate and quarantine or remove from service all identified product; and complete and return the attached CUSTOMER RESPONSE FORM via fax to designated persons at: (404) 678-4395 or mail to: Plasma Surgical Inc., 1125 Northmeadow Parkway, Suite 100, Roswell, Georgia 30076 (even if they no longer have any of the product in their inventory). A Plasma Surgical Representative will contact the customers to arrange for the return of the products and product replacement, should they have any of the physical inventory (product).Should you have any queries concerning this matter, please do not hesitate to contact, Vice President of Sales and Marketing, imawhinney@plasmasurgical.com or call (678) 578-4390. |  | 36 | 36 units | Worldwide distribution: USA (nationwide) including states of: CA and NY; and country of: United Kingdom. | Terminated | November 01, 2012 | 188 | 36 | Not_Computer
Z-1796-2012 | 61717 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra Mozaik Putty |  |  | October 11, 2011 | 2011 | Integra LifeSciences Corp. | The Instructions for Use {HIFU} {Putty Part Number 7700713001 and Strip Part Number 770151300 l for CE marked Mozaik Strip and Putty products contains an incorrect translation of the term "bone marrow aspirate". The term is incorrectly translated in Spanish and French, | Employee error | Integra LifeSciences sent a Field Safety Notice letter dated October 11, 2011 to their sole foreign consignee. The letter identified the affected products, problem and actions to be taken. The consignee was asked to remove the existing Instructions for Use (IFU) and replace it with the corrected IFU. The consignee was instructed to complete and return the attached Distributor Rework and Acknowledgment Forms provided with the customer notification letter as documentation that instructions were followed and all required actions were completed. For any questions, call: +33 (0)4 37 47 5900 , fax: +33 (0)4 37 47 5999; or email: carole.robin@integralife.com. |  | 42 | 42 units | Worldwide Distribution | Terminated | February 14, 2013 | 492 | 64 | Not_Computer
Z-1802-2012 | 61945 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems, Inc. A VYGON Company PICC LINE TRAY Kit |  |  | May 11, 2012 | 2012 | Vygon Corporation | Manufacturer issued a voluntary recall on multiple lots of VY(TM)Flush Heparin Lock Flush Syringes therefore Vygon has subsequently initiated a recall on Vygon branded Heparin Lock Flush Syringes and kits containing the syringes. | Nonconforming Material/Component | VYCON sent an URGENT MEDICAL DEVICE RECALL letter dated May 10, 2012, to all affected customers. The firm expanded their recall on May 25, 2012 by sending a second letter with additional lots and products. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory and quarantine any affected product. Customers should contact Vycon's Customer Service Department at 1-800-473-5414 for a Return Merchandise Authorization number. Customers were instructed to enter the number on the space provided on the Recall Acknowledgment and Inventory Return Form, sign and date form and fax to Vycon at 215-672-6740. Customers should return the affected product as well as the original signed Recall Acknowledgment and Inventory Form using the instructions from Customer Service involving product pick-up. All inventory will be credited. If product was further distributed the Urgent Medical Device Recall letter should be provided to their customers.For questions regarding this recall call 215-672-6740. |  |  |  | Nationwide Distribution including AZ, CA, DE, FL, IL, IN, IA, KS, KY, MD, MA, MI, MN, MO, NJ, NY, OH, PA, UT, VA, WA, WV, and WI. | Terminated | March 11, 2013 | 304 |  | Not_Computer
Z-1490-2012 | 61355 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Sigma Spectrum VOLUMETRIC INFUSION PUMP, with or without Master Drug Library, Part Number 35700 -- Common/Classification Name: Infusion Pump -- Class 2 medical device -- SIGMA LLC, Medina, NY 14103 |  |  | July 06, 2011 | 2011 | Sigma | Recall expansion; the infusion pumps have the potential to fail causing inaccurate flow conditions during use. These conditions range from back flow to free flow, which could result in over-infusion. Sigma's evaluation of subsequent complaints involving failed bearings indicated that the loss or degradation of bearing lubricant to be a more significant contributor to bearing failure than initiall | Nonconforming Material/Component | Sigma sent an URGENT Medical Device Recall Notifications letter dated July 6, 2011 via UPS mail the consignees informing them of the expansion of the Spectrum Infusion Pumps recall. Some consignees were notified by telephone. The letter identified the affected product, problem, actions to be taken, clinical precautions and recommendations. Additional information regarding this voluntary recall is available on Sigma's web site www.sigmapumps.com/Recall. Customers were provided the following instructions: (1) Confirm that you have the affected serial number(s) ranging from 700000 to 794213 in inventory using the Customer Response Form. Refer to attached list of known affected devices in your possession. Revise this list as required. (2) Immediately remove affected pumps from neonatal use until B. Braun BC 1000 back check valves are employed. (3) Respond to this notification within five (5) working days of receipt by faxing the enclosed Customer Response Form to the number noted on the form, calling (866) 482-2893 between the hours of 8am and 8pm (EDT) or e-mailing a scanned version of form to sigma2407@stericycle.com. (4) Within three weeks of receipt of the Customer Response Form, (which verifies the device serial numbers in your facility), SIGMA will issue a customized recall report for each facility identifying the specific affected serial numbers, SIGMA internal contacts, and the plan for coordination of pump return and repair. For any adverse reactions experienced with the use of this product should be reported to SIGMA at (800)-356-3454 ext. 300 and to FDA MedWatch either by: Phone at 1-800-FDA-1088, by fax at 1-800-FDA-0178; Mail at MedWatch, FDA 5600 Fishers Lane, Rockville, MD 20857, or MedWatch website by linking to: https://www.accessdata.fda.gov/scripts/medwatch/medwatch-online.htm. |  | 63225 | 63,225 pumps total (corrected number): Domestic - 63,032 pumps; Canada - 29 pumps; Okinawa - 164 pumps | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of Canada and Okinawa | Open, Classified |  |  | 63225 | Not_Computer
Z-1538-2012 | 61830 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Practix Convenio Mobile X-ray system |  |  | October 19, 2011 | 2011 | Philips Healthcare Inc. | When the system is switched off or when the tube arm is parked in the down position (X-ray generator switch off), a radiation pulse (50 kV, 2.5 pGy) is generated inadvertently. .When the system is switched off or when the tube arm is parked in the down position (X-ray generatorswitch off), a radiation pulse (50 kV, 2.5 pGy) is generated inadvertently. The failure happened whencomponent supplie | Radiation Control for Health and Safety Act | Philips Healthcare sent a proprosed corrective action letter dated May 3, 2012, to FDA. The proposed corrective action plan is to bring the product into compliance with regulatory requirements by issuing a field corrective action which should be completed by Philips by August 1, 2012.For questions regarding this recall call 978-687-1501. |  | 11 | 11 - US | Nationwide Distribution | Open, Classified |  |  | 11 | Not_Computer
Z-1810-2012 | 61848 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | T4-Hytrel Zipper Toga |  |  | June 01, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | The clear tape, that is intended to aid in the prevention of patients fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. Stryker has received 5 complaints associated with the defective seams which allowed bleed through of bodily fluids. | Process control | Stryker sent the May 29,2012 URGENT MEDICAL DEVICE RECALL NOTIFICATION and the June 14, 2012 REVISED URGENT MEDICAL DEVICE RECALL Notice to all consignees. The letters identified the affected products, problem and actions to be taken. The Revised letter instructed customers to find, quarantine, make product unusable and destroy all implicated product . Replacement products may not immediately available. Customers were asked to complete and sign the Business Reply Form (BFR) and fax to Stryker Instruments Regulatory Department at 866-521-2762. If the affected products have been further distributed, forward this notification letter and the attached BRF to all affected locations. Questions should be addressed to Angela Ragainis, 269-389-4354; angela.ragainis@stryker.com |  | 16 | 16, 846 (last 3 years) 35,492 all years | Worldwide Distribution - US (nationwide) including the states of: AK, AL, AZ, CA, CO, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV and the country of SWITZERLAND | Terminated | April 14, 2014 | 682 | 16 | Not_Computer
Z-1812-2012 | 61737 | No | MKQ |  |  | Pathology | PMA | PrepStain GYN 480 Disposable Test Kit with PreCoat Slides (Catalog #490523 |  |  | March 22, 2012 | 2012 | Tripath Imaging, Inc. | The PrepStain Syringing Pipettes may not dispense the minimum volume of 7.5 mL of the sample due to leaking. This event resulted in an increase in complaints of leaking syringing pipettes | Component design/selection | BD Diagnostics sent a "BD PrepMate Syringing Pipette" letter dated March 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The firm encouraged customers to refer to the Operator's Manual and inspect the centrifuge tubes to verify that the correct amount of sample has been transferred. Customers experiencing this issue were instructed to discard the pipettes and contact BD Diagnostic Technical Support for a replacement. Contact the firm at 1-877-822-7771 for questions regarding this notice. |  | 46112784 | 46,112,784 for all products in Recall Event | Worldwide Distribution. | Terminated | December 03, 2012 | 256 | 46112784 | Not_Computer
Z-1818-2012 | 61926 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Software, software release 2.0.2 and 2.0.2SP1 including mirror icon |  |  | May 22, 2012 | 2012 | Philips Healthcare Inc. | With Digital Diagnost (DIDI) R 2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "right patient side" marker, although this can appear on the left patient side | Software change control | Philips Healthcare issued a Field Safety Notice on 5/22/12 advising users of the error. A software update (Rel 2.1.2) will issue to correct the problem. Contact your local Philips representative. 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 752 | 752 units | Worldwide Distribution - USA (nationwide) and the countries of:NationwideForeign: CanadaAFGHANISTANALBANIAALGERIAANGOLAARGENTINAARMENIAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHAMASBAHRAINBANGLADESHBARBADOSBELARUSBELGIUMBENINBHUTANBOLIVIABOSNIA&HERZEGOVINA;BRAZILBRUNEI DARUSSALAMBULGARIABURUNDICAMBODIACAMEROONCANARY ISLANDSCHADCHILECHINACOLOMBIACONGOCONGO,DEMOCRATICCOSTA RICACOTE D'IVOIRECROATIACUBACYPRUSCZECH REPUBLICDENMARKDJIBOUTIDOMINICAN REPUBLICECUADOREGYPTEL SALVADOREQUATORIAL GUINEAERITREAESTONIAETHIOPIAfake test ssdFINLANDFRANCEFRENCH GUIANAGABONGAMBIAGEORGIEGERMANYGHANAGREECEGUADELOUPEGUAMGUATEMALAGUINEAGUYANAHONDURASHONG KONGHUNGARYICELANDINDIAINDONESIAIRANIRAQIRELANDISRAELITALYJAMAICAJAPANJORDANKAZAKSTANKENYAKUWAITKYRGYZSTANLAO PEOPLE'S DEM.REPLATVIALEBANONLIBERIALIBYAN ARAB JAMAHIRILIECHTENSTEINLITHUANIALUXEMBOURGMACAUMACEDONIAMALAWIMALAYSIAMALIMALTAMARTINIQUEMAURITANIAMAURITIUSMETRACOMMEXICOMOLDOVAMONGOLIAMOROCCOMOZAMBIQUENAMIBIANEPALNETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORTH KOREANORWAYOMANPAKISTANPALESTINEPANAMAPARAGUAYPERUPHC BESTPHILIPPINESPMCC BU CanadaPMCC BU USAPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIARWANDASAUDI ARABIASENEGALSERBIA&MONTE; NEGROSIERRA LEONESINGAPORESLOVAKIASLOVENIASOMALIASOUTH AFRICASOUTH KOREASPAINSRI LANKASUDANSURINAMESWEDENSWITZERLANDSYRIAN ARAB REPUBLICTAIWANTAJIKISTANTANZANIATHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYTURKMENISTANUGANDAUKRAINEUnidentified CountryUNITED ARAB EMIRATESUNITED KINGDOM URUGUAYUZBEKISTANVENEZUELAVIET NAMYEMENZAMBIAZIMBABWE | Open, Classified |  |  | 752 | Computer
Z-1816-2012 | 62130 | No | GAC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Accolade Offset Rasp Handle |  |  | December 29, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of the potential for the above noted Accolade Rasp Handle Assembly to fracture upon use. | Device Design | Stryker Orthopaedics sent an email notification to all consignees on December 29, 2011, which was followed by an Urgent Product Recall letter dated January 4, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the attached Product Recall Acknowledgement Form and fax back to 201-831-8089 within 5 days. Customers were also asked to contact the hospitals in their territory that have the affected product to arrange return of the product. Customers were instructed to return the affected product to the attention of Regulatory Compliance, Stryker Orthopedics, 325 Corporate Drive, Mahwah, New Jersey 07430 with the fluorescent orange PRODUCT REMEDIATION sticker attached indicating the Product Remediation # RA 2011-161. Mark the outer box with the words "Product Recall."For questions customers were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5970. |  | 41 | 41 units | Nationwide Distribution | Terminated | June 13, 2013 | 532 | 41 | Not_Computer
Z-1787-2012 | 61639 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aestiva/5 7900 SmartVent, anesthesia machine. Used in Datex-Ohmeda Aestiva/5 Anesthesia Systems. |  |  | April 10, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the interlock slide mechanism missing from the 3-vaporizer manifold of your Aestiva 7900 Anesthesia Device. This Potential for two (2) vaporizers delivering agent at the same time. This could manifest as over delivery of a single agent if both vaporizers contain the same agent or as delivery of more than one agent. This m | Process control | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 10, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact GE Healthcare Customer or Technical Support at 800-345-2700 for questions regarding this notification. |  | 8 | 8 | Nationwide Distribution - including the states of AL, FL, GA, MO, and VA. | Terminated | July 09, 2012 | 90 | 8 | Not_Computer
Z-1822-2012 | 62120 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Peak SE Primer Kits (4 syringes per kit) manufactured by Ultradent Products, Inc, South Jordan, UT, Self Etch Primer. The product is contained within prefilled, labeled syringes in a transparent plastic container identified with an aqua, red, orange, blue and black label with black and blue print writing. Product code: 5135, 872.3200 KLE, liquid and powder mixture is a syringe-in-a-syringe delivery system. |  |  | May 24, 2012 | 2012 | Ultradent Products, Inc. | Ultradent is recalling lot B6Z4K of Peak SE Primer as a result of a field complaint, dated May 10, 2012, stating that a syringe of Peak SE Primer was mislabeled as Peak LC Bond Resin. These products are used by Dental professionals for bonding needs in restorative dental procedures. | Labeling mix-ups | Ultradent Products Inc. sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 5, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discard any affected product from their 5135 kit. Customers were asked to complete and return a Recall Return Response form to the firm by fax at 801-553-4609. |  | 172 | 172 kits (4 syringes per kit) | Worldwide Distribution. | Terminated | August 15, 2012 | 83 | 172 | Not_Computer
Z-1821-2012 | 61877 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Philips Easy Diagnost |  |  | May 11, 2012 | 2012 | Philips Healthcare Inc. | If a lead ruler is not positioned in the same plane as the anatomical image detail to be measured, measured object distances (using the ruler) can be wrong by a corresponding factor. If, additionally, the measurement tool is not used to check the ruler, or is used in the right way (ie: calibration), such calibration error is not detected. As a result of this, a surgical intervention may be execut | Device Design | Philips sent an Urgent - Field Safety Notice letter dated May 11, 2012 to their affected customers. The letter identified the affected product, what the issues are, under what circumstances they may occur and the actions to be taken to avoid or minimize the occurrence of the issue Philips plans to supply an Addendum to the Instruction for Use. A Philips Service Engineer will contact the affected customers as soon as the Field Action Kit is ready to be implemented. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 6 | 6 | 6 US customers | Open, Classified |  |  | 6 | Not_Computer
Z-1820-2012 | 61868 | No | IRO | Therapeutic vibrator. | Physical Medicine | Physical Medicine | 510(K) Exempt | USA-002 Rev A brochure is a two sided tri-fold printed marketing brochure describing the approved labeling indications for use, precautions and general operational information of the Exer-Rest device, Models AT3800 and AT4700. |  |  | May 07, 2012 | 2012 | Non-Invasive Monitoring Systems, Inc. | Non-Invasive Monitoring Systems Inc (NIMS) is voluntarily recalling promotional materials which contain unsubstantiated direct and indirect claims. | Labeling False and Misleading | NiMS sent an Urgent Promotional Material Recall letter dated May 29, 2012, to all affected customers. NIMS contacted the 18 facilities by phone starting on May 29, 2012 and sent the 73 end users letters dated May 29, 2012. NiMS instructed customers to immediately examine their inventory and quarantine any of the US-001 or USA-002 brochures. If customers further distributed brochures customers were instructed to notify their customers at once by including a copy of the recall notification letter. Customers were requested to destroy all Exer-Rest promotional materials. For any questions customers were instructed to call 305-575-4204.For questions regarding this recall call 305-575-4201. |  | 91 | USA-002 Rev A brochure: 91 | Nationwide Distribution | Terminated | July 16, 2012 | 70 | 3756 | Not_Computer
Z-1778-2012 | 61843 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | NIDEK Medical NUVO LITE MODEL 525 OCSI, 115V -60Hz - 330W. |  |  | May 01, 2012 | 2012 | Nidek Medical Products Inc | Capacitor failure may result in a fire hazard and loss of supplemental oxygen. | Component design/selection | Nidek Medical Products Inc sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 1, 2012 to affected customers. The letter identifies the product, problem, and the actions to be taken by the customers.The letter included a response form for customers to complete and return to the firm. Contact the the firm at 205-856-7200 ext. 242. |  | 730 | 730 units | Worldwide Distribution-USA (nationwide) and the countries of ALBANIA, ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ENGLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, KAZAKHSTAN, KINGDOM OF SAUDI ARABIA, KUWAIT, LEBANON, LIBYA, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, MYANMAR, NAMIBIA, NEPAL, NEW ZEALAND, NIGERIA, PAKISTAN, PERU, PHILIPPPINES, PORTUGAL, REPUBLIC OF CHINA, REPUBLIC OF SOUTH AFRICA, RUSSIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH AUSTRALIA, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, THE NETHERLANDS, TOGO, TUNISIA, TURKEY, UGANDA, UNITED ARAB EMIRATES, and VIETNAM. | Terminated | March 21, 2017 | 1785 | 46972 | Not_Computer
Z-1824-2012 | 57878 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | HEPARIN LOCK FLUSH SOLUTION, USP, 10 USP Units/mL, 1 mL Multiple Dose Vial, plastic vial, vial size 3 mL, NDC 63323-544-01, Product Code 504401, Rx only |  |  | May 20, 2010 | 2010 | App Pharmaceuticals Llc | CGMP Deviations: Incomplete documentation associated with test results. | Release of Material/Component prior to receiving test results | APP sent an URGENT DRUG RECALL LETTER and Response Form dated May 20, 2010, to all affected consignees. The consignees contacted were at the retail/ medical facilities/hospital IeveIs, including intermediate wholesale IeveIs. Direct distributor consignees were instructed to notify their customers immediately of this recall action, and direct them to discontinue distributing or dispensing the affected lots, plus return the affected products and lots to APP Pharmaceuticals, LLC in Bensenville, IL. Direct consignees with affected lots on hand were instructed to immediately discontinue distributing or dispensing the affected products and return them to APP Pharmaceuticals, LLC in Bensenville, IL 60106. Questions regarding the information on how to return product can be made to the firm's Quality Assurance Department at 1-866-716-2459.For questions regarding this recall call 847-969-2700. |  | 133150 | 133,150 vials | Nationwide Distribution and Puerto Rico | Terminated | June 20, 2012 | 762 | 637651 | Not_Computer
Z-1823-2012 | 62125 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow International, Inc. 7FR 3-Lumen Arrowg+ard Blue Plus(R) Pressure Injectable Central Venous Catheter - Catalog number CDC-42703-XP1A |  |  | June 07, 2012 | 2012 | Arrow International Inc | Arrow is recalling the product due to the possibility of voids in the catheter hub which could lead to an interlumen crossover. | Nonconforming Material/Component | Arrow International sent a Urgent Medical Device Recall Notification letter dated June 11, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following action:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog and lot numbers listed.2. To return product complete the enclosed Recall Acknowledgement Form and fax to 1-800-343-2935, Attn: Customer Service. This will allow us to document the amount of the product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization ( RGA) Number and will provide instructions for the return of product to Arrow International.3. If you have no affected stock, please complete t he enclosed Recall Acknowledgment Form and fax it to 1-800-343-2935, Attn: Customer Service. This will allow us to document your receipt of this letter.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 1085 | 1,085 | USA (nationwide) AL, AZ, AR, CA, CO, FL, GA, IL, IN, KS, LA, MA, MI, MO, NV, NH, NC, OH, PA, TX, VA, WA, and WI. | Terminated | April 03, 2014 | 665 | 1085 | Not_Computer
Z-1828-2012 | 61376 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Implant, Dental, Endosseous, Acid Etched, ICE, Osseotite, Internal Connection. . |  |  | November 04, 2011 | 2011 | Biomet 3i, LLC | On November 3, 2011 Biomet 3i, Palm Beach Gardens, FL initiated a recall of their Osseotite Certain Implant, Model # IOSM313, Lot #2010111529. Dental implant lot may not have an internal thread. | Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated November 04, 2011 via e-mail, facsimile, telephone or postal mail to all consignees affected. The letter identified the affected product, problem and actions to be taken. Each customer was instructed to check their respective inventory for the affected product and return any unused item(s) to Biomet 3i for prompt replacement. For question or concerns contact Biomet 3i at 1-800-342-5454. |  | 219 | 219 units in total | Worldwide Distribution - US(nationwide) and the countries of: Brazil, Canada, Columbia, Italy, Lebanon, Spain, Switzerland, and UK. | Terminated | June 25, 2012 | 234 | 219 | Not_Computer
Z-1835-2012 | 61742 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | SARNS FLEXIBLE ARTERIAL CANNULA: 6.7 MM (20 FR) 00 WITH 3/8" CONNECTOR,SUTURE RING, 9.5" (24 CM) LONG. |  |  | May 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During production of one lot of the Sarns High-Flow Aortic Arch Cannula, Terumo Cardiovascular Systems (Terumo CVS) found a foreign substance on the outer surface of some cannulae connectors.Terumo CVS' preliminary investigation found that the substance can be dislodged from the cannula surface, the substance was likely deposited during the molding process, but the exact composition of the parti | Nonconforming Material/Component | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 9, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter advises customers to discontinue use of the affected product and to return the product to the firm. A Response Form was attached for customers to complete and return. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 133793 | 133,793 (each) | Worldwide Distribution-USA (nationwide) including the states of AL , AR , AZ, CA , CO , CT , FL , GA , HI , IA , IL , IN , KS , KY , LA , MA , MD , ME , MI , MI , MN , MO , MT , NC , ND , NE , NH , NJ , NM , NV , NY , OH , OK , OR , PA , RI , SC , SD , TN , TX , UT , VA , WA , WI, and WY and the countries of Australia, Belgium, Brazil, Canada, Chile, Germany, Honduras, Hong Kong, India, Japan, Mauritius, Nicaragua, Philippines, Saudi Arabia, Singapore, South Africa, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates (UAE), and Vietnam. | Terminated | July 11, 2013 | 428 | 133793 | Not_Computer
Z-1831-2012 | 61913 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | The product is Pathfinder SL, version 1.6.0.3057 software used with the Spacelabs Healthcare Pathfinder SL Holter Analyzer, version 1.6.0. |  |  | May 11, 2012 | 2012 | Spacelabs Healthcare, Llc | The firm received reports of the Pathfinder LS software crashing after editing a batch of Holter Analyzer records. The software will stop with an error message requiring a restart. In addition, QTc values will not update after moving the marker. | Software design | Spacelabs Healthcare, Llc sent a "MEDICAL DEVICE CORRECTION" letter dated May 11, 2012 via mailed, return receipt requested, to US customers. The customer letter (translated as necessary) was emailed to all international subsidiaries and distributors of record. The letter identified the affected product, problem and actions to be taken. Consignees are advised to turn the power off and then on to restore normal functionality. Updated software will be provided upon completion of validation. For questions contact Spacelabs Healthcare at 1-800-522-7025 and select 3 for Technical Support. |  | 114 | 114 | Worldwide Distribution - US (nationwide) including the states of Nevada and Vermont, and the countries of AUSTRALIA, AUSTRIA, CANADA, CYPRUS, GERMANY, ICELAND, ISRAEL, POLAND, ROMANIA, SOUTH AFRICA, SWITZERLAND, THAILAND, TURKEY and UNITED KINGDOM. | Terminated | October 22, 2012 | 164 | 114 | Computer
Z-1829-2012 | 62052 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Multi Diagnost Eleva II with swivel cable |  |  | May 21, 2012 | 2012 | Philips Healthcare Inc. | MultiDiagnost Eleva X--Ray system may have damage to System Cables from repeated movement | Device Design | Philips Healthcare Imaging issued Urgent Field Safety Notice dated May 21, 2012, to accounts informing an inspection will be performed of the unit. The Cable guiding plate will be replaced. If the cables show any damage the complete cable set wll also be replaced. No action is be taken by the customer. For further information customers should call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 433 | 433 units | Worldwide Distribution - USA (nationwide)and the countries of:NationwideForeign: AustraliaAustriaBelgiumBharainCanadaChinaColumbiaCzech RepublicDenmarkDutch AntillesEgyptEstoniaFinlandFranceGermanyHong KongIrelandIsraelItalyLatviaLebanonLuxembourgMalaysiaMexicoNetherlandsNew ZealandNorwayPhilippinesQatarRussian Fed.Saudi ArabiaSerbiaSingaporeSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandSyriaTaiwanThailandTurkeyUnited Kingdom | Open, Classified |  |  | 433 | Not_Computer
Z-1843-2012 | 61946 | No | MXM |  |  | Neurology | PMA | Olympic Cool-Cap System is intended to provide treatment for |  |  | May 21, 2012 | 2012 | Natus Medical Incorporated | Olympic Cool-Cap Cooling Module's power supply failed during treatment. The power supply failure can happen without prior indication and results in the unavailability of the system to begin or continue cooling treatment. | Component design/selection | Natus Medical Incorporated sent a URGENT: MEDICAL DEVICE FIELD CORRECTIVE ACTIONS Olympic Cool-Cap System letter dated May 21, 2012, to all affected customers. Fed-Ex to all U.S. customers, specifically to Biomedical department, NICU, with affected devices. The letter identified the product the problem and the action needed to be taken by the customer.In addition, Natus also sent the Urgent Field Safety Notice-Olympic Cool-Cap System (DOC-007459) dated May 2012 to all US customers and International Distribution Partners who were affected by both Frozen Screen and Power Supply field corrective action.Customers were instructed to: " During all routine checks of infants undergoing treatment, look at the clockdisplayed on the screen to verify that the seconds counter on the clock isadvancing. The time displayed should continue to advance if the system isoperating correctly." If the system experiences a screen-freeze during a cooling session, the Olympic Cool-Cap system can be rebooted and cooling resumed." Ensure that all the personnel in your facility that use the Olympic Cool-Capsystem are advised of this Field Safety Notice (see attached).The U.S. customers are advised of the issue specifically stated in Action #2 in the letter and they are also informed that Natus will contact them within the next 30 to 120 days to schedule the power supply replacement in their affected system(s).Natus requires that the Distribution Partners to notify their customers with copies of the Natus recall letter and the Field Safety Notice (DOC-007459). Specifically these two documents should be addressed to both the Biomedical Department and the Neonatal Intensive Care Unit (NICU).Customers with questions about this recall should call Natus Technical Service at 888-496-2887 or email to TService@natus.com. |  | 181 | 181 units | Worldwide Distribution -- USA (nationwide) including the states of : AK; AR; AZ; CA; CO; DE; FL; GA; HI; ID; IL; IN; KY; MA; MI; MN; MS; ND; NE; NJ; NY; OH; PA; SC; SD; TN; TX; and WA. and the countries of : Argentina; Australia; Austria; Canada; Egypt; Germany; Hong Kong; Indonesia; Iran; Italy; Kuwait; Poland; Russia; Saudi Arabia; Singapore; Slovenia; South Africa; South Korea; Spain; Taiwan; Thailand; Turkey; United Arab Emirates; and United Kingdom. | Terminated | February 21, 2014 | 641 | 181 | Not_Computer
Z-1844-2012 | 61942 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien ROTICULATOR 55-3.5 Single Use Stapler |  |  | May 09, 2012 | 2012 | Covidien LP | Potential for the sterility barrier to be compromised. | Package design/selection | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 16, 2012 via federal express to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify and quarantine all affected inventory (note: units from affected lots may have been incorporated into Covidien BEST PRACTICE KITS; evaluate kits for affected lots). All affected ROTICULATOR units must be returned. The customers were also instructed to complete and return the attached ROTICULATOR Single Use Stapler Recalled Product Return form via fax to: (800) 895-6140 or email to: SDFeedback@Covidien.com, even if you do not have any of the affected units, and include a copy of completed form with products being returned. Customers who received product directly from Covidien should contact Covidien Customer Service at SDFeedback@Covidien.com or (800) 722-8772, option 1, prior to returning affected products to Covidien, Field Returns Department, at 195 McDermott Road, North Haven, CT 06473; if received from distributor, return form and contact distributor directly.If you have any questions or concerns, please do not hesitate to contact Covidien Representative or Covidien Quality Assurance at (203) 492-5232. |  | 50698 | 50698 units | Worldwide distribution: USA (Nationwide) and countries including; Africa, Argentina, Australia, Belgium, Brazil, Canada, China Costa Rica, CROATIA, Czech Republic,Dominion Republic, ESPAÂżA France, Germany, Greece Guadeloupe, Guatemala, Honduras Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Korea ,Malaysia, Mexico, Peru, Portugal, Puerto Rico, Spain, Sri Lanka, Switzerland, Taiwan , Thailand, and UK. | Terminated | February 12, 2014 | 644 | 123856 | Not_Computer
Z-1849-2012 | 61741 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic 0-arm Imaging System with Software Version 3.1.1, Catalog Numbers: BI-700-00027-120R and BI-700-00027-120. The 0-armÂż Imaging System is a mobile x-ray system designed for 2D fluoroscopic and 3D imaging. |  |  | October 12, 2010 | 2010 | Medtronic Navigation, Inc. | Software Version 3.1.1 does not meet navigational accuracy. | Software design | Medtronic Navigation contacted customers by telephone on 10/12/2010 and with "Urgent Field Safety Notice" follow up letters dated 10/14 or 10/15/2010. Sites were visited by Medtronic Service Representative to complete navigational accuracy testing and either replace or correct the unit. |  | 3 | 3 units | Nationwide Distribution including the states of CO, IN, and MN. | Terminated | June 21, 2012 | 618 | 3 | Computer
Z-1850-2012 | 61935 | No | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Profile Single-Use Pediatric Snare 11mm |  |  | May 16, 2012 | 2012 | Boston Scientific Corporation | Difficulty in extending snare loop from the catheter | Process control | Boston Scientific issued an Urgent Medical Device Recall Action letter via Federal Express Priority mail on May 16, 2012. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. Scan and fax to 508-683-5578.For questions regarding this recall call 508-683-8000. |  | 2760 | 2760 units | Worldwide Distribution - USA (nationwide) and the countries of Canada, Australia, Austria, Belgium, France, GB, Germany, Italy, Netherlands, Qatar, Spain, Sweden, and Switzerland | Terminated | February 28, 2013 | 288 | 6070 | Not_Computer
Z-1853-2012 | 61761 | No | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow P2Y12 Assay, Part Number: 85064, All reagent lots. |  |  | February 06, 2012 | 2012 | Accumetrics Inc | The recall was initiated by Accumetrics because of a field correction to the VerifyNow System; specifically the results reported when running the VerifyNow P2Y12 Test. Currently the VerifyNow P2Y12 Test reports three values: PRU, BASE and % Inhibition. The change implemented is the elimination of the BASE and % Inhibition results. There will be no change to the reporting of the PRU result whic | No Marketing Application | Accumetrics, Inc. sent an Urgent Field Correction notification letter and an Ackknowledgement - fax back form to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter informed customers that an Accumetrics representative will becontacting them to schedule an appointment to update their VerifyNowinstrument software. The letter stated that Accumetrics realizes that procedures and report format may need to be revised as software changes is implemented, and will provide assistance to help in making this transition. For any questions or concerns, call at 1-800-643-1640 ext. 2; oremail at: support@accumetrics.com. |  | 26075 | 26,075 units | US Nationwide Distribution | Terminated | November 27, 2012 | 295 | 26075 | Computer
Z-1856-2012 | 62127 | No | MCN |  |  | Obstetrics/Gynecology | PMA | seprafilmÂż Single Site ADHESION BARRIER |  |  | June 13, 2012 | 2012 | Genzyme Corporation | Sterility of product may be compromised due to packaging defect | Packaging process control | Genzyme issued an Urgent: Medicatical Device Product Recall Notification Letter dated June 13, 2012 to all affected customers. Stericycle Expert Solutions, will be acting on behalf of Genzyme to coordinate notifications, communications, and logistics for returns of all product subject to this recall. The letter identified the affected products, problen and actions to be taken. Customers are requested to check their inventory and return all affected products per the instructions provided. Customers are instructed to complete the enclosed Business Reply Form and return via fax transmission to 866-912-2519. For questions call Stericycle at: 877- 473-7069. |  | 1285 | 1285 units | Worldwide Distribution - US (nationwide) and the country of S. Korea | Terminated | July 24, 2013 | 406 | 5165 | Not_Computer
Z-1861-2012 | 61712 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Rose Bed Model (FL14E) |  |  | May 17, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | An adverse trend of scale and bed exit service reports/complaints were detected for the GoBed+, GoBed II, Rose, and MA204 model beds. Investigation found an abnormally high failure rate of loadcells, an integral component in scale and bed exit function. The scale faults are manifested as a fluctuating scale reading, a scale error causing the scale system to be disabled until serviced, or an inac | Component design/selection | The firm, Stryker Medical, sent an "URGENT- Medical Device Correction" letter dated May 17, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to observe the beds for any malfunctions according to Preventive Maintenance Checklists in maintenance manual; report scale or bed exit malfunctions to Stryker (1-800-327-0770, option 2) to arrange for field service or to order service parts, at no charge. The customers were also instructed to forward this notice and advise Stryker of any subaccounts if the product was further distributed. A postage paid return postcard is to be filled for tracking and to make arrangements for repair. If you have any urgent questions or concerns, call 269-389-6466, Monday-Friday 8a.m-5p.m. (ET). |  | 999 | 999 | Worldwide distribution:USA (nationwide) and countries including: Argentina, Bertec, Brasil, Brazil, Canada, Chile, China, India, Latin America, Mexico,Osteonics SA, Polska, Portugal, Puerto Rico, and Singapore. | Terminated | May 12, 2014 | 725 | 27319 | Not_Computer
Z-1858-2012 | 61899 | No | NAB | Nonresorbable gauze/sponge for external use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hartmann Sterilux Ref 56910000 Sterile Premium gauze sponge 4"x4" 12-ply 24 boxes/case 1200 total sponges |  |  | May 14, 2012 | 2012 | Hartmann USA, Inc | Micro holes were detected in the laminate of the sterile barrier system of sterilized Hartmann Econolux and Hartmann Sterilux gauze products | Material/Component Contamination | The firm, HARTMANN USA, sent an "URGENT: RECALL NOTICE" dated May 14, 2012 to its Consignees/Customers. The notice described the product, problem and actions to be taken. The customers were instructed to examine their inventory, block further shipments of affected items and retrieve additional affected stock from their customers; to confirm initial receipt of the notice by emailing recall@harmanninfo.com; complete and return the Affected Products and Stock Levels form via fax to: HARTMANN USA customer service at (1-803-325-7606); contact customer service to obtain a Returned Goods Authorization; and take adequate measures to ensure this notice is distributed appropriately.If you have any further questions or issues related to this recall, please contact HARTMANN USA customer service or your sales representative at 1-800-243-2294 or email: recall@hartmanninfo.com. |  | 604711272 | Lot 102147399 - 421 Cases; Lot 102250396 - 360 Cases; 200105390 - 522 Cases; 200206394 - 390 Cases | Nationwide distribution: USA including states of: AZ, CA, FL, GA, KY, MO, MS, NC, PA, SC, TN, TX, and VA. | Terminated | October 06, 2014 | 875 | 1313561759 | Not_Computer
Z-1865-2012 | 62016 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Uroview 2800, Model number A349855. |  |  | November 20, 2006 | 2006 | GE OEC Medical Systems, Inc | Recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system Lok-up, incorrect image display, incorrect dose area calculated air kerma in "film mode" and "digital spot mode", and the system does not reboot. | Software design | On 11/20/06 two written notifications were sent to consignees addressed to the Facility Administrator and the Radiology Administrator. The notifications advised them that GE Healthcare had identified several intermittent potential safety issues that may occur with the product. Interim solutions were provided for each issue. The letter stated that GEHC OEC would contact customers and remedy the issues free of charge when a permanent solution became available. An Urgent Recall Notice, dated April 13, 2007, was also sent to customers. Customers were to contact the service team at 800-874-7378 for further information, if there were any questions or concerns. |  | 717 | 717 units | Worldwide Distribution. | Terminated | June 25, 2012 | 2044 | 717 | Not_Computer
Z-1864-2012 | 62331 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Advanta 2 Bed. |  |  | June 18, 2012 | 2012 | Hill-Rom, Inc. | During an evaluation of Advanta 2 siderail samples returned from the field it was found that the screws used to attach the siderail were being stripped from the plastic cavity of the siderail. | Process design | Hill-Rom sent an "URGENT FIELD SAFETY NOTICE" dated May 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm. Contact Hill-Rom Technical Support at 800-445-3720 for questions regarding this notice. |  | 7843 | 7843 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada, France, and Singapore. | Terminated | June 06, 2014 | 718 | 7843 | Not_Computer
Z-1866-2012 | 61955 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Formula Shaver Handpiece (with buttons); |  |  | May 21, 2012 | 2012 | Stryker Endoscopy | Stryker Endoscopy determined that there may not have been adequate determination of required routine checks to ensure accuracy and fitness for use of specific serial numbers of Formula and 180 Shaver Hand Control. | Nonconforming Material/Component | The firm, Stryker Endoscopy, sent an "Urgent: Device Removal" letter dated May 21, 2012 to all affected customers via Federal Express. Stryker Sales Reps were notified by e-mail on the same day. Stryker international sites were notified by e-mail on May 24, 2012. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory for part and serial number; quarantine any of the affected product, call Stryker Endoscopy's Customer Service at 1-800-624-4422 for replacements and return affected product.If you have any questions regarding this letter please contact us at shavers@stryker.com or call 1-800-624-4422. |  | 34 | 34 devices | Worldwide distribution: USA (nationwide) including states of: CA, FL, IA , MI, NV, NC, OK, TX, and WI; and countries of: China and Great Britain. | Terminated | September 20, 2012 | 122 | 34 | Not_Computer
Z-1826-2012 | 62108 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius NaturaLyte Liquid Acid Concentrate |  |  | March 29, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | Risk of Alkalosis with acetate containing dialysis acidconcentrates | Labeling False and Misleading | Fresenius Medical Care sent an "IMPORTANT PRESCRIBING INFORMATION" letter dated March 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Reply Form was enclosed for customers to complete and return via fax to 781-699-9635. Contact Customer Service at 1-800-323-5188 or Medical Information at 1-855-616-2309 for questions regarding this recall. |  |  |  | Nationwide Distribution and the country of Guam. | Terminated | July 20, 2017 | 1939 |  | Not_Computer
Z-1867-2012 | 61768 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | BCS(R) Behring Coagulation System, Device catalog No. OVIO03 |  |  | December 01, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed that if the Cleaner SCS vial becomes empty after processing a test for Antithrombin assay using the INNOVANCE(R) Antithrombin kit, and the test immediately following is an APTT based clotting assay several results may be shortened. | Device Design | An Urgent Field Safety Notice dated December 2010 was issued to all BCS(R) System owners. Firm conducted a voluntary field corrective action to notify BCS(R) System owners and to install a software "patch" to resolve the issue. Until the software patch was installed on the BCS(R) System, customers were instructed in the Safety Notice that "the potential issue of carryover can be avoided if the Cleaner SCS vial is kept filled and not allowrd to run empty." |  | 186 | 186 | Worldwide Distribution - US (nationwide) including the states of: AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IO, KS, KY, LA, MD, MA, MI, MS, MN, MT, NE, NH, NJ, NY, NC, ND, OH, OK, OR, PA, TN, TX, UT, VA, WA, WV, and WI and the countries of: Australia, France, Canary Islands, Spain, Belgium, Germany, Switzerland, Netherlands, Italy, Guadeloupe, Ceuta, Austria, Slovenia, Croatia, Saudi Arabia, Turkey, Greece, and San Marino. | Terminated | March 05, 2014 | 1190 | 186 | Computer
Z-1830-2012 | 61915 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Two-Lumen Central Venous Catheterization Set with Blue FlexTip(R) ARROWgard Blue(R) Catheter |  |  | May 03, 2012 | 2012 | Arrow International Inc | Lidstock of product does not contain chlorhexidine contraindication and contains wording "contains no medication" on the label. Product insert does reflect medicated status. | Labeling mix-ups | Arrow International sent an Urgent Field Corrective Action letter dated May 3, 2012 to their US consignee. The letter identified the affected product, problem and actions to be taken. The letter instructed the consignee to quarantine any affected inventory and relay the information to their Arrow Sales Representative. The letter stated that the Arrow Sales Representative would affix new labels to each kit within their inventory. The consignee was instructed to sign the Acknowledgement Form provided by the Sales Representative which would then be submitted to Arrow. An Urgent Field Safety Notice letter dated May 15, 2012 was sent to foreign consignees with field safety corrective action instructions to follow. If you require additional information or clarification regarding this matter contact your sales representative. |  | 634912 | 634,912 | Worldwide Distribution - United States (nationwide) include the state of Texas. and the countries of: ARMENIA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, HUNGARY, INDIA, ITALY, IRELAND, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LEBANON, LITHUANIA, LUXEMBOURG, MONTENEGRO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, PORTUGAL, QATAR, RÂżUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, TURKEY, UKRAINE, , UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | March 22, 2013 | 323 | 634912 | Not_Computer
Z-1868-2012 | 62238 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | MRCTT80041 MedComp 8F Plastic Pro-Fuse(R) CT Low Profile Port with Attachable Chronoflex(R) Polyurethane Catheter. |  |  | April 02, 2012 | 2012 | Medical Components, Inc dba MedComp | Product was packaged with the incorrect port and does not match the label. The kit is labeled as an 8F Low Profile Pro-Fuse CT Port. The kit contains an 8F Midsize Dignity CT Port. | Packaging process control | MedComp sent a "PRODUCT ALERT" letter dated April 2, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The firm requested the return of all affected product. Contact the firm at 215-256-4201 ext. 2225 for questions regarding this notice. |  | 30 | 30 | Worldwide Distribution-USA (nationwide) including the states of GA and WA and the countries of Ireland and Turkey. | Terminated | May 20, 2013 | 413 | 30 | Not_Computer
Z-1896-2012 | 62333 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx monitor/defibrillators |  |  | June 20, 2012 | 2012 | Philips Healthcare Inc. | HeartStart MRx Unexpected Pads/Paddles ECG Failure | Process control | Philips issued an Urgent Voluntary Medical Device Correction/Field Safety Notice letter on June 20, 2012. The letter informs customers of the issue, identifies details (including models and software versions applicable) of units affected, gives instructions on actions to be taken until the correction is available, and identifies what action Philips plans to eliminate or remedy the issue. The letter states that a Philips Healthcare representative will contact customers to arrange for the replacement of the Power PCA in all the affected devices.Philips will complete the installation free of charge. For questions or support concerning this issue contact your local Philips representative at 1-800-722-9377. |  | 2914 | 2914 units | Worldwide Distribution - US (nationwide) and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHINA, COLOMBIA, EL SALVADOR, ETHIOPIA, FRANCE, GERMANY, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, LEBANON, LIBYAN ARAB JAMAHIRIYA, MEXICO, MOROCCO, NETHERLANDS, NETHERLANDS ANTILLES, NICARAGUA, NORWAY, PERU, PHILIPPINES, POLAND, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIET NAM. | Terminated | March 16, 2017 | 1730 | 2914 | Computer
Z-1869-2012 | 61889 | No | CDD | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) System B12 Flex(R) Reagent Cartridge (K6442) |  |  | May 08, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that the presence of Intrinsic Factor Blocking Antibody in some patients can result in a false and significant elevation in the reported concentration of Vitamin B12 on the Dimension Vista System. This issue affects all lots of B12 Flex reagent cartridges on the Dimension Vista System. | Device Design | Siemens Healthcare Diagnostics Inc. sent a Urgent Device Recall letter to all affected customers dated May 2012. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed customers to immediately discontinue use of the device, quarantine and dispose of all inventory, ensure that your Medical Director reviews this letter as well as the attached physician letter. Please complete the attached Field Correction form and fax to 302-631-8467. Further questions please call 302-631-6311. |  | 59735 | 59,735 | Worldwide Distribution -- USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Great Britian, Italy, Netherlands, Norway, Portugal, Slovakia, South Africa, Spain, Switzerland, United Arab Emirates, Canada, New Zealand, and South Korea. *** Center Recommended Depth is Retail*** | Terminated | April 03, 2014 | 695 | 59735 | Not_Computer
Z-1898-2012 | 62272 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Adult Divided Cannula end tidal C)2 sampling w/male luer-lock connector. |  |  | May 30, 2012 | 2012 | SALTER LABS | A complaint was received, and later verified, that product labeled as SO-1296 with male luer lock connector actually contained product with a female luer lock connector. | Packaging | The firm, Salter Labs, called their customers to discuss the recall and followed up with a "PRODUCT RECALL" letter dated May 29, 2012. The letter was distributed to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to examine their inventory and return any of the lots in their possession to Salter Labs; contact Salter Labs Customer Service at 1-800-235-4203 to arrange for return of the product, and complete and return the attached form/record via fax to: 661-854-6816 or Toll Free 1-800-628-4690.Should you have any questions, please contact Salter Labs at 1-800-235-4203. |  | 242 | 242 cases | Nationwide distribution: USA including states of: FL, IN, KS, PA and TX. | Terminated | July 27, 2012 | 58 | 242 | Not_Computer
Z-1902-2012 | 62017 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | MULTIGENT Acetaminophen for use with Architect/Aeroset |  |  | May 29, 2012 | 2012 | Sekisui Diagnostics Llc | Discoloration of the acetaminophen enzyme reagent (R1), generates a high calibration factor and erratic control recovery. | Storage | Sekisui Diagnostics Llc sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated May 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately stop using the affected product and to destroy any remaining inventory. A Confirmation of Notification Form was attached for customers to complete and return via fax to 902-628-6504. Contact the firm at 902-628-0984 for questions regarding this notice. |  | 1600 | 1600 kits | Illinois | Terminated | November 14, 2012 | 169 | 1600 | Not_Computer
Z-1897-2012 | 62056 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Ethicon Surgiflo Hemostatic Matrix Kit with Thrombin Plus Flex Tip |  |  | April 03, 2012 | 2012 | Ethicon, Inc. | The lyophilized thrombin component of the Surgiflo Hemostatic Matrix kit with Thrombin has the potential to be non-sterile due to an undetected air leak which occurred during the manufacturing process. | Packaging process control | Ethicon Inc. sent a Urgent Voluntary Product Recall letter dated April 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you have any product subject to this voluntary recall, please remove it from your inventory immediately and return it per the following instructions.Return product to Stericycle using the enclosed pre-paid shipping label.Shipping fees wil be charged directly to Ethicon through the use of this shipping label. No reimbursements will be made for the shipping charge if this label is not used. Extra shipping labels may be obtained by calling 1-866-338-2611. Credit will be issued for impacted product that is returned.Further questions please call 1-866-338-2611, or contact your Ethicon Biosurgery Sales Representative.Thank you for your coorporation and immediate assistance.. |  | 27396 | 27,396 kits | USA (nationwide) | Terminated | April 16, 2014 | 743 | 27396 | Not_Computer
Z-1871-2012 | 61670 | No | MMD | Powered toothbrush. | Dental | Dental | 510(K) Exempt | Colgate Adult Motion Whitening Battery Toothbrush |  |  | March 23, 2012 | 2012 | Colgate Palmolive Company | Brush battery cap was forcibly blown off at the end of the brush handle. | Nonconforming Material/Component | Colgate-Palmolive Company sent an "URGENT: MEDICAL DEVICE RECALL" letter dated March 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter requests the termination of sale and segregation of the affected product for return. Contact your Colgate-Palmolive representative to coordinate return of the affected product or if you have any questions concerning this matter. |  |  |  | Nationwide Distribution. | Terminated | June 19, 2013 | 453 |  | Not_Computer
Z-1895-2012 | 61952 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/Flatow The Complete Shoulder Solution Trabecular Metal Glenoid Instrument Drill with Stop, 6 mm; the products are medical instruments intended for use when performing surgery of the shoulder. |  |  | May 12, 2012 | 2012 | Zimmer, Inc. | Bigliani/Flatow The Complete Shoulder Solution Trabecular Metal Glenoid Instrument Drill with Stop, 6 mm. Catalog number 47-4301-031-00 was actually packaged and labeled as catalog number 47-4307-031-00, lot 62071190. | Mixed-up of materials/components | Zimmer sent URGENT DEVICE REMOVAL NOTIFICATION emails to distributors on May 12, 2012 to remove the product from all locations in their territories. This was an initial notification to prevent further distribution and use of the product. Zimmer Inc. sent follow-up notifications dated June 7, 2012 to the distributors, surgeons and hospital user accounts describing the product issue and providing recommended actions. Customers were directed to (877) 946-2761 for assistance or questions. |  | 87 | 87 units | Nationwide: TX, FL, MS, MI, OH, KS, TN, KY, OK, UT, CO, CA, IL, NC, NJ, ND, and WI. | Terminated | September 12, 2013 | 488 | 87 | Not_Computer
Z-1875-2012 | 61731 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Antegrade Cardioplegia Cannula: ROOT INFUSTION VENT CATHETER WITH 12 GAUGE BLUE NEEDLE RECTANGULAR FLANGE, 6" (15 CM) LONG |  |  | May 09, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During production of the Sarns Antegrade Cardioplegia Cannula, Terumo Cardiovascular Systems (Terumo CVS) identified the presence of a foreign substance on the inner surface of some cannulae tips.Terumo CVS' preliminary investigation found that the substance can be dislodged from the cannula surface. The substance was likely deposited during the molding process, but the exact composition of the | Material/Component Contamination | Terumo sent an Urgent Medical Device Recall letter dated May 9, 2012 to All US first consignees will receive a removal notification letter via Federal Express. The letter identified the affected products, reason for correction and actions to be taken. Customers were instructed to discontinue use of affected products and return for replacement. The letter advised customers to assure that all users are aware of this notice and to fill out and return the attached Response Form. For questions call Terumo CVS Customer Service at 1-800-521-2818 or by fax at 1-800-292-6551. |  | 53139 | 53,139 (each) in total | Worldwide Distribution - US (nationwide) including the states of: AL, AR , AZ , CA, CT , FL, IA , IL, IN , KS , LA , MI , NM, MO , NC, NE , NJ , NV, NY , OH, OK , PA, RI , SC , TN, TX, WA , WI, WV and the countries of: BELGIUM, BRAZIL, Byelorussian SSR, CAMBODIA , Dominican Republic, Honduras, Japan, Malaysia, Mexico, Nicaragua, Saudi Arabia, Singapore, South Africa, Thailand, Trinidad and Tobago, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | April 24, 2013 | 350 | 53139 | Not_Computer
Z-1903-2012 | 61966 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels Creatinine test strips. |  |  | May 15, 2012 | 2012 | Polymer Technology Systems, Inc. | During a routine product investigation, a creatinine sample with a value less than 1.0 mg/dl was tested on a version 2.60 CardioCek PA meter and the meter provided a result of 0.08mg/dl, which is not the reportable measuring range for the CardioChek PA meter for the creatinine assay. | Software design | Polymer Technology Systems, Inc sent a URGENT FIELD CORRECTION letter dated May 15, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed the customers to discontinue use and dispose of the product. PTS Inc. will provide appropriate replacement for any product that is properly identified on the attached Customer Acknowledgement and confirmation letter,If you have any questions please contact Customer Service toll-free at 1-877-870-5610, or +1-317-870-5610 ( outside the US ) |  | 1650 | 1650 vials | Worldwide Distribution -- Nationwide Distribution including the states of CA and VA., and the countries of China, Poland, Ireland and France. | Terminated | July 20, 2012 | 66 | 1650 | Computer
Z-1901-2012 | 61943 | No | BZE | Breathing system heater. | Anesthesiology | Anesthesiology | 510(k) | AirLife Isothermal Breathing Circuit Accessory - Infant Patient Y Connector; an Rx device for single patient use; 30 connectors per box; CareFusion, 22745 Savi Ranch Parkway, Yorba Linda, CA 92887 USA, Made in Mexico |  |  | May 29, 2012 | 2012 | Carefusion 2200 Inc | The Y adapter within the breathing circuit may spontaneously crack, causing a leak in the closed ventilation system. This leak could lead to a decrease in the tidal volume delivered to the ventilated patient. | Equipment maintenance | The firm, CareFusion, sent an "URGENT RECALL NOTICE" dated May 29, 2012 to all of the affected customers on the same date via registered return receipt mail. The notice described the product, problem and actions to be taken. The customers were instructed to immediately destroy any affected product in-stock at their facility and to NOT further distribute these products. NOTE: The notice stated "Carefusion is removing from the the market all lot numbers less than lot #0000414257 or begin with the letter 'Y' for all Infant Breathing Circuit product codes listed in the provided attachment".The customers were also instructed to not return the product to Carefusion. Return the product directly to: Owens and Minor Health Care Logistics at 1651 California Street, Suite C, Redlands, CA 92374, Attn: Recall Coordinator, for immediate replacement.NOTE: The notice the stated "Failure to follow these important measures may result in a potential patient safety risk".In addition, the customers were instructed to forward the recall notice to any customers to whom these products were sold, notifying them of the potential risk, advising them to discard or return any CareFusion AirLife Infant Breathing Circuits with the noted product/lot numbers for replacement; and to complete and return a enclosed reply form indicating the amount of inventory in stock, amount destroyed and amount returned. The distributors were requested to complete and return the attached Acknowledgement Form indicating that they have received the recall notice and that they communicated the Recall Notice to all customers who have purchased the affected breathing circuits via Fax to (847) 473-7774 or email to David.Wehrheim@Carefusion.com. They were also instructed to contact CareFusion Channel Management at (866) 802-6320 for further details on return and/or replacement of stock.If you have any questions related to this information, please call (847) 473-7495 or email David.Wehrheim@Caref |  |  |  | Worldwide distribution: USA (nationwide) including Puerto Rico and internationally to: Canada, Egypt, France, Germany, Great Britain, Italy, Kuwait, Mexico, Portugal, Saudi Arabia, Spain, Taiwan and United Arab Emirates. | Terminated | January 22, 2016 | 1333 |  | Not_Computer
Z-1924-2012 | 61749 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LUSTER HIP SZ 10 |  |  | May 16, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a Field Safety Notice for the LUSTER Cemented Polished Femoral Hip Stems and the ULTIMA LX Cemented Polished Femoral Hip Stems that were implanted with any metal-on-metal (MoM) articulation (including a head of CoCr and an acetabular cup where the cup is of an all metal construction (monoblock or modular)).The LUSTER Cemented Polished Femoral Hip Stems and the U | Device Design | DePuy sent an "URGENT RECALL INFORMATION-FIELD SAFETY NOTICE" letter dated May 16, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 574-372-7333 for questions regarding this notice. |  | 69 | 85) LUSTER Cemented Polished Femoral Hip Stems in 69 lots | Nationwide Distribution-including the states of CA, CO, IA, NC, OH, TX, WA, WI, and WV. | Terminated | March 05, 2013 | 293 | 5064 | Not_Computer
Z-1905-2012 | 61555 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX SP Power Wheelchair, TDX SR Power Wheelchair, TDX SI Power Wheelchair. Model Number(s): TDXSI, TDXSI-2, TDXSI-2-S, TDXSI-CG, TDXSI-HD, TDXSI-HD-S, TDXSIV, TDXSIV-2, TDXSIV-2-S, TDXSIV-HD, TDXSIV-HD-S, TDXSP, TDXSPBASE, TDXSP-CG, TDXSPEURO, TDXSP-MCG, TDXSP-MCG-GT, TDXSR, TDXSR-CG, TDXSR-CG-HD, TDXSR-MCG, FDX, FDX-CG, FDXEURO-14, FDX-MCG. Mrf. By: Invacare Corporation, 1200 Taylor Street, Elyria, OH 44035. |  |  | April 06, 2012 | 2012 | Invacare Corporation | Invacare Corporation decided to recall the product because of a potential risk of fire and serious injury. The post fuse assemblies could potentially have a short between the positive ring terminal and the terminal battery bracket which would create an alternate current path around the fuse. | Employee error | Invacare sent an "URGENT: RECALL INFORMATION" letter dated June 7, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. THe letters instructs customers to examine their inventory for the affected product and quarantine any product found. AProvider Response Form was enclosed for customers to complete and return via fax to 888-345-0289. Contact the firm at 888-641-9738 for assistance with this recall. |  | 680 | 665 affected power wheel chairs and 15 post-fuse assembly service parts have been identified. | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Puerto Rico and the countries of Australia, Germany, New Zealand. | Terminated | May 20, 2014 | 774 | 680 | Not_Computer
Z-1904-2012 | 61979 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ARCHITECT CA 19-9XR Reagent Kit; ARCHITECT CA 19-9 Reagent Kit, list 291-20 (4 x 100 tests), 2K91-25 and 2K91-27 (1 x 100 tests); |  |  | May 30, 2012 | 2012 | Abbott Laboratories | The six ARCHITECT CA 19-9XR reagent lots are demonstrating a shift up in patient results. The Abbott Controls however do not detect the upward shift and will test within range. | Nonconforming Material/Component | The firm, Abbott Laboratories, sent a "Product Recall Immediate Action Required" letter dated May 30, 2012 to all Architect CA 19-9XR customers who received the affected reagent lots. The customers were directed to immediately discontinue use and destroy any remaining inventory of the lots listed; to review the patient results obtained with the affected lots with the treating physician or heath care provider to ensure they are aware of the potential analytical issue and can determine whether re-evaluation of the patient is necessary; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; and to retain a copy of the letter for their records and to complete and return the Customer Reply Form via fax to: 1-800-777-0051 or e-mail: QAGCO@abbott.com. For questions were call your local Customer Service representative or call the field action coordinator at 847-938-1923. |  | 8672 | 316 kits - USA, 8,356 kits ROW | Worldwide Distribution - US (nationwide) the states of: Alabama, Arkansas, Arizona, California, Florida, Iowa, Indiana, Louisiana, Michigan, Minnesota, New Jersey, Nevada, New York, Ohio, Oklahoma, Pennsylvania, South Dakota, Texas, Virginia, Washington, including Puerto Rico, and the countries of Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Barbados, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, Estonia, Finland, France, Gaza &amp; Jericho, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kyrgyzstan, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Open, Classified |  |  | 8672 | Not_Computer
Z-1817-2012 | 61253 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System Console. Model: TX1 |  |  | February 22, 2012 | 2012 | American Optisurgical Inc | Two reasons for recall.1. Incorrect pressure relief valve installed inside console, leading to a remote chance of a risk to health. 2. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect. | Nonconforming Material/Component | The firm, American Optisurgical Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter via email on February 23, 2012. An updated "URGENT: MEDICAL DEVICE RECALL" recall letter dated March 1, 2012 was sent via email on March 1, 2012 to its customer. This updated recall letter described the product, problem and actions to be taken. The letter notified the two reasons for recall: " An incorrect pressure relief valve has been installed inside the console. Should the regulator that controls pressure delivered to the inflation cuff fail, the pressure relief valve would not open until a pressure of 10psi was achieved, which may present a health risk". "The Operator's Manual included with the system, Rev. 1, has a typographical error. Specifically, in Section E. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect". The updated recall included the above second reason for recall. The firm instructed the customer, for replacement of the pressure relief valve, and to return the revision 1 operator manual to the company's address using a pre-paid FedEx account. For questions call 949.580.1266 M-F, 7:00am to 5:00pm PST. |  | 5 | 5units | Nationwide distribution: MN only. | Terminated | June 15, 2012 | 114 | 5 | Not_Computer
Z-1793-2012 | 61332 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Tibia Plateau (component of the Columbus REVISION Knee System) |  |  | February 21, 2012 | 2012 | Aesculap Implant Systems | During an internal quality inspection, the tibia plateau was found to be out of specification. There is a potential for misalignment between the tibia plateau and attachment of the extension stem. | Nonconforming Material/Component | Aesculap Implant Systems (AIS) initiated their recall of this product by first telephoning consignees and then on February 21, 2012 by sending an IMPORTANT RECALL NOTIFICATION letter dated February 21, 2012 to all consignees. The letter identified the affected product, problem and actions to be taken. The notification told consignees not to use the affected product and to return any stock to AIS (USA) promptly. For questions contact Scott Stephan at (610) 984-9239 or Christian Gabriel at (610) 984-9300. |  | 8 | 8 | Nationwide Distribution including PA and TX. | Terminated | January 15, 2014 | 694 | 8 | Not_Computer
Z-0857-2012 | 60853 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | Arm and Hammer Spinbrush Dual Action Powered Toothbrush Pro-Clean, various models. |  |  | December 22, 2011 | 2011 | Church &amp; Dwight Inc | Church and Dwight plans to take actions to best convey to the public the potential small parts choking hazard with using a Spinbrush beyond the recommended 3-month brush head life. | Device Design | Church &amp; Dwight Co., Inc. proposed a consumer education media campaign starting on December 22, 2011 with the following statement:"Safety Notice: Please remember to replace your brush head after 3 months of use, or if the brush is damaged or the parts become loose. Extended usage, loose parts or excessive wear could lead to brush head breakage, generation of small parts and possible choking hazard. Inspect the brush for loose parts before use."The campaign will be conducted in the following media and anticipate that the combined consumer reach of all of these vehicles to be approximately 144 million people an average of 4 times.- Television campaign will be run during the entire Q1 of 2012 (January -March). -Print ads that will appear in the December 30th issues of People, Star, US Weekly and Entertainment Weekly. -Church &amp; Dwight Website will contain the Safety Alert notice starting December 22nd. (www.spinbrush.com)Consumer Relations IVR (Interactive Voice Response) to consumers calling the Spinbrush toll-free numbers). (800-353-3384 and 800-5610752).4-30-121) A print advertisement (Attachment A) containing the Sonic Safety Notice that will appear in each magazine (Star, People, US Weekly, and Entertainment Weekly) as noted:Five total print insertions are planned (on sale dates listed):I. 4/30 -Star2.5/14 -Us Weekly3.5/28 -People4.6/11 -Entertainment Weekly5.6/25 -Star2) A disk containing the 15 second commercial containing the safety notice.3) A document (Attachment B) of the National TV slots showing when the commercials will run in the quarter, and another with the shows during which they will air. TV ads will begin running the week of April 16th. The DRTV schedule is determined weekly and we are informed after the fact when the ads where shown. The DRTV stations are included at the bottom of the National TV listing." In total, combined print and TV are estimated to reach 163MM impressions (measurement target: women |  | 22783000 | 22.783 million units | Worldwide Distribution. | Terminated | April 14, 2015 | 1209 | 41417300 | Not_Computer
Z-1927-2012 | 62283 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | AirLifeÂż MiniSpacerÂż MDI Dispenser Adapter. Dual-Spray MiniSpacerÂż Metered Dose Inhaler (MDI) Dispenser/Adapter with 15 mm connections. |  |  | June 06, 2012 | 2012 | Thayer Medical Corporation | Internal data identified a molding defect in the 15 mm O.D. / I.D. MiniSpacer meter dose inhaler (MDI) adapter that results in an incomplete fill of the aerosol dispensing pathway. | Device Design | Thayer Medical sent an Urgent Medical Device Recall letter dated June 7, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Customers were also instructed to destroy and/or dispose of the product in a proper manner and certify to Thayer Medical that they have done so by complelting the attached Certificate of Destruction. The Certificate of Destruction should be sent to Thayer Medical via fax or email. If customers prefer to return the product to Thayer Medical they should contact Customer Service at info@thayermedical.com or fax 1-520-790-5854 for instructions.For questions regarding this recall call 520-790-5393, ext 2210. |  | 19340 | 19340 | Worldwide Distribution - USA including CA, KY and Europe. | Terminated | January 15, 2013 | 223 | 19340 | Not_Computer
Z-1936-2012 | 61998 | No | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics 217 Osseo Avenue North, St. Cloud, MN 56303, Gram Stain Control Slides, SLO3-10, ATCC Licensed Derivative, MediMart Europe, 11 rue Emile, Zola B.P. 2332, 38933 Grenoble Cedex 2 France. |  |  | May 29, 2012 | 2012 | Microbiologics Inc | Microbiologics is recalling a number of lots of KWIK-QC Gram Stain Slides. This product gram negative control organism (E.coli) is not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could potentially lead to misdiagnosis and improper treatment of a patient. | Nonconforming Material/Component | Microbiologics sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 30, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers not to use or distribute the affected product and to quarantine and return to the firm any remaining product. A Customer Response Form was attached for customers to complete and return to the firm. Contact Customer Service at 320-229-7057 for questions regarding this recall. |  | 13720 | 13,720 slides (10 slides per box) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, KY, ME, MA, MI, MN, MS, NJ, NM, NY, PA, PR, SC, TX, VT, and VA. and the countries of SWITZERLAND, ITALY, ISRAEL, IRELAND, COLOMBIA, SWEDEN, INDIA, SINGAPORE, SERBIA, SAUDI ARABIA, CANADA, TAIWAN, SOUTH AFRICA, FRANCE, BRAZIL, MALTA, MACEDONIA, MALAYSIA, FINLAND, ECUADOR, VENEZUELA, REPUBLIC OF KOREA, NETHERLANDS, PERU, UNITED ARAB EMERITES, TURKEY, POLAND, and LEBANON. | Terminated | January 08, 2013 | 224 | 13720 | Not_Computer
Z-1941-2012 | 57350 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Victory 9 with Power Seat (SC609PS) |  |  | November 26, 2010 | 2010 | Pride Mobility Products Corp | The main harness connector was manufactured without the required back shield sealing protection. As a result, there is the potential for water to seep through the nonconforming harness. | Nonconforming Material/Component | Three (3) pieces of information was distributed in the notification. They are: (1) Letter of Correction to the Provider (2) Letter of Correction for the Consumer (provider to issue to consumer) (3) Product Insert for Victory 9 Controller Interface Harness Replacement. The letter of Correction for Victory 9 with Power Sear (SC609PS), dated November 2010, identified the affected product and the reason for the correction. Customers were to record the serial number(s) and contact their provider's office to make an appointment to have the harness replaced. |  | 186 | 186 | Worldwide Distribution | Terminated | March 22, 2013 | 847 | 186 | Not_Computer
Z-1933-2012 | 62137 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinder NXT Pivoting Percutaneous Rod Holder |  |  | June 05, 2012 | 2012 | Zimmer, Inc. | Complaints have been reported where, during surgery, the surgeon was unable to loosen orremove the inserted rod from the rod holder instrument. The complaint investigation showedthe hex mating feature of the set screw was broken and material missing. There were noreports of fragments of the set screw being left in the patient or that there was any patient oruser injury, but these events we | Device Design | The firm, Zimmer Spine, sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 5, 2012 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the notification and ensure they are aware of the content; to immediately remove and discontinue use of Pivoting Percutaneous Rod Holder (part number 3564-1) from Pathfinder NXT instrument sets in their possession and follow the PathFinder NXT Surgical Technique; to notify their Zimmer Spine customer service representative to obtain return authorization; and return the complete Surgeon Contact Certification Form within 10 working days from receipt of notification to ronald.musselman@zimmer.com or Fax to 1-512-258-0995.If you have any questions, call Zimmer Spine Customer Service at 866-774-6368. |  | 77 | 77 units | Worldwide distribution: USA (nationwide) and countries of: Australia, France and Germany. | Terminated | February 14, 2013 | 254 | 77 | Not_Computer
Z-1934-2012 | 62240 | No | HWW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S4 Cervical (S4C) Revision Instruments (FW099R and FW101R). The S4C Revision Instruments assist in the removal of an occipital cervical screw during revision surgery. |  |  | March 07, 2012 | 2012 | Aesculap Implant Systems LLC | Aesculap Implant Systems, LLC has initiated a voluntary removal of the S4 Cervical (S4C) Occiput Revision Instrument that is part of the S4C Occipital Instrument Set. The instruments may not thread into the screw properly during a revision surgery due to tolerance issues between the two instruments. | Component design/selection | The firm initiated their recall of this product by issuing an "Important Product Removal Notification" letter to all consignees dated March 7, 2012. The letter described the problem, advised them not to use the product, and instructed consignees to return the instruments to the recalling firm with the attached Distribution Inventory Sheet. Customer Service may be contacted if questions at 1-866-229-3002. |  | 15 | 15 | Nationwide distribution, including the states of CA, CO, IL, LA, NV, OH, and TX . | Terminated | August 08, 2012 | 154 | 15 | Not_Computer
Z-1940-2012 | 61874 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Depuy Spine Side Fire Needles 13G X 4IN Side-Fire Needle, part of the Confidence Spinal Cement System |  |  | May 04, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Additional Instructions for Use and Surgical Technique for the Confidence Needle due to reports of breakage with the side fire needles | Labeling design | Depuy Spine sent a "CORRECTION NOTICE" to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Notification Form was attached for customers to complete and return via fax to 508-828-3070. For questions about Confidence Needles, please contact your local DePuy Spine Sales Representative. If you need assistance in locating your local Sales Representative please contact JJHCS CIS @ 1 -800-255 -2500 . |  | 6979 | 6979 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Austria, China, Croatia, Czech Republic, France, Germany, Greece, Hungary, Italy, Lebanon, Netherlands, New Zealand, Poland, Russian Federation, Singapore, South Africa, Spain, Switzerland, Turkey, UAE, and UK. | Completed |  |  | 8896 | Not_Computer
Z-1935-2012 | 62447 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Vacuette Quickshield Complete Plus, 21G x 1 (0.8 x 25 mm), Greiner bio-one. |  |  | April 02, 2012 | 2012 | Greiner Bio-One North America, Inc. | : The Quickshield Complete Plus may have a defect (split at the base), which can cause the green safety shield attachment of the Quickshield holder to be loose and unstable. | Process control | greiner bio-one sent a notification letter dated April 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached Product Disposition Site Confirmation form and fax to 800-726-0052 or email to patech@gbo.com. This form will confirm the affected product has been discarded/destroyed. The firm will replace the product after the completed form is returned to them. Customers with questions should call 800-515-8112.For questions regarding this recall call 704-261-7823. |  | 263.5 | 19.5 cases (5,850 pieces) and 244 cases (73,200 pieces)= 263.5 (73258 pieces total) | Nationwide Distribution including IL, MS, OH, PA, TN and VA. | Terminated | February 27, 2013 | 331 | 263.5 | Not_Computer
Z-1943-2012 | 61525 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Single Shot Epidural Trays, Sterile, Integra Pain Management, West Valley City, Utah 84119, each kit contains sterile instruments, drugs and other medical materials that will support the administration of spinal anesthesia for use by a surgeon, nurse or other licensed medical professional. Model number is 3403644. |  |  | June 07, 2012 | 2012 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is recalling one lot of Single Shot Epidural Trays as the product was inadvertently shipped missing injectable components. | Packaging process control | Integra sent an Urgent - Medical Device Recall letter dated June 7, 2012, to one domestic consignee via Federal Express with a recall acknowledgement form followed up with a telephone call. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the Recall Acknowledgement Form including quantity of affected product in their inventory. Customers were also instructed to contact Integra Pain Management Customer Service to arrange for returns of all affected product. Product reconciliation will be performed to account for all affected products returned. Returned product will be reprocessed to correct for the missing components.Customers with questions were instructed to call Integra Pain Management Customer Service at 1-800-241-2210.For questions regarding this recall call 609-936-2311. |  | 40 | 4 cases x 10 units per case | Nationwide Distribution including MA | Terminated | August 07, 2012 | 61 | 40 | Not_Computer
Z-1942-2012 | 62270 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips Xcelera Connect, Software R2.1 L 1 SP2, an interface engine for data exchange. |  |  | June 14, 2012 | 2012 | Philips Healthcare Inc. | Xcelera Connect R2.1 L 1 SP2 , incomplete information arriving from unformatted reports interface | Software design | Philips Healthcare sent an "URGENT-FIELD SAFETY NOTICE" letter dated June 1, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Philips representative will contact customers when the upgrade becomes available. Contact your local Philips Representative for questions concerning this issue. |  | 226 | 226 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Chile, Denmark , Estonia, Finland, France, Germany ,Hong Kong, Ireland, Italy, Malaysia, Netherlands, Poland, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Taiwan, United Kingdom, and Utd. Arab. Emir. | Terminated | March 14, 2017 | 1734 | 226 | Not_Computer
Z-1957-2012 | 60341 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Central Control Module for System 1: System 1 Base, 100/120V |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo CVS has received reports of the Central Control Monitor (CCM) for Terumo Advanced Perfusion System 1 losing partial or full functionality. Symptoms that have been associated with a malfunctioning CCM include: full or partial loss of display, loss of control functions; inability to distinguish various status, alert, alarm conditions, the inability to operate the touch screen, or the CCM | Device Design | The firm, TERUMO, sent an "URGENT MEDICAL DEVICE RECALL" notice dated June 22, 2012 and addendum to the Operator's Manual via Federal Express to its customers. The letter described the product, problem and actions to be taken. The notice states to the customers "Please note that we are not recommending that you stop using your Central Control Monitor". The customers were instructed to review the Medical Device Recall notice; assure that all users are aware of this notice; place the Addendum at the beginning of the Central Control Monitor Section (page 4-1) of the Operator's Manual and complete and return the attached Customer Response Form to the fax number indicated on the form. If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818; customer service hours are from Monday-Friday, 8 AM - 6 PM ET. |  | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries including: Argentina, Belgium, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates (UAE), and Vietnam. | Terminated | March 06, 2013 | 257 | 1640 | Not_Computer
Z-1950-2012 | 62470 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Drager PT 4000 Phototherapy System |  |  | July 02, 2012 | 2012 | Draeger Medical Systems, Inc. | A routine internal quality check found that the Drager PT 4000 Phototherapy System is fitted with a 16 Amp rated fuse instead of the specified 1.6 Amp rated fuse. This may mean that the internal fuse would not trip promptly in the event of an internal short-circuit. The possible consequence is the wall outlet circuit breaker would trip first on the device may severely overheat. There have not b | Other | Drager Medical Systems sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 2012 to all affected customers. The letter identifies the product, problem, and actions to bee taken by the customers. In addition, a Customer Reply Card was included for customers to complete and return. Contact the firm at 800-543-5047 for questions regarding this recall. |  | 65 | 65 | Nationwide Distribution-USA (nationwide) and Puerto Rico. | Terminated | July 01, 2013 | 364 | 65 | Not_Computer
Z-1945-2012 | 61807 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Dual Incu i, (Atom Infant Incubator Model 100) Atom Medical Corporation 3-18-15 Hongo, Bunkyo-Ku, Tokyo, Japan 113-0033 Usage: An incubator for infants. |  |  | May 01, 2012 | 2012 | Atom Medical Corporation | There is a potential for the device's canopy to move unintentionally when using the incubator mode or radiant warmer mode. | Device Design | Atom Medical sent an Urgent Field Safety Notice dated May 1, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately identify all affected product in their inventory and quarantine them until the rework was complete. Customers were also asked to identify any customers to whom they further distributed the product and to contact their customers at once and arrange a time in whcih the filed service technician could perform the rework. Customers were asked to return the Recall Response Forms to Atom Medical (kirautc@atomed.co.jp, tsugino@atomed.co.jp,or matsumoto@atomami.co.jp). All forms should be completed and returned by January 31, 2013.For questions regarding this recall call 512-327-9997. |  | 377 | 377 units | Worldwide Distribution - USA including PA. Internationally to Japan, Kuwait, Saudi Arabia, Lenanon, Italy, Spain, China, East Europe, Slovakia, Poland, Czech Republic, Ukraine, Egypt, Russia, Algeria, Colombia, Germany, Malaysia, Switzerland, Australia, Europe, Portugal, Iran, Israel, Philippines, UAE, Korea, and Netherlands | Terminated | January 31, 2013 | 275 | 377 | Not_Computer
Z-1946-2012 | 61832 | No | DSK | Blood pressure computer. | Cardiovascular | Cardiovascular | 510(k) | Safedraw Venous Close Loop System, Blood Sampling Set, packaged in a tyvek bag, Model/Part Number 688695, Argon Critical Care Systems, Singapore Pte. Ltd., 198 Yishun Avenue 7, Singapore 768028. |  |  | May 04, 2012 | 2012 | Argon Medical Devices, Inc | Some of the kits may exhibit a potential defect with the pouch seal. The pouch issue may compromise the sterile barrier. | Packaging | Argon Medical Devices sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Recall Response Form was attached to the letter for customers to complete and return to the firm via fax to 903-577-9393. Contact the firm at CRS@argonmedical.com for questions regarding this recall. |  | 10 | 10 pieces | Distributed only in Missouri. | Terminated | January 14, 2013 | 255 | 10 | Not_Computer
Z-1956-2012 | 61793 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | CIVCO VirtuTRAX Instrument Navigator, Reusable non-sterile general purpose driveBAY, trakSTAR, electromagnetic sensor (8mm) with VirtuTRAX Instrument navigator and (8.9 tapered to 1.3 x 147cm) telescopically-folded CIV-Flex sensor covers (5), Ref. 610-1080. The firm name on the label is CIVCO, Kalona, IA. The kit contains sensor component part number 410-240. |  |  | May 07, 2012 | 2012 | Civco Medical Instruments Inc | The general purpose electromagnetic sensor was incorrectly calibrated causing potential variability in the projected position of the needle path with the ultrasound device. | Component design/selection | CIVCO sent an Urgent Medical Device Recall letter dated May 3, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return all general purpose sensors as the product may not function as intended. To return the affected product customers should contact the OEM Sales Coordinator at 319-248-6521 to schedule the return of their device. Upon receipt of the snesor, CIVCO will arrange for recalibration. Once recalibration is completed, the sensor will be returned to the customer at no charge. Customers were asked to forward this communication to all their customers who purchased the affected product. Once CIVCO has received the customer list, they will contact the end-user to request they return their general purpose sensor directly to CIVCO for recalibration. For questions customers should call 800-441-6741.For questions regarding this recall call 319-656-4447. |  | 24 | 24 kits | Worldwide Distribution - USA including WI, CA, MA, IN, WA, and NJ. Foreign distribution was made to Germany, Canada, Denmark, Hong Kong, Ukraine, Japan, Italy, Korea, and France. | Terminated | January 07, 2015 | 975 | 384 | Not_Computer
Z-1951-2012 | 61529 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AEM Disposable Electrodes, Device, Electrosurgical, Cutting and Coagulation and Accessories. AEM Disposable Electrodes used as electrosurgical accessories intended for ablation, removal, resection and coagulation of soft tissue in open, endoscopic and laparoscopic surgical procedures channeling monopolar high-frequency electrical current from compatible electrosurgical generators. The device is model FC03XX series. |  |  | June 15, 2012 | 2012 | Encision, Inc. | Encision is performing a field correction to their AEM Disposable Electrodes due to two consumer complaints reporting unanticipated alarms from the AEM monitor while using the device. This issue could result delays or complications during surgical procedures. | Device Design | Encision sent an Urgent Medical Device Recall letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to lubricate the contact on the instrument cord with a conductive gel (to be supplied by Encision) prior to making the connection of the cord to the instrument. Instructions for applying the gel were attached. This corrective action was temporary; allowing customers to use their inventory of devices and for new product in the short term, until a design correction can resolve the problem. Customers would be contacted by their Encision Sales Representative to review this information for their specific situation and train their staff on its correction. For questions customers should call 303-339-6917 or Customer Service at 800-998-0986.For questions regarding this recall call 303-339-6917. |  | 48 | 48 boxes | Worldwide Distribution - USA including CA, TX, GA, NY, NC, OH), and New Zealand. | Terminated | August 07, 2012 | 53 | 48 | Not_Computer
Z-1925-2012 | 62118 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | FLOW-I Anesthesia System |  |  | February 20, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | A technical alarm may be generated on the FLOW-i system when using the MAN/AUTO switch to change ventilation modes from manual to automatic or from automatic to manual if the switch is not fully engaged in the "on" or "off" position but remains in the "in-between position. | Device Design | The firm, Maquet Getinge Group, sent an "URGENT DEVICE FIELD CORRECTION" letter dated February 20, 2012 to its customers. The letter describes the product, problem and action to be taken. The letter state that a Maquet Service Representative will upgrade the customers software in the FLOW-i Anesthesia System to the current revision and also provide the customers with the corresponding user's manual. Note: This upgrade includes other minor improvements. The customers were instructed to sign a document to verify satisfactory completion of the work once the software upgrade was completed.If you have questions or require additional information, please contact your local MAQUET representative or Technical Support at 888-627-8383.A press release was issued on 7/6/2012. |  | 284 | 284 units | Worldwide distribution: USA (nationwide) state of GA only and countries of: Argentina, Australia, Austria, Belgium, China, Czech Republic, Ecuador, Finland, France, Germany, India, Ireland, Israel, Italy, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Serbia, Singapore, South Africa, Spain, Sweden incl Maquet Critical Care, Switzerland, Thailand, United Arab Emirates, and United Kingdom. | Terminated | September 09, 2015 | 1297 | 284 | Computer
Z-1929-2012 | 61927 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage Cardiac Panel PN 97000HS |  |  | May 22, 2012 | 2012 | Alere San Diego, Inc. | Certain lots of the affected products may have significantly decreased precision relative to the package insert, which could result in an increased frequency of false positive or false negative results. These false positive and false negative results are unpredictable within lots and may not be detected by Quality Control testing. | Process design | Phase 1 of the recall included a recall communication being initiated on May 22, 2012 with Alere forwarding a recall letter to all their customers who purchased the Triage CardioProfiler Panel, Triage Cardiac Panel, Triage Profiler SOB Panel, Triage BNP, and Triage D-dimer. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to discontinue use of the affected product, and to complete and fax the enclosed verification form within 10 days. Phase 2 of the recall included a recall communication letter being forwarded to customers on 6/11/12. The second letter included additional lots of Triage products that go beyond the lots included in an earlier recall notification letter. The letter informed the customers that the additional lots have significantly decreased precision relative to the package insert which could result in an increased frequency of false positive or false negative results. A third letter dated June 12, 2012 included 3 more lots for the Triage Cardio Profiler Panel PN 97100CP and Alere Triage Profiler SOB Panel PN 97300. The letter informed the customers that the additional lots have an increased frequency of Troponin I results &gt;0.05 ng/mL for samples which are found to be below 0.05 ng/mL upon additional testing.Customers with questions about the information contained in the notifications, were instructed to contact Alere San Diego at (877) 308-8287 or at 9975 Summers Ridge Road, San Diego, CA 92121. |  | 6044 | 6,044 kits | Nationwide Distribution | Terminated | October 18, 2013 | 514 | 60074 | Not_Computer
Z-1948-2012 | 62116 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Chorus Mobile Patient Lifter; manufactured by Arjo Ltd., St. Catherine Street, Gloucester, GL1 2SL, England, U.K.; Models KKB5060-21US, KKB5062-21US and KKB5080-21US. |  |  | June 29, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | ArjoHuntleigh has received a limited number of customer reports concerning injuries related to an exposed metal support bracket when the proactive knee pad assembly is removed. | Labeling False and Misleading | ArjoHuntleigh sent an "URGENT FIELD CORRECTION RECALL" notice dated June 25, 2012, to all affected customers via UPS Ground on June 29, 2012. The notice informed the accounts of the potential for severe patient injury if the proactive knee pad support is removed from the lift without the removal of the knee pad support bracket. The accounts were reminded that the entire knee pad and support bracket assembly must be removed from the lift when required to do so for exercise. An additional warning statement with diagrams was included in the notice. The accounts were requested to disseminate the information to all users of the affected lifts, to include a copy of the notice in the Instruction for Use Manuals for their lifts, and to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 3234114 | 1,989 units in the U.S., 1,626 units in the rest of the world | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Belgium, Bermuda, Canada, Cyprus, Czech Republic, Denmark, Dubai, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Hong Kong, Iceland, Ireland, Israel, Italy, Japan, Kenya, Korea, Luxembourg, the Netherlands, New Zealand, Northern Ireland, Norway, Poland, Portugal, Quatar, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland and Turkey | Open, Classified |  |  | 22492034 | Not_Computer
Z-1944-2012 | 62302 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information Systems interfaced to Elekta Treatment machines. |  |  | June 08, 2012 | 2012 | Impac Medical Systems Inc | Impac Medical became aware of a problem with the recording of the Wedge Motor Units in MOSAIQ when it resumes after a machine terminates/faults, the remaining Wedge MU was being erroneously set to zero. | Software design | A User Notice began being distributed on June 8, 2012. This will be followed by contact at each site by IMPAC support to determine if the site is affected by this issue, and if so, the upgraded version will be suggested. The Notice identified the affected device, the problem, the clinical impact, along with the interim mitigations and the final resolution. Users are encouraged to monitor daily delivery and recordings of treatments to be sure that the Wedge Motor Units treatments are properly delivered. The letter also recommends that treatment charts be checked weekly and that users not override any parameters in MOSAIQ until the machine setup has been completed. Users are to contact Elekta Support for the upgrade. |  | 11 | 11 affected sites | Worldwide Distribution -- USA, including the states of VA, CT, GA, NJ and NY and the countries of Belgium, Denmark, France, and the United Kingdom. | Terminated | June 19, 2013 | 376 | 11 | Not_Computer
Z-1979-2012 | 62389 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical HOTLINE 3 Blood and Fluid Warmer, |  |  | June 27, 2012 | 2012 | Smiths Medical ASD, Inc. | HOTLINE 3 Blood and Fluid Warmer( Reorder No. HL-390 ) disposable is not correctly installed , the microswitch may become damaged and unusable | Device Design | Smiths Medical sent an Urgent Medical Device Correction -Field Safety Notice letter dated June 21, 2012 via U.S. Mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter included an Urgent Medical Device Field Safety Notice Acknowledgement Form ("Confirmation Form") that customers were instructed to complete and return to Smiths Medical for carrying out the correction and affectivity tracking purposes. The letter stated that a Smiths Medical Representative will visit the affected sites to correct the units with a shelf upgrade kit, which includes instructions and all the parts necessary to make the correction to their affected devices. Customers were provided with a form to record the serial numbers of the devices corrected so that this can be returned to Smiths Medical. For questions regarding this information call Smiths Medical Quality Department at 1-781-763-9335.. |  | 2247 | 2247 units | US Nationwide Distribution | Terminated | February 20, 2014 | 603 | 2247 | Not_Computer
Z-1974-2012 | 62319 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat Portable Clinical Analyzer; Abbott Point of Care, Princeton, NJ 08540 |  |  | September 08, 2011 | 2011 | Abbott Point Of Care Inc. | Some i-Stat Analyzers become uncomfortably hot to touch in the are of the battery. This is a recall extension. | Device Design | Abbott Point Care sent an Urgent Product Recall letter dated September 2011 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to confirm the number of affected product at their facility that were using the Old Battery Carrier. Fill out the affected Buesiness Reply Card and submit their facility information as per the instructions on the card. Upon the receipt of the New Battery Carriers, replace the Old Battery Center with the New Battery Center. Customers were instructed to immediately dispose of all Old Battery Carriers per local regulations. If customers forward the affected product to another facility, they should provide a copy of the recall letter to them and advise them to replace the carriers if required. For any questions call 800-366-8020, Option 1 or your Abbott Point of Care representative.For questions regarding this recall call 609-454-9272. |  | 857258 | 06F16-02-1295 UNITS; 210000-1076 UNITS; 210001-447 UNITS; 210002-5359 UNITS; 210003-9051 UNITS | Worldwide Distribution - USA (nationwide) | Terminated | January 28, 2013 | 508 | 2577875 | Not_Computer
Z-1967-2012 | 52194 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Rigid Suction Wand, sterile. Model numbers S033 and SPC2081. |  |  | June 06, 2012 | 2012 | Edwards Lifesciences, LLC | Edwards is recalling Suction Wand models SPC2081 and S033 because the plastic tip attached to the metal suction wand may become detached during routine use. | Process control | Edwards Lifesciences sent an Urgent Field Safety Notice Product Recall letter dated June 2012 to all affected customers. The letter identified the affected product, description of the issue and actions to be taken. Customers were instructed to examine their inventory for the affected product, quarantine and return any remaining in stock and complete the enclosed acknowledgment form with fax instructions to assure that the notification was received. For question call Edwards Customer Service at (800) 424-3278 from 8:00AM-4:30PM Pacific Time. |  | 1524 | 1524 units | Worldwide Distribution - US (nationwide) including the states of CT, TX, CA, AZ, MO, and MD and the countries of Europe and Japan. | Terminated | September 10, 2012 | 96 | 1524 | Not_Computer
Z-1964-2012 | 61231 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received two complaints related to the Gas System for the Terumo Advanced Perfusion System 1. Both were reported instances of users experiencing difficulty adjusting the System 1 Gas System local flow control knob. | Error in labeling | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL-SAFETY ADVISORY" letter dated June 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 992 | 992 Units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | February 19, 2013 | 242 | 2735 | Not_Computer
Z-1970-2012 | 62373 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 100/120V |  |  | June 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of various roller pump malfunctions of the Terumo Advanced Perfusion System 1 that could result in the inability to use the pump and would then require replacement with back-up equipment. Description of malfunctions:" Unplanned pump stop or pause" Changes in pump speedIf it is necessary to replace a roller pump during cardiopu | Device Design | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 4, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Customer Response Form to the firm via fax at 1-800-292-6551. Contact the firm at 1-800-521-2818 for questions regarding this recall. |  | 992 | 992 units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 01, 2013 | 249 | 1640 | Not_Computer
Z-1972-2012 | 62441 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | enFlow IV Fluid Warmer (510k name: EFLOW IV FLUID WARMER, MODEL |  |  | February 20, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | The EnFlow Warmer Strapdoes not meet the requirements of the biocompatibility standard (ISO 10993) for products that mayhave contact with skin for less than 24 hours. | Device Design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included customer instructions and a Customer Response Form to be completed and returned to the firm. Contact Customer Service at 1-800-932-0760 for questions regarding this recall. |  | 11803 | 11803 cases | Nationwide Distribution. | Terminated | August 19, 2013 | 546 | 11803 | Not_Computer
Z-1960-2012 | 61591 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Advantage CT/PET Fusion for Advantage Windows Workstation with Volume View CT Colonography/Navigator 2 software. (Version 4.0, 4.1,or 4.2) |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to incorrect alignment of fused CT and PET images associated with the CT/PET fusion capabilities of your Volume Viewer and CT/PET Fusion application. Volume Viewer allows two studies (one PET and one CT) that are registered in the same spatial domain to be viewed simultaneously in fused mode. The two images should be aligned to th | Software design | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter provides safety instructions for customers to use until a software upgrade is available. Contact the firm at 800-437-1171 for questions regarding this recall. |  | 70 | 70 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IL, MA, NJ, NY, SC, and TX and the countries of TURKEY, TUNISIA, TAIWAN, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, KENYA, JAPAN, ITALY, IRELAND, INDIA, GREAT BRITAIN, GERMANY, FRANCE, EGYPT, CHILE, CHINA, CANADA, BRAZIL, BELGIUM, AUSTRIA, and ARGENTINA. | Terminated | December 06, 2012 | 196 | 70 | Computer
Z-1969-2012 | 62367 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 220/240V |  |  | June 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received reports of various roller pump occlusion malfunctions that could result in the inability to use the pump and would then require replacement with back-up equipment. Occlusion malfunctions includestiff/difficult to adjust occlusion, loss of occlusion setting, unequal roller-to-roller occlusion and inability to adjust the occlusion.If it is | Equipment maintenance | The firm, Terumo, sent an "URGENT MEDICAL DEVICE RECALL" notification dated June 25, 2012 and an update to the Operator's Manual for the Terumo Advanced Perfusion System 1 via Federal Express to its customers. The notification describes the product, problem and actions to be taken. The customers were instructed to review this Medical Device Recall notice; assure that all users are aware of this notice; place the addendum at the beginning of the Roller Pumps Section (page 5.1) of the Operator's Manual, and complete and return the attached Customer Response Form to the fax number/email address indicated on the form. Note: Terumo is continuing its investigation into the root cause of the malfunctions, and may initiate further action pending results of the investigation.If you have any questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818; customer service hours: Monday-Friday, 8 AM - 6 PM (ET). |  | 648 | 648 units | Worldwide distribution: USA (nationwide) and countries including: Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | September 18, 2015 | 1180 | 1640 | Not_Computer
Z-1961-2012 | 62346 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker brand Arthroscope and Hardware Set Tray (sterilization tray); Model/Part Number: 233-032-116; Product is manufactured and distributed by Stryker Endoscopy, San Jose, CA |  |  | June 19, 2012 | 2012 | Stryker Endoscopy | The gravity steam, Ethylene Oxide (EO) and STERRAD 100S sterilization methods are being removed from the "Instructions for Use"; due to the firm's evaluation that identified original validation was not done at worst case conditions. | Process control | Stryker Endoscopy sent a Urgent Device Correction letter dated June 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to:1. Discard any old Instructions for Use for your Arthroscope and Hardware Set Tray.2. Download the new IFU from the URL link: http://stryker.com/arthrotray .Complete and sign the enclosed acknowledgement of receipt form and fax to(408) 754-8378 or scan in and email to trays@stryker.com.* It is important to send the self addressed confirmation form back to Stryker.Stryker is committed to delivering the highest quality products to its customers. We sincerely apologize for any inconvenience this action may cause.If you have any questions regarding this letter please contact us at 1-800-624-4422 or via email at trays@stryker.com.. |  | 6388 | 6388 units | Worldwide Distribution--USA (nationwide) | Terminated | October 23, 2012 | 126 | 6388 | Not_Computer
Z-1965-2012 | 61480 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base, 100/120V |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Advanced Perfusion System 1 battery charge is being depleted faster than expected during periods of system shutdown. The System 1 may falsely indicate a longer battery life than actual battery time remaining if not subjected to regular use, such as would occur with a backup system. This may result in unexpected system shutdown. Prolonged battery depletion may cause permanent battery damage. | Device Design | TERUMO sent an Urgent Medical Device Recall letter dated June 22, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to review the Medical Device Recall Notice, assure that all users are aware of the notice. Consignees should place the Addendum at the beginning of the System Base Section (page 3-1) of the Operator's manual, confirm receipt of this communication by faxing the attached Customer Response Form to the fax number indicated on the form. All System 1 product is on quality hold. Terumo CVS will release product including batteries as needed only in the event that customers require service. All affected System 1 customers will receive a safety notification and an addendum to the Operator's Manual. There will be a field correction but no product returned for this recall.Customers with questions should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 993 | approx. 993 units | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV.Internationally to: Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam, | Terminated | November 08, 2013 | 504 | 1625 | Not_Computer
Z-1959-2012 | 62406 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | CSI Stealth 360-degree Orbital Atherectomy Device (OAD), Part Number 70056-11, Model Number PRD-SC30-125S. Sterilized with Ethylene Oxide. The Stealth 360-degree Orbital Atherectomy Device (OAD) is part of the Stealth 360-degree Orbital PAD system. |  |  | June 18, 2012 | 2012 | Cardiovascular Systems, Inc. | CSI has initiated a recall on specific lot numbers of Stealth 360-degree Orbital Atherectomy devices which have the potential for an insufficient crown solder bond that may result in the crown becoming detached from the driveshaft. | Employee error | Phone calls to affected sites began June 15, 2012 to quarantine the product. These same sites were also notified with an Urgent Medical Device Recall letter, dated June 18, 2012, via FedEx and e-mail (where available) on June 18, 2012. The letter notified each customer to remove from use and return the recalled product. The customer was instructed to complete a Customer Acknowledgement Form and return it with the device(s) being returned in the Returned Goods Authorization kit provided. Customers with questions can contact customer service at 1-877-274-0901.For any questions regarding this recall call 651-259-2805. |  | 27 | 27 | Nationwide distribution including CA, FL, IL, MI, MN, PA, and TX. | Terminated | July 19, 2012 | 31 | 27 | Not_Computer
Z-1983-2012 | 61151 | No | LGW |  |  | Neurology | PMA | Eon Mini Neurostimulation (IPG) System (Model 3788); |  |  | December 19, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | The firm has received 112 complaints of the Eon Mini IPGs that lost the ability to communicate or recharge due to a workmanship issue resulting in loss of pain relief and subsequent explant. The firm has determined the cause of these reports to be related to process variances in the positioning of the internal battery and printed circuit board, causing a short, and therefore, prematurely depletin | Process design | St. Jude Medical sent an Important Medical Device Recall letter dated December 19, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions:For unimplated inventory of affected product customers should contact their SJM Representative to have the device returned to St. Jude Medical. A replacement device will be provided at no additional cost to them.For implanted affected product it is recommended that customers do not unnecessarily explant the devices associated with this advisory if the IPGs are functioning as intended. If there is a sudden loss of power or if the duration between recharges becomes significantly shorter, customers should contact their St. Jude Medical Representative. If device replacement is required due to IPG failure related to this voluntary recall notice, St. Jude Medical will provide a replacement IPG at no charge.For questions regarding this recall call 972-309-2154. |  | 25255 | 25, 255 devices | Worldwide Distribution - USA (nationwide) and internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Puerto Rico, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | July 24, 2015 | 1313 | 25255 | Not_Computer
Z-1980-2012 | 62325 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Pathfinder NXT Fixed Percutaneous Rod Holder, Part numbers 3573-1 (angle handle) , Stainless Steel / Silicone Rubber, Rx only, non sterile. |  |  | June 05, 2012 | 2012 | Zimmer, Inc. | Zimmer is recalling Pathfinder NXT Fixed Percutaneous Rod Holder, because complaints have been received that the rod holder tip may crack or break when the rod is manipulated during surgery. | Device Design | Zimmer sent an Urgent Medical Device Correction and Removal notification dated June 5, 2012 to distributors and all affected accounts of Zimmer Spine. The notice identified the affected products, problem and actions to be taken. Distributors were asked to hand deliver Recall Notices to Surgeons and return the completed Surgeon Contact Certification Form within 10 working days to ronald.musselman@zimmer.com or FAX to 1-512-258-0995. Notices to Surgeons and User facilities included revised directions/precautions for the Instructions for Use and attached Surgical Techique. Devices will be removed once newly design parts become available.For any related questions or assistance about this please contact Zimmer Spine Customer Service at 866-774-6368. |  | 249 | 249 ( angle handle and straight) in total | Worldwide Distribution - USA (nationwide) and the countries of: Austria, Australia, Belgium, France, Gelmany, Italy, Spain, Switzerland and the United Kingdom. | Terminated | September 03, 2014 | 820 | 249 | Not_Computer
Z-1982-2012 | 61923 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Definition PM Femoral Hip Stem Catalog Numbers 6265-3-099 through 6265-3-107. Howmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430. A subsidiary of Stryker Corp. Stryker France, ZAC Satorlas Green Pusignan, Av De Satolas Green 69881 Meyzeiu, Cedex France. Made in Ireland. |  |  | December 29, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of the potential for cracking of the PMMA coating on their PM COCR Hip Stem product. | Employee error | Stryker Orthopaedics sent an Urgent Product Recall letter dated January 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Stryker Sales Representative to arrange for return of any affected product. Customers were also asked to fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069. For any questions customers should call 201-972-2100.For any questions regarding this recall please call 201-831-5970. |  | 16913 | 16,913 units | Nationwide Distribution | Terminated | March 19, 2015 | 1176 | 16913 | Not_Computer
Z-1984-2012 | 62138 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Pudenz Flushing Valve with Integral Connectors, catalog number NL850-1330. Integra Neurosciences, Plainsboro, NJ. Device is used for the treatment of hydrocephalic patients when shunting cerebrospinal fluid from the ventricles of the brain. |  |  | April 04, 2012 | 2012 | Integra LifeSciences Corp | The pouch label of catalog NL850-1330 indicates an incorrect model and closing pressure of the device. The incorrect label reads: "Model: 12 mm-Medium Pressure" instead of "Model: 16 mm-Low Pressure". | Labeling Change Control | The firm initiated a voluntary recall by notifying the two consignees initially via email and then with an "URGENT: MEDICAL DEVICE RECALL" dated April 5th or 9th, 2012. The letter informed them of the labeling error and provided instructions regarding affected product. The International Customer Service number was provided 609-936-5400. |  | 39 | 39 units | International Only Distribution, including the countries of Thailand and Japan. | Terminated | August 02, 2012 | 120 | 39 | Not_Computer
Z-1989-2012 | 62397 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Axcess Cranial Screw (sometimes referred to as the Axcess Cranial Pin) is a self-tapping, self-drilling screw and is an accessory used in conjunction with the Axcess Cranial Tri-Pin Kit. |  |  | November 08, 2006 | 2006 | GE OEC Medical Systems, Inc | GE Healthcare is aware of a few instances where the Axcess Cranial Screw has broken during insertion. There is a possibility that breaking of the screws could lead to a health hazard. | Device Design | GE Healthcare sent an "URGENT REMOVAL NOTICE" dated November 8, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact GE Healthcare Customer Service at 1-800-874-7378 for questions regarding this recall. |  | 15 | 15 kits | Distributed in Washington, D.C., the states of AL, CA, FL, NY, OR, and VA | Terminated | July 11, 2012 | 2072 | 15 | Not_Computer
Z-1988-2012 | 62107 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal software. |  |  | May 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issuing a software update to the Intellispace Portal systems to address a number of non-conformances. The most frequent and severe non-conformance identified involves the portal setting of "PAC Integration". This setting has a feature that allows the user to launch images of multiple patient studies into an application, but does not display a warning message informing the user that diff | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated May 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact your local Philips representative or the Customer Care Solutions Center at 1-800-722-9377, option 5, for questions regarding this recall. |  | 96 | 96 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, FL, GA, IA, IN, MI, NC, NJ, NY, OK, PA, SC, TN, TX, and WA and the countries of Australia, Austria, Canada, China, Czech Republic, Denmark, and France. | Terminated | March 28, 2013 | 309 | 96 | Computer
Z-1990-2012 | 62332 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Button Switchpens, w/Extendable Needle Electrode 2 mm Exposure Single Use Electrosurgical./Extendable Pencils, 510(k) #K936304, Model Number: 75820. The device is packaged in a sterile Tyvek peel pouch within a cardboard box which contains 24 pouches per box. |  |  | June 05, 2012 | 2012 | PSC Industries Inc | The carton label for this device has the manufacturer symbol where the expiration date symbol should be. Firm officials met with FDA on 5/17/2012. At which time, they reported various labeling errors for their Button Switch Pen Electrode Electrosurgical /Extendable Pencils. | Labeling mix-ups | Olsen Medical sent an "URGENT:MEDICAL DEVICE RECALL" letter dated June 5, 2012. The letter identifies the product, problem, and actions to be taken by the customers. The letter asks that the customer immediately examine their inventory for the presence of the suspect product and quarantine the product, if found. Customers who may have further distributed the recalled product are asked to contact their (sub-account) customers and notify them of the recall. The customers are also asked to call a toll-free telephone number in order to received a Return Material Authorization (RMA) number and so identify if they would like credit or replacement inventory upon returning the product. The customer is also asked to complete and return the enclosed 'Recall Response Form' and to fax the completed form to the fax number on the form or to send the completed from back to the recalling firm via parcel post. |  | 50 | 50 units | Nationwide Distribution-including DC and the states of CA, GA, IN, KY, LA, MT, TX, UT, and WI. | Terminated | January 17, 2014 | 591 | 650 | Not_Computer
Z-1987-2012 | 62005 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal software and Extended Brillance Workspace software |  |  | May 23, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The Intellispace Portal software and Extended Brillance Workspace software may have significantly different results in the colors and size of the affected area when the customer uses the "Time Sensitive" and "Time Insensitive" methods. | Software design | Philips Healthcare sent a Urgent Medical Device Correction letter dated May 31, 2012, to all affected customers. The letter was sent via certified mail to each customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. For any further information or support, customers can contact their local Philips representative or local Philips Healthcare office at (1-800-722-9377 Option 5 : Entersite ID or follow the prompts).Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. The user was informed that a field service representative will visit the facility to install the new software.Philips apologizes for any inconveniences caused by this problem. |  | 38 | 38 | Worldwide Distribution -- Nationwide Distribution including the states of AZ, OR, PA and TX and the countries of Australia, Belgium, Canada, Denmark, Finland, France, Germany, India, Indonesia, Israel, Russia, South Africa, South Korea, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | September 30, 2013 | 495 | 38 | Computer
Z-1986-2012 | 62342 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx M3536A Defibrillator/Monitor with Software R.03.03 |  |  | June 20, 2012 | 2012 | Philips Healthcare Inc. | Defibrillator/monitor cannot analyze an ECG report during acquisition of 12 lead ECG | Nonconforming Material/Component | Philips Healthcare issued an Urgent Voluntary Medical Device Correction letter on June 20, 2012, to all affected customers. The Notification informs customers of the issue, identifies details (including models and software versions applicable) of units affected, instructions on actions to take while awaiting the correction, and identifies what action Philips plans to eliminate or remedy the issueA Philips Healthcare representative will contact customers to arrange for the updated R.03.03 software installation in all the affected devices.For questions customers were advised to call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 35 | 35 units | Worldwide Distribution - USA including FL, IN, MA, M, NY, OH, TN, TX, and VA. Internationally to AUSTRALIA, HONG KONG, and INDIA | Terminated | August 14, 2013 | 420 | 35 | Computer
Z-2000-2012 | 62128 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal software |  |  | May 22, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | In the Multi-Modality Tumor Tracking application, when changing the contours of the lesion using the editing mode "edit contour" tool, the measurements are not recalculated causing discrepancies with the lesion measurements. | Software design | Philips Healthcare sent an Urgent-Medical Device Correction letter dated May 23, 2012, via certified mail to each consignee. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the communication. The letter informed affected customers that a Philips Service Engineer will contact them for implementation of the software update on the affected systems free of charge. For any further information or support, customers can contact their local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 44 | 44 | Worldwide Distribution - US (nationwide) including the states of: AR, LA, MI, NC, NY, OH, OK and PA and the countries of: Australia, Canada, Denmark, France, Germany, Israel, Netherlands, Spain and Switzerland. | Terminated | March 28, 2013 | 310 | 44 | Computer
Z-1997-2012 | 61958 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bayonet Forceps 8.25 (209.6 mm) 0.5mm Tips, Sterile Single Use Surgical/Bipolar Forceps, Model #20-13420, 510 (k) #K884175. The surgical device is package one (1) per Tyvek peel pouch with ten (10) Tyvek peel pouches per box. |  |  | June 01, 2012 | 2012 | PSC Industries Inc | Firm officials reported to CIN-DO that for this device, in the labeling on the inner Tyvek peel pouch, the letter 'E' in Electrosurgical was turned into an 'X' and the number '5" of the (mm tips) was turned into a '0'. This affected the description and some of the text of the labels. However, it did not t affect the Part Number and Lot Number on the label. | Labeling mix-ups | The firm, Olsen Medical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 5, 2012 to their customers via Certified mail along with representative labeling to illustrate the specific labeling issue associated with the various types and/or sizes of Bayonet Forceps surgical instruments that the customers received. The letter described the product, problem and actions to be taken. The letter instructed the customers to immediately examine their inventory for the presence of the suspect product and quarantine the product, if found; to discontinue use immediately and promptly call (800) 251-3000 for a 'Return Material Authorization (RMA) number; to complete and return the enclosed 'Recall Response Form' and fax the completed form to the fax number on the form or to send the completed form back to the recalling firm via parcel post, and for customers who may have further distributed the recalled product are asked to contact their (sub-account) customers and notify them of the recall.If you have any questions call (615) 964-5515 M-F 8:00 am - 5:00 pm CST. |  | 10 | 10 units | Worldwide distribution: USA (nationwide) including states of: CA, CT, FL, GA, MD, MI, NC, OR, SC, and TX and countries of: Canada and Philippines. | Terminated | September 15, 2015 | 1201 | 267 | Not_Computer
Z-2007-2012 | 62434 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco |  |  | June 21, 2012 | 2012 | Elekta, Inc. | Incorrect patient shift directions when the Setup reference dialog is printed out when the DICOM coordinates option is enabled for reports. | Software design | Elekta sent an Important Safety Notice to affected customers. The notice identified the affected product, problem, clinical impact and actions to be taken. The notice inform customers of how the issue will be resolved and instructions to follow when it is available. |  | 38 | 38 | US Nationwide Distribution - including the states of: CA, CT, FL, IL, IN, KY, MI, MN, MO, NJ, NY, NC, OH, OK, OR, PA, WA, and WI | Terminated | April 01, 2014 | 649 | 38 | Not_Computer
Z-1985-2012 | 62112 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrak with Multiple Dataset Navigation, 892.2050 System, Image Processing System, Model Number IT3500 Plus. |  |  | September 15, 2006 | 2006 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. On October 11, 2006, GE Healthcare recalled GE OEC InstaTrak 3500 Plus System with Software version 5.2, the surgical Navigation and Visualization Application due to software related issues. | Software design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC sent an Urgent Safety Advisory Notice dated October 11, 2006 to all affected customers. The notice identified the affected product, problem and actions to be taken. Customers were informed that a GEHC OEC would be contacting their facility to verify the current software version installed at their site. If the issue affects their facility a software upgradewill be provided at no cost. Customers were advised to excercise caution when using the affected product until the 5.2.1software upgrade is completed. For question contact GE Healthcare Customer Service at 1-800-874-7378, option 1 or 1-801-536-4688, option 1.For questions regarding this recall call 801-536-4516. |  | 86 | 86 units | Worldwide Distribution - US (nationwide) | Terminated | August 21, 2012 | 2167 | 86 | Computer
Z-2002-2012 | 62433 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal SIM |  |  | June 21, 2012 | 2012 | Elekta, Inc. | Incorrect patient shift directions when the Setup reference dialog is printed out when the DICOM coordinates option is enabled for reports. | Software design | IMPAC Medical Systems, Inc. sent an Important Safety Notice LINFFM0001/1.0 on June 21, 2012, to consignees on concerning a patch to the software. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the DICOM coordinates option for Reports when printing the Setup Reference Diaglog. Customers were instructed go to www.elekta.com, select the SupportPlus Login at the tope of the page, enter their portal, select Downloads/Updates to download the latest patches.For questions regarding this recall call 770-670-2548. |  | 38 | 38 units | Nationwide Distribution including CA, CT, FL, IL, IN, KY, MI, MN, MO, NJ, NY, NC, OH, OK, OR, PA, WA, and WI. | Terminated | April 01, 2014 | 649 | 38 | Computer
Z-2001-2012 | 62381 | No | OGK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | Enforcement Discretion | superDimension Systems AAS00016-xx with Software Version 4.0-4.9. |  |  | June 18, 2012 | 2012 | superDimension, Inc | superDimension, Inc., is conducting a recall of certain superDimension Systems operating Software Version 4.0 - 4.9. The affected systems are sold under the product code AAS00016-xx.The products being recalled have been found to malfunction when connected to the superDimension Locatable Guide (LG). Specifically, the Systems using Software Version 4.0-4.9 will not recognize the Locatable Guide | Device Design | superDimension sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 15, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter advised customers to discontinue use of the system and requested consignees to complete and return the acknowledgement receipt form via fax to 763-210-4098. Contact the firm at 763-210-4057 for questions regarding this recall. |  | 99 | 99 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of Alabama, Alaska, Arizona, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New Mexico, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Vermont, Virginia, Washington, West Virginia, and Wisconsin and the countries of Canada, Australia, Austria, Belgium, Czech Republic, Denmark, France, Germany, Hungary, Italy, Netherlands, South Africa, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | February 28, 2013 | 255 | 99 | Computer
Z-2012-2012 | 62281 | No | OYG | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Presage ST2 Assay, catalog numbers BC-1065 and BC-1065E. |  |  | February 28, 2012 | 2012 | Critical Diagnostics | The recall was initiated because Critical Diagnostics has confirmed that it is possible to obtain inaccurate test results for specimens when using the Presage ST2 Assay kits. | Component design/selection | Critical Diagnostics sent a Medical Device Recall letter dated March 5, 2012, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.Customers were instructed to please use Federal Express (FedEx) to return the kits, and charge to account: 305599069 Replacement kits will be sent to you.Customers with questions about this notice were instructed to contact Critical Diagnostics' customer support at (877) 700-1250 ext 3. |  | 33 | 33 units | USA including the state of California and the countries of Austria, Spain, United Kingdom, and Italy. | Terminated | July 16, 2012 | 139 | 33 | Not_Computer
Z-2008-2012 | 62457 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom surgical kits and trays: |  |  | June 16, 2012 | 2012 | DeRoyal Industries Inc | Surgical kits and trays contain Stryker Hytrel Togas which were recalled because the clear tape, that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. | Nonconforming Material/Component | DeRoyal initiated the recall on 6/16/2012 by notifying customers by overnight delivery of letters describing the problem, instructing the customer how to handle inventory on hand, and requesting the customer to destroy the product, fill out the provided notice of destruction form, and return the form to DeRoyal.For questions customers were instructed to call 865-362-1020 or 865-362-1034.For questions regarding this recall call 865-362-2334. |  | 303 | 303 units | Nationwide Distribution including SD and NE. | Terminated | November 09, 2012 | 146 | 303 | Not_Computer
Z-2011-2012 | 62024 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bipolar Forceps model number: S02-XXX, S03-XXX, P02-XXX, and P03-XXX. "Instructions For Use" (QF7.5.1A.R01 thru QF7.5.1A.R06 Care and Handling of Bipolar Forceps. |  |  | May 23, 2012 | 2012 | Stingray Surgical Products, LLC | On 05/23/2012, Stingray Surgical Products, LLC., Boca Raton, FL initiated a recall initiated due to "Instructions For Use" of Bipolar Forceps not being cleared by FDA and incorrect "Instructions For Use" . The sterilization instructions have been altered without the clearance of the FDA. | Packaging | Stingray Surgical Product sent an URGENT Instructions for Use Recall letter dated Mary 23, 2012, to all distributors. The letter identified the product, the problem, and the action to be taken by the client. Distributers were requested to convey the recall to their users along with fowarding copies of the firm's recall notification letter. IFU's will be destroyed or discarded by either user or distributor since returns will not be necessary. Distributors were asked to complete the form at the bottom of the letter and return via e-mail or fax. For questions clients were instructed to call 561-210-7582.For questions regarding this recall call 561-210-7582. |  | 6326 | 6326 | Worldwide Distribution - USA (nationwide) | Terminated | August 06, 2014 | 805 | 6326 | Not_Computer
Z-2010-2012 | 61763 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 |  |  | June 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received reports of several issues that manifest themselves as communications or power failures in the system network for Terumo Advanced Perfusion System 1. Issues that affect a single pump or module are:" The appearance of a red X on a single pump or module icon on the CCM " The appearance or a single '?' symbol on a pump or module icon on the CCM" The fai | Nonconforming Material/Component | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 21, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. CUSTOMER INSTRUCTIONS1. Review this Medical Device Recall notice.2. Assure that all users are aware of this notice.3. Confirm receipt of this communication by faxing, or emailing the attached Customer Response Form to the fax number/ email address indicated on the form.Terumo CVS Customer Service-1-800-521-2818Fax-1-800-292-6551.Customer Service Hours:Monday - Friday, 8 AM - 6 PM EST |  | 648 | 648 units | Worldwide Distribution-USA (nationwide) including the states of , AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | June 10, 2013 | 354 | 1640 | Not_Computer
Z-2014-2012 | 61463 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | EtacÂż Ono Walker, EtacÂż Avant Walker, and EtacÂż Salsa Walker. |  |  | April 06, 2012 | 2012 | Etac Supply Center Ab | The recall has been initiated because there is the potential that users may fall and become injured, with the potential for serious injury, including bone fracture. One death occurred in a user fall incident, but it is unclear whether the fall was caused by the loss of a wheel. Among the walkers distributed worldwide, EtacÂż has received a total of 49 reports from Europe and Japan of wheels falling | Device Design | Etac sent an "URGENT DEVICE RECALL " letter dated March 20, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to immediately discontinue selling and distributing the affected product. A Reply Postcard was enclosed for customers to complete and return. Contact the firm at 847-277-7705 for questions regarding this recall. |  | 3187643123 | 204,323 units of the Etac Ono Walker, 15,601 units of the Etac Avant Walker, and 6,771 units of the ETAC Salsa Walker | Worldwide Distribution-USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, China, Colombia, Germany, Denmark, Estonia, Spain, Finland, France, Great Brittain, Greenland, Greece, Hungaria, Ireland, Israel, Iceland, Italy, Japan, South Korea, Kuwait, Lithuania, Luxembourg, Latvia, Malaysia, Netherlands, Norway, Poland, Portugal, Rumania, Russia, Saudi Arabia, Sweden, Slovenia, Slovakia, and Thailand. | Terminated | June 18, 2013 | 438 | 3187643123 | Not_Computer
Z-2013-2012 | 62329 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Molecular Bar Code Scanner User's Guide; instructions for setting up the Bar Code Scanner LN 6L89-01, a component of the Abbott m2000 Instrument System; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list number 06L88-02. |  |  | January 31, 2012 | 2012 | Abbott Molecular | In the Abbott Molecular Bar Code Scanner User's Guide, List Number 6L88-02/09, the section Modifying the Keyboard Format for Bar Code Scanner LN 6L89-01 contains incorrect bar codes that can prevent the correct configuration of the bar code scanner language format.The Abbott m2000 System is intended for use in performing nucleic acid testing in clinical laboratories. It is comprised of the Ab | Error in labeling | Abbott Molecular Inc. sent an Urgent Field Safety Notice letter dated January 31, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to replace the affected pages of the Bar Code Scanner User's Guide with the replacement pages provided with the letter, and immediately perform the Modifying the Keyboard Format section with these replacement pages. The accounts were requested to complete the enclosed reply form and fax it to (847) 775-6728 to acknowledge receipt, understanding and implementation of the information provided in the notice. |  | 100 | 100 units | California, Florida, Maryland, Oregon and Virginia, and internationally to Belgium, Brazil, China, France, Germany, Hong Kong, Japan, Malaysia, Mexico, Qatar, Switzerland and United Kingdom | Terminated | December 04, 2012 | 308 | 100 | Not_Computer
Z-2015-2012 | 62123 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Invasive Blood Pressure Care Cables. |  |  | June 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware through complaints of a potential safety issue with its Invasive Blood Pressure Cable made for use with a Spectramed / BD invasive blood pressure transducer. Due to use of an incorrect O-ring, the mating of the cable with the transducer is faulty. This situation inhibits blood pressure measurement from being monitored correctly. | Nonconforming Material/Component | GE Healthcare sent a Urgent Medical Device Correction letter dated June 12, 2012, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Inforamtion. Advised customers to discontinue use, further distribution and to contact their customers. Please return any defective Invasive Blood Pressure Cables by isolating all affected product and completing the attached Confirmation form/fax back per its instructions. This step is required to confirm receipt of communications with all customers.If you have any questions or concerns regarding this notification, please contact Customer Service at +1-800-932-0760 (Domestic &amp; International). Hours of Operation: 8:00 am EST to 6:00 pm EST. . |  | 279 | 279 | Worldwide Distribution -- USA (nationwide) including the states of AK, CA, IL, MS, NJ, PR, SD, TN, TX, and VA. and the countries of ARGENTIAN, AUSTRALIA, BELGIUM, CHINA, GERMANY, ALGERIA, ECUADOR, EGYPT, FINLAND, UNITED KINGDOM, JORDAN, JAPAN, MOROCCO, MAURITUS, NETHERLANDS, OMAN, PERU, SEBIA, SUDI ARABIA, SINGAPORE, EL SALVADOR, TURKEY, URUGUAY and SOUTH AFRICA. | Terminated | December 05, 2013 | 541 | 279 | Not_Computer
Z-2006-2012 | 62349 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe Ventilator |  |  | June 19, 2012 | 2012 | CareFusion 203, Inc. | CareFusion identified that the EnVe Ventilator may not hold the set Positive End Expiratory Pressure (PEEP) value either intermittently or continuously or it may alarm for a Patient Circuit Fault. The unit will activate either a Low PEEP or Patient Circuit Fault Alarm both audibly and Visually to alert the healthcare professional. Ventilation delivery to the patient may be compromised. | Device Design | CareFusion sent an Urgent Product Recall letter with a return receipt dated June 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Required Action for Users: If the ventilator resumes normal ventilation (intermittent leak) the audible alarm stops. To clear the alarm indicator on the ventilator display, enter Alarm Messages tab, push alarm Reset to clear display. If the ventilator does not resume normal ventilation (continuous leak) provde an alternative method of ventilation to the patient. CareFusion will follow-up with a phone call to coordinate the correction of the devices. For questions customers should call 1-800-554-8933. |  | 266 | 266 | Nationwide Distribution including AZ, CA, CO, FL, IL, KY, ME, MI, MO, MN, MT, ND, NJ, NM, NV, NY, OK, PA, SC, SD, TX , UT, WI, and WV. | Terminated | February 11, 2013 | 237 | 266 | Not_Computer
Z-2016-2012 | 62051 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Arthrodesis Nailing System. Manufacturer: Stryker Trauma GmbH, Professor-Kuntscher-Str. 1-5, 24232 Schonkirchen Germany. Distributed in the USA by Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | January 26, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has become aware that there exists a potential for breach of the sterile barrier of packaging associated with certain lots of T2 Knee Arthrodesis Nails. | Packaging | Stryker Orthopaedics sent Recall letters/Customer Response forms dated 1/26/2012 and 1/30/2012 on 1/27/2012 and 1/31/2012 via FEDEX to their branches, hospitals and surgeon accounts respectively. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were instructed to contact their Stryker representative to arrange for return of any affected product. Customers were asked to fax back the attached Product Recall Acknowledgment Form as soon as possible to 201-831-6069..For any questions customers should call 201-972-2100.For questions regarding this recall 201-831-5970. |  | 532 | 532 units. | Nationwide. | Terminated | July 17, 2013 | 538 | 532 | Not_Computer
Z-2019-2012 | 62437 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Xprezzon Bedside Monitor, Model 91393, installations with Option C. It is a patient monitor with arrhythmia detection or alarms. |  |  | June 12, 2012 | 2012 | Spacelabs Healthcare, Llc | There were reports that a certain combination of key presses will cause the module to disconnect from the Xprezzon bedside monitor, model 91393, and the patient parameters will drop from the display. The parameters return after power cycling the monitor. | Software design | Spacelabs Healthcare sent an Urgent Medical Device Correction letter dated June 29 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to advise their staff of the potential failure of the XPREZZON Bedside Monitor. Spacelabs will contact customers to schedule a convenient time for their hospital to upgrade their affected monitors, at no cost. On July 6 2012, a customer letter (translated as necessary) URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare XPREZZON Bedside Monitor will be emailed to all international subsidiaries and distributors of record. Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support.For questions regarding this recall call 800-522-7025, ext 2. |  | 21 | 21 units | Worldwide Distribution - USA including Florida, Iowa, Oklahoma, South Caroline and the countries of India, Germany and Panama. | Terminated | November 28, 2012 | 169 | 21 | Not_Computer
Z-2017-2012 | 62451 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroCool Breathable Impervious Surgical Gown, (Large and X-Large), Sterile, Made in Honduras. |  |  | September 17, 2009 | 2009 | Kimberly-Clark Corporation | Complaints were received of gowns pilling or balling, linting and abrades, blood strikethrough and fibers falling off the gowns. | Process control | Kimberly Clark sent an Urgent Medical Device Recall letter dated September 17, 2009. to all affected consignees. The Recall notification of the expanded recall was sent to consignees on November 17, 2009. Customers were instructed to destroy all affected product remaining in their inventory. For questions regarding this recall call 770-587-7400. |  | 79941 | 79,941 units | Nationwide Distribution including AL, FL, GA, KY, LA, MS, NC, OH, SC, TN, and VA | Terminated | July 19, 2012 | 1036 | 79941 | Not_Computer
Z-1992-2012 | 62115 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Infusomat(R) Space Pump (US Version) with software version 686G030103 (commonly termed 'G03'), 686E030003, 686F030005, 686F030007, 686G030002, and 686G030102. |  |  | March 23, 2012 | 2012 | B. Braun Medical, Inc. | Firm became aware of the potential for breakage of the anti free flow clip catch, located on the inside of the pump door, when the IV set anti free flow clip is incorrectly inserted into the pump and the pump door forced closed. If the clip catch is broken and the door opened, free flow protection is still ensured. However, after the set is manually removed from the pump without the closure of t | Device Design | B. Braun Medical, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated March 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to identify the affected product in their inventory and to complete and return the Product Acknowledgement Form via fax to (610)849-1197. Contact the Customer Support Department at (800) 626-7867 for assistance with returning any product or if you have questions regarding this notice. |  | 38225 | 38,225 pumps | Nationwide Distribution-including Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Terminated | November 03, 2014 | 955 | 38225 | Not_Computer
Z-2022-2012 | 62337 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | LMA Esophageal Intubation Detector (EID), Catalog Number EID100 |  |  | May 23, 2012 | 2012 | LMA North America Inc | The recall was initiated because LMANA has received a complaint that the Esophageal Intubation Detector (EID100) has a potential to create a false positive if used on a patient during confirmation of intubation. | Process control | LMA North America Inc. sent a recall letter dated May 23, 2012, for US customers, and May 24, 2012, for International customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to fill out the bottom portion of the recall letter and promptly fax it back to LMA North America at (858) 622-4130. If customers have units of EID100 from the affected lots remaining in stock, customers were instructed to arrange for immediate return by contacting LMANA's customer service department at (800) 788-7999, menu selection #2. For further questions please contact Customer Service Representative at (800) 788-7999. |  | 18799 | 18,799 units | Worldwide Distribution--USA (nationwide) and the countries of Canada, United Kingdom, ZA, and Australia. | Terminated | December 26, 2012 | 217 | 18799 | Not_Computer
Z-2023-2012 | 62020 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EkoSonic MACH4 Endovascular Device (Catalog #500-55106, 500-55112; 500-55118; 500-55124; 500-55130; 500-55140; 500-55150; 500-54106; 500-56130; 500-56140; 500-56150) is a catheter, continuous infusion. it consists of the Intelligent Drug Delivery Catheter (IDDC) and the MicroSonic Device (MSD). |  |  | May 24, 2012 | 2012 | EKOS Corporation | EKOS Corporation have recently received several complaints from users where the radio-opaque marker bands on the Intelligent Drug Delivery Catheter (IDDC) were dislodged during a procedure and report of MicroSonic Device (MSD) breaking during insertion or operation. | Device Design | Starting May 25, 2012 to May 31, 2012, EKOS Corporation sent "New warnings in EkoSonic Endovascular Device Instruction for Use" letter, dated May 24, 2012, to all affected customers.Attached with the notification letter is the revised EKOS MACH4 Endovascular Device Instructions for Use in Peripheral Vasculature. Customers are informed of the complaints about the radio-opaque marker bands on the Intelligent Drug Delivery Catheter (IDDC) were dislodged during a procedure and reports of the MicroSonic Device (MSD) breaking during insertion or operation.Customers with questions about the revised Instruction for Use, please contact EKOS Corporation at 425-415-3132. |  | 21892 | 21,892 units | Worldwide Distribution - USA including OH, KY, and TN and the countries of Switzerland, Germany, Austria, Czech Republic, Slovakia Republic, Poland, Netherlands, Belgium, Luxembourg, United Kingdom, Republic of Ireland, Slovenia, Costa Rica, Argentina, Brazil, Dominican Republic, Taiwan, Korea, Greece, Mexico, Italy, Uruguay, Paraquay, Yugoslavia, New Zealand, India, Venezuela, South America, Norway, Sweden, Denmark, Turkey, Great Britain, Israel, Spain, Portugal, Greece, Korea, China, and Japan. | Terminated | November 20, 2012 | 180 | 21892 | Not_Computer
Z-2021-2012 | 61753 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base, 220/240V. |  |  | June 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | TerumoÂż Advanced Perfusion System 1 has displayed a Red X error message on multiple module icons indicating a loss or fluctuation of internal power on one side of the system's base. Terumo Cardiovascular Systems' initial investigation has determined that there may be more than one cause for the malfunction and is investigating further to determine the root cause.Potential risks with the loss or | Process control | Terumo sent an "URGENT MEDICAL DEVICE RECALL" letter dated June 21, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to place the addendum in the Operator's Manual for Terumo System 1, at the beginning of the Central Control Module Section. Additionally, a Customer Response Form was attached for customers to complete and return to the firm. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this notice. |  | 648 | 648 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV. and the countries of ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 26, 2013 | 278 | 1640 | Not_Computer
Z-2031-2012 | 61769 | stains, microbiologic | JTS | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | BD Gram Stain Kit, containing 1- 250 mL bottle of Gram Crystal Violet, and additional reagents: Gram Iodine (Stabilized), Gram Decolorizer, Gram Safranin. |  |  | April 12, 2012 | 2012 | Becton Dickinson &amp; Co. | Potential for increase level of artifacts to appear as large, spherical and easy to distinguish or small &amp; more difficult to distinguish from gram positive cocci. | Material/Component Contamination | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter dated April 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use and discard the affected product. An Acknowledgement Form was attached for customers to complete and return via fax to 410-316-4258. Contact the BD Customer Service Department for questions regarding this recall. |  | 539 | 539 kits | Nationwide Distribution. | Terminated | December 10, 2012 | 242 | 3084075 | Not_Computer
Z-2030-2012 | 61691 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMASpectrum Infusion Pump Master Drug Library (MDL) Editor software. |  |  | April 23, 2012 | 2012 | Sigma | Sigma is a safety alert for the Master Drug Library (MDL) Editor software associated with the SIGMA Spectrum Infusion Pump System because it may result in a facility drug library that does not represent common clinical practices. | Software design | Sigma sent an Urgent: Customer Safety Alert dated April 23, 2012 via UPS delivery to all consignees requesting both acknowledgement of notification and confirmation that recommended actions have been taken. Customers are asked to follow the instructions contained in the recall letter. Customers with questions may contact the firm at 800-356-3454 ext 300. |  | 2048 | 2,048 CD's | Worldwide Distribution - US (nationwide) including Puerto Rico, and Canada. | Open, Classified |  |  | 2048 | Computer
Z-2028-2012 | 62464 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Integra(TM) Jarit(R) Take-Apart Endoscopic Instruments. |  |  | May 23, 2012 | 2012 | Integra Limited | The Integra(TM) Jarit(R) Take-Apart Laparoscopic Instrument shafts were released with internal component out of specification. | Nonconforming Material/Component | Integra sent a Urgent Field Corrective Action letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter instructed customers and sales reps to return the instruments identified on the attached part number list and complete a "Field Corrective Action Acknowledgement and Return Form" to be completed and returned to the firm.We apologize for any inconvenience this issue may cause you. Please call Integra Surgical Customer Service at 800-431-1123 if you have any questions. |  | 108 | 108 | USA (nationwide) including the states of AZ, CA, IL, MA, MN, NJ, OK, TN, TX, VA, WA and WI. | Terminated | March 19, 2013 | 300 | 108 | Not_Computer
Z-2029-2012 | 62408 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink. |  |  | April 09, 2012 | 2012 | Gulf Fiberoptics, Inc. | The GMF-RC65 Replacement Surgical Headlight Cable for Cogent Microlink falls outside of the firm's 510k due to light transmitting fiber material difference. | Device Design | The firm, Gulf Fiberoptics, Inc., sent an "RECALL NOTIFICATION-PLEASE DO NOT DISCARD" letter dated April 09, 2012 to its customer via (traceable letter &amp; email). A notification was put on the firm's website. The letter described the product, problem and actions to be taken. The distributors were instructed to determine the status of the product; remove existing product from service immediately and return to their company in a timely manner; send all returned GMF-RC65 cables to Gulf Fiberoptics; forward copies of the recall letters to their customers, ultimately the end user, and complete and return the attached Recall Worksheet Form for each GMF-RC65 cables regardless of its status via fax to: +1 813.855.6627; email to: difigueoa@gulffiberoptics.com or mail to : Gulf Fiberoptics, 448 Commerce Blvd, Oldsmar, FL 34677. The customers were instructed to remove the cables from service immediately, if they think they may have one of the listed cables; contact customer service at +1 813.855.6618 for serial number verification and for instructions to return this cable to Gulf Fiberoptics.If you have any questions regarding this voluntary recall, please contact your Customer Service representative at: +1813.855.6618 or by email at: difigueroa@gulffiberotpics.com. |  | 169 | 169 units | Worldwide distribution: USA (nationwide) including states of: AL, CA, CO, FL, MI, NC, NJ, NY, OH, PA, TN, TX and WI and country of: South Africa. | Terminated | March 07, 2013 | 332 | 169 | Not_Computer
Z-2027-2012 | 62404 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 8800 Flexiview designed to provide fluoroscopic and spot-film imaging of the patient during diagnostic, surgical and interventional procedures. |  |  | November 19, 2007 | 2007 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC recalled imaging devices OEC 9900, OEC 9800, and OEC 8800 as a result of an FDA inspection identifying that the vertical lift column power supply in the mainframe C-arm is defective and subject to early life failure. | Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE Healthcare sent an"URGENT RECALL NOTICE" dated February 22, 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 800-874-7378 for questions concerning this recall. |  | 1349 | 1349 units | Worldwide Distribution. | Terminated | January 28, 2013 | 1897 | 13175 | Not_Computer
Z-2034-2012 | 61147 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | epoc Blood Analysis BGEM Test Cards for use with the epoc(R) Blood Analysis System |  |  | June 28, 2011 | 2011 | Epocal Inc | Customer reported problem while performing incoming QC of newly received lot of test cards. Liquid quality control results were found to be out of range. Further investigation by Epocal determined that a portion of the test card lot (approximately 3000 cards) could potentially report low glucose results. | Device Design | Epocal Inc. sent a Voluntary Product Recall letter dated June 27, 2011 to the affected consignee. The letter identified the affected products, problem and actions to be taken. The letter ask that all products affected by this recall be sent back to Epocal and a replacement may be obtained. For questions call 1-613-738-6912. |  | 8650 | 8,650 | US nationwide Distribution in the state of CA. | Terminated | December 18, 2013 | 904 | 8650 | Not_Computer
Z-2052-2012 | 61479 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Sarns Modular Perfusion System 9000, 100V with color screen. |  |  | June 18, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received several reports of malfunctions of the Air Bubble Detection System used with its heart-lung machines over the last three years. The malfunctions were expressed in various failure modes (over sensitivity, false alarms and were found to have various root causes. The malfunctions occurred in systems that were excluded from prior corrective actions or in syst | Nonconforming Material/Component | Terumo CVS sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 30, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers are asked to complete and return a Customer Response Form via fax to 1-800-292-6551. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions regarding this recall. |  | 31 | 31 units | Worldwide Distribution-USA (nationwide) and the countries of ARGENTINA, BANGLADESH, BELGIUM, Bolivia, BRAZIL, BRISBANE, CANADA, CHILE, China, COLOMBIA, COLOMBIA, Costa Rica, DENMARK, Dominican Republic, Ecuador, Egypt, FRANCE, GERMANY, Guatemala, HOLLAND, Honduras, Hong Kong, India, INDIA, Indonesia, Israel, ITALY, Jordan, KOREA, Kuwait, Lebanon, LEICESTERSHIRE, Malaysia, Mexico, NETHERLANDS, New Zealand, Nicaragua, NORWAY, Pakistan, Panama, PERU, Philippines, Russia, Saudi Arabia, SINGAPORE , SOUTH AFRICA, South Korea, SPAIN, Sri Lanka, SWEDEN, SWITZERLAND, Taiwan, THAILAND, Turkey, UNITED ARAB EMIRATES (UAE), Uruguay, Venezuela, Vietnam, WEST GERMANY and Yemen. | Terminated | April 12, 2013 | 298 | 4689 | Not_Computer
Z-2039-2012 | 62293 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Alere Triage Tox+MTD Drug Screen, PN 94400 |  |  | May 14, 2012 | 2012 | Alere San Diego, Inc. | A recall was initiated because Alere San Diego has confirmed that the Triage TOX Drug Screen may have significantly decreased precision and accuracy than claimed. An increased frequency of false positive and/or false negative results has been observed for Triage TOX Drug Screen. | Nonconforming Material/Component | Alere sent an Urgent Medical Device Recall letter dated June 11, 2012 to all affected customers. The letter informed the customers of the affected product, problem identified and actions to be taken. Customers were instructed to discontinue all use of the recalled lots, immediately inform clinicians of this recall at the testing site(s) and complete and fax the enclosed customer verification form and list of affected lots within 10 days to confirm receipt of the letter. Customers with questions about the information contained in the recall letter, can contact Alere San Diego at 9975 Summers Ridge Road, San Diego, CA 92121 or call (877) 308-8287. |  | 9997 | 9,997 kits | US Nationwide Distribution | Terminated | December 04, 2012 | 204 | 9997 | Not_Computer
Z-2036-2012 | 62284 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI Concha Therm Neptune Humidifier, Rx Only, |  |  | June 12, 2012 | 2012 | Teleflex Medical | Speaker Failure. If the speaker fails and an event occurs which would generate an audio alarm, the alarm will not sound and could cause a potential delay in patient treatment. | Device Design | Teleflex sent Urgent Medical Device Notification letters dated June 12, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to complete the enclosed Acknowledgement Form and return per the instructions provided. Once the form is received, Teleflex Medial will schedule for a technician to come to their facility and exchange the affected product. For questions call your local sales representative or Customer Service at 1-866-246-6990. |  | 14794 | 14,794 each | Worldwide Distribution - U.S. (nationwide) and the countries of Argentina, Australia, Brazil, Canada, Chile, China, Europe, India, Japan, Mexico, South Korea, Taiwan and UK. | Terminated | January 28, 2014 | 595 | 14794 | Not_Computer
Z-2041-2012 | 62420 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian Eclipse Treatment Planning System for radiotherapy; Model Number: H48; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 18, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning (for radiotherapy treatments) where changing the prescription in a calculated VMAT or RapidArc plan may lead to Monitor Units which do not reflect the new prescription. If the dose prescription of a VMAT or RapidArc plan is changed after the initial Dose Volume Optimization, the Monitor Units and dose distribution may not reflect | Software design (manufacturing process) | The firm, VARIAN Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION; URGENT FIELD SAFETY NOTICE" dated June 18, 2012, to all affected users. The notice described the product, problem and actions to be taken. Varian is developing a technical correction for this issue. The customers will be contacted by a Varian Customer Service representative when this correction is available to schedule its installation.If you require further clarification, please feel free to contact your local Varian Medical Systems Customer Support District or Regional Manager.In USA and Canada: 1.888.VARIAN5 (888.827.4265)Europe: +41 41 749 8844; email: support-emea@varian.comEmail: North America: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: support-sea@varian.com; China/Asia: support-china@varian.com; Japan: support-japan@varian.com and Latin America: soporte.al@varian.com. Internet: Oncology systems customer site - www.myvarian.com and Varian Medical Systems public site - www.varian.com. |  | 7113 | 7113 units | Worldwide distribution: USA (nationwide) and countries including of: Austria, Australia, Bahrain, Belgium, Brazil, Canada, China, Chile, Colombia, Denmark, Egypt, Finland, France,Germany, Greece, Hungary, Hong Kong, India, Indonesia, Israel, Ireland, Italy, Japan, Latvia, Malaysia, Mexico, Morocco, Nepal, Netherlands, Norway, Panama, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, United Kingdom, Taiwan, Turkey, Ukraine, United Arab Emirates and Venezuela. | Terminated | August 06, 2014 | 779 | 7113 | Not_Computer
Z-2035-2012 | 62498 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Brilliance 6, 16, 40, 64, Big Bore, iCT and iCT SP are Computed Tomography X-Ray Systems. |  |  | July 02, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare was notified that several of the sixteen bolts securing the brackets which support the carbon table top on various models of their Brilliant CT scanner have been found to be defective and the bolt heads may break off. | Process control | Philips sent a Urgent Medical Device Correction letter dated June 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will visit each site to inspect and correct the issue, and Philips will be tracking the progressof the correction. For further information, Customers were asked to contact their local Philips representative or Philips Healthcare Office. North American and Canadian consignees should contact the Philips Customer Care Solutions Center at 1-800-722-9377 |  | 524 | 524 | Worldwide Distribution -- USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, MJ, NM, NV, NY, OH, OK, OR, PA, PR, TN, TX, VA, VT, and WI. and the countries of Canada, Argentina, Australia, Austria, Azerbaijan, Brazil, Chile, China, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, France, Germany, Greece, Iceland, India, Indonesia, Iraq, Italy, Jordan, Kazakhstan, Kazakhstan, Libyan Arab Jamahiriya, Malaysia, Mexico, Mozambique, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Palestine, Peru, Philipines, Poland, Portugal, Russia, Russia Fed, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Syrian Arab Republic, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | December 24, 2013 | 540 | 524 | Not_Computer
Z-2033-2012 | 61723 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker AVS Navigator Trials 6mm-15mm |  |  | January 13, 2012 | 2012 | Stryker Spine | Beginning in February 2011, Stryker Spine began receiving reports related to the shaft of the AVS Navigator trial failing at the junction of the trial head. | Device Design | Stryker sent an "URGENT PRODUCT RECALL" letter to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Customer Response Form was included for customers to complete and return via fax to Regulatory Compliance at 201-760-8370. Contact the firm at 201-760-8298 for questions regarding this recall. |  | 273 | 273 units | Nationwide. | Terminated | July 29, 2013 | 563 | 273 | Not_Computer
Z-2037-2012 | 62372 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi Scenaria CT System Software |  |  | May 04, 2012 | 2012 | Hitachi Medical Systems America Inc | Hitachi discovered a software error in the Reconstruction Status function that creates a potential risk that a patient study will not be fully reconstructed. When multiple reconstructions are queued for processing, if a queue is suspended and the task order changed, the second queue may not reconstruct the full range of slices. This may result in missing diagnostic data. | Software design | Hitachi sent a "Device Correction Notice" dated May 4, 2012 to their customers. The notice identified the affected product, problem and actions to be taken. Customers were instructed to "not use the Reconstruction Status" function to suspend active reconstructions until the software problem can be corrected. They also informed their customers that a representative will visit their facility to install the software update to permanently resolve the issue at no cost to the facility. For further information regarding this notice call 800-800-3106 x3720. |  | 6 | Six (6) units | US Nationwide Distribution to the following states of: MD, MI, MX, ND, NY, OH and TX. | Terminated | February 19, 2013 | 291 | 6 | Computer
Z-1993-2012 | 62527 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima Mobile X-ray System. The GE Automatic Mobile X-Ray |  |  | April 04, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the mobile radiographic product, Optima XR200amx, ond Optima XR220amx related to Visual indication as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-1-3. | Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication tothe user or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the firmware on the system to address the issue.Your proposed corrective action plan (CAP) (Field Modification Instructions 10865) describes the rework plan that GEHC Engineers will make to the imaging system to return it to compliance with the applicable performance standard. It appears to adequately address the problem and is hereby approved. |  | 254 | 254 units have been shipped to/installed in the U.S. | Worldwide Distribution (USA) nationwide | Terminated | January 10, 2013 | 281 | 254 | Computer
Z-2055-2012 | 61242 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 Base 100/120V |  |  | June 22, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received eight complaints since September 2008 of drive motor failure of the Sarns Centrifugal System of the Terumo Advanced Perfusion System 1. A failure of the drive motor for the Terumo System 1 centrifugal system triggers error messages and alarms; therefore it is easily recognizable by the user.The potential hazards are dependent on the application | Equipment maintenance | TERUMO sent an Urgent Medical Device Recall letter dated June 22, 2012, and an update to the Operator's Manual via Federal Express to inform consignees of the reported malfunction and potential hazards and to raise awareness of the benefits of using a retrograde valve and the importance of backup system availability. Consignees were instructed to review the Medical Device Correction notice, assure that all users are aware of the notice. Place the Addendum at the beginning of the Centrifugal System Section of the Operator's Manual, Confirm receipt of the communication by faxing or emailing a completed Customer Response Form as indicated on the form. For any questions customers should call 1-800-521-2818. |  | 992 | 992 devices | Worldwide Distribution - USA (nationwide) and the countries of Nationwide and worldwide: ARGENTINA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatamala, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | March 06, 2013 | 257 | 1632 | Not_Computer
Z-2057-2012 | 62351 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics PediLoc Locking Plate System, 4.5mm Contour Locking Compression Femur 8 Hole Plate Right. |  |  | June 14, 2012 | 2012 | OrthoPediatrics Corp | The device is a Left 8 hole plate but is incorrectly labeled and etched as a Right device. | Error in labeling | On 15-June-2012, a follow-up, URGENT Recall of MIS-BRANDED PediLoc Locking Plate e-mail was sent to all Distributors and direct representatives to provide them with a written notification of the voluntary recall. The e-mail identified the affected product and the reason for the recall. It also instructed them to quarantine all identified devices from the affected lot number immediately. A certified Voluntary Recall notice was also sent as a follow up along with a return prepaid pouch for the return of the recalled device(s) and Mandatory Reply Form. |  | 15 | 15 devices | Worldwide Distribution -- USA, including states of IN, GA, MI, FL, MO, OH, and SC and country of Australia. | Terminated | April 30, 2013 | 320 | 15 | Not_Computer
Z-2054-2012 | 62324 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Universal Spinal System Lamina Hooks |  |  | May 31, 2012 | 2012 | Synthes (USA) Products LLC | USS Lamina Hooks, Medium, Right are being recalled because the hooks are etched with the wrong part number - 498.321 (Lamina Hook, Medium, Left). The Lamina Hooks should be etched with part number 498.320 (Lamina Hook, Medium, right). | Employee error | SYNTHES sent a Medical Device Recall letter dated May 31, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees and sales reps remove recalled product from their inventory, call the firm at 1-800-479-6329 to receive an authorization number and complete the enclosed verification form and return to:Credit/Returns, Synthes1101 Synthes AvenueMonument, Colorado 80132.If customers do not have any of the affected product they were instructed to complete the Verification Section at the end of the letter by checking the appropirate box indicating that no affected porudct has been located. Return the document to Synthes by fax to 610-251-9005 or Scan/Email: FieldAction@synthes.com.For any questions please call 610-719-5450. |  | 12 | 12 units | Nationwide Distribution including GA, PA, SC, TX, and UT and Canada. | Terminated | August 31, 2015 | 1187 | 12 | Not_Computer
Z-2053-2012 | 62114 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica M525 F50 Surgical Microscope; Article No. 10338422; Leica Microsystems (Schweiz) Ag, 201 Max Schmidheiny Strasse, Heerbrugg, Switzerland |  |  | June 04, 2012 | 2012 | Leica Microsystems, Inc. | Leica Microsystems evaluated data showing that a malfunction resulting in an illumination reduction can be observed when Leica M525 F50 surgical microscope is used in combination with a remote control connected to a remote video adapter or motorized dual video adapter. | Software design | Leica Microsystems sent an Urgent Medical Device letter dated June 4, 2012 to the affected customers. The letter described the product, problem and actions to be taken. The customers were informed that the Leica surgical microscope M525 F50 required a software upgrade. Customers were instructed to NOT USE a remote video adapter or motorized video adapter until the required software upgrade has been installed by their Leica Microsystems representative. The customers were requested to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the serial numbers of the units at their facilities and to fax it back to Leica Microsystems at (847) 236-3747. |  | 43 | 39 Units - USA , 4 Units - Canada | Worldwide Distribution - US (nationwide) in the states of: California, Colorado, Florida, Georgia, Illinois, Indiana, Massachusetts, Michigan, Minnesota, New York, Oklahoma, Oregon, Pennsylvania, Texas, Utah, West Virginia, including Puerto Rico, and the countriy of Canada. | Terminated | October 29, 2012 | 147 | 43 | Computer
Z-2058-2012 | 62520 | No | KMJ | Patient lubricant. | General Hospital | General Hospital | 510(K) Exempt | Oto Ease by Westone, labeled as a bacteria free, greaseless lubricant. Sold in 0.5 oz consumer sized semi transparent flexible plastic bottles. |  |  | July 13, 2012 | 2012 | Westone Laboratories, Inc. | Westone is recalling all Oto Ease ear lubricant in all packaging configurations and sizes due to potential microbial contamination of the product. | Material/Component Contamination | Westone Laboratories sent an Urgent Recall letter and Press Release to distributors and customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory, quarantine product subject to recall and discard or return the product to Westone. In addition, if the affected product have been further distributed please identify your customers and notify them at once of this product recall. Consumers with questions may contact Westone Laboratories at 1-800-357-3240 between the hours of 8:00am and 5:00pm MST or email at otoeaserecall@westone.com. |  | 115 | Approx 100k sold annually x 15 years. | Worlwide Distribution - US (nationwide) and to countries of: Australia, Belgium, Brazil, Canada, France, Germany, Greece, Hong Kong, Ireland, Japan, Korea, Mexico, Panama, Philippines, Russian Federation, and Singapore, Sweden, Switzerland and Thailand and United Kingdom. | Terminated | December 14, 2012 | 154 | 115 | Not_Computer
Z-2068-2012 | 61910 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5600 Integrated System Software |  |  | May 21, 2012 | 2012 | Ortho-Clinical Diagnostics | software automatically switches the lot when a vancomycin (VANC) or valproic acid (VALP) reagent is depleted or the operator manually changes the lot in use causing lack of verification of proper quality assessment. | Software design | Ortho Clinical Diagnostics notified its foreign direct account consignees via email on May 21, 2012 and their direct account US consignees were sent an Urgent Distributor Safety Notice on May 22, 2012 via US Postal Service Priority Letter. The recall letter notified the consignees of the situation and that Ortho will replace their remaining inventory of non-expired GENs 14 or 15 for VITROS VALP Reagent only. Until then the consignees can continue using the VITROS VALP Reagent, GENs 14 or 15 until the replacement arrives. And to note that VITROS VANC Reagent, GENs 19 &amp;20 are now expired. --- The letter also requested the consignees to complete and return the enclosed Confirmation of Receipt, no later than May 31, 2012. And to review the quality control beginning with the Calibration date of the VITROS VANC Reagent ( GEN 21 or above) or VITROS VALP (GEN &amp; above) so that the QC results can be verified to be within acceptable ranges. --- For customers using VITROS VALP Reagent GEN 14 or 15, they were to respond as to the quantity of the non-expired GENs 14 and 15 remaining inventory; to load and use only one GEN of VITROS VALP Reagent at a time on the system until the replacement arrives. This notification letter should be posted by each VITROS System using the VITROS VALP Reagent or VITROS VANC Reagent or with the user documentation. And in addition, if these products were distributed then this information should be forwarded as well.For questions regarding this recall call 585-453-4110. |  | 1387 | 1387 | Worldwide Distribution - USA (nationwide) and the countries of Australia, New Zealand, Brazil, Canada, Chile, Argentina, Paraquay, Ecuador, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, UK, Northern Europe, France, Germany, Italy, Spain and Portugal. | Open, Classified |  |  | 3508 | Computer
Z-2059-2012 | 62307 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | 4.5mm TI Multiloc Screw Length 38MM-Sterile |  |  | May 22, 2012 | 2012 | Synthes USA HQ, Inc. | Two lots of 4.5mm multi-loc screws were recalled due to incorrect labeling of screw size. | Error in labeling | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated May 22, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Verification Form was enclosed for customers to complete and return via fax to 610-251-9005. Contact the firm at 610-719-5450 for questions regarding this recall. |  | 20 | 20 | Nationwide Distribution-including the states of AZ, CA, FL, GA, IL, LA, MD, MN, NJ, NY, OH, PA, TN, TX, and VA. | Terminated | September 10, 2015 | 1206 | 20 | Not_Computer
Z-2024-2012 | 62393 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics Trilogy 100, 200, and 202 Ventilators |  |  | April 27, 2012 | 2012 | Respironics, Inc. | The specified Trilogy units are being recalled because of the potential for a faulty component in the power supply. | Component design/selection | Respironics, Inc. began contacting its affected US customers by phone beginning on Friday, April 27, 2012, and requested that they quarantine any affected devices within their possession and that they retrieve any devices with patients and quarantine those as well. Replacement devices were to be provided to all affected customers. For questions customers should call 724-387-7651.For questions regarding this recall call 877-387-3311.. |  | 173 | 173 | Worldwide Distribution - USA including AZ, CA, IL, IN, LA, MI, MN, MO, NJ, NY, NC, PA, SC, TN, TX, and WI and the countries of Canada, Denmark, Estonia, Finland, France, Great Britian, India, Italy, Japan, Korea, Saudi Arabia, and Spain. | Terminated | November 01, 2012 | 188 | 173 | Not_Computer
Z-2073-2012 | 62519 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE or ONCOR Digital Linear Accelerators with syngo RT Therapist version 4.2/MOSAIQ OIS Systems and Multiple-X, Manufactured by SIEMENS AG. The ARTISTE and ONCOR family of Linear Accelerators are used to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | June 25, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Potential safety issue when using the IMMEDIATE RESUMPTION feature after a Control Console interlock during the delivery of flattening-filter-free (Multiple-X) beams or segments. In a very rare scenario, the RT Therapist will send again the previously completed beam or segment for delivery to the control console, thus repeating the already delivered Monitor Units of the same beam or segment. | Software design | The firm issued "Urgent: Medical Device Correction" letters delivered by the Siemens Customer Service Group or by certified mail to all its direct consignees, informing them of the affected product and providing instructions on the recall. The firm stated that release of this UI TH018/12/S to distribute the notifications began on June 25, 2012. |  | 18 | 18 | Worldwide Distribution, including the state of Florida (USA) and the countries of Germany, India, Poland, Thailand, and the Republic of Korea. | Terminated | October 18, 2012 | 115 | 18 | Not_Computer
Z-2070-2012 | 61969 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Product is labeled in part: "***OPTETRAK***TIBIAL TRAY***TRAPEZOID CEMENTED***Ti Alloy*** SIZE 1F/1T, 2F/1T***FOR CEMENTED USE ONLY***Exactech, Gainesville, FL 32653-1630.***REF 204-04-21***". |  |  | September 23, 2011 | 2011 | Exactech, Inc. | Exactech, Inc. Gainesville, FL is recalling their Optetrak Cemented Trapezoid Tibial Tray, Size 1F/1T, 2F/1T and the Optetrak Rotating Bearing Knee Tibial Tray, Finned, Cemented Tibial Tray, Size 2F/1T. Both trays have the potential to be mislabeled. | Labeling mix-ups | Exactech, Inc. sent an "IMPORTANT MARKET WITHDRAWAL NOTICE" on September 23, 2011, to all affected customers. The customers were instructed to: 1. Immediately cease distribution or use of the products. 2. Extend the information to accounts that has possession of the products. 3. Verify if they have any of the affected products listed. 4. Fax back the attached form. Customers were instructed to contact Exactech inventory representative to confirm quantities at their location. For questions regarding any inventory restocking, the customers were directed to contact Kaya Davis 1-800-392-2832. |  | 16 | 16 | Worldwide Distribution-USA-including the states of CA, MA, OH, and OK, and the countries of Brazil, Colombia, France, Italy, Japan, Luxembourg, and Spain. | Terminated | June 11, 2014 | 992 | 28 | Not_Computer
Z-2069-2012 | 62534 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Reporting |  |  | July 10, 2012 | 2012 | AGFA Corp. | When users selected "Left stenosis" in the "Graft Duplex Conclusion" section in IMPAX CV Reporting Non-Invasive(NIV) module, "Right stenosis" appeared in the sentence display within the printed representation of the report. | Error in labeling | AGFA Healthcare sent an Urgent Field Safety Notice dated July 10, 2012, to all affected consignees indicating that a representative from the IMPAX CV product management or professional services team would be contacting them to perform a review of their system. For questions customers were asked to call 401-604-2180.For questions regarding this recall call 864-421-1754. |  | 6 | 6 | Nationwide Distribution including IN, SC, OH, TX | Terminated | October 29, 2013 | 476 | 6 | Not_Computer
Z-2074-2012 | 62450 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BridgePoint Medical, Stingray Catheter, REF M - 1000, Sterilized with Ethylene Oxide Gas, Prescription Only, Manufacturer BridgePoint Medical Inc. 13355 10th Ave N, Suite 110, Plymouth, MN 55441. |  |  | June 26, 2012 | 2012 | Bridgepoint Medical | BridgePoint Medical Inc. is voluntarily recalling Stingray Catheter, Model M1000 lots BP20121440078 and BP20121560089. BridgePoint Medical, Inc. recently determined the potential of compromised sterility assurance of the product due to a faulty pouch seal. In no case was there any reported adverse clinical event related to the problem. | Process control | The firm, BridgePoint Medical, sent an "Urgent Field Safety Notice" letter dated June 26, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory, quarantine the affected lots and complete and return the enclosed "Verification Form" via fax to 763-225-8718 ,attn: BridgePoint Medical, Inc., Upon receipt of form, BridgePoint will supply the customers with the appropriate materials to return the products.If you have any questions regarding this notice, please call BridgePoint Medical, Vice President of Quality Assurance, at 763-225-8511 or e-mail at cschlawin@bridgepointmedical.com. |  | 95 | 95 (57 USA, 38 OUS) | Worldwide distribution: USA (nationwide) including: AL, AZ, CA, CO, GA, IL, KY, MD, MN, MO, NC, NY, OR, PA, WA, and WI; and countries of: Australia and Sweden. | Terminated | September 06, 2012 | 72 | 95 | Not_Computer
Z-2040-2012 | 62382 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 6800 Miniview, Image-intensified fluoroscopic x-ray system used during diagnostic, surgical and interventional procedures. |  |  | November 20, 2006 | 2006 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The OEC 6800 Miniview was recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system Lock-up, incorrect image display, incorrect dose area calculated air kerma in "film mo | Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. On 11/20/2006, two written notifications were sent to 1513 identified consignees, addressed to the Facility Administrator and the Radiology Administrator for a combined total of 3026 letters. The Urgent Recall Notice, dated 11/8/2006, advised customers that GE Healthcare had identified several intermittent potential safety issues that may occur with the identified affected product. The notification also provided interim or permanent solutions for each of the problems described. For the issues that did not have permanent solutions at the time, the firm was to contact customers and remedy the issues free of charge. For the issues that did have a permanent fix at the time, the GEHC OEC service team was to contact customers to install the correction. If customers had questions or concerns, they were to contact the service team for further information at 800-874-7378 option 8. On 04/23, 2007, a second written notification was sent via certified mail to the same 1513 consignees providing further clarification on the issue of "incorrect Image Display" and a temporary remedy to the issue. |  | 1933 | 1,933 units | Worldwide Distribution | Terminated | May 13, 2013 | 2366 | 1933 | Not_Computer
Z-2065-2012 | 62387 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 8800 Flexview, 892.1650 Image-intensified fluoroscopic x-ray system |  |  | November 20, 2006 | 2006 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The OEC 8800 Flexview was recalled in November 2006 due to several intermittent potential safety issues including navigation tracking inaccuracy, system lok-up, incorrect image display, incorrect dose area calculated air kerma in "film mod | Device Design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE Healthcare sent two Urgent Recall Notice letters to the Facility Administrator and the Radiology Administrator. The letters identified the affected product, problem and actions to be taken. The letters advised consignees that GE Healthcare is actively working on solutions that will permanently resolve each of the issues. For questions or concerns contact GE Healthcare service team at 800-874-7378 option 8. |  | 1257 | 1257 units (472 domestic systems) | Worldwide Distribution - US Nationwide | Terminated | May 13, 2013 | 2366 | 1257 | Not_Computer
Z-2038-2012 | 51222 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800; OEC FluoroTrak 9800 Plus; OEC 9800 Plus; OEC 9800 MD Motorized C-arm System, 892.1650 Image-intensified fluoroscopic x-ray system. |  |  | February 22, 2008 | 2008 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The vertical lift column power supply in the mainframe C-arm of the device is defective and subject to early life failure. | Software design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. The firm, GE Healthcare, Surgery, sent an "URGENT RECALL NOTICE" dated November 8, 2006 via certified mail with return receipt request to consignees/customers with specific instructions. Customers were advised "to stop using the system WITH NAVIGATION until this replacement has been successfully completed. OEC@ FluoroTrak 9800 Plus system may be used as a standard C-Arm providing you review the other issues as listed below." GEHC OEC is actively working on solutions that will permanently resolve each of the noted issues.If you have any questions or concerns regarding these issues, please do not hesitate to contact the service team for further information at 800-874-7378 option 8. Information is available at this number 24 hours per day, 7 days a week. |  | 8214 | 8214 units | Worldwide distribution. | Terminated | February 19, 2013 | 1824 | 8214 | Not_Computer
Z-2077-2012 | 62326 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | .***ANSPACH***Sterile. REF SHD-4820DS***QTY: 1, 4.8 mm x 2mm Twist Drill with Depth Stop. Use with SHORT-HD Attachment ***RX Only |  |  | September 19, 2011 | 2011 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL is recalling their Single Use Sterile Bone Cutting Burrs used with XMax, microMax, microMax Plus, eMax, eMax2 and eMax2 Plus High Speed Drill Systems. It is possible for the cutter to come apart. | Device Design | The firm, Anspach, sent an "URGENT: Medical Device Product Removal" letter dated December 15, 2011, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following:1) Screen their inventory immediately for any of the batches listed in the letter.2) If they have any of the products listed, contact Anspach Customer Support at (800) 327-6887 to arrange for their immediate return and replacement.3) Complete and return the attached Customer Reply form indicating if the item(s) are being returned and confiming their receipt of this letter, and fax to 1-800-327-6661. If they have distributed any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitate to contact Anspach Customer Support at (800) 327-6887. |  | 5 | 5 each | US Nationwide Distribution including the states of: Arizona. and California. | Terminated | July 25, 2013 | 675 | 5 | Not_Computer
Z-2075-2012 | 62582 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Hospira LifeShield Latex-Free 100 mL Burette Set, Convertible Pin, 77 inch with 2 Prepierced Injection Sites and Option-Lok Microdrip Soluset with Calibrated Burette and Precision Drip Chamber; an Rx sterile medical device used to administer fluids; 20 individually wrapped sets per case; Made in Costa Rica, Hospira, Inc., Lake Forest, IL 60045 USA; list 12722-65. |  |  | July 13, 2012 | 2012 | Hospira Inc. | The float valve in the burette sticks to the burette wall and does not open or close properly. | Employee error | The firm, Hospira, sent an "URGENT DEVICE RECALL" letter dated July 13, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to check their inventory and immediately quarantine any affected product; complete and return Reply Form via fax to 1-877-722-7019 or e-mail to: hospira4921@stericycle.com; return affected product to Stericycle using the label provided with the letter; and if they have distributed the product further, notify their accounts that received the product . Please contact Hospira Customer Care at 1-877-946-7747 or your Hospira representative for information regarding product availability. |  | 4420 | 4,420 sets | Worldwide Distribution-USA (nationwide) including the states of California, Colorado, Connecticut, Florida, Illinois, Kentucky, Maryland, Massachusetts, Minnesota, Oklahoma, Texas, Utah, Virginia, Washington, West Virginia and Wyoming, and the countries of Barbados and Canada. | Open, Classified |  |  | 4420 | Not_Computer
Z-2081-2012 | 54724 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | The Handi-Fil Disposable Syringe Fill Tube is a sterile plastic tube, approximately 10 inches in length, with a "J" shaped curve that projects upward approximately two inches. |  |  | February 01, 2010 | 2010 | Mallinckrodt Inc | A report was received from a Japanese distributor regarding a complaint from a hospital customer stating that ten packages of Handi Fil were not sealed. | Packaging process control | Covidien send an Urgent Device Recall letter dated February 1, 2010, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product and discontinue use of the product immediately. Customers were asked to complete the attached business form and report the quantity of affected product currently in their stock. Once the form is completed customers should fax the form to 314-654-8206. Once the business notification form is received, an RGA number will be sent to the customer for return of the product. Customers will receive credit for all returned product. For questions regarding this recall call the Product Monitoring Department at 800-778-7898. For questions about credit for the product, please call Customer Service at 888-774-1414, menu option 1, then menu option 2.For questions regarding this recall call 513-948-5719. |  | 75 | 75 cases (100 units per case) | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, GA, IL, KS, KY, LA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA, WI, &amp; WV and countries including: AUSTRALIA, BELGIUM, BRAZIL, CANADA, JAPAN &amp; SINGAPORE. | Terminated | August 03, 2012 | 914 | 75 | Not_Computer
Z-2084-2012 | 62365 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tibial Punch Tower Assembly. Non-Sterile. CAT NO. 8000-1089. Howmedica Osteonics Corp., 59 Route 17, Allendale, NJ 07401. |  |  | December 22, 2011 | 2011 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report indicating that a Tibial Keel Punch could not be inserted into the slot of the Tibial Punch Tower. The slot that guides the stem of the keel punch into the channel of the tower was later determined to be undersized. | Nonconforming Material/Component | Stryker Orthopaedics sent an Urgent Product Recall/Recall Acknowledgement letter dated January 10, 2012, to all affected customers. The letter identified the product(s) the problem and the action needed to be taken by the customer.Customers were instructed to contact their Stryker Sales Representative to arrange for return of the affected product. Customers were asked to fax back the attached Product Recall Acknowledgement Form as soon as possible to 201-831-6069. For questions customers should call 201-972-2100.For questions regarding this recall call 201-831-5970. |  | 152 | 152 units | Worldwide Distribution | Terminated | June 13, 2013 | 539 | 152 | Not_Computer
Z-2078-2012 | 62462 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza. |  |  | February 24, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential malfunction when using syngo.plaza with software version VA20C_HF01. For datasets with distance measurements it can happen, that all images are not loaded to Viewer but no error message is displayed to user. | Software design | Siemens sent a Customer Safety Advisory Notice letter dated February 24, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to ensure that the prefix defined in the viewer settings for "Content Creator's Initials" is not empty. Then all images will be display correctly. This has to be checked for every user. Further questions please call 610-219-6300. |  | 2 | 2 | (USA) nationwide including the states of CA, FL, GA, MA, MN, MO, NJ, NY, OH, WA, and WI. | Terminated | January 14, 2014 | 690 | 2 | Computer
Z-2083-2012 | 62481 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Slice CT Scanner System is a computed Tomography X-ray system, and has a Part number 989605858501. |  |  | June 08, 2012 | 2012 | Philips And Neusoft Medical Systems Co., Ltd. | One of the four screws that secure the Oil-accumulator to the X-Ray Tube pulled out in the NeuViz 16 CT scanner, which caused the Oil-accumulator to detach. | Device Design | Philips Healthcare North America Company sent an "Urgent Device Correction NeuViz 16 X-Ray Tube Oil-accumulator Detached" letter and attachments of Annex A Method to Identify Affected Products and Annex B Instruction for Mechanic, dated June 14, 2012, to all affected customers. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to follow the instructions on Annex A and Annex B.Philips and Neusoft Medical Systems planned to release the mandatory Field Change Order (FCO) and FCO kits to all customers by the end of June 2012.If you have any questions about this corrective action program, please contact the Service Support Department, Neusoft Medical Systems, Co., LTD by e-mail to NMS-SERVICE@NEUSOFT.COM |  | 15 | 15 units distributed in the US | Nationwide Distribution -- North Carolina, Ohio, and Texas. | Terminated | September 17, 2012 | 101 | 15 | Not_Computer
Z-2086-2012 | 62117 | No | MKQ |  |  | Pathology | PMA | PrepStain Upgrade Kit: 1.3.0.3/1.2 DiTi Cone. |  |  | December 08, 2011 | 2011 | Tripath Imaging, Inc. | Some Prep Stain Kit, 1.2 DiTi Cones have been found to be leaking and the PrepStain Preparation kits contain incorrect preventive maintenance (PM) instructions. | Nonconforming Material/Component | BD Diagnostics sent a letter dated March 2012 to all affected customers. The letter identifies the product, problem and actions that will be taken by the firm. Contact BD Diagnostics Technical Support at 1-877-822-7771, option 2 for questions regarding this letter. |  | 1274 | 1274 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WY, and WV and the countries of Belgium, Brazil, Canada, China, Columbia, Hong Kong, India, Japan, Korea, Peru, Singapore, Taiwan and Victoria (Australia). | Terminated | November 08, 2013 | 701 | 2659 | Not_Computer
Z-2076-2012 | 62007 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 |  |  | May 25, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because the Alaris PC unit model 8015 has a component on the PC unit power supply board that is causing an error code(120.4630) "System Error" or "Missing Battery Error" at start up. The error code is accompanied by both an audible alarm and visual error messages on the PC unit screen. | Component design/selection | Carefusion sent an Urgent: Medical Device Recall Notification letter dated June 27, 2012 to all their customers who purchased the Alaris PC unit (model 8015). The letter identified the affected units, explanation of the problem, potential risk and actions to be taken. Customers were instructed to not return their devices. If customers observe either of the error messages identified in the letter, then they were instructed to remove the PC unit from use and contact the Carefusion Support Center. Carefusion will contact their facility by phone within 60 days of receiving this notification to schedule a visit to replace their affected power supply board on their PC unit. Customers were instructed to promptly complete and return the enclosed customer response card to expedite the corrective action process. For questions and support contact:CareFusion Support Center at 888-562-6018 for (Recall Related Questions), Customer Advocacy at 800-854-7128, Option 1, Option 1, Option 3 or email customerfeedback@carefusion.com for (Adverse Event Reports), Technical Support at 888-812-3229 for (Technical Questions Regarding the Alaris System). |  | 112912 | 112,912 units total (106,888 units in US) | US Nationwide Distribution | Terminated | May 13, 2014 | 718 | 112912 | Not_Computer
Z-2082-2012 | 62453 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Affinity Four Birthing Bed P3700 |  |  | April 02, 2012 | 2012 | Hill-Rom, Inc. | During review of our last braking system via that field corrective action, the system still did not meet our expectations, which is to eliminate the hazards of patient falling. Bed movement immediately prior to, or during the birthing process may still create distractions and delays, resulting in patient or caregiver falls, or contribute to a deviation in aprocedure, (e.g. IV insertion, epidural | Device Design | Hill-Rom sent an URGENT FIELD SAFETY NOTICE letter dated April 2, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were informed that Hill-Rom will visit their facility to modify the affected products, however, until then follow the actions to be taken outlined. For questions concerning this request or the procedure as outlined contact Hill-Rom Technical Support at 800-445-3720 (U.S.). |  | 12726 | 12,726 devices | Worldwide Distribution - US Nationwide and countries including: Argentina , Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Brunei, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guam, Guyana, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, Turks and Caicos Islands, United Arab Emirates, United Kingdom, United States, US Virgin Islands, Venezuela and Vietnam | Terminated | July 21, 2015 | 1205 | 12726 | Not_Computer
Z-2079-2012 | 61974 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge brand --- large capacity, floor loading jet spray, washers for mechanical washing, intermediate level, thermal disinfection and drying of moisture and temperature stable Hospital case carts, sterilization containers, material handling carts, stands and utensils. --- Model numbers Getinge 9120, 9122, 9125, and 9128 |  |  | May 25, 2012 | 2012 | Getinge USA Inc | Getinge Disinfection AB is voluntarily correcting Washer Disinfectors, Getinge Electrically Heated 9100-Series manufactured between 2009-05-04 through 09-26-2011. This action is based on a report from the ground fault breaker manufacturer of a production error. | Nonconforming Material/Component | Getinge USA issued a Device Correction Notice dated June 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to perform a trip test to determine if they have a functioning ground fault breaker. Customers will be contacted by a Getinge representative to schedule an appointment to exchange the earth fault breaker. Customers were provided a Customer Response Form to complete and fax back to 585-272-5033. For questions customers should call 800-475-9040, ext. 5270.For questions regarding this recall call 585-475-1400. |  | 8 | 8 washers | Worldwide Distribution - USA including Maryland, Indiana, New Mexico, Massachuesetts, Texas, Kansas and the country of Canada | Open, Classified |  |  | 8 | Not_Computer
Z-2090-2012 | 62456 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Rejuvenate Modular Stems |  |  | April 30, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has updated the instruction for use (IFU) (product correction)for the ABGII Modular and Rejuvenate Modular Hip Systems. This is based on a reported rate of less than one percent for revisions potentially associated with fretting and or corrosion at or about the modular neck junction. On June 29, 2012, Stryker has decided to remove the affected product from the field since there may be an | Device Design | Stryker sent an "URGENT PRODUCT CORRECTION" letter dated April 23, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter included a Product Correction Acknowledgment Form for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-831-5970 for questions regarding this notice.Stryker sent and "Urgent Update Product Recall" Letter dated January 3, 2013 to all affected customers. The new letter includes updated corrective actions for patients. Another letter was sent out on 7/25/2014. |  | 481092399 | (NLS/NLV) - 23193 units and (SPT) - 20,743 units | Nationwide Distribution | Open, Classified |  |  | 481101402 | Not_Computer
Z-2100-2012 | 33854 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD Cyclesure Biological Indicator (BI), P/N 14324 are sold in cases, packaged as part of STERRAD System Validation kits and STERRAD test packs. |  |  | June 28, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling certain lots of the STERRAD CYCLESURE 24 Biological Indicator (BI) product because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | Incorrect or no expiration date | ASP has posted the STERRAD CYCLESURE Biological Indicator customer letter on their external facing website at: wwww.aspjj.com/us/news/cyclesureletter on July 3, 2012. The letter informed customers who purchased STERRAD CYCLESURE 24 Biological Indicator (BI) product that Advanced Sterilization Products (ASP) is recalling certain lots because ASP has determined that it does not have adequate data to support the entire duration of the product's labeled shelf-life. The letter also provided the customers with what actions are needed for the recall. Customers were instructed to return all unused STERRAD CYCLESURE 24 BI product that exceeds the newly determined expiration dated using the enclosed prepaid UPS return label to Stericycle, 2670 Executive Drive Suite A, Indiapolis, IN 46241. Customers were instructed to count their inventory of product that exceeds the newly determined expiration date and record the data on the enclosed Business Reply Card and packing slip that are included with the customer recall letter. Customers were instructed to complete the requested information on the enclosed postage paid Business Reply Card and return it to Stericycle. Customers were instructed to contact Stericycle at (877) 907-7030 if they need additional packing labels. If customers further distribute the affected products to other locations, then they were instructed to contact Stericycle to arrange for return of the product. Customers were instructed that new shipments of STERRAD CYCLESURE 24 BI product will be shipped on July 9, 2012 that has accurate expiration dates that do not require conversion. Customers wil be able to differentiate between current product and the newly shipped product by the lot number. Current product has a lot number containing 6 characters, eg. 123456.Newly shipped product will have a lot number containing 8 characters, eg. 123456EE.Customers with any additional questions were instructed to contact (877) 907-7030 or contact their A |  | 38693 | 38,693 units | Worldwide Distribution - USA (nationwide) | Terminated | August 13, 2013 | 411 | 38693 | Not_Computer
Z-2105-2012 | 62553 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | KIT,DISPOSABLE ACCESSORY,4ARM,IS3000,5 PACK; |  |  | July 13, 2012 | 2012 | Intuitive Surgical, Inc. | Specific lots of the Instrument Arm Drapes were manufactured with a sterile adaptor that may have difficulty engaging an instrument. | Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Correction letters dated July 16, 2012 via Fed Ex to all affected customers in the US. The customers outside of the US were notified starting July 17, 2012. The letter identified the affected products, problem and actions to be taken. Customers were instructed to Identify affected lots, order replacement products and return both full and partial boxes of the affected products for credit as soon as unaffected products are available. All staff should be made aware of the issue. Customers were instructed to complete and return the acknowledgement form immediately.Questions and concerns should be directed to Intuitive Customer Service at 1-800-876-1310.EXPANSION:Intuitive discovered an additional lot (Lot DA121075) of Instrument Arm Drape IS2000, 20 pack that should have been included in the recall and expanded the recall on September 4, 2012. Letters to customers were sent by Fed Ex on September 5, 2012. |  | 92390 | 92390 total instrument Drapes | Worldwide Distribution - US Nationwide and the countries of: Australia, Spain, India, Italy and Saudi Arabia. | Terminated | June 12, 2013 | 334 | 92390 | Not_Computer
Z-2091-2012 | 61764 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Nextgen Complete Knee Solution Femoral and Provisional Impactor/ Extractor, REF 00-5901-026-00, Non-Sterile |  |  | April 30, 2012 | 2012 | Zimmer Inc. | The design of the spring clip for the jaw assembly was updated in March 2011 due to fractures of the spring clip. Zimmer has received complaints for the spring clip fracturing on the updated design. The spring clip may fracture or fall off of the instrument during use, and a device fragment may be left in the surgical site. The risk of disassembly or fracture occurs when the instrument is typicall | Component design/selection | Zimmer sent an Urgent Medical Device Recall letter dated April 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Requests for alternative instruments should be requested by calling (574) 372-4470. Fax Back Inventory Return Certification Form is to be sent back to Zimmer at (574) 372-4265. All recalled instruments available at the distributor are to be returned to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. The Distributor notice included instructions to notify surgeons immediately of the recall and requested a list of hospitals and surgeons receiving the device be sent back to Zimmer via E-mail. Zimmer is continuing distribution of the replacement jaw ( 005901-026-40) until alternative tools are available. The distributors are asked to identify and deliver preferred alternative instruments to affected accounts. Beginning May 7, 2012 Urgent Recall notices are being sent to hospitals and surgeons identified with instructions to acknowledge the recall communication.For questions regarding this recall call 574-372-4807. |  | 1662 | 1662 units | Worldwide Distribution - USA (nationwide) and the countries of Singapore, China, Hong Kong, Canada, Australia, Japan, Europe, Germany, India, and Finland. | Terminated | June 04, 2014 | 765 | 1727 | Not_Computer
Z-2108-2012 | 62590 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Getinge Group Operating Table System Magnus Carbon-Fibre Table Top Version 1180.16 |  |  | April 13, 2012 | 2012 | Maquet Medical Systems, Usa | As a result of a small number of customer complaints, MAQUET conducted an investigation and identified a potential problem with specific MAGNUS Carbon-Fibre Table Tops manufactured between February 23, 2010 and February 28, 2012 and distributed between November 11, 2010 and April 6, 2012. There is a potential issue with the motor brakes for the longitudinal shift drive, and the tilt drive. | Device Design | MAQUET GETINGE Group sent An "Urgent Device Field Correction" notification, dated 13 April 2012 via Federal Express to all affected customers. An acknowledgement form is included with the customer letter which can be mailed or faxed back to Maquet Quality by the notified customers. If the completed Acknowledgement form is not received within 15 days, up to three follow ups will be made and documented via telephone call to each customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and notify their MAQUET Service Representative immediately. The MAQUET Service Representative will contact customers to make arrangements to replace both te logitudinal shift and tilt drives of their table top(s). For questions customers should call 888-627-8383, press option 2, followed by option 1.For questions regarding this recall call 201-995-8968. |  | 191 | 164 Units located in 27 Countries | Worldwide Distribution - USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Germany, Denmark, Spain, France, Hong Kong, Ireland, Israel, India, Italy, Japan, South Korea, Netherlands, Poland, Russia, Saudi Arabia, Sweden, Singapore, and Thailand. | Terminated | September 04, 2013 | 509 | 191 | Not_Computer
Z-2120-2012 | 62566 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Octane-C Cervical, 6Âż Lordotic |  |  | June 15, 2012 | 2012 | Exactech, Inc. | Exactech, Inc. initiated a recall of Equinoxe Replicator Plates, Equinoxe Locking Screws, Equinoxe Compression Screw Kits, Equinoxe Fracture Adapter Tray Screws, and Octane-C Intervertebral fusion devices due to the firm determining that the device labeling identifies the affected parts as having a ten year shelf life; however, the packaging configuration that was used it qualified to a five year | Labeling False and Misleading | Exatech sent an Important Voluntary Recall Notice dated June 18, 2012, to all affected customers. The Notice identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to immeidately case distribution or use of the affected product. Verify that they have the affected product. Contact their Exatech Inventory Representative for further instruction on how to return their products. The firm sent out recall notification to consignees. Customers were asked to complete and return the Voluntary Recall Fax Notice to 1-352-378-2617. Customers with questions should call 1-800-392-2832.For questions regarding this recall call 352-377-1140. |  | 54877 | 54,877 units for all products in Recall Event | Worldwide Distribution - USA including FL, MO, MA, AL, KY, IN, NC, LA,PA, VA, OH, NY, ME, MI, TX, KS, TN, OK, NJ, WA, CO, GA, CA, MN, IA, OR, and IL and Internationaly to Austria, Luxenbourg, Sweden, Germany, Spain, France, Greece, Italy, Nederland, United Kingdom, Canada, India, Australia, Switzerland, and Sao Palo. | Terminated | July 03, 2013 | 383 | 54877 | Not_Computer
Z-2113-2012 | 62493 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Padded Swing Away Arm Rest with Receiver; |  |  | June 26, 2012 | 2012 | Sunrise Medical (US) LLC | Reports of injury when the Padded Swing Away Armrest is used for full body weight. Owner's manual states that armrest is only designed for resting of forearm. | Labeling design | Sunrise Medical sent a Safety Notice letter dated June 25, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter request that clinicians/suppliers post this safety notice in a conspicuous location, making customers and therapists aware of the issue. The letter also encourage end users to read the owner's manual and follow the recommenations for proper usage of the recalled product. Customers are instructed to complete and return the Supplier Notification and Response Form. For questions contact Sunrise Medical Regulatory Affairs at 888-208-4901. |  | 14035 | 14, 035 | Worldwide Distribution - US (nationwide) and the countries of Argentina, Brazil, Columbia, Costa Rica, Mexico, Panama, Uruguay and Venezuela. | Terminated | November 01, 2012 | 128 | 14035 | Not_Computer
Z-2110-2012 | 62526 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | Beekley A-Spots Floral Skin Markers -Order Number: 751 |  |  | July 09, 2012 | 2012 | Beekley Corporation | Digital Skin Mammography Marker separating from the backing of the spot material | Unknown/Undetermined by firm | Beekley Medical issued an Urgent Medical Device Recall letter dated July 9, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were advised to contact Beekley Medical's Customer Service Team at 1-800-233-5539 to return affected lots for replacement. Customers were instructed to complete and return the attached reply form. |  | 32640 | 32,640 units | US Nationwide Distribution | Terminated | July 09, 2013 | 365 | 342910 | Not_Computer
Z-2114-2012 | 62626 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap SIBD 4 degree Trial Implant (SJ768T-US) and Aesculap SIBD 9 degree Trial Implant (SJ780T-US) |  |  | July 12, 2012 | 2012 | Aesculap Implant Systems LLC | The 4 degree and 9 degree trial instruments (SJ768T-US/SJ780T-US) are correctly dimensioned, but the part number etched on the instrument may be incorrect. The 4 degree trial instrument (SJ768T-US) may be marked with the incorrect part number of the 9 degree trail instrument (SJ780T-US) and the 9 degree trial instrument (SJ780T-US) may be marked with the incorrect part number of the 4 degree tria | Employee error | Aesculap sent an Important Product Removal Notification letter dated July 11, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product by contacting their local sales representative or the customer service department at 1-866-229-3002 for a Returned Goods Request )RGR) number. Customers were asked to complete and return the Distribution Inventory Sheet. For questions call 610-984-9260 or 610-984-9274. |  | 5 | 5 | US Nationwide Distribution - including the states of CO, OH, OK and TX. | Terminated | February 22, 2013 | 225 | 5 | Not_Computer
Z-2123-2012 | 61655 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | IntegraÂż Advansys Medial Lisfranc Plate (MLP), Integra, 311 Enterprise Drive, Plainsboro, NJ 08536 |  |  | February 09, 2012 | 2012 | Integra LifeSciences Corp. | Incorrect technique applied may cause potential rupture of the tibialis anterior tendon during surgery. | Labeling design | INTEGRA sent an Urgent Medical Device Correction Notice dated February 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to identify members of surgical team that ues the affected product and make them aware of the Medical Device Correction Notice. Customers were asked to sign and return the enclosed Medical Device Correction Notice Acknowledgement and Return Form and to maintain a copy of the notification and signed copy of the acknowledgment form for their records. For any questions regarding this recall call 609-936-6832.. |  | 2437 | 2,437 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NM, NJ, NY, NV, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA &amp; WI and internationally to Belgium; Danemark; France; Germany; IsraÂżl; Italy; Lebanon; Luxemburg; Netherlands; Norway; Portugal; Saudi Arabia; Spain; Switzerland; Turkey &amp; United Kingdom. | Terminated | February 19, 2013 | 376 | 2437 | Not_Computer
Z-2126-2012 | 62449 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | OEC MiniView 6800 Mobile; MDL number A349855. |  |  | March 27, 2007 | 2007 | GE OEC Medical Systems, Inc | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC had discovered an issue with the hard disk drives that were installed on OEC device models: UroView 2800; 9800 Plus Digital; MiniView 6800; and FlexiView 8800 Mobile C-Arm. The devices were manufactured between October 15, 2004 and | Software design | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being classified by the Agency. GE OEC sent an Urgent Safety Advisory Notices dated October 11, 2006 and February 21, 2007 via certified mail delivery letter to affected customers. The notices identified the affected products, problem and actions to be taken. Customers were informed that a GEHC OEC field service engineer would be contacting their facility to arrange to have a replacement hard disk drive installed on the affected system(s). Customers were instructed to ensure that images on the hard disk drive were stored via an alternate permanent media as the hard disk drive approaches the 170 image capacity until a GEHC OEC field service engineer had replaced the drive. For questions regardiug this issues contact the service team at 800-874-7378. |  | 82 | Part of a total of 82 units distributed. | Worldwide Distribution - US (nationwide) and the countries of: Canada, France, Japan, and China. | Terminated | August 02, 2012 | 1955 | 82 | Not_Computer
Z-2107-2012 | 61247 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical, MedfusionÂż Syringe Infusion Pump, Model 4000, Software Version V1.0 &amp; V1.1, Reorder Numbers 4000-0100-50, 4000-0100-95, 4000-0101-50, 4000-0101-51, and 4000-0101-78. |  |  | February 14, 2012 | 2012 | Smiths Medical ASD, Inc. | Smiths Medical has identified a software anomaly in the Medfusion 4000 Syringe Infusion Pump that causes a device history log corruption and triggers a Watchdog Fail-Safe alarm. When the device exhibits this failure mode, visual and audible alarms will sound and the device ceases operation. | Software design | Consignees of MedfusionÂż 4000 pumps with software version 1.0 &amp; 1.1 were visited by a Smiths Medical representative on 2/14/12 informing them that their pumps require a software upgrade. A record of this software upgrade was documented as part of the records for this action. |  | 607 | 607 pumps (257 USA, 350 Canada) | Worldwide Distribution -- USA, including the state of CA, and country of Canada. | Terminated | November 14, 2012 | 274 | 607 | Computer
Z-2124-2012 | 62555 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IRISpec" CA/CB two-part urine chemistry control twin set, Iris Diagnostics part number 800-0074. |  |  | July 09, 2012 | 2012 | Iris Diagnostics | The firm recalled after it received complaints of IRISpec CA/CB two part Urine Chemistry Control, 101131-12, where the nitrite results read negative for the part B control, which should be positive. | Nonconforming Material/Component | The firm initiated the recall on 07/09/12 by sending out a Medical Device Recall notification to customers via e-mail. The recall notification identified the affected product, reason for recall, and instructions on the recall. Customers were to confirm how many sets of unopened or partially opened IRISpec CA/CB that was in their inventory, dispose of the affected product, and provide written confirmation of the product's disposition. The letter also stated that Iris Diagnostics would replace any unused inventory that was returned at no charge to the customer. Contact information for the recall was also provided. The communications also included a fax response sheet for return/replacement. |  | 862 | 862 | Nationwide Distribution | Terminated | January 27, 2014 | 567 | 862 | Not_Computer
Z-2122-2012 | 62557 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | The Reliance 130 and 130L Cart Washer/Disinfector is a high capacity mechanical washer intended for use in the washing, low-level disinfecting, and drying of carts, utensils, beds and other miscellaneous reusable items used in the care of patients at healthcare facilities / Crated for shipping. |  |  | July 14, 2012 | 2012 | Steris Corporation | To ensure that users are properly operating the Reliance 130 and 130L Cart Washers, a door close confirmation button and door open delay will be installed on your cart washer. | Labeling design | Steris Corp. sent a Urgent Field Correction Notice dated July 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To ensure users are properly operating the Reliance 130 and 130L Cart Washers a 'door close' confirmation button and a door open delay will be installed on the affected cart washer(s). A STERIS Service Representative will contact the affected facilities to arrange for STERIS to upgrade their Reliance 130 and 130L Cart Washer(s). For further information or if affected customers have any questions regarding the STERIS visit to their facility, they can contact STERIS Field Service Dispatch at 1800-333-8828 |  | 367 | 367 | Worldwide Distribution -- USA (nationwide) including the states of : AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY. and the countries of : Australia, South Korea, France, Italy, Lebanon, Russia, Saudi Arabia, Spain, UAE, United Kingdom and Canada. | Terminated | June 06, 2014 | 692 | 367 | Not_Computer
Z-2121-2012 | 62512 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Perkin Elmer Genetic Screening Processor (GSP) |  |  | June 27, 2012 | 2012 | Perkinelmer | Contamination from GSP Shaker unit results in decreased signal activity of assay measurements for Genetic Screening for Neonatal IRT, TSH and IRT | Component design/selection | PerkinElmer sent an Urgent: Medical Device Recall letter dated June 27, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement units of the recalled product will be installed as soon as they are available. Customers were instructed to complete and return the response form attached. For further information contact your local PerkinElmer representative or GS. Technical.Support@PerkinElmer.com |  | 78 | 78 units | Worldwide Distribution - US (nationwide) and the countries of Australia, Belgium, Brazil, Egypt, France, Germany, Netherlands, Norway, Saudi Arabia,Switzerland, | Terminated | April 18, 2014 | 660 | 78 | Not_Computer
Z-2080-2012 | 62432 | No | IZH | Mammographic x-ray system. | Radiology | Radiology | 510(k) | Bella Blanket Protective Coverlet, for Mammography, 18cm x 25cm or 24cm x 30cm, Product Code: 311. The purpose of the product is to remove the cold by placing this sheet between the patients breast and the imaging receptor plate. |  |  | July 02, 2012 | 2012 | Beekley Corporation | Imaging artifact can occur on mammogram result | Nonconforming Material/Component | Beekley Medical issued "Urgent: Medical Device Recall" notifications dated June 29, 2012 that were sent via certified mail USPS on July 2, 2012. The product issue was was described and customers were provided the recommended actions. Customer Service number is 1-800-233-5539. |  | 86000 | 86,000 units | International Distribution including Nationwide(USA), and the countries of Japan and Canada. | Terminated | July 09, 2013 | 372 | 86000 | Not_Computer
Z-2129-2012 | 62538 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Cardiac Catheterization Waste Management System w/1000mL Waste Bag, 2 Fluid Admin Sets |  |  | April 17, 2012 | 2012 | ICU Medical, Inc. | The firm initiated the recall due to a molding defect in a high pressure stopcock where a smallpercentage of the side ports could become totally occluded. | Nonconforming Material/Component | ICU Medical sent an Urgent Medical Device Recall letter dated April 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to inspect inventory, destroy affected products according to hospital protocol, complete and return verification form and contact ICU Customer Service to order replacement products. The letter states that ICU Medical will credit customer for any product destroyed once the completed verification form is received. This recall should be carried out to the user level and passed on to all those who need to beaware within your organization or any other organization the device may have been transferred to.Should you have any questions or require assistance relating to this recall please contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com |  |  |  | Nationwide distribution | Terminated | August 23, 2012 | 128 |  | Not_Computer
Z-2132-2012 | 62528 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Extraneal and/or DIANEAL LOW CALCIUM Peritoneal Dialysis Solution in Plastic Container, PL-146 used with Automated Peritoneal Dialysis System/Automated PD Set with Cassette; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. The HomeChoice dialysis system is a personal cycler automated peritoneal dialysis system for pediatric and adult renal patients with fill volumes ranging from 60 mL to 3000 mL. |  |  | July 06, 2012 | 2012 | Baxter Healthcare Renal Div | Baxter discovered that a Baxter technical service representative was providing incorrect instructions for inspection of solution bags for a blocked frangible prior to connection to the Automated PD Set with Cassette. | Employee error | Baxter Healthcare Corp. sent "Important Product Information" letters dated July 6, 2012 to the affected home patients and their clinicians, via first class mail on the same date. The letter identified the product, problem and actions to be taken. The customers were instructed to refer to the Extraneal and/or Dineal package insert, Directions For Use - Preparation for Administration and the HomeChoice APD Systems At-Home Guide for proper procedures to connect to a solution line. Customers are instructed to retain a copy of the letter. The customers were also requested to complete the attached customer reply form and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Global Technical Services, Renal Technical Support at 1-800-553-6898, option 1, 24 hours a day.For questions regarding this recall call 1-800-422-9837. |  |  |  | Nationwide Distribution includingArizona, California, Ohio, Pennsylvania and Virginia. | Terminated | October 15, 2012 | 101 |  | Not_Computer
Z-2135-2012 | 62504 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | RF Cannula, 100 mm,Curved 10 MM, 20 G, REF 0406-630-125 RX only, Disposable ( 10/pk) |  |  | July 17, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Stryker has received 6 complaints that the 5mm needle was placed in 10 mm package and 10 mm needle was placed in the 5 mm package. If the clinician does not recognize that the active tip on the cannula is 10 mm instead of 5mm and the cannula is used in a cervical procedure there is a potential for a thermal skin lesion due to a larger lesion creation. If the 10 mm active tip cannula, as opposed | Labeling mix-ups | Stryker sent an Urgent MEDICAL DEVICE Recall Notifications dated July 16, 2012 to all consignees. The letter identified the affected products, problem and actions to be taken. Customers/distributors were instructed to check their inventory and quarantine any affected product found, complete and fax the business reply form to Stryker Instruments Regulatory Department, 866-521-2762 and return all affected product using the pre-paid shipper provided. Upon receipt of the recalled product, a replacement will be issued.For questions regarding this recall please contact Stryker Instruments (Recall Coordinator) Monday-Friday 8am-5pm (EST) at 269-389-4354. |  | 226 | 226 | Worldwide Distribution - US Nationwide and the countries of England, France, Netherlands, and Switzerland. | Terminated | May 02, 2013 | 289 | 436 | Not_Computer
Z-2133-2012 | 62371 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | I-STAT EG7+ Cartridge Abbott Point of Care, Princeton, NJ. |  |  | January 01, 2011 | 2011 | Abbott Point Of Care Inc. | i-Stat EG7 cartridges may generate falsely depressed PCO2 results for samples above 26 mm hg when stored at room temperature for four weeks or longer. | Nonconforming Material/Component | Abbott Point of Care, Inc sent a Urgent Recall Notice letter dated January 14, 2011 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of Cartridges from the identified lot. All unused cartridges will be credited and are to be returned as instructed on the enclosed Business Reply Card.If you have any questions regarding this information please contact your Abbott Point of Care Technical Support at 800-366-8020 Option 1, or your Abbott Point of Care representative. |  | 52750 | 52,750 cartridges | USA ( nationwide ) | Terminated | May 22, 2013 | 872 | 52750 | Not_Computer
Z-2138-2012 | 62514 | No | FIN | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | TANGO3 Water Storage Tank with Ozone Disinfection System, Tango3, LLC, 132 Citizens Blvd., Simpsonville, KY 40067. Models are WOZ-250-PE, WOZ-240-SS, WOZ-180-SS. The TANGO3 Water Storage Tank with Ozone Disinfection System is intended to be used for disinfection of the water distribution system of a dialysis facility. |  |  | March 26, 2012 | 2012 | Tango3 LLC | During an FDA inspection, the firm was notified of Quality System deficiencies, which included lack of procedures for design control and customized modifications that were either not reviewed or received proper clearance from FDA. There was no assurance that all systems in the field were working properly and operating at a minimum of 0.2ppm of ozone. | Process control | On 3/26/2012, the firm sent "Urgent: Customer Notification" letters and Acknowledgement Forms to their customers. The issue was described and recommended actions were provided. Customers were given the number 866-782-7286 to call if questions. |  | 44 | 44 units | Product was shipped to the following states: CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, MA, MD, ME, NH, NJ, NM, OH, PA, RI, TX and VA. | Terminated | March 01, 2013 | 340 | 44 | Not_Computer
Z-2145-2012 | 62617 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew 5.5 mm Dyonics Bonecutter ,Electroblade |  |  | July 14, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | Device labeled with an expiry period of 48 months instead of 36 months; an additional 12 months for which product and package reliability has not been validated. | Error in labeling | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification Letter dated July 2012 via Federal Express to all affected customers. The letter identified the affected product, potential risk and actions to be taken. Customers were instructed to examine inventory, quarantine immediately, complete the form attached indicating the lot numbers and quantities. The letter also instructed customers to contact Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) for a Return Authorization (RA) or send an email to endo.andreturns@smith-nephew.com. They will provide customers with instructions for returning affected products and receiving a credit. Each OUS affiliate shall be notified via (email or equivalent) of the recall by Smith &amp; Nephew Inc.; the affiliate must then notify each facility affected by the recall.For question concerning this recall call 405-917-8585 or 978-749-1630. |  | 108 | 108 units | Worldwide Distribution - US Nationwide and the countries of : Australia, Canada, Korea, Turkey, UAE, and Japan. | Open, Classified |  |  | 448 | Not_Computer
Z-2137-2012 | 62583 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | SPIDER Surgical Instrument: Flexible Clip Applier |  |  | July 02, 2012 | 2012 | Transenterix, Inc | Some clip applier units may be inadequate to achieve clip closure. | Process control | TransEnterix sent an Urgent Device Recall letter date July 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product, remove the affected product from their inventory and either return it or dispose of it to prevent usage. Customers were instructed to complete the attached Inventory Reconciliation Formand return the form even if they no longer have any affected units in their inventory. Customers were instructed to ship the affected product to TransEnterix using the enclosed shipping label. If returning a form only, they should e-mail to adunn@transenterix.com. For questions customers were instructed to call 1-888-879-4111.For questions regarding this recall call 919-765-8420. |  | 89 | 89 boxes(6 units per box) | Worldwide Distribution - USA (nationwide) and the countries of Austria, Germany, Italy, Qatar and Switzerland | Terminated | February 27, 2013 | 240 | 89 | Not_Computer
Z-2143-2012 | 62023 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab with SA INST versions: 3.1.6.12, 4.0.1.26-4.0.1.32, 4.0.2.21-4.0.2.51, and 4.0.3.5-4.0.3.13. |  |  | October 17, 2011 | 2011 | SCC Soft Computer | On 10/17/2011 SCC Soft Computer, Clearwater, FL initiated a correction on SoftLab with SA INST versions: 3.1.6.12, 4.0.1.26-4.0.1.32, 4.0.2.21-4.0.2.51, and 4.0.3.5-4.0.3.13. A client reported an issue which caused two different sets of results from the Sysmex series of Hematology instruments to display under the same order on the instrument menu. | Software design | SCC Soft Computer's proprietary Task Management System sent a correction communication letter to all consignees with affected software. The letter identified the affected product, problem and corrective actions to be taken. The letter also recommended alternative working solution if the corrective actions cannot be used. Consignees with affected software are scheduling receipt of the software correction. For questions call 727-789-0100. |  | 88 | 88 | US Nationwide Distribution | Terminated | September 03, 2014 | 1052 | 88 | Computer
Z-2151-2012 | 62665 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PopLok Punch labeled in part: |  |  | July 06, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled one lot of their PopLok bone Punch 3.5MM because the metal shaft diameter is 4.5mm when it should have been 3.5mm. | Process change control | The recall is being conducted to the end user level. ConMed Linvatec Urgent Medical Device Recall Notifications were sent by FedEx next day delivery or equivalent method to notify all direct consignees. The notifications identified the affected product along with the reason for the recall and health risk. All accounts were asked to immediately check their facility for the affected product, stop using any of the product with specified lot number, and to segregate the product. The letter also asked customers to return all affected product to ConMed Linvatec and respond using a Reply Form, even if there is no affected product available. If customers have any additional questions, they should contact the Customer Service department at 800-237-0169. |  | 2 | 2 | Worldwide Distribution -- US and Korea. | Terminated | November 21, 2013 | 503 | 2 | Not_Computer
Z-2149-2012 | 62563 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Plasma Pump Model PP-04a |  |  | July 09, 2012 | 2012 | Apheresis Technologies, Inc. | On 07/09/2012, Apheresis Technologies, Inc. initiated a recall on the Plasma Pump Model PP-04a due to a complaint received regarding a rotor of the PP-04a pump that stopped during a therapeutic plasma exchange treatment causing interruption of flow of both the plasma discard line and replacement fruit line. | Process control | APHERESIS TECHNOLOGIES, INC. sent a letter dated July 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the device for repair. For questions customers should call 727-787-5616. |  | 13 | 13 | Natonwide Distribution | Terminated | December 16, 2013 | 525 | 13 | Not_Computer
Z-2148-2012 | 62683 | No | FWH | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Chromovision High Definition (HD) Wireless Camera, Pt # CC-9715900. |  |  | July 10, 2012 | 2012 | Berchtold Corp. | Defective camera: There is a possibility that the connection of the wires to the connector may become loose, causing image distortion or loss. | Nonconforming Material/Component | BERCHTOLD Corporation sent a letter dated July 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. BERCHTOLD will replace the customer's cameras at no cost to the customer. Customers with questions should call 843-569-6100, ext 322.For questions regarding this recall call 843-569-6100. |  | 17 | 17 cameras | Nationwide Distribution including FL, MI, MS, and NH. | Terminated | June 06, 2013 | 331 | 17 | Not_Computer
Z-2146-2012 | 62561 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | djo surgical Screw 6.5 mm, Sz 25mm, low profile. |  |  | July 16, 2012 | 2012 | Encore Medical, Lp | One lot of the 010-55-030, 6.5 mm Cancellous Bone Screw was recalled because a product complaint identified a mislabeling of the device. Although the package label indicates Size 30mm, the screw dimension is actually 25mm. | Packaging | DJO Global sent an Urgent Field Safety Notice dated July 16, 2012, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the notice to all thoe who need to be aware within their organization or to any other organization where the potentially affected product was transferred, contact customers to place a replacement order and receive an RMA number, and return all affected devices using the RMA number. For questions customers should call 512-834-6255.For questions regarding this recall call 512-832-9500. |  | 18 | 18 devices | Nationwide Distribution including AR, CA, FL, ID, OK, PA, TX, and UT. | Terminated | December 04, 2012 | 141 | 18 | Not_Computer
Z-2150-2012 | 62556 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Finapres Medical Systems Portapres with NiMh battery pack |  |  | July 06, 2012 | 2012 | FMS, Finapres Medical Systems BV | Battery may become hot and leak | Process control | Finapres Medical Systems BV sent a Field Safety Notice letter dated July 6, 2012 to notified US accounts of the recall. The letter identified the affected product, description of the problem and actions to be taken. The letter states that the recalling firm recommends not using the affected battery packs. A replacement battery was included in shipment. Customers were requested to return recalled battery. For questions call +31 (0)20 609 0974 or email info@finapres.com. |  | 3 | 3 units US | US Nationwide Distribution - including the states of: OH and NY | Terminated | September 12, 2012 | 68 | 3 | Not_Computer
Z-2147-2012 | 62395 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CrossFT Punch, Product Number PFT-45M. Manufacturer: CONMED LINVATEC. The intended use is to create a hole to facilitate insertion of the CrossFT suture anchor into bone. |  |  | August 12, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled lot number #156920, PFT-45M, CrossFT Punch because there is a possibility that the device was etched incorrectly listing the item number as PFT-00M and includes incorrect depth lines for insertion. | Use error | The firm sent an "URGENT: Medical Device Recall Notification" letter dated August 12, 2011, to its customers after initial phone notifications. The letter describes the product, problem and actions to be taken. The customers were instructed to: 1) Immediately check their facility's inventory for the product number and lot number listed in the letter. If the product is at their facility, they were instructed to discontinue the use of the product. 2) To return the products to ConMed Linvatec for credit and return the products and Reply Form. 3) If the product was transferred to another facility, complete the Reply Form documenting the transfer of the product, and return i by mail or fax. For questions, contact 727-399-5205. |  | 2 | 2 | Nationwide Distribution, including the states of Colorado and Pennsylvania. | Terminated | November 21, 2013 | 832 | 2 | Not_Computer
Z-2152-2012 | 62535 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 731 EMV+ Uni-Vent, Impact Instrumentation Inc., 27 Fairfield Place, West Caldwell NJ 07006. Indicated for use in the management of acute or chronic respiratory failure or during resuscitation by providing continuous positive -pressure ventilation. |  |  | March 19, 2012 | 2012 | Impact Instrumentation, Inc. | Ten units of Model EMV+ Devices were shipped out in error before the verification and validation was completed for a new integrated circuit (IC) on the CPU PCB. | Process control | Impact Instrumentation Inc. notified their sole customer via email on 3/15/12 and then emailed an "Urgent Device Product Recall" letter and tracking form dated 3/19/2012. The letter described the reason for the recall and provided recommended actions and the contact number for Customer Service. |  | 10 | Ten Units | US Distribution to the state of Illinois. | Terminated | August 15, 2012 | 149 | 10 | Not_Computer
Z-2155-2012 | 62648 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LUXOR RETRACTOR BASE; Catalog Number: 48250030; Manufactured by: STRYKER Spine SAS; Z.I MARTICOT-33610-CESTAS-FRANCE; Distributed in the USA by: Stryker Spine. 2 Pearl Ct. Allendale, NJ 07401. |  |  | April 23, 2012 | 2012 | Stryker Spine | Stryker received three reports involving blades of the Luxor Retractor becoming disassembled from the base causing the retractor to collapse insitu. | Nonconforming Material/Component | Stryker sent Urgent Product Recall letters/return response forms dated April 23, 2012 via FedEx to Stryker branches/agencies and hospitals where product might have been distributed. The letter identified the affected product, discussed the issue, potential hazards, risk mitigations, and the return process. Customers are to examine their inventory and reconcile any of the affected product with the information provided. In addition, the Customer Response Form should be completed and faxed back to Regulatory Compliance at 201-760-8370 within 5 days of receipt of the letter. The affected product should also be returned following the instructions provided. Questions should be directed to Dan Roberts, Regulatory Affairs Compliance Specialist at 201-760-8298. |  | 30 | 30 devices | Worldwide Distribution -- US, Switzerland, Netherlands, Brazil, Hong Kong, Germany, and Taiwan. | Terminated | November 13, 2013 | 569 | 30 | Not_Computer
Z-2154-2012 | 62595 | No | GIM | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | 3mL 9NC Coagulation Sodium Citrate 3.2%, Item #454334, |  |  | July 02, 2012 | 2012 | Greiner Bio-One North America, Inc. | Part of the lot received a lower concentration of citrate solution which may cause falsely lower values for common coagulation tests. | Nonconforming Material/Component | Distributors were notified by letter on 07/02/2012 and customers were sent a letter dated 07/05/2012. The letters identified the affected product and the reason for the recall. Distributors were asked to stop distributing the affected product and to complete the attached Product Disposition form. The form along with the affected product was to be returned to GBO. The firm will provide replacement stock or issue a credit once product return and receipt of the Product Disposition form has been confirmed. Customers were also asked to complete the Product Disposition Site Confirmation form and fax it back to Greiner Bio-One North America, Inc. at the number provided. The letter states that the firm will replace any affected product once the form is received. Questions should be directed to 800-515-8112. |  | 838.9 | 838.9 cases (1,006,700 pieces) | Nationwide Distribution | Terminated | June 11, 2013 | 344 | 838.9 | Not_Computer
Z-2158-2012 | 62375 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | ***REF IFNT611***Full OSSEOTITE Tapered Certain Implant 6 x 11.5mm***Sterile by radiation***Rx Only***BIOMET 3i Dental Iberica S.L. WTC Almeda Park, Ed, 1, Planta 1 Pl. de la Pau s/n08940 Cornella de Llobregat (Barcelona) Spain. Lot #2011110798. |  |  | March 23, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i is recalling their product Full Osseotite Tapered Certain Implant IFNT611. The depth of the implant's internal hex is too shallow. The condition, if present, may prevent the driver; cover screw or abutment from fully engaging the implant. | Process control | BIOMET 3i sent an Urgent Medical Device Recall letter on March 19, 2012, to all affected customers via email or fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers are to return any unopened and unused product for replacement at no cost. In addition, customers should contact Customer Service at 1-800-342-5454 for assistance if they encounter issues with the mating component, or if they have any questions or concerns. Customers are to respond to the notification regardless of whether they have product to return. |  | 240 | 240 | Worldwide Distribution -- United States, UK, France, Canada, Mexico, Poland, Lebanon and Hong Kong. | Terminated | December 16, 2013 | 633 | 240 | Not_Computer
Z-2161-2012 | 62575 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Straumann Narrow CrossFit (NC) Screw for CARES Abutment ZrO2. Endosseous dental implant |  |  | July 11, 2012 | 2012 | Straumann USA, LLC | Incorrect Screw (RC) included with NC CARES Abutment Set was not a Narrow CrossFit (NC) Screw for CARES NC Abutment ZrO2 | Packaging process control | Straumann notiifed accounts by telephone on July 10, 2012 and follow-up Field Safety Correction Action letter dated July 11, 2012, via United Parcel Service. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to discontinue use of the affected product and follow the instruction provided until replacement is received. For questions call 978-747-2514. |  | 14 | 14 Distributed in US | US Nationwide Distribution including the countries of FL, IN, LA, NY and TX. | Terminated | September 05, 2012 | 56 | 14 | Not_Computer
Z-2139-2012 | 62558 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | AUTOMIX High Speed Compounder System, product 2M8077; automated nutrition compounders; |  |  | July 13, 2012 | 2012 | Baxter Healthcare Corp. | After fluid ingress into the keypad of the AUTOMIX Compounder's (AUTOMIX) control module, the AUTOMIX may generate an incorrect device response to an AUTOMIX operator's key press. Also, the AUTOMIX system may experience an intermittent electrical failure, in which unexpected pump motor operations will occur. Also, the AUTOMIX system may intermittently experience an electrical failure where thecompounder may stop compounding without an alarm and not be detected by the user. | Device Design | Baxter Healthcare Corp. sent an Urgent Product Recall letters dated July 13, 2012, via first class mail to all affected customers, informing personnel that Baxter was executing a voluntary controlled product withdrawal of the Automix system from the U.S. market by June 1, 2013. Until their units can be replaced, customers were advised to exercise care when entering data, and verify the accuracy of the keypad response; to inspect the umbilical cable of the Automix for any visible cuts, nicks deformation, or physical damage to the cable, discontinuing use of the device if such damage is found; and to return the device to Baxter if the device stops compounding without an alarm and without a visible green complete light, discarding the solution bag. The customers were requested to convey this information to all users of the Automix compounders, and a wall chart was provided to hang over the device to remind users of the mitigations to take. The customers were also requested to complete the attached customer reply form listing the model and serial number of each of the Automix compounders at the facility and confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to their local Baxter representative, and technical questions were directed to Baxter's Global Technical Services at 1-800-626-2667. |  | 270 | 270 units | Worldwide Distribution - USA - Nationwide including Puerto Rico, and internationally to Brazil, Canada, Colombia, Costa Rica, Czech Republic, El Salvador, Egypt, Germany, Guatemala, Honduras, Luxembourg, Mexico, Panama, Saudi Arabia, Singapore, Switzerland, Taiwan, Trinidad, Turkey, United Arab Emirates and United Kingdom. | Terminated | August 11, 2017 | 1855 | 5166 | Not_Computer
Z-2163-2012 | 62503 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The ABL software version 6.12 for the ABL800 and ABL700 analyzer. The analyzers are intended for the in Vitro testing of: |  |  | July 05, 2012 | 2012 | Radiometer America Inc | A customer has reported an error on an ABL8xx configured to a Hospital Information System using the data transfer protocol HL7 version 2.5. If the instrument transmits results to a HIS system and the connection is broken the results are queued for later transmission. If the ABL and HIS system are configured with HL7 2.5, any results in the queue will be mixed up when the instrument transmits t | Software design | The firm, Radiometer, sent two "Priority Level: Urgent" letters one dated July 9, 2012 and the other dated July 24, 2012 to its affected customers via FedEx with tracking. The letter identifies the product, problem, and actions to be taken by the customers. The customers were instructed to complete and return the attached Recall Response Fax Form via fax to 800-736-0601. The letter informed the customers that a local service representative will be contacting them shortly to install a software update that will prevent this from occuring (in the event that you utilize HL7 version 2.5 or intend to at some future date.) If you have any questions regarding this letter or the Recall, please contact Radiometer America Technical Support at 1-800-736-0600 opt. 4. |  | 6861 | 6854 ABL 7xx units | Worldwide distribution: USA (nationwide) including the states of AK, AL, AR, AZ, BM, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Canada, Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cuba, Cyprus, Czech republic, Denmark, Ecuador, Egypt, Eritrea, Estonia, Finland, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Malta, Mauritius, Mexico, Moldova, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Peru, Philippines, Poland, Puerto Rico, Qatar, Republic of Korea, Romania, Russia, Saudi Arabia, Switzerland, Singapore, Slovenia, South Africa, Spain, Sweden Syria, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Yugoslavia, and Zimbabwe. | Terminated | August 20, 2013 | 411 | 20315 | Computer
Z-2157-2012 | 62536 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | I-STAT PT/INR cartridge Abbott Point of Care Inc. Abbott Park, IL 60064 USA |  |  | March 01, 2012 | 2012 | Abbott Point Of Care Inc. | The i-STAT PT/INR cartridges have the potential to exhibit incorrectly elevated results. | Process control | Abbott Point of Care Inc, sent a URGENT RECALL NOTICE dated March 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were provided in the letter with instructions in the letter on how to handle the affected product. Please complete and return the enclosed Business Reply card to acknowledge receipt of this communication and installation of the interium software CLEW and also 1) Update all handhelds 2) download appropriate assignment sheets. For further questions please call (609) 454-9000. |  | 3106704 | 2,950,296 cartridges US; 156,408 cartridges OUS | Worldwide Distribution -- (USA) nationwide | Terminated | February 08, 2013 | 344 | 3106704 | Computer
Z-2159-2012 | 62500 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas KRAS Mutation Test for In Vitro Diagnostic Use |  |  | February 24, 2012 | 2012 | Roche Molecular Systems, Inc. | During the real-time stability testing for the cobas KRAS Mutation Test on the commercial lot PO6778, it was identified that the kit was just within specifications for the codon 61 peak height ratio, and that the kit has trended towards the upper specification limit at the 6 month time point but did passed the 6 month stability testing. The kit is labeled for 9 months shelf life stability. The d | Nonconforming Material/Component | Roche issued a Safety Board Notice (Prooduct Support Notice 2012-01) and Product Bulletin 2012/03 dated February 24, 2012 to all affected customers. The notice idenfied the affected product, description of situation and actions to be taken. Customers were instructed to discontinue use of lot P06778 by March 4, 2012 and to quarantine and discard any remaining kits from lot P06778. The notice stated that a new lot will be made available as a replacement. |  | 3432 | 264 kits to 13 EU countries; 2 kits to Canada and 33 kits to 11 countries in the Rest of the World | Worldwide Distribution - including the countries of Austria, Belgium, Czech Republic, Denmark, Germany, Italy, Luxemburg, Netherlands, Poland, Slovakia, Slovenia, Spain, United Kingdom, Canada, Chile, Greenland, Hong Kong, Israel, Malaysia, New Zealand, Philippines, S. Africa, Switzerland, Taiwan and Turkey. | Terminated | July 16, 2013 | 508 | 3432 | Not_Computer
Z-2153-2012 | 62562 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Varian ARIA Oncology Information System, Import Export |  |  | July 02, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | A software anomaly has been identified with the Import Export Application version 11.0 whereMLC leaf positions for plans using BrainLab m3 and mMLC are incorrectly imported. | Software design | Varian Medical Systems sent a " Urgent Medical Device Correction Urgent Field Safety Notice" dated July 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed until a correction is deployed all treatment plans using the BrainLab m3 mMLC must be imported using the Import Wizard only.Varian Medical Systems is developing a technical correction for this issue. You will be contacted by a Customer Service Representative when this correction is available to schedule its installation on your system.For further questions please call 1.888.827.4265. |  | 2 | 2 Units. | Nationwide Distribution including the states of New York and Ohio. | Terminated | July 01, 2013 | 364 | 2 | Computer
Z-2160-2012 | 62645 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | The finger cots are made out of natural rubber. They are disposable products used to cover a fingertip bandage when working with liquids and are intended to keep the fingertip bandage dry. Cots are packed in clear plastic boxes, with 60 cots in a box |  |  | February 08, 2012 | 2012 | Afassco Inc | The firm is initiating a field correction because the product(s) were missing the Latex Warning Label. Afassco Inc. is implementing a voluntary field correction for MULTI-COLORED FINGER COTS because product lacked the warning label:"CAUTION: This product contains Natural Rubber Latex which may cause allergic reactions." | Packaging process control | Afassco sent a Field Correction Notice to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Labels have been included with the notices that the firm will be sending to their distributors. The distributors were asked to examine their inventory and affix the caution label to the boxes of affected product. If additional warning labels are needed, they may contact Afassco (800-441-6774) for additional labels.If they have further distributed these products, they should notify consignees of the field correction. Their notification to their customers may be enhanced by including a copy of the notice sent to them. The recall should be carried out to the user level. For any questions customers may call 1-880-441-6774.For questions regarding this recall call 775-783-3555. |  | 1956 | 1956 Boxes | Nationwide Distribution including CA, FL. GA. HI, IL, IN, MI, NH, NV, TN, TX, and WV. | Terminated | January 14, 2014 | 706 | 1956 | Not_Computer
Z-2156-2012 | 62551 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips MultiDiagnost Eleva with Flat Detector (MD Eleva FD)Model: 708034 |  |  | July 11, 2012 | 2012 | Philips Healthcare Inc. | When doing a 3 Dimensional Rotational Angiography (3DRA) the 3D reconstruction does not always start, even when all images have been acquired. | Software design | Philips issued an Urgent Field Safety Notice dated July 2, 2012 to all affected customers. The letter identified the affected product, problem description, actions to be taken by the customer/user in order to prevent risks for patients or users and actions planned by Philips to correct the problem. For question or support concerning this issue contact your local Philips representative. |  | 41 | 41 units | Worldwide Distribution - US (nationwide) including the states of CA, CT, MI, MN, MN, NC, NY, NY, VA and the countries of Belgium, Columbia, Czech Republic and Denmark.FranceGermanyItalyLatviaLuxembourgNetherlandsRussian Fed.SpainSwedenSwitzerlandThailandUnited Kingdom | Terminated | March 26, 2013 | 258 | 41 | Not_Computer
Z-2136-2012 | 62414 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Radiological Image Processing System |  |  | April 02, 2007 | 2007 | GE OEC Medical Systems, Inc | The FluoroTrak Spinal Navigation Application on the OEC 9900 EliteNAV could result in an incorrect position of the navigated instrument(s) versus the displayed reference image. | Software design | GE Healthcare sent an Urgent Recall Notice dated April 2, 2007, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were asked to call 800-874-7378. For questions regarding this recall call 801-536-4952, |  | 17 | 17 | Worldwide Distribution - USA including MO, IL, LA, FL, NH, TX, IN, PA, NJ and the countries of France and China | Terminated | September 28, 2012 | 2006 | 17 | Not_Computer
Z-2181-2012 | 62596 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Kimberly Clark Patient Warming System - Model 1000 Control Unit. |  |  | April 27, 2012 | 2012 | Kimberly-Clark Corporation | A defect of the "fuse holder" can lead to a Control Unit failure that may cause the device to lose its ability to warm and maintain set temperature and may also result in excessive heat with the fuse assembly that could lead to the generation of smoke. | Device Design | Kimberly-Clark sent an "URGENT: FIELD CORRECTIVE ACTION NOTICE" dated April 27, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Corrective Action Response Form was attached for customers to complete and return. Contact the Field Corrective Action Coordinator at 770-587-7400 for questions regarding this notice. |  | 129 | 129 units | Nationwide Distribution-including the states of AZ, CA, CO, DE, FL, GA, IA, ID, IN, IL, KS, KY, LA, MA, MD, ME, MO, MS, NE, NM, OH, OR, PA, RI, SC, TN, TX, VA, WI and WV. | Terminated | September 05, 2012 | 131 | 129 | Not_Computer
Z-2172-2012 | 62448 | No | NSX |  |  | General Hospital | Contact ODE | EasyLink(TM) Informatics System, Part Number 1000034805. EasyLink(TM) Informatics System is classified as a data calculator and processing module for use with In-Vitro diagnostics devices. |  |  | June 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | There is a potential for the EasyLink system to display and transmit to the Laboratory Information System networks an incorrect result following the EasyLink systems application of the configured result number format. The problem is encountered when the result number format configured on the Sysmex CA-7000 or CA-1500 coagulation analyzer does not match the EasyLink system settings on the following | Software design | The firm initiated this recall on June 7, 2012 by issuing an "Urgent Field Safety Notice" to all customers who are known to have the EasyLink system and who also own and operate Sysmex CA-1500 and CA-7000 instruments. The Notice described the problem and provided recommended actions. Customers can contact 800-441-9250 for questions. |  | 237 | 237 | Worldwide distribution, including the countries of Australia, Austria, Belgium, Canada, Czech Republic, France, Germany, Italy, Netherlands, New Zealand, Slovakia, Spain, Switzerland and the United States, with nationwide distribution. Center Recall Depth is Retail. | Terminated | August 26, 2014 | 810 | 237 | Not_Computer
Z-2179-2012 | 62747 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medstream 81" (206 cm) 20 Drop Standard bore Universal Administration Set with Y-Site, Sterile |  |  | August 03, 2012 | 2012 | Medical Specialties Distributors, Inc. | Devices leaking at the Y-site | Process control | Medical Specialties Distributors LLC sent an Urgent Medical Recall letter dated August 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and remove all affected product from use immediately. Further distribution of the affected product should be discontinued immediately. Customers were instructed to contact their customers and inform them of the recall; complete and return the acknowledgement form and fax to 781-344-8320 to arrange for a return and to contact the Customer Service Department at (800) 967-6400 with questions. For questions regarding this recall call 781-344-6000. |  | 472 | 472 cases (50/case) | Nationwide Distribution | Terminated | January 23, 2015 | 903 | 472 | Not_Computer
Z-2165-2012 | 62525 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew 5.5 mm TWINFIX Ultra Ti Suture Anchor with two #2 ULTRABRAID sutures (blue, blue-COBRAID) with Needles |  |  | July 03, 2012 | 2012 | Smith &amp; Nephew, Inc. Endoscopy Division | Distal part of the anchor may break on insertion into bone during surgery | Device Design | Smith &amp; Nephew sent an Urgent Product Recall 1st Notification letters dated June 2012 via Federal Express to all affected customers. The letter identified the affected product, potential risk and actions to be taken. Customers were instructed to inspect inventory and quarantine them immediately, complete and return the attached form indicating the Lot# and quantities that need to be returned and contact Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) or email to endo.andreturns@smith-nephew.com. for instructions for returning affected products and receiving a credit. For questions call 1-508-261-3731. |  | 913 | 913 units | Worldwide Distribution - US Nationwide and the countries of: Argentina, AUSTRALIA, AUSTRIA, Belgium, CANADA,CHILECOLOMBIA, Denmark, DUBAI, FINLAND, FRANCEGDC BAAR Distribution,GERMANY, GREECE, INDIAITALY, KOREA, MALAYSIA, MEXICO, NETHERLAND, NORWAYPORTUGAL, PUERTO RICO, SHANGHAI, SOUTH AFRICASPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | April 18, 2016 | 1385 | 12189 | Not_Computer
Z-2176-2012 | 62415 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Reliance AL Implant Blocker (Various sizes) |  |  | March 21, 2011 | 2011 | Stryker Spine | Stryker Spine has received six reports regarding the breakage of one or more prongs at the extremity of the instrument. | Component design/selection | Stryker sent an Urgent Product Recall letters/return response forms dated April 5, 2012, via FedEX to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital location to identify the product. Reconcile any SLIDE Implant Blockers with catalog and lot numbers referenced above. Retrieve and return the affected product using the enclosed pre-paid mailing label to Stryker Spine 59 Route 17 South, Allendale, New Jersey, 07401. For questions customers were instructed to call 201-760-8298. |  | 109 | 109 Lots | Nationwide Distribution including NC, AR, TN, MA, IL, GA, NY, LA, MN, MO, OH, KS, NM, TX, KS,and IN. | Terminated | March 21, 2013 | 731 | 109 | Not_Computer
Z-2175-2012 | 62609 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Vancomycin Flex(R) Reagent Cartridge (VANC, Catalog # K4086). |  |  | July 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has received complaints of customers experiencing slopes &gt;1.05 when calibrating Flex(R) reagent cartridge lot 12095AD. This results in a slope failure and prevents the VANC assay from being run on the Dimension Vista(R) System. | Nonconforming Material/Component | The firm initiated their recall of this product on July 17, 2012 by sending an Urgent Field Safety Notification via FedEx next day delivery to all consignees. The Notice informed consignees of the high calibration slopes with the affected device. Affected consignees were requested to immediately discontinue use and discard any remaining Dimension Vista(R) VANC Flex(R) reagent cartridges, lot 12095AD. Affected customers were provided with replacement cartridges. Customers were to also complete the attached form and fax it back to the number provided. The notification letter should be forwarded to anyone, whom customers may have further distributed the affected lot. Questions and concerns should be directed to Siemens Technical Solutions Center at 800-441-9250. |  | 1894 | 1,894 cartons | Worldwide Distribution -- US, including Washington, DC and the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WV, and the countries of Canada, New Zealand, Germany, and Puerto Rico. | Terminated | January 14, 2014 | 546 | 1894 | Not_Computer
Z-2182-2012 | 62611 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) Reagent Management System (RMS) - Power Cord. |  |  | July 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed that partial insertion of the power cord on the Dimension Reagent Management System (RMS) power input module can cause the plug to overheat and potentially create a fire hazard. RMS without a power cord retention clip may be affected. | Device Design | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated June 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an attached Field Correction Effectiveness Check Form and fax it to (302) 631-8467. Contact Siemens Technical Solutions Center at 800-441-9250 for questions regarding this notice. |  | 2554 | 2,554 | Distributed in the US including Washington, DC, Puerto Rico and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WI, WV, and WY and the Virgin and Pacific Islands. | Terminated | April 03, 2014 | 625 | 2554 | Not_Computer
Z-2180-2012 | 62341 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Angiographic catheter --- Soft-Vu (JB-1) Non-braided. --- Size: 5F; Length: 100cm; designed to accept guidewire: .035 in. --- Catalog No. [REF]: 10734201. Packaged in boxes of 5 catheters each, with manufacturer listed as AngioDynamics, Inc., Queensbury, NY USA. |  |  | June 15, 2012 | 2012 | Angiodynamics Worldwide Headquarters | AngioDynamics is recalling one lot of Soft-Vu Angiographic catheters because one of the catheter pouches within the outer box may be mislabeled. | Labeling mix-ups | The firm, AngioDynamics, Inc., sent an "Urgent Medical Device Recall - Immediate Action Required letter dated June 15, 2012 to all consignees/customers via FedEx. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to carefully read this recall notification in its entirety; IMMEDIATELY DISCONTINUE USE OF AND SEGREGATE RECALLED PRODUCT in a secure location for return to AngioDynamics, Inc.; immediately remove recalled product from their inventory; immediately forward a copy of this recall notification to all sites to which they have distributed affected product; and promptly complete, sign and return the Reply Verification Tracking Form (even if you do not have any product to return) via fax to Soft-Vu Angiographic Catheter Recall Coordinator at 1-518-798-1360 or email rdenino@agniodynamics.com. If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00 a.m. and 7:00 p.m. (Monday - Friday: Eastern Standard Time) to obtain a replacement. The consignees/customers were also instructed to package and return the recalled product, write the RMA number on the box and send to AngioDynamics, Inc., 603 Queensbury Avenue, Queensbury, NY 12804; Attn: Soft-Vu Angiographic Catheter Recall Coordinator.If you have any questions about this recall action, contact your local Sales Representative or AngioDynamics, Inc. Customer Service Manager at 800-772-6446 or by email at customerservice@angiodynamics.com |  | 245 | 7 boxes/35 units | Worldwide distribution: USA (nationwide) including states of: CA, KS, MO, and NC and countries of: Czech Republic and Switzerland. | Terminated | February 10, 2016 | 1335 | 245 | Not_Computer
Z-2171-2012 | 62533 | No | MYT | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-Stat CK-MP cartridges |  |  | February 01, 2012 | 2012 | Abbott Point Of Care Inc. | The frequency of suppressed results for i-Stat cTnl, BNP and CK-MP cartridges is affected by atmospheric pressure. | Device Design | Abbott Point of Care (APOC) sent an URGENT RECALL NOTICE dated February 2012 to all affected customers via e-mail and also contacted by phone. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that if their facility was performing testing at more than approximately 7500 feet above sea level, they should have an alternate test method available should they experience an issue. If they have forwarded any cartridges to another facility, they should provide a copy of the Urgent Recall Notice to them. For questions customers were instructed to call 800-366-8020, Option 1 or contact their Abbott Point of Care representative.For questions regarding this recall call 609-454-9000. |  | 400874 | 368,149 cartridges US; 32,725 OUS | Worldwide Distribution - USA including AZ and CO and internationally to Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Chile, China, Colombia, Costa Rica, Dominican Republic, El Salvador, Finland, France, Georgia, Germany, Greece, Honduras, Hong Kong, India, Ireland, Italy, Korea, Kuwait, Mauritius, Mexico, Netherland (Antilles), Nigeria, Norway, Oman, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, R.O.C., Thailand, UAE, United Kingdom, and Yemen. | Terminated | February 08, 2013 | 373 | 9289349 | Not_Computer
Z-2177-2012 | 62444 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 1 base 100/120V |  |  | July 26, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems sent a Safety Advisory (AA-2010-007-C) on December 2, 2010 (Recall# Z-0882/0883-2011 RES 57416) regarding malfunctions of the Occluder and Air Bubble Detection (ABD) Modules to a confined number of Terumo Advanced Perfusion System 1 users. The investigation of the malfunction found a faulty component on the modules' application board caused the failures. Terumo CVS re | Component design/selection | The firm, Terumo CVS, sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 26, 2012 to its customers. The letter described the product, problem and actions to be taken. To customers were instructed to review this Medical Device Recall notice and addendum; assure that all users are aware of this notice and addendum; place the addendum in the front of Section 7: Modules (p 7.1) and confirm receipt of this communication by faxing the attached Customer Response Form to the fax number/email address indicated on form. Terumo CVS is issuing a Safety Advisory to alert all Terumo System 1 users of the potential for failure and providing an addendum for the Operator Manual with instruction on how to respond to the failure.If you have any question or concerns contact Terumo CVS Customer Service at 1-800-521-2818. Customer Service hours are Monday-Friday, 8AM - 6PM ET. |  | 992 | 992 units | Worldwide distribution: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Egypt, Germany, Guatemala, Honduras, Hong Kong, India, Indonesia, Japan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | November 16, 2012 | 113 | 1641 | Not_Computer
Z-2173-2012 | 62403 | No | NSX |  |  | General Hospital | Contact ODE | Siemens syngo(R) Lab Data Manager. |  |  | June 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | The recalling firm has discovered that when using a Sysmex(R) CA-7000 or CA-1500 coagulation analyzer with the syngo Lab Data Manager (LDM), there is a potential for the syngo LDM system, to display and transmit to the LIS an incorrect result following the syngo LDM's application of the configured result number format. The problem is encountered when the result number format configured on the Sys | Software design (manufacturing process) | Siemens Healtcare Diagnostics sent a Urgent Field Safety Notice dated June 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised to discuss the contents of this letter with your Laboratory Director regarding the need to review any previous test results.If you have any technical questions or concerns, please contact the Seimens Technical Solutions Center at 1-800-441-9250, for further assistance. |  | 15 | 15 | USA (nationwide) including the states of AZ, CA, MS, MO, NJ, NC, OH, OR, PA, and TX. | Terminated | April 01, 2014 | 663 | 15 | Not_Computer
Z-2187-2012 | 62499 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The NCB Polyaxial Locking Plate System NCB screws 26mm: 02.03155.026, non-sterile |  |  | July 13, 2012 | 2012 | Zimmer, Inc. | Zimmer investigation following a complaint found one lot of NCB Polyaxial Locking Plate System NCB screws, 26mm, distributed in the United States contained incorrect 32 mm screws. | Labeling mix-ups | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL" letter dated July 13, 2012, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate the implicated lot in their inventory and user accounts; return all packaged units of this lot; and provide Zimmer with the name, complete mailing address, and phone number of each user facility and surgeon that have used the product via E-mail to dawn.kindle@zimmer.com. Zimmer will send a follow-up Urgent Recall notification to these sub-accounts with instructions to verify the screw length of any screws already in the implant rack inventory prior to use in surgery. All implicated product is to be returned. The customers were also instructed to complete and return the Inventory Return Certification form to Zimmer via fax at +1-574-372-4265. If you have any questions about this correction and/or product removal, contact Zimmer at +1-877-946-2761. |  | 31 | 24 to US Distributor with 7 further distributed | Nationwide distribution: USA including states of: AZ, IL, IN, ME and NJ. | Terminated | September 24, 2013 | 438 | 31 | Not_Computer
Z-2186-2012 | 62598 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TAUT, Cholangiogram Catheter Intraducer, Peritoneal Catheter, Teleflex Medical, Research Triangle Park, NC. |  |  | July 20, 2012 | 2012 | Teleflex Medical | Product is missing the label on the internal packaging. | Packaging process control | Consignees were notified by letter on 07/20/2012 |  | 611 | 611 units | AK, AL, AZ, CA, CO, FL, GA, IA, IL, IN, KS, LA, MA, MI, MO, MS, NC, NY, OH, PA, SD, TX, UT, VA, WI, WV. | Terminated | February 12, 2015 | 937 | 611 | Not_Computer
Z-2184-2012 | 62646 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular (CV) Reporting |  |  | July 27, 2012 | 2012 | AGFA Corp. | Content entered into the "Conclusions" free text box on the Report Writer screen was not represented on the final printed report. | Software change control | AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE dated July 27, 2012, to all affected customers. The letter described the safety alert and mitigation. Customer were asked to complete the Acknowledgment form and fax back to 864-421-1664 that the information was received and understood.For questions customers were instructed tol call 401-604-2180.For questions regarding this recall call 864-421-1754. |  | 280 | 280 | Worldwide Distribution - USA (nationwide) and Canada | Terminated | January 13, 2015 | 900 | 280 | Not_Computer
Z-2174-2012 | 62559 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Reusable Cannula Low Pressure Cuffed Tracheostomy Tubes, model numbers (Worldwide) 8LPC, 8FEN; (Outside U.S.) 8FENJ, 8FENJ-S, 8LPC-S, 8FEN-S. |  |  | July 18, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is recalling certain cuffed Shiley tracheostomy tubes due to consumer complaints related to the connections and functions of the inner and outer cannula which were observed during mechanical ventilation of patients. | Device Design | Covidien sent an Urgent Medical Device Recall letter dated July 18, 2012 to all affected consignees via mail (FedEx) notification. The customer communication identified the affected product , problem and actions to be taken. Customers were instructed to examine their inventory, segregate product with the identified affected product codes and lot numbers and complete the enclosed verification form and fax to 925-463-4600. Customers were asked to return unused products from the affected product codes and lots. The firm also recommended the replacement of affected tubes that were already in use as soon as clinically appropriate. The letter stated to return the affected product for credit or question contact Technical Services group at 800-635-5267 option 3, option 1 and reference the Return Goods Authorization number. |  | 337227 | 337,227 units | Worldwide Distribution - US Nationwide | Terminated | January 22, 2015 | 918 | 337227 | Not_Computer
Z-2183-2012 | 61855 | [6 empty] | BeautyAngel Health Fitness Natural Collagen Stimulation Anti Aging |  |  | May 07, 2012 | 2012 | JK Products &amp; Services, Inc | The Beauty Angel was marketed without a 510k. | No Marketing Application | JK Products sent a Mandatory Field Correction Notice - Immediate Action Required letter dated February 10, 2012, to all affected customers. The letter identified the product the problem, and the action to be taken by the customer. The firm instructed customers that they would be replacing the operating instructions for the affected product with new ones that meet the requirements defined by the regulatory authorities. Customers were instructed not to make claims that indicate that using the affected product will result in medical or health benefits, or that directly or indirectly imply that the affected product is intended to cure, treat, diagnose or pevent an illness, or disease. Customers were asked to complete the new owner information request on the return post care enclosed and return to the firm. If product was further distributed those accounts should be contacted immediately and should be advised of this Mandatory Field Correction Notice. For questions customers should call 870-935-1130. For questions regarding this recall call 870-268-2702. |  | 598 | 598 units | Worldwide Distribution - USA (nationwide) and Canada | Terminated | January 07, 2015 | 975 | 598 | Not_Computer
Z-2188-2012 | 62278 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Swan-Ganz Thermodilution VIP Catheters Model 131F7J. |  |  | May 11, 2012 | 2012 | Edwards Lifesciences, LLC | The device is being recalled because the firm received multiple complaints related to passage of a J-tip guidewire when inserted through the distal lumen hub of the catheter. | Process control | Edwards Lifesciences sent a "PRODUCT RECALL" letter dated May 8, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Confirmation Form was attached for customers to complete and return via fax to 1-800-422-9329. Contact Customer Service at 1-800-424-3278 for questions regarding this notice. |  | 32145 | 32,145 units for all products in Recall Event | Worldwide Distribution-USA (nationwide) and the countries of Canada, Japan, EMEA (Europe, Middle East and Africa), Asia Pacific, (Hong Kong, Indonesia, Malaysia, Australia, and Taiwan, Thailand), Latin America, (Chile, Colombia, Panama, Mexico, and Honduras). | Terminated | June 17, 2014 | 767 | 32145 | Not_Computer
Z-2202-2012 | 62310 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-1 System Battery --- Made in U.S.A. for Carestream Health, Inc., Carestream Health, Inc., 150 Verona St, Rochester, NY 14608. --- The battery may be sold with the DRX-1 imaging device or as a stand-alone accessory. |  |  | June 12, 2012 | 2012 | Carestream Health, Inc. | Carestream DRX-1 System batteries manufactured before June 2012 due to two occurrences of lithium-ion battery overheating, which may cause burns. | Nonconforming Material/Component | Firm sent consignees "Urgent: Customer Notification and Field Corrective Action" letters on June 8, 2012. The letters described the issue and provided recommended actions. Service dealers will contact all affected consignees to evaluate the battery. U.S. customers with questions can contact Carestream Customer Care at 800-328-2910. Customers outside the United States should contact their local Carestream service number. |  | 9971 | 9971 units | Worldwide Distribution, including Nationwide (USA). | Open, Classified |  |  | 9971 | Not_Computer
Z-2200-2012 | 62588 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | 13.5 mm x 50 mm BrainPath Sheath |  |  | July 17, 2012 | 2012 | Nico Corp. | BrainPath sheaths were mislabeled. Sheaths which measure 50mm were labeledas 60mm. The length callout on the label does not match the actual length of the sheath within thepackaging resulting in an improper fit as further explained below. This mislabeling wasdiscovered on 29-June-2012 during a clinical case. | Labeling mix-ups | NICO Neuro and Spine sent an Urgent Recall letter to all affected customers on July 18, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the recall response form.Affected sheaths distributed within the US will be obtained and discarded by NICO. Affected sheaths shipped to Canada will be obtained and discarded by NICOs distributor, Galen Medical.For questions customers should call 888-632-7071. |  | 100 | 20 devices (4 boxes, 5 devices per box) | Worldwide Distribution - USA including MI and Canada | Terminated | November 07, 2012 | 113 | 175 | Not_Computer
Z-2203-2012 | 61614 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Advantage Workstations, Advantage Worksation version 4.2 and 4.3 with Volume Viewer 2 (version 6.4-6.8) provided with Advance Vessel Analysis (AVA) and Advantage Workstation 4.4 with Volume Viewer 2 (version 6.8 or before) or Volume Viewer 3 (version 7.2 -7.3) provided with Advanced Vessel Analysis (AVA), Model 2378698-2 |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare is informing you of incorrect measurements when using the Aorta analysis protocol or customized protocols. Until a new software version is installed they are asking you not to use the aorta analysis protocol or the customized protocols for an aneurysm.A second issue was also discovered internally. When loading PET data into the Volume Viewer, ifslices are missing in a PET serie | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 23, 2012 to consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the affected product details, safety issue, safety instructions and product correction. For more information on this topic, contact a GE Application Specialist or your GE Field Service representative.For questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative |  | 1055 | 1055 | Worldwide Distribution - US (nationwide) including the states of : AL, AZ, AR, CA, CO, DE, DC, FL, GA, GU, HI, IK, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO,MT, NE,NH,NJ, NM, NY, NC, ND, OH, OK, PA, PR, SC, SD, TN, TX, VA, WA, WV, WI and the countries of YUGOSLAVIA, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKEY, TUNISIA, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAÂżDE ARABIA, RUSSIAN FEDERATION, ROMANIA, REPUBLIC OF MALTA, PORTUGAL, POLAND, PHILIPPINES, NORWAY, NIGERIA, NEW ZEALAND, MEXICO, KOREA, KENYA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, HUNGRY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Terminated | December 11, 2012 | 201 | 1055 | Computer
Z-2199-2012 | 62589 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SenTec, V-Sign"2 Sensors |  |  | July 09, 2012 | 2012 | Sentec AG | SenTec is recalling a number of V-Sign"2 Sensors due to a problem regarding configuration parameter writing during manufacturing. | Device Design | SenTec sent a Field Safety Notice dated July 12, 2012, via email to all affected customers. It identified the product the problem and the action needed to be taken by the customer. Customers were instructed the affected Sensors in the field must be replaced.Please locate the affected Sensors and return them to SenTec for analysis immediately. SenTec will provide you with replacement Sensors within the next three months.This notice needs to be passed on to all those who need to be aware within your organization. Please transfer this notice to other organizations on which this action has an impact.Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. Further questions please call (262 ) 594-2114. |  | 8 | 8 | Worldwide Distribution -- Nationwide including the state of Missouri,. and the countries of Germany, Spain, France, Netherlands, Australia, Belgium, Thailand. | Terminated | October 11, 2012 | 94 | 8 | Not_Computer
Z-2185-2012 | 55121 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module model 8100. |  |  | June 29, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion identified a potential risk associated with the Alaris Pump module model 8100. Carefusion has received reports of customers experiencing motor stalls during infusion with Alaris Pump Module (model 8100) manufactured between August 2010 and July 2011. | Under Investigation by firm | CareFusion sent an Urgent Medical Device Recall Notification letter dated July 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Required action for users instructed high risk patients undergoing infusions at high rates, consider having additional devices as back up. Clinicians should weigh the risk/benefit to the patients before continuing to use the device. Customers with recall related questions were instructed to contact Carefusion support center at (888) 562-6018. Customers with adverse event report questions were instructed to contact customer advocacy at (800) 854-7128, option 1, option 1, option 3 or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system were instructed to contact Technical support at (888) 812-3229.For questions regarding this recall call 858-617-4000. |  | 102248 | 102,248 units | Worldwide Distribution - USA (nationwide) and the countries of: South Africa and Canada.Additional Countries: Netherlands. | Terminated | August 26, 2015 | 1153 | 102248 | Not_Computer
Z-2209-2012 | 62586 | No | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS AmpliPrep Sample Tube Input Barcode Clips, catalog number 03137040001. Input S-tube barcode clips hold sample input tubes in the COBAS AmpliPrep Instrument (CAP) sample racks and have a barcode to identify the sample. |  |  | February 01, 2012 | 2012 | Roche Molecular Systems, Inc. | Certain lots of COBAS AmpliPrep (CAP) S-tube Input (Box of 12 bags 12x24 S-tube and 12x24 S-clip) can cause the sample input tubes to stick in the barcode clips. As a consequence, the CAP instrument will lift both the sample input tube together with the barcode clip from the sample rack, resulting in an aborted run. No erroneous results would be produced if the issue occurs, but samples in proce | Nonconforming Material/Component | On 1/17/2012 an "Analyzer Bulletin" notification was sent to Roche affiliates and on 2/1/2012 an "Analyzer Bulletin" notification was sent to all affected US Consignees. The notification described the product issue and provided recommended actions. Customers were referred to Roche Technical Support at 800-526-1247 for questions. |  | 2132 | 2,132 boxes | Nationwide Distribution (USA). | Terminated | July 16, 2013 | 531 | 2132 | Not_Computer
Z-2211-2012 | 62612 | No | LFM | Digitoxin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products DGXN Slides --- [REF] 834 3386. Product is packaged in a box with multiple slides. The lot and expiration codes are found on the side of the package. |  |  | June 28, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics revised the VITROS Chemistry Products DGXN Slides Instructions for Use (IFU) and removed heparin plasma as a recommended specimen type for VITROS DGXN Slides | Process control | Ortho Clinical Diagnostics sent an Urgent Field Safety Notice dated June 28, 2012, via US Postal Service Priority Mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Immediately discontinue using heparin plasma samples and use only serum samples when processing VITROS DGXN Slides; Notify their laboratory and medical staff of this change and forward the information if the product have been distributed outside of your facility; Replace the pages of their VITROS Instructions for Use Manual with the enclosed Instructions for Use and Table of Contents; and Discontinue using any copies of the VITROSÂż DGXN Slide IFU (Version 7.0).--- On 28 June 2012 an Urgent Field Safety Notice was sent via US Postal Service Priority Mail to all affected Distributors. They were advised of the issue and instructed to discontinue distributing and discard their remaining inventory of VITROS DGXN Slides. --- Foreign affiliates were informed by email on 28 June 2012 of the issue and instructed to notify their consignees of the issue and actions. The letters included a Confirmation of Receipt form to be completed and return. For questions contact Customer Technical Servives at 1-800-421-3311 |  | 13338 | 13338 sale units | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 13338 | Not_Computer
Z-2061-2012 | 61747 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 1 Silver Rover, Part 0700-003-000 Stryker Instruments Neptune Rover |  |  | June 08, 2012 | 2012 | Stryker Instruments Div. of Stryker Corporation | Stryker has received two reports of serious injury as a result of tissue damage associated with the use of the Neptune 2, including an event in which one customer connected the Neptune 2 System to a passive chest drainage tube post operatively, resulting in a fatality. The current IFU did not specifically warn against connecting all Neptune Waste Management devices, which is a high vacuum/high flow device, to a passive drainage tube. June 12, 2012, customers were instructed to review the revised IFU, distribute to affected departments, and educate users of the Neptune on this warning. Warning Labels, which are pending distribution, are required on all Neptune devices. Further investigation found the Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) lacked FDA 510(k) clearance and therefore are not legally marketed without determination of their safety and effectiveness. FDA advises Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) not be used. However, customers who do not have an alternative device to use should weigh the risks and benefits associated with continued use of these devices. If customers choose to continue use of the Neptune Silver, Neptune 2 Ultra (120V) or Neptune 2 (230V), they must complete a Certificate of Medical Necessity and return it to Stryker by October 12, 2012. | Labeling False and Misleading | On June 5, 2012, Stryker notified customers that it was recalling the IFUs for the above products. The current IFU did not specifically warn against connecting the Neptune Rover, which is a high vacuum/high flow device, to a passive drainage tube. Customers were instructed to review the revised IFU, distribute to affected departments, and educate users of the Neptune on this warning. Customers must confirm with Stryker via business reply form that they have completed these actions. Customers who have the Neptune 1 Gold, Neptune 1 Gold International or Neptune 2 Bronze will receive a follow up mailing containing warning labels for the device and instructions detailing how to apply them. Customers may continue to use the Neptune 1 Gold, Neptune 1 Gold International, and the Neptune 1 Bronze without any further actions.On September 18, 2012, Stryker notified customers via letter delivered by FedEx overnight delivery that it is expanding the recall on the Neptune 1 Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) because FDA has also advised Stryker that these devices require, but do not currently have, 510(k) clearance. FDA is therefore unable to determine whether these devices are as safe and effective as their legally marketed predicate, the Neptune 1 (Gold) Waste Management System (510(k) K012992). As such, Stryker has ceased distribution of the Neptune Silver, Neptune 2 Ultra (120V) and Neptune 2 Ultra (230V) devices until FDA clears these devices.At this time, FDA does not consider the Neptune Silver, the Neptune 2 Ultra (120V) or the Neptune 2 Ultra (230V) to be legally marketed devices because their safety and effectiveness have not yet been determined. As such, FDA advises that the devices not be used. However, customers who do not have an alternative device to use should weigh the risks and benefits associated with continued use of these devices. If customers choose to continue use of the Neptune Silver, Neptune 2 Ultra (120V) or Nept |  | 2657 | 2657 | Worldwide Distribution -- USA (nationwide) and country of: Canada.. | Terminated | August 23, 2016 | 1537 | 20329 | Not_Computer
Z-2210-2012 | 62471 | No | NCP | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Getinge Group Pediatric Arterial Cannula |  |  | May 29, 2012 | 2012 | Maquet Inc. | Certain pediatric arterial cannulas may have limited flexibility due to an increased length of the "hardened" cannula tip. | Device Design | Maquet issued Urgent Device 2, Recall letters/Urgent Product Recall Forms dated May 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their current stock and identify any affected product and remove from their inventory. Customers were instructed to contact the customer service department at 800-777-4222 (option #2, followed by option #2). Customers were also instructed to completed the enclsoed Recall Form to acknowledge receipt of the notification and to document the current status of the devices shipped to their facility. For questions regarding this recall call 973-709-7660. |  | 142 | 20 devices US; 122 devices foreign | Worldwide Distribution - USA including MI, MN and Internationally | Terminated | April 03, 2014 | 674 | 142 | Not_Computer
Z-2208-2012 | 62568 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Philips Easy Diagnost Classic-Stationary x-ray system |  |  | July 11, 2012 | 2012 | Philips Healthcare Inc. | Inclination switch in the stand may have been mounted incorrectly, as a result, the switch could not stop the tilting movement of the table base due tol hardware or software errors. | Equipment maintenance | Philips issued an Urgent Field Safety Notice dated July 9, 2012, to notify users that the inclination swtich may be mounted incorrectly. A Field Safety Engineer will be dispatched to verify if a switch was installed properly, and in the event it was not installed properly they will re-install it properly. The Field Service Engineer will add a label as described in the spare part replacement manual, which is part of the FCO kit.. Customers were instructed review the information contained in the Field Safety Notice and to retain a copy of the notice with the equipment Instruction for Use. For further questions customers were instructed to call their Phillips representataive Technical Support Line at 1-800-722-9377. For questions regarding this recall call 978-687-1501 |  | 12 | 12 units | Worldwide Distribution - USA including IN, MA, MI, MO, NC, NY PA, SC, and WI and to the countries of Canada, HongKong, New Zealand. | Open, Classified |  |  | 12 | Computer
Z-2207-2012 | 62650 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Essenta DR, X-Ray System, Diagnostic, general-purpose |  |  | July 25, 2012 | 2012 | Philips Healthcare Inc. | Essenta DR C-Arm may fall | Equipment maintenance | Philips Medical sent an "URGENT FIELD SAFETY NOTICE" dated July 09, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers are asked to inspect the C-Arm unit and discontinue use if not parallel to the column, do not use the system and call for a Service Technician. A Field Change Order (FCO) is being released that requires an inspection of C-ARM connection and, if necessary, repair of systems. The bolts will be disconnected and thread inserts will be checked. If not the thread inserts are no longer secured, additional holes will be drilled to use in place of the original ones. For further information or support concerning this issue, contact your local Philips representative at 1-800-722-9377. |  | 7 | 7 units | Worldwide Distribution-USA including the states of CA, NC, and TN and the countries of China, Germany, and Hungary. | Terminated | March 01, 2013 | 219 | 7 | Not_Computer
Z-2215-2012 | 62540 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | 00-4309-029-01 - Trabecular Metal Reverse Shoulder Liner 36 mm 65ÂżÂż Neck Angle |  |  | July 18, 2012 | 2012 | Zimmer, Inc. | Lot specific Trabecular Metal (TM) Reverse Shoulder Liner Impactors are being recalled due to reports of alignment peg fractures. All lots of TM Reverse Shoulder Liner Impactors manufactured from 455 stainless steel are being retrieved, as all reported fractured instruments were manufactured with thismaterial. The TM Reverse Shoulder Liner Impactor units that were manufactured with 13-8stain | Device Design | Zimmmer sent an Urgent Medical Device Recall letter dated July 18, 2012, to all affected customers. The letter identified the product, the problem, and thea action to be taken by the customer. Customers were instructed to inspect all TM Reverse instrument sets for affected lot(s)(instruments. If affected product was found customers were instructed to immediately place an order to backfill the identified product with Customer Service. If an instrument was was further distributed customers were to provide a copy of the Notice to ensure their customers were aware of the removal. Customers were also instructed to provide a list of all hospitals that currently use the affected product. For questions customers were instructed to call 800-348-2759.For questions regarding this recall call 574-371-1689. |  | 212 | 212 (179 USA) | Worldwide Distribution - USA (nationwide) and internationally to Germany, Belgium, Denmark, Finland, France, Italy, Netherlands, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | April 17, 2014 | 638 | 724 | Not_Computer
Z-2217-2012 | 62766 | [6 empty] | DePuy Custom Implant Devices |  |  | July 31, 2012 | 2012 | DePuy Orthopaedics, Inc. | In connection with a Warning Letter received from the U.S. Food &amp; Drug Administration on December 9, 2011, DePuy Orthopaedics is taking steps to retrieve custom device implants that surgeons may have ordered from DePuy Orthopaedics, but not implanted. DePuy Orthopaedics received a Warning Letter from the FDA on December 9, 2011 regarding custom device implants. DePuy Orthopaedics has discontinue | PMA | DePuy sent a written communication letter dated July 31, 2012 to all consignees. The communication to the hospital customers who received the devices includes a "Surgeon Letter" with instructions to provide the surgeon letter to the Surgeon users. All DePuy Distributors will be notified via email. The sales representatives notified the hospital and surgeons either by mail or in person with a written communication. The letters identified the affected product, problem and actions to be taken. Customers are instructed to complete the attached Roponse form and return to fax number 574-371-4939 or email: kseppa@its.jnj.com. The sales representatives are expected to aid customers in returning of unused Custom Device implants, if needed. The devices can be returned through the normal DePuy Returns process, to attention of "Returns" and marked "Custom Returns" on the outside of the box.Questions Regarding the RemovalQuestions regarding the removal may be directed to the following DePuy representatives -Questions about removal information provided: Katie Seppa, Manager of Post-Market Surveillance, 574-372-7333 (M-F; 8 am  5 pm EDT)-Clinical related questions from surgeons: Direct to DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am-5 pm EDT) |  | 8338 | 8338 | US Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, HI, IA, ID, IL, IN, KS, KY, LA, MA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | January 13, 2014 | 531 | 8338 | Not_Computer
Z-2216-2012 | 62344 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory SMART &amp; Sunquest Laboratory SMART Select |  |  | April 28, 2011 | 2011 | Sunquest Information Systems, Inc. | In Online Entry (OEx) there are two scenarios where a Test result from one container may file to a Test on another container. In order for this to occur:1) Site parameter SIT 1,2,1,26 (3,217) [Restrict resulting of tests to owner container ( | Software design | Sunquest sent an Urgent Product Safety Notice dated April 27, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. For the 103 sites that still have this issue Sunquest distribute a new maintenance release. Customers were instructed to use LabAccess Results Workstation (LARS) instead of Online Entry when entering results for instrument method codes that are defined to use replicate tests that are typically ordered on the same accession. Customers were also instructed to share this notification with appropriate laboratory personnel. Customers were questions should call 520-570-2000. For questions regarding this recall call 520-570-2252. |  | 103 | 103 sites | Nationwide Distribution | Terminated | September 04, 2012 | 495 | 103 | Not_Computer
Z-2218-2012 | 62732 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | System 8000 Base |  |  | June 20, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | During a retrospective review of complaint events, TCVS identified on 06/15/12 that three (3)Sarns" System 8000 bases were manufactured and distributed from 11/13/06-11/29/06 with anincorrect incorrect amperage rating identified on the name plate label. The mislabeling wasidentified in January 2007 a set up process of the automark label printer. All units were correctedin the field in Marc | Labeling mix-ups | Terumo field correction/service records were completed in 2007 and all affected units were corrected in the field in March 2007 by replacing the incorrect label at the customer location. For questions about this report call 734-741-6056. |  | 3 | 3 | US Nationwide Distribution - including the state of New York | Terminated | November 07, 2012 | 140 | 3 | Not_Computer
Z-2206-2012 | 62592 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Synthes(R) Hemostatic Bone Putty, Part numbers 08.901.001.97S, 08.901.001.98S, 08.901.001.99S, 08.901.001D, and VB1025.10S. |  |  | July 05, 2012 | 2012 | Synthes USA HQ, Inc. | There is the potential for Hemostatic Bone Putty to ignite if contacted with electrosurgical cautery systems under certain conditions during surgery. | Labeling design | The firm initiated their recall of this product on July 5, 2012 by sending a "Notice: Medical Device Recall" letter to their consignees. The letter identified the affected product along with the reason for recall; and, requested that consignees immediately cease using the product, examine their inventory, and remove the product from inventory. The letter further instructed consignees to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, and to complete and return the verification form that was included with the letter. Questions should be directed to 610-719-5450. |  | 8853 | 8,853 packs | Worldwide Distribution -- US, including Washington, DC and states of AK, AZ, CA, CT, FL, GA, ID, KS, LA, MD, MA, MI, MN, MO, NE, NH, NY, NC, OH, OR, PA, SC, TX, VA, WA, and WI, and the country of Canada. | Terminated | November 06, 2015 | 1219 | 8853 | Not_Computer
Z-2220-2012 | 62649 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Altrx" Altralinked" Polyethylene liners |  |  | July 24, 2012 | 2012 | DePuy Orthopaedics, Inc. | In the recalled lot, the products were not machined according to the production router, a size 52 cup (122136052) was machined instead. This resulted in a size 52 liner being etched and labeled as a size 54 liner. | Employee error | Depuy sent an URGENT MEDICAL DEVICE RECALL notification dated July 24, 2012, to all affected consignees. The recall is extended to the Depuy Distributor level and hospital level. All DePuy Distributors were notified via email. The sales representatives notify the hospital either by mail or in person with a written communication.Distributors and customers were instructed to cease use or distribution of recalled products and to return them to DePuy in Warsaw, IN. The sales representatives are expected to aid customer in returning of recalled products. Distributors and sales representatives are to return the recalled products through the normal DePuy Returns process, to attention of Returns and marking H12-16 on the outside of the box. Reconciliation forms should be returned to the DePuy Sales Representative or faxed to 574-372-7567. For clinical questions from Surgeons contact DePuy's Scientific Information Office at 1-888-554-2482. For questions regarding recall information call 574-372-7333. |  | 20 | 20 liners | Worldwide Distribution - USA including PA and the country of Canada | Terminated | April 30, 2013 | 280 | 20 | Not_Computer
Z-2219-2012 | 62581 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Version 9.5 |  |  | July 13, 2012 | 2012 | Siemens Medical Solutions, USA, Inc | If multiple uncompressed images/clips are sent to the syngo Dynamics server, around the same time, it is possible for an image from a patient to be placed in another patient's study. | Software design | On 07/13/2012, Siemens issued a "Customer Safety Advisory Notice" to affected consignees according to Update Instruction SY062/121S. The notice informs the customers of the potential issue with the identified affected product and provides details regarding the resolution. A software update to relieve the symptom of the issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instruction SY063/121S. A software update to correct this issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instruction SY064/121S. |  | 30 | 30 | Nationwide Distribution -- MA, GA, MO, IA, CA, PA, OR, CO, NY, NJ, WA, UT, OH, WI, MI, KY, and NC. | Terminated | July 15, 2013 | 367 | 30 | Computer
Z-2222-2012 | 62576 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Conformis iTotal CR-Cruciate Retaining Knee Replacement System |  |  | July 17, 2012 | 2012 | Conformis Inc | Addendum Labeling iTotal CR Surgical Technique with Detailed Instructions of Total CR Poly Insertion Technique | Device Design | ConforMIS sent an Urgent Field Safety Notice letter dated July 5, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. An Addendum is included that provided updates to the Surgical Technique Guide. The updated technique guide will be delivered to all iTotal surgeons along with an on-site vist from ConforMIS sales representative. Customers were instructed to place the Surgical Technique Guide Addendum alongside the surgical plan, it puts the information in a highly visible location for the surgeon to review prior to the surgical procedure. Customers were asked to sign the communication verification form attached. For any questions regarding this notice, please contact your local ConforMIS representative or MDSS GmbH, Tel.: +49-511-62628630, info@mdss.com. |  | 1427 | 1427 units | Worldwide Distribution - US Nationwide and the country of Germany | Terminated | February 27, 2013 | 225 | 1427 | Not_Computer
Z-2205-2012 | 62642 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump Module Model 8100 (formerly Medley Pump Module) |  |  | June 15, 2012 | 2012 | CareFusion 303, Inc. | The recall was initiated because Carefusion has identified a potential risk associated with the Alaris Pump module model 8100. The pump module door keypad overlay may separate from the keypad assembly | Device Design | CareFusion sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 20, 2012. The letter identified the product, problem, and actions to be taken by the customers. The potential risk identified in the letter was that when the pump module door keypad overlay has delaminated (loose, peeled away or separated from the door assembly) this could cause a potential for fluid ingress which could lead to a keypad malfunction causing termination of infusion with alarm. Termination of infusion could result in serious injury or death. Customers were informed that CareFusion will contact their facility with in 60 days of receiving the notification to schedule a visit to replace the door assembly of their affected pump module. Customers with recall related questions were instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports questions were instructed to contact Customer advocacy at (800) 854-7128, option 1, option 1, option 3 or e-mail at customer feedback @carefusion.com. Customers with technical questions were instructed to contact technical support at (888) 812-3229. Customers were instructed to promptly complete and return the enclosed Customer Response Card to expedite the corrective action process. |  | 43862 | 43,862 units total (40,780 units in the US) | Worldwide Distribution-USA (nationwide) and the countries of Australia, European Union, and South Africa. | Terminated | May 13, 2014 | 697 | 43862 | Not_Computer
Z-2223-2012 | 62641 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 64 and Ingenuity CT are whole body Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | July 18, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | 1. The current system design of software version 4.0.0xx379 does not prevent the system from scanning if the dose level exceeds a set upper dose limit for CT Brain Perfusion scanning, which previous versions did. However, the current version does have the Dose Check (NEMA XR-25) feature, which provides dose notifications and alerts, but it is possible for the user to disable Dose Check after ente | Software design | Philips Healthcare sent an URGENT - Medical Device Correction letter date July 18, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Phillips will be installing a software update version 4.0.0xx516. A Philips Service Engineer will contact the customer for implementation of the software update on the affected systems free of charge.For further information or support concerning this issue, customers may contact their local Philips representatives or Philips Healthcare office. For North America and Canada, they may also contact the Customer Care Solutions Center (1-800-722-9377, option 5)For questions regarding this recall call 440-483-7000. |  | 8 | 8 | Worldwide Distribution - USA including IN and OH, also, the countries of Australia, Belgium, China, Israel, and Sweden | Terminated | February 22, 2013 | 219 | 8 | Computer
Z-2224-2012 | 62745 | device, prosthesis alignment | IQO | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | Oxford Gap Gauge SML 3/4mm. |  |  | July 27, 2012 | 2012 | Biomet U.K., Ltd. | A drafting error during an unrelated design change has led to 4 misleading dimensions on the drawing. The CAD modelremained the same thickness but the problem dimensions were accidentally struck from an incorrect point result in anincrease of 1.0mm in the value of each dimension | Device Design | Biomet sent a "FIELD SAFETY CORRECTIVE ACTION" letter dated July 24, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Response Form was attached for customers to complete and return. Contact the firm for questions regarding this notice. |  | 119 | 119 units | Worldwide Distribution-USA-including the states of IN, SD, and TX and the countries of Austria, China, Japan, Netherlands, Sweden, Turkey, and UK. | Terminated | February 15, 2013 | 203 | 119 | Not_Computer
Z-2221-2012 | 62798 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC 8800 Mobile Fluoroscopic X-ray System, GE Healthcare, Surgery, Salt Lake City, UT. |  |  | October 19, 2007 | 2007 | GE OEC Medical Systems, Inc | X-ray units did not meet manufacturer's specification for linearity. | Nonconforming Material/Component | GE updated their user manual and changed their internal specifications to match those of the CFR In addition, a software code correction was to be made to forward productions in order to correct the reported defect.*** This recall issued was previously classified under an incorrect recall number. It is being reclassified with a new number. *** |  | 175 | 175 units | Worldwide Distribution - USA - Australia, Brazil, Cameroon, China, Colombia, Cyprus, The Czech Republic, Finland, France, Germany, Greece, Italy, Japan, Korea, Malaysia, The Netherlands, New Zealand, Norway, Russian Federation, Spain, Sweden, Taiwan, Turkey, UAE, UK, and Canada. | Terminated | August 17, 2012 | 1764 | 175 | Computer
Z-2225-2012 | 62755 | No | CKG | Adrenocorticotropic hormone (ACTH) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel DxH Slidemaker Stainer Instrument System, PN 775222. |  |  | December 13, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the DxH Slidemaker Stainer Floor Stand shipped with Slidemaker Stainer instruments were missing vinyl caps, exposing sharpt edges on the DxH Slidemaker Stainer Reagent Cabinet (Floor Stand) drawer slides. The sharp edges on the drawer slides may expose operators to physical injury when replacing reagents and waste containers whic | Process change control | Beckman Coulter initiated a "PRODUCT CORRECTIVE ACTION (PCA)" which was communicated via phone calls starting the week of December 19, 2011 to all customers who purchased UniCel DxH Slidemaker Stainer Coulter Cellular Analysis System. The phone calls provided the customers with an explanation of the problem identified, potential impact should the problem occur and an action to be taken. Customers were instructed to share the PCA information with their laboratory staff and if they had any questions, they should contact Beckman Coulter at (800) 526-7694. |  | 41 | 41 units (15 units in the US) | Worldwide Distribution-USA (nationwide) including the states GA, NY, IL, TX, WA, PA, IN, MS, NC, and MO and the countries of Belgium, France, Hong Kong, Italy, Spain, and Switzerland. | Terminated | August 20, 2012 | 251 | 41 | Not_Computer
Z-2227-2012 | 62791 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Total T3 Reagent, PN: 33830 |  |  | October 11, 2011 | 2011 | Beckman Coulter Inc. | A recall was initiated because Beckman Coulter has identified a negative bias in test results when comparing elevated Total T3 patient samples diluted with Sample Diluent A to samples diluted with Access Total T3 Calibrator S0. | Device Design | Beckman Coulter sent a Product Correction letter dated October 11, 2011, to all affected customers. The PCA letter informed the customers of the problem identified and the actions to be taken. Customers were informed to stop SDA usage as an alternative dilutent and that the S0 dilutent is to be used for manual dilution of the patient sample for Total T3. They will also be informed to no longer use the On Board Dilution (OBD) option for Total T3. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions regarding this notification were instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-961-4941. |  | 117385 | 117,385 units total (33,918 units in US) | Worldwide Distribution - USA (nationwide), Canada, and internationally. | Terminated | March 14, 2013 | 520 | 117385 | Not_Computer
Z-2228-2012 | 62640 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | Siemens ADVIA CentaurÂż Syphilis (SYPH) Assay , 200 ReadyPack |  |  | July 17, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Syphilis Assay Interference with HBsAg Assay | Component design/selection | Siemens Healthcare sent Urgent Device Recall Notice letters dated July 2012 via e-mail to all affected customers inside and outside the US. The letters identifed the affected product, reason for voluntary recall and risk to health. The letters provided actions to be taken by the customers and if there are any remaining inventory of the affected product, follow the guidelines outlined. Customers were instructed to complete and return the Field Correction Effectiveness Check form. For questions contact your local Technical Support Provider or Distributor. |  | 2050 | 2050 Kits | Worldwide Distribution - US Nationwide and the countries of:Foreign: Brazil, China, Denmark, Hong Kong, Malaysia, and New ZealandViet Nam | Terminated | November 22, 2013 | 493 | 2050 | Not_Computer
Z-2230-2012 | 62684 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Serial interface module RS-232 |  |  | July 26, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems has received reports of malfunctions involving the loss of data transfer capabilities using the data transfer capabilities using modules on the Terumo Advanced Perfusion System 1. Some reports stated that the odor of smoke was as associated with the failures. The malfunctions occurred during set-up, during cardiopulmonary bypass, and during routine testing of the modu | Device Design | Terumo sent an URGENT MEDICAL DEVICE Recall letter dated July 24, 2012 to all affected customers. The letter identified the affected product, reason for correction, potential hazard and actions to be taken. The letter stated that Terumo CVS will replace all affected modules. Customers were informed that a Terumo VS Representative will contact the user to schedule a service call for the replacement upon receipt of the Response Form attached. For questions contact Terumo CVS Customer Service at 800-521-2818. |  | 579 | 579 units | Worldwide Distribution - US Nationwide including the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI, and WV and the countries of: AUSTRALIA, BELGIUM, CANADA, CHILE, COLOMBIA, Costa Rica, Dominican Republic, Germany, Hong Kong, Indonesia, Japan, Jordan, Malaysia, Singapore, South Africa, South Korea, Taiwan, Thailand, UNITED ARAB EMIRATES (UAE), and Vietnam | Terminated | June 27, 2014 | 701 | 1427 | Not_Computer
Z-2234-2012 | 62134 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OPTETRAK LOGIC Proximal Tibial Spacer &amp; Screws (2) 2.5 LT BLUE Size 2.5, 8mm***STERILE***Match insert/spacer/tray color. Use with LOGIC Tibial Components. REF 02--012-42-2508. |  |  | October 17, 2011 | 2011 | Exactech, Inc. | Exactech, Inc. of Gainesville, FL recalled their Optetrak Proximal Tibial Spacer after the devices were determined to be dimensionally incompatible to mate as intended with the Logic Tibial Tray. | Device Design | The firm, Exactech, Inc., sent a "IMPORTANT PRODUCT RECALL NOTICE" date October 17, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately cease distribution or use of these products; extend this information to their accounts that may have possession of this product; verify if they have any of the subject Logic size 2.5 Proximal Tibial Spacer products, complete and return the attached PRODUCT RECALL FAX NOTICE form within five (5) working days via fax to: 352-337-3915, and to contact the firm inventory representative to confirm quantities at the location. For questions regarding inventory restocking, please call 1-800-392-2832. |  | 70 | 70 | Nationwide distribution: USA including states of: AL, AZ, CO, FL, GA, IL, LA, MD, MN, NC, NJ, NY, OH, OK, PA, PR, TN, VA, and WA. | Terminated | November 07, 2013 | 752 | 70 | Not_Computer
Z-2229-2012 | 62452 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory SMART &amp; Sunquest Laboratory SMART Select |  |  | August 12, 2011 | 2011 | Sunquest Information Systems, Inc. | In GUI inquiry, searching by ID (CID) may return the wrong patient demographic data if the accession number associated to the CID is no longer inactive. | Process design | Sunquest sent an Urgent - Product Safety Notice letter dated August 12, 2011 to all affected customers. The letter identified the affected product, summary of the problem, actions to eliminate the problem and actions to be taken by clients. Sunquest will distribute a new maintenance release.Client Support can be reached at 1-877-239-0337 for U.S and Canadian sites and at +44 (0) 101-335-0562 for international sites. |  | 47 | 47 sites | US Nationwide Distribution | Terminated | September 04, 2012 | 389 | 47 | Not_Computer
Z-2244-2012 | 62571 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-armÂż Imaging System. Product Numbers: BI-700-00027-100, BI-700-00027-100R, BI-700-00027-120, BI-700-00027-120R, BI-700-00027-230, BI-700-00027-230R. |  |  | July 11, 2012 | 2012 | Medtronic Navigation, Inc. | A fault in a high voltage relay in the 0-armÂż Imaging System might cause the relay to overheat and burn. | Component change control | Medtronic Navigation issued an "Urgent Field Safety Notice" to consignees dated 7/11/12. The problem was described and recommended actions were provided. For each affected O-armÂż Imaging System, a Medtronic service representative will visit the site and replace the 500VDC relay on the system and test the O-armÂż Imaging System. |  | 17 | 17 units | Worldwide Distribution - USA, including the states of AL, CO, FL, IA, MI, IA, NC, NE, NY and the countries of South Africa, Italy, and Spain. | Completed |  |  | 17 | Not_Computer
Z-2238-2012 | 62567 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging system. |  |  | May 29, 2012 | 2012 | Mckesson Medical Imaging Group | There is a potential to save reports in RadReport with an Incorrect Final Status. | Software design | McKesson sent an Urgent Field Safety Notice letter dated May 29, 2012 to all affected consignees. The letter identified the affected product, circumstances under which the issue occurs and immediate actions to be taken. Customers were advised to notify the Customer Support department that they have read and understood this Advisory Notice at 1-800-663-2533. For questions contact the Support Manager at the Customer Support department. |  | 516 | 516 units | Worldwide Distribution - U.S. Nationwide and the countries of: Canada, France, Ireland and United Kingdom. | Terminated | September 12, 2013 | 471 | 516 | Not_Computer
Z-2239-2012 | 62892 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Positron Emission Tomography (PET) and X-ray Computed Tomography, GEMINI TF 16 PET/CT, GEMINI GXL 16 PET/CT, GEMINI TF Big Bore PET/CT, Philips Medical Systems, Cleveland, OH. |  |  | August 08, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips became aware of an issue relating to a UPS battery cabinet that appeared to have burned a hole through the casing on one of the batteries. Supplier of the UPS, Chloride/Emerson, determined the batteries, housed within the UPS battery cabinet, entered thermal runaway causing the battery to overheat. | Material/Component Contamination | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated August 8, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Field Service Engineers will be implementing a Field Change Order (FCO) to replace the batteries of the affected UPS units. Contact the Philips Healthcare Customer Care Solutions Center at 1-800-722-9377 for questions regarding this notice.. |  | 7 | 7 units | Nationwide Distribution-including the states of CO, FL, MD, MN, NE and PA. | Terminated | April 28, 2014 | 628 | 7 | Not_Computer
Z-2246-2012 | 61671 | No | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | Powered Mobility Offers Versatility (MÂżV)TM Chair |  |  | June 27, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Paramed Corporation (d/b/a, Descent Control Systems), a manufacturer and distributor of patient transport devices, merged with an affiliate of Stryker Corporation in August 2009. As part of the due diligence process, Paramed's powered Mobility Offers Versatility (MÂżV) chair was reviewed to identify potential compliance issues in relation to product documentation and classification. | PMA | Stryker sent a Urgent Medical Device Recall notice dated June 18, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to " Return the enclosed post card to confirm receipt of this notification." Locate the chairs listed in this notice." Call 269-389-6750 to arrange shipment of the chairs to Stryker Medical." If you have loaned or sold any of the chairs listed in this letter, please, forward a copy of this notice to the new users and advise us of their new location." If you have disposed of any of the affected chairs and they are no longer in use, please, advise us of their obsolescence by providing us with their serial numbers.We apologize for any disruption that this advisory notice may cause. We strive to make products that meet our customer expectations for quality and durability.For further questions please call (269) 389-6927 |  | 159 | 159 units | Worldwide Distribution -- USA (nationwide) including the states of AK, AL, AZ, CA, FL, ID, IN, KS, KY, LA, MA, MD, MN, MO, NC, NJ, NY, OR, PA, TN, TX, UT, UT, VA, WA, WI, and WV, and the country of Canada. | Terminated | December 30, 2013 | 551 | 159 | Not_Computer
Z-2240-2012 | 62691 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | 892.1650 Image-intensified fluoroscopic x-ray system; the device is designed to provide fluoroscopic and spot film imaging of the patient during diagnostic surgical and interventional procedures. The systems include features specifically designed for use in diagnostic and interventional cardiac imaging procedures, and are also intended for cholangiography, endoscopic, urologic, orthopedic, neurologic, vascular, critical care and emergency room procedures. The systems may be used for other imaging applications at the physician's discretion. |  |  | November 06, 2009 | 2009 | GE OEC Medical Systems, Inc | GE OEC is recalling certain 9800 Image Intensified Flouroscopic x-ray systems due to customer complaint analysis showing that a replacement Hitachi 160 GB Hard Drive, when installed on certain models is susceptible of causing unanticipated system shut downs, no boots, data loss or data mix, or unexpected hard drive corruption. | Device Design | GE OEC sent a Urgent Recall Notice dated November 6, 2009, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The notification advised them that GE Healthcare had identified a potential safety issue associated with a "Significant increase on Data Loss, Data Mix, and System file corruption of the Hard Drives" on the 9800 Systems.The field correction consisted of installing a "Smart Power Switch" between the power control PCB and the system power switch on 9800 Systems with a Celeron single board computer and software version 29. This will remedy the potential for data loss, data mix, and system file corruption of the hard drives.For further questions please call 800-874-7378 Option 8. |  | 3595 | 3,595 units | Worldwide Distribution -- USA (nationwide) | Terminated | January 28, 2013 | 1179 | 3595 | Computer
Z-2237-2012 | 62410 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT cTnI cartridges |  |  | March 07, 2012 | 2012 | Abbott Point Of Care Inc. | Abbott Point of Care has determined that i-Stat CTNL cartridges have the potential to exhibit incorrectly elevated cardiac troponin I results. This recall was expanded on July 19 2012 to include additional information regarding user documentation. | Device Design | Abbott Point of Care (APOC) sent an Urgent Recall Notice letters, via Fed Ex on March 07, 2012 to all affected customers. APOC sent an important product information update on July 19, 2012 to all of their affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to review the information included as it may pertain to their cartridges. Customers were instructed that if the affected product have been forwarded to another facility, provide a copy of this letter to them. For additional information, contact APOC at 800-366-8020 option 1 or your Abbott Point of Care representative. |  | 7830075 | 7,057675 cartridges US; 772,400 cartridges OUS | Worldwide Distribution - US Nationwide, International | Terminated | February 19, 2013 | 349 | 7830075 | Not_Computer
Z-2241-2012 | 62729 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE HC OEC 9900 Elite and GE OEC 9800 Fluoroscopes Snubber Boards are a PCB with a 70 amp fuse to provide fused protection of internal system components from the high voltage system components, such as the X-ray tube or High Voltage Transformer. System model number A349855, Snubber Board number 00-880405-03 rev 5. This product is used for image intensified fluoroscopic x-rays for interventional radiologic procedures. |  |  | April 26, 2010 | 2010 | GE OEC Medical Systems, Inc | GE OEC recalled certain 9800 Fluoroscope &amp; 9900 Elite Models because the Snubber Board did not conform to the IEC 60601-1 standard for clearance and creepage of voltage isolation. | Component change control | No customer communication took place. On April 26, 2010, GE began opening service cases for field engineers to begin replacing the snubber boards. This action has been completed as of October 29, 2010. |  | 845 | 845 units | Worldwide and Nationwide (USA) distribution. | Terminated | August 22, 2012 | 849 | 845 | Not_Computer
Z-2243-2012 | 62751 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Encompass |  |  | April 13, 2009 | 2009 | Sunquest Information Systems, Inc. | Sunquest is recalling the Sunquest Encompass versions 2.4 and later because the specimen collect time is displaying and printing one hour later than the actual collect time only when Daylight Saving Time is active. | Software design | The firm, Sunquest, sent a "Urgent-Product Safety Notice: PSN-09-01" letter dated April 13, 2009 all their customers who purchase the Sunquest Encompass versions 2.4 and later. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to request software correction and to create a service request via client support webpage at www.sunquestinfo.com/supportweb. Then choose Submit a Service Request. Select Issue Type and select Software Request. When Software Request is selected, the Work Queue field automatically fills in with "Software Request." Promptly advise appropriate personnel.Customers with questions can contact Client Support and Services at (877) 239-6337 for US and Canadian sites and at +44(0)0800-960-972 for international sites. |  | 23 | 23 clients | Worldwide distribution: USA (nationwide) and country of: Bahamas. | Terminated | August 22, 2012 | 1227 | 23 | Computer
Z-2242-2012 | 62460 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatic System is included with the Dimension Vista System. The Dimension Vista Integrated System is an in vitro diagnostic device intended tot duplicate manual analytical procedures such as pipetting, mixing, heating, and measuring spectral intensities to determine a variety of analytes in human body fluids. |  |  | May 29, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | When using the EasyLink Quality Control Package the system may not perform as intended for Multi-Rule QC violations [e.g., 2(2S), 4(1s) and 10x], unless configured using the instructions send with the Urgent Field Safety Notice. When this situation occurs, QC status may not be flagged appropriately. If auto-verification rules are defined to hold patient results when QC status is out, the patient | Software design | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated May 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Form was attached for customers to complete and return via fax to 302-631-8467. Contact Siemens Technical Solutions Center at 800-441-9250 for questions regarding this notice. |  | 637 | 637 | Worldwide Distribution-USA (nationwide) and the countries of Canada, France, Germany, Spain, Portugal, Italy, Japan, Norway, Australia, Austria, New Zealand, Belgium, Switzerland, Netherlands, Malaysia, Slovenia, South Korea, Denmark, Saudi Arabia, Slovakia, Taiwan, Philippines, Czech Republic, and Singapore. | Terminated | April 10, 2014 | 681 | 637 | Not_Computer
Z-2252-2012 | 62796 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AGFA IMPAC CV DICOMStore with Media Purge Daemon (MPD) and IMPAX CV DICOMStore with Cardiovascular Purge Services (CPS), Medical Image Storage Device |  |  | August 09, 2012 | 2012 | AGFA Corp. | Loss of patient data can occur under certain circumstances due to misconfiguration of DICOMstore configured with Agfa's Medical Image Storage Devices - Media Purge Daemon (MPD) and Cardiovascular Purge Service (CPS). | Process design | On 8/9/2012 AGFA Healthcare issued an "URGENT FIELD SAFETY NOTICE" letter to consignees. This letter describes the problem and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood was also required from the consignees |  | 535 | 535 | USA Nationwide, Virgin Islands and Canada | Terminated | December 08, 2015 | 1216 | 535 | Not_Computer
Z-2245-2012 | 62790 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 6, 10, 16, 40, 64 and CT Big Bore are Computed Tomography X-Ray Systems These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | August 03, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The patient support may move in an unintended manner if the foot switch cover becomes damaged and impinges on the foot switch. | Device Design | Philips Medical Systems (Cleveland), Inc. sent an Urgent - Medical Device Correction letter dated August 1, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taked. The letter contained important information for the continued safe and proper use of your equipment. The letter stated that, a Field Service Engineer will visit each affected site to install the hardware upgrade. For further information or support concerning this issue contact your local Philips representative or local Philips Health office. For North America and Canada contact Customer Care Solution Center (1-800-722-9377), option 5: Enter Site ID or Follow the prompts). |  | 126 | 126 units | Worldwide Distribution - US Nationwide including the states of: AL, AR, AZ, CA, DC, FL, GA, KY, MD, MI, MN, MS, ND, NM, NY, OH, OK, SC, TN, TX, VT, WA, and WI and the countries of: Austria, China, Denmark, France, Germany, Iceland, Japan, Netherlands, South Africa, Sweden, Switzerland, United Kingdom. There were two government consignees. There were no Canadian consignees. | Terminated | July 23, 2015 | 1084 | 126 | Not_Computer
Z-2235-2012 | 62663 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | The Instructions For Use for the following: |  |  | April 25, 2012 | 2012 | Stryker Spine | The front cover of Instructions For Use (IFU) #NOLI135B03Rev02 was improperly marked with a "Sterile" logo in the lower right corner. The products these IFUs were packed with are sold as "Non-Sterile" and require sterilization before use in surgery. | Labeling mix-ups | Stryker sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter to affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Customer Response Form was attached for customers to complete and return via fax to 201-760-8370. Contact the firm at 201-760-8298 for questions regarding this notice. |  | 40 | 40 units. | Distributed in Atlanta, Georgia. | Terminated | February 27, 2013 | 308 | 40 | Not_Computer
Z-2236-2012 | 62744 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | OEC InstaTrak 3500, picture archiving and communications system intended as an aid to locate anatomical structures during open percutaneous surgical procedures. It is indicated for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. |  |  | October 11, 2006 | 2006 | GE OEC Medical Systems, Inc | GE Healthcare had recalled certain OEC InstaTrak 3500 Carts due to the potential for the cart to tip over when the arm of the imaging device is extended during use. | Component design/selection | GE Healthcare sent an "URGENT SAFETY ADVISORY NOTICE" dated October 11, 2006 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. The letter stated that GEHC OEC Field Service Engineers would be contacting their facility to place a caution label on the articulating monitor arm of the cart and to provide a revised operators manual that will contain additional information on the safe usage of the cart. Until a GEHC Engineer had taken the actions described, they advised their consignees to exercise caution when fully extending the articulating monitor arm on the device to avoid potentially tipping over the cart. Contact the firm at 800-874-7378 for questions regarding this notice. |  | 99 | 99 units | Nationwide Distribution. | Terminated | August 22, 2012 | 2142 | 99 | Not_Computer
Z-2259-2012 | 62706 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 660i, 680i, 860i, 880i SYNCHRON Access Clinical Systems, |  |  | October 13, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has received seven reports of a software anomaly for the UniCel Dxl Immunoassay Systems. If undetected, this anomaly may cause erroneous test results. | Software design | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated October 13, 2011 with attached PCA Response Form to all customers who purchased the UniCel Dxl Immunoassay Systems and DxC Clinical Systems which includes all software versions. The letter describes the product, problem and actions to be taken. The customers were instructed verify the system temperatures immediately following each initialization; do not begin processing samples until this verification is complete; refer to the appendix included with the letter for detailed instructions about verifying system temperatures; review their quality control (QC) performance history; and share this information with their staff and retain this information as part of their quality system documentation, and to complete and return the enclosed response form within 10 days to Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318; Attn: Regulatory Affairs or fax to: 1-786-639-4000. Customers with any questions regarding the notice, were instructed to contact Customer Support Center via our website, http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representatives. |  | 2442 | 2442 units total (784 units in US) | Worldwide distribution: USA (nationwide) and countries of: Algeria, Angola, Argentina, Australia, Austria, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Czech Republic, Finland, France, French Polynesia, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea,Republic of; Kuwait, Liechtenstein, Malaysia, Mayotte, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | August 23, 2012 | 315 | 2442 | Computer
Z-2254-2012 | 62870 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon MRI System/Oasis MRI, Software. These are software controlled MRI devices. |  |  | June 15, 2012 | 2012 | Hitachi Medical Systems America Inc | Hitachi discovered a software error with a feature called Multi-Planar Reconstruction(MPR). MPR takes an original volume dataset reconstructed in any plane view and creates a new image set from another view (i.e. axial to coronal, coronal to sagittal, etc.). The error causes the first image in the dataset to have a right/left reversed orientation. The software error can occur randomly when execu | Software design | Hitachi sent a Device Correction Notice letter dated June 18, 2012 to all affected customers. All letters are recorded by tracking number and date received by the consignee. The letter identified the affect product, problem and actions to be taken. All sites will be visited by Hitachi Service to install the software update to permanently resolve the issue at no cost to the facility. Software correction is targeted to start by September 1, 2012. For question contact the Manager of Regulatory Affairs at 800-800-3106 x3720 |  | 165 | 165 units | US Nationwide Distribution - including the states of:: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY and Mexico. | Terminated | February 19, 2013 | 249 | 165 | Computer
Z-2261-2012 | 62871 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Visual-ICEÂż Cryoablation System, Model FPRCH 6000 |  |  | July 20, 2012 | 2012 | Galil Medical, Inc. | Galil Medical is recalling a number of Visual Ice Cryoablation Systems because the gas regulators may fail to regulate the gas pressure, rendering the devices inoperable. | Device Design | Galil Medical sent a "Voluntary Medical Device Field Correction" letter dated July 20, 2012, via UPS to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that a Galil Field Service team will be contacting customers to schedule a service visit the week of July 23rd. Customers were instructed to forward and communicate this field correction notice and instructions to individuals in their organization and/or to other organizations affected by this action. For questions call Galil Medical Customer Service. |  | 4 | 4 | Worldwide Distribution - US nationwide including the state of MN and the countries of United Kingdom and Netherlands | Terminated | September 13, 2012 | 55 | 4 | Not_Computer
Z-2262-2012 | 62746 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC Injection Cable for MedRad Provis V Contrast Injector. |  |  | October 10, 2008 | 2008 | GE OEC Medical Systems, Inc | GE initiated a correction to the instructions/use due to the potential for erroneous auto injection of contrast dye. | Device Design | An Urgent Safety Advisory Notice was sent to consignees on October 10, 2008. The letter identified the affected product and notified customers of the situation, along with instructing them to discontinue use until proper validation and instruction update for cable use was completed. GEHC OEC was to have contacted customers when the solution became available and resolved it at no cost to the customer. |  | 149 | 149 units | Nationwide distribution. | Terminated | January 10, 2013 | 1553 | 149 | Not_Computer
Z-2266-2012 | 62851 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba CT systems, Asteion and Aquilion, that are equipped with a PC based computer, or to those older systems (SGI based) that received an improved table by an upgrade purchase. |  |  | February 27, 2008 | 2008 | Toshiba American Medical Systems Inc | The systems all have a similar table control board. It hasbeen found that this board may cause the table position to be incorrect if the table cradle is movedmanually. This is referred to as "Free" mode in the operation manuals. If this error does occur it ispossible that the patient could be scanned at an incorrect position. | Process design | Toshiba sent a "URGENT: MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to please avoid moving the table in the "Free" or manual mode as much as possible. If it is necessary to move the table in this manner please verify the table position prior to initiating the scan sequence.If it is noted that the position is not correct, exit the eXam plan and start a new one beginning with the scanoscopy. Please share this information with all users of the equipment and your clinical or biomedical engineering departments.Please find the attached customer reply form. The purpose of this form is for FDA tracking. Please take a moment to complete the form and fax it to the toll free number on the form.". The contact information for this recall is:"(800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968.". |  | 757 | 757 | USA ( Nationwide Distribution ) | Terminated | September 04, 2012 | 1651 | 757 | Not_Computer
Z-2268-2012 | 62881 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Mini-Infuser Model 150XL Single Speed Infusion Pump, an Rx battery operated syringe infusion pump; |  |  | March 04, 2011 | 2011 | Baxter Healthcare Corp. | Should a battery with the minimum contact diameter (.51 inches) be placed on the battery spring (.53 inches), an electrical circuit will not be completed and the Mini-Infuser will not operate. The reason being that the Mini-Infuser Battery Spring has a flat contact that measures .53 inches Âż.01 inches on the inner diameter of the spring coil. The IEC 60086-2 and ANSI 08 standards specify a minimum | Nonconforming Material/Component | Baxter sent a Safety Alert letter on January 11, 2011 for the InfusO.R. Pump. Baxter expanded the scope of the recall to include the Mini-Infusers System and sent an Urgent Device Correction letter dated March 4, 2011, via First Class mail to all affected customers. The letters informed the customers that the Mini-Infusers use the same size battery and incorporate the same battery contact interface as the InfusOR pump, and that the same feature that inhibits electrical contact in over-the counter Energizer C size batteries may also be incorporated by other battery manufacturers. The customers were requested to verify that the InfusOR and Mini-Infuser will power-on before use; if the device does not power on, replace the batteries with a different brand of size C batteries and verify that the device will power-on before use; and to place a copy of the letter with each operator's manual for the InfusOR and Mini-Infuser System for future reference. For Technical questions call Baxter Global Technical Services at 1-800-843-7867. For question regarding this communication contact the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 26485 | 26,485 units | 45,865 units | Terminated | August 27, 2012 | 542 | 45865 | Not_Computer
Z-2270-2012 | 62708 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory Versions 6.4 and later using Microbiology Result Entry |  |  | December 05, 2011 | 2011 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest has confirmed that the Sunquest Laboratory versions 6.4 and later while using Microbiology Result Entry (MRE), the specimen ribbon containing the Accession Number (AN) and the Battery Name may be missing or contain the wrong information. | Software design | Sunquest sent a recall letter dated December 5, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your reply to this email message serves as verification you have received, read, understood and/or followed instructions outlined in the attached Product Safety Notification. Your reply assists Sunquest Information Systems to satisfy our obligation to Regulatory Agencies.Please note in your email reply if your site will not be requesting the software correction. If your site plans to request the software correction, please follow the directions noted on the attached PSN.Further questions please call (520) 570-2347. |  | 127 | 127 sites | Worldwide Distribution -- USA (nationwide) and the country of Canada. | Terminated | August 27, 2012 | 266 | 127 | Computer
Z-2267-2012 | 62864 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviCide, Part Numbers: 13-1025, 13-1000, 13-1024, MC-1000. |  |  | February 06, 2012 | 2012 | Metrex Research, LLC. | Metrex Research is recalling Cavicide containing the Spring Fresh Fragrance because the affected lot of product containing this fragance has caused an increase in reported health reactions in some individuals such as headache, nausea, and respiratory discomfort. | Component change control | Metrex sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 16, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Mandatory Recall Response Form was attached for customers to complete and return via fax to 734-947-6725.Contact Customer Care Center at (800) 841-1428 for questions regarding this recall. |  | 17690 | 17,690 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, and YE. | Terminated | August 24, 2012 | 200 | 17690 | Not_Computer
Z-2271-2012 | 62578 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Low Profile Abutment, Non-Hexed Castable Cylinder Model LPCCC1. Product is labeled in part: "***REF LPCCC1***Low profile Abutment Hexed Castable Cylinder***Cylinder; Cylindre; Zylinder; Cilindro; Cilindro; Cilindro***RX only***NON STERILE***Caution: Law prohibits dispensing without prescription***Consult instructions for use*** |  |  | May 31, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i conducted a recall on the device Low Profile Abutment, Non-Hexed Castable Cylinder due to some packages labeled LPCCC1, Hexed Castable Cylinder possibly containing LPCCC2, Non-Hexed Castable Cylinders. | Labeling mix-ups | Biomet 3i sent a "MEDICAL DEVICE ADVISORY NOTICE" dated May 31, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-342-5454 for questions regarding this recall. |  | 611 | 611 units | Worldwide Distribution-Distributed in the USA and the countries to Spain, France, Germany, Switzerland, and Australia. | Terminated | November 21, 2013 | 539 | 611 | Not_Computer
Z-2272-2012 | 62714 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory LabAccess Results Workstation (LARS) |  |  | October 05, 2010 | 2010 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest has confirmed that the Sunquest Laboratory LabAccess Results Workstation (LARS) may file results to the wrong patient by incorrectly combining results. | Software design | An "Urgent - Product safety notice " was issued on 11/4/10 to all customers who purchased the Sunquest Laboratory LabAccess Results Workstation (LARS). The safety notice informed the customers of the problems identified and the actions to be taken. US customers may contact 1-877-239-6337. |  | 151 | 151 sites | Worldwide Distribution, including Nationwide (USA) and the countries of Canada, Denmark, Ireland, the United Arab Emirates and the United Kingdom. | Terminated | August 24, 2012 | 689 | 151 | Not_Computer
Z-2263-2012 | 62770 | No | NSX |  |  | General Hospital | Contact ODE | Cisco brand HealthPresence, version 2.1; Model/Catalog Number: CHP-ENTPRSYR2.I-K9; Product is manufactured and distributed by Cisco Systems, Inc. San Jose, CA |  |  | July 23, 2012 | 2012 | Cisco Systems Inc | Software anomaly may result in incorrect values and interpretations. The initial vital signs reading will display accurately on the Cisco Health Presence Vitals tab and the OnePlace tab. However, subsequent vital signs readings will only update on the Cisco HealthPresence Vitals tab and within the OnePlace database, but will not update under the OnePlace tab, and thus, will not provide current vit | Software design | Cisco notified their one customer by phone on July 25, 2012, with a follow up email and instructed the customer to stop using the OnePlace vitals tab and to use the CHP only. On July 27th, 2012, the vendor of the OnePlace application, fixed the bug remotely.On July 31, 2012, Cisco's staff physician contacted the customer to confirm that the customer was aware of the fix. Validation of the correction was completed by Cisco, along with customer on August 9, 2012. Until validation was satisfactorily completed, the integrated system lie dormant at customer's site. |  | 1 | 1 unit | Nationwide Distribution including the state of TN. | Terminated | August 24, 2012 | 32 | 1 | Computer
Z-2265-2012 | 62828 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Canon Solid State X-ray Imager (Flat Panel/Digital Imager), Digital Radiography, Models CXDI-55C, CXDI-55G and CXDI-70C. These devices are intended to replace radiographic film/screen systems in all general purpose diagnostic procedures. |  |  | November 01, 2010 | 2010 | Canon Inc. | Screws holding the 55REF circuit board are loosening and roaming inside of the sealed sensor. | Nonconforming Material/Component | Canon Inc. notified sales companies including Canon USA on 9/15/2010 with Service Bulletin 10-017, and on 8/25/2011 with Service Bulletin 11-011, and on 1/20/2011 with Service Bulletin 11-001. Sales representatives were to visit customers and check and replace the screws. |  | 468 | 468 units | Nationwide Distribution (USA) | Terminated | January 07, 2014 | 1163 | 468 | Not_Computer
Z-2264-2012 | 62678 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS. |  |  | July 18, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction when using syngo Imaging XS, version VA70A or higher. In some situations it can happen that a miscalculation is done for the grey scale values in the functions "Regions of Interest", "Pixel Lens", "Edge Enhancement" and "Histograms". | Software design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated July 12, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. For questions concerning this notice contact your local Siemens Service Rep.. |  | 52 | 52 | Nationwide Distribution-CA, CO, CT, FL, GA, IL, IN, IA, KS, MD, MN, MO, NE, NV, NJ, NY, NC, OH, PA, TN, TX, and WI. | Terminated | November 06, 2014 | 841 | 52 | Computer
Z-2273-2012 | 51437 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TRESTLE LUXEÂż Anterior Cervical Plating System, Product code KWQ. 510(k) K102820. |  |  | May 29, 2012 | 2012 | Alphatec Spine, Inc. | On May 9th 2012 Alphatec Spine received a complaint which reported that the distal tipof the Trestle Luxe Slide Alignment Tool (Part Number 71721), does not properly fit intothe hexalobe screw on the Trestle Luxe Plate. | Device Design | Alphatec Spine sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" letter dated May 29, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A form was attached for customers to complete and return via fax to 760-431-0289. Contact Alphatec Spine's Customer Service Department at customerservice@alphatecspine.comfor questions regarding this recall.) |  | 38 | 38 | Nationwide Distribution. | Terminated | September 04, 2012 | 98 | 38 | Not_Computer
Z-2274-2012 | 62564 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Apelo Modular Screwdriver labeled in part: "APELO PEDICLE SCREW SYSTEM***Modular Screwdriver***Atlas Spine, Inc.***www.AtlasSpine.com*** 1555 Jupiter Park Drive #4***Jupiter, Florida 33458***561-741-1108***Tissue Sleeve***MFG Date: XX/XX/XXXX***Lot #: XXXX***Weldment***MFG Date: 03/23/2009***Sleeve Assembly***MFG Date: XX/XX/XXXX***Lot #: XXXX***Qty: 1 Rx Only***Non-Sterile. Sterilize Prior to Use*** Rev. A***. The Modular Screwdriver is used to insert modular pedicle screw implants which are part of the Apelo System. |  |  | June 04, 2009 | 2009 | Atlas Spine, Inc. | Atlas Spine, Inc. recalled their Apelo Modular Screwdriver, a component of the Apelo Pedicle Screw System Instrumentation, because of complaints received about the tip breaking. If the tip is fractured inside of the modular screw, the tip could become implanted in the body or if the entire screw is removed, it could cause a delay in the surgery. | Component design/selection | Consignees were notified by Telephone and visit. All products will be returned, quarantined and then destroyed. |  | 4 | 4 pieces | USA Distribution to the state of Florida only. | Terminated | July 16, 2014 | 1868 | 4 | Not_Computer
Z-2276-2012 | 62919 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX OPA, High Level Disinfectant Solution, Glutaraldehyde-free (0.55% ortho-phthaladehyde) solution. The CIDEX OPA solution is intended for automated and manual cleaning and disinfection. |  |  | June 01, 2012 | 2012 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products has discovered instances of missing labels on bottles of CIDEX OPA Solution, Part # 20390 manufactured between May 2011 and April 2012. | Process control | An "Urgent Medical Device Recall" letter dated August 10, 2012 was sent to all customers who purchased the CIDEX OPA solution, part #20390 (manufactured between May 2011 and April 2012). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to notify their customers if they have distributed bottles of CIDEX OPA Solution to other locations. Customers were instructed to call Stericycle at (866) 629-6182 to arrange for return of the product. |  | 1310844 | 1,310,844 units | Nationwide (USA) and Puerto Rico. | Terminated | September 11, 2013 | 467 | 1310844 | Not_Computer
Z-2277-2012 | 62780 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 16 Fr HemoStar XK Long-Term Hemodialysis Catheter, implanted Standard kit containing an AirGuard introducer. The device is used for patients requiring hemodialysis. |  |  | August 07, 2012 | 2012 | Bard Access Systems | Bard Access Systems is recalling one lot of 16 Fr HemoStar XK Long Term Hemodialysis Catheters because they were identified with an incorrect expiration date, 03/2015 rather than 10/2013. | Incorrect or no expiration date | Bard Access Systems sent an "URGENT: *16 FR HEMOSTAR* LONG-TERM HEMODIALYSIS CATHETER RECALL" letter dated August 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Reply Form and Inventory Reconciliation Form was enclosed for customers to complete and return via fax to 1-801-522-5674. Contact BAS Customer Service at 1-800-290-1689 for questions regarding this recall. |  | 120 | 120 units | National US distribution only: MO, KS, AZ, OK, OR, AL, MA, TX, FL, MI, OH, LA, NJ, NC, WA, AR. | Terminated | January 07, 2013 | 153 | 120 | Not_Computer
Z-2275-2012 | 62679 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo.plaza, Model number 10592457. |  |  | July 18, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a potential malfunction in which a miscalculation occurs for images with pixel depth values greater that 12 bit in certain functions. | Software design | The firm initiated their recall/correction of this device on July 18, 2012 by sending a "Customer Safety Advisory Notice" letter, dated July 12, 2012, to their consignees. The letter identified the affected product along with the malfunction. It also discussed the steps the user can take to avoid the issue. Users were asked to not use the mentioned functions for images from those modalities until the issue was resolved and the system was updated with the software version containing the solution. Users were to contact their local Siemens Service in order to plan for the updated software version that will resolve the issue. Additionally, users should place the safety advisory notice in the system's instructions for use, and personnel should be instructed accordingly and maintain awareness. If the device was sold and is no longer in the users' possession, the notice should be forwarded to the new owner and the firm should be notified of the new owner. |  | 14 | 14 | Nationwide Distribution -- CA, FL, GA, MA, MN, MO, NJ, NY, OH, WA, and WI. | Terminated | January 24, 2014 | 555 | 14 | Computer
Z-2286-2012 | 61982 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is labeled in part: "***MEDLINE***Extended PTFE Needle 6.5" (16.51cm) Length***Reorder: ES0016***Contents: 1 Each***SINGLE USE ONLY**latexfree***RX Only***STERILE EO***www.medline.com Manufactured for Medline Industries, Inc. MEDLINE is a registered trademark of Medline Industries, Inc. 1-800-MEDLINE***Mundelein, IL 60060 USA Made in USA.***" |  |  | June 16, 2010 | 2010 | Unimed Surgical Products, Inc. | On 6/16/2010 Unimed Surgical Products, Inc., 10401 Belcher Road, Largo, FL 33777 initiated a recall of the Medline Industries, Inc. Electrosurgical electrodes used for cutting and coagulation of soft tissue due to concerns of sterility. | Process control | Unimed Surgical Products, Inc. issued a Medical Device Removal letter dated June 4, 2010 to Medline Industries, Inc. The letter identified the affected product, problem and actions to be taken. Medline was instructed to quarantine any undistributed product and notify its customers of the recall. For further information, contact Unimed Surgical Product, Inc. by phone 800-531-2627 or FAX 800-886-1225. |  | 1440 | 1440 ea | US Nationwide Distribution - including the state of IL. | Terminated | August 30, 2012 | 806 | 25440 | Not_Computer
Z-2281-2012 | 61171 | No | LGW |  |  | Neurology | PMA | The Eon Mini Neurostimulation (IPG) System (Model 3788) is a 16-channel, rechargeable, multi-program system that allows the connection of one or two leads with 4, 8, and 16 electrode leads. |  |  | December 19, 2011 | 2011 | Advanced Neuromodulation Systems Inc. | As of 11/30/2011, St. Jude Medical has received 110 patient complaints of warmth or heating at the implantable pulse generators (IPGs) implant site during charging for the Eon IPG and 116 reports of similar symptoms for the Eon Mini IPG, respectively. | Labeling design | ST. Jude Medical sent an Important Medical Device Correction letter dated December 19, 2011 to all affected customers. The letter identified the affected product, problem and recommendations to be followed. For questions contact your St. Jude Medical Neuromodulation Division Representative. |  | 30560 | 30,560 units | Nationwide within the US, and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Isreal, Italy, Japan, Mexico, Netherlands, Puerto Rico, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | July 24, 2015 | 1313 | 56948 | Not_Computer
Z-2283-2012 | 62712 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Siemens Stratus(R) CS Acute Care(TM) D-dimer Calibrator (DDM CalPak) lots 502058002 and 501318002 |  |  | June 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the Stratus(R) CS Acute Care(TM) D-Dimer CalPak calibrator lots 502058002 and 501318002 may exhibit SLOPE or CALCV error messages (instrument errors), which indicate calibration failures. | Nonconforming Material/Component | SIEMENS sent an Urgent Field Safety Notice dated June 7, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included a response form to be faxed to the recalling firm at 302-631-8467 by the consignee. If customers experienced calibration failures, they were instructed to contact the Siemens Technical Solutions Center at 800-405-6473 for assistance. Customers were also asked to forward the notification to anyone to whom they may have distributed the affected product. For questions regarding this recall call 302-631-6311. |  | 786 | 786 cartons | Nationwide Distribution - USA including AK, AL, AR, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, WA, WI, WV, and WY. | Terminated | January 16, 2014 | 588 | 786 | Not_Computer
Z-2251-2012 | 62715 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | OEC IT3000; IT2500 (EnTrak); IT2500 Plus (EnTrak Plus); IT3500 (InstaTrak); IT3500 Plus (InstaTrak Plus); ConneCTstat; and ConneCTstat Plus picture archiving and communications systems. The systems are an aid to locate anatomical structures during open or percutaneous surgical procedures. It is intended for use in medical conditions that may benefit from the use of stereotactic surgical technique. The system provides a reference to rigid anatomical structures such as sinus, skull, long bone, or vertebra, which are visible on medical images such as CT, MRI, or X-ray. |  |  | June 22, 2009 | 2009 | GE OEC Medical Systems, Inc | GE Healthcare Surgery had discovered that using the Inverted Headset Placement is not a validated configuration with the IT 3000, 2500, 2500 plus, 3500, 3500 plus, ConneCTstat, and ConneCTstat Plus Surgical Navigation equipment. | Labeling False and Misleading | GE Healthcare sent an Urgent Safety Advisory Notice dated June 22, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised not to use the ENT registration Headset in the inverted position for any CT scans or surgical procedures and to remove all references in their Naviation System Operator's Manual that discuss an inverted headset placement, and replace it with the enclosed updated instructions for the Navigation System Operator's Manual. For any questions customers should contact the service team for further information at 800-874-7378, option #8.For questions regarding this recall call 801-536-4516. |  | 768 | 768 units | Worldwide Distribution - USA (nationwide) and the countries of India, Japan, Canada, Switzerland, Italy, United Kingdom, Germany, Austria, Hong Kong, France, and Denmark. | Terminated | September 28, 2012 | 1194 | 768 | Not_Computer
Z-2253-2012 | 62769 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | ON-Q pump with ONDEMAND bolus button, ON-Q C-bloc with ONDEMAND; ON-Q C-bloc with Select a Flow and ONDEMAND Model/Reference #: 7000180, 7000181, 7000182, 7000367, 7000397, CB003, CB005, CB006, P400X1-7+5/60, P400X2-14+5/30, P400X5+5/60, and PMB01. |  |  | May 08, 2012 | 2012 | I-Flow Corporation | It was determined that in a small quantity of ON-Q pumps with ONDEMAND, the bolus button may not lock in the down position when depressed and/or the orange bolus refill indicator may stay in the lowermost position. As a result, the patient may receive continuous infusion at a rate greater than expected. | Device Design | An Important Voluntary Recall Notice letter dated 5/8/12 was sent to all their customers who purchased the ON-Q pump with ONDEMAND bolus button. The letter provided customers with the problem indicated and the actions to be taken. Customers were instructed to complete and sign the attached Recall Response Form and fax to (920)969-4565 or scan and email to IflowProductInquiry@kcc.com.Once I-Flow receives the completed FAX BACK form, an I-Flow Customer Service Representative will immediately contact the customers to provide instructions on product return and credit. Customers with questions regarding the recall letter, were instructed to contact I-Flow Customer Service between 8am to 5pm (PST) at (800) 443-3569. Or customers can contact I-Flow 24 hour Technical Support line at (800) 444-2728. |  | 80646 | 80,646 units total (65,926 units in US) | Worldwide Distribution -- USA, Canada, Australia, the EU (Germany), United Arab Emirates, and Cyprus. | Terminated | April 25, 2013 | 352 | 80646 | Not_Computer
Z-2279-2012 | 62880 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | V Series Patient Monitors; Mindray DS USA, Inc. |  |  | May 07, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified two software anomalies contained in the V Series Monitoring System. There have been no reports of injuries associated with these anomalies. These anomalies affect V Series Systems distributed between September 31, 2010 and January 9, 2012. | Software design | Mindray sent an Urgent: V Series Monitoring System Correction letters dated May 7, 2012 via certified mail to all affect customer. The letter identified the affected product, problem and actions to be taken. Customers are directed that VPS modules should not be used between systems until a software upgrade has been completed. This work will be done at no cost to the customer. For questions contact your Mindray Serviice Representative at 1-800-288-2121. |  | 979 | 979 units | Worldwide Distribution - US Nationwide including the states of: NY, SD, NE, IL, MO, CA, OH, OK, WA, IN, WY, CO, TX, FL, NC, AL and Puerto Rico. and the countries of: Australia, Canada, Columbia, New Zealand, Pakistan, Russia, Saudi Arabia, United Kingdom and Venezuela. | Terminated | July 17, 2013 | 436 | 979 | Computer
Z-2285-2012 | 62572 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | PHADIA Varelisa Cardiolipin IgM |  |  | August 02, 2012 | 2012 | Phadia US Inc | Customers were not following the labeled assay procedure. This was due to a letter issued to them in 2004 from Phadia GmbH, which indicated that customers could omit the pre-wash step. The omission of the pre-wash step, when combined with a change in material source in kit lot #0067, caused elevated results. | Labeling Change Control | On 8/2/2012 an URGENT PRODUCT CORRECTION INFORMATION FOR PHADIA VARELISA CARDIOLIPIN IGM, Article Number 15696 letter was sent to all consignees the received the Direction for Use modification letter in 2004. |  | 67254624 | 837 kits (96 assays per kit), total of 80,352 individual assays | USA Nationwide. | Terminated | December 12, 2012 | 132 | 67254624 | Not_Computer
Z-2284-2012 | 62776 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | QKC1692E |  |  | August 10, 2012 | 2012 | Steris Corporation | STERIS has been actively seeking devices for quick connect revalidation projects. STERIS was unable to obtain two (2) of the needed devices; these two devices represent the design of seven (7) models in one quick connect. Due to the unavailability of these two validation devices, STERIS will remove 7 models from the labeling of QKC1692E and will re-label all QKC1692Es in the field.. | Device Design | The firm, Steris, sent a "Voluntary Field Correction: Re-labeling of QKC1692E" letter on August 10, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter to re-label the affected QKC1692E(s) and return the completed form by September 30. 2012 to: Regulatory Affairs Specialist via Kate_Beveridge@STERIS.com or via fax at (440) 350-7082.If you have any questions, contact Director, Regulatory Affairs at 440-392-7519. |  | 64 | 64 units | Nationwide distribution: USA including states of: AR, AK, CA, CO, DE, FL, IL, MA, MD, NC, NE, NH, NY, OH, PA, SD, TN, TX, WI, and WV. | Terminated | July 05, 2013 | 329 | 64 | Not_Computer
Z-2282-2012 | 53055 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Whole Body X--ray Scanner |  |  | July 13, 2009 | 2009 | Toshiba American Medical Systems Inc | 1. It has been found that in raw data processing for Go&Return; helical scanning, when the raw data acquired in the first scan (Go) and that acquired in the second scan (Return) are selected and reconstructed at one time, incorrect couch positions and scan times are displayed on the reconstructed images for the data acquired by Return scanning, while all related information is displayed on the reco | Software design | The firm, Toshiba America Medical Systems, Inc., sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated July 30, 2009 to its Consignees/customers. The letter described the product, problem and actions to be taken. The letter informed them of the above mentioned reason for recall and telling them that a corrective measure will be installed to prevent the occurrence of the above problems. They were told that their Toshiba service representative will contact them to install the measures.Furthermore Toshiba instructed them to adopt the following temporary measures to avoid these phenomena until corrective measures have been taken.1. When performing raw data processing for Go&Return; helical scan, select the raw data acquired for Go and that acquired for Return separately in order to perform reconstruction2. When using the Sure Exposure 3D function, confirm that the error message shown in the preceding page is not displayed on the monitor at the time of setup (until the Scan Start button lights after the Confirm button is clicked).3. When executing an eXam Plan in which Sure Exposure 3D is selected as the actual scan for Sure Start, confirm that the error message show in the preceding page is not displayed on the monitor when monitoring scan is switched to the actual scan. If the error message is displayed, click [OK] in the error message dialog to continue processing.The letter should be shared with facility's Biomedical or Clinical Engineering Department. Also the customers were instructed to complete and return the Customer Reply Form via fax to: Regulatory Affairs at 877-349-3054.If you have any questions regarding this letter, contact Director, Regulatory Affairs at (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968 with any questions regarding the letter. |  | 24 | 24 Systems | Nationwide distribution: USA including states of: AL, AR, CA, CO, FL, GA, KY, MA, MD, MT, ND, NJ, NV, NY, OH, PA and Puerto Rico. | Terminated | September 04, 2012 | 1149 | 24 | Not_Computer
Z-2278-2012 | 62853 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | EXCELART Vantage, Atlas, and Titan MRI Systems, MRT-1503 and MRT-1504. |  |  | May 13, 2008 | 2008 | Toshiba American Medical Systems Inc | The status display in the sequence queue window may not change from "Current" to "Done". | Process design | A recall notification letter titled: "URGENT: MEDICAL DEVICE CORRECTION," was sent to customers. The letter stated the problem, corrective action, and (instructions) request to customers. Customers were instructed to perform the temporary recovery procedures if the problem occurs. Additionally, the acknowledgement sheet should be completed and faxed back. The information should be shared with all users, including the bio-medical or clinical engineering departments. Questions pertaining to the information for this recall were to be directed to (800) 421-1968. |  | 54 | 54 | Nationwide Distribution | Terminated | September 04, 2012 | 1575 | 54 | Not_Computer
Z-2300-2012 | 49060 | No | FTN | Plastic surgery kit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Various SRI Surgical Disposable Accessory Packs. |  |  | July 16, 2008 | 2008 | SRI Surgical | Distribution of a product that did not meet specifications. (labeling error) | Component change control | SRI sent a "MEDICAL DEVICE CORRECTION" letter dated July 16, 2008 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 813-891-9550 ext. 3124 for questions regarding this recall. |  | 802135 | 802,135 | Nationwide Distribution-including the states of CA, FL, GA, KY, MD, MI, NC, NJ, OK, PA, SC, TN, TX, and UT. | Terminated | September 10, 2012 | 1517 | 802135 | Not_Computer
Z-2301-2012 | 62705 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DGPHP RFA High-Power Single Use Grounding Pads and Cool-Tip RFA Electrode Kits (containing DGPHP RFA High-Power single use grounding pads) |  |  | August 01, 2012 | 2012 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a recall of various production lots of DGHP RFA High Power single use grounding pads and Cool-tip RFA Electrode Kits due to the potential for foil degradation on the DGP-HP single use grounding pad which may result in an electrical burn to the patient at the application site. | Process control | The firm, Covidien, sent an "URGENT MEDICAL DEVICE RECALL" letter dated August 6, 2012, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine and discontinue use of the device; immediately advise all DGPHP RFA High-Power single use grounding pads and Cool-tip RF Electrode kit users of this recall; complete and return the attached Recalled Product Return Form in its entirety via fax to (800) 895-6140; return affected product to Covidien Attn: Field Returns Department, 195 McDermott Road, North Haven, CT 06473; if purchased directly from Covidien include completed form via fax or by email to : feedback.customerservice@covidien.com; if purchased from distributor complete form and contact your Distributor directly (completed form and all affected units must be returned through the Distributor); and reply to Covidien WHETHER OR NOT you have affect product at your site. If you have any questions or concerns, please do not hesitate to contact your Covidien representative or Quality Assurance, Monday through Friday, 8:30-5:00, directly at 203-492-5232. |  | 184934 | 184,934 units | Worldwide distribution: USA (nationwide) and countries including: Australia, Brazil, Canada, Columbia, Deutchland, France, Japan, Mexico, Panama, Puerto Rico, Singapore, and Uruguay. | Terminated | September 06, 2013 | 401 | 184934 | Not_Computer
Z-2304-2012 | 60514 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX (with FLEX software) and ABL80 FLEX CO-OX (with CO-OX software) - pH Blood Gas and Electrolytes Analyzer System. |  |  | June 06, 2012 | 2012 | Sendx Medical Inc | The firm recalled because the glucose measurements from a patient sample that the customer felt was too low and did not reflect the patient's clinical condition. | Software design | Radiometer sent a Field Safety Notice letter dated June 06, 2012 via email to all affected customers. The letter identified the affected product, problem and actions to be taken. Also included is a Note to Users to be placed at the front of customers existing Operator's Manual. Customers were instructed to provide confirmation to Radiometer that all customers have received the letter and the Note to Users. For questions contact your Radiometer representative. |  | 3599 | ABL80 FLEX is 1680, ABL80 CO-OX is 1759 | Worldwide Distribution - US Nationwide including the state of OH. | Terminated | September 04, 2012 | 90 | 3599 | Not_Computer
Z-2302-2012 | 62801 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert - HXLPE-VE, Sterile, Material: Highly cross linked vitamin E PE, Sz 2/9mm. |  |  | August 09, 2012 | 2012 | Encore Medical, Lp | The firm is recalling one lot of 392-09-706, 3DKnee e+ Tibial insert, 9 mm, Size 6, Right. The shelf life labeling of specific lots for a component were found to be incorrect and were found to be expired. | Incorrect or no expiration date | DJO Global sent an Urgent Field Safety Notice dated August 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Notice to all those who need to be aware within their organization or to any organization where the potentially affected product has been transferred. Customers were instructed to contact the firm to place a replacement order and receive an RMA number. For questions customers were asked to call 512-834-6255.For questions regarding this recall call 512-832-9500. |  | 7 | 7 units | Worldwide Distribution including Venezuela. | Terminated | January 25, 2013 | 169 | 7 | Not_Computer
Z-2303-2012 | 62865 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips SureSigns VSi/VS2+ patient monitors |  |  | August 15, 2012 | 2012 | Philips Healthcare Inc. | Philips SureSigns VSi/VS2+ patient monitors may not be properly secured and may fall from the wall mount | Process design | Philips sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated August 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Customers using the VSi/VS2+ with the GCX wall mount are instructed to complete and fax a Customer Reply Wall Mount Order Card to the Business Unit or Key Market indicating the number of wall mounts in use. Upon receipt of the Customer Reply Wall Mount Order Card, Philips will send a replacement of the mounting hardware. Customers should contact their local Philips representative for questions regarding this recall. |  | 2314 | 2314 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, DENMARK, EL SALVADOR, FRANCE, GERMANY, ICELAND, INDIA, IRELAND ISRAEL, ITALY, JORDAN, KENYA, LEBANON, MALAYSIA,NAMIBIA NETHERLANDS, NORWAY, POLAND, PORTUGAL, RWANDA SAUDI ARABIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | March 13, 2014 | 575 | 2314 | Not_Computer
Z-2305-2012 | 62709 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | syngo(R) Lab Data Manager and syngo(R) Lab Process Manager, all versions. |  |  | May 29, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | When using the syngo Quality Control package the system may not perform as intended for Multi-Rule QC violation [e.g., 2(2S), 4(1s) and 10x], unless configured using the instructions on page 2 of the recall notification. QC status may not be flagged appropriately. If auto-verification rules are defined to hold patient results when QC status is out, the patient results will not be held as expecte | Software design | Siemens sent a Urgent Field Safety Notice letter dated May, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed the customers to re-set the Multi-Rule option and instructions are provided on how to do so. This recall notification also included a Field Safety Effectiveness check form to please fax to the Siemens Technical Solutions Center at 302-631-8467 BY June 15th. |  | 44 | 44 | USA ( nationwide ) including the states of AZ, CA, FL, IL, KY, MD, MA, MN, MS, MO, NE, NH, NJ, NY, NC, OH, PA, TX, VA and WV. | Terminated | April 10, 2014 | 681 | 44 | Not_Computer
Z-2307-2012 | 62537 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Poly stat Flu A &amp; B Test |  |  | March 09, 2012 | 2012 | Princeton Biomeditech Corp | There have been reports of dual positive (Flu A + Flu B positive) results being encountered. | Nonconforming Material/Component | OraSure Technologies, Inc. sent an "IMPORTANT NOTICE FOR ORASURE QUICKFLU RAPID FLU A + B TEST" letter dated March 13, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact the firm at 610-882-1820 for questions concerning this recall. |  | 103 | 103 kits (22 tests per kit) | Nationwide Distribution-including the states of NJ, NY, and PA. | Terminated | October 14, 2014 | 949 | 563 | Not_Computer
Z-2260-2012 | 62848 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | V60 Ventilator, Continuous Ventilator, Model# V60, PN 1056490 |  |  | July 31, 2012 | 2012 | Respironics California Inc | Respironics is recalling the V60 ventilator because certain blower motor assemblies in the V60 ventilators may not meet specifications. Specifically, the impeller may not be properly seated on the blower motor shaft. This could subsequently cause the blower to cease functioning which would result in the failure of the V60 ventilator to deliver therapy to the patient due to the loss of ventilatio | Nonconforming Material/Component | Respironics California a division of Philips Healthcare sent a Medical Device Correction letter dated July 31, 2012 to all affected customers. The letter identified the affected product, problem identified and actions to be taken. The letter states that a Phillips Respironics Field Service Engineer or distributor will be contacting customers to schedule a no-cost replacement of the blower motor assembly on all affected V60 ventilators. For information or support concerning the recall issue, contact Philips Respironics at (800) 722-9377. |  | 197 | 197 units total (116 in the US) | Worldwide Distribution - US (Nationwide) and the countries of: Australia, Belgium, Canada, China, Germany, Italy, Japan, Kuwait, Netherlands, Norway. Qatar, Saudi Arabia, Turkey, and United Kingdom. | Terminated | March 22, 2013 | 234 | 197 | Not_Computer
Z-2309-2012 | 61451 | Yes | NIO |  |  | Cardiovascular | PMA | Zilver 518 Vascular Self Expanding Stent |  |  | January 05, 2012 | 2012 | Cook Ireland Ltd. | Cook Ireland received a product complaint from Japan on December 22nd, 2011. The complainant reported that a 5 French Zilver Vascular Stent (part number ZIV5-18-125-8-80 from Lot# CF694429) was inserted into a sheath over a wire guide where it became lodged in the sheath hub. After further examination, the user determined that the product was a 6 French Zilver Vascular Stent (part number ZIV6-35 | Packaging process control | The firm, Cook Medical, sent a letter dated February 13, 2012 to foreign consignee. The letter described the product, problem and actions taken. The firm stated that re-training of the relevant personnel has been carried out; they will continue to monitor to identify potential emerging trends and further corrective and preventive action will be considered in an effort to avoid the recurrence of similar events.If you have any question contact in US: Customer Relations Manager at 800.346.2686.ext 2222 or email: rita.harden@cookmedical.com; OUS: Customer Quality Supervisor at +353 61 239343. |  | 15 | 15 units of ZIV5-18-125-8-80 LOT# CF694429 | Worldwide distribution: USA (nationwide) and country of: Japan. | Terminated | February 04, 2013 | 396 | 15 | Not_Computer
Z-2250-2012 | 62856 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 and HT70 Plus Ventilators, Model Number: HT70 and HT70 Plus. |  |  | August 03, 2012 | 2012 | Newport Medical Instruments Inc | May emit a continuous high priority alarm and the ventilator may stop ventilating, due to a component failure on the control board. | Nonconforming Material/Component | Newport Medical sent an Urgent Field Correction Notice letter dated August 3, 2012 to all affected customers, via DHL and FedEx courier services. The letter identified the affected product, reason for recall and actions to be taken. Customers were instructed to immediately stop using the affected ventilator and retrieve the replacement kit provided free of charge. In addition, the letter states that if the product was further distributed, to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Replacement form as replacements are completed. For questions contact Newport Medical Technical Support Department at 800-451-3111, extension 282, for U.S. customers; 714-427-5811, extension 500, for international customers.". |  | 72 | 72 | Worldwide distribution: USA (nationwide) and countries of: Brazil, Chile, Egypt, France, Hong Kong, India, Japan, Norway, Philippines and South Africa. | Terminated | April 14, 2014 | 619 | 72 | Not_Computer
Z-2310-2012 | 61984 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | OPTETRAK LOGIC TIBIAL INSERT POSTERIOR STABILIZED***Match Component Colors 6 YELLOW, 11mm, Use with LOGIC Femoral and Tibial Components***REF 02-012-35-6011. |  |  | August 19, 2011 | 2011 | Exactech, Inc. | Exactech, Inc. of Gainesville, FL is recalling their Optetrak Logic Posterior Stabilized Tibial Insert, Size 6 after the device was determined to be out of dimensional specification. | Device Design | Exactech, Inc., sent an "IMPORTANT PRODUCT MARKET WITHDRAWAL NOTICE" letter dated August 19, 2011, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to: 1. Immediately cease distribution or use of the products. 2. Extend the information to accounts that are in possession of the products. 3. Verify if they had any of the affected products on the list. 4. Fax back the attached form. In addition, they were instructed to contact the Exactech inventory representative to confirm quantities at their location. If you have any questions regarding inventory restocking issue, please call -800-392-2832. |  | 193 | 193 | USA ( nationwide) WA, OH, OK, TX, NC, NY, CO, VA, MA, IL, TN, MN, NJ and FL. | Terminated | November 07, 2013 | 811 | 193 | Not_Computer
Z-2312-2012 | 62982 | No | IZD | Nuclear uptake probe. | Radiology | Radiology | 510(K) Exempt | Capintec Captus 3000 Thyroid Uptake System |  |  | July 24, 2012 | 2012 | Capintec Inc | Customers who received a Capintec Captus 3000 Thyroid Uptake System with Software Revisions 1.22, 1.23, 1.24, 1.25, and 1.26 may encounter an error while using the thyroid uptake protocol ?Measure Liquid and Subtract Residual." This protocol was a new feature added to Software Revision 1.22 in January 2011. The error occurs only if the user does NOT measure the residual dose immediately after me | Software design | The firm, CAPINTEC, INC., sent a "SOFTWARE ERROR ALERT" letter dated July 20, 2012 to all affected consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the four measuring steps outlined in the letter; DO NOT PRINT, EXIT THE THYROID UPTAKE MODULE, OR MEASURE ANOTHER PATIENT BEFORE YOU HAVE COMPLETED THESE FOUR MEASUREMENT STEPS; and follow the attached addendum for step-by step instructions. Capintec will send corrected software to all affected customers in approximately 2-3 weeks. The customers are instructed to contact Capintec Customer Support Representative if you have any questions or concerns regarding this notification, or if you do not receive your replacement software. Capintec Customer Support can be reached by phone at 1-800-631-3826 or by email at getinfo@capintec.com or by visiting our website at www.capintec.com. |  | 256 | 256 | Worldwide distribution: USA (nationwide) and countries of: Argentina, Belgium, Brazil, Canada, China, Greece, Hong Kong, India, Israel, Malaysia, Philippines, Spain, South Korea, Thailand, and Turkey. | Terminated | December 10, 2013 | 504 | 256 | Computer
Z-2311-2012 | 62593 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Acrobat 3000 (AC3000) Spring Arm as support arm for iLED operating lights, TRUMPF Medizinsysteme GmbH + Co KG 82178 Puchheim. |  |  | May 29, 2012 | 2012 | Trumpf Medical Systems, Inc. | Cracks can arise on the welded seam on the rear joint. | Process control | TRUMPF sent an Urgent Safety Information letter dated February 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that within their organization, all users of the affected product be informed about the Urgent Safety Information. If affected product was further distributed, customers should forward a copy of the information or inform the necessary parties. Customers were asked to confirm receipt of the Urgent Safety Information by completing and returning the confirmation of receipt.For questions regarding this recall call 888-474-9359. |  | 402 | 402 units | Worldwide Distribution - USA including CA, OH, OK, NY, PA, TX, WY and the country of Canada | Terminated | April 11, 2013 | 317 | 402 | Not_Computer
Z-2315-2012 | 62928 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Toshiba Ultra-Short Magnetic Resonance Imaging System, EXCELART Vantage. MRT-1503/P5-Vantage MRI System. |  |  | February 21, 2012 | 2012 | Toshiba American Medical Systems Inc | The firm recalled due to a potentail problem with the position of the hold in the lug teminals on the power cables connected to the terminal blocks on the gradient coil. | Component change control | Toshiba America Medical Systems, Inc, sent a Urgent Medical Device Correction letter dated February 21, 2012, via USPS return receipt mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed if you subsequently find that images are abnormally noisy compared to images acquired under the same conditions in the system before this modification was performed, contact your Toshiba Service Representative.Please share this information with all users and reviewing radiologist as well as clinical engineering or biomedical group at your facility. Please complete and return the attached form and fax it to the toll free number at the top of the form. This form can also be sent via email toraffairs@tams.com.Contact information was also provided (800) 421-1968 or contact your local Toshiba Representative at (800) 521-1968. |  | 3 | 3 units | US (nationwide) distribution including the states of Ohio and Texas. | Terminated | September 07, 2012 | 199 | 3 | Not_Computer
Z-2316-2012 | 62505 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm MRSA ACP Assay Catalog #441637, box 48 tests labeled in part***GeneOhm Sciences Canada, Inc., 2555 boul, du Parc-Technologique, Quebec, QC, Canada G1P 4S5*** |  |  | May 22, 2012 | 2012 | Becton Dickinson &amp; Co. | Leakage in Cepheid SmartCycler reaction tubes. | Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter via UPS overnight delivery on May 22, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:Discontinue use and discard any affected product. BD would send replacement tubes for the discarded material.Complete the attached form whether or not they have any inventory remaining so that the firm may acknowledge their receipt of the notification.Fax the form to:Attention: Regulatory ComplianceFAX: 410-316-4258Email: RegulatoryComplianceFax@bd.comIf further assistance is needed, customers should contact BD Technical Services Department at 1-800-838-8663.For questions regarding this recall call 410-316-4054. |  | 38 | 38 kits | Nationwide Distribution including CA, DE, FL, GA, IL, IN, KS, KY, ME, MT, NC, NJ, NM, NY, OH, OK, PA, SC, TX, VA, WA, WI, and WV. | Terminated | November 02, 2012 | 164 | 200 | Not_Computer
Z-2321-2012 | 62891 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AmeriWater MRO Dialysis RO+ System, AmeriWater 1303 Stanley Avenue, Dayton, OH 45404 (800) 535-5585 www.AMERIWATER.com. |  |  | August 09, 2012 | 2012 | AmeriWater Inc | The firm was notified that a malfunction of the device was caused by an incorrect fuse 32 volt 15 amp fuses instead of 250 volt 15 amp fuses, that was installed in the water system. The incorrect fuse may result in excessive heat in the fuse holder eventually resulting in failure of the device to operate. | Nonconforming Material/Component | The firm, AmeriWater, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2012, to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the product subject to recall; remove affected devices from service until the correction is completed to prevent malfunction of the device during patient treatment; and complete and return the attached response form to AmeriWater within 30 days via fax to 937/461-1988 Attn: Quality Assurance. AmeriWater dealer will be contacting their customers to schedule replacement of the fuses.If you have any questions, contact your local AmeriWater dealer or AmeriWater customer service at 1-800-535-5585. |  | 80 | 80 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, CO, CT, FL, GA, IL, IN, KS, MA, MD, MO, MS, NC, NY, OH, OK, SC, TN, TX and VA; and country of: Canada. | Terminated | April 17, 2014 | 616 | 80 | Not_Computer
Z-2322-2012 | 63042 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett Mira. |  |  | August 21, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, Siemens has become aware of a potential issue with the Mobilett Mira system. During normal clinical operation the switch on the motor drive handle of the unit may become loose and slide horizontally to the right and become trapped under the cover. This may cause the system to continuously move. In response, Siemens has issued Update Instruction XP038/12/S to r | Device Design | Siemens initiated their correction of the MOBILETT Mira on August 21, 2012 by sending an "UPDATE INSTRUCTIONS XP038/12/S to all affected customers. The Instruction identifies the product, problem, and work steps for customers to follow. A XP038/12/S Completion Protocol letter was included for customers to complete and file with the instruction once the update is completed. |  | 20 | 20 | Nationwide Distribution- including the states of GA, ID, IL, IA, KS, NC, OH, SD, TX, and WI. | Terminated | April 03, 2014 | 590 | 20 | Not_Computer
Z-2326-2012 | 58247 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Polar Dressing, Knee/MU/Shldr Intelli-Flo-Breg, PN 10630 |  |  | August 11, 2010 | 2010 | Breg Inc, An Orthofix Company | This recall has been initiated due to confirmation that the Polar Dressing, Knee/MU/Shldr Intelli-Flo-Breg does not meet the required thickness as specified. The dressing in the sterile pouch measures a thickness of 0.010" instead of 0.020". This change in thickness makes the dressingthinner, less durable and subject to tearing. The product is mislabeled. | Nonconforming Material/Component | The firm, Breg Orthofix, called their customers on August 11, 2010 sent a letter dated August 12, 2010 and emails to all customers. The call and letter described the product, problem and actions to be taken. The customers were instructed to take immediate action to identify and quarantine the product at their facility. Breg established a return materials authorization and is contacting their customers to provide a Return Authorization to return the dressings. Breg is also providing replacement dressings. If you have any questions regarding communication or the Intelli-Flo Sterile Dressing, contact Breg Customer Relations at (800) 321-0607. |  | 502 | 502 units | Worldwide distribution: USA (nationwide) and country of: Nova Scotia. | Terminated | September 06, 2012 | 757 | 502 | Not_Computer
Z-2313-2012 | 62702 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Multi-Diluent 1 - Used for automated dilutions for the ADVIA Centaur BNP and TSH3 Ultra assays when on-board the ADVIA Centaur or ADVIA Centaur XP systems. |  |  | July 25, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Decrease in onboard dilution recovery when using Multi-Diluent 1 that has been stored on board the ADVIA Centaur and ADVIA Centaur XP systems for for BNP and TSH3 Ultra assays | Nonconforming Material/Component | Siemens sent an "URGENT DEVICE RECALL NOTICE" dated July 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Response Form was included for customers to complete and return via fax to (302) 631-7597. Contact your local Sales Representative for questions regarding this notice. |  | 7414 | 2 Packs = 15879 and 6 Packs = 7414 | Worldwide Distribution-USA (nationwide) and the countries of Canada, Argentina, Australia, Brazil, Chile, China, Croatia, Denmark, Ecuduar, Egypt,Guatemala, Hong Kong,India,Ireland, Israel,India, Japan,Korea, Mexico, Malyasia, Peru, Pakistan, Saudi Arabia,Singapore, Thailand,Uruguay, Viet Nam, and South Africa. | Terminated | November 04, 2015 | 1197 | 7414 | Not_Computer
Z-2314-2012 | 62949 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | TxRx Knee 15 Coil 3T. |  |  | July 27, 2012 | 2012 | Quality Electrodynamics LLC | On July 27, 2012, it was reported to the firm that the Siemens 3T 15ch Knee Coil was not functional because the coil identification code for the device was incorrect. | Error in labeling | Since all of the devices were located at various Siemens distribution locations, all communications were internal to Siemens concerning bringing the coils from one warehouse to another. |  | 17 | 17 units | Worldwide distribution: USA (nationwide): Tennessee only and country of: Germany | Terminated | February 19, 2013 | 207 | 17 | Not_Computer
Z-2323-2012 | 62757 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | VersaTREK Windows Software, Version 5.4.3 is a component of the VersaTrek System. The Software CD will be individually packaged using a standard CD plastic case and placed into a standard CD mailer |  |  | July 23, 2012 | 2012 | Trek Diagnostic Systems | Review of instrument software databases and investigation of two customer reports determined that on rare occasions (less than 0.3%), following the removal of a positive signaling bottle from the VersaTREK instrument, the positive test result inadvertently reverted to a negative result in the instrument database. | Software design | ThermoFisher Scientific sent an Important Medical Device Product Correction Notice dated July 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact the Technical Service Department at 800-642-7029 to schedule their upgrade and to complete and return the Correction Notice Customer Acknowledgement by fax or in the enclosed return envelope. Customers were also asked to return the checklist. For questions customers were instructed to call 800-642-7029.For questions regarding this recall call 913-985-4185. |  | 333 | 333 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO , CT, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, RI, SD, TN, TX, VA, VT, WA, WI, WV, and WY. Internationally to the countries of Bahamas, Bangladesh, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, India, Ireland, Italy, Japan, Jordan, Kenya, Kuwait, Malaysia, Mexico, Pakistan, Peru, Philippines, Portugal, Puerto Rico, Russian Federation, South Africa, Spain, Thailand, Turkey, and United Kingdom | Terminated | June 18, 2013 | 330 | 333 | Computer
Z-2325-2012 | 62852 | No | LRT | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 100 ft roll of tubing inside a cardboard box. Medical devices are sealed into the tubes prior to ethylene oxide sterilization. Tube maintains sterility of the enclosed medical devices until opened. |  |  | August 02, 2012 | 2012 | Steris Corporation | On 7/20/2012 the firm became aware that a lot of tubing was distributed by the supplier after it was rejected by Steris due to failing process indicators. | Nonconforming Material/Component | Steris sent a Urgent Voluntary Recall Notice letter dated August 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Steris has learned that the VIS-U-All tubing is not performing to product specifications. Upon receipt of replacement product we ask that you destroy any remaining inventory of the affected lot in your possession. In the interim period, should you receive a failing result following a sterilization cycle, the results must be handled in accordance with your internal hospital procedures.We apologize for any inconvenience this matter may cause you. For further questions please call 1-(800) 548-4873. |  | 77 | 77 packs | Worldwide Distribution--USA (nationwide) including the states of AL, AZ, CA, FL, IL, IN, KY, LA, MO, MS, NJ, OH, PA, SD, TX, and VA., and the countries of Japan, Peru, and Suriname. | Terminated | September 04, 2013 | 398 | 77 | Not_Computer
Z-2324-2012 | 62967 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Brand name: RayStation. Generic name: Treatment Planning System. Software version and build numbers are 2.5.0.144 and 2.5.1.89. Product lifetime is 5 years. |  |  | May 09, 2012 | 2012 | RAYSEARCH LABORATORIES AB | The firm decided to recall when an issue was found with the dose calculation in RayStation 2.5, Raystation 2.5.1. The problem has not caused any patient mistreatment or other incidents. However, the user must be aware of this information to avoid incorrect dose calculations during treatment planning. The dose calculation is not correctly normalized in the unusual case of preventing the optimizatio | Device Design | The firm, RaySearch Laboratories, sent an "Urgent Field Safety Notice, Medical Device Correction" letter dated May 9, 2012 to its customers. The letter describes the product, problem and actions to be taken. The letter informed the customers that the "workaround is to not use segment MU optimization" and that there "is in general no need to manually adjust the segment shapes to such an extent there is need for a separate optimization of the segment MU." Additionally, the customers were instructed -Do not uncheck the "Segment shapes" check box; educate planning staff and all users about this workaround; inspect your product and identify all installed units with the noted software version numbers; and complete and return the Reply Form via fax to: RaySearch Americas Inc. at 888-501-7195; email: freddie.cardel@raysearchlabs.com and/or email to: Raysearch Laboratories AB at niclas.borglund@raysearchlabs.com.For further information and return of the Reply Form contact For the Americas: Director of Customer Support, RaySearch Americas at +1877 778 3849 or freddie.cardel@raysearchlabs.com; and For Europe, Asia and rest of the world: Director of Service, RaySearch Laboratories AB at +46 8 5450 6130 or niclas.borglund@raysearchlabs.com. |  | 19 | 19 units | Worldwide distribution: USA (nationwide) including states of: FL, MA, NJ, and NY; and countries of: Canada Germany, Netherlands and Switzerland. | Terminated | January 25, 2013 | 261 | 19 | Computer
Z-2338-2012 | 62925 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassette, Batch # 12A078. |  |  | June 19, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling (removing) the Sterrad 100S and Sterrad NX System cassettes because the camera used to verify printing on the Sterrad System cassettes was inadvertently disabled. The STERRAD 100S and STERRAD NX System cases were shipped without verification of proper printing. | Error in labeling | Advanced Sterilization Products sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL) letter dated July 16, 2012 to all their customers who purchased the STERRAD 100S and STERRAD NX System cassettes. The letter informed the customers of the problem identified and the actions to be taken. For the STERRAD NX System cassettes, there should be no impact to the functionality of these affected cassettes. Although printing may be smeared, illegible, or missing, the packaging has multiple references to the information, and the cassettes should be usable despite any printing error.For the STERRAD 100S System cassettes, there is a chance that the barcode will be unreadable and the STERRAD 100S Sterilizer will not accept the cassette.Customers were instructed to complete and return the business reply card and packing slip that are included in the letter. Customers were instructed to contact Stericycle at (888) 406-9303 for additional labels. Customers with questions about the recall letter were instructed to contact Stericycle at (888) 406-9303. |  | 7550 | 7,550 units | Nationwide Distribution. | Terminated | November 08, 2013 | 507 | 15045 | Not_Computer
Z-2335-2012 | 62895 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | The SYSTEM 1E Liquid Sterilant Processing System. |  |  | August 01, 2012 | 2012 | Steris Corporation | The firm initiated a recall after complaint investigations revealed that when the units are left on and not in use at night and/or on weekends the hose disconnects/separates resulting in water leakage that can cause damage to property. | Nonconforming Material/Component | STERIS sent a Urgent Voluntary Field Correction Notice letter dated August 1, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.STERIS has received no reports of injuries associated with the disconnection of a SYSTEM IE water hose.STERIS Corporation will instill new hoses and connections on your SYSTEM IE Liquid Chemical Sterilant Processing System. For further information or if you have any questions regarding our visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 6051 | 6,051 units | Worldwide Distribution -- USA (nationwide) and the countries of Bahrain and Hong Kong. | Terminated | June 06, 2014 | 674 | 6051 | Not_Computer
Z-2334-2012 | 61869 | No | LEJ | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SUN¦DASH RADIUS 252 tanning system; 52 lamps, Ultraviolet Radiation. |  |  | July 08, 2005 | 2005 | JK Products &amp; Services, Inc | The firm has tracked a systematic failure in the door cord on the Sun¦Dash Radius 252 product line. In some instances when the door cord is installed with little or no slack outside the upper decor piece, the cord has been known to fail causing a short in the wiring at the pivot point in the cord. | Equipment maintenance | JK Global Services sent out Mandatory Inspection/Modification letters to their Distribution Service Teams dated 07/08/2005, 01/12/2006, 09/15/2010, 10/04/2010 and 03/01/2012. A letter dated September 15, 2010, was sent to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to contact a JK representative or their Distributor as soon as possible to schedule an inspection to the affected product in their possession. For any questions customers were instructed to call 1-800-445-0624 or send an e-mail to support@jkamerica.com. |  | 1948 | 1948 units | Worldwide Distribution - USA (nationwide) and Canada | Terminated | January 13, 2017 | 4207 | 1948 | Not_Computer
Z-2332-2012 | 62768 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory |  |  | March 15, 2012 | 2012 | Sunquest Information Systems, Inc. | Sunquest is recalling Sunquest Laboratory versions 6.1.0125 through 6.1.0139, 6.2.1, 6.2.2, 7.0, and 7.0.1 because when the Windows-based merge application is used and the "Event merge across patients" option is used, a receiving system (for example a HIS or CoPath system) may interpret the event merge as a full patient merge. This occurs because a trigger from the Laboratory system for the event | Software design | The firm, Sunquest, sent an "Urgent- Product Safety Notice: PSN-12-02" dated March 15, 2012 to all their customers who purchased the Sunquest Laboratory versions 6.1.0125 through 6.1.0139, 6.2.1, 6.2.2, 7.0, and 7.0.1. The notice described the product, the problems and the actions to be taken. Customers were instructed to request the software correction by creating a Case via the Client Support webpage at: www.sunquestinfo.com/supportweb. Customers with questions were instructed to contact Client support at (877) 239-6337 for US and Canadian Sites and at +44 (0) 161 335 0562 for international sites. |  | 126 | 126 sites | Worldwide distribution: USA (nationwide) and countries of: England, United Kingdom and United Arab Emirates. | Terminated | October 02, 2012 | 201 | 126 | Computer
Z-2336-2012 | 62970 | No | EWL | Laryngeal prosthesis (Taub design). | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | BLOM-SINGER indwelling TEP Occluder, catalog IN2004-TO |  |  | March 08, 2012 | 2012 | Helix Medical LLC | Helix Medical LLC is recalling the BLOM-SINGER indwelling TEP Occluder due to an invalid expiration date. | Labeling mix-ups | A recall letter dated March 8, 2012, was sent to customers who purchased the Blom-Singer Rapid Response Voice Prosthesis TEP Occluder 4mm 20Fr. (IN2004-TO). The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to contact Customer Service Center at (800) 477-5969 if they have products to be returned. Customers were instructed to return Blom-Singer Rapid Response Voice Prosthesis TEP Occluder 4mm 20Fr. (IN2004-TO) to:Helix Medical, LLCAttn to: Customer Service-IN2004-TO.1110 Mark AvenueCarpinteria, CA 93013For questions regarding this recall call 805-684-3304. |  | 2 | 2 devices | Worldwide distribution: USA (nationwide) including state of: Ohio and country of United Kingdom | Terminated | April 02, 2013 | 390 | 2 | Not_Computer
Z-2331-2012 | 62875 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CMS-2981-R3 Arthrogram Tray containing 2 oz. bottle PVP Prep Solution |  |  | September 30, 2011 | 2011 | Custom Medical Specialties, Inc. | The firm saw on the FDA web site a recall by the manufacturer due to no microbial testing in place at the manufacturer and product did not meet proper finished goods specifications. | Nonconforming Material/Component | Custom Medical Specialties sent an Urgent Medical Device Recall letter dated August 24, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory for the affected product and to discontinue the use of any affected product. Customer were instructedto complete and return the Return Response form, acknowledging they have received the letter and checked their inventory. Any product still in the customers inventory should be returned to the firm. Customers should call (919) 202-8462, ext 205 to obtain a Return Authorization Number and the shipping number.For questions regarding this recall call 919-202-8462. |  | 15 | 15 cases | Nationwide Distribution including GA, MI, NY, and PA. | Terminated | September 10, 2012 | 346 | 38 | Not_Computer
Z-2333-2012 | 62842 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Aesculap S4 Element 4.5mm Polyaxial Screws (ST240T, ST241T, ST242T, ST243T, ST244T, and ST245T). |  |  | July 25, 2012 | 2012 | Aesculap Implant Systems LLC | There is a possibility that the screw channel that holds the rod in place may be out of tolerance. When the depth of the channel is not deep enough, the rod may not be able to provide ample compression on the insert preventing the screw head from locking down to the bone screw. | Nonconforming Material/Component | Aesculap sent an "IMPORTANT PRODUCT REMOVAL NOTIFICATION" letter dated July 25, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. A Distribution Inventory Sheet was attached for customers to complete and return with the affected product. Contact the firm at (610) 984-9275 for questions regarding this recall. |  | 5 | 5 | Nationwide Distribution-including the states of IL, MN, NV, OK, and TX. | Terminated | March 18, 2014 | 601 | 5 | Not_Computer
Z-2327-2012 | 62774 | No | EZC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Foley procedural trays packaged in breather bags. The product is packaged ten units in each shipper box, and are sold sterile for single use. |  |  | June 29, 2012 | 2012 | C.R. Bard, Inc., Urological Division | Incomplete seal on sterile packages of product. | Process control | C. R. Bard Inc, sent a Urgent Product Recall Notification letter dated June 27. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This notification is being sent because there is a potential for a limited number of units within the subject lots (noted above) that may contain an open seal of the sterile packages. The issue is visually obvious even before the package is opened. No patient injury or adverse events have been reported. Do not use or further distribute any of the affected products.Once the product affected by tlle recall has been removed from your inventory 1.Fill out the Recall &amp; Effectiveness Check Form. Be sure to state the quantities affected of the recalled product that you have in stock. It is extremely important that we receive this information. 2. Fax the Recall &amp; Effectiveness Check Form to BMD at 1-770-784-6469. The Recall Coordinator will issue you with a Return Authorization (RCL) to facilitate the expedient return and replacement of the product. 3. A mailing label is enclosed for your convenience to return the affected product. Please mark the outside package as "RECALLED PRODUCT" and include the RCL number. All products should be retumed to the following shipping address that is in the letter. |  | 56694 | 56,694 units | Nationwide distribution: USA | Terminated | February 07, 2013 | 223 | 56694 | Not_Computer
Z-2346-2012 | 62909 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Sterile Empty Vial and Injector (30mL PCA vial), Rx only, One/List No. 6021-03, Hospira, Inc., Lake Forest, IL 60045. |  |  | August 02, 2012 | 2012 | Hospira, Inc. | Reports of leaking during filling and administration. | Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter dated August 3, 2012, to all affected customers. The letter identifed the product, the probem, and the action to be taken by the customer. Customers were instructed to first consider the use of Prefilled Morphine PCA Vials wherever possible. In some instances, the treating physician may require an alternative to morphine.Customers may continue using the affected product by following the instructions provided in the Correction Letter. Customers were asked to complete the attached Reply Form and fax it to 1-888-879-8144.For questions regarding this recall call 877-946-7747. |  |  | unknown | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, NK, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY and the countries of Canada, Costa Rica and the Virgin Islands. | Terminated | May 08, 2014 | 644 |  | Not_Computer
Z-2340-2012 | 62915 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Multileaf Collimator |  |  | July 27, 2012 | 2012 | Elekta, Inc. | Information released with the product Agility did not originally contain information on compatibility with the add-on Dynamic Multileaf Collimator Apex. | Error in labeling | ELEKTA sent an Important Field Safety Notice, FCO Ref: 20001405000 dated July 7, 2012, to current Agility users and a separate notice FCO Ref: 20001407026 for current Apex users who may be in the process of considering an upgrade to Agility. Both notices made it clear that Apex has not been integrated or classified as compatible for use with Agility. Users are advised not to use Apex with Agility.For questions regarding this recall call 770-300-9725. |  | 2 | 2 | Nationwide Distribution including DC &amp; MD | Terminated | October 06, 2014 | 801 | 2 | Not_Computer
Z-2347-2012 | 62328 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx monitor/defibrillators |  |  | June 20, 2012 | 2012 | Philips Healthcare Inc. | Therapy Cable Connection wear may lead to a malfunction in detection of defibrillation pads/paddles therapy Cables | Device Design | Philips notified customers on June 20, 2012, by an Urgent Medical Device Correction notification/Field Safety Notice and a revised Instructions for Use addendum for each HeartStart MRx unit. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of units affected and and identifies the action. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Carefully read the attached HeartStart MRx Instructions for Use Addendum. Immediately implement ongoing therapy connection inspection on all of your MRx devices to detect wear from higher than expected levels of stress. The HeartStart MRx Instructions for Use Addendum describes how a user can identify wear. If wear is detected, remove affected devices immediately from use and contact Philips to arrange for service.WARNING: The service life of your therapy cables/external paddles is up to three years. To maintain reliable performance and reduce the possibility of failure during patient use, replace them every three years from the time they were initially placed into service or if they fail the inspection criteria in the attached Instructions for Use AddendumThe addendum provides the customer with complete user information regarding ongoing inspections of port. In addition, service support will be provided for Therapy Port and Cables requirement replacements.Further information or support concerning this issue,contact your local Philips representative at 1-800-722-9377.For questions regarding this recall calll 978-687-1501. |  | 58701 | 58,701 units | Worldwide Distribution - USA (nationwide) and the countries of : AFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBANGLADESHBELGIUMBOLIVIABOSNIA AND HERZEGOWINABRAZILBRUNEI DARUSSALAMBULGARIACAMBODIACANADACHILECHINACOLOMBIACOSTA RICACROATIA (local name: Hrvatska)CYPRUSCZECH REPUBLICDENMARKECUADOREGYPTEL SALVADORESTONIAETHIOPIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGICELANDINDIAINDONESIAIRAQIRELANDISRAELITALYJAPANJORDANKENYAKOREA, REPUBLIC OFKUWAITLAOSLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIALUXEMBOURGMACAUMACEDONIAMALAYSIAMALTAMEXICOMOROCCONAMIBIANETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORWAYOMANPAKISTANPANAMAPAPUA NEW GUINEAPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIAN FEDERATIONSAUDI ARABIASINGAPORESLOVENIASOUTH AFRICASPAINSRI LANKASWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATANZANIA, UNITED REPUBLIC OFTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMUZBEKISTANVENEZUELAVIET NAMYEMEN | Open, Classified |  |  | 58701 | Not_Computer
Z-2344-2012 | 63008 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Interfaces subject to recall: |  |  | September 27, 2010 | 2010 | Medtronic Xomed, Inc. | In January 2010, Medtronic issued a recall of Medtronic NIM 3.0 Patient Interfaces after Medtronic identified a trend of open channel conditions (signal loss from the electrode). When the NIM 3.0 detects the signal loss from the electrode (i.e. open channel), it fails the electrode check and provides an audible and visual warning to the user which could also create a delay in the surgical procedu | Device Design | The firm, Medtronic USA, Inc., sent an "URGENT Product Recall Notification" letter dated September 27, 2010, to its U.S. and International customers. The letter described the product, problem and actions to be taken. The customers were instructed to verify their receipt of this notice and complete and sign the Product Replacement checklist form indicating that their sales representative has completed the replacement of their devices and fax to 1-651-367-8055.U.S. Customer Units - replacement kits with their contents will be provided to each Area Sales Manager (ASM) for each customer along with documentation of serial number units to be replaced. Each ASM will be responsible for delivering a replacement unit(s) and picking up and returning each of their customer's units using the pre-labeled replacement kit boxes.International Customer Units- International units, depending on the geography will either be replaced or repaired and returned. If replacement is planned, the International Business Manager (IBM) will receive replacement kits as described for the U.S. In those geographies where devices will be repaired and returned, kits with the following contents will be sent to IBM's along with documentation of serial number units to be repaired.Foreign distribution: In the larger territories (i.e. Australia, Japan, Canada, EMEA-Europe, Middle East, Africa), repairs/rework were handled at the Medtronic facilities in those geographic areas. In the smaller geographic areas (Singapore, Thailand, Hong Kong, Korea, Latin America ), where there were limited sales, devices were returned to the U.S. In all cases (large and small geographic areas), devices were returned via a Medtronic entity (sales office or distribution center).If you have any questions regarding the subject Actron or the content of this letter, contact the Distinguished Regulatory Affairs Advisor at 904-279-7532. |  | 7 | 7 | Worldwide distribution: USA (nationwide) and countries of: Australia, Austria, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Lebanon, Libya, Luxembourg, Netherlands, New Zealand, Norway, Portugal, Saudia Arabia, Singapore, Slovak, Spain, SRI Lanka, Sweden, Turkey, United Arab Emirates, and United Kingdom. | Terminated | November 26, 2013 | 1156 | 524 | Not_Computer
Z-2345-2012 | 63102 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano brand Intravascular Ultrasound, Volcano s5/s5i Family Software Version 3.2.1 or 3.2.2, Models/Part Numbers: 435-0607.01, 435-0607.04, 435-0602.01, 435-0602.02; Product is manufactured and distributed by Volcano Corporation, Rancho Cordova, CA. |  |  | August 24, 2012 | 2012 | Volcano Corporation | The display on the system has the potential to change. In house testing eventually identified the unique sequence of events that must occur in order for this issue to manifest itself. | Software design | Volcano Corporation sent an "IMPORTANT FIELD CORRECTIVE ACTION" letter dated August 24, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. Contact your local Volcano sales representative or Volcano Customer Service for questions regarding this recall. |  | 1020 | 1020 units | Worldwide Distribution. | Terminated | November 25, 2014 | 823 | 1020 | Not_Computer
Z-2339-2012 | 62807 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Tip Cover for 8mm Monopolar Curved Scissors (Hot Shears)Accessory, Disposable, 8mm MCS, Box of 10. |  |  | August 09, 2012 | 2012 | Intuitive Surgical, Inc. | There is potential for the sterility of the product to be compromised. | Employee error | Intuitive Surgical sent a "MEDICAL DEVICE CORRECTION" letter dated August 9, 2012 to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return via fax to (408) 716-3040. Contact Customer Service at 1-800-876-1310 for questions regarding this recall. |  | 11121 | 11,121 Boxes | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Canada, and Lebanon | Terminated | December 31, 2012 | 144 | 11121 | Not_Computer
Z-2341-2012 | 62988 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Uroskop Access. |  |  | April 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens has discovered a possible hazard to patients or hospital personnel when tilting the Uroskop Access system starting with serial number 6076. When tilting the table on the Uroskop Access table, the motor may become disconnected and the entire table assembly may fall on the floor. | Nonconforming Material/Component | Siemens sent a Update instructions letter dated April 2012, to all affected customers. The safety instruction letter identified the product the problem and the action needed to be taken by the customer.The update has been completed successfully. The customer has been informed of the benefits of improvements resulting from this update.Further questions please call (610) 219-4834. |  | 5 | 5 | Nationwide distribution: USA including states of: IN, MI, MS, TN and TX. | Terminated | September 20, 2012 | 143 | 5 | Not_Computer
Z-2367-2012 | 62989 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett Mira mobile x-ray system |  |  | April 30, 2012 | 2012 | Siemens Medical Solutions USA, Inc | During regular product monitoring, firm became aware of a potential issue with the Mobilett Mira system, (serial numbers 1034, 1038, 1041, 1042, 1043, 1050, 1053, and 1054. During normal clinical operation the center screw on the driving wheel may become loose resulting in system immobilization. There is no risk of the system tipping over or collapsing on the floor, however the system may tilt t | Device Design | Siemens has issued an Update Instruction XP018/12S in April 2012 to all affected consignees notifying them of the problem. The update instruction provides a correction of the wheel fixation. For questions call Siemens at 610-219-4834 or 610-448-4634. |  | 9 | 9 | US Nationwide Distribution including the states of: GA, IA, KS, NC and TX. | Terminated | September 25, 2012 | 148 | 9 | Not_Computer
Z-2362-2012 | 62140 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 101-05-30***NOVATION MODULAR DRILL BIT LENGTH: 30mm, DIAMETER 3.2mm***Use with 30mm Screw***STERILE***EXACTECH 2320 NW 56th Court, Gainesville, FL 32653***. |  |  | January 24, 2012 | 2012 | Exactech, Inc. | Exactech Inc. of Gainesville, FL is recalling their Novation Modular Drill Bits after the devices were reported to have been used by end user while in a non-sterile condition. | Finished device change control | The firm, Exactech, Inc., called the lone customer on January 24, 2012 to discuss the product, problem and actions to be taken. The customer was instructed to return all unused devices and associated packaging. |  | 2 | 2 | Nationwide distribution: USA state of: Arkansas only | Terminated | September 26, 2012 | 246 | 2 | Not_Computer
Z-2368-2012 | 62926 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Enduramesh Corpectomy Spacer System |  |  | July 15, 2012 | 2012 | Lucero Medical LLC | The firm voluntarily recalled the device after they discovered that the sterilization recommendations on the product insert did not pass sterilization validation. | Error in labeling | Lucero Medical sent a Field Correction Notice letter dated July 15, 2012 to it's sole consignee. The notice identified the affected product., problem and actions to be taken. The customer was instructed to complete the response card and submit back to their local Lucero Medical Representative. |  | 169 | 169 | US Nationwide Distribution | Terminated | February 20, 2013 | 220 | 169 | Not_Computer
Z-2348-2012 | 62321 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Safety monitor included with System 8000 Base |  |  | August 12, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received complaints from our Affiliate, TerumoDubai that the safety monitors were tripping off. Terumo CVS' investigation determined that 8monitors were manufactured with an underrated circuit breaker due to an error where an associatepulled an underrated circuit breaker from an incorrect inventory location. Since this time, inventorylocations hav | Mixed-up of materials/components | TERUMO Cardiovascular Systems communicated with their affiliate, Terumo Dubai, in 2009 at the time the complaints were reported. No further communication was warranted.For questions regarding this recall call 734-741-6173. |  | 8 | 8 | Worldwide distribution: International only-country of: Terumo Dubai, UAE | Terminated | January 29, 2013 | 1266 | 8 | Not_Computer
Z-2369-2012 | 54659 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 V1.05(FG-5400-00). |  |  | November 17, 2009 | 2009 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster has become aware of information regarding a potential hazard that may occur during the operation of a Bloom EP Stimulator device in conjunction with RF generators and other electrophysiology (EP) equipment, including the CARTO 3 Electropysiology Mapping System. | Labeling mix-ups | Biosense Webster sent an " URGENT FIELD SAFETY NOTICE" dated November 20, 2009 to all affected customers who purchased the Carto 3 Electrophysiology Mapping System. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to continue use of the CARTO 3 System and that it was safe and no further action is needed. Customers with questions about the notice were instructed to contact Biosense Webster representatives.On January 12, 2010 another letter was sent from Biosense to all their customers recommending to avoid any connection of an EP stimulator's pacing pins to the direct ports on the CARTO 3 System when applying RF energy, unless the set up is indicated by the stimulator's instructions for use. The letter also informed the customer that the user manual for the CARTO 3 System has been formally revised to include a warning. Customers were instructed to complete an acknowledgement form to confirm receipt of the revised user manual. |  | 71 | 71 units total (30 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Denmark, France, Germany, Israel, Italy, Netherlands, and United Kingdom. | Terminated | September 12, 2012 | 1030 | 71 | Not_Computer
Z-2376-2012 | 62760 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2121IQ, Innova 3131IQ Cardiovascular X-ray imaging systems. |  |  | August 16, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Innova 2121IQ &amp; Innova 3131IQ Systems. While performing a fluoroscopic acquisition with the Innova 2121IQ or Innova 3131IQ systems, there is a potential for loss of usable imaging mode which requires the system to restart in order to recover. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated August 16, 2012, to all affected consignees. The letter described the Safety Issue, safety Instructions, Affected Product Details, Product Correction, and Contact Information. Consignees were instructed that hospital staff should wait for Photo or Image to appear on the display (when store photo function is performed) before start of the next short duration fluoroscopy acquisition. For questions customers were instructed to call:United States (800) 437-1171Japan: 0120 - 055 - 919For other countries consignees should contact their local GE Healthcare Service Representative.For questions regarding this recall call 262-513-4122. |  | 240 | 240 (133 US and 107 OUS) | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, FL, HI, IK, IL IN, IA, KS, LA, MD, MA, MI, MN,MO, NV, NH, NH, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI,.and internationally to:AUSTRALIA, AZERBAIJIN, BAHRAIN, BELARUS, BULGARIA, CANADA, CHILE, CHINA, CZECH REPUBLIC, EGYPT, FRANCE, GERMANY, INDIA, INDONESIA, IRAQ, IRELAND, ITALY, JAPAN, KUWAIT, LEBANON, LIBYA, MALAYSIA, MEXICO, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAÂżDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, TAIWAN, THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM, and YEMEN. | Terminated | October 20, 2015 | 1160 | 240 | Not_Computer
Z-2372-2012 | 62912 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Omnilife Science Apex Knee System Modular Tibia Augment, Size 3 x 4mm |  |  | August 09, 2012 | 2012 | Omni Life Science | The Apex Knee System Modular Tibia Augment may breach the inner and outer sterile pouches, compromising the sterility | Packaging | Omnilife Science contacted affected customers by phone and sent an "URGENT MEDICAL DEVICE RECALL" e-mail on August 9, 2012. The e-mail identifies the product, problem, and actions to be taken by the customers. Customers with questions regarding product return should contact Customer Service at 800-448-6664. Customers with questions regarding the recall should contact the firm at 774-226-1845. |  | 310 | 310 units | Worldwide Distribution-USA (nationwide including the states of CO, NY, OK, PA, UT, VA, and WI, and the countries of Spain and the Netherlands. | Terminated | January 10, 2013 | 154 | 1396 | Not_Computer
Z-2256-2012 | 62946 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Optima XR200amx. Mobile general purpose radiographic imaging of the human head and body. |  |  | July 10, 2012 | 2012 | GE Healthcare, LLC | Software issue: a violation of 21 CFR 1020.31(a). After an image is taken, when the user selects and clears an additionally displayed screen the system inhibit is also cleared, allowing the user to take an exposure without the technique factors being displayed. | Radiation Control for Health and Safety Act | The firm will issue Important Electronic Product Radiation Warning letters that identify the affected product along with a description of the defect and related hazards. Users are instructed to ensure that the acquisition mode is always selected prior to attempting an exposure to confirm the techniques selected. The firm will remedy the issue or bring the device into compliance free of charge. A GE Healthcare Service Representative will update the software on the system to address the issue. Field Modification Instructions (FMI) 10865(2) describes the rework plan the GEHC Engineers will make to the imaging system to return it to compliance with the applicable performance standard. It appears to adequately address the problem and is hereby approved. Users are to contact 800-437-1171 in the US, if they have any questions or concerns. |  | 254 | 254 total in the US | Worldwide Distribution -- USA (nationwide) Distribution | Terminated | January 10, 2013 | 184 | 254 | Computer
Z-2377-2012 | 62844 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Capsule Neuron Docking Station, a component in the DataCaptor Connectivity System |  |  | August 10, 2012 | 2012 | Capsule Tech Inc. | Capsule Neuron Docking Station may experience a lost connection, resulting in loss of data or loss of power to the Caspsule Neuron | Device Design | The firm, Capsule Tech, sent a recall letter dated August 10,2012 to its customers. The letter described the product, problem and actions to be taken. Capsule is sending field correction kits to customers with instructions. Capsule's Professional Services team will contact the customers to schedule shipment of their kit. Final shipments of the field correction kits will be completed by Oct 2012. If you have any questions or concerns, contact your professional services representative and/or call in US: 978-482-2340 or France: +33 153-34-14-00. |  | 12448 | 12,448 units | Worldwide distribution: USA (nationwide) and countries of: Belgium, Canada, Denmark, France, Germany, Netherlands and UK. | Terminated | April 22, 2015 | 985 | 12448 | Not_Computer
Z-2382-2012 | 63007 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL 91493 Integrated Module Housing. |  |  | August 27, 2012 | 2012 | Spacelabs Healthcare, Llc | There is a potential failure of the Integrated Module Housing causing the patient module parameters to drop off the monitor display or the patient module will fail to sign on when inserted into the housing. | Mixed-up of materials/components | Spacelabs Healthcare sent an "URGENT- MEDICAL DEVICE CORRECTION" letter dated August 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-522-7025 and select 2 for assistance regarding this recall. |  | 527 | 527 units total; 350 US and 177 International | Worldwide Distribution. | Terminated | April 15, 2013 | 231 | 527 | Not_Computer
Z-2384-2012 | 62884 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost Systems with M Cabinet |  |  | August 14, 2012 | 2012 | Philips Healthcare Inc. | Creepage Distance between the two primary phases in the GEO Transformer is 4.5 mm.does not meet (IEC) International standard of 5.5 mm. | Device Design | Philips Healthcare Imaging issued an Urgent - Field Safety Notice dated August 8, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The unit will be exchanged with a new unit that complies with the IEC requirements. Please reference Field Change Order (FCO) 71200081 when contacting I your local Philips representative. 1-800-722-9377.Customer were instructed to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter and to maintain a copy of the letter with the equipment instruction for Use. For further assistance, customers should contact their local Phillips representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 4 | 4 systems | Worldwide Distribution - USA including OH and WA and the country of Canada | Open, Classified |  |  | 4 | Not_Computer
Z-2387-2012 | 62894 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | Instrumentation Laboratory ACL TOP (Base) |  |  | August 03, 2012 | 2012 | Instrumentation Laboratory Co. | Low Abnormal Fibrinogen Sample Reports as Normal and High Fibrinogen Sample Reports as Low | Software design | The firm, Instrumentation Laboratory, sent a "Mandatory Test Parameter Upgrade" letter dated June 7, 2012 with upgrade kits P-14.8 via USPS on August 3 thru August 10 to US and Canadian customers and notified International Affiliates via e-mail on August 2, 2012. The customers were ask to complete and return the Mandatory Test Parameter Upgrade Tracking Form via fax: 781-861-4207 or email: ra-usa@ilww.com. If you have any questions, call 781-861-4204. |  | 1074 | 1074 units | Worldwide distribution: USA (nationwide) and countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, REPUBLIC OF LUXEMBOURG, MACAU, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SLOVAKIA (Slovak Republic), SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TURKEY, UNITED KINGDOM, URUGUAY and VIETNAM. | Terminated | October 02, 2015 | 1155 | 4297 | Not_Computer
Z-2378-2012 | 63110 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The product affected by this notice is sold under the trade name TomoTherapy SharePlan 1.1 . To determine if the version you are using is affected, bring up the About box in the SharePlan application and check if the build number reported there is 1.3.1.10 . |  |  | May 30, 2011 | 2011 | RAYSEARCH LABORATORIES AB | This field notice concerns two behaviors of the SharePlan (Multiple Fraction Groups and Depth Offset) that may be unexpected and are not clearly described in the user manual. None of these behaviors have caused patient mistreatment or other incidents. However, the described user actions must be adhered to for best agreement between treatment plan evaluation and the clinical outcome of the delivere | Labeling False and Misleading | RaySearch Laboratories sent an Urgent Field Safety Notice Medical Device Correction dated May 30, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consigness were instructed to completely avoid affecting the erroneous behaviour, verify in TomoTherapy Hi-Art that the plans sent to TomoTherapy SharePlan 1.1 only have one single fraction group with a single fraction size,used for all fractions, and do not attempt to import treatment plans having more than one fraction size into the product. To completed avoid affecting the erroneous behaviour the user can either offset the measured curves before importing them into beam commissioning or add the depth offset tothe calibration depth.For further information consigness should contact RaySearch Laboratories AB Quality and Regulatory Affairs at +46 722 366 110 or the Tomo Therapy support at support@tomotherapy.com.North America 1 866 368 4807Belgium 0800 38783France 0805 631 565Germany 0800 000 +541Italy 800-986 399Netherlands 0800 0201364Spain 800 3000049Switzerland 0800 001927UK 0868 238 6035China/Northern 10 800 712 1701China/Southern 10800 120 1701Hong Kong 800 967912Japan 0044 22 132374Singapore 800 1204683South Korea 0079 81 4800 7204All other locations + 1 608 824 2900 or +32 2 40 44 44 |  | 15 | 15 units | Worldwide Distribution - USA including CA, KY, MI, NY, OK, and OR. Internationally to Canada, Belgium, United Kingdom, Australia, France, Germany, and Poland. | Terminated | November 16, 2012 | 536 | 15 | Not_Computer
Z-2381-2012 | 62728 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 2.7mm Drilling Guide is a reusable instrument used in conjunction with a drill to prepare oriented holes in bone for insertion of a Standard SurfixÂż Locking Screw with plates for the Tiblaxis, Uni-CP, Advansys MLP and Advansys DLP compression plate systems. |  |  | June 14, 2012 | 2012 | Integra Life Sci. | Integra' s quality system noted an adverse trend in complaints for breakage of the drilling guides during use. None of the complaints resulted in patient injuries but in some cases resulted in an increase in the time to complete the surgical procedure. A physician was consulted to assess patient risk and concluded that any potential for risk would be transient, self-limiting or minor, or result | Device Design | Integra sent a "SUBJECT: MEDICAL DEVICE RECALL" letter dated June 14, 2012 to all affected customers. The letter identified the product, problem, and action to be taken by the customers. Customers were asked to complete a Recall Acknowledgement and Return Form whether or not they were affected by the recall, and return all their affected inventory to Integra LifeSciences.Customers were asked to contact their regional managers or the firm at 609-936-6832 for questions regarding the recall. |  | 2389 | 2,389 units, Catalogue No. 219635ND; 2.7mm Drilling Guides | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI and the countries of Austria, Czech republic, Denmark, France, Germany, Ireland, Israel, Italy, Kingdom of Saudi Arabia, Lebanon, Norway, Portugal, Switzerland, South Africa, Spain, Sweden, and United Kingdom. | Terminated | February 20, 2013 | 251 | 2389 | Not_Computer
Z-2379-2012 | 62867 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HALO Flex Energy Generators |  |  | August 03, 2012 | 2012 | Covidien LLC | Potential failure of HALO Energy Generator to enter "stand by" mode when initially powered on; and therefore can not be used for patient care. | Software design | COVIDIEN sent an Urgent Medical Device Correction letter dated August 3, 2012, to all affected customers via Federal Express. Additionally, Field Service Representatives will be contacting their accounts to notify them of the correction.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following steps prior to initiating patient treatment and specifically, prior to administering conscious sedation and performing upper endoscopy:1. Power-on the HALO Flex Energy Generator and confirm that the generator passes self-test, enters standby mode and displays "Ready, Connect Catheter."2. Confirm that the HALO Flex Energy Generator is functioning properly at power-up and enters the standby mode then the procedure can begin.3. Do not use the HALO Flex Energy Generator if the generator does not pass the power on self-test. In this situation, please call Covidien, GI Solutions customer service at (888) 662-2779 or your local territory manager (sales representative) so that the firm can provide them with a replacemenent generator.4. Customers were asked to reply to Covidien, GI Solutions (formerly BARRX Medical) by signing and returning the attached Product Correction Form via fax to the number provided on the form. Customers with questions were instructed to call 888-662-2779.For questions regarding this recall call 408-328-7300. |  | 617 | 617 generators | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, Colombia, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Israel, Jordan, Netherlands, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Venezuela and UK. | Terminated | September 30, 2013 | 423 | 617 | Not_Computer
Z-2386-2012 | 63014 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice CT Scanner System, models 728131 and 728132. The MX 16-slice CT Scanner System can be used as a whole body computed tomography X-ray system. |  |  | August 24, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | On August 24, 2012 the firm recalled the MX 16-slice after discovering that a bolt which attaches the oil accumulator within the gantry may fail, resulting in the oil accumulator detaching from the tube housing assembly within the gantry. | Process design | An "Urgent Medical Device Correction" Field Safety Notice dated August 24, 2012 was sent to all US consignees via certified mail. The problem was described and recommended actions were provided. Philips Key Market representatives distributed the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Orders. US Customer contact number: 1 800-722-9377. |  | 328 | 328 units | Worldwide distribution, Nationwide (US) including the states AL, AR, CA, FL, GA, IL, IN, MS, NJ, NV, NY, OH, PA, and VA and the countries of Albania, Argentina, Australia, Belgium, Brazil, Chile, China, Colombia, Dominican Republic, Ecuador, France, Germany, Greece, India, Indonesia, Ireland, Italy, Jordan, Lebanon, Malawi, Mexico, Nepal, Netherlands, Oman, Palestine, Philippines, Poland, Romania, Russia, Saudi Arabia, South Africa, South Korea, Spain, Switzerland, Turkey, Uganda, and Venezuela. | Terminated | August 28, 2013 | 369 | 328 | Not_Computer
Z-2383-2012 | 63033 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Trilogy, Novalis Tx, Unique, linear accelerators; Model Number: H14, H18, H29, Reference/FSCA Identifier: CP-08881; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA. |  |  | August 14, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Some X-jaw (lower collimator jaw) carriers were made using the incorrect metal alloy, which can cause the jaw carrier to crack. If both carriers on the same X-jaw crack, the jaw would be able to move freely in the closed direction. The position readout interlock circuitry will not detect this jaw position deviation. This may lead to treatment with an under-dose of the intended target volume. | Nonconforming Material/Component | Varian sent an "URGENT MEDICAL DEVICE CORRECTION URGENT FIELD SAFETY NOTICE" dated August 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-888-827-4265 for questions regarding this recall. |  | 221 | 221 Units | Worldwide Distribution. | Terminated | December 04, 2013 | 477 | 221 | Not_Computer
Z-2385-2012 | 62899 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Sterile Distraction Screws of 3 different sizes: 12mm, 14mm and 16mm sold under TSI, 3COR, DB Surgical, and American Surgical Products brands and Sterile XL Screws of 7 different sizes: 90mm, 100mm, 110mm, 120mm, 130mm, 140mm, 150mm sold under TSI brand |  |  | June 11, 2012 | 2012 | Tedan Surgical Innovations Llc | This letter is to notify you that several of our sterile screw products are being voluntarily recalled. Product sterility may be compromised due to a mechanical failure of the protective packaging tube. Tube end plugs may dislodge creating the potential for the screw to come out of the protective tube and then the sharp end of the screw pierce the outer sterile barrier pouch. | Packaging | The firm, TeDan Surgical Innovations (TSI), sent an "URGENT: Medical Device Recall" letter dated June 11, 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine all products subject to recall and return product to TSI. In addition, if they have further distributed this product, identify all customers and notify them of this product recall at once in writing; and complete and return the enclosed Mandatory Response Form as soon as possible and no later than June 25, 2012 via mail: 12675 W. Airport Blvd., Suite 200, Sugar Land, TX 77478; fax: (713) 726-0846 or email: thernandez@tedansurgical.com.If you have any questions, contact the Business Manager at 713-726-0886, Monday thru Friday, from 8am - 5pm or via email at THERNANDEZ@TEDANSURGICAL.COM. |  | 12065 | 12,065 Sterile Distraction Screws (2413 boxes); 384 Sterile XL Screws (128 boxes) in distribution | Worldwide distribution: USA (nationwide) and countries of : Australia and Taiwan. | Terminated | July 31, 2013 | 415 | 12065 | Not_Computer
Z-2380-2012 | 62939 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Ultraview DM3 Monitor |  |  | June 06, 2012 | 2012 | Zoe Medical Incorporated | Unit fails to power up, resulting in an equipment alarm | Software design | The firm, Zoe Medical, coordinated with the Distributor via email on June 7, 2012. The email described the product, problem and actions to be taken. Zoe will remove product the site and return the units for correction.If you have any questions, contact Manager, Global Technical Support at 425-657-7200 ext 5508. |  | 35 | 35 units | Nationwide distribution: WA only. | Terminated | September 18, 2012 | 104 | 35 | Not_Computer
Z-2394-2012 | 63105 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, Pre-Filled Flush Solution, 10 mL in 12 mL Single Use Syringe, Rx only, List no. 1078-20, made in Nashville, TN, manufactured for Hospira, Inc., Lake Forest, IL 60045 USA. |  |  | August 15, 2012 | 2012 | Amsino Medical USA | One lot of 0.9% Sodium Chloride Flush syringe, 3 mL in 12 mL may be mislabeled as 0.9% Sodium Chloride Flush syringe, 10 mL in 12 mL. | Employee error | The firm, Amusa, notified their sole consignee of the mislabeled product via e-mail and telephone on August 15, 2012. An "URGENT-CLASS II MEDICAL DEVICE RECALL" letter dated August 22, 2012 was sent to the customer. The letter identified the product, problem, and actions to be taken by the customer. A Response Form was included for the customer to complete and return. Contact the firm at 615-521-2723 for questions regarding this recall. |  |  | unknown | Distributed to one customer in Illinois. | Terminated | August 13, 2013 | 363 |  | Not_Computer
Z-2395-2012 | 63078 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Empty Evacuated Container 250 mL and 1000 mL. |  |  | August 28, 2012 | 2012 | Hospira Inc. | Embedded iron oxide glass defect which may have the potential to break off into the solution. | Material/Component Contamination | Hospira sent an "URGENT DEVICE RECALL" dated August 28, 2012 to all affected customers. The letter identified the product, problem,and actions to be taken by the customers. The letter included an Urgent Device Recall Reply Form for customers to complete and return via fax to 1-888-237-5582. Contact Hospira Customer Care at 1-877-946-7747 for questions regarding this recall. |  | 191094 | 191,094 units | Worldwide Distribution-USA (nationwide) and the countries of Saudi Arabia, Canada and Barbados. | Terminated | September 18, 2013 | 386 | 191094 | Not_Computer
Z-2402-2012 | 54705 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | A mobile, electro-hydraulically operated, fluoroscopic/C-Arm compatible general surgical table. |  |  | February 17, 2010 | 2010 | Steris Corporation | Customers may be storing objects on the base and/or around the column of the surgical table. This prohibited practice and misuse known by the customers can damage the override switch assembly located at the top of the column cover. Depending upon the nature and severity of the damage to the electronics of the table, it is possible that the table could move without prompting by an operator . | Employee error | Steris sent an Urgent Field Correction Notice letter dated February 17, 2010 to their customers. The letter identified the affected product, description of the problem and actions to be taken. Customers were instructed to examine their units for any structural damage. If found they are to discontinue use and arrange for a field representative to visit the facility to assess the damages and arrange for repair of the unit. For information or questions regarding STERIS visit to your facility contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 11044 | 11,044 Units | Worldwide Distribution - US Nationwide including District of Columbia and Puerto Rico and the following countries: American Samoa, Argentina, Australia, Bahrain, Bangladesh, Bermuda, Brazil, Canada, China, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, Germany, Greece, Guam, Hong Kong, India, Indonesia, Israel, Italy, Japa, Jordan, Kazakhstan, Korea, Kuwait, Lebanon , Libya , Malaysia , Mexico, Morocco, New Zealand, Panama, Peru, Philippines , Poland, Portugal, Qatar, Russia, Saudi Arabia , Singapore, South Africa, Spain, Sri Lanka, Taiwan, Thailand, Trinidad and Tobago, Turkey &amp; United Kingdom. | Terminated | March 13, 2013 | 1120 | 11044 | Not_Computer
Z-2401-2012 | 63004 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll X Series Defibrillator/Pacemaker/Monitor |  |  | August 16, 2012 | 2012 | ZOLL Medical Corporation | Shipped with incorrect software | Process control | Zoll Medical telephoned affected customers on August 16, 2012. Customers were informed of the affected product, problem and actions to be taken. Customers were instructed to remove and return the affected product for replacement. For question contact ZOLL's Technical Support. |  | 11 | 11 units | Worldwide Distribution - US Nationwide including the states of: LA, MA, ME, NC, ND, NE, OH and the country of Germany. | Terminated | September 19, 2012 | 34 | 11 | Computer
Z-2396-2012 | 63047 | [6 empty] | InfinityÂż sampling device, Cytology Brush, United States Endoscopy Group, Inc., 5976 Heisley Road, Mentor, OH 44060 |  |  | December 22, 2011 | 2011 | US Endoscopy Group Inc | US Endoscopy received five complaints which noted difficulty in deploying the cytology brush through the catheter. | Nonconforming Material/Component | US Endoscopy sent an Addendum and Correction to the Instructions for Use letter dated December 22, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to rinse the brush head with sterile water immediately prior to use (reference the attached Addendum. Customers were also instructed to retract and deploy the brush to ensure that is is functioning properly prior to use. Customers were asked to review the instruction for use for this product along with the Addendum to ensure proper usage and performance.For questions customers should call 1-410-639-4494 or their local Endoscopic Product Specialist.For questions regarding this recall call 440-639-4494, ext 378. |  | 393 | 393 Devices | Worldwide Distribution - USA including CA, FL, GA, IA, IL, MA, ME, MN, MO, NC, NH, OH, OK, PA, TN, TX, VA, WA &amp; WV and the countries of Australia, Canada, Columbia, Finland, Germany, India, Israel, Italy, Netherlands, New Zealand &amp; United Kingdom. | Terminated | March 22, 2013 | 456 | 393 | Not_Computer
Z-2397-2012 | 62494 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS 7-color Setup beads |  |  | June 29, 2012 | 2012 | BD Biosciences, Systems &amp; Reagents | Incorrect version of the CD was included with BD FACS 7 Color Setup Beads. | Employee error | BD Biosciences sent a Important Product Recall Information letter beginning July 3, 2012, to all affected consignees via Fed Express. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed that the CD enclosed in their recent purchase of BD FACS 7-Color Setup Beads, lot number 33412, contain CD version 2.0 instead of the newer version 3.0. The letter goes on to explain the the potential risk and include a table of listed range values to determine if a potential false pass or false fail for the listed detectors has occurred. Customers are asked to check the CD with the kit to ensure that it is version 3.0 (the kit lot number ID is on the CD label). If it is not the correct version, discard version 2.0 and do not use it with this bead kit. Call BAD customers service for a replacement.If the 2.0 kit was already uses, letters request that clinicians inspect the FITC and APC-CY7 stains in the plots of the laboratory reports for the clinical applications in BD FACSCanto clinical software for compromised resolution between negative and dim populations.The verification form/effectiveness check should be completed and returned as requested. Questions or concerns may be directed to BD Customer Support at 1-855-236-2772. |  | 96 | 96 kits | Worldwide Distribution -- USA (nationwide) including the countries of Australia, Belgium, Canada, Chile, Korea, Mexico, Singapore and Taiwan. | Terminated | March 25, 2013 | 269 | 96 | Computer
Z-2399-2012 | 62391 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | JUGGERKNOT INÂżGUIDE PUNCH, 2.9MM. |  |  | August 14, 2012 | 2012 | Biomet, Inc. | Investigation found that the current design of the product allows the press-fit of the knob/shaft interface to be overcome by the impaction force of a mallet, causing the knob to potentially unfasten and slide down the shaft. | Device Design | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated August 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Response Form was included for customers to complete and return via fax to 574-372-1683. Questions related to the recall should be directed to 574-371-3758, Monday through Friday, 8 a.m. to 5 p.m. |  | 62 | 62 | Nationwide Distribution-including the states of WA, IN, AK, PA, GA, SD, KS, NY, VA, AR, AL, OH, WI, MA, and CA. | Terminated | April 18, 2013 | 247 | 62 | Not_Computer
Z-2400-2012 | 63098 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Guardian II Hemostasis Valve, Model 8210 and Guardian II NC Hemostasis Valve, Model FH102. |  |  | August 30, 2012 | 2012 | Vascular Solutions, Inc. | Vascular Solutions, Inc., became aware of a potential problem involving pouches missing a seal from Guardian II (Model 8210) and Guardian II NC (Model FH102) Hemostasis Valve which poses a risk of product contamination. | Process control | The firm, Vascular Solutions, sent an "URGENT - Medical Device Field Safety Notice" dated August 24, 2012 and a follow-up "Urgent Medical Device Removal - Lot Specific" letter dated August 30, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check their inventory and to follow steps outlined in the letter which included to complete and return the "VSI Account Inventory Form" via email to: customerservice@vasc.com or fax: to (763) 656-4251.If you have specific questions or concerns regarding this matter, please contact your local Vascular Solutions Account Manager; call (763) 656 4300 or email: Senior Regulatory Operations Associate at SOuellette@vasc.com and/or email: www.vascularsolutions.com. |  | 3500 | 3500 units | Worldwide distribution: USA (nationwide) including states of: AZ, AR, CA, CT, DE, FL, GA, ID, IL IN , KY, LA,MD, MN, MO, NV, NM, NY, OH, PA, SC, TN, TX, VA, WA, and WV; and countries of: Italy and Korea.Italy, Korea | Terminated | December 04, 2012 | 96 | 3500 | Not_Computer
Z-2398-2012 | 62850 | No | LOM |  |  | Microbiology | PMA | The Bio-Rad MONOLISA Anti-HBs EIA is for In Vitro Diagnostic Use and it has the catalog number 25220. It is a qualitative and quantitative enzyme immunoassay for the detection of antibody for hepatitis B surface antigen in human serum and EDTA or citrated plasma. |  |  | July 27, 2012 | 2012 | Bio-Rad Laboratories | Prolonged exposure of the MONOLISA Anti-HBs EIA Conjugate Diluent component to light may result in elevated Optical Densities (OD) in some samples such that a negative sample could elevate to a grey zone or low positive result. | Packaging | The firm, Bio-Rad Laboratories, Inc., sent an "Important Product Information Bio-Rad MONOLISA Anti-HBs EIA Catalog #25220", letter dated 27 July 2012, to their domestic consignees/customers via FedEx. The International subsidiaries will be notified by e-mail. The consignees/customers were informed of the important storage and handling information regarding the MONOLISA Anti-HBs EIA Conjugate Diluent component . The consignees/customers were instructed to follow the storage and handling instructions as stated in the Package Insert, and compare their laboratories practices to the storage and handling instructions stated by Bio-Rad.Customers who have questions can call Bio Rad Technical Support at 1-800-224-6723, option 2, then option 3. |  | 7400 | 7,400 (approximately)- All lots since distribution date May 9, 2007 | Worldwide distribution: USA (nationwide) and countries of: Canada, South Africa, and Zimbabwe. | Terminated | December 14, 2012 | 140 | 7400 | Not_Computer
Z-2403-2012 | 62738 | No | OYT |  |  | Hematology | 510(K) Exempt | Amplilink (AL) Software 3.2.3 for use with COBAS AMPLICOR, CAP/CA, COBAS TaqMan,and COBAS AmpliPrep/COBAS TaqMan. The software can be used with a Laboratory Information System. |  |  | April 23, 2012 | 2012 | Roche Molecular Systems, Inc. | When using a Laboratory Information system (LIS) with automated systems utilizing AMPLILINK software v3.2 series (AL v3.2) there is a potential for wrong results to be assigned to an order. | Device Design | Roche Molecular Systems Inc. sent "Urgent Field Safety Notice" letters and Faxback forms on 4/23/2012 via UPS to all affected customers. The issue was described and recommended actions were provided. A Product Advisory Notice and an Updated Field Safety Notice were issued as well. |  | 1292 | 1292 instruments | Nationwide (USA) Distribution | Terminated | February 12, 2014 | 660 | 1292 | Computer
Z-2405-2012 | 62917 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Precise Digital Accelerator |  |  | August 06, 2012 | 2012 | Elekta, Inc. | There have been a small number of instances in hospitals where the joint between the drive gear and the gantry base drive wheel has become loose, leading to fatigue failure of some of the securing bolts. | Nonconforming Material/Component | Elekta sent an Important Field Safety Notice dated August, 2012 to inform users that inspection and possibly bolt replacement needs to be carried out. Users should return the signed acknowledgement to Elketa, and an inspection should be carried out under FCO20002103053.For questions customers should call +44(0)1293 654200.For questions regarding this recall call 770-300-9725. |  | 110 | 110 | Nationwide Distribution including AL, AZ, CA, CO, CT, FL, GA, IL, IA, MD, MA, MI, MN, NE, NV, NJ, NM, NY, ND, OH, OK, OR, PA, VA, SC, TN TX, WA, WV, WI and Puerto | Terminated | October 08, 2014 | 793 | 110 | Not_Computer
Z-2410-2012 | 62897 | [6 empty] | ADVIA Centaur FSH Assay: ADVIA Centaur Calibrator B |  |  | August 07, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Prolactin Assay and FSH, Calibrator B fails the analytical sensitivity stated in the IFU for lots below 40 | Nonconforming Material/Component | Siemens Healthcare issued an Urgent Field Safety Notice letter dated August 2012 by e-mail to all affected Siemens Healthcare Diagnostics Branch offices on August 7, 2012. Customers were instructed to discontinue using the ADVIA Centaur Calibrator B lots 40 and below following their obtaining Calibrator 8 lot 41 or higher. Customers were asked to complete and fax-back the response form attached. |  | 15445 | 15,445 kits | Worldwide Distribution - US Nationwide and the countries of: Canada, Argentina, Australia, Bangaldesh,Brazil, Chile, China, Columbia, Croatia, Ecuduar, Egypt, Hong Kong, India, Israel,Japan, Korea, Maylasia, New Zealand, Mexico, Peru, Philipines, Pakistan, Saudia Arabia, Singapore, Thailand, Taiwan, Viet Nam, South Africa and UK. | Terminated | February 26, 2014 | 568 | 30900 | Not_Computer
Z-2404-2012 | 63006 | [6 empty] | The Brilliance CT 16 System |  |  | August 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was notified that the system logout in software version 2.3.6 is now longer and may sometimes fail to logout, requiring a forced system restart or hard shutdown with the power button. | Software change control | Philips sent an Urgent - Medical Device Correction letter dated August 24, 2012 via certified mail to all US consignees. Philips Key Market representatives distributed the letter internationally. The letter identified the affected product, problem and actions to be taken by customer/user. Philips is implementing a software update to correct the issue free of charge. A Philips representative will contact affect sites to schedule the software installation. For further information or support contact your Philips representative or local Philips Healthcare office. |  | 93 | 93 units | Worldwide Distribution - US Nationwide including the states of: AL, CA, GA, MA, MD, NC, OH, PA, TX, and WA; and to the following countries of: Algeria, Oman, Australia, Brazil, China, Saudi Arabia, United Arab Emirates, and India. | Terminated | December 24, 2013 | 483 | 93 | Computer
Z-2415-2012 | 63001 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Kimberly-Clark MIC Safety PEG Kit - Pull Type |  |  | June 22, 2010 | 2010 | Kimberly-Clark Corporation | The placement guidewire diameter is too large and does not pass-through the introducer needle during the placement procedure of the PEG Feeding Tube. | Nonconforming Material/Component | Kimberly-Clark Health Care sent an Urgent Voluntary Medical Device Withdrawal letter dated June 22, 2010, to all affected customers. The letter included a Product Withdrawal Response Sheet which was to be returned to Kimberly-Clark.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine and destroy all affected products and complete the attached Product Withdrawal Response Sheet and Fax the form to 920-225-4182.For questions customers should call 770-587-7362.For questions regarding this recall call 770-587-8393. |  | 277 | 277 units | Nationwide Distribution including AL, AK, AZ, CA, CO, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, ME, MA, MI, MN, MS, MO, NE, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI. | Terminated | February 07, 2013 | 961 | 1043 | Not_Computer
Z-2423-2012 | 63146 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Asahi KASEI, REXEED(TM)- 15 LX, HEMODIALYZER |  |  | September 21, 2012 | 2012 | Asahi Kasei Medical Co., Ltd., Okatomi Plant | Asahi became aware of a negative trend for rate of the adverse events that may be associated with a manufacturing change to certain lot numbers of the products, though there is no sufficient information to date to conclusively link these adverse events to the products. | Process change control | AsahiKASEI Medical Co., LTD sent a Medical DeviceRecall letter on September 21, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to confirm that they have no inventory and to forward the letter to all clinics who may have received the product requesting they check their stock and arrange for the return of any recalled product. Non-responding consignees will be contacted by telephone.For questions regarding this recall call Asaki Kasei Medical America, Inc. at 888-362-6105. |  | 24 | 24 units | REXEED SX: Worldwide Distribution - USA including TN and the countries of Japan and GermanyREXEED LX: Worldwide Distribution - USA including TN and Germany | Terminated | February 05, 2013 | 137 | 325008 | Not_Computer
Z-2417-2012 | 63129 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | VerifyÂż Bowie-Dick Test Card, Model #S3098, Steris Biological Operations, 9325 Pinecone Drive, Mentor, OH 44060 |  |  | July 03, 2012 | 2012 | Steris Corporation | STERIS has learned that the Verify Bowie Dick Test Cards are not performing to product specifications. Specifically, test cards have resulted in false fail results when sterilizer performance is within acceptable ranges. | Nonconforming Material/Component | Steris Corporation sent a Urgent Voluntary Recall Notice dated July 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine their inventory and compare the lot numbers printed on the box with the listed lots. If they have products from the affected lot numbers to contact STERIS Customer Service for account credit. After notifying STERIS of the amount of recalled product on hand they are to immediately destroy the product. Distributors are to also notify their customers of the recall. They are to send a copy of the Customer Recall Notice. They are to inventory their stock and notify STERIS of any remaining recalled product for replacement product. All recalled product should be destroyed.Please contact your STERIS Account Manager or STERIS Customer Service at 1-800-548-4873. |  | 629 | 629 cards | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CO, CT, FL, GA, IA, IN, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK,OR, PA, RI, SC, TX, UT, VA, WA, WI, WV and WY. and the country of Canada. | Terminated | March 27, 2013 | 267 | 629 | Not_Computer
Z-2413-2012 | 63069 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Ventralight ST Mesh. - Ellipse 7" x 9" (17.8 cm x 22.9 cm) |  |  | September 04, 2012 | 2012 | Davol, Inc., Subs. C. R. Bard, Inc. | Mislabeled: The 8" x 10" package contains 7" x 9" VENTRALIGHT ST Mesh product and the 7" x 9" package contains 8" x 10" VENTRALIGHT ST Mesh | Mixed-up of materials/components | Davol issued an "URGENT: MEDICAL DEVICE RECALL" letter dated September 4, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter advises customers to examine their inventory and quarantine affected product for return to the firm. Customers were asked to complete and return a Recall Effectiveness Check Form. Contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. |  | 49 | 49 units | Nationwide Distribution. | Terminated | September 03, 2013 | 364 | 49 | Not_Computer
Z-2430-2012 | 62736 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Model 324JL On-Board Suction System |  |  | March 30, 2012 | 2012 | Impact Instrumentation, Inc. | A shipment mix-up took place. | Process control | The firm, Impact Instrumentation Inc., contacted their customers via phone and email notification on April 9, 2012. Customers were provided the correct packing slip and asked to acknowledge the correction.If you have any questions, contact the Regulatory Affairs Manager at 973-882-1212 or email: aglordano@impactii.com. |  | 2 | 2 devices | Nationwide distribution: USA including states of: FL and OH. | Terminated | July 21, 2015 | 1208 | 2 | Not_Computer
Z-2425-2012 | 62577 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ULTRA Fabric Reinforced Surgical Gown, X-Large, Sterile, Test File. |  |  | December 31, 2010 | 2010 | Kimberly-Clark Corporation | The impacted product was not intended for commercial distribution to end users. | Packaging process control | Kimberly-Clark sent a Urgent Medical Recall Notice dated on December 31, 2010. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to: Evaluate their Distribution Facilities Inventory of Ultra Fabric Reinforced surgical Gown to determine if any of impacted product remained within their facility. If present, to discontinue distribution of the impacted product and place in a quarantined location. Complete the attached Distributor Product Withdrawal Response Form and fax it to KC within five days. Destroy the impacted product according to their facility procedures. And to review their distribution records to identify all customers who were shipped any of the impacted products. Notify each of their customers who were shipped any of the impacted products by sending them a copy of the attached Customer Letter with the Customer Withdrawal Response Form. A minimum of three (3) attempts to contact each non-responding customer within interval of approximately four to six weeks will be performed. Once all customers who were notified have responded or after approximately four weeks after the third attempt to contact the last non-responder, a recall Termination will be initiatedIf your Kimberly-Clark Representative has not already contacted you, please contact the Ultra Fabric Reinforce surgical Gown Coordinator at (770) - 587-7700. |  | 32 | 32 cases | US (nationwide) including the states of CA, CO, LA, NM, OK, TX and WA. | Terminated | November 01, 2012 | 671 | 32 | Not_Computer
Z-2411-2012 | 63011 | No | HPT | Perimeter. | Ophthalmic | Ophthalmic | 510(K) Exempt | Octopus 101 Perimeter; PeriTrend Software Analysis |  |  | August 17, 2012 | 2012 | Haag-Streit USA Inc | The firm recalled the device after learning of a possible deviation in the stimulus luminance of the Perimeter OCTOPUS 101. | Equipment maintenance | The firm, HAAG-STREIT USA, Inc., sent a "Field Safety Notice 2012-01/01" dated August 10, 2012 via certified mail or onsite visit to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to make all users of the product and relevant personnel aware of this Field Safety Notice; forward a copy to third parties, if affected product has been passed on; keep this notice at least until the completion of the corrective measure, and complete and return the enclosed Confirmation of Receipt Form via fax to: 513-229-3867 or scanned and emailed to : fieldsafetynotice@haag-streit-usa.com.If you have any questions, call 513-336-7255. |  | 64 | 64 units | Worldwide distribution: USA (nationwide) including states of: AR, CA, CT, FL, GA, IA, IL, LA, MA, MI, MN, MS, NJ, NV, NY, OH, OR, PA, SC, SD, WA, WI, and WV; and to the countries of Jamaica, Puerto Rico, and Mexico. | Terminated | May 29, 2015 | 1015 | 64 | Not_Computer
Z-2427-2012 | 63089 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 12/14 TAPER HIGH OFFSET NECK USE WITH SMITH &amp; NEPHEW STEMS AND 12/14 FEMORAL COMPONENTS ONLY CO-CR, REF 71352112, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA. |  |  | August 22, 2012 | 2012 | Smith &amp; Nephew Inc | Modular necks were manufactured out of specification by a vendor; the taper provide is undersized. | Employee error | Smith &amp; Nephew sent an "URGENT:CLASS II RECALL-HIGH OFFSET MODULAR NECK R1213 letter dated August 22, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Customers were instructed to cease use and distribution of the product and quarantine for return. An Inventory Return Certification Form was included for customers to complete and return via fax to 901-566-7975. |  | 70 | 70 units | Worldwide Distribution-USA (nationwide) including the states of CA, FL, GA, IL, IN, MA, NC, NJ, NY, OH, OK, PA, TN, TX, and WA, and the country of Canada. | Terminated | February 27, 2013 | 189 | 70 | Not_Computer
Z-2431-2012 | 63016 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Sekisui Spectrolyse PAI-1, In-Vitro Diagnostic for Quantitative determination of Plasminogen Activator Inhibitor Type- 1 (PAI-1) activity in human plasma. |  |  | August 28, 2012 | 2012 | Sekisui Diagnostics Llc | SPECTROLYSE PAI-1 activity assay producing lower than expected absorbance values with "0" standard. This will reduce the slope of the standard curve requiring the user to make more frequent sample dilutions. | Process control | The firm, Sekisui Diagnostics, LLC, sent an "URGENT PRODUCTS RECALL NOTICE" dated August 28, 2012 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to immediately stop using the product; to discard any remaining inventory and complete and return the Customer Acknowledgement Form by September 4, 2012 documenting the # of kits destroyed via FAX to Quality Assurance at 203-602-5553. The customers were also instructed to return the form even if they do not have the affected product currently in inventory. Sekisui will replace kits.If you have any questions contact, Customer Service Department at 800-637-3375, ext 10, 12, or 31, or email to recall coordinator at MKHAN@AMDIAG.COM. |  | 699 | 699 kits | Nationwide distribution: USA including states of: FL, IN, IL, and WA. | Terminated |  |  | 699 | Not_Computer
Z-2360-2012 | 62132 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | BR980-9600 Hysteroscopy Sterile Procedure Kit containing 2 oz. bottle PVP Prep Solution |  |  | September 30, 2011 | 2011 | Custom Medical Specialties, Inc. | The firm saw on the FDA web site a recall by the manufacturer of the Povidone Iodine Prep Solution due to no microbial testing and product did not meet proper finished goods specifications. | Material/Component Contamination | Custom Medical Specialties sent an Urgent Medical Device Recall letter dated September 27, 2011 and a follow-up letter dated September 4, 2012 to all affected customers. The letter identified the affected product, problem, and actions to be taken. The letter requested that customers inspect inventory, discontinue use and return any remaining affected products. A Return Authorization Number and shipping number will be given to customer by calling 919-202-8462 ext 205. Customers were instructed to complete and return the attached Acknowledgement and Return Product form. |  | 135 | 75 cases &amp; 60 Bottles for reps | US Nationwide Distribution including the states of CO, FL, KS, MA, MI, NJ and PA. | Terminated | April 29, 2013 | 577 | 994 | Not_Computer
Z-2429-2012 | 62978 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Integra HRAIM Head Ring Assembly with Intubation Mounts |  |  | August 09, 2012 | 2012 | Integra LifeSciences Corporation | T-handle screw is used on a complementary product to the CRW, the Integra HRAIM, which is the CT-compatible intubation Head Ring Assembly. When the T-handle on the HRAIM intubation hoop is tightened, the T-handle could stop in a vertical position that prevents the CRWPRECISE and certain models of the CRW-ASL from seating properly on the head ring although it appears to be fully seated | Device Design | The firm, Integra, sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" dated August 10, 2012 to US &amp; Canada consignees/customers via traceable courier service and Non-US &amp; Non-Canada consignees/customers - through written notification delivered either by traceable courier, email, facsimile, or through a directed visit by an Integra Sales Representative. The notification describes the product, problem and actions to be taken. The Integra sales representative will replace the screws according to the attached rework instructions. If for some reason it is not possible to replace the screws during the visit, the sales representative will leave the screws and customers can then replace them by following the attached rework instructions and discard the old screws immediately after replacement. The customers were instructed to complete and return the Product Recall Acknowledgement Form via the sales representatives visit; fax at 1-609-750-7999 or email to: csnirecalls@integralife.com and include the number of devices to be identified and replaced . Note:The T -Handle screws that are replaced will be disposed of at the consignee's/customers location.Please feel free to contact our service hotline at 888-772-7378 should you have any additional questions. |  | 188 | 188 units | Worldwide distribution: USA (Nationwide) and countries of: Argentina, Australia, Belgium, Canada, China, Czec, Denmark, France, Hong Kong, Hungary, Italy, Japan, Malaysia, Philippines, Singapore, Spain, South Africa, and UK. | Terminated | March 26, 2014 | 594 | 363 | Not_Computer
Z-2426-2012 | 63142 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS 5085 and 5085 SRT Surgical Tables, |  |  | April 06, 2012 | 2012 | Steris Corporation | STERIS has learned from Customer feedback and field service experience that damage from external impact to the table hydraulic column can cause the stainless steel shroud sections covering the table hydraulic column to become misaligned; if the damage is severe, the up/down movement of the table may be impeded. In addition, the table hand control may indicate that the table floor locks are unlock | Device Design | Steris sent an Urgent Voluntary Field Correction Notice dated April 6, 2012 to all affected customer. The letter identified the affected product, description of the problem and actions to be taken. Customers were informed that a STERIS Service Representative will contact affected facility to arrange for the replacement of the stainless steel shroud covering the hydraulic column and relocation of the floor lock pressure switches on their 5085 and/or 5085SRT table(s). For further information or if you have questions regarding the visit to your facility, contact STERIS Field Service Dispatch at 1-800-333-8828.If you have questions regarding this matter, call the Director of Marketing Tables and Case Goods at 440-392-7148 or STERIS Customer Service at 1-800-548-4873. |  | 379 | 379 tables | Worldwide Distribution - US Nationwide including the states of: AL, AR, CA, CO, CT, FL, HI, ID, IL, IN, IA, KY, LA, MA, MD, ME, MI, MN, MO, MS, NV, NYH, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WI &amp; WV and countries of: Australia, Canada, China, India, Korea, Mexico, Panama and Thailand. | Terminated | September 04, 2013 | 516 | 379 | Not_Computer
Z-2414-2012 | 62765 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | S-RomÂż Total Hip System Poly Liner |  |  | August 22, 2012 | 2012 | DePuy Orthopaedics, Inc. | Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of four lots of the S-ROM" liner because the Carton label has an Incorrect Color Code Series Designation. We have received one complaint regarding this issue. | Labeling Change Control | DePuy sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated August 22, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Reconciliation Form was included in the letter for customers to complete and return via fax to 574-371-4939. Contact your local Depuy Orthopaedics sales representative for questions related to this recall. |  | 12 | 12 liners | Nationwide Distribution including the states of MA, TX, OH, CA, and NC. | Terminated | October 09, 2013 | 413 | 12 | Not_Computer
Z-2440-2012 | 62606 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BP7561 OPTI CCA-TS B-Lac Cassette (25 per box) |  |  | July 03, 2012 | 2012 | OPTI Medical Systems, Inc | The values reported for PCO2 by BP7561 B-Lac cassettes are biased at high concentrations when used at altitudes above 2800 feet (853 meters). | Process design | OPTI Medical sent a Urgent Field Correction Notice dated July 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to apply correlation factors to correct the PCO2 measurement reported by the device. They were instructed to determine the elevation (altitude above sea level) for their location and to use the tables provided in the recall notification to determine the correct correlation factor. A website was provided if the consignee did not know their elevation. Steps were provided to enter the correlation factor into their OPTI CCA TS analyzer.Please contact OPTI Medical Technical Support at 1-770-510-4444 or 1-800-490-6784. |  | 895 | 895 boxes (25 cassettes per box) | Worldwide Distribution -- US (nationwide) including the states of AZ, ID, TX and UT., and the countries of Argentina, Bolivia, China, Ecuador, France, Germany Greece, India, Indonesia, Iran, Ireland, Japan, Kuwait, Lithuania, Luxembourg, Malaysia, Mexico, Morocco, Nigeria, Poland, Romania, Russia, Saudi Arabia, Singapore, South Africa, Sudan Switzerland, Turkey, United kingdom, Venezuela, and Viet Nam. | Terminated | April 11, 2013 | 282 | 895 | Not_Computer
Z-2434-2012 | 63015 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Duet TRS 45 4.8 mm Universal Straight Single Use Loading Unit. |  |  | August 21, 2012 | 2012 | Covidien LP | Duet TRS may have the potential to injure adjacent anatomical structures which may result in life threatening post-operative complications. | Device Design | Covidien issued an "URGENT MEDICAL DEVICE RECALL" letter dated August 21, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The firm sent a revised customer recall notification letter via FedEx overnight delivery to all customers who purchased Duet TRS" devices from August 2009- August 2012. This letter went to all customers whether they had been sent the August 21st letter or not. Users were advised that Covidien was recalling all lots of the Duet TRS Universal Straight and Articulating Single Use Loading Units (SULU). On August 23rd 2012, Covidien notified, via e-mail, perspective surgeons that were planning to participate in the "Duet TRS Patient Registry" of the recall and that as a result the Duet TRS registry program was discontinued. Customers were asked to complete and return a Recalled Product Form via fax to (800) 895-6140. Contact Covidien Customer Service at (800) 722-8772, option 1 for questions regarding this recall. |  | 4867 | 4,867 units | Worldwide Distribution. | Open, Classified |  |  | 1078939 | Not_Computer
Z-2442-2012 | 62927 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Sleek ***OTW ***Ref ***426-1501X***Inflated Balloon Dia/mm 1.5 Inflated Balloon Length/mm 15 Usable Catheter Length/cm 150***Lot 50026688***2013-09***Cordis***PTA Dilatation Catheter***Cordis***Johnson Johnson Company***Sterile***EO***Method of Sterilization: ETO***ClearStream***Clearstream Technologies, Ltd.***Moyne Upper, Enniscorthy Co Wexford, Ireland***Phone: +353-53-92-37111***Fax: +353-53-92-37-100***Made in Ireland***Distributed by: Cordis Cashel***Cahir Road, Cashel, Co. Tipperary, Ireland***". |  |  | January 31, 2011 | 2011 | Cordis Corporation | On 1//20/2011 Cordis Corporation , Miami Lakes, FL initiated a recall of their Sleek OTW PTA Dilatation Catheter Model #4261201X Lot #50027149 and Model #4261501X Lot #50026688. Product is incorrectly labeled. | Error in labeling | Cordis (Johnson &amp; Johnson Company) sent an Urgent Medical Device Field Corrective Action Notice dated January 31, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review, sign, and return the enclosed Acknowledgment Form directly to Cordis at the fax number on the form, or to their local sales representative. Pass the revised labeling notice to all who need to be aware within their facility or organization. Also, customers were asked to maintain awareness of this notice until all affected product in their facility has been consumed. In addition to their local Cordis representative, they may contact the normal HCS Customer Service number at 1-800-327-7714, option 1, option 3.For questions regarding this recall call 786-313-2365. |  | 1 | 1 unit | Worldwide Distribution - USA and Germany | Terminated | November 26, 2013 | 1030 | 6 | Not_Computer
Z-2445-2012 | 62440 | No | GQX | Varicella-zoster virus serological reagents. | Microbiology | Microbiology | 510(k) | Trinity Biotech, Captia VZV IgG. Product code 2325600: 96 Tests; product code: 2325601: 480 tests. |  |  | May 29, 2012 | 2012 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | The Positive Control for the Captia Varicella-Zoster Virus (VZV) IgG Kit has been reported as running high out of its assigned ISR range. | Nonconforming Material/Component | Trinity Biotech sent a "PRODUCT RECALL LETTER" dated May 29, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Product Recall Fax Back Form was included in the letter for customers to complete and return via fax to (716) 488-1990. Customers were instructed to discard any of the affected product. Contact the Trinity Biotech Help Desk at 800-325-3424, option 2 for questions regarding this recall. |  | 810 | 810 kits | Nationwide Distribution-including Puerto Rico and the states of AL, CA, IL, IN, MA, MD, NJ, OH, PA, TN, TX. | Open, Classified |  |  | 810 | Not_Computer
Z-2444-2012 | 63186 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SharePlan 1.1, aka t-RayAutoplan version 1.2.1 (build number 1.3.1.10). Manufacturing name RayAutoplan, aka t-RayAutoplan, commercial name (TomoTherapy) SharePlan |  |  | October 10, 2011 | 2011 | RAYSEARCH LABORATORIES AB | This notice concerns behavior of the SharePlan that may be unexpected and not clearly described in the user manual. This behavior has not caused patient mistreatment or other incidents. However, the described user actions must be adhered to for best agreement between treatment plan evaluation and the clinical outcome of the delivered treatment. The behavior appears, when importing RFA-files (.asc) | Labeling mix-ups | RaySearch Laboratories sent an Urgent Field Safety Notice Medical Device Correction letter dated October 10, 2011, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. To completely avoid effects of the erroneous behaviour, RFA-file has to be modified to have the crossline profiles sgtored in the X column of the file and the inline profiles in the Y column. Upon request, RaySearch Laboratories will provide help to the customers by transferring the RFA-files to the correct format through a scrip. Customers were instructed to pass on the notice to all those who need to be aware within their organization and to maintain awareness on this notice until the new labeling version has been received to ensure effectiveness of the corrective action.For questions regarding this recall call +46 8 54 50 61 42. |  | 20 | 20 units | Worldwide Distribution - USA including CA, KY, MI, NY, OK, OR, SD, and WI. Internationalyl to Canada, Sweden, France, Germany, Australia, Poland, Belgium, and United Kingdom. | Terminated | January 25, 2013 | 473 | 20 | Not_Computer
Z-2443-2012 | 62313 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 ASSEMBLY FLOW POD |  |  | September 10, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) received two reports of the false back flow alarms for the TerumoÂż Advanced Perfusion System 1 when there was no tubing in the flow sensor. Terumo CVS' investigation found that the voltage level on a circuit board was improperly set in a population of flow modules for Terumo System 1, the sensor operates correctly if the tubing is installed and it is highly unli | Release of Material/Component prior to receiving test results | Terumo CVS issued an URGENT MEDICAL DEVICE Recall letter dated September 10, 2012 to alert all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that a Terumo CVS representative will contact all affected US account.to schedule an appointment to replace the affected product. The affected product will be return to Terumo for destruction. Non-US Terumo Affiliates will be instructed to remove the devices from the field and provide a record of destruction. For questions call the Terumo CVS Customer Service at 1-800-521-2818 |  | 150 | 150 units | Worldwide and US Nationwide Distribution | Terminated | January 13, 2014 | 490 | 150 | Not_Computer
Z-0573-2 | 3529 | [6 empty] | Brand name: Olympus Bronchoscopes |  |  | November 30, 2001 | 2001 | Olympus America Inc. | Loosened bronchoscope biopsy channel port &amp; microbial contamination of the port. |  | Please be aware that this is not a new recall. The firm took action in 2001; but, due to changes in FDA's recall reporting system this recall is now being posted by the Agency. |  | 5135 | 5135 units (revised final #): 4663 units (revised final #) in USA, 472 units to Canada &amp; South America. | Medical facilities nationwide &amp; Puerto Rico. The devices were also distributed to U.S. Government/military accounts and 53 foreign accounts. Distribution was provided to CDRH &amp; OE on 3/4/02.Foreign distribution: Canada, Mexico, Dominican Republic, Argentina, Brazil, Panama, Chile, Peru, Ecuador, Colombia, Venezuela, Paraguay, Uruguay, Costa Rica, El Salvador. | Completed |  |  | 5135 | Not_Computer
Z-2451-2012 | 63044 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray Systems, (Brilliance 6, Brilliance 16, MX8000 Dual v. Exp, GEMINI Dual), Philips Healthcare Systems, Cleveland, OH. |  |  | August 21, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | A bolt which attaches the oil accumulator within the gantry may fail, resulting in the oil accumulator detaching from the tube housing assembly within the gantry. | Nonconforming Material/Component | Philips Healthcare sent an Urgent Medical Device Correction letter dated August 22, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that no action needed to be taken at that time b the customer or user. A Philips Field Service Engineer would schule a time to inspect and repair all the affected systems within approximately 6 months. For further information customers were instructed to contact their local Philips representative or contact the Philips Healthcare Customer Care Solution Center at 1-800-722-9377 and follow the recorded menu options toreach a Customer Solutions Engineer.For questions regarding this recall call 440-483-7000. |  | 112 | 112 Units | Worldwide Distribution - USA including AL, AZ, CA, CO, FL, GA, IL, KS, KY, MI, MS, NJ, NM, NV, NY, OH, TX, WA &amp; WI and internationally to: Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, India, Italy, Japan, Jordan, Mexico, Saudi Arabia, South Korea, Spain, Thailand, Turkey &amp; United Kingdom. | Terminated | December 24, 2013 | 490 | 112 | Not_Computer
Z-2449-2012 | 62322 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Flow Module |  |  | September 20, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (CVS) received one report of a flow module for the Terumo Advanced Perfusion System 1 not displaying blood flow information on the Central Control Monitor or the Centrifugal Control Unit. The tricolor LED on the flow module was not illuminated. Terumo CVS' investigation found that the malfunction was caused by an improper component on a circuit board used in a populat | Process control | TERUMO sent an Urgent Medical Device Recall letter dated September 20, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Medical Device Recall notice; assure that all users are aware of the notice; and confirm receipt of the communication by faxing the attached Customerr Resonse Form to 1-734-741-6149. Terumo CVS will send a representative to each affected US account to replace the affected devices and return the affected devices to Terumo for destruction. Non-US Terumo Affiliates will be instructed to remove the devices from the field and provide a record of destruction.For question customers should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 72 | 72 units | Worldwide Distribution - USA (nationwide) and the countries of AUSTRALIA, Singapore, Jordan, Japan, Vietnam, Germany, Russia, Saudi Arabia, BELGIUM, Mexico, and CANADA. | Terminated | January 13, 2014 | 480 | 72 | Not_Computer
Z-2453-2012 | 63180 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Hologic Fluoroscan Mini C-Arm Model with InSight- FD Software Version 5.0.1and 5.0.2 |  |  | September 14, 2012 | 2012 | Hologic, Inc. | Fluoroscan Mini C-Arm models -InSight2, version 4.0 and 4.0.1 - DICOM images stored in PACS may be flipped 180Âżaround vertical axis.InSight FD, version 5.0.1 and 5.0.2 - DICOM images stored in PACS may be lipped 180Âż around vertical axis as well as 180Âż around a horizontal axis. | Software design | Hologic issued an Urgent: Medical Device recall letter dated September 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Halogen corrected the anomaly and they provided the customer with the revised version of the software. Customers were asked to update their systems in accordance with the instructions provided with the software included with the letter. Customers were also asked to complete the attached acknowledgement form and fax or mail it to the firm at their earliest convenience. For any questions customers were instructed to call Hologic's Help Desk at 1-800-321-4659.For questions regarding this recall call 781-999-7300. |  | 46 | 46 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia,Belgium, Canada, China, Germany, Ecuador, Spain, France, Great Britain, Greece, Hong Kong, Italy, Japan Korea, Lithuvania, Neherlands, qatar, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Thailand, Taiwan, South Africa, and Switzerland. | Terminated | January 10, 2014 | 483 | 460 | Computer
Z-2456-2012 | 63009 | No | GCT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ***REF LS7700***Qty 1*** EC REP*** ConMed Linvatec Xenon Universal Light Source |  |  | May 20, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec, Largo, FL initiated a recall of their products LS7700 Xenon Universal Light Source and LS7600 Xenon Light Source are intended to be used with an endoscope to provide illumination during endoscopic procedures. | Nonconforming Material/Component | Linvatec Corporation sent an "URGENT Medical Device Recall Notification letter dated May 20, 2011 to its customers via FedEx service. Firm advised their suppliers of the Light Sources potentially has unapproved components and does not meet the finished product requirement. The distributors and end user level are advised if any affected product is in their possession can be used until returned for advance exchange. The customers were instructed to: 1) locate the identified product by serial number at your facility and complete the Reply Form attached; 2) Contact ConMed Linvatec Customer Service (800-237-0169) to receive SR Number Advance Exchange affected product only. Once Customer Service authorizes the SR Number your Advance Exchange Light Source will be shipped to your facility; 3) Once replacement unit is received, utilize the packaging material to return the recalled Light Source to ConMed Linvatec using the enclosed FedEx Airway Bill. Mail or FAX copy of Reply Form to ConMed Linvatec; 4) Complete and return a copy of the Reply Form. For questions contact your local ConMed Linvatec office or email Custseryl@linvatec.com, Customer Service at 800-237-0169 or fax to 727-319-5701. |  | 57 | 57 | Worldwide Distribution - US Nationwide and the countries of Canada, DK, HK, IN, IT, KR, MY, RU, CO, VN, ZA. | Terminated | September 26, 2012 | 495 | 58 | Not_Computer
Z-2446-2012 | 63100 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Embolectomy Catheters (Models A44XX, A4FXX, CE0XXXST, and CE0XXX): |  |  | August 09, 2012 | 2012 | Applied Medical Resources Corp | Applied Medical is recalling specific lot numbers of its Vascular Catheters because they may have a potential for packaging particulate matter to reside on the product. | Material/Component Contamination | Applied Medical sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to complete the attached Recall Notification Confirmation Form and to return it to the firm via e-mail to recall60476587@appliedmedical.com or fax to (949) 713-8832. For product return questions, customers were instructed to contact Sales Operations Administrator at (949) 713-8652.For questions regarding this recall contact the firm at (949) 713-8041. |  | 40238 | 40,238 units for all products in Recall Event | Worldwide Distribution-USA (nationwide) and the countries of United Arab Emirates, Austria, Australia, Belgium, Canada, Germany, Ecuador, Estonia, Spain, France, United Kingdom, Hong Kong, Ireland, Israel, Italy, Republic of Korea, Libyan Arab Jamahiriya, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Portugal, Qatar, Serbia, Saudi Arabia, Taiwan, and South Africa. | Terminated | March 22, 2013 | 225 | 40238 | Not_Computer
Z-2450-2012 | 63092 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Armada 35/ Armada 35 LL PTA Catheters. The device is intended for dilatation of lesions in the vasculature. |  |  | August 20, 2012 | 2012 | Abbott Vascular | Abbott Vascular is recalling the Armada 35 and Armada 35 LL PTA Catheters because they have discovered that some devices may exhibit difficulty inflating and/or deflating. | Process design | An "Urgent Device Recall" letter dated 8/20/2012 was sent to all customers who purchased the Armada 35 and Armada 35 LL PTA Catheters. The letter informed the customers of the problem identified and the action to be taken. A list of lot numbers and an effectiveness check form was sent to customers with the customer notification letter. Customers were instructed to call Abbott Vascular Customer Service at (800) 227-9902. |  | 61799 | 61,799 units | Nationwide(US) Distribution and Puerto Rico. | Terminated | July 15, 2013 | 329 | 61799 | Not_Computer
Z-2454-2012 | 63156 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Progel Pleural Air Leak Sealant, Model #: PGPS002 |  |  | August 13, 2012 | 2012 | Neomend Inc | Product was distributed with an incorrect expiration date. Product was labeled with anexpiration date of 2013-04 and should have been labeled with 2012-09. | Incorrect or no expiration date | Neomend sent a Product Recall Notice letter dated August 13, 2012 to all affected customers. The notification letter identified the product affected, problem and actions to be taken. Customers were instructed to check their stock supply and ask the Central Supply Department to determine whether they have any unused product from the affected lot. If you do have unused units of this product follow the instructions provided. If you do not have any product with the affected lot number complete and return the enclosedTrackingVerification Form. For questions contact your Neomend sales representative. |  | 95 | 95 units | US Nationwide Distribution | Terminated | September 25, 2012 | 43 | 95 | Not_Computer
Z-2460-2012 | 63091 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | Alere Triage TOX Drug Screen 9 Panel, P/N: 94402. |  |  | July 12, 2012 | 2012 | Alere San Diego, Inc. | Alere San Diego is recalling Alere Triage TOX Drug Screen because there is a possibility that after installing the code chip provided with the affected lots: AMP, BAR, and PCP, the threshold ranges may not be consistently set correctly in all meters. | Nonconforming Material/Component | Alere sent a 'NOTICE OF CORRECTION" letter dated July 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers with questions were instructed to contact Technical Services at (877) 308-8287. Customers were instructed to complete the Customer/Distributor Verification Form Notice of Correction even if they do not have any involved product and fax back to Technical Services at (858) 805-8457 or scan and email to responses.ts@alere.com. |  | 591 | 591 kits | Nationwide Distribution. | Terminated | February 01, 2013 | 204 | 591 | Not_Computer
Z-2464-2012 | 63197 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens Magnetom Skyra, Verio, Spectra and Biograph MMR |  |  | September 12, 2012 | 2012 | Siemens Medical Solutions USA, Inc | When positioning the Flex Large 4 Coil (part No. 8625761) off-center in left-right direction as it is used for example for hip and shoulder imaging and, in addition at an off-center position with respect to the head-foot direction, the coil can heat up in the area of the electronic housing (white plastic lids). | Device Design | Siemens Healthcare sent a Customer Safety Advisory Notice dated September 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included updated instruction to all customers.For questions regarding this recall call 610-219-6300. |  | 370 | 370 | Nationwide Distribution - USA including Washington, DC and AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI and Puerto Rico. | Terminated | January 14, 2014 | 489 | 370 | Not_Computer
Z-2458-2012 | 62711 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The MAYFIELDÂż Horseshoe Gel Pads are used in conjunction with the MAYFIELDÂż Base Units and MAYFIELDÂż Swivel Adaptors for neurosurgeries in the prone or supine positions. Those devices provide cranial stability and immobility for various craniotomies and spinal procedures. |  |  | July 27, 2012 | 2012 | Integra Limited | An investigation of customer complaints identified that for two product lots, Left Horseshoe Gel Pads were packaged as a Right Horseshoe Gel Pads and vice versa. None of the complaints resulted in a report of patient injuries or an increase in the time to complete the surgical procedure. A physician was consulted to assess patient risk and concluded that there is no potential harm to patients, onl | Packaging process control | Integra Lifesciences sent an "URGENTURGENT: VOLUNTARY MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 27, 2012 to all affected customers. The letter identifies the affected products, problem and actions to be taken by the customers. Customers were instructed to check their inventory for affected products and quarantine them for shipment back to Integra. Customers were asked to complete the Recall Acknowledgement and Return Forms ASAP whether or not they intended on returning any recalled product. Contact Integra Customer Service at 1-877-444-1114, Option 3 for Technical Support. |  | 94002 | 18 from Lot 093984 for Right Gel Pads | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DC, GA, IA, IL, IN, NC, and NY and the countries of: Canada and Belgium. | Terminated | February 20, 2013 | 208 | 1033003 | Not_Computer
Z-2462-2012 | 63093 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Eclipse Treatment Planning System, Model Number: H48; Eclipse 8.9 and 10.0 |  |  | August 17, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the Eclipse Treatment Planning where an outdated assigned Hounsfield Unit [HU] value can be used for dose calculation when the clock on the Eclipse Client Workstation is not synchronized with the Database/System Server. Treatment of the patient using these values can therefore lead to under or over-dose. | Software design | VARIAN Medical Systems sent an Urgent Medical Device Correction Urgent Field Safety Notice on August 17, 2012, to all affected customers with a description of the problem and user corrective action steps. Customers were instructed to ensure that the workstation clock be synchronized with the DatabaseSystem server clock. Customers were also instructed to contact Varian service for assistance with time synchronization if required. Customers were to complete the attached Proof of Notification Form once they have read the document and return it to Varian Medical Systems. For questions customers should call the following numbers:USA and Canada 1-888-827-4265.Europe +41 41 749 8844.For questions regarding this recall call 650-424-5731. |  | 6087 | 6087 | Worldwide Distribution - USA (nationwide) and Internationally... | Terminated | August 06, 2014 | 719 | 6087 | Not_Computer
Z-2459-2012 | 63141 | No | DRA | Steerable catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Z Flex 270, Steerable Sheath, Sterile EO, Rx only, Distributed by Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537, Manufactured by Greathbatch Medical, 2300 Berkshire Lane North, Minneapolis, MN 55441. |  |  | September 06, 2012 | 2012 | Greatbatch Medical | During the investigation of four returned Z Flex 270 Steerable Sheaths that were used with cardiac cryoablation balloon catheters, the internal PTFE sheath liners were found to display varying degrees of damage that could pose an embolic risk to a patient. | Nonconforming Material/Component | Greatbatch Medical sent a "SAFETY ALERT-Z-FLEX-270 STEERABLE SHEATH" letter dated September 6, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. The letter advised the customer to communicate with their customers to assure the that the recalled sheath model is not used in conjunction with cryoablation catheters. Contact the firm at 763-951-8214 for questions regarding this issue. |  | 787 | 787 | Distributed in MA. | Terminated | January 22, 2015 | 868 | 787 | Not_Computer
Z-2463-2012 | 62830 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DEXIS Imaging Suite - catalog number 1.008.7932, version numbers 10.0.0 through 10.0.3 |  |  | July 30, 2012 | 2012 | Imaging Sciences International, LLC | A condition can occur during renumbering of the patient database using the DEXCopy feature. When renumbering is being performed using the DEXCopy feature and multiple patients with an identical name are present, the software will assign all subsequent records to the first patient record, resulting in mixed patient records. In order for this problem to occur several conditions need to be met: 1) | Software design | DEXIS sent a Medical Device Corrective Action letter dated July 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. An Acknowledgement Form was included with the notification that is to be returned by the consignee to the recalling firm. Customers were instructed to call DEXIS Technical Support at 888-883-3947 as soon as possible to have Technical Support review and potentially correct any issues that may have resulted from its use. To request a free-of-charge upgrade kit, customers should email a request to orders@dexis.com.For questions regarding this recall call 215-997-5666, |  | 366 | 366 | Worldwide Distribution - USA including Washington, DC, AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and the country of Canada. | Completed |  |  | 366 | Computer
Z-2461-2012 | 62965 | chair and table, medical | KMN | Medical chair and table. | General Hospital | General Hospital | 510(K) Exempt | InvacareÂż I-Fit" Shower Chair (Models 9780, 9781, 9781-1), Invacare Corporation, Elyria, OH 44035. |  |  | August 06, 2012 | 2012 | Invacare Corporation | On August 6, 2012, Invacare voluntarily recalled the InvacareÂż I-Fit" Shower Chair (Models 9780, 9781, 9781-1) based on the return, quality and adverse event data received. | Device Design | Invacare sent an "URGENT: RECALL INFORMATION" letter dated August 31, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of any affected product and to call the Recall Support Line at 877-496-5041 for assistance. |  | 21168 | 21,168 chairs | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of AUSTRALIA, CANADA, GUAM, ISRAEL, NEW ZEALAND, THAILAND, TRINIDAD AND TOBAGO. | Terminated | February 11, 2014 | 554 | 21168 | Not_Computer
Z-2468-2012 | 63166 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301A/2 Aquilion One CT System. |  |  | May 13, 2008 | 2008 | Toshiba American Medical Systems Inc | The SureStart function may not operate in the usual manner, and you may obtainabnormal analysis results when using the cerebral blood flow analysis function. | Software design | Toshiba American Medical Systems, Inc. sent an "URGENT: MEDICAL DEVICE CORRECTION" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.customers: The firm advised customers to take the following temporary measures to avoid issues until a corrective action has been taken:I. In SureStart:" Perform scanning with Observation Direction set to VFF when the PatientInsertion Direction and Observation direction are FF." Perform scanning with the Observation direction set to VFH when the PatientInsertion Direction and Observation Direction are HF.2. When acquiring images for 4D perfusion analysis:" Perform scanning with the Scan Direction set to OUT when the Patient Insertion Direction and Scan Direction are HF.Perform scanning with the scan direction set to IN when the Patient InsertionDirection and Scan Direction are FF.Contact the firm at (800) 421-1968 for questions regarding this issue. |  | 3 | 3 units | Nationwide distribution: USA including states of: MA, MD, and NV. | Terminated | September 27, 2012 | 1598 | 3 | Not_Computer
Z-2472-2012 | 63188 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Marathon Flow Directed Micro Catheter, Model #105-5055. |  |  | September 23, 2011 | 2011 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | The recall was initiated because ev3 Neurovascular has identified one lot of UltraFlow Flow Directed Micro Catheters and Marathon Flow Directed Micro Catheters may have been mislabeled. | Error in labeling | Micro Therapeutics Inc, Dba Ev3 Neurovascular sent a recall notification letter dated September 23, 2011 to all their affected customers. The letter identified the affected product, problem, and actions to be taken. The notification letter also provides the customer (physician) with appropriate information regarding the affected products. The letter instructs customer to stop using the affected product, segregate from inventory and return for replacement with conforming product. The notice requests the customer to complete and return the reply form indicating use or return of devices in the affected lot(s). Customers with questions regarding the recall letter were instructed to contact Kim Salceda at (949) 680-1305. |  | 23 | 23 units total | Worldwide Distribution - US Nationwide including the states: MN, FL, OR, TX and the countries of: Brazil, Belgium, Czech Republic, Finlind, France, Germany, Netherlands, Poland, and United Kingdom. | Terminated | September 28, 2012 | 371 | 59 | Not_Computer
Z-2467-2012 | 62948 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY LIQUIBEADS Vaginal Moisturizer; Distributed By: McNeil-PPC, Inc. Skillman, NJ 08558-9418; UPC 380041087704; 380040087958. |  |  | July 02, 2012 | 2012 | Mc-NEIL-PPC, Inc. | Johnson &amp; Johnson Healthcare Products (JJHP) is voluntarily recalling this product to address FDA concerns that the product should have been cleared under a separate medical device application. This voluntary recall is not being conducted on the basis of any consumer safety or product performance issues related to K-Y LIQUIBEADS. JJHP received a Warning Letter on May 23, 2012 indicating that thi | No Marketing Application | Johnson &amp; Johnson sent Recall Letters/Response forms dated June 29, 2012 on July 2, 2012. The recall notification included affected lots and instructions for wholesale distributors and retailers that handle the product. The recalled product is being handled by Stericycle. The notification included phone and fax numbers in the event assistance is required. Customers were asked to complete the enclosed business Reply Form and return immediately; identify all warehouse/wholesale inventories of products and return immediately per the attached "Return Authorization Form." For any questions regarding the form, customers were asked to call Stericycle at 888-406-9309. Customers could also fax questions to 877-877-0318. Once customers have collected all returns, they should send back immediately per the enclosed instructions.For questions regarding this recall call 908-874-1000. |  | 818712 | 818,712 | Nationwide Distribution | Terminated | September 27, 2013 | 452 | 818712 | Not_Computer
Z-2470-2012 | 62419 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | This voluntary medical device correction is related to Alcon products that present a 23G nonvalved trocar cannula and infusion cannula together. These items are available as components within Accurus 23G TotalPLUS Pak, Accurus 23G Combined Procedure Packs (0.9mm and 1.1mm), Constellation 23G Total PLUS Pak, Constellation 23G Combined Pack (0.9 mm and 1.1 mm) and 23G Small Parts Kit. |  |  | May 29, 2012 | 2012 | Alcon Research, LTD. | This medical device correction is related to Alcon products that present a 23G non-valved trocar cannula and infusion cannula together because of an issue regarding the loose fit between the 23G non-valved trocar cannula and infusion cannula that could potentially cause a detachment between the non-valved trocar cannula and infusion cannula during surgery. | Device Design | Alcon sent an Urgent a Voluntary Medical Device Correction letter on May 9, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed NOT to use the affected product, instead customers should use a 1-Count Valved Trocar Cannula and to complete and immediately return the Medical Device Correction Response Form to Alcon via fax at: (817) 916-9087 or email to: ReplacementCannulas@alconlabs.com. For questions customers should call 1-800-862-5266. |  | 93319 | 93,319 (30, 413 distributed domestically &amp; 62,906 internationally) | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, United Kingdom, Greece, Singapore, Germany, Switzerland, Kuwait, United Arabian Emirates, and Austria. | Terminated | February 12, 2013 | 259 | 93319 | Not_Computer
Z-2465-2012 | 63119 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost-Mobile Detector Holder for Digital Diagnost (Single Detector, Dual Detector, Release 3) |  |  | September 06, 2012 | 2012 | Philips Healthcare Inc. | The Wireless Portable Detector (WPD) is not fixed securely inside the Mobile Detector Holder and may falldown | Device Design | Philips Healthcare sent a Urgent Field Safety Notice dated August 13, 2012, to all affected customers. Notifications began the week of September 6, 2012, for US and Worldwide Customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips field service engineer will Secure the 4 countersunk screws, which are used to attach the lock mechanism, with thread locker (loctite). Further information or support concerning this issue, contact your local Philips representative. 1-800-722-9377. |  | 93 | 93 Units | Worldwide Distribution--USA (nationwide) and the countries of Australia, Austria, Canada, Chile, Finland, France, Germany, Hungary. Italy, Japan, Latvia, Netherlands, New Zealand. Oman, South Korea, Switzerland and United Kingdom. | Open, Classified |  |  | 93 | Not_Computer
Z-2466-2012 | 62979 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | BariAir Therapy System |  |  | September 17, 2012 | 2012 | KCI USA, Inc. | KCI has issued a medical device correction for the BariAir Therapy System Hand Pendent. The BariAir Therapy System is not designed to allow the patient to place the bed into the patient exit position. KCI has received complaints indicating that patients used the hand pendent to inadvertently place the BariAir Therapy System into a pseudo-exit position. In a small number of instances, this allow | Device Design | KCI sent an Urgent - Voluntary Medical Device Correction letter dated September 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers would be contacted by KCI to coordinate a shipment of a replacement product. For questions customers would instructed to call 1-800-275-4524, Select Option 3 followed by Option 2.For questions regarding this recall call 1-800-275-4524. |  | 575 | 575 | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, United Kingdom, Singapore, Germany, Switzerland, Kuwait, United Arabian Emirates, and Austria. | Terminated | July 29, 2013 | 315 | 575 | Not_Computer
Z-2469-2012 | 63107 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Catheter Irrigation Set labeled in part: ***24K Arthroscopy Outflow/Suction Tubing Set*** |  |  | July 01, 2011 | 2011 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec recalled their 24k Arthroscopy Outflow/Suction Tubing Set (Product Number 24k100)because there is a possibility that the product may have a breach in the Tyvek seal that could potentially compromise the sterility of the contents. | Employee error | ConMed Linvatec sent a Urgent Medical Device Recall Notification via FedEx to all affected customers. This notice should be received July 5, 2011. The letter identified the product the problem and the action needed to be taken by the customer. All returned product under this removal action will be identified, segregated, and quarantined per ConMed Utica, N.Y. procedures. Returned product will be dispositioned in accordance with ConMed Utica, N.Y. procedures. This can include repackaged and resterilization with new lot number or product scrap. ConMed Utica, N.Y. will maintain control of all individual lot numbers returned and dispositioned.Firm is striving for 1005 effectiveness. If a Reply form is not received within approximately 30 days of the first mailing, ConMed Linvatec will make two additional contacts by mail approximately 30-35 days apart.Should you have any additional questions please contact our Customer Service at 800-237-0169 or fax at 727-319-5701. |  | 30064 | 30064 units | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA ,WI and Hawaii., and the countries of Australia Canada, China, Switzerland, Germany, Czech Republic, Denmark, Egypt, Spain, France, South Africa, United Kingdom, Israel, Italy, Netherlands, New Zealand, Poland, Portugal, Russia and Saudi Arabia. | Terminated | July 25, 2013 | 755 | 30064 | Not_Computer
Z-2072-2012 | 62621 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS X-ray Medical Diagnostic , X-ray Equipment System |  |  | June 19, 2012 | 2012 | Eos Imaging Inc | It was discovered of several generator failures of the EOS System due to failure of its anode controller board that drives the rotation of the X-ray tube anode. | Radiation Control for Health and Safety Act | EOS Imaging notified users by letter on 6/19/12 and that an on site correction would be performed by an EOS Imaging Field Service engineer.EOS indicated in its notification that although this specific defect has not been observed in EOS systems installed in the U.S, they will replace the anode driver boards with new boards equipped with high current capacitors. Similarly, the generator firmware will be upgraded with the software patch that improves the cooling fans' controls. This will be replaced at no charge to the customers, during regularly scheduled preventative maintenance of the user's EOS system. This should be completed by the end of calendar year 2012. Your proposed CAP appears to adequately address the issues you raised in your notification, and is hereby approved.Your proposed customer notification letter as revised, appears to adequately provide instructions with respect to the use of the product pending the correction of the defect and it appears to provide a clear evaluation in nontechnical terms of the potential hazards related to the defect. Your proposed notification letter is hereby approved for distribution to all affected parties. Further questions please call 678-564-5400. |  | 12 | 12 EOS Systems installed in the US | Nationwide distribution | Terminated | September 16, 2013 | 454 | 12 | Computer
Z-2494-2012 | 63165 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost with patient carriage |  |  | September 12, 2012 | 2012 | Philips Healthcare Inc. | Hook does not securely hold the footplate in vertical position, causing the footplate to disengage and may fall on patient or operator | Device Design | Philips Healthcare sent an Urgent Field Safety Notice letter dated August 15, 2012 to all affected customers. The letter identified the affected product, problem and the actions to be taken to avoid or minimize the occurrence of the issue. Customers were advised to follow the instruction for Use and that Philips plans to replace the old hook with one that has a new design. For questions contact your local Philips representative 1-800-722-9377. |  | 536 | 536 units | Worldwide Distribution - US Nationwide and the countries of :Australia, Austria, Bahrain, Belgium, Canada, Chile, China, Czech Republic, Denmark, Finland, , France, Germany, Ireland, Italy, Japan, Netherlands, New Zealand, Norway, Poland, Qatar, Romania, Russia, Saudi Arabia, Sweden, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan. Thailand, Turkey,United Arab Emirates and United Kingdom. | Terminated | July 26, 2013 | 317 | 536 | Not_Computer
Z-2493-2012 | 63125 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Cemented Tibial Drill |  |  | September 07, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a recall of the Cemented Tibial Drill for the Zimmer Persona Personalized Knee System due tothree reports being received as of August 29. 2012 for tibial cortex perforation during preparation of the tibial bone.The investigation into the reports that have been received found that the perforations are occurring while using theCemented Tibial Drill. The perforations have o | Device Design | Zimmer sent an Urgent Medical Device Recall dated August 2012 to distributors via electronic mail and Hospital risk managers and surgeons had notifications hand delivered by sales representatives. The recall notice describes the reason for the recall, emphasizes surgical technique to be used, includes instructions for delivering notifications to surgeons and risk managers, instructions for replacing the affected product, and instructions for returning the affected product. Customers were instructed to locate the affected product and contact their sales representative for replacement. Consignees are asked to fill out the Certificate of Delivery, scan it and send back per the instruction provided. |  | 60 | 60 | Worldwide Distribution - US Nationwide including the states of: AZ, CA, CO,FL, MN, IL, NJ, NY, OH, OR and the countries of Austrailia, Utd. Arab Emir., France, Germany, India, and Italy. | Terminated | April 17, 2014 | 587 | 60 | Not_Computer
Z-2492-2012 | 61796 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Tabs Professional Voice + Monitor, with wire bracket, Model 25223. Stanley Security Solutions, Inc., Lincoln, NE. |  |  | May 04, 2012 | 2012 | Stanley Security Solutions, Inc. | The monitor may fail to sound an alarm | Device Design | Stanley Security Solutions, Inc. sent an Important Recall Notice dated May 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease using any of the affected product. The customers were informed, however, that they may continue to use the units as long as they were powered with a 9-volt alkaline battery in accordance with current instructions because the issue only occurs when using the 12V DC adaptor. The letter also informs the consignee to expect a follow-up notice no later than 6/1/12 that will provide additional instructions.The recalling firm issued additional letters dated 6/5/12 to distributors and end users with updated instructions. The letter informed them further testing revealed the possibility that a subset of affected monitors may also fail to alarm when powered by a 9V battery if weight is rapidly applied and removed from the pressure pad. Therefore, the recalling firm decided to recall and repair this subset of affected monitors and provided the serial number ranges that must be returned for repair. The letter also provided a list of serial number ranges that could be corrected in the field by permanently blocking access to the 12V DC jack on the monitor. Instructions were enclosed to show the customer how to disable to the jack. In addition the distributor letters request the customer to provide the recall information to their downstream customers. A response form was included to report the amount of affected product in the customers' inventory.For questions customers were instructed to call 888-224-2460.For questions regarding this recall call 407-271-2097. |  | 12111 | 12,111 units | Worldwide Distribution - USA (nationwide) and the countries of Canada and Great Britain. | Terminated | September 24, 2013 | 508 | 23938 | Not_Computer
Z-2486-2012 | 63073 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Optilock T15 AO Driver (3.5mm). |  |  | August 23, 2012 | 2012 | Biomet, Inc. | Field complaints indicate that recent shipments of the 14-400795 and 36505 drivers do not properly mate("lock") into some of the specified screw heads. Drivers near MMC (Maximum Material Condition) or that are misshapen may not properly mate with screws at MMC. Design dimensions and tolerances of the drivers are not properly specified to prevent an incorrect fit between the drivers and the screw | Mixed-up of materials/components | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE'' dated August 23, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. A Fax Back Response Form was included for customers to complete and return via fax to 574-372-1683. Contact the firm at 574-372-1570 for questions related to this recall. |  | 363 | 363 units | Worldwide Distribution-USA (nationwide) including the states of AL, GA, MT, TA, AZ, WA, CA, TA, IN, NY, NJ, VA, MA, MI, MD, PA, KY, NC, UT, NH, FL, OK, CO, LA, OH, MO and SD, and the countries of Japan, Netherlands, Australia, and Costa Rica. | Terminated | August 27, 2013 | 369 | 363 | Not_Computer
Z-2487-2012 | 62594 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Restoration Modular Cone Body Trials |  |  | April 02, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has become aware of a burr, or a metal shaving, observed in the internal threaded portion of the Cone Body Trial. The typical size of the shaving observed was 1-2 mm in length and 0.037 mm squared in cross section. | Device Design | Stryker Orthopaedics sent a Urgent Product Field Action - Product Inspection Branch Acknowledgement Form on April 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.All inspections must be completed immediately as per the provided Product Correction Bulletin. Please document the results of your inspection on the Tracking Log, indicating conforming devices as "Not Affected" and non-conforming devices as "To be returned." If the devices listed on your tracking log have previously been returned or are otherwise not on hand, indicate the appropriate disposition following our usual process.Please assist us in meeting our regulatory obligation by faxing back the attached Product Acknowledgment Form and Tracking Log within 10 business days to 855-251-3635. Any further questions please call 201-831-5523. |  | 10983 | 10, 983 units | Nationwide distribution | Terminated | June 13, 2013 | 437 | 10983 | Not_Computer
Z-2480-2012 | 62110 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | 00780302601 MIS 12/14 Slotted Femoral Head Provisional -3.5X26 |  |  | April 01, 2011 | 2011 | Zimmer, Inc. | Complaints have been received reporting sub-optimal fit between provisional femoral heads (VerSys and MIS) and the mating components (ie stems and provisional femoral necks), in some cases alleging the fit was too loose and in other cases, alleging the fit was too tight. Intraoperative dislodging of the provisional head and the provisional head being left in-vivo after separating from mating comp | Device Design | Zimmer, Inc. sent an E-mail letter dated April 21, 2011, to all affected customers in the US. A June 2011 E-mail was sent to all customers outside the US. The letters requested Inventory &amp; Operations Managers to check product for proper fit /function, complete the return form with worn provisional heads to Zimmer Inc, Attention Product Service, 1777 West Center St. Warsaw, IN 46580.For questions regarding this recall call 574-372-4753. |  | 231641 | 231,641 | Worldwide Distribution - USA (nationwide) and internationally to Canada, Mexico, South America, Europe, Puerto Rico, China, and Korea. | Terminated | May 16, 2014 | 1141 | 231641 | Not_Computer
Z-0001-2013 | 62773 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | MIRS for MATRIX Spine System. Minimally Invasive Reduction Screw (MIRS) for percutaneous instrumentation of the thoracolumbar spine. |  |  | June 18, 2012 | 2012 | Synthes USA HQ, Inc. | Synthes is initiating a Medical Device Labeling Correction following a detailed review of the current labeling in response to reported complaints. Synthes has identified technique measures and has updated the labeling associated with this system. The labeling reflects updated technique recommendations and emphasizes the importance of avoiding rod mis-alignment. | Labeling design | SYNTHES sent an Urgent Medical Device Labeling Correction letter dated June 18, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to replace the previous version of the labeling with the new labeling provided. In addition, they can access an electronic version of the labeling by accessing Product Literature at the following link: http://us.synthes.com. The new version can be identified by code J11031-B located at the bottom right corner of the back cover. To confirm their receipt and understanding of the information contained in the letter, the recall package also included a response form to be completed by consignee and returned to Synthes.by Fax: 610-251-9005 or Scan/email: FieldAction@Synthes.com.Customers with questions were asked to call 1-800-620-7025, ext 5403 or contact their Synthes Spine Sales Consultant. For questions regarding this recall call 610-719-5377.The firm initiated this recall action on June 18, 2012 by sending an Urgent Medical Device Labeling Correction letter to all accounts. The updated labeling was Included with the notification letter. |  | 1602 | 1602 | Nationwide Distribution including CA, FL, ID, IN, MD, MI, MT, LA, NY, OH, PA, TX, VA and WA. | Terminated | September 10, 2015 | 1179 | 1602 | Not_Computer
Z-0003-2013 | 63152 | No | DHR | Rheumatoid factor immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent, PN 447070 |  |  | January 27, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports of QC shifts, lot-to-lot variation, and an increase in the percent of patient samples recovering &gt;20 IU/mL using the Lots# M012376 and M101865 of IMMAGE RF reagent. The lots identified for the RF reagent might produce false abnormal results. | Nonconforming Material/Component | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated January 27, 2012 with an attached PCA Response Form to all customers who purchased the IMMAGE Immunochemistry Systems Rheumatoid Factor (RF) Reagent. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue use of IMMAGE RF reagent lots M012376 and M101865; call Order Assistance at (800) 526-3821 in the United States, (800) 463-7828 in Canada or contact your local Beckman Coulter representative for replacements; and complete and return the enclosed FAX BACK RESPONSE FORM within 10 days via mail at: Beckman Coulter, Inc., 250 S. Kraemer Blvd, Mail Stop A2.0246, Brea, CA 92822-8000; Attn: Regulatory Affairs or fax to: (714) 961-4234. Note: Beckman Coulter has added additional release testing that will assure future lots RF will meet the reference interval claim of the IMMAGE RF Instructions for Use (IFU).If you have any questions concerning this letter, contact Beckman Coulter Support Center (call Center) on the Beckman Coulter Website at: http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter Representative. |  | 925 | 925 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Terminated | May 30, 2013 | 489 | 925 | Not_Computer
Z-0002-2013 | 63134 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COULTER PrepPlus, Part Number: 286600 |  |  | January 20, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Coulter PrepPlus may experience intermittant reagent level sense failures where the instrument may not produce a warning that a reagent vial is low or out of reagent. | Device Design | Beckman Coulter sent an Urgent Product Correction letter dated January 20, 2012 to all their customers who purchased the affected product. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were instructed to complete and return the Response Form with in 10 days. Customers with questions regarding the PCA letter were instructed to contact Beckman Coulter Service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter Representative. |  | 36 | 36 units total (15 units in the US) | Worldwide Distribution - US Nationwide and the countries of: Australia, Botswana, Canada, France, Germany, Japan, Malawi and Sweden. | Terminated | October 02, 2012 | 256 | 36 | Not_Computer
Z-0004-2013 | 63241 | No | EZD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Male, 16Fr, Straight-tip Intermittent Urinary Catheters M16. The product is distributed individually or in intermediate boxes containing 30 individual, sterile packaged catheters (PA/PE film and dialyzing paper pouches). |  |  | September 19, 2012 | 2012 | Cure Medical LLC | The firm initiated the product recall because labeling on some of the intermediate boxes of catheters identify the contents to be M16C, a coudÂż-tip catheters while the actual contents are M16Fr male, 16Fr, straight-tip catheters with the pouch and the case box labeled as M16Fr male, 16Fr, straight-tip catheters. | Labeling mix-ups | Cure Medical LLC sent an Urgent Medical Device Correction and Removal Notification letter dated September 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter instructed the customers to:a) Check their shelves for M16C with lot number 120425 and quarantine all affected product.b) Verify the contents (pouch label) are M16C and correspond with the intermediate box label.c) If all match, have the dealer sign and return a supplied verification record.d) If any mismatch is found, contact Cure Medical for assignment of a return goods authorization.e) Cure Medical will replace any incorrect product along with processing costs for this situation to the dealer.If affected product was further distributed customers were asked to identify those customers who received them and notify them of the problem. For questions customers were instructed to call 1-800-570-1778. |  | 4200 | 420cases (= 4,200 boxes) | Nationwide Distribution | Terminated | February 07, 2013 | 141 | 4200 | Not_Computer
Z-0005-2013 | 61780 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | IR1200, IR1250(IR1200 Plus), 2020 and One Touch(R) Ping(R) Insulin Pump glucose management system. |  |  | July 16, 2012 | 2012 | Animas Corporation | Wear of the keypad that may lead to damage such as tearing or peeling of the edge of the keypad from the pump. | Nonconforming Material/Component | The firm, Animas Corporation, sent an "IMPORTANT PATIENT NOTIFICATION" letter dated July 16, 2012 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to examine their insulin pump for wear or damage on a regular basis, as part of their routine process of inspecting their pump. If keypad is worn or damaged, the buttons may not respond as intended. If consignees/customers see any damage or have issues with the keypad, they are to contact Animas through their website at www.animas.com/contact-animas or telephone Animas Customer Technical Support at 1-855-230-7574, or the phone number provided on the back of their insulin pump. Animas carefully investigates all complaints related to the insulin pump keypad.If you have any questions contact Animas Customer Technical Support at 1-855-230-7574. |  | 236470 | 236,470 | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, V, VT, WA, WI, WV, WY and Washington, DC and countries of: Austria, Australia, Canada, Czech Republic, France, Finland, Germany, Hungary, Israel, Italy, Mexico, New Zealand, Puerto Rico, Spain, and United Kingdom. | Terminated | January 15, 2014 | 548 | 236470 | Not_Computer
Z-0007-2013 | 62966 | No | JOL | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | BriteField illuminated Suction Device; |  |  | May 12, 2012 | 2012 | Invuity, Inc. | Invuity is notifying customers to return all obsolete BriteField Illuminated Suction Devices from the field. It has not been distributed since March 2011 as the next generation product was introduced at that time. | Other | Invuity sent E-mails to those locations thought to still be in possession of the old model of the device in July 2012. Distributors had been e-mailed in May 2012 to make sure that they had returned any of the device that was in stock.For questions regarding this recall call 415-655-2100. |  | 197 | 197 devices | Nationwide Distribution including AR, GA, SC, CA, OR, UT, and WA. | Terminated | April 23, 2013 | 346 | 197 | Not_Computer
Z-0008-2013 | 63253 | No | HGI | Gynecologic electrocautery and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sterile, single-use, disposable electrosurgical electrode holder. It connects the electrode to the foot activated electric generator. Label reads in part "***GYNEX*** Electrode Handpiece, Foot Activated*** Manufactured for GYNEX***" |  |  | September 06, 2012 | 2012 | Gynex Corporation | The wrong electrical connector pin was attached to the cord during assembly by the manufacturer. this electrical connector is too big to fit into the port on the foot activated generator. this makes the device unusable. | Nonconforming Material/Component | Consignees were first notified by phone with a follow up e-mail or fax. The consignees were asked to complete the response section regarding the destruction/return of the device.For questions regarding this recall call 425-882-1179. |  | 228 | 19 boxes (containing 12 electorde hand pieces each) | Nationwide Distribution | Terminated | October 22, 2012 | 46 | 228 | Not_Computer
Z-0006-2013 | 63196 | No | FGO | Ureteral stone dislodger. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Endoscopic retrieval device. |  |  | November 14, 2011 | 2011 | US Endoscopy Group Inc | When the outer catheter is either coiled or in a contorted configuration outside of the endoscope and the patient, and excessive force is applied to the handle, the jaws may become inoperable and may become stuck in the open condition. | Device Design | US Endoscopy initiated a Voluntary Product Recall letter initiated November 11, 2011. The letter identified the product the problem and the action needed to be taken by the customer.US Endoscopy will provide a free replacement or grant your facility a credit toward a future purchase for every device that is returned to US Endoscopy or verified as having been used or destroyed.For further questions please call 440-639-4494 ext. 378. |  | 1153 | 1153 | Worldwide Distribution -- USA (nationwide) including the states of AL, AR, AZ, CA, FL, GA, HI, IA, IL, KY, LA, MD, ME, MI, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, and WI; and to the following countries: Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Caribbean, Chile, Columbia, Croatia, Czech Republic, Denmark, Ecuador, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, India, Ireland, Israel, Italy, Jordan, Lithuania, Malta, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, and United Arab Emirates. | Terminated | July 02, 2013 | 596 | 1153 | Not_Computer
Z-0009-2013 | 63064 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wound Pro Apex Negative Pressure Wound Therapy Powered Suction Pump; Accuro Medical Products LLC, 2854 Hitchcock Ave, Downers Grove, IL 60515; part number 84635400182, sold inside the Apex Negative Wound Therapy system, part number 84635400181 |  |  | August 31, 2012 | 2012 | Accuro Medical Product LLC | Accuro has received reports of some Apex Pumps either failing to operate or malfunctioning, at times without notifying the user of such failure or malfunction on the unit's display screen. | Device Design | The firm, Accuro, sent an "Urgent: Medical Device Recall" letter dated August 30, 2012 to the customers via e-mail on August 31, 2012. The letter describes the product, problem and actions to be taken. The customers were instructed to remove all devices from use immediately; to immediately complete and return the enclosed Accuro response form via email to; (customer.service@accuromedical.com) or fax to: (1-800-866-0545) . Note: the form still needs to be returned even if you no longer have or use the device. Upon receipt, Customer Service will call the account with instructions to arrange for all devices to be returned to Accuro. General questions were directed to Customer Service at 1-800-669-4757. |  | 919 | 919 units | Worldwide distribution: USA (nationwide) and country of: Canada. | Open, Classified |  |  | 919 | Not_Computer
Z-0011-2013 | 63116 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL M4735A Defibrillator/Monitors, used with the M3516A batteries |  |  | September 11, 2012 | 2012 | Philips Healthcare Inc. | HeartStart XL Defibrillator/Monitor (M4735A) used on a patient when AC power is not in use, with a battery that may fail, the device may shut down unexpectedly without providing the user with a low battery warning message or audible tone. | Nonconforming Material/Component | The firm, Philips Healthcare, sent an "URGENT-MEDICAL DEVICE RECALL/Field Safety Notice" letter dated September 6, 2012 to its customers. The letter describes the product, problem and actions to be taken by the customer and also identifies what action Philips plans to take to remedy the issue. The customers were instructed to operate the XL and AC power instead of battery power alone; use a battery from a different date of manufacture (must not have the R-2011-12 date of manufacture code) if they need to operate the XL on battery power alone; and use an alternative source for treatment or therapy (i.e., backup defibrillator/monitor) if unable to operate with AC power alone or locate a battery with a different date of manufacture than that identified. Phillips will provide a replacement battery(s) to customers as soon as new production is available.If you need any further information or support concerning this issue, contact your local Philips representative or call Quality Specialists,Recalls at 978-687-1501 or email: debra.levasseur@philips.com. |  | 2500 | 2500 units | Worldwide distribution: USA (Nationwide) and countries of: AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, FINLAND, FRANCE, GERMANY HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, REPUBLIC OF KUWAIT LEBANON, MACAU, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NORWAY, OMAN, PERU, PHILIPPINES, POLAND, PORTUGAL, ROMANIA RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VIET NAM. | Open, Classified |  |  | 2500 | Not_Computer
Z-0010-2013 | 63136 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 8 Fr. Bardex Pediatric All-Silicone Foley Catheter with a 3cc balloon. The product is packaged individually with 12 units per corrugated case. The product is sold sterile for single use and is used to provide continuous urinary drainage from the bladder. |  |  | September 07, 2012 | 2012 | C.R. Bard, Inc., Urological Division | The 8 Fr. Bardex All-Silicone Foley Catheter with 3cc balloon had an incorrect inflation volume (10 ml instead of 5 ml) printed on the valve cap. | Employee error | An "Urgent: Product Recall Notification" was sent to all customers and distributors in September 2012 via Federal Express. The letter described the product issue and provided recommended actions. Customers can contact Bard at (770) 784-6471 if questions. |  | 4032 | 4,032 units | Nationwide distribution: USA and Puerto Rico | Terminated | March 21, 2013 | 195 | 4032 | Not_Computer
Z-0016-2013 | 63140 | No | CIB | Angiotensin I and renin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active Renin IRMA Kit, Part Number DSL25100 |  |  | January 20, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the Active Renin IRMA Test Kit may exhibit a decrease in stability of the coated tubes. As a result, quality control and patient sample results may be falsely elevated by up to 72%. | Nonconforming Material/Component | Beckman Coulter sent an Urgent Production Correction letter dated January 10, 2012 to all their customers who purchased the Manual Immunoassay Active Renin IRMA Test Kit. The letter provides the customers with an explanation of the problem and actions to be taken. Customers were instructed to discontinue use and discard all remaining affected inventory. Customers were asked to complete and return the enclosed Response Form with in 10 days. Customers with questions contact Customer Service Support at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633. |  | 216 | 216 units total (6 units in the US) | Worldwide Distribution - US Natiowide including the state of: OH, CT, CA and the state of Argentina, Austria, Belgium, Brazil, Canada, China, Columbia, Czech Republic, France, Germany, Greece, Hungary, Italy, Kuwait, Netherlands, Poland, Portugal, Slovakia, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | October 05, 2012 | 259 | 216 | Not_Computer
Z-0013-2013 | 63059 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Buretrol Solution Set, 117" (3.0 m), with 150 mL Burette (Ball-Valve Drip Chamber) (Non-DEHP); a sterile Rx IV fluid pathway device; |  |  | September 07, 2012 | 2012 | Baxter Healthcare Corp. | Baxter has determined that the ball-valve feature of the Buretrol Solution Sets may not function as expected, allowing air to flow past the valve and into the tubing at the completion of dose within the burette. | Device Design | Baxter Healthcare Corp. sent an Urgent Product Recall letter dated September 7, 2012, via first class mail to all affected customers, informing them of the problem with the ball-valve, and instructing them to discontinue use of the affected product and to contact Baxter for instructions to return the affected product.. The customers were also requested to complete the attached customer reply form, confirming their receipt of the letter, and fax it to Baxter at the number provided on the form. Any questions regarding the communication were directed to Medical Information Services at Baxter at 1-800-933-0303. |  | 1995040 | 1,995,040 units | Worldwide Distribution - US Nationwide and Internationally to Canada, Costa Rica, Australia, New Zealand, Hong Kong, and the United Kingdom. | Open, Classified |  |  | 5027417 | Not_Computer
Z-0025-2013 | 63154 | No | LGW |  |  | Neurology | PMA | The Eon Mini Implantable Pulse Generator (IPG) is a 16-channel, rechargeable, multi-program system that allows the connection of one or two leads with four, eight, and 16 electrode leads. It is powered by a hermetically sealed battery within a titanium case. It uses microelectronic circuitry to generate constant current electrical stimulation. Stimulation programs can be delivered as either single stimulation or MultiStim programs depending on patients need. |  |  | September 04, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | Eon Mini IPGs manufactured in April 2012 could potentially exhibit a sudden, brief surge in stimulation that would be felt by the patient as uncomfortable or painful. SJMs routine testing, monitoring, and analysis of manufacturing processes identified three lots of Eon Mini IPGs that the internal battery has the potential to come into contact with the internal microcontroller board. | Nonconforming Material/Component | ST. Jude Medical sent an important Medical Device Advisory letter dated September 4, 2012 to all affected customers. The letter identified the affected product, problem and recommended actions to be taken. For questions contact your St. Jude Medical Neuromodulation Division Representative or call the Technical Services Department at 1-800-727-7846. |  | 56 | 56 units | US Nationwide Distribution | Terminated | May 12, 2014 | 615 | 56 | Not_Computer
Z-0024-2013 | 63211 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Reference materials "RETINAL OCT Analysis and Interpretation Method" shipped with iVue System (K091404) and the "Retina Poster" shipped with iVue System (K091404) and RTVue Systems (K101505); Models: iVue-100, RTVue-100; |  |  | September 07, 2012 | 2012 | Optovue Inc. | Devices labeled for an intended use not included in the existing 510(k) or PMA. Specifically, "The Retina" poster provided with the systems, and the book "RETINAL OCT Analysis and Interpretation Method" by Bruno Lumbroso and Marco Rispoli which are not labeled within 51O(k) cleared product description. | No Marketing Application | Optovue Inc., sent a Voluntary Field Correction and Removal letter dated September 6, 2012, to all affected customers. The letter identified the product the problem and the action to be taken by the customer .Please sign and return the attached form indicating your acknowledgement that that the Affected Materials have been returned or destroyed. If you have any further questions please contact Optovue, at 1-866-344-8948. |  | 795 | 795 units | Worldwide Distribution -- USA (nationwide) and Internationally. | Terminated | March 04, 2013 | 178 | 795 | Not_Computer
Z-0020-2013 | 63174 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Newport Portable Package, Model #50000, Serial # |  |  | April 29, 2012 | 2012 | Accutron Inc | Accutron is recalling the Flowmeter (Analgesia Gas Machine) because when it flows Nitrous Gas without Oxygen gas being turned on, it functions in an incorrect manner. | Device Design | Accutron sent an Urgent Advisory Notice letter dated April 6, 2012 to all their customers who purchased the Ultra PC% Cabinet Mount Flowmeter. The letter informed the customers of the problem identified and actions to be taken. Customers were instructed to discontinue use or to continue use of the device only with extreme care until a replacement can be installed. Extreme care includes the use of the titration method for nitrous oxide administration coupled with close monitoring of both the patient and the actual gas flow through the flowtubes. For question contact Accutron at (800) 531-2221 or by email at jheadrick@accutron-inc.com. |  | 220 | 220 units total | US Nationwide Distribution | Terminated | February 01, 2013 | 278 | 220 | Not_Computer
Z-0021-2013 | 63158 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System, Part Number 629029. |  |  | April 03, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH 800 Coulter Cellular Analysis System because the NRBC, Differential or Reticulocyte mix chamber drain ports may become plugged by rubber pieces from specimen tube caps, and result in biohazard or chemical exposure and/or the potential for generating incorrect results. | Device Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated April 3, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an enclosed Response Form. Customers were notified that Beckman Coulter is in the process of developing and testing a design change that will prevent the mix chamber waste port from plugging. Customers with questions were instructed to contact Beckman Coulter Customer service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter representative. |  | 1281 | 1281 units total (693 units in the US) | Worldwide Distribution-USA (nationwide) and the counties of Australia, Belgium, Canada, Colombia, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Republic of Kuwait, Macao, Malaysia, Mexico, Netherlands, Norway, Oman, Panama, Philippines, Puerto Rico, Qatar, Romania, Russian, Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | June 20, 2013 | 443 | 1281 | Not_Computer
Z-0023-2013 | 63019 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PREVI Isola System, Catalog numbers 29500 and 29500R, bioMerieux, Inc., Durham, NC. |  |  | August 28, 2012 | 2012 | Biomerieux Inc | The instrument may not dispense the specimen onto the pre-poured media resulting in a false negative result. | Labeling design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" letter dated August 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return. Contact Customer Service at 1-800-682-2666 for questions regarding this notice. |  | 350 | 350 systems | Worldwide Distribution-USA (nationwide)-including DC and the states of CA, CO, CT, FL, GA, IL, IN, KY, LA, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TX, and WI, and the country of Canada. | Terminated | September 04, 2014 | 737 | 350 | Not_Computer
Z-0026-2013 | 54725 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoHelix Part #HCS-056-32 3.2MM Cannulated Hex Driver NON STERILE MAT: 465 SS |  |  | January 29, 2010 | 2010 | Orthohelix Surgical Designs Inc | The tips of the drivers are breaking while being used to drive the screws into patients. | Software design (manufacturing process) | The firm, OrthoHelix Surgical Designs, Inc., sent an "IMPORTANT RECALL NOTICE: Mini MaxTorque HCS-056-32 &amp; HCS-056-38 Drivers" dated January 29, 2010 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to discontinue the use of these drivers immediately and return them to OrthoHelix Surgical Designs Inc, 1065 Medina Rd., Suite 500, Medina, OH 44256; ATTN: Driver Recall.Should you have any questions about system, call 330-247-1145 or the Director of Regulatory Affairs &amp; Quality Assurance at 330-247-1444. |  | 47 | 47 Pieces | Nationwide distribution: USA including the states of CA, CO, FL, GA, IL, MO, NC, NH, NJ, NM, NY, OH, OR, PA, SD, TN, TX, VA, WA, and WI. | Terminated | February 20, 2013 | 1118 | 167 | Not_Computer
Z-0022-2013 | 63214 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M Steri-Vac Gas Sterilizer/Aerator models 5XL and 8XL. |  |  | September 13, 2012 | 2012 | 3M Company - Health Care Business | 3M determined that the electrical grounding mechanism for the Steri-Vac 5XL and 8XL gas Sterilizer/Aerator did not meet all of the requirements of the standard in that the electrical wiring, which is insulated and contained within the sealed door should have a stronger connection to the protective electrical grounding mechansim. | Component design/selection | 3m Healthcare Service Center sent a Urgent Medical Device Field Correction letter dated September 19. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.3M will contact you to schedule an appointment for a field correction on any unit in the range of serial numbers provided in the table. 3Ms goal is to service the affected units with minimal disruption to your facility.We apologize for any inconvenience this field correction may cause. Please be assured that 3M is committed to the highest level of safety and customer service and we are proud of our longstanding safety record for 3M" Steri-Vac" Gas Sterilizers.For Further Information If you have any questions, please contact the 3M Health Care Service Line at 1-800-688-5888 or your local 3M representative.. |  | 707 | 707 units (61 US, 646 OUS) | Worldwide Distribution--USA (nationwide) including the states of AL, CO, CT, GA, IN, KY, MA, MI, MN, MO, MS, NJ, OH, PA, SC, TN, VA, WA and WI., and the countries of Argentina, Barbados, Chile, China, Colombia, Costa Rica, Ecuador, Germany, India, Indonesia, Japan, Korea, Kuwait, Lebanon, Malaysia, Mexico, Morocco, Pakistan, Paraguay, Peru, Philippines, Poland, Republic of Guinea, Russia, South Africa, Taiwan, Thailand, Trinidad, UAE, Venezuela, Vietnam and Zimbabwe. | Terminated | January 29, 2013 | 138 | 707 | Not_Computer
Z-0028-2013 | 63070 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Automatic Cutting Needles, Soft Tissue Biopsy Needle, NAC-1820M, Sterile, Remington Medical, Inc., 5830 Meadowridge Court, Alpharetta, GA 30005. |  |  | August 29, 2012 | 2012 | Remington Medical Inc. | Sterility of the product may be compromised. | Material/Component Contamination | The firm, Remington Medical, Inc., sent an "URGENT - MEDICAL DEVICE RECALL" letter dated August 29, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and quarantine the products; complete and return the FaxBack form via Fax to: Quality Engineer Remington Medical, Inc., at 770-887-0732 or email: tracyw@remmed.com; and call the Quality Engineer or customer service at (800) 989-0057 ext 233 to obtain a RGA number and have a call tag issued.Should you have any questions, please do not hesitate to call the Quality Engineer at (800) 989-0057 ext 233. |  | 800 | 800 needles | Nationwide distribution: USA including states of: AL, AZ, CA, FL, IN, MA, NC, NY, OH, PA and TX. | Terminated | January 15, 2013 | 139 | 800 | Not_Computer
Z-0042-2013 | 62968 | No | NEO | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IDS-iSYS N-MidÂż Osteocalcin, Catalog Number IS-2900. Packaged in cardboard Box containing plastic bottles. |  |  | August 10, 2012 | 2012 | Immunodiagnostics Systems Ltd | It has been noted by staff at IDS Inc. that stock of the affected products held at IDS Inc. are labeled with IVD labels although they are not registered for sale as IVDs in the United States. The products should have been labeled as Research Use Only or Investigational use only dependent on their usage. | Labeling False and Misleading | Immunodiagnostics Systems sent Customer notification letters via email and notified customers of the recall. The notification included the reason for recall, product information with codes, lot number, manufacturing/distribution date, and exp date. The notification included instructions to customers: Actions to be taken by the customer:Customers should complete the attached Recall and Acknowledgement Form and take the following actions:1) Check your stock for the products included within the scope of the recall. Cease use and distribution and quarantine all affected product lots immediately.2) Complete the enclosed Recall &amp; Acknowledgement Form and immediately fax back to Immunodiagnostic Systems. This will allow us to document your receipt of this letter and the amount of product you have on hand for return.3) Return any affected product freight collect, along with the original completed Recall &amp; Acknowledgement Form to your local IDS office.Contact information was also provided.For questions regarding this recall call +44 (191) 5195212. |  | 5 | 5 | Nationwide Distribution including NY, CA, ME and CT. | Terminated | December 17, 2012 | 129 | 56 | Not_Computer
Z-0037-2013 | 63157 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ATLANTIS(R) Anterior Cervical Plate System FIXED ANGLE SCREW MODULE, REF 176-504, QTY: 1 EA, Rx only, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132. |  |  | August 30, 2012 | 2012 | Medtronic Sofamor Danek USA Inc | Lines on the screw gauge may not placed on the caddy properly which could result in improper measurement of screws. An incorrect gauge could measure a screw to be approximately 1 mm shorter than its actual length. | Nonconforming Material/Component | Medtronic conducted a conference call on August 30, 2012 and sent an "URGENT VOLUNTARY RECALL" notice dated September 3, 2012 to all affected customers. The notice identified the product, problem, and actions to be taken by the customers. A Recall Questionnaire Form was attached to the letter for customers to complete and return. Contact your Medtronic Sales representative or the firm's Global Quality Department at 800-876-3133, ext. 3197 for questions regarding this notice. |  | 1014 | 1014 units | Worldwide Distribution-USA (nationwide) and the countries of India, Netherlands, Hong Kong, S. L Korea, Singapore, Canada, Australia, Japan, and Mexico. | Terminated | August 15, 2013 | 350 | 2308 | Not_Computer
Z-0031-2013 | 63085 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Caremor Cliner"***Manufactured By Winco***Ocala, FL 34474 www.wincomfg.com***Model No. 536 Serial No. 536A-101475***" |  |  | July 10, 2012 | 2012 | Winco Mfg., LLC | Winco Mfg., LLC in Ocala, FL is recalling the Caremor Cliner model numbers 5351 and 5361. The recall was initiated due to a report that a back frame did not "snap" onto a chair frame properly. | Process control | Winco notified consignees via phone except for the one located in Saudi Arabia, who was contacted by e-mail. All consignees have been sent a follow-up e-mail. The notification identified the product, the problem, and the action to be taken to be taken by the customer. Customers should remove the chair from use and the upholstered back will not be required to be returned. Winco will send out new fully assembled upholstered backs. For questions customers should call 352-854-2929, ext 123. |  | 23 | 23 chairs | Worldwide Distribution - USA including Arkansas, California, Maryland, Mississippi, New York, Pennsylvania and the country of Saudi Arabia | Terminated | May 23, 2013 | 317 | 43 | Not_Computer
Z-0047-2013 | 63169 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Folate WHO Calibrators, Part Number: A98033. |  |  | July 24, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access Folate Calibrators because it may produce falsely elevated results that could mask a folate deficiency. | Nonconforming Material/Component | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated July 24, 2012 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Customers were instructed to complete and return the enclosed Response Forms within 10 days. Customers with questions were instructed to call (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers were instructed to contact their local Beckman Coulter representative. |  | 5266 | 5,266 units total (3,214 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Austria, Australia, Bangladesh, Belgium, Botswana, Brazil, Canada, China, Colombia, Czech Republic, Croatia, Ecuador, Ethiopia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Oman, Pakistan, Panama, Philippines, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Switzerland, Sweden, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, and Zimbabwe. | Terminated | October 17, 2013 | 450 | 5266 | Not_Computer
Z-0033-2013 | 63161 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, Part Number 6607013. |  |  | May 22, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has determined that the specimen and prepared sample stability claims stated in the product labeling for tetraCXP software and CYTO-STAT reagents (CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5 and CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5) need to be updated for alignment with current standards and clinical decision-making associated with lymphocyte immunophenotyping | Process control | Beckman Coulter sent a Product Corrective Action (PCA) letter dated May 22, 2012, with attached PCA Response Form to all customers who purchased the tetraCXP Software, CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5, CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5. The letter provides the customers with an explanation of the problem identified and actions to be taken. Customers were informed that the tetraCXP SYSTEM Guide and CYTO-STAT tetraCHROME CD45-FITC/CD4-RD1/CD8-ECD/CD3-PC5 &amp; CYTO-STAT tetraCHROME CD45-FITC/CD56-RD1/CD19-ECD/CD3-PC5 product labeling will be revised to reflect the updated claims. Customers were instructed to complete and return the enclosed Response form within 10 days. Customers with questions were instructed to contact Beckman Coulter Customer Service at (800) 526-7694 in the US and Canada or contact their local Beckman Coulter representative.For questions regarding this recall call 714-961-4483. |  | 137445 | 137,445 units total (26,155 units in US) | Worldwide Distribution - USA (nationwide) and the following countries:AlbaniaAustraliaAustriaBahrainBangladeshBarbadosBelarusBelgiumBosnia and HerzegovinaBulgariaBurundiCanadaChileChinaColombiaCroatiaCzech RepublicDenmarkEcuadorEl SalvadorFinlandFranceFrench GuianaFrench PolynesiaGermanyGreeceGuadeloupeHong KongIndiaIndonesiaIraqIrelandIsraelItalyJapanKenyaKorea, Republic ofKuwaitLibyan Arab JamahiriyaLithuaniaMacaoMalaysiaMexicoMonacoMoroccoNamibiaNetherlandsNew ZealandNigeriaNorwayOmanPakistanPanamaPolandPuerto RicoQatarReunionRomaniaRussian FederationSaudi ArabiaSingaporeSlovakiaSloveniaSouth AfricaSpainSudanSwedenSwitzerlandTaiwanTurkeyUnited Arab EmiratesUnited KingdomUruguay | Terminated | October 10, 2012 | 141 | 137445 | Computer
Z-0030-2013 | 62832 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL) |  |  | July 30, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed highter imprecision with the recalled lots on Beckman Coulter AU(R) Systems. Also, if the Syva(R) Emit(R) 2000 Phenobartbital Assay (4D019UL) lots D4, D5, D6, and/or D7 as an application on AU(R) Systems is used, the issue may be experienced. Firm has found that within-run imprecison of the recalled lots when run on an AU(R) System may be higher that what would normally be obs | Nonconforming Material/Component | Siemens sent an Urgent Field Safety notice letter dated July 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to share the content of the letter with laboratory director regarding the need to review previous Phenobarbital results, conduct patient follow up and/or repeat testing. For a no charge replacement of remaining inventory customers were instructed to contact Siemens Customer Service Center at 888-588-3916. For questions contact Siemens Technical Solutions Center at 800-227-8994. |  | 79744 | 79,744 cartons | United States Nationwide Distribution in the states of: CA, FL, KS, MA, MN, MS, NM, NY, NC, OR, PA, TN, TX, UT, VA, and WA. | Terminated | August 27, 2014 | 758 | 79744 | Not_Computer
Z-0029-2013 | 63175 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Triage TOX Drug Screen Panel, Model# 94400. |  |  | May 28, 2009 | 2009 | Alere San Diego, Inc. | Biosite is recalling the Triage TOX Drug Screen kits where the box containing test devices because it was mislabeled. | Error in labeling | Biosite sent a Notice of Correction letter dated May 28, 2009 to all their customers who purchased the Triage TOX Drug Screen. The letter informed the customers of the problem identified and actions to be taken. Customers are instructed that performance of Triage TOX Drug Screen test devices is not affected in any way by the misprinted box flap, provided the test is used with a urine sample as per the instructions for use. Customers are instructed to complete and fax the enclosed verification form within 10 days. In the US, customers were instructed to contact Biosite at (877) 441-7440 or email productadvisory@invmed.com. Customers in UK are instructed to contact European Representative, Lesley Paice, lesley.paice@invmed.com. |  | 45638 | 45,638 units | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Brazil, Canada, Colombia, Czech Republic, Finlind, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, India, Indonesia, Italy, Lebanon, Mexico, Netherlands, Norway, Puerto Rico, Romania, Saudi Arabia, South Africa, South Korea, Spain, Switzerland, Taiwan, Tunisia, Turkey, United Arab Emirates, and Uruguay. | Terminated | October 10, 2012 | 1231 | 45638 | Not_Computer
Z-0048-2013 | 63045 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza radiological image processing system. |  |  | June 06, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using syngo.plaza. In syngo.plaza within the online "Short Time Storage" (STS), some studies are marked with an archived icon, however, the entry in the Archive Database referring to the physical file may be missing. | Software design | The firm initiated this correction on June 6, 2012 by issuing update instructions SY129/12/S to consignees. A script will be distributed to all potential affected customer systems which will prevent data loss by setting the delete protection on all affected studies. |  | 11 | 11 | Nationwide Distribution including the states of MO, FL, MA, NY, WI, GA, CA and WA. | Terminated | April 13, 2013 | 311 | 11 | Not_Computer
Z-0049-2013 | 63079 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Battery, 100/115V (Includes charger and connecting cable) |  |  | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2004 of variousmalfunctions of the battery backup system for the Sarns" Centrifugal System.The reports can be broadly categorized as follows:" Failure to charge" Failure to hold charge" Failure to operateIn the most severe case, a battery malfunction could result in an unexpected loss of arterial flow. In | Component design/selection | TERUMO sent an Urgent Medical Device Recall letter dated June 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. On 09/25/2012 an URGENT MEDICAL DEVICE RECALL notification and addendum was sent to all consignees. TCS issued this notification as a Safety Advisory to alert all users of the Sarns centrifugal system of the potential for malfunction and remind them to review the appropriate sections of the Operators Manual for the Sarns Delphin Battery. Providing an addendum to the Operators Manual for the Sarns Delphin Battery with additional instruction on the use and charging of a discharged battery.Customers were instructed to review the Medical Device Recall notice; assure that all users are aware of the notice; place the addendum in the Operator's Manual; and confirm receipt of the communication by faxing or emailing the attached Customer Response Form.For questions customers were instructed to call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 1076 | 1076 | Worldwide distribution: USA including: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, UT, VA, WA, WI, and WV. Puerto Rico and internationally to:ALEXANDRIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CENTRO AMERICA, CHILE, COLOMBIA, Dominican Republic, Ecuador, FRANCE, GERMANY, GREECE, Guatamala, HOLLAND, HONG KONG, India, ITALY, Japan, KL SEGUNDO, KOREA, LEICESTERSHIRE, MALAYSA, Malaysia, Mexico, Mexico, Mexico, NETHERLANDS, Panama, Peru, Philippines, RALEIGH, Republic of Georgia, SANTIAGO CHILE, Saudi Arabia, SEOUL KOREA, SHANGHAI, CHINA, Singapore, South Africa, SPAIN, SWITZERLAND, TAIWAN, Thailand, UNITED ARAB EMIRATES (UAE), Uruguay, USA, and Venezuela, | Terminated | January 15, 2014 | 477 | 1365 | Not_Computer
Z-0055-2013 | 63362 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | 3 PORT "OFF" MANIFOLD w/WASTE BAG, SYRINGE and |  |  | September 28, 2012 | 2012 | ICU Medical, Inc. | ICU Medical, Inc. has identified a potential comingling of components where a 60 drop micro dripchamber may have been included on a 20 drop macro drip set. | Mixed-up of materials/components | ICU Medical sent an "URGENT:MEDICAL DEVICE RECALL" letter dated September 28, 2012 to its customer. The letter identified the product, problem, and actions to be taken by the customer. The customer has been instructed to inspect their inventory for the specific lot numbers, examine for the presence of a metal cannula, and certify destruction of any affected devices. The customer letter included a Verification Form to be completed and returned. Contact ICUCustomer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com" for questions regarding this notice. |  | 60 | 60 | Nationwide distribution: Virginia only. | Terminated | February 19, 2013 | 144 | 60 | Not_Computer
Z-0051-2013 | 62059 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal System Control module, 100/115V |  |  | September 24, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) received one report of loss ofpower to the control module and subsequently the drive motor for the Sarnscentrifugal system during cardiopulmonary bypass. | Component design/selection | TERUMO sent an Urgent Medical Device Recall dated April 13, 2012, to all affected consignees. An Addendum was sent on September 25, 2012. All US Consignees will receive a Safety Advisory and an update to the Operator's Manual via Federal Express to inform them of the reported malfunction and potential hazards and to raise awareness of the benefits of using a retrograde valve and the importance of back-up system availability. TERUMO will include a customer response form and follow up with all affected users to obtain a response to the recall.For questions customers should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 847 | 847 units | Worldwide Distribution - USA (nationwide) and the countries of Bangladesh, Seoul Korea, Chile, Brazil, Venesuela, Thailand, Mexico, India, Taiwan, Canada, Dominican Republic, Guatemala, Ecuador, Germany, Republic of Georgia, Hong Kong, Japan, United Arab Emirates (UAE), Japan, Belgium, Phillipines, Singapore, and South Africa. | Terminated | November 13, 2013 | 415 | 1186 | Not_Computer
Z-0054-2013 | 63111 | No | HXS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TORQ Sternal Closure Device. The common name is TWISTER, WIRE. The model number is TQ01A. |  |  | August 28, 2012 | 2012 | Kardium | The TORQ Sternal Closure Device, Lot 062711, is being recalled due to probability one of the metal parts may be dislodged during use. | Employee error | Kardium Inc. sent an "URGENT: TORQ STERNAL CLOSURE DEVICE RECALL" e-mailed dated August 30, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to discontinue use of the TORQ Sternal Closure Device lot 062711 immediately. Contact the firm at 604-248-8891, x285 for any questions about this recall. |  | 480 | 480 units | Nationwide Distribution-including the states of FL, GA, IL, and NH. | Terminated | February 05, 2013 | 161 | 480 | Not_Computer
Z-0058-2013 | 62703 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Belmont Rapid Infuser, a Fluid Management System, |  |  | September 25, 2012 | 2012 | Belmont Instrument Corporation | Leak from a saline bag caused saline to drip along the back of the Rapid infuser close to the Power Entry Module causing an electrical flash at the power entry area. | Device Design | Belmont Instrument initiated recall on 9/25/12. A "Field Correction" notice was provided that described the product issue and recommended actions. US-based sales representatives visited all facilities to provide the modified power cord/shield. For units sold outside the United States, the modified power cords were sent to distributors who would ensure the new power cords were provided to all facilities within their accounts. |  | 2913 | 2,913 units | Worldwide Distribution including Nationwide (USA) and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, England, Estonia, France, Germany, Hong Kong, India, Israel, Italy, Jordan, Mongolia, Netherlands, New Zealand, Norway, Pakistan, Peru, Puerto Rico, S. Korea, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Turkey, UK, United Arab Emirates, and Vietnam. | Terminated | February 28, 2014 | 521 | 2913 | Not_Computer
Z-0056-2013 | 63155 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AutoMate 1250 and AutoMate 2550, Part Numbers ODL25125 and ODL25255. |  |  | April 11, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports of secondary tubes being labeled with two different barcode labels on the AutoMate 2500 Family (previously referred to as OLA2500 Systems) units which can cause misidentification of a sample when software versions lower than version 4.1 is used. | Labeling False and Misleading | Beckman Coulter initiated an Urgent Product Correction letter dated April 11, 2012, with attached PCA Response Form to all customers who purchased the affected product. . The letter identified the product, the problem, and the action taken by the customer. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers with questions were instructed to contact their local Beckman Coulter representative. Customers with questions should call 714-993-5321.For questions regarding this recall call 714-961-4941. |  | 468 | 468 units total (51 units in US) | Worldwide Distribution - USA (nationwide) and the following countries:ArgentinaAustraliaAustriaBelarusBelgiumChinaCroatiaCzech RepublicDenmarkFinlandFranceGermanyHungaryIndiaIrelandIsraelItalyJapanKorea, Republic ofKuwaitMacaoMalaysiaNetherlandsNew ZealandNorwayPanamaRussian FederationSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited Kingdom | Terminated | February 11, 2013 | 306 | 468 | Computer
Z-0059-2013 | 63191 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Unicel Closed Tube Aliquotter (UCTA) for UniCel DxC 660i, UniCel DxC 680i, UniCel DxC 860i, UniCel DxC 880i Synchron Access Clinical Systems. Part Number A45853. |  |  | July 06, 2012 | 2012 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the current sample syringe pump drive assemblies installed on DxC Integrated ystem with UniCel Closted Tube Aliquotter (UCTA) serial numbers between 547 and 574 may cause a hardware motion error (ID0x08000501). | Nonconforming Material/Component | The recall notification was initiated on July 6, 2012 with Beckman Coulter forwarding a Product Corrective Action (PCA) letter with attached PCA Response Form to all customers who purchased the UniCel DxC 880i, UniCel DxC 860i, UniCel 680i, UniCel DxC 660i Clinical System Integrated Workstations UCTA Sample Syringe Pump Drive Assembly. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were notified that Beckman Coulter will schedule a Service Call to replace the affected sample syringe pump drive assemblies. Customers were instructed to complete and return the enclosed PCA Response Form within 10 days. Customers with any questions regarding the PCA, were instructed to call Customer Support at (800) 854-3633 in the US and Canada. Outside the US and Canada, customers were instructed to contact Beckman Coulter Representative. |  | 18 | 18 units | Worldwide Distribution and US (nationwide) and the countries of China, France, Italy and Taiwan. | Terminated | November 15, 2012 | 132 | 18 | Not_Computer
Z-0057-2013 | 62879 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hemorrhoidal Circular Stapler and Accessories 33mm |  |  | August 03, 2012 | 2012 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery initiated this voluntary global recall for specific production lots of PROXIMATEÂż PPH Hemorrhoidal Circular Stapler and Accessories 33mm (PPH01 and PPH03) and TranstarÂż Circular Stapler Procedure Set (STR10) due to difficulty firing the device which may result in incomplete firing stroke that may result in an incomplete staple formation. | Process change control | Ethicon Endo-Surgery sent an Urgent Device Recall letter dated August 3, 2012 via overnight, Federal Express UPS to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, discontinue use, quarantine the material and return per the instructions provided. Customers were advised to complete the Business Reply Form (BRF) and fax it to Stericycle at 1-888-628-0733 within 3 business days, even if you do not have affected product. The letter states that Ethicon Endo-Surgery will issue credit for recalled products returned by November 5, 2012. Any product returned after November 5, 2012 will not be eligible for credit. For clinical or product support call 1-800-873-3636, option 6 or contact your local sales representative and reference Event 7283. |  | 157881 | 157,881 devices | Worldwide Distribution - US (nationwide) in the states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the countries of: ARGENTINA, AUSTRALIA, BAHRAIN, BELGIUM, BRAZIL, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, MALAYSIA, MEXICO, NEPAL, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, PUERTO RICO, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, and YEMEN. There was one Canadian consignee and 43 U.S. Government consignees. | Terminated | July 12, 2013 | 343 | 157881 | Not_Computer
Z-0061-2013 | 63124 | [6 empty] | CPT Hip System Femoral Stem Petite, Item No. 00-8011-000-05. |  |  | September 13, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc. is initiating a removal of the CPT Modular Stem (item 00-8011-001-00), CPT Stem Petite (item 00-8011-000-05; ) and CPT Femoral Stem 12/14 Neck Taper Extended Offset,(item 00-8114-001-10 ) because devices were packaged in incorrect cartons and distributed internationally and domestically. Zimmer received five complaints of this packaging issue. | Packaging change control | Zimmer Inc. sent an "URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC" letter dated September 13, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Consignees were asked to return all recalled product with a completed Inventory Return Certification Form. Contact the firm at 1-800-348-2759 for questions regarding this notice. |  | 20 | 20 | Worldwide Distribution-USA (nationwide) including the states of Indiana and Oklahoma and the countries of Australia, Thailand, Malaysia, Germany, India, Bulgaria, Egypt, UK, Russian FED, Slovak, and Brazil. | Terminated | January 15, 2014 | 489 | 138 | Not_Computer
Z-0066-2013 | 63115 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | "***PressFT***2.1 w/ Two #0 (3.5 metric) Hi-Fi***Sutures***NP212***2009-02***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA*** |  |  | June 26, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec in Largo, FL is recalling PressFT HI-FI Sutures (items NP211, NP212, NP261, NP262, NP211H, NP261H-all lot numbers) due to the potential for the distal tip of the driver to become separated from the driver shaft after successful insertion of the implant. | Process control | ConMed Linvatec sent an Urgent Medical Device Recall Notification letter on June 26, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their facility for the recalled products, segregate them for return to ConMed Linvatec. The customers were also told to return the REPLY FORM to the Regulatory Department at the firm. For questions customers were instructed to call 800-237-0169 or email at Custserv1@linvatec.com.For questions regarding this recall call 727-392-6464. |  | 879 | 879 units for all products in recall event | Worldwide Distribution - USA including AZ, CA, CO, FL, GA, LA, MA, ME, MI, MN, NJ, NM, NY, PA, TX, UT, and WA. Internationally to BE, CA, ES, MY, NZ and PL | Terminated | October 03, 2014 | 829 | 879 | Not_Computer
Z-0071-2013 | 63225 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Backplate Replacement CXFVR 1500 Holder |  |  | October 02, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | Backplate reservoir holders were manufactured with an incorrect configuration of connector supports. The devices were manufactured and configured with two 3/8 inch and one Âż inch connector supports instead of one 3/8 inch and two Âż inch connector supports (See location of supports as # 7 and # 8 on backplate attachment). The backplates were manufactured as described in manufacturing instructions a | Labeling design | The firm, Terumo, interviewed the Sr. Product Manager from the Marketing Department of the company on August 29, 2012. The interview included the description of the product, problem and action taken. Replacement connectors were provided to users and all units were corrected.If you have any questions, call 734-663-4145 or Manager, Regulatory Affairs at 734-741-5576. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including states of: CA, MI, NC, NJ, PA, and WI;and country of: EUROPE. | Terminated | January 29, 2013 | 1215 | 33 | Not_Computer
Z-0114-2013 | 63027 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare innova 3100-IQ, Mobile X-ray system |  |  | December 20, 2011 | 2011 | GE Healthcare, LLC | It was discovered of a potential failure of the spectral filter mechanism inside the collimator of the innova 3100, innova 3100-IQ and innova 3131-IQ systems. | Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services. A GE Healthcare Service Representative will replace the collimator to address the issue.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. Further questions please call 262-513-4122. |  | 199 | 199 installed in the US | Nationwide distribution | Terminated | September 18, 2013 | 638 | 199 | Not_Computer
Z-0072-2013 | 63359 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ***This recall is being conducted due to an incomplete recall for RES 57829*** Curlin Infusion Administration Set, REF 340-4165, Sterile EO, |  |  | September 26, 2012 | 2012 | MOOG Medical Devices Group | This recall is the same issue for the FDA processed recall initiated by Moog Medical Devices on 1/25/11 which was related to customers reporting unclearable Air-In-Line alarm with specific lot numbers of Curlin Administration Sets used with Curlin Ambulatory Infusion Pumps. The lots associated with this recall were inadvertently omitted from the original recall which has since been terminated by | Process control | MOOG sent a Medical Device Recall Notification letter dated September 26, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove product from their inventory and return to the firm. Customers were also instructed to contact their distributor to arrange for replacement product. For questions customers were instructed to call 1-800-970-2337, prompt #7.For questions regarding this recall call 801-264-1001. |  | 12000 | 12,000 sets | Nationwide Distribution including California | Terminated | April 05, 2013 | 191 | 12000 | Not_Computer
Z-0073-2013 | 62516 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory : |  |  | November 15, 2011 | 2011 | Sunquest Information Systems, Inc. | Online calculations may not be performed as expected or may yield results that are different than expected when the calculation module is called from the Online Entry (OEM) application for a subtest whose subcolumn number on the instrument online string is divisible by 10. | Software design | Urgent - Product Safety Notice: PSN-11-06,dated, 15 November 2011, was sent to customers. Costumers should follow instructions in the notification letter. To request the software correction, create a Case via the Client Support webpage at: www.sunquestinfo.com/supportweb.Return receipt acknowledging PSN-11-06 is required .Client Support can be reached at 1-877-239-6337 for U.S. and Canadian sites and at +44 (0) 161 335 0562 for International sites. |  | 127 | 127 sites | Nationwide distribution | Terminated | October 19, 2012 | 339 | 127 | Computer
Z-0079-2013 | 63336 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LeVeen Standard Needle Electrode System, Sterile (3.5/12), (3.5/15), (4.0/15), (3.5/25), (5.0/15), (5.0/25), and (4.0/25). |  |  | September 17, 2012 | 2012 | Boston Scientific Corporation | Leveen Standard and Leveen Coaccess Needle Electrode System, Directions for Use(DFU) in Portuguese, are incorrect. all other languages in the DFU are correct. Consignees in Brazil, Portugal,and Hong-Kong are being notified. | Labeling Change Control | Boston Scientific sent notifications, dated September 17, 2012, to each account in Brazil, Portugal and Hong Kong affected by the recall. The affected accounts were provided the corrected pages of the Portuguese DFU as part of the communication to ensure that those users that rely on the Portuguese DFU have the correct information. |  | 19 | 19, 716 total all units | Distributed in Brazil, Portugal and Hong Kong. | Terminated | July 17, 2013 | 303 | 19 | Not_Computer
Z-0087-2013 | 62569 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Grafton DBM Crunch |  |  | April 20, 2012 | 2012 | Osteotech Inc | Multiple lots of GRAFTON Âż and XPANSE Âż, which are used for bonegrafting procedures are being recalled because there is a possibility that sterility of the outer surface of the inner pouch may becomecompromised. | Device Design | Osteotech sent out recall letters/questionairres dated April 20, 2012, to Risk Managers/service representatives. The recall was expanded on August 1 2012. The letter identified the product, the problem, and the action to be taken by the customer. Sales Representatives would facillitate the return of any affected product they may have or provide cautionary labels for any affected product and/or support in the completion of the attached questionnaire. Customers were asked to disseminate this information to additional personnel within their facility as appropriate and return the questionnaire to Medtronic.. Customers with questions should contact their Medtronic Sales Representative at 800-876-3133, extension 3197.For questions regarding this recall call 732-542-2800.Osteotech sent an additional letter stating that The OPTIUMTM (DBM) product line was added in error. |  | 3541 | 3,541 units | Worldwide Distribution - USA (nationwide) and Internationally. | Terminated | January 17, 2014 | 637 | 79207 | Not_Computer
Z-0078-2013 | 63365 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hem-o-lok Large Polymer Ligating Clips, REF #544240, Rx only, Sterile, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | September 25, 2012 | 2012 | Teleflex Medical | Product is being recalled due to the possibility that the tray may contain pinholes, compromising product sterility. | Packaging process control | Teleflex sent a Urgent Medical Device Recall Notification dated September 25, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their inventory and quarantine any affected inventory found. They were to complete the enclosed Field Corrective Action Acknowledgement Form and return it to the Teleflex Medical sales representative who will then remove the affected product from the facility and ensure its return to Teleflex Medical.We sincerely apologize for any inconvenience this action may cause your operations. If you have any questions please contact your local sales representative or Customer Service at 1-866-246-6990. |  | 126 | 126 units | Nationwide Distribution including the state of Nevada. | Terminated | November 08, 2012 | 44 | 126 | Not_Computer
Z-0082-2013 | 63223 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Columbus 4-IN-1 Femoral Cutting Guide M3 Instrument (NQ083R) |  |  | May 09, 2012 | 2012 | Aesculap, Inc. | The manufacturer informed the recalling firm of a potential laser marking error in size identification on two Columbus 4-IN-1 Femoral Cutting Guide Instruments. The guides are correctly dimensioned, but the laser markings have been incorrectly etched. | Process change control | Both consignees were notified verbally on May 15, 2012 and asked to inspect their inventory for the recalled items, remove the mislabeled instruments and return them to the recalling firm. Both mislabeled instruments were returned to the firm. Customer questions were directed to (610) 984-9074. |  | 2 | 2 | Within the US, product was distributed to PA and TN. | Terminated | January 15, 2014 | 616 | 2 | Not_Computer
Z-0074-2013 | 63086 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | "***Electric Systems Foot Control****Rx Only***" |  |  | January 09, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. in Palm Beach Gardens, FL is recalling the foot control/foot pedal due to the Foot Control Cable lacking the shield connection from the cable to the connector body. | Nonconforming Material/Component | Anspach sent an Urgent Product Removal letter dated January 9, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to contact their sales representative to facilitate the removal action by providing a replacement for the affected product. For question contact your local sales representative or the Product Support department. |  | 16 | 16 units | Worldwide Distribution - US (nationwide) and the countries of: Italy, Japan, Switzerland, Turkey and Taiwan. | Terminated | September 16, 2014 | 981 | 67 | Not_Computer
Z-0077-2013 | 63255 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Helion and Xenion surgical lights, 501/01 series KLM lights. |  |  | September 19, 2012 | 2012 | Trumpf Medical Systems, Inc. | Incidents regarding fractures of the front joint of the spring arm have been reported. | Device Design | Trumpf sent an "URGENT RECALL NOTICE" dated September 19, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Confirmation of Receipt Form was attached for customers to complete and return. Contact the firm at 888-474-9359 for questions regarding this notice. |  | 275 | 275 units | Nationwide Distribution-USA (nationwide) including the states of AR, CA, CO, CT, FL, GA, IL, KS, MA, MI, MO, NC, NJ, NY, OH, SC, TX, UT, VT and WI. | Terminated | January 21, 2014 | 489 | 275 | Not_Computer
Z-0070-2013 | 63056 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq Two Channel Infuser; an Rx medical device infusion pump used to administer I.V. fluids; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; List Number 16027-01 |  |  | August 29, 2012 | 2012 | Hospira Inc. | The Symbiq pump touchscreen may not respond to user selection, may experience a delayed response, or may register a different value from the value selected by the user. Failure of the touchscreen to respond to user input could result in a delay in therapy or result in over delivery or under delivery of medication if the user does not confirm the programmed values on the pump's confirmation screen prior to starting the infusion. | Device Design | Hospira sent an Urgent Device Correction letter dated August 29, 2012, to all affected customers. Customers have been informed of reports that the Symbiq pump touchscreen may not respond to user selection, may experience a delayed response, or may register a different value from the value selected by the user. Failure of the touchscreen to respond to user input could result in a delay in therapy or result in over delivery or under delivery of medication if the user does not confirm the programmed values on the pump's confirmation screen prior to starting the infusion. The users were instructed to complete and fax the enclosed reply card to 1-888-943-5180 or email it to Hospira194@stericycle.com indicating the number of infusers at the facility and if they notified their users throughout the facility. Wholesalers were asked to notify their customers. If they have further distributed the affected product, they should notify their accounts and ask them to contact Stericycle at 1-877-272-6141 to receive a reply form. For questions regarding this recall call 224-212-2000. |  | 14067 | 14,067 units | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada | Terminated | March 15, 2017 | 1659 | 33681 | Not_Computer
Z-0095-2013 | 61603 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Seno Advantage 1.x, Seno Advantage 2.0, Seno Advantage 2.1 |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the magnification images acquired on non-GE digital mammography systems and your Seno Advantage Workstation that may impact patient safety. The measurement values provided by the Seno Advantage Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems. If not noticed by the caregiv | Software design | GE Healthcare sent a Urgent Medical Device Correction letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will install new software to correct this issue with the release of Field Modification Instruction (FMI) 12129 in the near future. This correction will not affect the performance or any other functionality of your current equipment. A GE Healthcare service representative will contact you to arrange for this modification. This activity will be performed at no cost to you.If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171, Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 188 | 188 | Worldwide Distribution--USA (nationwide) including the states of CA, CO, CT, DC, FL, GA, KY,LA, MA, MO, NV, NJ, NY, NC, OH, OK, PA, SD, TN, TX, VA and WI., and the countries of UNITED KINGDOM, TUKEY, TAIWAN, SPAIN, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL, NEW ZEALAND, NETHERLANDS, MEXICO, MALAYSIA, KOREA, JAPAN, ITALY, HUNGARY, GREECE, GERMANY, FRANCE, CHINA, CANADA, BRAZIL, AUSTRIA and ALGERIA. | Terminated | March 26, 2013 | 306 | 188 | Computer
Z-0108-2013 | 63276 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Pulset Arterial Blood Gas (ABG) sampling kit. |  |  | January 18, 2011 | 2011 | Westmed Inc | Westmed is recalling Westmed ABG convenience kits containing TRIAD brand alcohol pads because TRIAD is recalling TRIAD brand alcohol pads due to potential contaminants. | Material/Component Contamination | Westmed sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 18, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were also given the TRIAD GROUP: "Urgent Drug Recall" and "Facts: Alcohol Recall" both dated January 3, 2011. Westmed customers were instructed to remove the TRIAD alcohol pad from the kit and dispose of the alcohol pad only. Customers were instructed to complete and return the enclosed response form as soon as possible. Customers with questions were instructed to contact the firm at (520) 264-7987 x1233. |  | 1259200 | 1,259,200 units total | Worldwide Distribution-USA (nationwide) and the countries of Canada, El Salvador, United Arab Emirates, Holland, Costa Rica and Lebanon. | Terminated | October 22, 2012 | 643 | 1259200 | Not_Computer
Z-0110-2013 | 63130 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron Systems Synchron Calibrator Multi Calibrator Diskettes, Part Number 442600. |  |  | December 15, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed a mislabeling error, some Synchron Systems Multi Calibrator Diskettes (REF 442600, lot M002642G) may actually be Thyroxine (T4) Reagent Calibrator Diskettes. | Labeling mix-ups | Beckman Coulter sent an Urgent Product Correction letter dated December 15, 2011, withan attached Fax Back Response Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. If customers are currently using Multi Calibrator lot M002642, then they were instructed to verified diskette and discard all Multi Calibrator diskettes lot M002642 and M002642G without a dark green dot. Customers were instructed to complete and return the enclosed Response Form within 10 days. Customers were instructed to contact Beckman Coulter Support Center at http:www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. Outside of the US or Canada, customers were instructed to contact their local Beckman Coulter Representatives.For questions regarding this recall call 714-961-4483. |  | 5746 | 5746 units total (2729 units in US) | Worldwide distribution: USA (Nationwide) and Internationally to: Afghanistan Angola Australia Bangladesh Belgium Bhutan Brazil Brunei Darussalam Bulgaria Canada China Colombia Costa Rica Cote d'Ivoire Czech Republic Denmark Ecuador Egypt El Salvador Eritrea Ethiopia Finland France French Polynesia Germany Ghana Greece Guatemala Hong Kong India Indonesia Ireland Israel Italy Japan Jordan Kenya Kuwait Lebanon Lesotho Libyan Arab Jamahiriya Lithuania Macao Malawi Malaysia Mexico Morocco Namibia Netherlands New Zealand Nigeria Oman Pakistan Panama Paraguay Philippines Poland Puerto Rico Qatar Reunion Romania Russian Federation Rwanda Saudi Arabia Serbia Singapore Slovakia South Africa Spain Swaziland Sweden Switzerland Taiwan Thailand Tunisia Turkey Uganda United Arab Emirates United Kingdom Viet Nam Zambia Zimbabwe | Terminated | December 28, 2012 | 379 | 5746 | Not_Computer
Z-0112-2013 | 63215 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx GNR Traffic Light PNA FISH, In Vitro Diagnostic for Culture Identification Kit for Escherichia coli,and/or Klebsiella pneumoniae and/or Pseudomonas aeruginosa |  |  | September 14, 2012 | 2012 | Advandx Inc | GNR Traffic Light PNA Fish may have false positive results with Pseudomonas putida and Pseudomonas fulva | Device Design | AdvanDx sent a "GNR Traffic Light PNA FISH (KT011) Package Insert Replacement" letter dated September 14, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. All customers received an Acknowledgement Form to complete and return to AdvanDx so that the effectiveness of the replacement can be evaluated. Contact AdvanDx Technical Support at 1-866-376-0009 for questions regarding this notice. |  | 187 | 187 kits | Worldwide Distribution-USA (nationwide) and the countries of Austria, Denmark, Italy, Germany, Finland, and UK. | Terminated | March 05, 2013 | 172 | 187 | Not_Computer
Z-0117-2013 | 63353 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | HiRes 90K devices with Helix Electrodes Model CI 1400-02H |  |  | September 14, 2007 | 2007 | Advanced Bionics Corporation | Advanced Bionics is recalling the the HiRes 90K cochlear implant with HiFocus Helix Electrode because of loose-fitting insertion tool assemblies. | Nonconforming Material/Component | The firm, Advanced Bionics, sent a " RECALL NOTIFICATION" letter dated September 14, 2007 to all their customers who purchased the HiRes 90K devices with Helix Electrodes Model CI 1400-02H. The letter identified the product, problem, and actions to be taken. The customers were instructed to immediately locate and return the affected product to Advanced Bionics. The firm's Service representative will contact the customers within 48 hours to retrieve the affected product and send a replacement as needed. The customers were also instructed to complete and return the enclosed Notification Acknowledgement form via fax to: (661) 362-1511.If you have any questions regarding this letter, call Advanced Bionics at (877) 829-0026 between the hours of 5 AM and 5 PM Pacific Time, Monday through Friday. |  | 17 | 17 units | Worldwide distribution: USA (nationwide) including state: IL and country of: Europe. | Terminated | October 22, 2012 | 1865 | 17 | Not_Computer
Z-0111-2013 | 63254 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | EVICEL Fibrin Sealant (Human) Rx Only. |  |  | July 02, 2012 | 2012 | Ethicon, Inc. | Ethicon and Omrix Biopharmaceuticals have received reports of air or gas embolism occuring in association with the use of spray devices employing a pressure regulator to administer EVICEL; these events appear to be related to the use of the spray device at higher-than-recommended pressures and/or in close proximity to the tissue surface. The field correction initiated is to provide the recommend | Employee error | Ethicon sent a notification letter "Urgent Product Information EVICEL Fibrin Sealant (Human) All Lots and Product Codes" dated June 27, 2012, to all affected customers. Notifications were sent via UPS on July 2, 2012. The notification informed customers of the identified risks and included instructions for use. The notification included a business reply card consignees were requested to fill out and return.Customer were notified by a notification titled "Urgent Product Information Omrix Pressure Regulator All Product Codes" that was sent by Ethicon on August 1, 2012, as a follow up to the letter dated June 27, 2012. The notification informed consignees of the identified risks and included instructions for use. Customers were notified that a company representative would visit them to confirm receipt of the notification and to ensure the enclosed label has been affixed correctly on all Omrix Pressure Regulators.For further questions please call 1-877-384-4266. |  | 6300 | 4800 pressure regulators distributed worldwide; 1500 distributed within the US | Worldwide Distribution -- USA (nationwide) and the countries of Canada, European Union, Singapore, Israel, Mexico and Brazil. | Terminated | September 15, 2015 | 1170 | 6300 | Not_Computer
Z-0116-2013 | 63199 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx E. coli/P. aeruginosa PNA FISH, In Vitro Diagnostic Culture Identification Kit for Escherichia coli and Pseudomonas aeruginosa |  |  | September 14, 2012 | 2012 | Advandx Inc | Some strains of Pseudomonas putida cross-react with the P. aeruginosa PNA probe used in E. coli/P. aeruginosa PNA FISH (KT007) due to a sequence similarity resulting in a cross-reaction that may produce a red-positive signal | Process design | AdvanDx notified customers on September 14, 2012 by email along with a PDF file of the new package insert to each contact followed by written notification and a paper copy of the new package insert via FedEx courier. The notification described the product, problem and actions to be taken. The customers were instructed to destroy the package insert in any existing inventory of the kit and replace it with the enclosed package insert. "Note: All shipments of E. coli/P.aeruginosa PNA FISH as of September 14, 2012, will contain this new package insert." The customers were also instructed to complete and return the Acknowledgement of Receipt form to AdvanDX: Regulatory Affairs; 400 TradeCenter, Suite 6990, Woburn, MA 01801-7476; Fax to: +1 (781) 376-0111; bsc@advandx.com. If you have any questions or requests for further information, please contact AdvanDx Technical Support: US (Toll-Free): 1-866-376-0009, techsupport@advandx.com. |  | 33 | 33 units | Worldwide distribution: USA (nationwide) including: CA, FL, IL, MI, and NJ; and country of: Great Britain. | Terminated | March 05, 2013 | 172 | 33 | Not_Computer
Z-0119-2013 | 63313 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite Tapered Certain Implant 5 X 10mm Model INT510 |  |  | August 03, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i recalled their Osseotite Tapered Certain Implant, Model # INT510 due to a small percentage of the devices are missing the internal threading. | Process control | Biomet 3 i sent an Urgent Medical Device Recall letter dated August 3, 2012, to all affected customers via e-mail, facsimile, telephone, or postal mail. Each customer was instructed to check their inventory for the affected product and provide feedback to Biomet 3i. Product is to be returned to:Biomet 3iP.M. Regulatory Compliance4555 Riverside DrivePalm Beach Gardens, Florida 33410For questions customers were instructed to call 1-800-342-5454. For questions regarding this recall call 561-776-6700. |  | 227 | 227 | Worldwide Distribution - USA including PA, GA, OR, KS, NY, VA, CT, NH, IL, OK, and MA and internationally to Australia, Brazil, Columbia, France, Germany, Ireland, Italy, Japan, Paraguay, Portugal, and Spain. | Terminated | December 20, 2012 | 139 | 227 | Not_Computer
Z-0121-2013 | 63424 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Discovery NM/CT670; GE Discovery NM 630 |  |  | October 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue associated with thetable pallet of your Discovery NM/CT 670 and/or Discovery NM630 system which may impactoperation safety.At scan setup, during gantry rotation towards scanning position, a potential contact between the table-pallet-edge and the detector was identified. Such occurrence may damage the pallet front-edge and | Software design | GE Healthcare, LLC sent an " Urgent Medical Device Correction" letter dated October 12, 2012 to all affected customers. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Mangers, and Mangers of Radiology / Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions or concerns contact your local GE Healthcare Service Representative. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other countries: 262-896-2890. |  | 210 | 210 | Worldwide distribution: USA (Nationwide) in the states of: AL, CA, FL, GA, IL, IN, KS, LA, MD, MA, MN, MS, MO, NJ, NY, NC, OH, OK, PA, SC, TN, TX, VA, WI and WV. AUSTRALIA, BELGIUM, CANADA, CHINA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GUADELOUPE, HONG KONG, INDIA, IRELAND, ISRAEL,ITALY, JAPAN, REPUBLIC OF KOREA, NETHERLANDS,NORWAY, POLAND, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, TAIWAN,THAILAND, TURKEY, UNITED KINGDOM. | Terminated | November 14, 2013 | 398 | 210 | Not_Computer
Z-0118-2013 | 63405 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra External Beam 4.1. |  |  | August 21, 2012 | 2012 | Elekta, Inc. | During the planning process the beam weights can be changed by the customers. When the beam weights are changed and renumbered and this happens during the dose calculation it is possible to save and reopen the plan without any error messages. | Software design | Nucletron sent a Customer Information Bulletin 555.00248 to all affected customers. The bulletin identified the product, problem, and actions to be taken by the customers. The bulletin contained information on how an incorrect dose association can be prevented by not renumbering beams during recalculation. A Read Confirmation Form was included for customers to complete and return. Contact your local Nucleton Service Representative for questions regarding this notice. |  | 52 | 52 units | Nationwide Distribution-including the states of AL, CA, DC, KS, LA, MO, NB, NY, NC, ND, OH, OR, PA, RI, SC, TX, and WA. | Terminated | November 03, 2014 | 804 | 52 | Not_Computer
Z-0120-2013 | 63296 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | The brand name of the device is Damon 3MX Bicuspid Bracket, an orthodontic bracket. Part Number 494-4491, Description -Damon 3MX Bicuspid Bracket, Lot Number - 011265039 |  |  | April 24, 2012 | 2012 | Ormco/Sybronendo | The firm initiated the recall as a result of a misidentification of the brackets. Damon 3MX Bicuspid Brackets are produced with a purple colored ID dot to identify them as maxillary bicuspid brackets; however the affected lot of product listed below was produced with dark blue colored ID dots,which identified them as maxillary central incisor brackets. | Labeling mix-ups | Ormco/Sybronendo sent an Urgent Medical Device Recall letter via USPS 1~1 class mail, on April 24, 2012 to all consignees. Consignees were sent an updated customer notification via USPS lSI class mail on June 28, 2012. The letters identified the affected product, problem and actions to be taken. Consignees were instructed to complete the Acknowledgement/Return Form and to return any affected product in their inventory. Customers were instructed to contact Ormco Customer Care at 1-800-854-1741 to receive an RMA number. The RMA will allow for a quick return and replacement or credit. |  | 148 | 148 | Worldwide Distribution - US Nationwide and the countries of Canada, China, Costa Rica, Ecuador, French Polynesia, India, Kuwait, Malaysia, Philippines, Singapore, South Korea, Thailand and UAE. | Terminated | December 12, 2012 | 232 | 148 | Not_Computer
Z-0125-2013 | 61612 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.2 workstations with software version 22_01 and 22_02 |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of inaccurate Scale factor annotation provided by SenoAdvantage 2.2 workstations that may impact patient safety. This issue may result in a potential miscalculation of size. | Software design | Consignees were sent on 5/24/12 a GE Healthcare "Urgent Medical Device Correction" letter dated May 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction. Customers should verify if their workstation is affected. GE Healthcare will correct all affected workstations by providing a software upgrade and Operator Manual update. A GE Healthcare service representative will contact you to arrange this correction. This activity will be performed at no cost to you. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. |  | 12 | 12 | Worldwide Distribution -- USA (nationwide) including states of: CA, CT, NY, TX, and MS and countries of TAIWAN, ITALY, CANADA, and BAHRAIN | Terminated | December 11, 2012 | 201 | 12 | Computer
Z-0123-2013 | 63143 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCelÂż DxC 600/800 with |  |  | September 03, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter has confirmed that the Shuttle to Barcode Reader and Cap Piercer alignmentprocedure cannot be performed on certain DxC configurations with v5.0.11 software installed. If the Shuttle to Barcode Reader becomes misaligned, the Operator will not be able to perform thisalignment procedure on the DxC systems listed above. Software v5.0 .11 does not allow access to theShuttle to Bar | Software change control | Beckman Coulter sent an Urgent Field Safety Notice dated August 30, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that their Beckman Coulter Service Representative would contac them to reinstall the software. Customers were asked to complete and return the enclosed Response Form. For questions customers were instructed to call their Customers Support Center on the Beckman Coulter website at www.BeckmanCoulter.com/CustomerSupport/SupportRequest.For questions regarding this recall call 714-993-5321. |  | 7 | 7 units total (both US and International) | Worldwide Distribution - USA including TX and OH. Interntionally to China and Turkey | Terminated | February 14, 2014 | 529 | 7 | Computer
Z-0122-2013 | 63081 | No | ONN | Stereotaxic instrument. | Neurology | Neurology | 510(k) | "****ORTHO SENSOR**12-123-A Knee Trial for Trithion***1560 Saw grass Corporate Pkwy***4Th floor***Sunrise, FL 33323***www.orthosensor.com***" |  |  | February 04, 2012 | 2012 | Orthosensor | Ortho sensor in Sunrise, FL is recalling the Graphics User Interface Software 1.8 (which was updated to version 1.82, validated on 02/13/12). Recall was initiated due to customer feedback which included the line of site between the knee ba lancer and the Link station was obstructed by OR metal stands or trays or the knee balancer device was placed atop of or near the metal stands or trays, interru | Software design | Orthosensor did not provide a written customer notification. As the recall was reported to the FDA well after all nine of the products were corrected (the recall stemmed from an initial inspection of the firm), then no customer notification was requested. The removal and replacement of the previous version of the Graphic User Interface Software is being handled internally by members of Operations and the Sales force. The items the firm is recalling is software, and the software cannot be returned; the old version (version 1.8) would be erased. The new version 1.82 would be installed and downloaded into the equipment at the site during servicing. This does not involve the customer of which none are consignees. Effectiveness checks associated with the update of the software is being handled internally by the Operations and the Sales force.Further questions please call (813) 645-2855. |  | 9 | 9 | Nationwide Distribution including the states of CA, FL, MI and NY. | Terminated | December 12, 2012 | 312 | 9 | Computer
Z-0131-2013 | 62315 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Perfusion System 8000, base 5 pump, 100V safety monitor included |  |  | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2007 of various malfunctions of the power distribution and battery back-up systems for the Sarns Modular Perfusion System 8000. Very few of the reports occurred during cardiopulmonary bypass; the majority of the malfunctions occurred during set-up or were reported during routine preventative maintenance. The malfunction | Labeling design | TERUMO sent an URGENT MEDICAL DEVICE letter dated September 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed review the Medical Device Recall notice, assure that all users are aware of the notice and place the addendum in the Operator's Manual before page 1.13. Customers were also instructed to confirm receipt of the Urgent Medical Device letter by faxing, the attached Customer Response Form to 1-800-292-6551. For questions customers were instructed to call 1-800-521-2818. |  | 81 | 81 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV, and WY and the countries of ARGENTINA, AUSTRALIA, BANGLADESH, BELGIUM, Bolivia, BRAZIL, CALLO PERU, CANADA, CAROLINA PR, CHILE, China, COLOMBIA, Costa Rica, DEKALB, DENMARK, Dominican Republic, Ecuador, Egypt, FRANCE, Germany, GUAYQUIL ECUADOR, Hong Kong, India, Indonesia, ITALY, Japan, , JARAMA MADRID, Lebanon, LEICESTERSHIRE, MADRID SPAIN, MADRID SPAIN, Malaysia, MEDLEY, Mexico, D.F., MIAMI, MUMBAI 26 INDIA, MUMBAI, New Zealand, Nicaragua, Pakistan, Panama, Philippines, REPUBLICA, SANTIAGO, Saudi Arabia, SELANGOR, MALAYSA, SEOUL, KOREA, Singapore, SOUTH AFRICA, South Korea, Sri Lanka, SWITZERLAND, TAIPEI, TAIWAN, Thailand, TUJUNGA, UNITED ARAB EMIRATES (UAE), URUGUAY, Venezuela, Vietnam, VIGNATE ITALY, WOODDALE, and Yemen. | Terminated | August 21, 2015 | 1060 | 2256 | Not_Computer
Z-0132-2013 | 63316 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath. Small Curve catalog no. D140010, Large Curve catalog no. D140011. The MobiCath Bi-Directional Guiding Sheath is a flexible tipped percutaneous catheter introducer (or guiding sheath) designed for gaining access to the coronary systems. The device features adjustable tip geometry through use of a rotating handle to deflect the catheter. This catheter is assembled with a hemostasis valve and a sideport infusion line. |  |  | September 24, 2012 | 2012 | Greatbatch Medical | We are notifying you of a potential interaction between therapeutic catheters and MobiCath Bi-Directional Guiding Sheaths. We became aware of these interactions while investigating recent complaints regarding catheters "self-advancing" while being used in conjunction with the MobiCath Sheaths. | Device Design | The firm, Greatbatch, Ltd., sent an "URGENT CUSTOMER NOTIFICATION Medical Device - Voluntary Field Safety Notice" dated Monday, September 24, 2012 to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to read the "Description of the Problem" and "Recommendations for Clinical Use" sections carefully; complete and return the attached Field Action Acknowledgement Form via mail to: Biosense Webster, Incl, 15715 Arrow Highway, Irwindale, CA 91706; Attn: Recall Coordinator, or Fax to: (909) 839-8514; pass on this notice to anyone in your facility that needs to be informed; maintain awareness of this communication until information has been incorporated into the device labeling, and to avoid the catheter inside the MobiCath Bi-Directional Guiding Sheath from "self-advancing", physicians should reposition the catheter handle when manipulating the catheter in and out of the sheath. Greatbatch, Ltd. is updating the MobiCath Bi-Directional Guiding Sheath Instructions For Use (IFU) accordingly. Supplemental information will be provided with future shipments of these devices, until the revised IFU is available.For questions related to the issue please contact your Biosense Webster representative or call (800) 729-9010, Monday through Friday from 9:00 AM to 6:00 PM EST. For questions related to the Acknowledgment Form and its return, please contact Field Action Coordinator at (909) 839-8468. |  | 9040 | 9040 | Nationwide distribution: CA only. | Terminated | May 15, 2014 | 598 | 9040 | Not_Computer
Z-0133-2013 | 63283 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DePuy Mitek VAPR Hook Electrode 3.5mm Hook Electrode with Integrated Handpiece |  |  | March 12, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product may have wrong large product identification label with symbols on the box indicating a S90 VAPR Electrode instead of the VAPR 3.5 hook electrode | Error in labeling | The firm, Depuy Mitek, sent a "Medical Device Labeling Alert" letter dated March 14, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect your inventory to determine if any inventory of product code 227305 lot with the number M1108086 that is incorrectly labeled; review the following label information and the attached examples of correct and incorrect labels; and complete and return the Business Reply Form via fax to: 1-508-828-3750. To return the product call 1-877-379-4871 for an RGA and replacement product and return to DePuy Mitek, Attn: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324.Contact your local DePuy Mitek Representative or DePuy Mitek Customer Service directly at 1-877-379-48 |  | 376 | 376 units | Nationwide distribution: USA including states of: AL, CA, CT, FL, GA, IA, ID, IL, IN, KY, MD, MI, MN, NJ, NY, OH, OR, TN, TX, UT, VA, and WA. | Terminated | June 05, 2015 | 1180 | 376 | Not_Computer
Z-0137-2013 | 63302 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | AquaLiner Hydrophilic Guidewire 0.035" X 260cm, Straight/ Stiff. |  |  | July 27, 2012 | 2012 | Nipro Medical Corporation | Nipro Medical Corporation, Miami, FL recalled their AquaLiner Hydrophilic Guidewire because the wires contain insufficient lubricity which causes them to get stuck in the housing and become nonfunctional. | Equipment maintenance | Nipro Medical Corp. sent notifications to their consignee via e-mail and fax on July 27, 2012 and September 13, 2012 to their affected customer. The letters identified the product, problem, and actions to be taken by the customer. The letter instructs the customer to place all affected product in stock on hold. A Aqualiner-Recall Response form was attached for the customer to complete and return. Contact the firm at 305-599-7174 ext. 249 for questions regarding this notice. |  | 300 | 300 | Worldwide Distribution-Distributed in the state of New York and the countries of Spain, Germany, Argentina, Panama, Mexico, Chile, Colombia, Ecuador, Peru, and Brazil. | Terminated | May 25, 2013 | 302 | 1500 | Not_Computer
Z-0155-2013 | 63121 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SpiraLok 6.5 w/out Needles with violet and blue Orthocord |  |  | September 06, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Fracture of SPIRALOK Anchors post op, requiring patient revision | Device Design | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated September 6, 2012, to all affected customers. Dr Doctor letters were included with the recall letter. Recall is coordinated with Stericycle , contact 1-888-202-3694 The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the enclosed Business Reply Form, even if they have no product remaining in their inventory. Customers were instructed to isolate all inventory of the affected product and return to Stericycle at the following address or fax to 1-888-912-2189 with the completed Business Reply Form:Stericycle2670 Executive Drive, Suite AIndianapolis, IN 46241Attn: Event #8874Customers with questions were instructed to call 1-888-202-3694.For questions regarding this recall call 508-880-8100. |  | 4321 | 4,321 units | Worldwide Distribution - USA (nationwide) and the countries of AE, AR, AT, AU, BE, BR, CA,, CH,CL .CN, CO, CZ,DE, DK, EC, ES, FI, FRGB, GP, GR, IE, IL, IN, IT, KR, LU, LY, MX, NC NL, NO, NZ, PE, PLPT, RE, RU, SG, SI, TH, TN, TR, UG ,UY , VE, and ZA. | Terminated | June 03, 2015 | 1000 | 78594 | Not_Computer
Z-0158-2013 | 63361 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Gish 3/8 x 1/2" Reducer Connector with Luer fitting, Item Numbers: EC2180S, Lot Numbers: 1108900099, 1119900091, 1203700129, 1208600013, 1219200107, Exp. Dates: 3/31/14, 7/31/14, 2/28/15, 3/31/15, 7/31/15. |  |  | September 14, 2012 | 2012 | Sorin Group USA, Inc. | Due to a molding irregularity, certain lots of connectors have the potential to leak fluid or entrain air at the interface of the luer fitting and mating cap. Entrained air could result in patient injury. | Process control | Sorin Group USA, Inc. sent a Product Recall letter to all affected customers by certified mail on September 14, 2012. The letter identified the affected products, problem and actions to be taken. Customers were instructed to discontinue use of the affected product, return all unused product to the firm and to complete and return a Response Form via fax or email. For further questions, contact the firm at 1-800-650-2623. |  | 1380 | 1,380 units | US Nationwide Distribution in the states of: AL, AR, CA, CO, CT, FL, IL, MI, MN, MO, MS, NC, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA and WI. | Terminated | April 23, 2013 | 221 | 3020 | Not_Computer
Z-0156-2013 | 63379 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Repeater Pump, Class II device used for facilitating repeatable drug dosage distribution and reconstitution in hospital pharmacies. The device are Fluid Transfer Tube Sets, Sterile/EO, Std Volume Trifurcated, BAXA Corporation, Englewood, CO. |  |  | October 03, 2012 | 2012 | Baxa Corporation | The firm is issuing a field correction to the operators manual included with the 099 Repeater Pump because the Technical Manual was distributed with the device as opposed to the Operator Manual and the instruction for use contradict each other with regards to multiple source container use. | Packaging process control | Baxa Corporation sent a "REPEATER PUMP PRODUCT NOTICE" to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. All customers received a CD containing a copy of the correct Operator Mannual, 5300-0754. Contact Technical Support at 800-678-2292 for questions regarding this notice. |  |  |  | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Brazil, Canada, Columbia, Guam, Japan, Malaysia, Mexico, Panama, Puerto Rico, Taiwan, UK, and Venezuela. | Terminated | February 15, 2013 | 135 |  | Not_Computer
Z-0139-2013 | 63321 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 6.5mm Cancellous Bone Screw |  |  | September 28, 2012 | 2012 | Encore Medical, Lp | A product complaint was received where a size 20 screw was labeled as size 30. After receipt, measurement of the screw determined it to be a size 25. Further investigation also determined that there were other screws labeled as size 25, but a size 30 screw was inside of the package. It was thought that all of these screws were found in-house and were contained, but it was determined later that on | Labeling mix-ups | Encore Medical sent a Urgent Field Safety Notice dated September 28, 2012, to all affected customers.This notice needs to be passed on all those who need to be aware withinyour organization or to any organization where the potentially affecteddevices have been transferred.Please contact Customer Service to place a replacement order and receive an RMA number. Return all affected devices using the RMA number.For further questions please call (512) 834-6255 |  | 1 | 1 unit | Nationwide Distribution only to CA. | Terminated | January 25, 2013 | 119 | 1 | Not_Computer
Z-0160-2013 | 62947 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX iD HMOD 2030 Diffusion Membrane Oxygenator; QUADROX-i ADULT HMO 70000 / HMO 71000; QUADROX-iD Pediatric HMOD 30000 |  |  | June 12, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | It has come to firm's attention that, in some rare cases, the blood outlet connector on some models of MAQUET Oxygenators may detach from the oxygenator. This event has occurred in about 0.01% (1:10,000) of units shipped. | Nonconforming Material/Component | Maquet sent a Field Safety Notices and Confirmation Forms dated June 13, 2012 to all affected customers. A Follow-up communication dated July 20, 2012 was mailed to all affected US customers. The Field Safety Notice identifies the problem, affected product, risk factors and actions to be taken by the customer. Customers were instructed to complete the Confirmation form and return in the self-addressed prepaid envelop. For questions call 201-995-8968.Maquet issued an updated Field Safety Notice/Field Safety Notice Confirmation Form dated 2/14/13 via Fed Ex on 2/18/2013.Maquet sent out an Updated Field Safety Notice/field Safety Notice Confirmation Form dated 4/4/2013 in order to notify customers they would be providing safety clamps. |  | 23203 | 23,203 units | US Nationwide Distribution | Open, Classified |  |  | 23203 | Not_Computer
Z-0159-2013 | 62386 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm), Reorder Number: PMILK40, Product Sterile in Unopened/Undamaged pouch, Cotton Tip - X-Ray Detectable, 1 Pc/Pouch, 10 Pouches/Box, 20 boxes/case. |  |  | October 28, 2011 | 2011 | Progressive Medical Inc | The tip of the product could come loose or unravel. | Packaging | The recalling firm sent a letter dated October 28, 2011 to the customer who reported the problem and calls were initiated 6/26/12 to the remainder of the customers. The reason for recall was explained and the customer were asked if any of the lot remained in their inventory. If inventory remained, the recalling firm sent them a return goods authorization number and informed them the product would be replaced. For questions call 1-800-969-6331. |  | 124 | 124 boxes | US Nationwide Distribution in the states of: CO, IL, IN, KS, LA, MA, MN, MO, NE, OH, OK, TN, TX, VA, and WI. | Terminated | October 26, 2012 | 364 | 124 | Not_Computer
Z-0124-2013 | 63139 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartSine Samaritan(R) 300/300P PAD devices with serial numbers 0400000501 to 0700032917; 08A00035000 to 10A0070753; 10C00200000 to 10C00210106 |  |  | September 13, 2012 | 2012 | Heartsine Technologies, Limited | Certain Samaritan 300/300P PAD devices may experience one or both of the following conditions that could affect the ability to deliver therapy to a patient in an SCA event if needed: Issue 1 (On/Off Issue): the device may turn itself on and off without input from the user. When this occures, the normal sequence of audible instruction prompts that occurs when the device is turned on will be emit | Device Design | HeartSine sent an Urgent: Medical Device Recall dated September 2012 to all affected customers. The letter identifed the affected product, problem and actions to be taken. A response form was included with the letter to be completed by customers and returned to the recalling firm's representative. In addition, HeartSine is providing a software upgrade to bring all users up to a more recent version of the software. For questions call 1-877-877-0147 or your service representative.The company issued anadditional press release on September 13, 2013 in a further attempt to reach out to end customers thathad not yet responded to the companys efforts. |  | 81982 | 81,982 in total | Worldwide Distribution and US (nationwide) in the following states: AK, AL, AP, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | July 10, 2015 | 1030 | 81982 | Computer
Z-0162-2013 | 63167 | No | MOI | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens ADVIA Centaur BR Assay: Catalog Number: |  |  | September 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | BR Assay for CA 27.29 Calibration Interval Change due quality control (QC) results outside of the expected ranges. | Under Investigation by firm | Siemens sent two "URGENT FIELD SAFETY NOTICE" letters dated August and September 2012 to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. A Field Correction Effectiveness Check Form was included for customers to complete and return. |  | 15455 | 15,455 kits | Worldwide Distribution-USA (nationwide) and the countries of Australia, Brazil, Canada, Columbia, Japan, Korea, Mexico, New Zealand, Peru, Saudi Arabia, and Spain. | Open, Classified |  |  | 15455 | Not_Computer
Z-0165-2013 | 63112 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Myocardial Heart Wires |  |  | July 25, 2012 | 2012 | Oscor, Inc. | Oscor, Inc. in Palm Harbor, FL is recalling myocardial heart wires, models TME60S, TME64S, TME64C, TME64S-3, TME65S, and TME66S. The recall was initiated due to the spacing between anchoring zig-zag to electrode is out of specification. | Process control | Oscor sent a customer notification letter dated July 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to send the listed models and lots in their inventory back to the recalling firm for immediate replacement. They were also instructed to call Oscor Customer Service at 727-937-2511 to obtain a Return Goods Product (RGA) number.For questions regarding this recall call 727-937-2511, ext. 133. |  | 1425 | 1425 | Worldwide Distribution - USA including OR, FL, Al, TX, NM, CA, AL, PA, KY, IN, OH, and IL and internationally to Canada, Russia, and New Zealand | Terminated | December 17, 2013 | 510 | 1425 | Not_Computer
Z-0164-2013 | 63090 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***All in one CrossFT***Punch***PFT-00M***4.5mm to 6.5mm***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***" |  |  | April 09, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling the product due to the punch being mis-assembled with the incorrect handled. The correct handle is a clear anodized silver colored handle. | Process control | Linvatec Corp sent a Urgent Medical Device Recall Notification letter dated April 9, 2012, to all affected customers. The letter was sent via FEDEx and by e-mail. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their inventory for the recalled product and segregate them for return to the recalling firm. Furthermore, return the products along with the "REPLY FORM".Should you have any further questions please contact our Customer Service department at (800) 237-0169. |  | 2 | 2 units | Worldwide Distribution--US (nationwide) including state of Minnesota and the country of Thailand | Terminated | November 21, 2013 | 591 | 2 | Not_Computer
Z-0168-2013 | 63385 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ILLlCOÂż MIS Posterior Fixation System, Bone-screw Internal Spinal Fixation System, |  |  | September 25, 2012 | 2012 | Alphatec Spine, Inc. | The firm initiated the recall because the stainless steel guide-wire of theIllico MIS Fixation System did not completely fit/pass through the tip of certain Illico TiCannulated Polyaxial Screws, therefore rendering such screws nonfunctional. | Nonconforming Material/Component | Alphatec Spine sent an Urgent Medical Device Recall Notification letter dated September 25, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Alphatec Spine's Customers Service Department (customerservice@alphatecspine.com) immediately for instructions on how to return the affected product. Customers were asked to fill out the lower portion of the letter and fax to (760) 431-0289 to confirm that they have read the notification and will take all necessary removal actions. For questions customers were instructed to call 1-800-922-1356. For questions regarding this recall call 760-431-9286. |  | 209 | 209 screws | Nationwide Distribution including AZ, IL, CA, SC, TN, TX, GA, CT, NY, FL, OK, LA, MO, NV,and UT. | Terminated | March 29, 2013 | 185 | 209 | Not_Computer
Z-0166-2013 | 63381 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | SCP Pump Control Panel, Product Code: 60-02-15, Serial Numbers: 60S09121, 60S09123-60S09129, 60S09136-60S09139. |  |  | September 20, 2012 | 2012 | Sorin Group USA, Inc. | Sorin Group USA, Inc. has issued an Urgent Field Safety Notice regarding certain serial numbers of the Stockert Centrifugal Pump (SCP) Pump Control Panel . An out of specification circuit board component within the Panel may lead to a malfunction consisting of a blank or flashing display. The flow and pressure display information may also be incorrect after the malfunction. | Mixed-up of materials/components | The firm, Sorin Group USA, Inc., sent an "URGENT FIELD SAFETY NOTICE" dated September 20, 2012 to all affected customers by certified mail. The notice described the product, problem and actions to be taken. Customers were instructed to discontinue use of and segregate affected product for return and replacement to the firm; to complete and return the attached Customer Response Form via fax to: 303-467-6502 or email to: yvonne.feyerherm@sorin.com; and if they have transferred the affected products to a third party, to pass this information to the Sorin Group.Further questions may be directed to Sorin Group Customer Service at 1-800-650-2623. |  | 12 | 12 units (3 units USA and 9 units OUS) | Worldwide distribution: USA (nationwide) including states of : MI and NY; and countries of: Belgium, Canada, Japan, India, Russia, United Arab Emirates. | Terminated | January 28, 2013 | 130 | 20 | Not_Computer
Z-0163-2013 | 63272 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aviator Drill Bit (10 mm, 12 mm, 14 mm, 16 mm, |  |  | July 05, 2012 | 2012 | Stryker Spine | Aviator drill bits are 10 mm longer than the Reflex Hybrid drill bits and are not interchangeable which could lead to over-drilling by as much as 10 mm. | Device Design | Stryker Spine sent an Urgent Medical Device Field Correction Notification letter beginning July 5, 2012, to all affected customers. Teh letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the drill bit length in each respective drill guide prior to surgery to ensure the correct length protrudes through the guide. Customers with questions should call 201-760-8298. Customers were asked to report any adverse events or product quality problems to 1-888-457-7463. |  | 6154 | 6154 units | Nationwide distribution. | Terminated | January 28, 2013 | 207 | 6154 | Not_Computer
Z-0169-2013 | 63335 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | September 18, 2012 | 2012 | Elekta, Inc. | A patient was seriously mistreated after a user made and saved an inadvertent change to the definition of a treatment filed in MOSAIQ. | Device Design | The firm, MOSAIQ, sent an "IMPORTANT SAFETY NOTICE- LINMSQ0003" was sent to consignees/customers on September 18, 2012. The notice described the product, problem and actions to be taken. The customers were instructed to check their Department Setup configuration settings, and re-enable the Field Edit Delta Display feature if they have disabled it. Note: sites can take this action without an upgrade/patch. The customers were also instructed to complete and return the CONFIRMATION OF RECEIPT form via fax to: 702-992-5002 in US or +44 (0)1293 654401 in OUS, Attn: Elekta OIS Support; scan and email to: support@impac.com or europe.support@impac.com; Attn: Elekta OIS Support and/or mail to: C/O Elekta OIS Support, 2310 Corporate Circle, Suite 275, Henderson, NV, USA; or C/O Elekta OIS Support, Elekta Limited, Linac House, Fleming Way, Crawley, West Sussex, RH10 9RR, UK.(Note: Product upgrade/patch is not mandatory or even necessary to reduce the related risk.)For questions, concerns, and requests for upgrade, email: support@impac.com (North America and Rest of World) or europe.support@impac.com (Europe) or call Vice President of RA/QA RNA at 770-670-2548. |  | 300 | 300 | Worldwide distribution: USA (nationwide) and country of Canada. | Terminated | June 06, 2013 | 261 | 300 | Not_Computer
Z-0161-2013 | 60030 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Medpro Elastomeric Infusion Pump. |  |  | July 09, 2011 | 2011 | First Medical Source LLC | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being reclassified by the Agency as a Class I. The recall was initiated because First Medical Source has confirmed that these lots may have a higher flow rate than specified. The use of this product may lead to over-administration of drug solutions to the patients. The pr | Nonconforming Material/Component | Please be aware that this is not a new recall. The firm has taken action; but, due to administrative issues this recall is now being reclassified by the Agency as a Class I. First Medical Source received an email from Medpro International with the following lnstruction to Customer - Informed customers to remove product from sales/discontinue use, cease distribution, remove the products from inventory and quarantine, inform their customers to retum the unused products to them, and destroy the product. Distributors' customer service will contact the customer to coordinate the return of the goods in question where required.First Medical Source emailed to their customer, Williams Medical Technologies, to informed them of the recall of the Medpro Elastomeric infusion pumps. Customer was instructed to inform them of the date and method of disposal of recalled products. |  | 500 | 500 units | US Nationwide distribution in the state of IL | Terminated | September 22, 2014 | 1171 | 500 | Not_Computer
Z-0182-2013 | 63387 | No | NOP | West Nile virus serological reagents. | Microbiology | Microbiology | 510(k) | West Nile Virus IgG DxSelect" kit Catalog No. EL0300G |  |  | October 03, 2012 | 2012 | Focus Diagnostics Inc | The firm recalled due to higher reactivity with some samples in the effected lots, which may reflect ahigher false positive rate. | Nonconforming Material/Component | Focus Diagnostics Inc., sent a Urgent Medical Device Recall letter dated October 3, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you received at least one of the effected lots in the month of August 2012. Please document the number of kits in your inventory on the table below, even if zero (0), and fax to 562-240-6526.For further questions please call (714) 220-1900. |  | 425 | 425 units | Worldwide Distribution--US (nationwide) and the country of: Canada, Deuche, France, Greece, Mexico, Israel, and Japan. | Terminated | December 04, 2012 | 62 | 425 | Not_Computer
Z-0184-2013 | 63338 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Knee Components: |  |  | September 05, 2012 | 2012 | Wright Medical Technology Inc | There may be inaccuracies on the labeled shelf life on all products packaged at the Toulon, France facility. | Packaging process control | Wright sent an Urgent Product Recall Notification letter dated September 19, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for affected product and immediately return to Wright Medical Technology, Inc. Customers were asked to complete and return the enclosed notification form by fax to 901-867-7401. For questions customers should call 901-867-4324.For questions regarding this recall call 901-867-4655. |  | 729 | 729 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Austria, Belgium, Bulgaria, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Latvia, Lithuania, Norway, Poland, Portugal, Serbia, Slovakia, Slovenia, Spain, The Netherlands, Argentina, Brazil, Chile, China, Japan, Taiwan, Turkey, United Arab Emirates, Australia, South Africa, Russia, and United Kingdom. | Terminated | January 07, 2014 | 489 | 33135 | Not_Computer
Z-0175-2013 | 63347 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | AirFlow Resuscitator, Adult Resuscitator, AF1000 Series, Single use, Manufactured by: Ventlab Corporation, 155 Boyce Drive, Mockville, NC 27028 |  |  | July 11, 2012 | 2012 | Ventlab Corporation | Possible volume leakage through the inlet valve during compressions of the resuscitator bag. | Component design/selection | VentLab Corporation sent an Attention: Recall Notification letter dated July 10, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product, complete and return the enclosed response form by fax to 336-753-5002, indicating whether or not their facility has any of the affected product. Customers can also scan and email the response form to CSR@Ventlab.com. For questions regarding this recall call 336-753-5000. |  | 14602 | 14,602 total for all units | Nationwide Distribution including IL, OH, NY and TN. | Terminated | June 30, 2013 | 354 | 14602 | Not_Computer
Z-0186-2013 | 63104 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Software for Vitrea, Vitrea fX, and Vitrea Enterprise Suite (VES): |  |  | August 13, 2012 | 2012 | Vital Images, Inc. | Vital Images has found two potential errors in calcium score values in restored snapshots of studies generated through use of Vitrea or Vitrea Enterprise Suite. This issue affects Version 3.1 of Vitrea fX; Version 5.2 of Vitrea; Versions 6.0, 6.1 and 6.2 of Vitrea and Vitrea fX; and Versions 1.3, 6.0, 6.1 and 6.2 of Vitrea Enterprise Suite; and all updates to those versions. | Software design | Vital Images sent an "URGENT DEVICE CORRECTION" letter dated August 13, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter gave advise on how to resolve the two issues and advised consignees to immediately inform all potential users. Contact Customer Support at 1-800-208-3005 for questions regarding the notice. |  | 2088 | 2088 (1133 USA) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, BULGARIA, CANADA CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, INDIA, INDONESIA, IRAQ, IRELAND, ITALY, IVORY COAST, JAPAN, JORDAN KAZAKHSTAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MALI, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TURKEY, UKRAINE, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, AND YEMEN | Terminated | March 25, 2015 | 954 | 2088 | Computer
Z-0187-2013 | 63471 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA Radiation Oncology, Version 11. Varian Medical Systems, Palo Alto, CA. The ARIA Radiation Oncology product is a treatment plan and image management application. |  |  | September 24, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | An anomaly has been identified with the ARIA Oncology Information System Version 11 where Monitor Units of one fraction may be incorrect, when partial treatment is being continued. | Software design | "Urgent Medical Device Correction/ Urgent Field Safety Notice" letters were sent on September 24, 2012 by Certified Mail with return receipt to all affected users. The problem was described and recommended user actions were provided. The notification letters were to also be distributed to Varian Sales, Marketing, and Service organizations. US customers may contact the Varian help desk at 1-(888) 827-4265. |  | 205 | 205 | Worldwide Distribution - Nationwide(USA) and including the countries of Australia, Denmark and the United Kingdom. | Terminated | December 05, 2013 | 437 | 205 | Computer
Z-0188-2013 | 63310 | No | OWD |  |  | Molecular Genetics | PMA | cobas 4800 BRAF V600 Mutation Test for In Vitro Diagnostic Use Roche Molecular Systems, Inc., |  |  | July 31, 2012 | 2012 | Roche Molecular Systems, Inc. | It was identified that BRAF controls dispositioned as "Internal Use Only" were incorrectly packaged into cobas 4800 BRAF V600 Mutation Test CE-IVD kit lot R03607 and released for distribution. Two control batches manufactured during two process validation studies are involved:PV7711: BRAF Mutant (MUT) ControlPV0853: BRAF Wild Type (WT) Control | Mixed-up of materials/components | Roche sent a Safety Board Notice - Product Bulletin 2012-013 on 31-Jul-2012 to affected customers. The notice identifed the affected product, description of situation and actios to be taken. Customers were instructed to discontinue use, quarantine and discard any remaining kits of the affected product in inventory. Customers were advised to contact their affiliate sales office for replacement. The notice states a new cobas 4800 BRAF V600 Mutation Test lot (R07749) is anticipated to be available on 03-Aug-2012. For questions call +01.908.253.7569. |  | 44 | 44 kits | Worldwide Distribution - Australia, Austria, Belgium, China, Hungary, Turkey and United Kingdom. | Terminated | September 11, 2013 | 407 | 44 | Not_Computer
Z-0190-2013 | 63096 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Boule Con-Diff Hematology Control |  |  | March 24, 2011 | 2011 | Clinical Diagnostic Solutions | Clinical Diagnostic Solutions, Inc. in Plantation, FL recalled Boule Con Diff US Tri Pack (Product Part 501-605, Lot #1102-575 and 1102-576; Product Part 501-607, Lot #1102-576; and Product Part 502-012, Lot #1102-574) which is used for in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. The recall was initiated due to the receipt of a complai | Nonconforming Material/Component | Clinical Diagnostic Solutions, Inc. sent a letter dated March 24, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm istructed the accounts to replace the assay sheet they were currently using with the new assay sheets enclosed. For an unused product they would need to discard. To correct the ranges in their instrument, they would need to scan in the new bar codes specific for their instrument type. For any questions customers were instructed to contact CDE Technical Support atl 800-453-3328.For questions regarding this recall call 954-791-1773. |  | 901 | 901 vials | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MS, ND, ND, NE, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WI, WV,and WY. Foreign: Brazil and Colombia | Terminated | August 13, 2014 | 1238 | 1721 | Not_Computer
Z-0192-2013 | 63309 | No | JJT | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cliniqa Liquid QCTM Cardiac Marker Control for i-STAT Level 1 (List Number 06F12-09; Lot Number 071039) |  |  | July 19, 2012 | 2012 | Abbott Point Of Care Inc. | Abbott Point of Care (APOC) has determined that for the Cliniqa Liquid QC Cardiac Marker Control for i-STAT Level 1 (List #06F12-09; Lot Number 071039); i-STAT Level 2 (List #06F12-10; Lot Number 081039); and i-STAT Level 3 (List #06F12-11; Lot Number 091039) results may be generated below the lower value assignment range for some vials. | Material/Component Contamination | The firm, Abbott Point of Care, Inc. (APOC), sent an "URGENT RECALL NOTICE" dated July 2012 via FED-EX on July 19, 2012. The recall notice identified the product, problem and the action needed to be taken by the customer. The customers were instructed to discontinue use of the control lot numbers; return all unused controls on the enclosed Customer Reply Form; If another institution was provided the controls, provide a copy of this notice to them. Note: if any used vials gave acceptable product performance, a review of the cartridge performance and associated patient results is not necessary)Note: APOC is shipping replacement products to customers facilities to minimize any disruption this might cause.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care representative. |  | 154 | 154 (06F12-09 (74) boxes; 06F12-10 (35) boxes and 06F12-11 (45) boxes.) 1 box contains six, 3ml plastic bottles. | Worldwide distribution: USA (nationwide) including states of: AZ, GA, ID, IL, LA, MD, MO, MN, NC, NV, OK, PA, TN, TX, VA and WI; and country of: Italy. | Terminated | June 20, 2013 | 336 | 154 | Not_Computer
Z-0206-2013 | 63430 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 VCC P/N 4598 001 41341 REV A, Philips Medical Systems (Cleveland), Inc., Fitchburg, WI. |  |  | October 05, 2012 | 2012 | Philips Medical Systems (Cleveland), Inc. | Philips Medical Systems Cleveland), Inc. is recalling Pinnacle3 Virtual Client Connection software version 3.1 due to an improper release and subsequent distribution. | Process control | The firm, Philips Healthcare, sent a "Customer Information Letter" dated 2012 Oct 05. The letter listed the affected products, problem description, how to identify affected products, advise on actions by customer / user, actions planned by Philips and further information and support. The letter stated that no action is needed on the part of the customer. The field service engineer will be scheduling a site visit to complete the replacement. Philips Healthcare will be delivering the correct VCC kit.For customers in North America if you need any further information or support concerning this issue, please contact our Customer Care Solutions Center at 1-800-722-9377. Select option 5 for "All Imaging Systems". Enter your site ID # (If you do not have a site ID#, simply pause for a moment.). Select option 5 for "Nuclear Medicine" and finally select option 3 for "Pinnacle" support. |  | 11 | 11 | Nationwide distribution: USA including states of: ME, MI, NY, and WI. | Terminated | February 21, 2013 | 139 | 11 | Computer
Z-0203-2013 | 62945 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | "***Sharpoint IQ Geometry Slit Knife Angled, Bevel Up***Surgical Specialties Corporation Reading, PA***REF 72-3040***" |  |  | August 15, 2012 | 2012 | Surgical Specialties Puerto Rico Inc. dba Angiotech PR Inc. | Potential for bent tips. | Process control | Angiotech sent an Urgent Product Recall Notification letter dated August 15, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the current inventory, segregate the affected lot numbers, and to return the identified product. Customers were asked to complete the attached Customer Acknowledgement Form and fax to the Quality Assurance department at 787-658-1073. If product was further distributed customers should notify their customers to whom they sold the affected product. Customers with questions should call 787-658-1800 option 2 ext. 1775 or 787-658-1800 option 2 ext 1807.For questions regarding this recall call 787-658-1800. |  | 117948 | 117,948 units (total) | Worldwide Distribution - USA including DE,IL, NC, CA, CO, GA, KY, PA, MD, MN, NY, and MA and internationally to Austria, UAE, China, Argentina, Italy, Taiwan, Czech Republic, Australia, UK, India, Germany, Spain, France, South Africa, Malaysia, Japan, Lebanon, Korea, Greece, Saudi Arabia, Switzerland, Pakistan, Tunisia, Serbia, Morocco, Netherlands, Portugal, Truk, Belgium, and New Zealand | Terminated | January 13, 2014 | 516 | 117948 | Not_Computer
Z-0205-2013 | 63412 | No | DZN | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Vector TAS Modular Driver, a manual driver for Vector TAS dental screws. |  |  | December 12, 2011 | 2011 | Ormco/Sybronendo | The Instructions For Use (IFUs) for Vector TAS Modular Driver for some foreign language sections incorrectly lists the sterilization temperature as 130ÂżC. The correct sterilization temperature should be listed as 132ÂżC and/or 270ÂżF. | Error in labeling | "Urgent - Medical Device Correction" letters dated 12/12/2011 were sent via USPS 1st class mail to all customers and again on 2/21/2012 for recall expansion. The device issue was described and recommended actions were provided. Customers were instructed to acknowledge and complete the Return Form. Questions were referred to (800) 854-1741. |  | 4543 | 4,543 | Worldwide distribution, including Nationwide (USA) and the countries of Canada, Austria, Belgium, Bulgaria, Switzerland, Germany, Denmark, France, Estonia, Spain, Finland, Iceland, United Kingdom, Greece, Croatia, Hungary, Ireland, Italy, Luxembourg, Netherlands, Monaco, Poland, Romania, Serbia, Slovenia, Turkey, Jersey, and Martinique. | Terminated | November 05, 2012 | 329 | 4543 | Not_Computer
Z-0195-2013 | 63368 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX CO-OX analyzer, model #393-841 (all software versions prior to 1.35). |  |  | January 06, 2010 | 2010 | Sendx Medical Inc | The recall was initiated because SenDx Medical Inc. became aware of a software issue in the ABL80 FLEX CO-OX analyzer related to the reporting of the hemoglobin fractional parameters. | Software design | SenDx Medical sent a recall letter to all their foreign customers on January 6, 2010, to informed them of the ABL80 FLEX CO-OX recall via the Radiometer Field Action Notification (FAN) system. The Field Action Notes informed the customers of the problem identified and the action to be taken. Customers were instructed to submit confirmation fax#1 to confirm that the customer letter has been distributed to all ABL80 CO-OX customers by January 31, 2010. Customers were instructed to submit confirmation fax #2 to confirm that all ABL80 CO-OX analyzers have been upgraded to version 1.35 by 6/30/10.In the U.S., the customer was provided with a software upgrade CD along with a letter advising them to the issues associated with the recall. Customers with questions were instructed to contact TechLine at (800) 736-0600 opt 2. |  | 767 | 767 units | Worldwide Distribution-- US Distribution including the state of Ohio and the countries of Switzerland, Germany, France, Poland, Austria, Italy, Japan, Australia, Denmark, China, Sweden, Norway, Spain, United Kingdom, Netherlands, and Czech Republic. | Terminated | November 05, 2012 | 1034 | 767 | Computer
Z-0207-2013 | 63396 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Trochar Scalpel Instrument (Surgical Punch) |  |  | September 13, 2012 | 2012 | OrthoPediatrics Corp | A complaint received from sales representative that a replenishment instrument was too short. Upon investigation, it has been determined that the device was not manufactured to design specifications and will not function as intended. | Reprocessing Controls | OrthoPediatrics Corp sent an Urgent Recall e-mail dated September 14, 2012, to all affected customers. The e-mail identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine all identified devices from the affected lot numbers immediately. Customers were also asked to complete the Mandatory Reply Form. Customers with questions should call 574-268-6379. |  | 78 | 78 | Nationwide Distribution including AL, IL, NY, FL, TX, MI, KY, NC, CO, and GA. | Terminated | September 18, 2013 | 370 | 78 | Not_Computer
Z-0193-2013 | 63094 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Boule Con Diff US Tri Pack, 6 Vials. |  |  | August 04, 2010 | 2010 | Clinical Diagnostic Solutions | Clinical Diagnostic Solutions, Inc. in Plantation, FL recalled the Boule Con Diff US Tri Pack (Product part 501-605, Lot #1005-549 and Product part 501-607, Lot #1005-549) which is used for in vitro diagnostic use as a control to monitor the performance of multi-parameter hematology instruments. The recall was imitated due to receipt of customers stating the MCV values and RDW% recovered for Boule | Process control | Clinical Diagnostic Solutions, Inc. sent a "PRODUCT NOTIFICATION:BOULE CON DIFF TRI-LEVEL LOT 1005-549" letter dated August 4, 2010 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions regarding this notice, contact CDS Customer Service at 1-800-453-3328. |  | 532 | 532 packs | Nationwide Distribution. | Terminated | August 14, 2014 | 1471 | 1222 | Not_Computer
Z-0196-2013 | 63506 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The touchscreen display (model 94267) designed for use with the XPREZZON bedside monitor (model 91393) is width-452 mm (17.2 inches) and height- 368 mm (14.5 inches). The depth without the pedestal --- 67.7 mm (2.67 inches). The display is an integrated LCD 19-Inch. |  |  | October 05, 2012 | 2012 | Spacelabs Healthcare, Llc | Potential failure of Spacelabs Healthcare Touchscreen Display, Model 94267 where the alarm lights failed to illuminate upon installation. | Vendor change control | Spacelabs Healthcare sent a URGENT - MEDICAL DEVICE CORRECTION letter dated October 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are advised to advise their staff of this potential failure and advise them to ensure the alarm lights are functional. The firm will contact the consignee at the earliest possible date to schedule a convenient time for the hospital to upgrade the affected product at no cost.Customers can call Technical Support at 1-800-522-7025 and select 2 for any questions about this recall. |  | 169 | 169 units distributed (98 in US and 71 Internationally) | Worldwide Distribution--US (nationwide) including states of Alaska, California, Florida, Illinois, Michigan, South Carolina, and Texas., and the countries of Canada, Chile, France, Libya, Panama, Qatar, Saudi Arabia, Switzerland and United Kingdom | Terminated | March 25, 2013 | 171 | 169 | Not_Computer
Z-0253-2013 | 62833 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista(R) Homocysteine Flex(R) Reagent Cartridge |  |  | August 03, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Cartridges are leaking. The leaking is observed when opening the wrapper that holds the reagent cartridge. | Packaging | Siemens sent a Urgent Field Safety Notice letter dated July 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.At this time Siemens Healthcare Diagnostics is conducting a voluntary recall for Dimension VistaÂż HCYS FlexÂż reagent cartridge lots 12062MA and 12066MA. Please immediately discontinue use and discard any remaining inventory of Dimension VistaÂż HCYS (K7044) lots 12062MA and 12066MA. Please complete the attached form and fax it to (302)631-8467 to indicate that you have received this information. We ask that you forward this notification to anyone to whom you may have distributed this lot.Contact your Siemens Customer Service Center to report your no-charge replacement needs at 888-588-3916.If you have technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance.We apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support. |  | 583 | 583 cartons | USA (nationwide) including the states of AL, CA, CO, CT, DC, DE, FL, IA, IL, IN, LA, MA, MD, MI, MO, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA and WI. | Terminated | January 15, 2014 | 530 | 583 | Not_Computer
Z-0250-2013 | 63369 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory, version 7.0 |  |  | November 03, 2011 | 2011 | Sunquest Information Systems, Inc. | Sunquest is recalling Sunquest Laboratory version 7.0 because when saving an Accession to a non-current Episode/Event, a User-Defined field can be incorrectly replaced by the current Episode/Event User-Defined data. | Software design | Sunquest Information Systems, Inc. sent a recall letter dated November 3, 2011, to all affected customers. Customers were informed of the problem identified and the actions to be taken. Customers were informed that a change request CR11-00174907 has been created to address this issue. This will be available as part of Sunquest Laboratory version 7.0.1. To request the software correction, customers were instructed to create a Case via the Client Support webpage at www.sunquestinfo.com/supportweb. Then choose a new case, and finally select work queue and select software request. Customers with questions were instructed to call (877) 239-6337 for US and Canadian sites and at +44 (0) 161 335 0562 for international sites.For questions regarding this recall call 520-570-2252. |  | 14 | 14 sites | Worldwide Distribution - USA (nationwide) and Internationally to Ireland | Terminated | November 06, 2012 | 369 | 14 | Computer
Z-0252-2013 | 63270 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | STA Satellite Automated Multi-Parameter Analyzer |  |  | June 15, 2012 | 2012 | Diagnostica Stago, Inc. | Following a single customer complaint that incorrect INR values were occasionally printed for PT tests, even though the primary units (seconds) were correct, the firm investigated and confirmed the potential for printing incorrect secondary units when using an external printer and the customer printout mode on the STA SatelliteÂż. | Software design | Diagnostica Stago sent an Urgent Field Safety Notice dated June 12, 2012, to all affected customers via email return receipt requested on June 15, 2012. Customers were instructed to follow the instructions in the notification and complete the Customer Verification form and return to the Hotline by fax to 973-644-9348. For additional information, customers were instructed to contact the Hotline at 800-725-0607.For questions regarding this recall call 973-631-1200, ext 2044. |  | 237 | 232 units US; 5 units Canada | Worldwide Distribution - USA (nationwide) and Canada | Terminated | July 19, 2013 | 399 | 237 | Not_Computer
Z-0251-2013 | 63380 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Application Interfacing Outbound Phlebotomy Point of Care Interface, v3.0.0_07222004 and later. |  |  | November 19, 2009 | 2009 | Sunquest Information Systems, Inc. | The recall was initiated because Sunquest Information System has confirmed reports of an Outbound Phlebotomy Point of Care (POC) interface is combining patient orders inappropriately. | Nonconforming Material/Component | Sunquest sent an "URGENT-PRODUCT SAFETY NOTICE: PSN-09-06" letter dated November 19, 2009 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were informed that Sunquest Application Interfacing would be contacting their facility for scheduling to correct the issue. |  | 9 | 9 units | Nationwide Distribution. | Terminated | November 06, 2012 | 1083 | 9 | Not_Computer
Z-0243-2013 | 63103 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | NCBÂż cancellous screw |  |  | September 04, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc., is initiating a correction to the patient labels of products manufactured before March 20I0 by the Zimmer GmbH production site in Winterthur, Switzerland, after receiving four complaints of a discrepancy between product patient labels and product labels with respect to the reference and lot numbers of the product. While the information printed on the product label is correct, the info | Error in labeling | Zimmer, Inc. sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated September 4, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Certificate of Acknowledgement Form was attached for customers to complete and return via fax to Zimmer, Inc. at 574 372-4265. |  | 2 | 2 | Nationwide Distribution. | Terminated | April 16, 2014 | 589 | 23110 | Not_Computer
Z-0255-2013 | 62580 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Sysmex XE-5000 Automated Hematology System; an automated differential cell counter; Sysmex Corporation, Kobe, Japan; Sysmex America, One Nelson C. White Parkway, Mundelein, IL 60060; catalog #063-7591-7 |  |  | June 01, 2012 | 2012 | Sysmex America, Inc. | There is a variation of reticulocyte counts between XE Series (XE-2100, XE-5000) and XT-Series (XT-2000i, XT-4000i) hematology analyzers that may result in a retic counts that are an average of 21% lower on the XE-Series because of a low bias. | Device Design | Sysmex America, Inc. sent a "PRODUCT NOTIFICATION" letter dated May 2012 to all affected customers on June 1, 2012. The letter identifies the product, problem, and actions to be taken. The letter instructs customers to review patient results to determine if their Reference Intervals need adjustment, and to review QC results via Insight and document any change that may be linked to the date of calibration if appropriate. Any questions concerning the information contained in the notification were directed to the Sysmex Technical Assistance Center at 1-888-879-7639. |  | 763 | 763 units | Nationwide distribution: USA including Puerto Rico. | Open, Classified |  |  | 1931 | Not_Computer
Z-0209-2013 | 52796 | Yes | CBF | Hyperbaric chamber. | Anesthesiology | Anesthesiology | 510(k) | Sechrist Monoplace Hyperbaric Chamber with Electronic Control Module Model 3600E/ER |  |  | August 28, 2008 | 2008 | Sechrist Industries Inc | Sechrist conducted a field upgrade on the hyperbaulic chambers equipped with Electronic Control Modules in order to correct a software bug that could potentially affect the performance of the chamber. | Software design | Sechrist Industries contacted consignees by phone to notify them of a software upgarde and to schedule appointments to perform the required field upgrades. An approximate summary of the phone notification is provided below."I am contacting you to notify you of a field upgrade we are performing on ourhyperbaric chambers equipped with Electronic Control Modules in order to correct a software bug that could potentially affect the performance of your chamber. The upgrade includes the installation of a new motherboard with revised software, a reset switch, and the exchange of other control panel components. The process will take approximately 2.5 hours and will be performed at no charge to you."For questions regarding this recall call 714-579-8400. |  | 65 | 65 units total (61 units in the US) | Worldwide Distribution - USA (nationwide) and the countries of Bulgaria, and Canada | Terminated | November 06, 2012 | 1531 | 138 | Computer
Z-0259-2013 | 63540 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS AdVance" Male Sling System |  |  | October 19, 2012 | 2012 | American Medical Systems, Inc. | American Medical SystemsÂż (AMS) is issuing a voluntary recall on a number of lots of the AdVanceÂż Male Sling System (AdVance).AMS has recently received a small number of complaints regarding the resorbable suture in the AdVance sling not lasting for the entire 15 month shelf life of the product. An internal investigation has been conducted and confirmed that some sutures degraded prior to the | Device Design | American Medical Systems sent an "URGENT MEDICAL DEVICE RECALL" dated October 18, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter instructs customers to return any existing stock of the recall product and to complete and return the Product Recall Acknowledgment Form. Contact AMS Customer Service at 1-800-328-3881 for additional information or questions.. |  | 4872 | 4872 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of Austria, Australia, Brazil, Canada, Germany, Estonia, Spain, Finland, France, Great Britain, Hong Kong, Croatia, India, Morocco, Mexico, Malaysia, Netherlands, New Zealand, Panama, Poland, Russian Federation, Solomon Islands, Singapore, Thailand, Turkey, Venezuela, South Africa, Chile, Colombia, Costa Rica, Slovenia, Argentina, China, and Taiwan. | Terminated | April 11, 2013 | 174 | 4872 | Not_Computer
Z-0258-2013 | 63598 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroCool Breathable Gown with Towels, X-Larg, X-Long, Sterile |  |  | July 05, 2012 | 2012 | Kimberly-Clark Corporation | Package seam of some units may not remain properly sealed to ensure the contents might not be sterile as labeled. | Packaging process control | Kimberly Clark sent an Urgent Voluntary Product Recall letter dated July 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to determine if they have any affected product in their inventory. Any affected product should be quarantined and distribution discontinued. Distributors were also asked to conduct a sub-recall for any affected product further distributed. Customers were asked to complete the attached Distributor Product Recall Response Form and fax to 1-920-225-3522. Upon completion of the response form customers were asked to destroy any affected product.For questions regarding this recall call 770-587-7835. |  | 933 | 933 ((30 units per case) | Nationwide Distribution | Terminated | June 13, 2013 | 343 | 90789 | Not_Computer
Z-0270-2013 | 63083 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | AC Powered adjustable hospital bed, Prime Bed TL500, Distributed by: Primus Medical LLC, 8401 Southern Blvd, Boardman, OH 44512. |  |  | February 02, 2011 | 2011 | Primus Medical LLC | The "Retaining washers &amp; Clevis Pins" provided with the AC -powered adjustable hospital beds may fail under certain circumstances. The pins at the top of the leg were disengaging from the channel. This may cause the bed frame to come off tracks or collapse. | Equipment maintenance | The firm, Primus Medical LLC, sent an "Important Notification Relating to Primus Medical Product TL500" letter dated March 29, 2011 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to replace the "Retaining washers &amp; Clevis Pins" with the "Clevis Pins &amp; Bow Tie Cotter Pins" in the TL500-001 Service Pack provided with the letter in order to avoid any potential failure and/or bed collapse and to return the" Retaining washers &amp; Clevis Pins" after the replacement procedure is performed. If you have any questions, or need assistance in performing the replacement, call 330-965-9681. |  | 342 | 342 Beds | Nationwide distribution: USA including states of: AZ, IL, KY, MD, MI, OH, OR, PA, TX and VA. | Terminated | May 14, 2014 | 1197 | 342 | Not_Computer
Z-0268-2013 | 63297 | No | ELW | Impression material. | Dental | Dental | 510(k) | The brand name of the device is AlgiNot, a dental impression material. |  |  | February 21, 2012 | 2012 | Kerr Corporation | The firm recalled the product because it may cure harder than expected within theshelf life, which could result in a material that may be more difficult to remove. Use of this product onpatients with loose teeth may result in possible tooth fracture or tooth extraction. | Nonconforming Material/Component | Kerr Corporation sent "URGENT: MEDICAL DEVICE RECALL" letters to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. US, Canadian, and Rest of World consignees were notified via USPS 1st class mail on June 19, 2012, and Australianconsignees on July 5, 2012. Consignees were instructed to complete and return the Acknowledgement Form. Notification of the recall expansion was sent, via USPS 1st class mail, on February 21, 2012 to US, Canadian, Australia and European, consignees; however, once further investigation determined that all unexpired lots of the product were affected, the Alginot recall was expanded a second recall letter was sent to clarify that all lots of AlgiNot FS Volume recall and AlgiNot recall expansion are affected by this expandedrecall. Letters were mailed to US, Canadian and Rest of World (ROW) customers on June 19, 2012; Australian and European consignees were notified of the expansion on July 5, 2012.All consignees were instructed to complete the Return Form to acknowledge the updated instructions for use and the added cautionary statement. Consignees not responding are to be contacted at least three times via telephone/fax/e-mail as follow up on the recall. |  | 132 | 45 kits and 87 refills | Worldwide Distribution-USA (nationwide) and the countries of Canada, Europe, and Australia. | Terminated | November 15, 2012 | 268 | 132 | Not_Computer
Z-0260-2013 | 62271 | No | IZW | Diagnostic x-ray beam-limiting device. | Radiology | Radiology | 510(K) Exempt | R-30H , Diagnostic X-Ray Beam-Limiting Device. |  |  | October 28, 2011 | 2011 | Shimadzu Medical Systems | Shimadzu is recalling Diagnostic X-ray Beam-Limiting Device, Mobile x-ray system and Diagnostic x-ray tube assembly because they do not comply with the Product Performance Standards, 21CFR 1020.30 (m) Beam quality --(1) Half-value layer (HVL), which was revised in 2006. Certain devices might not comply with the requirements for them in some x-ray generating condition after the revision in 2006. | Nonconforming Material/Component | Shimadzu sent an Urgent Voluntary Medical Device Recall - Field Safety Notice letters dated September 19, 2012 to all the affected customers. The letter identifed the description of non-compliance, how to identify the affected products, corrective action plan by Shimadzu and actions to be taken customers. Customers were instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at Shimadzu Medical Systems National Technical Support (800) 228-1429. |  | 665 | 665 units | US Nationwide Distribution | Open, Classified |  |  | 1402 | Not_Computer
Z-0274-2013 | 63383 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Mast Roller Pump 150, Item Number: 10-88-00 |  |  | September 19, 2012 | 2012 | Sorin Group USA, Inc. | Sorin Group USA, Inc. has issued an Urgent Field Safety Notice regarding certain serial numbers of the Sorin Group S5 Perfusion System due to reports of a "Fault in motor controller" alarm message displayed on the system panel. This could lead to unintended pump stoppage which may result in permanent impairment or life-threatening injury. An out of specification electrical component was determin | Mixed-up of materials/components | The firm, Sorin Group, sent an "URGENT FIELD SAFETY NOTICE" dated September 19, 2012 to all affected customers by certified mail. The notice described the product, problem and actions to be taken. The customers were instructed to follow the provided instructions for the continued safe use of the Perfusion systems until they have been serviced or replaced by Sorin Group. The Sorin Group Service Team will contact the customer to schedule servicing and replacement of the affected product. The customers were also instructed to complete and return the Customer Response Form via fax to 303-467-6502 or by email to yvonne.feyerherm@sorin.com; assure this notice is distributed to all personnel who need to be aware of this notice; and if they have transferred the affected products to a third party, pass this information with them as well as with Sorin Group Customer Service at 1-800-650-2623. For questions regarding this notice, contact Sorin Group USA Customer Service at 1-800-650-2623. |  | 2 | 2 units | Worldwide distribution: USA (nationwide) including states of: AZ, GA, MI, MS, NC, NY, PA, SC, and TX; and countries of: Austria, Australia, Belgium, Canada, People's Republic of China, Columbia, Denmark, Finland, Germany, India, Iran-Islamic Republic of, Ireland, Japan, Lebanon, Morocco, Pakistan, Poland, Russian Federation, Saudia Arabia, Spain, South Africa, Thailand, Turkey and United Kingdom. | Terminated | January 18, 2013 | 121 | 47 | Not_Computer
Z-0266-2013 | 63273 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Telescopic Strut Short |  |  | July 03, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Three complaints have been filed where the Hoffman LRF Telescopic Struts have broken during load bearing application by patients having a body weight greater than 250 LBS. | Device Design | Stryker Orthopaedics sent an "URGENT MEDICAL DEVICE FIELD REMOVAL" letter dated July 2, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Business Reply Form was included for customers to complete and return to via fax to 1-865-252-3635. Contact the firm at 1-866-626-7747 for assistance. |  | 165 | 165 units | Nationwide Distribution. | Terminated | April 17, 2014 | 653 | 165 | Not_Computer
Z-0269-2013 | 63599 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TI Standard Polyaxial Screw Assembly 6.5MM x 55MM. Part # 62065-55. |  |  | December 09, 2010 | 2010 | Alphatec Spine, Inc. | The recall was initiated because Alphatec Spine confirmed that the TI Standard Polyaxial Screw Assembly was laser marked with the incorrect length of 45 MM. | Error in labeling | Alphatec Spine distributors who received the TI Standard Polyaxial Screw Assembly were contacted by phone and by email on December 9, 2010. A phone script was used as a recall notification to contact their three distributors who received the recalled products. The recall notification informed the customers of the problem identified and the action to be taken. Customers were informed that they will receive replacements within 2-3 days. Customers with questions were instructed to call (800) 922-1356.For questions regarding this recall call 760-431-9286. |  | 11 | 11 units | Nationwide Distribution including TN, WI, and FL. | Terminated | November 09, 2012 | 701 | 11 | Not_Computer
Z-0267-2013 | 53337 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Quadra S Offset Broach Handle |  |  | October 13, 2008 | 2008 | Medacta Usa Inc | Medacta USA Inc has initiated a recall on the Quadra S Offset Broach Handle because of the potential of the rivet that locks the metallic wire to the lever may become loose during a surgery and make it impossible to lock/unlock the broaches. | Nonconforming Material/Component | Medacta USA, sent a Urgent Safety Notification letter dated October 13, 2008, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.The letter informed the customers of the lot numbers on recall and the reference numbers. The recall was initiated because the offset broach handles contain a lever mechanism that is connected to the locking axis by means of a metallic wire. The rivet that locks the metallic wire to the lever may become loose during a surgery and make it impossible to lock/unlock the broaches. For further questions please call (858) 705-0350. |  | 12 | 12 units | Worldwide Distribution -- US including the states of Texas, Nevada and Colorado., and the country of Belgium. | Terminated | November 09, 2012 | 1488 | 12 | Not_Computer
Z-0276-2013 | 63445 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Progressive Medical PMI Endoscopic Dissector with Kittner Tip (5mm x 40cm), Reorder Number: PMILK40, Product Sterile in Unopened/Undamaged pouch, Cotton Tip - X-Ray Detectable, 1 Pc/Pouch, 10 Pouches/Box, 20 boxes/case. Exclusively Distributed in the U.S. Market by: Progressive Medical, Inc., St. Louis, MO, Made in China. |  |  | December 09, 2011 | 2011 | Progressive Medical Inc | Some of the sterile packages had broken seals. | Packaging | Progressive Medical Inc contacted their customers via telephone on December 9, 2011, through December 15, 2011.Customers were informed the reason for recall was explained and the customer was asked if any of the lot remained in their inventory. If inventory remained, the recalling firm sent them a return goods authorization number and informed them the product would be replaced. For further questions please call (314) 961-5786. |  | 70 | 70 boxes | USA (nationwide) including the states of CO, IN, KS, MA, MN, MO, NE, OH, SC, TN, TX, VA, and WI. | Terminated | November 14, 2012 | 341 | 70 | Not_Computer
Z-0287-2013 | 53336 | bath, flotation, tissue | IDY | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | PT Link Waterbath, Model#PT100. |  |  | May 01, 2008 | 2008 | Dako North America Inc. | Dako is recalling the PT Link Waterbath because of a possible electrical hazard. Specifically, if the heater pad under the reagent tank becomes wet, it may present an electrical shock hazard when the lid is open. | Device Design | The firm, Dako, sent a "DAKO PT LINK ADVISORY NOTICE" dated May 1, 2008 to customers who purchased the PT Waterbath along with a customer acknowledgement document informing the users of the potential electrical hazard. The notice described the product, problem and actions to be taken. The customers were instructed to follow the attached Safety Instructions; place the enclosed warning labels directly on the instrument as shown in the attached instructions and adhere to the warning labels as described. Customers with questions were instructed to call Dako Products' Technical Support at (800) 424-0021. |  | 136 | 136 units | Nationwide distribution: USA including states of: AK, CA, CO, CT, DC, FL, GA, HI, IA, IL IN, KS, KY, MA, MD, MI, MN, MO, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, and WA. | Terminated | November 13, 2012 | 1657 | 136 | Not_Computer
Z-0288-2013 | 63586 | No | MIH |  |  | Cardiovascular | PMA | TriVascular brand Fill Polymer Kit, Endovascular Graft, Aneurysm Treatment System, European Model Number: TVFP14-B, (in European distributor inventory only); |  |  | October 19, 2012 | 2012 | Trivascular, Inc | Selected lots of the Fill Polymer Kit (European Model Number: TVFP14-B) in European distributor inventory only that potentially contain leaking stopcocks due topossible cracks in the stopcock component. TriVascular indicated that these Fill Polymer Kits are for use with the Ovation or Ovation Prime Abdominal Stent Graft Systems that may display leaks upon initiation of mix but prior to use in | Nonconforming Material/Component | Trivascular sent an Urgent Field Safety Notice letter dated October 19, 2012 to all affect consignees. A verbal (telephone) communication was also initiated. The forms communication identified the affected product, problem and actions to be taken. Consignees were instructed to immediately quarantine and return affected products. A return response form was included for consignees to complete and returned via the e-mail address or fax number provided. Consignees were instructed to further notify its (sub-level) consignees, and have them return affected product directly to the recalling firm. For questions call 1-707-543-8732. |  | 41 | 41 units | Worldwide Distribution - US Nationwide and the countries of Belgium, Cyprus, Italy, Spain, Switzerland and Turkey. TriVascular stated that no purchased devices are in hospital inventory in Europe or the U.S. | Terminated | December 17, 2012 | 59 | 41 | Not_Computer
Z-0282-2013 | 62733 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Perfusion System 8000, base 4 pump, 220V/240V safety monitor included |  |  | September 25, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) has received multiple reports since 2007 of various malfunctions of the arterial and cardioplegia monitors for the Sarns ModularPerfusion System 8000. The malfunctions can be broadly categorized as follows monitor power malfunctions, inability to measure pressure and inaccurate and/or false pressurealarms.If the malfunction of the arterial or cardio | Device Design | The firm, Terumo CVS, sent an "URGENT MEDICAL DEVICE RECALL" notice/Safety Advisory Alert dated September 24, 2012 and an addendum to the Operator's Manual to alert all affected Customers. The notice described the product, problem and actions to be taken. The customers were instructed to review this medical device recall notice; assure that all users are aware to this notice; place the addendum in the Operator's Manual just before page 4.1 and complete and return the attached Customer Response Form via Fax to: 1-734-741-6149 or email to:www.termuo-cvs.com.If you have any questions or concerns contact Terumo CVS Customer Service at 1-800-521-1818; Monday - Friday, 8 AM - 6 PM ET. |  | 466 | 466 | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY; and countries of: ARGENTINA, AUSTRALIA, BANGLADESH, BRAZIL, CALI COLOMBIA, CALLO PERU, GERMANY, GUAYQUIL ECUADOR, HONG KONG, INDIA, INDONESIA, JAPAN, LEBANON, MADRID SPAIN, MEXICO, NEW ZEALAND, SANTIAGO, SANTIAGO CHILE, SINGAPORE, SOUTH AFRICA, SOUTH KOREA, SRI LANKA, TAIWAN, THAILAND, and UNITED ARAB EMIRATES (UAE). | Terminated | July 11, 2013 | 289 | 2256 | Not_Computer
Z-0277-2013 | 63446 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes External Fixation System Tube to Tube Clamp, Part Number 390.007. |  |  | September 28, 2012 | 2012 | Synthes USA HQ, Inc. | The insert provided with the product for the specific lot number indicates that this product is MR safe. This product is MR Conditional and is etched as MR Conditional. The insert is incorrect. | Labeling mix-ups | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated September 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Verification Form via fax to 610-251-9005. Contact the firm at 610-719-5450 for questions regarding this notice. |  | 8 | 8 | Nationwide Distribution-including the states of CA, MN, NC, and TN. | Terminated | September 21, 2015 | 1088 | 8 | Not_Computer
Z-0278-2013 | 63344 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | HX2 Temp Manage SYS 120V |  |  | October 16, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Based on a retrospective review of quality data, AC Control boards used in the manufacture of theHX2 Temperature Management System were identified as being manufactured using the incorrectflux material. Use of the incorrect flux may contribute to premature board failure. A search ofTCVS complaints revealed 3 complaints received related to the HX2 failure. Each failure occurredduring set up | Nonconforming Material/Component | TERUMO sent an Urgent Medical Device Recall letter dated October 15, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Terumo has determined that a field correction is required to notify affected customers that the boards may fail, instruct the users that in a case of failure to swich to the second channel on the device or use a back-up unit, and that TCVS will replace the affected boards. Customers with questions were instructed to call 1-800-621-2818. For questions regarding this recall call 734-741-6173. |  | 19 | 19 units | Nationwide Distribution including CA, HI, KS, LA, MI, MO, OR, PA, UT, and WI | Terminated | July 11, 2013 | 268 | 34 | Not_Computer
Z-0286-2013 | 63275 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Special Orthopaedic Solution |  |  | July 16, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has become aware that insufficient taper torsional strength may result when a PAR extension piece is used with a forged CoCr stem. | Device Design | Stryker Orthopaedics sent an Urgent Product Recall letter dated July 16, 2012 via Fed Ex with return receipt to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter requested customers to contact their Stryker Sales Representative to arrange for return of the affected product. Customers were instructed to fax back the attached Product Recall Acknowledgement Form to 201-821-6069. For questions call 201-972-2100. |  | 301 | 301 units | Worldwide Distribution - US Nationwide | Terminated | September 09, 2015 | 1150 | 301 | Not_Computer
Z-0289-2013 | 63536 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Creatinine 120 CP, Model Number Creatinine Part # A11A01868, Lot Numbers: 203273706, 203273707, 203273808, 203273709. |  |  | September 14, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | This recall was initiated because some customers were experiencing an alarm during Creatinine calibration "REAG RANGE HIGH". This alarm caused the calibration to fail and prevented the user from using the Creatinine assay. | Nonconforming Material/Component | HORIBA Medical sent customer communication letters dated October 24, 2011, to all affected customers via Fax with the exception of two communications that were emailed. The response or confirmation of receipt was conducted via telephone and also documented on attached Notification List. Customers were given a list of procedures to follow and asked to complete and fax the Acknowledgement of Receipt Form to +33 (0)4 67 14 15 17.For questions regarding this recall call 949-453-0500, ext 208. |  | 8503 | 8,503 units | Nationwide Distribution | Terminated | November 29, 2012 | 76 | 8503 | Not_Computer
Z-0295-2013 | 63603 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Select Medical Products Vinyl Powder-Free Exam Gloves, Medium Model # 6370 and Large Model # 6371, 100 per box; 10 boxes per case. |  |  | June 20, 2012 | 2012 | Gulf South Medical Supply Inc | Gulf South Medical Supply recalled their Select Medium and Large size Vinyl Exam Gloves because they were inadvertently shipped prior to U.S. Customs clearing them for sale. | Employee error | The firm, PSS World Medical, sent an "URGENT RECALL NOTICE" dated June 20, 2012 via USPS to all customers that received the product. The notice described the product, problem and actions to be taken. The customers were instructed to examine their inventory, quarantine the affected lot numbers listed in this recall and return the product; If the product was further distributed, to inform their customers of the recall and request they return the affected product directly to them; and contact their Gulf South Customer Care team via telephone at 1-800-347-2456 or email to customercare@gsms.com to request a return goods authorization.If you have any questions, contact the Product Integrity Leader Quality Assurance at 904-380-4501. |  | 292 | 275 cases and 17 boxes | Nationwide distribution: USA including states of: DE, KY, LA, MD, NJ, NY, PA, TX, VA, and WV. | Terminated | August 19, 2014 | 790 | 292 | Not_Computer
Z-0296-2013 | 63458 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ATS 3000 Automatic Tourniquet System |  |  | September 05, 2012 | 2012 | Zimmer, Inc. | Zimmer is conducting a voluntary removal of some Zimmer A.T.S. 3000 tourniquets after receiving complaints of both out-of-box failures and failures during surgical procedures on some of the affected units. During failure the unit will alarm, lock into a non-operational mode, and the display screen will go blank. Based on investigation, this problem appears to be due to a counterfeit chip on the | Counterfeit | Zimmer sent an Urgent Device Removal Notice letter dated September 5, 2012, to all affected customers. Zimmer Distributors and Sales Leadership were notified via E-mail of the pending recallthe product under recall beginnig September 11, 2012. The notification listed the reason for the recall, risks involved, and requested return of the unit for repair To: Zimmer Surgical, Attn: QAlRA Dept. - Recall, 200 West Ohio Avenue, Dover, Ohio 44622 USA The notice asked to ensure all users be made aware of the recall, remove the affected ATS 3000 tourniquest sytem for repair and return the fax back form/acknowledgement certification to 1-888-429-7380 or email it to zimmer3611@stericycle.com. Questions should be directed to 1-888Âż943Âż5167 between 8 a.m. and 5 p.m. EDT.For questions regarding this recall call 574-372-4807. |  | 204 | 204 | Worldwide Distribution - USA (nationwide) and Internationally to Belgium, Czech Republic, Egypt, France, Germany, Morocco, Netherlands, Romania, Saudi Arabia, Spain, Switzerland, Turkey, and UAE. | Terminated | April 17, 2014 | 589 | 204 | Not_Computer
Z-0294-2013 | 63660 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | EnteraLite Infinity Enteral Feeding Pump identified with the following list numbers: INFKIT2 and INFORNG. The device is used with disposable PVC administration tubing sets to provide a means for the volumetric delivery of fluids used in enteral applications. |  |  | November 08, 2012 | 2012 | MOOG Medical Devices Group | MOOG Medical Devices Group is conducting a voluntary for a limited set of serial numbers of the EnteraLite Infinity and Infinity Orange Small Volume Enteral Feeding Pumps because certain diodes on the PCB assembly may result in insufficient voltage to a portion of the pump resulting in an error 13 message and shutting the pump down entirely. | Component design/selection | MOOG Medical Devices Group sent a Urgent Medical Device Recall letter dated November 8, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to contact Moog Medical Devices Group to discussthe process for product replacement at 1-800-970-2337, prompt #7.See Attachment A to determine if you have a serial number that is related to a suspect pump. Determine the number of affected pumps. Notify all of your affected customers of the recall and the protocol for getting a replacement pump. Provide serial numbers, number of product(s) to be returned and a shipping address for replacement product. Execute the pump replacement process discussed in the conversation with Moog Customer Advocacy. Package and return product via shipping label. |  | 8736 | 8,736 | Worldwide Distribution - USA (nationwide) including the countries of Canada and Lithuania. | Terminated | January 30, 2014 | 448 | 8736 | Not_Computer
Z-0198-2013 | 63597 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima XR 220amx |  |  | September 20, 2012 | 2012 | GE Healthcare, LLC | It was discovered GE Healthcare has become aware of a firmware issue on the mobile radiographic product, Brivo XR285amx, Optima XR200amx, Optima XR22Oamx and Optima XR200amx with Digital Upgrade, related to Technique Accuracy as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-2-54. | Radiation Control for Health and Safety Act | GE Healthcare will remedy the defect or bring the product into compliance with each applicable Federal and IEC standard. A Service Representative will update the firmware on the system to address the issue.Field Modification Instructions (FMI) 10870, 10872, and 10873 are describing the rework plan the GEHC Engineers will make to the imaging systems compliant with the applicable performance standard.Further questions please call (262) 513-4122. |  | 364 | 364 total units installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 441 | 364 | Not_Computer
Z-0292-2013 | 63322 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | ETHICON SURGIFLO HEMOSTATIC MATRIX KIT Plus FlexTip with Thrombin distributed by ETHICON ETHICON Inc. Somerville, NJ 08876 USA |  |  | August 21, 2012 | 2012 | Ethicon, Inc. | SURGIFLO Hemostatic Matrix Kit &amp; SURGIFLO Hemostatic Matrix Kit with Thrombin have an issue within the packing process where a cut could potentially breach the double Tyvek pouch of the packaging. | Packaging | Ethicon Inc. sent an Urgent Product Information letter dated August 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory to determine if they have any of the affected product. Affected product should be returned for replacement to Stericycle using the enclosed pre-paid shipping label. Customers with questions should call 1-877-384-4266 or contact their Ethicon Biosurgery Sales Representative.For questions regarding this recall call 877-384-4266. |  | 28276 | 28,276 | Worldwide Distribution - USA (nationwide) and the countries of Russia, China, and the European Union. | Terminated | September 11, 2014 | 751 | 73476 | Not_Computer
Z-0299-2013 | 63467 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | bioMerieux BacT/ALERT Control Module, catalog number 210147, Clinical instrument |  |  | October 16, 2012 | 2012 | Biomerieux Inc | The product may mislabel patient information on culture specimen bottles due to a software miscommunication. | Software design | BIOMERIEUX sent an Urgent Product Correction Notice to all affected customers starting on October 16, 2012, The letter identified the affected product, the problem, and the action to be taken by hte customer. The firm provided modified instructions for use to clarify how to properly load bottles and to provide appropriate cautions and warnings concerning bottle loading. Customers were instructed to ensure that the letter and guidance on how to properly scan and load bottles is distributed, and training is provided to all appropriate personnel within their organization. Check for lost bottles in their system; fill out and return the Acknowledgement Form in Attachment A by fax to confirm receipt of the notice. Customers were instructed to contact their local bioMerieux Customers Service Representative with any questions.For questions regarding this recall call 919-620-2682. |  | 70 | 70 modules | Worldwide Distribution - USA including AR, AZ, CA, FL, KS, LA, MD, MN, MT, NC, NV, NY, OR, PA, RI, SC, TX, VA and WA. Internationally to Argentina, Armenia, Austria, Bangladesh, Bosnia and Herzegowina, Brazil, Bulgaria, Burkina FASO, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DIVOIRE, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LITHUANIA, MALTA, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, PANAMA, PERU, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UNITED STATES, URUGUAY, VENEZUELA, VIETNAM, PALESTINE, AND SERBIA. | Terminated | October 28, 2013 | 377 | 6400 | Computer
Z-0331-2013 | 63581 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Superstat Modified Collagen Hemostatic Sponge.115mmx115mm Pledget, 1.5%, Part#9500-15. |  |  | October 24, 2012 | 2012 | Superstat Corp | The recall was initiated because Superstat Corporation has confirmed that no documentation are available to verify that quarterly dose sterility audits were performed for the Superstat Modified Collagen Hemostatic Sponge. | Process control | Superstat initially contacted customers by telephone and sent an Urgent - Device Recall Notification Letter on October 24, 2012, to all affected customers. The recall letter provided the reason for recall was due to inadequate documentation of quarterly dose sterility audits. The recall includes all lots and part numbers of Superstat produced between 2002 and 2012. Customers are instructed to complete and return the enclosed response form as soon as possible. Customers with questions are instructed to call customer service at (800) 487-3786.For questions regarding this recall call 310-605-1655. |  | 26422 | 26,422 units | Worldwide Distribution - USA (nationwide) and the countries of Saudi Arabia, Thailand, Mexico, Canada, Pakistan, Israel, Greece, Serbia, Canda, Isle of Man, Nova Scotia, and Singapore. | Terminated | February 06, 2013 | 105 | 26422 | Not_Computer
Z-0330-2013 | 63481 | No | HGP | Fetal scalp circular (spiral) electrode and applicator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Corometrics Qwik Connect Plus Spiral Electrode- Model #7000AA0 |  |  | October 16, 2012 | 2012 | Vital Signs Devices, a GE Healthcare Company | Thirty fetal scalp electrodes were found to have an incomplete package seal and lacked the printed lot number stamped on during sealing. | Employee error | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter on October 16, 2012 via Fed Ex to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: DO NOT use Qwik Connect Plus Spiral Electrodes part number 7000AAO with lot number S03123 or product with missing lot numbers; isolate all affected product; complete and return the Urgent Medical Device Correction Confirmation form via fax at 800-535-7923, return form even if no recall product is in inventory; to return only unused product, and if they have forwarded any affected lot numbers of this product to any other healthcare institutions, forward a copy of this letter to those institutions.If you have any questions or concerns regarding this notification, contact GEMSIT Customer Service at +1-800-588-7044 (Domestic) or Vital Signs Customer Service at +1-800-932-0760 (International). Hours of operation: 8:00 am EST to 6:00 pm EST. |  | 126 | 126 cases (50 pouches per case) | Worldwide distribution: USA (nationwide) including states: AZ, CA, CT, FL, GA, IL, KS, KY, MA, MD, MN, MO, NJ, NM, NY, OH, PA, SC,TX, WA, and WY; and countries of: Australia and Canada. | Terminated | April 16, 2014 | 547 | 126 | Not_Computer
Z-0297-2013 | 63618 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Product is Full OSSEOTITE 2 Tapered Certain Implant, Model Number XIFNT413, and size 4 x 13 mm. The product is sterile. |  |  | September 28, 2012 | 2012 | Biomet 3i, LLC | Biomet 3i recalled Full OSSEOTITE 2 Tapered Certain Implant and Full OSSEOTITE Parallel Walled Certain Implant due to a manufacturing condition where the depth of the implant's internal hex is too shallow. This condition may prevent the driver or abutment from fully engaging in the implant. | Nonconforming Material/Component | Biomet 3i sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 28, 2012. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-342-5454 for questions regarding this notice. |  | 191 | 191 devices | Worldwide Distribution-USA (nationwide) including the states of VA, SC, HI, LA, OH, FL, IL, NJ, NC, CA, CT, MI, TN, AK, NH, NY, MA, GA, PA, and KY and the countries of United Kingdom, Spain, Germany, Canada, Japan, Singapore, Taiwan, Israel, Poland, Hong Kong, Ireland, The Netherlands, Austria, and France. | Terminated | August 06, 2015 | 1042 | 267 | Not_Computer
Z-0291-2013 | 63314 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Mindray A5 Anesthesia Delivery System |  |  | August 13, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified the potential for a system leak resulting an issue related to the carbon dioxide absorber canister gasket on A3/A5 Anesthesia Delivery Systems. A potential system leak resulting from the canister gasket will likely present itself during the Automatic Circuit Leak and Compliance Test performed at startup and the Manual Leak Test recommended to be performed before each case. | Device Design | The firm, Mindray DS USA Inc., sent an "URGENT: A3/A5 ANESTHESIA DELIVERY SYSTEM CORRECTIVE ACTION" letter dated August 13, 2012 to their customers via Certified Mail return receipt requested. The letter described the product, problem and actions to be taken. The customers were instructed to route this letter to all potential uses of the product within their facility; increase fresh gas flow, should the leak occur during use, to compensate for any pressure loss resulting from the leak; alternatively, the absorber canister may be unlocked and locked again to reseat the gasket and verify that the canister and gasket are fully seated, a leak test is performed after replacement of the Pre-Pak or loose-fill absorbent and also be performed before each patient use until the canister gaskets are replaced on systems in their facility. A Mindray Representative will replace the canister gasket to correct the issue.If you have any questions or need to arrange for a replacement of the absorber canister gasket on the A3/A5 System(s) in your facility, contact your Mindray Service Representative, at 1-800-288-2121 (Monday-Friday 8:30am-5:30pm).A press release was issued by the firm on 11/14/2012. |  | 346 | 346 units (311 US - 35 Foreign) | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Australia, Chile, Colombia, and Venezuela. | Terminated | June 04, 2013 | 295 | 415 | Not_Computer
Z-0344-2013 | 62916 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | IO Fix IntraOsseous Fixation System; Extremity Medical. www.extremitymedical.com; Customer Service: 888.499.0079. |  |  | April 13, 2011 | 2011 | Extremity Medical LLC | Product Event Reports revealed that lag screw passes through washer (e.g. X-Post) intra-operatively. | Employee error | Extremity Medical sent an e-mail notification (Extremity Medical- Sales Bulletin- IO Fix) dated April 13, 2011 to distributors notifying them of a change made to the 4.0 mm Lag Screw. |  | 120 | 120 Units | Worldwide Distribution-USA (nationwide) including the states FL, MA, NC, CA, MD, VA, IN, AL, NJ, AZ, OK, MT, NY, OH, IA, MI, TX, WI, GA, IL, TN and NE, and the country of Switzerland. | Terminated | March 21, 2013 | 708 | 120 | Not_Computer
Z-0368-2013 | 52634 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 4.5mm x 35mm Noncannulated Screw Shank |  |  | January 05, 2009 | 2009 | NuVasive Inc | The tulip portion of the screw assembly can disengage from the shank after implantation. | Device Design | Telephone communications to the sales representatives of surgeons actively using the Posted Pedicle Screw System were made on 01/05/2009. The calls were unscripted and requested that all inventory be forwarded to NuVasive's San Diego office for quarantine. Surgeons who were actively using the system were also contacted and notified of the situation. Follow-up communications were sent via mail on 01/19/2009 to all surgeons who participated in the Alpha release. The letter informed the customers of the problems identified. Customers were instructed that the recalling firm does not believe any further action is required and they were instructed to monitor the progress of their patients and notify NuVasive if any issues are detected.Customers with questions or concerns are instructed to contact the firm at (858) 909-1964. |  | 240 | 240 units total | Nationwide Distribution-including the states of OR, MO, FL, WI, CA, and NC. | Terminated | November 19, 2012 | 1414 | 240 | Not_Computer
Z-0379-2013 | 62987 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Thermo Scientific MAS Omni IMMUNE and Omni IMMUNE PRO; |  |  | August 20, 2012 | 2012 | Microgenics Corporation | In house testing confirmed that CK-MS and Insulin claimed in both Omni IMMUNE and Omni IMMUNE PRO products are displaying reduced stability. | Nonconforming Material/Component | Microgenics Corporation sent a Recall Notification letter dated August 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed that immediate attention is required and identify the two affected constituents. The letter contains a list of the affected product with their catalog number, lot number and expiration date and request that the CK-MB and Insulin use be discontinued. The remaining components of the product may be used. Evaluate quality controls according to your laboratory procedures.Retain a copy of the letter for your laboratory records. If the product was forwarded to another laboratory, provide them with a copy of the letter.Complete the attached Recall Response form and return within 5 days.Questions may be directed to 1-800-232-3342. |  | 2076 | 2076 total in US , 596 total Foreign | Worldwide Distribution--USA (nationwide) and the countries of Germany, Canada and Australia. | Terminated | March 04, 2013 | 196 | 2076 | Not_Computer
Z-0345-2013 | 63505 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel Dxl 800 Access Immunoassay System, Part Numbers: 973100, A71456, and A71457. |  |  | August 01, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the UniCel Dxl and DxC Access Immunoassay Systems had field modifications to the aspirate and dispense probe assemblies. | Process design | Beckman Coulter sent a field modification was released in August 2011 which included:- Installing a new aspirate and dispense tubing routing bracket.- Replacing existing aspirate and dispense probe tubing with modified tubing to help eliminatemisrouting of aspirate and dispense probe tubing.During installation the FSE informed the customer:-That aspirate and dispense tubing are different lengths.- That probes are not interchangeable with other instruments once tubing has been modified.- To route probe tubing by referring to the labels on the tubing.An updated service document, Routine Maintenance Changes, A82294B was provided to every customer.For questions regarding this recall call 714-961-4483. |  | 1756 | 1,756 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, AE, AU, BG, BW, CH, CN, CZ, DE, ES, FR, GB, HK, HR, HU, IN, IT, JO, JP, LB, MA, MX, NL, PL, QA, RO, RU, SK, TN, TR, ZA, AO, AR, AT, BE, CO, DZ, FI, IE, IL, KR, KW, LI, NA, NO, NZ, PF, PH, PR, SE, SG, SZ, TH, TW, VE, VN, and YT. | Terminated | November 19, 2012 | 476 | 2188 | Not_Computer
Z-0383-2013 | 63077 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | "***SENTINEL***Cannulated Drill Bit***S8580***8 X 229mm (9")***CONMED LINVATEC 11311 CONCEPT BLVD LARGO, FL 33773-4908 USA***LOT***2012-03***". |  |  | April 10, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | Linvatec Corporation dba ConMed Linvatec in Largo, FL is recalling the S8580, 8.0MM SENTINEL Cannulated Drill Bit, lot numbers 322225 and 322226; and the S8585, 8.5MM SENTINEL Cannulated Drill Bit, lot number 322287. Recall was initiated due to the incorrect labeling on package. | Labeling Change Control | ConMed/Linvatec sent an Urgent Medical Device Recall Notificaiton letter via phone and FedEx on April 9, 2012 and again on May 11, 2012, to all affected customers. They also sent the same Notification on May 14, 2012, to international accounts. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete a reply formand fax it to 727-319-5701. Customers were also instructed to return all affected product to the firm. All returned product under this removal action will be identified, segregated and quarantined per firm procedures, etc.Customers with questions should contact Customers Service at Custserv@linvatec.com.For questions regarding this recall call 727-392-6464. |  | 70 | 70 total | Worldwide Distribution - USA including NJ, UT, GA, NY, PA, IL, LA, OH, CA, MS and Internationally to Canada, Japan, and Romania. | Terminated | November 21, 2013 | 590 | 70 | Not_Computer
Z-0385-2013 | 63607 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | The RIO (TGS 2.), Model # MAKO TGS 2.0 (Part No. 204000). |  |  | September 12, 2012 | 2012 | Mako Surgical Corporation | MAKO Surgical Group recalled their RIO System software, version 2.4 and is implementing software version 2.5 to address a software functional issue with the existing version of the system's software. Loss of tactile feedback constraining the cutting burr has been reported. | Software design | Mako Surgical Corporation sent an Important Medical Device Correction letter dated August 31, 2012, to all affected customers. The letter included specifics regarding the software issue and how the correction will take place. The letter informed the customers that their MAKOplasty Sales Specialist will coordinate with them to establish the appropriate time for the upgrade .Customers were asked to return by fax an Acknowledgement Form to 954-927-0446, ATTN: Quality Assurance. For questions customers were instructed to call 954-628-0447.For questions regarding this recall call 954-927-2044, ext 605. |  | 137 | 137 devices | Worldwide Distribution - USA including IL, OR, IA, AZ, CA, CO, FL, MS, OK, PA, WV, NC, TX, NY, GA, NJ, OH, WA, VA, MA, MD, WI, MI, NV, MO, AL, UT, RI, AR, LA, ID, IN, and TN and Internationally to Japan, Scotland, Italy, Singapore, Turkey, Hong Kong, and Korea. | Terminated | May 22, 2013 | 252 | 137 | Computer
Z-0387-2013 | 63358 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics |  |  | August 17, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has received reports from the field indicating that certain lots of 6.5mm Cancellous Bone Screws are either associated with a product mix or potentially associated with a product mix. Packaging associated with a 35 mm screw may contain a 25mm screw. Conversely, packaging associated with a 25mm screw may contain a 35mm screw. | Mixed-up of materials/components | The firm, Stryker Orthopaedics, notified Domestic locations of this action by e-mail on August 6, 2012 and foreign locations of this action by e-mail on August 7, 2012. Stryker follow-up by sending an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated August 17, 2012 to their consignees/customers via Fed Ex with return receipt. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine and return affected product to Attn: Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430 (attach fluorescent orange sticker and mark box with words "Product Recall"), or contact your Stryker Sales Representative to arrange for return of the product ,and complete and return the attached Product Recall Acknowledgment Form within 5 days via fax to 855-251-3635.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance at (201) 972-2100. |  | 72 | Approximately 72 units | Worldwide distribution: USA (nationwide) and countries including: Canada, Belgium, Finland, Germany, Greece, Hong Kong, Ireland, Italy, Japan, Poland, Portugal, Spain, Turkey and UK. | Terminated | April 10, 2014 | 601 | 430 | Not_Computer
Z-0382-2013 | 63386 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SPECIALTY AUDIBLE TORQUE WRENCH, XIA |  |  | August 09, 2012 | 2012 | Stryker Spine | Multiple complaints have reported the hex tip of the torque wrench fracturing during final tightening. | Nonconforming Material/Component | Stryker Spine sent an Urgent Medical Device Removal Notification letter/Product Accountability Forms dated August 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.For questions regarding this recall call 201-760-8150. |  | 28 | 28 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, France, Germany, Netherlands, Japan, Spain, Denmark, Switzerland, and the United Kingdom. | Terminated | September 30, 2015 | 1147 | 101 | Not_Computer
Z-0389-2013 | 63066 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Product 100789502202 VerSys Femoral Head Provisional 22mm +0 |  |  | October 10, 2010 | 2010 | Zimmer, Inc. | Surgical technique instructions for the Versys and Slotted Provisional Femoral Heads have been enhanced to inform users that the head may become difficult to visually locate if it becomes displaced deep or medial to the iliopsoas tendon and provide non-invasive imaging methods that can be used to locate the head. To reduce incidence of sub-optimal fit between provisional femoral heads and the mat | Device Design | The firm, Zimmer, sent notification by email to consignees beginning October 10, 2012 with a follow-up letter titled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2012 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete a spreadsheet of current users and contact information and immediately provide this information via email to: CorporateQuality.Postmarket@zimmer.com, and review and become familiar with the content and recommendations of the attached letter sent to the customers. For questions call 1-877-946-2761. |  | 128946 | 128,946 all lots, all sizes | Worldwide distribution: USA (nationwide) and countries of: Asia, Barbados, Canada, Europe, Mexico, Puerto Rico, San Juan, and South America. | Terminated | April 30, 2014 | 1298 | 128946 | Not_Computer
Z-0386-2013 | 63126 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trilogy II Constrained Liner |  |  | August 27, 2012 | 2012 | Zimmer, Inc. | Trilogy Longevity Constrained Liner devices were packaged in incorrect cartons. The product is packaged in the incorrect carton. Product labels and patient labels are correct, matching the product inside the carton. | Labeling Change Control | Zimmer contacted their customer beginning the week of August 27, 2012, via telephone. Zimmer informed the customer the reason of the recall. The customer was instructed to quarantine and return the device to Zimmer Product Service, 1777 W. Center St. Warsaw, IN 46580. The customer was asked to provide information if the product was further distributed. For further questions please call 574-372-4316. |  | 1 | 1 | USA (nationwide) including the states of Ohio, Carolinas and New England. | Terminated | August 27, 2013 | 365 | 1 | Not_Computer
Z-0406-2013 | 63596 | No | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | INTRABEAM Balloon Applicator Set; |  |  | October 25, 2012 | 2012 | Carl Zeiss Meditec, Inc. | A manufacturing defect has been identified which could results in the small tube becoming dislodged between the luer fitting and the check valve. | Nonconforming Material/Component | Carl Zeiss Meditec sent an Important Field Corrective Action letter dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return the affected product to Carl Zeiss Meditec, Inc. A Carl Zeiss Meditec representative will contact the customer to schedule a visit to support the removal of the affected product. For questions customers were instructed to call 925-580-5371 or by email at v.venturo@meditec.zeiss.com.For questions regarding this recall call 925-557-4616. |  | 36 | 36 total, all lots | Nationwide Distribution including IL and MD | Terminated | November 20, 2012 | 26 | 36 | Not_Computer
Z-0409-2013 | 62805 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoRest Delta Kinetic Therapy System |  |  | June 27, 2012 | 2012 | KCI USA, Inc. | The firm is initiating a voluntary medical device correction for the RotoRest Delta Therapy System head pack. KCI has received complaints alleging patients experiencing skin breakdown on the head and face due to the positioning of the head pack. In a few instances, the patient received injuries ranging from Stage III to Stage IV pressure ulcers. | Component design/selection | KCI sent an Urgent - Voluntary Medical Device Correction letter dated June 27, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that KCI would contact them soon to coorinate a shipment of a replacement of the affected product. For any questions customers were instructed to call 1-800-275-4524, Selection Option 3 followed by Option 2.For questions regarding this recall call 1-800-275-4524. |  | 282 | 282 units | Nationwide Distribution including AL, CA, CO, OK and GA. | Terminated | April 22, 2013 | 299 | 282 | Not_Computer
Z-0401-2013 | 63557 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | BioPlant, a dental bone grafting material. BioPlant |  |  | October 25, 2012 | 2012 | Kerr Corporation | The firm initiated the recall becaue disturbing the BioPlant product soon after it is placed may halt or delay the ossification process. | Labeling design | Kerr Corporation sent an Urgent Medical Device Correction letter and acknowledgement form dated October 25, 2012, to all affected customers via USPS First Class. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: "...review the table with product description , part numbers and lot numbers to determine if they have any of the affected product in their inventory and confirm that they were aware of the additional placement instructions necessary to achieve desired ossification results detailed in the paragraph in the letter. Customers were instructed to contact Kerr Customer Care at 1Âż800Âż537Âż7123 if they have any questions. Customers were also instructed to complete and fax back the enclosed Acknowledgement Form to 714-288-4609 in order to confirm their receipt of this correction regardless of whether they have any affected product in their inventory.For questions regarding this recall call 714-516-7400. |  | 22410 | 22,410 units | Worldwide Distribution - USA (nationwide) Canada, Europe, Australia and ROW consignees. | Terminated | September 23, 2013 | 333 | 22410 | Not_Computer
Z-0403-2013 | 63504 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | qube Compact Monitor (Patient Monitor) 91390. The monitor is 26.2 cm x 31.5 cm x 13.2 cm and weighs 4.1 kg (10.3 inches x 12.4 inches x 5.2 inches - 9 lbs). It has a 26.2 cm (12.1 inches) LCD display with 1024 x 768 resolution. |  |  | October 19, 2012 | 2012 | Spacelabs Healthcare, Llc | Spacelabs Healthcare has learned through severalreports that the input circuit may fail and the monitorwill go on battery mode. The monitor will functionnormally until the batteries become depleted and thenthe monitor will shut down. In addition the monitor mayfail to turn on due to an internal short. | Nonconforming Material/Component | Spacelabs Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated October 19, 2012. The letter identified the product, problem, and actions to be taken by the customers. Contact Spacelabs Healthcare at 1-800522-7025 option 2 for questions regarding this notice. |  | 742 | 742 total devices distributed; 530 with the U.S. and 212 Internationally. | Worldwide Distribution.On 4/16/13 the firm advised FDA that they added serial number 1390-000055 and 1390-000913. The two additional serial numbers were added because they were provided as replacements to a French customer whose devices had been returned to support the firm's investigation. The customer had already received the Urgent Medical Device correction, customer letter and required Spacelabs to provide these devices with the understanding that they would require the field corrective action | Terminated | May 22, 2013 | 215 | 742 | Not_Computer
Z-0407-2013 | 63532 | No | KAK | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Vital Signs Disposable General Purpose 9 French Temperature Probe |  |  | August 17, 2012 | 2012 | Vital Signs Colorado Inc. | There is a potential for epistaxis during use of the Disposable General Purpose 9 French Temperature Probe. | Device Design | The firm,GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter and Confirmation form on August 17, 2012 via Fed Ex to its consignees/customers. The letter describes the product, problem and action to be taken. The customers were instructed to do the following: Do not use the Disposable General Purpose 9 French Temperature Probe P/N 1016 or 1016EU; return all unopened product(s) to Vital Signs Devices (VSD) 11039 East Lansing Circle, Englewood, CO 80112-note: replacements and/or refunds discussed in letter are available only in exchange for the return of any unopened product(s); complete and return the URGENT MEDICAL DEVICE CORRECTION CONFIRMATION via fax to: 800-535-7923; forward a copy of the letter to any institutions they may have forwarded the product and for customers using the GXXXXX64 kit that includes the Disposable General Purpose 9 French Temperature Probe follow the directions listed in the letter.Note: VSD is removing p/n 1016 and 1016EU, Disposable General Purpose 9 French Temperature Probe, from the market and offering a replacement. If you have any questions or concerns regarding this notification, contact Customer Service at +1-800-932-0760 (Domestic &amp; International). Hours of Operation 8:00 am EST to 6:00 pm EST. |  | 243685 | 243,685 units | Worldwide distribution: USA (nationwide) including states of: AR, CA, FL,GA, IL, LA, MA, MI, MO, NC, NH, NY, OH, OK, PA, RI, SE,TX, and WA; and countries of: AUSTRALIA, BARCELONA, GERMANY, NEW ZEALAND, NETHERLANDS, SWITZERLAND, and UNITED KINGDOM. | Terminated | June 14, 2013 | 301 | 243685 | Not_Computer
Z-0402-2013 | 63512 | No | EBA | Dental operating light. | Dental | Dental | 510(K) Exempt | Freedom Cordless LED Light System. Freedom System, Standard Length, Part Number 910863-1, Serial Numbers 77001000-77003470 and Freedom System, Long Length, Part Number 910863-2, Serial Numbers 77001000-77003470. |  |  | November 17, 2011 | 2011 | Kerr Corporation | The firm recalled the Freedom Cordless LED Headlight because some units were assembled with an incorrect component in thecontrol board electronics that may cause some Freedom Cordless LED Headlight units to become warm to the touch; however, no skin damage will occur. | Labeling mix-ups | Orascoptic sent an "URGENT: MEDICAL DEVICE RECALL" letter dated November 17, 2011. The letter identified the product, problem, and actions to be taken by the customers. An Acknowledgement Form was included for customers to complete and return to the firm. Contact Orascoptic Customer Care at 608-831-2555 for assistance regarding this notice. |  | 1330 | 1330 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, and Europe. | Terminated | November 20, 2012 | 369 | 1330 | Not_Computer
Z-0410-2013 | 63122 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | MultipolarÂż Bipolar Cup Liner 22 mm LD.; |  |  | August 31, 2012 | 2012 | Zimmer, Inc. | Zimmer Inc. is initiating a removal of the MultipolarÂż Bipolar Cup Liner 22 mm J.D. (item 00-5001-040-22) due to reports that the split poly ring was incorrectly assembled. | Process design | Zimmer sent an Urgent Medical Device Recall letter to all affected consignees during the week of September 3, 2012. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately, quarantine the recalled product and return remaining inventory. Customers were advised to complete Inventory Return Certification form attached. |  | 195 | 195 | Worldwide Distribution - USA nationwide including the states of: AK, CA, IL, FL, HI, MI, MO, NC, NJ, NY, OH, PA, TX and the countries of Germany, Brazil, Canada, Japan, Taiwan, Thailand, India and Korea. | Terminated | September 23, 2013 | 388 | 195 | Not_Computer
Z-0411-2013 | 63629 | No | ETD | Tympanostomy tube. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | PARPARELLA-TYPE VENT TUBE, 1.02 MM, SILICONE, REF 240044, QTY 6, Rx Only, STERILE EO, GYRUS ACMI INC., |  |  | October 17, 2012 | 2012 | Gyrus Acmi, Incorporated | One lot of tympanostomy tubes were manufactured with silicon that was found to be contaminated with ethylene vinyl acetate copolymer. | Material/Component Contamination | The firm, Olympus, sent an "URGENT: Medical Device Recall" letter dated October 15, 2012 to their consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect their stock to confirm if any of the affected lot of product remains in their possession; immediately cease any further use of any affected product; remove it from their stockroom and quarantine it until it is shipped back to Gyrus ACMI, Inc.; call their Gyrus ACMI customer service representative (800-773-4301) to obtain a Returned Goods Authorization and promptly return any affected product to Gyrus ACMI, Inc., 2925 Appling Rd, Bartlett, TN 38133., and complete and return the RECALL REPLY FORM to Attn: Regulatory Affairs via fax to: (508) 804-2624, even if there is no affected stock left at your facility.If you have any questions, contact the customer service at 800-773-4301. |  | 600 | 10 boxes/60 units | Worldwide distribution: USA state of Kansas and country of Korea. | Terminated | April 09, 2013 | 174 | 600 | Not_Computer
Z-0408-2013 | 62835 | No | KLS | Theophylline test system. | Toxicology | Toxicology | 510(k) | Dimension Vista(R) Systems Vista(R) Theophylline THEO Flex(R) Regent Cartridge (K4071) |  |  | August 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has received complaints from customers of "Abnormal Reaction" errors occurring on calibrations, QC and patient samples. The Abnormal Reaction flag is being triggered by an internal check in the method for atypical particle reagent aggregation. | Nonconforming Material/Component | Siemens Healthcare Diagnostic sent a Urgent Field Safety notice dated July, 20. 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please immediately discontinue use and discard any remaining inventory of Dimension VistaÂż Theophylline FlexÂż reagent cartridge (THEO, Catalog # K4071) lot 12135AE and/or 12171BB.Please complete the attached form and fax it to (302)631-8467 to indicate that you have received this information and to facilitate your additional no charge replacement.We ask that you forward this notification to anyone to whom you may have distributed this lot.If you have technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. For questions concerning the status of your no charge replacement order please contact your Siemens Customer Service Center at 888-588-3916.We apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support. |  | 959 | 959 cartons | USA (nationwide) including the states of AL, AZ, CA, CO, DC, DE, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, MS, MT, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, VA, VT, WA and WI. | Terminated | January 08, 2014 | 517 | 959 | Not_Computer
Z-0422-2013 | 63580 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | KING LAD Standard ClearSeal SINGLE USE, King Systems ETO Sterilie10 devices/case |  |  | November 02, 2012 | 2012 | King Systems Corp. | The firm received two complaints from a user facility in New Jersey regarding color-coded EtO sterilization indicators on the primary device packaging that did not undergo the expected color change. The complainant reported devices that had pink color indicators when the product packaging specified that the color should change from pink to yellow (or brown) if the product had subjected to an EtO | Nonconforming Material/Component | King Systems sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated October 31, 2012 to all affected consignees via email. The letter identified affected product, problem and actions to be taken. The letter states that further use and/or distribution of any remaining product affected by this recall should cease immediately. The letter instructs customers to destroy all of the affected products and indicate the destruction by signing and returning the attached forms using the enclosed self-addressed envelope. For questions contact your King Systems customer service representative at 800-642-5464. |  | 263 | 263 cases of product. (2,630 devices) | US Nationwide Distribution - WI, TX, NY, MO, IN, KY, FL, NJ, PA, TN, and NC | Terminated | May 28, 2013 | 207 | 263 | Not_Computer
Z-0412-2013 | 63144 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Spine Frontier Inspan compressors |  |  | August 28, 2012 | 2012 | SpineFrontier, Inc. | Inspan Compressors may break at weld causing a loss of compression | Device Design | Spine Frontier issued e-mail notices to customers on May 14, 2012 advising of the potential for breakage. Units were replaced with weld free design. |  | 17 | 17 units | Worldwide Distribution-USA (nationwide) including PR and the states of FL, MA, MO, OK, and TX, and the country of Mexico. | Terminated | June 11, 2015 | 1017 | 17 | Not_Computer
Z-0413-2013 | 63633 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | CapLOX II Final Torque Driver, Indicator is a Pedicle screw spinal system. |  |  | September 20, 2012 | 2012 | Captiva Spine, Inc | Captiva Spine recalled their CapLOX II Final Torque Driver, Indicator due to the device prematurely breaking | Nonconforming Material/Component | Captiva Spine contacted the customer by phone to have them return the device.For questions regarding this recall call 561-277-9480. |  | 10 | 10 | Nationwide Distribution (Distributed to one customer located in MN) | Terminated | January 11, 2013 | 113 | 10 | Not_Computer
Z-0416-2013 | 63541 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Trilogy Acetabular System Shell with Cluster Holes Porous 56 mm O.D. Sterile |  |  | September 20, 2012 | 2012 | Zimmer, Inc. | Zimmer TrilogyÂż Acetabular System Shell with Cluster Holes, Porous,56 MM O.D., sterile; is missing polar hole threads. Zimmer received one complaint in which no injury was reported. | Process control | Zimmer sent a Urgent Medical Device Recall Notice dated September 20, 2012, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customer. Customers were instructed to locate quarantine and return affected product to Zimmer Product Service Department, 1777 West center St. Warsaw, IN 46580 for product in the USA. Customers were asked to complete the Inventory Return Certification Form and follow the instructions where to return this form.For further questions please call (574) 372-4807. |  | 95 | 24 USA; 71 worldwide | Worldwide Distribution--USA (nationwide) including the states of AL, CA, CT, KS, MA, MD, MN, MO, NY, SD, TX and WI and the countries of Bulgaria,Canada, Germany, Italy, Netherlands, Russian Fed, Sweden and UK | Terminated | April 18, 2014 | 575 | 95 | Not_Computer
Z-0428-2013 | 63500 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | IMRIS ORT100 (Part Number 110470-000) and ORT200 (Part Number 112216-00), are tables used in both the Neuro II-SE and Neuro III-SV Magnetic Resonance Imaging Systems. |  |  | October 19, 2012 | 2012 | Innovative Magnetic Resonance Imaging Systems Inc. | IMRIS Inc. initiated a recall of their OR Tables ORT100 and ORT200. When proper procedure is not followed, the foot end of the IMRIS intra-operative OR table may be attracted to the MRI magnet as the magnet approaches the table. This poses a potential risk of injury to both the patient and operators standing near the table. | Labeling design | Consignees were sent on 10/19/12 a IMRIS "Urgent Medical Device Correction" letter dated October 16. 2012. The letter described the product and the problem involved in the recall. It also provided changes to the Operator's Manual and labeling which addresses the issue. Customers may contact 1-866-475-0525. |  | 23 | 23 | Nationwide Distribution, including the states of FL, GA, IN, MD,MA, MI, MN, MO, TN, UT, and VA. | Terminated | December 11, 2012 | 53 | 23 | Not_Computer
Z-0429-2013 | 63539 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Minotrol 16, Model Numbers: 2043748, 2043752, 2043756, Lot Number: MX090(Levels: Low, Normal, and High). |  |  | September 14, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical Irvine Technical Support received reports from some customers that have experienced high recover values for Hemoglobin when using Minotrol 16 lot #MX090 (Level High). Customer communication was sent 10/01/2010. | Nonconforming Material/Component | HORIBA Medical sent a Product Information Notice letter dated November 5, 2010 via Fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the instructions included with the Product Information Notice. For questions customers were instructed to call the Technical Support Hotline at 1-888-903-5001.For questions regarding this recall call 949-453-0500, ext 208. |  | 22954 | 22,954 | Nationwide Distribution | Terminated | April 22, 2013 | 220 | 22954 | Not_Computer
Z-0414-2013 | 63516 | No | FSP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The brand name of the device is Laser Loupes, a dentalloupe |  |  | April 10, 2012 | 2012 | Kerr Corporation | The firm initiated the recall for Laser Loupes because the ink used to mark which laser frequency each laser Loupe protectsagainst may wipe off. The loupes work according to specification, but the marking may not permanently adhere tothe device. | Finished device change control | Customers were notified via USPS 1st class mail on April 10, 2012 to US and Canadian, and Rest of World (ROW) consignees, and on April 27, 2012 to European consignees. The notification letter identifies the product, reason for recall, and instructions to customers:"Please review the table below to verify that the laser loupes being used in your office has the correct filter color. ORASCOPTIC KINDLY REQUESTS YOUR COOPERATION IN COMPLETING AND FAXING BACK THE ENClOSED ACKNOWLEDGEMENT FORM IN ORDER TO CONFIRM YOUR RECEIPT OF THIS FIELD CORRECTION.If you are an authorized Orascoptic distributor, please inform your affected customers of this notification within forty-eight (48) hours of receipt.". Customer contact information was also provided:"Please contact Orascoptic Customer Care at 608.831.2555 or toll free in the United States and Canada at 1-800369-3698 for more information.". |  | 639 | 639 units | Worldwide Distribution--USA (nationwide) and the countries of Canada and EU. | Terminated | November 21, 2012 | 225 | 639 | Not_Computer
Z-0415-2013 | 63613 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Ranawat/Burnstein Acetabular Shell 54mm/Limited Holes/Porous Coated is a metallic, full-hemisphere, acetabular shell component. |  |  | October 04, 2012 | 2012 | Biomet, Inc. | The color code on label conflicts with the liner size (liner size 24 should be green, not yellow) | Labeling Change Control | BIOMET called distributors who received the affected product to return it to Biomet with follow-up Urgent Recall Notices sent via E-mail October 22, 2012.. The notice included the identity of the recalled device, reason for the recall, and directions to complete the FAX back form and return it to 574-372-1683. Product is to be returned to Return Goods, Recalls, Biomet, Inc.56 East Bell Drive, Warsaw, IN 46580. Questions related to the recall should be directed to Field Action Coordinator (574) 372-1570, Monday through Friday. 8 a.m. to 5 p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 16 | 16 (5 USA) | Nationwide Distribution including CA, IN, NC, OH, and UT. | Terminated | March 18, 2013 | 165 | 16 | Not_Computer
Z-0417-2013 | 63075 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | S-Lift Implant Guides (also known as S-LIF) |  |  | August 22, 2012 | 2012 | SpineFrontier, Inc. | When using the optional S-LIFT (also known as S-LIF) Implant Guides without stops, there is a known risk of moving the Implant Guides, and possibly the implant, too far into the disc space creating the potential for injury to the patient. The instructions for Use for the S-Lift Implant Guides will be amended to add the following warning: "Care is to be taken to not over insert the Implant Guides | Labeling design | Spine Frontier issued a Correction Notice Letter August 22, 2012 to all affected customers to update the Instructions for Use. The letter identified the affected product, problem and actions to be taken. Customers were Instructed to remove the Instructions for Use they currently have and replace them with the update. The letter instructs customers to complete and return the Response Form to 978-232-3990. |  | 19 | 19 units | US Nationwide including the states of: MA and TX. | Terminated | June 11, 2015 | 1023 | 19 | Not_Computer
Z-0430-2013 | 63479 | [6 empty] | 0.9% Sodium Chloride Injections, USP, 1000ml, VWR/Baxter |  |  | October 17, 2012 | 2012 | Biomerieux Inc | The firm inventoried and distributed 0.9% Sodium Chloride Saline bags as 0.45% Sodium Chloride Saline bags for use with the VITEK 2 system. | Employee error | Biomerieux sent an Urgent Product Removal Notice letter dated October 16, 2012 to all affected customers via certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected product and immediately discard any remaining inventory. The letter included instructions on receiving product replacement or credit. For questions or concerns contact your local BioMerieux representative. |  | 199 | 199 cases/14-1000mL bags | Worldwide Distribution - USA (nationwide) including the states of: AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, PA, RI, TN, TX, VA and WI. Distribution was also made to Canada, military facilities and other foriegn countries. | Terminated | December 05, 2013 | 414 | 199 | Not_Computer
Z-0432-2013 | 63554 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 4.5mm PLEO Osteotomy Guide. |  |  | October 18, 2012 | 2012 | OrthoPediatrics Corp | Complaint received from sales representative that the drill guide tubes would not screw into replenishment instrument. Upon investigation, it has been determined that the device was not manufactured to design specifications and will not function as intended. | Nonconforming Material/Component | OrthoPediatrics sent an "URGENT RECALL" e-mail of the Trochar Scalpel Instrument to all affected customers. The e-mail identifies the product, problem, and actions to be taken by the customers. The e-mail states that the instrument has been manufactured incorrectly and will not work (function as intended). The threads in the instrument are faulty and the guide tubes will not thread in correctly. Customers will receive follow up information. |  | 24 | 24 | Nationwide Distribution-including the states of OH, NC, CO, FL, IN, MI, and TX. | Terminated | September 18, 2013 | 335 | 24 | Not_Computer
Z-0434-2013 | 63631 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | SenTec V-Sign Membrane Changer Insert (VS-MCI), 5 pieces. The SenTec V-Sign Membrane Changer Insert is indicated for use with the V-Sign Membrane Changer only. |  |  | November 05, 2012 | 2012 | Sentec AG | A manufacturing defect was observed for the membrane changer insert that does not allow electrolyte to be present under the new membrane. The PCO2 performance of a sensor without electrolyte will be deteriorated. | Component design/selection | SenTec notified the one US distributor via e-mail on November 5, 2012. The distributor was advised to stop distribution of affected lot numbers. The communication to the distributor described the issue, identified affected product, and stated that SenTec will replace current stock at the end-customer level. The end user can discard the affected products. The distributor is asked to respond back with an e-mail to SenTec that they took appropriate action. |  | 265 | 265 boxes (5 per box) | Distributed only in MO. | Terminated | December 18, 2012 | 43 | 460 | Not_Computer
Z-0431-2013 | 62834 | No | LMF | Absorbable hemostatic agent and dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | PMA | Styker Orthobiologics Vita Plasma Separator, a component of Vitagel Surgical Hemostat. |  |  | July 26, 2012 | 2012 | Stryker Orthobiologics, Inc. | When preparing plasma using VitaPrep, it was reported that a small amount of blood collected in the bottom of the centrifuge cup following the centrifugation step. Firm has since discovered that if blood is present in the Luer connection of the VitaPrep following the blood draw step, this blood could be displaced to the outside of the syringe when the VitaPrep cap is reattached to the syringe. T | Device Design | Stryker sent a "MEDICAL DEVICE FIELD CORRECTION" dated July 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter included a Business Reply Form for customers to complete and return via fax to 1-484-323-8803. Contact the firm at 1-888-774-8870 for questions regarding this notice. |  | 3268 | 3,268 | Nationwide Distribution-including the states of AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV | Terminated | October 29, 2014 | 825 | 3268 | Not_Computer
Z-0437-2013 | 63544 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen CompleteKnee Solution Stemmed Nonaugmentable Tibial component |  |  | September 21, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the NexGenÂż Stemmed Nonaugmentable Tibial Component due to the devices being processed through a manufacturing cleaning operation that was operating outside of the validated parameters.As a result, the devices may contain residual particulate from the manufacturing process.Risks include: Allergic reaction due to the foreign material or late onset | Process change control | Zimmer sent an Urgent Medical Device Recall letter dated September 21, 2012, to all affected customers. The letter identifiedthe product, the problem, and the action to be taken by the customer. Customers were instructed to locate and quarantine the affected product. Customers were to return the affected product along with the completed Inventory Return Certificate Form. If product was further distributed the Recall Letter should be forwarded to their customers. Product was to be returned to USA accounts - Zimmer Product Service Department, 1777 West Center St. Warsaw, IN 46580. If product has been implanted, affected surgeons will be notified. For questions regarding this recall call 574-372-4807. |  | 123 | 123 worldwide (33 USA) | Worldwide Distribution - USA including AZ,CA, FL, IA, IL, IN, MO, NE, NJ, NY, OH, PA, WA. Internationally to Thailand, Canada, Germany, France, UK, Hungary, Netherlands, Russian Fed, and Sweden, | Terminated | April 17, 2014 | 573 | 123 | Not_Computer
Z-0440-2013 | 63681 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical Advincula Arch , Reusable Uterine Manipulator Handle |  |  | October 31, 2012 | 2012 | Cooper Surgical, Inc. | Excessive pressure exerted on the uterine tip when attached to the Arch may, under certain circumstances, cause the internal steel rod of the tip to protrude through the cleaning hole on the shaft of the handle | Device Design | CooperSurgical issued notification on October 31, 2012, to all affected customers Fed X . The notification identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return units for replacement. Customers with questions should contact Customer Serivce at 1-203-601-9818 or 800-243-2974. Customers were asked to fill out the attached Replacement of Recalled Product Form and fax it to 1-800-262-0105.For questions regarding this recall call 203-601-9818. |  | 1215 | 1215 total | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Czech Republic, Germany, Italy, Spain, France, United Kingdom, Japan, New Zealand, South Korea, Saudi Arabia, Sweden, Singapore, Slovenia, and Taiwan. | Terminated | March 04, 2014 | 489 | 1215 | Not_Computer
Z-0438-2013 | 62918 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | IO FiX X-Post Reamer, Extremity Medical, www.extremitymedical.com; Customer Service: 888.499.0079 |  |  | April 25, 2011 | 2011 | Extremity Medical LLC | Product Event Reports revealed a broken reamer experienced intra-operatively. Pre-drilling was not performed prior to reaming. | Employee error | Extremity Medical sent a notification by e-mail that describes the "Slight modification to the Surgical Technique" on April 25, 2011. For questions regarading this recall call 973-588-8980. |  | 156 | 156 units | Worldwide Distribution - USA including FL, MA, NC, CA, MD, VA, IN, AL, NJ, AZ, OK, MT, NY, OH, IA, MI, TX, WI, GA, IL, TN and NE. Internationally to Switzerland. | Terminated | March 21, 2013 | 696 | 156 | Not_Computer
Z-0436-2013 | 63543 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 220/240V AC, Advanced Perfusion System Platform (APS) |  |  | November 14, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular System (TCVS) has received reports of a situation where usersexperienced a total loss of functionality for some System 1 units. The reports indicate that theunits went blank and shut down with no sign of power and battery backup did not initiate. Theresult is all pumps stop, with no safety system functionality, and the battery would not beactivated. The user would be | Device Design | Terumo Cardiovascular Systems sent a Urgent Medical Device Recall Correction letter dated November 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Terumo CVS is alerting all users of Terumo System 1 of the reports of malfunction. Customers were instructed to: 1. Review this Medical Device Safety Advisory.2. Assure that all users are aware of this notice.3. Confirm receipt of this communication by faxing, or emailing the attached Customer Response Form to the fax number/email address indicated on the form.We encourage you to contact us with any questions or concerns:Terumo CVS Customer Service 1-800-521-2818, Recall Fax 1-734-741-6149Customer Service Hours: Monday thru Friday, 8 AM - 6 PM ET |  | 1647 | 1647 units total | Worldwide Distribution--USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | August 30, 2013 | 289 | 1647 | Not_Computer
Z-0444-2013 | 63583 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Pentron Clinical Metrex Research; Pinnacle Cure Sleeve, a protective barrier |  |  | September 20, 2011 | 2011 | Metrex Research, LLC. | The firm initiated the recall due to a labeling error. The boxes labeled as Pinnacle Cure Sleeve 4511, which fits 11mm light guides, actually contain Pinnacle Cure Sleeves that fit 7mm or 8mm light guides (part number 4500). | Error in labeling | The firm, Sybron Dental Specialties, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated September 20, 2011 to its consignees/customers. The consignees/customers were instructed to determine if they have any of the affected product in their inventory; if they have the affected product to contact Metrex Research Customer Care at 1-800-841-1428 to receive an RMA number; return any affected product in their inventory; and complete and return the Recall Response Form via mail to: Metrex Research, 29210 Wick Road, Romulus, Michigan 48174, Attn: QA Lab or fax to: 734-947-6725.If you have any questions regarding this recall, you can contact our Customer Care Center at 1-800-841-1428. |  | 3348 | 3,348 units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, and Europe. | Terminated | November 29, 2012 | 436 | 3348 | Not_Computer
Z-0442-2013 | 63711 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | Echo Bi-Metric Press-Fit Stems/Echo Bi-Metric 11mm x 135mm Collarless Porous Stem (Orthopedic Implant) |  |  | November 06, 2012 | 2012 | Biomet, Inc. | It was reported that the surgeon noticed that the Echo 11mm hip stem had an etching indicating it was 13mm stem. | Labeling Change Control | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated November 8, 2012 to its customers. The notice describes the product, problem and actions to be taken. The customers were instructed to take the following actions:Âż Immediately locate and remove the identified device(s) listed below from circulation.Âż Carefully follow the instructions on the enclosed FAX Back Response Form .Âż Fax a copy of the Response Form to 574-372-1683 prior to return of product.Âż Use priority carrier for your shipment.Questions related to notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 4 | 4 | Nationwide distribution: USA including states of: AZ, MD, and KS. | Terminated | August 28, 2013 | 295 | 4 | Not_Computer
Z-0446-2013 | 63675 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The brand name is MRidium Series 1000 MR Infusion Sets, Type 1058 MR IV Extension Set, Lot Code LBG001. The generic name is infusion administration set. |  |  | August 24, 2012 | 2012 | Iradimed Corporation | Iradimed Corporation is recalling MRidium Series 1000 MR Infusion Sets, Type 1058 MR IV Extension Set, Lot Code LBG001 because a section of the device does not meet specifications. | Nonconforming Material/Component | Iradimed Corporation contacted all affected customers directly by phone, and followed up with an email notification. Customers were instructed to return the product to Iradimed Corporation. For questions regarding this recall call 407-677-8022. |  | 719 | 719 sets | Nationwide Distribution including AZ, CA, FL, KY, MD, MA, NJ, NY, PA, and TX. | Terminated | July 19, 2013 | 329 | 719 | Not_Computer
Z-0449-2013 | 63212 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima XR 220amx 15kw Mobile X-Ray System |  |  | August 01, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the mobile radiographic product, Optima XR22Oamx and Optima XR200amx with Digital Upgrade. | Radiation Control for Health and Safety Act | GE Healthcare will without charge remedy the defect or bring the product into compliance with each applicable Federal and IEC standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication through a GE Healthcare field engineer site visit. Field Modification Instructions (FMI) 10868 describes the rework plan the GEHC Engineers will make to the imaging system compliant with the applicable performance standard.This CAP appears to adequately address the problem and is hereby approved.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 217 | 217 total Installed in the US | Nationwide distribution | Terminated | January 10, 2013 | 162 | 217 | Computer
Z-0451-2013 | 63649 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Kimberly-Clark Patient Warming System - Model 1000 Control Unit |  |  | September 16, 2010 | 2010 | Kimberly-Clark Corporation | Error 85 alarm on the Model 1000 Control Unit was occasionally triggered inappropriately. | Component design/selection | Kimberly Clark sent an Urgent Field Corrective Action letter dated September 16, 2010, to all affected customers. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were asked to confirm whether their facility as the affected product by faxing the completed Field Corrective Action Response Form to 920-380-6682. Kimberly-Clark will arrange to perform the upgrade at the customer's facility. Customers with questions should call the Field Corrective Action Coordinator at 770-587-7400.For questions regarding this recall call 770-587-8393. |  | 400 | 400 units | Nationwide Distribution including AZ, AK, CA, CO, DE, DC, FL, GA. ID, IL, IN, IA, KS, KY, LA, MD, ME MA, MS, MO, MT, NE, NJ, NM, NV, NY, NC, OH, OR, PA, RI, SC, TN, TX, VA, WV, and WI | Terminated | December 20, 2012 | 826 | 400 | Not_Computer
Z-0441-2013 | 63722 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The Vertex Max Drill Tube is a component within the Navigated Vertex Max Drill Guide Set (item 9733370). The Navigated Vertex Max Drill Guide Set is intended to be used with a Medtronic Computer Assisted Surgery System to help assist with screw placement during spine surgical procedures. The device is reusable and requires sterilization according to product labeling. |  |  | November 13, 2012 | 2012 | Medtronic Navigation, Inc. | Medtronic is recalling certain Vertex Max Drill Tubes which are used with the StealthStation System due to a defect related to the internal threading for component set up. | Process control | Medtronic Navigation sent a Urgent Medical Device Recall letter dated November 13 -14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed do not use the Vertex Max Cannulated Drill Tube clinically. Medtronic Navigation will remove the Vertex Max Cannulated Drill Tube. Medtronic Navigation will provide a replacement Vertex.Max Cannulated Drill Tube manufactured from a different lot which was inspected to verify the defect has been corrected.If you have any questions about this notification, please contact a Sales Representative in your area. For further questions please call (720) 890-3200 Ext. 3303 . |  | 12 | 12 Units | Worldwide Distribution--US (nationwide) including the states of CA, FL, MA, LA, and MI and the countries of Switzerland, Saudi Arabia, Lebanon, Canada and Korea. | Terminated | January 04, 2013 | 52 | 12 | Not_Computer
Z-0447-2013 | 61835 | No | MXM |  |  | Neurology | PMA | The product has the Catalogue/Part number 60010. |  |  | May 09, 2012 | 2012 | Natus Medical Incorporated | The Olympic Cool-Cap's Control Module has experienced a frozen screen during use. When this occurs, the on-screen information remains on display, but the system is no longer providing cooling treatment to the infants. The visible clock displayed at the upper right hand corner of the screen stops advancing. | Nonconforming Material/Component | The firm, Natus Medical Incorporated, sent its customers an "URGENT: MEDICAL DEVICE" letter dated May 16, 2012 with an attached Field Safety Notice. The letter describes the product, problem and actions taken. Customers can call Natus Technical Service at 888-496-2887 for any questions about this recall. |  | 324 | 324 units | Worldwide distribution: USA (nationwide) including states of: AK, AR, AZ, CA, CO, DE, FL, GA, HI, ID, IL, IN, KY, MA, MI, MN, MS, ND, NE, NJ, NV, NY, OH, OK, PA, SC, SD, TN, TX, UT, WA, Puerto Rico, and countries of: Argentina, Australia, Austria, Bahrain, Canada, Chile, China, Colombia, Ecuador, Egypt, Germany, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Italy, Kuwait, Latvia, Malaysia, Mexico, Morocco, Panama, Puerto Rico, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | January 27, 2015 | 993 | 324 | Not_Computer
Z-0443-2013 | 63235 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Devices includes a Multi Port Inflator or without Multi Port Inflator (MPI). All the devices are vascular blood flow machines. Devices are distributed with pressure cuffs of varying sizes as accessories. |  |  | August 08, 2012 | 2012 | Parks Medical Electronics, Inc | Parks Medical Electronics, Inc. was made aware of two incidents where an air supply tube from their PARKS PVR unit was inadvertently connected to an IV in the patient's arm instead of into the blood pressure cuff. | Employee error | The firm, Parks Medical Electronics, Inc., sent an "IMPORTANT SAFETY ISSUE SERIOUS HEALTH CONSEQUENCE" letter, dated September 10 2012 to their consignees/customers. The letter described the product, problem and actions to be taken. The customers are advised to order LUER UPGRADE KIT WITH MPI if the Parks device includes a Multi Port Inflator (MPI), and/or order the LUER UPGRADE KIT WITHOUT MPI the Parks device does not have a Multi Port Inflator. While waiting for the kit or its installation, consignees/customers should use extreme caution when hooking up the air supply tubes to the cuff's Luer fittings. For ordering your kit, please call 800-547-6427 option 0.For responses, questions or concerns, customer should contact Parks Medical Electronics at 888-356-9522 (Phone), 503-591-9753 (Fax) or send email to support@parksmed.com. |  | 349 | 349 units (Domestic: 270 units - Foreign: 79 units) | Worldwide distribution: USA (nationwide) and countries of: Australia, Brazil, Canada, Chile, Egypt, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Korea, Mexico, Peru, Republica Argentina, Spain, Switzerland, and Venezuela. | Open, Classified |  |  | 349 | Not_Computer
Z-0445-2013 | 63410 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | The Is-anti-Sm test kits****Catalog No: 720-240*** |  |  | August 06, 2012 | 2012 | Diamedix Corporation | Diamedix Corporation recalled their Is-anti-Sm Test Kit due to the Positive Control recovering high out of specification. | Nonconforming Material/Component | Diamedix sent a Customer Notification letter to all affected customers via Certified Mail. The letter identifed the affected product. problem and actions to be taken. The letter instructed customers to contact Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any remaining affected product. For questions call 305-324-2314. |  | 65 | 65 Test Kits | Worldwide Distribution - USA Nationwide including the states: CA, ID, NJ, NY, OH, PA, PR, RI, TN, and TX, FL, and the countries of Italy, Chile and Ecuador | Terminated | August 08, 2014 | 732 | 65 | Not_Computer
Z-0454-2013 | 63634 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | NeptuneÂż 2 Rover - ULTRA ~120V ~ 60 Hz 12A |  |  | May 24, 2011 | 2011 | Stryker Instruments Div. of Stryker Corporation | Beginning May 24, 2011 Stryker Neptune 2 (N2) Waste Management Systems castors were replaced with a new castor design because investigation found the old design had the potential for the casters to bind compromising the Rover's ability to roll. This can result in additional strain on the operator, due to multiple exposures resulting in back pain. | Component design/selection | Stryker Instruments sent a "STRYKER NEPTUNE 2 ULTRA 120V &amp; 230V WASTE MANAGEMENT SYSTEM" letter dated June 22, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return a Reply Form via fax to Stryker at 269-906-9121 or email to sales.service@stryker.com. Technical questions should be addressed to Neptune Field Service at 800-800-4236 ext. 5052. |  | 5865 | 5865 | Worldwide Distribution-USA (nationwide) and Pacific countries, and the countries of EEMEA, Canada, Latin America, and Japan. | Terminated | May 28, 2014 | 1100 | 6019 | Not_Computer
Z-0464-2013 | 62749 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Total Hip Pack, K12T-05537, Merit's Total Knee Pack, K12T05538 |  |  | June 08, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural trays contain Stryker Hytrel Togas which are being recalled because the clear tape that is intended to aid in the prevention of patient's fluids from contacting the surgeon/surgical staff through the sleeve seam of the toga, was not applied during the manufacturing process. | Process control | The firm, Merit Medical, sent an "URGENT PRODUCT RECALL NOTICE" letter dated June 8, 2012 to its customers by letter sent by US Postal Service - Certified Mail Receipt describing the problem, instructions as to actions that need to be taken, and requesting the customer to complete and sign the Customer Response Form and immediately fax a copy to this form to 1-804-416-1031. Customers were instructed to mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. Should the customer choose to return the product for rework, they are to return the affected products by shipping them back to Merit Medical Systems via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department RGA #NC-53417, Merit Medical Systems, Inc., Attn: Manager of Quality Assurance and Regulatory Affairs, 12701 Kingston Ave., Chester, VA 23836. For questions customers were instructed to contact Manager of Quality Assurance and Regulatory Affairs at 1-804-416-1048 or email: jbraxton@merit.com. |  | 83 | 83 packs | Nationwide distribution: US including states of: NE and SD. | Terminated | December 10, 2012 | 185 | 83 | Not_Computer
Z-0459-2013 | 63407 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | 18 Fr 6mm Sft Flow Angled Wire |  |  | November 21, 2012 | 2012 | Terumo Cardiovascular Systems Corporation | The presence of plastic flash was identified at the tip of certain lots of Sarns Soft-Flow Aortic Cannula. The plastic flash has the potential to detach and contribute to an adverse patient effect. While this potential failure has not been reported from the field as an observed defect, it was determined that the product does not meet specifications and this condition could result in an unacceptabl | Process design | TERUMO sent an URGENT MEDICAL DEVICE RECALL letter dated September 20, 2012, to all affected customers. The letter identified the product, the problem, and the aciton to be taken by the customer. Customers were instructed to remove this cannulae from circulation and return them for credit. Customers with questions should call 1-800-521-2818.For questions regarding this recall call 734-741-6173. |  | 5760 | 5760 total | Worldwide Distribution - USA including AL , CA , CO , DC , DE , FL , GA , IL , IN , KS , KY , LA , MA , MI , MO , MS , MT , NC , ND , NE , NH , NY , OH , PA , TN , TX , VA , and WI. Internationally to Australia, United Arab Emirates (UAE), Singapore, Malaysia, BELGIUM, Japan, and CANADA. | Terminated | April 23, 2013 | 153 | 5760 | Not_Computer
Z-0463-2013 | 63615 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Sekisui Diagnostics, Spectrolyse PAI-1 In-Vitro Diagnostic Use for the quantitative determination of Plasminogen Activator Inhibitor Type 1 in human plasma. |  |  | October 24, 2012 | 2012 | Sekisui Diagnostics Llc | Spectrolyse PAI-1 activity assay, producing lower than expected absorbance values 0 IU/ml standard reducing the slope of the standard curve requiring more frequent sample dilutions. | Nonconforming Material/Component | Sekisui Diagnostics issued an "URGENT PRODUCT RECALL NOTICE" dated October 24, 2012 to all consignees. The letter instructed consignees to immediately stop using the specified lot of product, discard any remaining inventory and return the enclosed response form. Contact the Customer Service Department at 800-637-3375, ext. 10, 12, or 31 for assistance with this recall. |  | 63 | 63 kits | Nationwide Distribution-including the states of FL, IN, NC, OH, UT, and WA. | Terminated | January 08, 2013 | 76 | 63 | Not_Computer
Z-0465-2013 | 63521 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza Radiological Image Processing System |  |  | June 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Firm has become aware of an unintended behavior when using syngo.plaza and issued an Update Instruction. The Update Instruction SY050/12/S provides a software update and addresses the following issues: - Patient Rename messages for the same patient are being continually processed for several days. This fills up the 'pcvdbsrv' log and potentially affects performance of the system. - Data could | Software design | The firm, Siemens, sent "Update Instruction SY050/12/S" on June 20, 2012 to its consignees/customers. The Update Instructions described the product, problem and actions to be taken. It also provides a software update.If you have any questions call 610-219-4834 or email: anastasia.mason@siemens.com. |  | 15 | 15 | Nationwide distribution: USA including states of: CA, FL, GA, MA, MN, MO, NJ, NY, OH, and WI. | Terminated | January 14, 2014 | 573 | 15 | Computer
Z-0453-2013 | 63642 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | OmniGuide BP-ROBOTIC ( BeamPath Robotic), Single Use |  |  | November 05, 2012 | 2012 | OmniGuide, Inc. | Failure of glue joint, allowed a portion of the single use device to shift inside, may separate and fall into the patient | Process control | OmniGuide issued letter notifiction on November 5, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that an OmniGuide sales representative would contact them immediately to recover the affected product. For any questions customers were instructed to call 617-551-8561. For questions regarding this recall call 617-551-8410. |  | 45 | 45 units | Nationwide Distribution including AZ, FL, MA, NY, OH, and OR. | Terminated | October 16, 2014 | 710 | 45 | Not_Computer
Z-0452-2013 | 63429 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) IRON Flex(R) reagent cartridge (K3085) |  |  | October 08, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed the potential for falsely elevated results with IRON Flex(R) reagent cartridges when tests are processed from a well set contaminated with iron from environmental sources. All lots of IRON Flex(R) reagent cartridges have the potential for this issue. If a well set is contaminated with trace amounts of iron, the entire well set is impacted, and the mag | Material/Component Contamination | The firm intiated their recall of the specified lots on October 8, 2012 by issuing a letter to all consignees informing them of the problems with the reagent cartridges. Firm recommended batching the processing of the tests on one instrument and provided instructions to disable normal flow of IRON and to run the QC on each IRON Flex(R) reagent cartridge well set. In addition, a a reply form and Frequently Asked Questions attachment were enclosed with the communication.For questions regarding this recall call 302-631-6311. |  | 19310 | 19,310 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, Washington DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Internationally to Canada, AU, NZ, JP, DE, SA, and MY | Terminated | January 14, 2014 | 463 | 19310 | Not_Computer
Z-0470-2013 | 63442 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics |  |  | July 27, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Offset Bushings associated with the lots identified were mismarked: the rotational reference numbers on the Tibial Offset Bushing were marked counterclockwise instead of clockwise; the Femoral Offset Bushing did not include rotational markings. | Error in labeling | Stryker Orthopaedics sent an "URGENT PRODUCT RECALL" letter dated July 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Product Recall Acknowledgement Form was attached for customers to complete and return via fax to 201-831-6069. Contact the firm at 201-972-2100 for assistance with this recall. |  | 8 | 8 | Worldwide Distribution. | Terminated | September 30, 2015 | 1160 | 15 | Not_Computer
Z-0472-2013 | 63612 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | OSS Reamer Sleeve REF132-472662 KNEE REAMER SLEEVE 13.0 thru 20.0 STAINLESS STEEL |  |  | November 01, 2012 | 2012 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the thin shaft reamer may get stuck and not slide into the sleeve. | Nonconforming Material/Component | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated November 1, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately locate, discontinue use of and remove the product from circulation; and return recalled devices with directions to distribute notices, if product was further distributed; and to complete and return the enclosed FAX Back Response Form within three (3) business days to FAX # 574-372-1683. Distributors were charged with locating and returning products to Biomet ATTN: Return Goods-Building B RECALLS RGA#: _, Biomet, Inc., 56 East Bell Drive, Warsaw, IN 46580. Questions should be directed to the Field Action Coordinator at (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m |  | 36 | 36 | Worldwide distribution: USA (nationwide) and countries of: Japan and Netherlands. | Terminated | April 29, 2013 | 179 | 82 | Not_Computer
Z-0467-2013 | 63646 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | BiOM Rechargeable Li Ion 22.2 V Battery and BiOM Lithium Battery Charger sold as an accessory to the BiOM lower leg system: |  |  | October 26, 2012 | 2012 | iwalk inc | Battery may over-heat during charging and smoke | Unknown/Undetermined by firm | iWalk notified customers directly by sales representatives by telephone on October 26, 2012, and instructedthem to stop use of the battery and charger and return them to iWalk. Accounts provided with Return Material Authorizations (RMA).For questions regarding this recall call 781-761-1560. |  | 30 | 21 Batteries; 9 Chargers | Nationwide Distribution including OH, NC, and NY. | Terminated | September 26, 2014 | 700 | 30 | Not_Computer
Z-0468-2013 | 63643 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression MRI Patient Monitoring System. |  |  | October 15, 2012 | 2012 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling the Expression MRI Patient Monitoring System because during transport of the systems cart with a docked Display Controller Unit (DCU), a caster (wheel) may become detached. If this occurs, the cart can become unbalanced and may fall. | Nonconforming Material/Component | Philips Healthcare sent an "URGENT-MEDICAL DEVICE RECALL" letter dated October 15, 2012 to all affected customers. The letter identified the product, problem, and actons to be taken by the customers. Customers were instructed to inspect all casters of the unit to ensure that they are all secured. If a caster is loose, customers were told to lock the caster in place, limit movement of the cart and contact their local Phillips Invivo Representative. Nevertheless, a Philips Invivo representative will contact the customer regarding their affected device. All affected devices will have new casters installed in order to correct the problem. Contact your local Philips Invivo Representative at 1-800-722-9377 for further information and support. |  | 543 | 543 units | Worldwide Distribution-USA (nationwide) including the states of MO; MS; MA; IL; OH; NY; KS; PA; WI; VA; SD; AZ; FL; OR; TX; MA; MO; FL; MA; MI; CT; CA; NJ; NC; ME; IL; MA; MA; WV; MA; MN; CA; AL; AL; WI; DC; OH; WV; CA; NC; NY; MT; NJ; CA; WA; IN; MI; PA; TX; NC; TN; TX; ID; NM; IL; PA; IN; IL; WA; CO; MN; CA; IN; MD; MI; DC; GA; DE; CO; CT; WA; NJ; MD; TX; TN; MN; WA; CA; FL; OH; PA; NY; IN; OK; ID; IA; NY; WI; VA; CA; HI; MS; TX; HI; CT; TN; OR; WA; PA; NH; CA; MA; LA; IN; PA; AZ; GA; UT; ND; GA; TX; TX; IN; CA; MI; CT; CA; NC; MS; VA; IN; MI; NE; NV; MI; WA; IL; KY; LA; OR; OK; LA; TN; KY; TX; CA; PA; NM; NC; CO; DC; AL; MI; TX; SD; PA; OR; ID; NM; SD; CO; MN; NJ; MI; ND; MN; MT; LA; CT; WA; TX; IA; NJ; WI; SD; and UT, and the countries of Canada, Italy, Thailand, Russian Federation, Saudi Arabia, Slovenia, Singapore, Republic of Korea, Uzbekistan, Iceland, Israel, Austria, United Kingdom, Afganhistan, Japan, South Africa, Germany, Qatar, Netherlands, Switzerland, Denmark, and Sweden. | Terminated | February 13, 2015 | 851 | 543 | Not_Computer
Z-0466-2013 | 63553 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz 16 Multi-Sliced CT Scanner System, part number 989605858501. NeuViz Dual series CT Scanner System, part number 989605651321 |  |  | October 18, 2012 | 2012 | Philips And Neusoft Medical Systems Co., Ltd. | The four screws that secure the patient table top to the carrier pulled out and caused the table top to detach. This issue may affect NeuViz Dual and NeuViz 16 scanners. | Device Design | Philips and Neusoft Medical Systems Co., Ltd. sent an Urgent Device Correction NeuViz Dual, NeuViz 16 letter, dated Oct. 18, 2012, to all affected US customers and distributors. The letter described the product, problem, and actions to be taken by the customers. Customers are advised to:1. Have a mechanic check the patient table top screws fixing. Detailedinstructions are attached as Annex with the letter.2. If it is found that table top screws are improperly assembled, stop usingthe product and contact local Service personnel for repair.Philips and Neusoft Medical Systems Co., Ltd is in the process of upgrading devices with aforementioned kits and will contact customers to arrange for an update to affected systems. If you have any questions about this corrective action program, please contact the Service Support Department, Neusoft Medical Systems Co. Ltd. at nms-service@neusoft.comFor questions regarding this recall call 425-487-7665. |  | 320 | 20 units of NeuViz 16 CT Scanner Systems and 25 units of NeuViz Dual series | Nationwide Distribution including North Carolina, Ohio, Texas, and Puerto Rico. | Terminated | July 08, 2013 | 263 | 320 | Not_Computer
Z-0473-2013 | 63577 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas 8000 Modular Analyzer Series, Model/Catalog/Part Number: 05641489001. The Cobas 8000 is a fully automated system for clinical chemistry analysis. |  |  | November 01, 2012 | 2012 | Roche Diagnostics Operations, Inc. | On cobas c 701 and cobas c 702 modules of the Cobas 8000 series analyzer, a mismatch between the reagent probe positioning and the aspirated reagent volumes under certain conditions can lead to insufficient reagent pipetting. This may affect patient sample and controls recovery for the following tests: Alkaline Phosphatase (ALP2), Alanine Aminotransferase (ALT), Bicarbonate (CO2-L), Creatinine (C | Software design | On 11/01/2012 an "Urgent Medical Device Correction" notification was sent via UPS Ground (receipt required) to consignees. The device issue was described and recommended actions were provided. A Roche Field Application Specialist will install the new e-barcode during a scheduled visit. Customers also will have an option to install the new e-barcodes themselves, following the instructions provided in the notification. A Software update will be released when available. Customers may contact Technical Support at 1-800-428-2336 for more information. |  | 78 | 78 analyzers | Nationwide Distribution, including the states of AL, AZ, CA, FL, GA, IL, IN, KS, KY, MA, MD, MI, MN, NC, NJ, NY, OH, VA, WA, and WI. | Terminated | January 08, 2014 | 433 | 78 | Computer
Z-0475-2013 | 63619 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost Software Version R2.0.2 |  |  | October 30, 2012 | 2012 | Philips Healthcare Inc. | With Digital Diagnost R2.0.2 a mirrored "R in a circle" is burnt into the image (if mirrored) in the lower right corner of the image. When such an image is mirrored back inside the PACS system, an unmirrored "R in a circle" appears in the lower left corner. This can be mistaken for a "Right patient side" marker, although this can appear on the left patient side. | Software design | Philips Healthcare initiated a recall on October 25, 2012 via Urgent - Field Safety Notice for the Digital Diagnost X-ray System. The letter identified the affected product, problem and actions to be taken. The letter informs customers that Philips will supply a software update kit to version 2.1.4. For questions please contact your local Philips representative 1-800-722-9377. |  | 623 | 623 units | Worldwide Distribution - USA (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, India, Indonesia, Iraq, Italy, Japan, Latvia, Lithuania, Libya, Malta, Mexico, Netherlands, New Zealand, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and UK. | Open, Classified |  |  | 623 | Computer
Z-0474-2013 | 63708 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bio-Modular Humeral Head 48mm x 24mm Ti 6AI4V/Ion Implanted (Orthopedic Implant) |  |  | November 01, 2012 | 2012 | Biomet, Inc. | The items in this lot are missing the ion implantation feature. The surface of the head may be softer and more prone to scratching, which in tum could potentially cause higher wear of the poly bearing. Increased wear could potentially lead to a shorter implant life. | Process control | Biomet, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTICE dated November 1, 2012 to all affected consignees. The letter identified the affected product, problem and actions to be taken. The notification letter instructed the surgeon to montior the patient as their is a potential for increased wear which have an unknown effect on implant life. For questions call 574-372-1570. |  | 4 | 4 | USA Nationwide Distribution in the state of WI. | Terminated | August 29, 2013 | 301 | 4 | Not_Computer
Z-0480-2013 | 63495 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilette Mira, intended use as Mobile x-ray system. |  |  | October 10, 2012 | 2012 | Siemens Medical Solutions USA, Inc | There is a risk of lost images during examination. | Software design | The firm initiated their recall/correction of this product on September 9, 2012 by sending a Field Safety Notification letter with Update Instruction XP054/12/S to all consignees. The letter described the steps to be taken to avoid the potential risk and included a response form to be returned to the recalling firm. |  | 21 | 21 | Nationwide distribution: USA including states of: GA, ID, IL, IA, NE, NC, OH, SD, TX, and WI. | Terminated | January 15, 2014 | 462 | 21 | Not_Computer
Z-0484-2013 | 63312 | No | NSX |  |  | General Hospital | Contact ODE | Data Station AMPLILINK software is an instrument user interface/ data management software used with COBASÂż AmpliPrep instrument, COBASÂż TaqManÂż analyzer, COBASÂż TaqManÂż 48 analyzer, COBASÂż AMPLICORÂż analyzer, and Cobas p 630 instrument for sample preparation, amplification and detection of PCR (Polymerase Chain Reaction) molecular diagnostic analytes. |  |  | July 16, 2012 | 2012 | Roche Molecular Systems, Inc. | Roche Molecular Systems, Inc. has become aware of a device-associated diagnostic software issue. | Software design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification letter dated July 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Molecular Systems, Inc. has disabled this software on affected Roche platforms and discontinued the distribution of the affected software. Contact Roche Diagnostic Technical Support, 24 hours a day, seven days a week at phone number : 1-800-428-2336 if you have questions regarding this notification. For questions regarding this recall call 908-253-7569. |  | 884 | 884 units | Nationwide Distribution | Terminated | July 22, 2015 | 1101 | 9250 | Computer
Z-0478-2013 | 63635 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | The GlideScope Video Laryngoscope (GVL) is a video laryngoscope that incorporates a high-resolution color camera, an LED light source, a rechargeable lithium battery, and video output to a remote monitor. It consists of a video laryngoscope and a portable video monitor. |  |  | October 25, 2012 | 2012 | Verathon, Inc. | Re-usable video laryngoscope blades of the GlideScope GVL3, GVL4, and GVL5 were recalled due to potential cracking. | Device Design | The firm, Verathon, sent an "URGENT - Medical Device Recall" letter dated October 25, 2012, to their consignees/cutomers. The letter described the product, problem, and actions to be taken. The customers were instructed to return the recalled blades back to Verathon and complete and return the Device Recall Notification Reply Form via fax to (425) 883-2896 or email: cservice@verathon.com. All returned products will be quarantined and will be destroyed later. Verathon will replace the affected video laryngoscope blade at no cost to customers. Customers in Canada and US with questions can call 800-331-2313. If you are located outside of the US or Canada, please call 425-867-1348. |  | 2710 | total 2710 units | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NM, NV, NJ, NY, OK, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY; and countries of: Argentina, Australia, Canada, China, France, Germany, Indonesia, Japan, The Netherlands, Pakistan, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, United Kingdom, and Venezuela. | Terminated | September 27, 2013 | 337 | 2710 | Not_Computer
Z-0481-2013 | 63518 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Single Channel Infusion Pumps with Plug-and-Play Modules that provide wired Ethernet and wireless local area networking; Hospira, Inc., Lake Forest, IL 60045; pump list number 11971, module list numbers 12101, 12102 and 12393; pump list number 11973, module list number 12380; pump list number 12391, module list numbers 12097 and 12680 |  |  | October 26, 2012 | 2012 | Hospira Inc. | The volume control knob on some Plum A+ single channel infusers (located on the back of the infuser) may not function as described in the System Operating Manual - the direction for Loud and Quiet may be reversed. The manual indicates a clockwise rotation of the knob decreases the volume; however, on some of the devices the alarm volume increases when the knob is turned clockwise. | Device Design | The firm, Hospira Inc. sent "URGENT DEVICE FIELD CORRECTION" letters dated October 26, 2012 to all affected customers. The letters identified the products, problem and action(s) to be taken. The customers were instructed to read the information provided carefully, complete the attached Reply Form and return it to the firm via fax at 1-877-650-8667. The letter states that the System Operating Manual will be updated to clarify the instructions for setting the volume, and that the updated manuals will be sent as soon as they are available. Customers were instructed to notify their consignees if they have further distributed the affected product. The customers were requested to contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, for additional information and technical assistance. |  |  |  | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of Australia, Bahrain, Belgium, Brazil, Canada, Chile, Colombia, Denmark, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Lavenia, South Africa, Spain, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates and Vietnam. | Open, Classified |  |  |  | Not_Computer
Z-0504-2013 | 63728 | No | EOQ | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | WANG Transbronchial Aspiration Needles Cat. No. SW-121, WANG Cytology Needles, Central, Carinal and Mediastinal Regions, 21 g x 15 mm long needle. --- Common Name: Bronchoscope (flexible or rigid) and accessories. The ConMed WANG Transbronchial Aspiration Needles are designed for use with a bronchoscope. They are used to puncture the tracheobronchial wall and to aspirate sufficient tissue and/or cell specimens to stage bronchogenic carcinoma. These are disposable sterile, single-use devices. |  |  | November 20, 2012 | 2012 | ConMed Corporation | ConMed received complaints from a user facility reporting that the WANG Transbronchial Aspiration Needles would not insert properly through the instrument channel of a 2.0 mm bronchoscope. If the device will not fit through the instrument channel of a 2.0 mm bronchoscope, it must be replaced with an alternate device and could result in minor inconvenience to the user. | Component design/selection | ConMed Corporation sent an Urgent Field Safety Notice letter dated November 30, 2012 with Response Form , to all affected customers via USPS Priority Mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete Attachment II and return it with any unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502United Sates of AmericaFor information regarding this recall call 315-624-3225. |  | 452 | Total 452 | Worldwide Distribution - USA including CA , FL, IA, IN, MA, MD, ME, MO, MS, NC, NM, OH, PA, TX. Foreign distribution to Spain, Portugal, Japan and Hungary. | Terminated | August 25, 2015 | 1008 | 452 | Not_Computer
Z-0495-2013 | 63709 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Information Center/Client and Philips IntelliVue Information Center iX |  |  | November 09, 2012 | 2012 | Philips Healthcare Inc. | Internal audio cable may have diminished retention force due to damaged contacts, which could lead to intermittent or loss of audible alarms. | Component design/selection | Philips sent an URGENT - Medical Device Correction letter dated November 9, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips is conducting a voluntary correction to replace the audio cable on affected devices. Customer were instructed to follow the "Action to be taken by CustomerUser" section of the instructions. Customers with wquestions were instructed to contact their Phililps representataive or contact the Customer Care Solutions Center at (800) 722-9377.For questions regarding this recall call 978-687-1501. |  | 3165 | 3165 | Worldwide Distribution - USA (nationwide) and Canada. | Terminated |  |  | 3165 | Not_Computer
Z-0501-2013 | 61209 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | Empi STIMC RE Self Adhesive electrodes. For use on TENS, NMES, FES. a DJO brand |  |  | February 09, 2012 | 2012 | Covidien LLC dba Uni-Patch | Covidien is conducting a voluntary recall regarding specific lots of TENS electrodes because it was determined that during the manufacturing process the surface of the wires in these lots may have been exposed to an employee who had a minor cut. | Employee error | Covidien sent An Urgent Product Recall letter, dated February 9, 2012, to all affected consignees. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the affected product if found. The letter asked customers to fill in teh quantity remaining in their inventory and fax it to 651-565-5888. A Covidien Customer Service rep would contact the consignee for returning product. Consignees were also requested to notify their customers of the recall.For questions regarding this recall call 508-261-8532. |  | 28138 | 28,138 pouches | Nationwide Distribution including CA, IL, MA, MN, and SD. | Terminated | January 03, 2013 | 329 | 37036 | Not_Computer
Z-0508-2013 | 63682 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Siemens Healthcare ADVIA Centaur CA 19-9 |  |  | November 14, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Operating range of the ADVIA Centaur systems is 18C to 30ÂżC (64.4ÂżF to 86.0ÂżF), a change in room temperature may increase or decrease assay results, depending on the specific assay and ADVIA Centaur system used. | Device Design | Siemens Healthcare issued An Urgent Medical Device Correction Notice (for US customers) and an Urgent Field Safety Notice (for customers outside the US) dated November 5, 2012, to all affected customers. These notices inform customers who used the ADVIA Centaur BNP, BR, CA 19-9, Digitoxin, Folate, and Theophylline 2 assays to reduce their instrument operating temperaturerange to 20 -25 degrees C and for the DHEAS assay to reduce theinstrument operating temperature range to 18 -26 degrees C when theseassays are being used. These notices also direct customers to review thecontents of the correction notice with their Laboratory Director or MedicalDirector. A confirmation fax-back form was included to ensure that allcustomers worldwide have been notified and have complied with themandatory action.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local technical support representative.For questions regarding this recall call 508-668-5000. |  | 50 | 250 test=525 and 50 test=12243 | Worldwide Distribution - USA (nationwide) and Internationally to the following countries: ArgentinaAustraliaAustriaAzerbaijanBelgiumBosnia Herzeg.BrazilCanadaChileCOLOMBIACroatiaCzech RepublicDenmarkEgyptEstoniaFinlandFranceGermanyGreeceHong KongHungaryIcelandIndiaIndonesiaIranIrelandItalyJapanKazakhstanLatviaLithuaniaMalaysiaMexicoNetherlandsNew CaledoniaNorwayOmanP.R. ChinaPhilippinesPolandPortugalQatarRep. of YemenRepublic KoreaReunionRomaniaRussian Fed.Saudi ArabiaSerbiaSlovakiaSloveniaSouth AfricaSpainSwedenSwitzerlandTaiwanThailandTurkeyU.A.E.United KingdomUruguayVietnam | Terminated | September 25, 2015 | 1045 | 25245 | Not_Computer
Z-0513-2013 | 63614 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM 6/7 Patient Monitor |  |  | August 11, 2010 | 2010 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is an issue with the DPM 6 and DPM 7 monitors where the following features may be unavailable : Full Disclosure, Drug, Hemodynamic Calculations, Renal Calculations, Oxygenation Calculations, and Ventilation Calculations. | Device Design | Mindray sent a "PRODUCT CORRECTIVE ACTION" letter dated August 11, 2010 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their Mindray Service Representative at 1-800-288-2121 for assistance with this issue. |  | 124 | 124 units | Nationwide Distribution-USA (nationwide) including the states of WA, IL, FL, MN, SC, NV, MI, TN, PA, LA, NE, TX, and IA. | Terminated | August 02, 2013 | 1087 | 124 | Not_Computer
Z-0515-2013 | 63496 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAVIG PORTEGRA 2 system operated in conjuction with a Siemens AX system |  |  | October 10, 2012 | 2012 | Siemens Medical Solutions USA, Inc | A particular component may not have been fitted during the installation of some MAVIG PORTEGRA 2 systems. | Equipment maintenance | Siemens sent a Customer Safety Advisory Notice on October 3, 2012, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were reminded that they should contact the service department in the event of damage or other irregularity. Customers were instructed to forward the information contained in the notice to their employees responsible for the operation of the affected product. Customers were also asked to forward the safety information to other organization affected by the product and to any users of the product.For questions regarding this recall call 610-219-4834. |  | 1315 | 1315 | Nationwide Distribution | Terminated | November 06, 2014 | 757 | 1315 | Not_Computer
Z-0520-2013 | 63758 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD CYCLESURE 24 Biological Indicator (BI) product, P/N 14324. |  |  | October 03, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling all lots of Sterrad Cyclesure 24 Biological Indicator (BI) manufactured between February 2008 and December 2011 because it may not have adequate data to support the entire duration of the labeled shelf-life. | Nonconforming Material/Component | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE RECALL (REMOVAL) letter dated November 2, 2012 to all customers who received the Sterrad Cyclesure 24 Biological Indicator (BI) product P/N 14324 manufactured between February 2008 and December 2011. The letter informed the customers of the problem identified and the actions to be taken. ASP determined that the 24 BI may not have adequate data to support the entire duration of the labeled shelf life of affected product manufactured during the above timeframe. The letter also informed customers that this is a different recall from the previous July 3, 2012 STERRAD CYCLESURE 24 BI recall. Customers with questions or need additional packing labels are instructed to call Stericycle at (877) 650-7692. |  | 38532 | 38,532 units | Worldwide Distribution-USA (nationwide) and the countries of AR, AU, BE, BR, CA, CL, CN, CO, EC, GT, HK, ID, IN, JP, KR, MX, MY, PA, PE, PH, SG, TH, VE, and RU. | Terminated | July 19, 2013 | 289 | 38532 | Not_Computer
Z-0519-2013 | 63530 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Quasar Power Pack contains a Baby Quasar and a Baby Blue. The Baby Quasar is an infrared LED lamp. The Baby Blue uses blue LED light to be used in combination with the Baby Quasar to treat acne. |  |  | October 25, 2012 | 2012 | Quasar Bio-Tech, Inc. dba Silver Bay LLC | Quasar Bio-Tech is recalling their Baby Quasar, Baby Quasar Pink, Quasar Power Pack, and Quasar MD because the device is unapproved. | No Marketing Application | The firm, Quasar Bio-Tech, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated November 16, 2012 to its customers. All domestic customers were notified by Certified mail. The letter described the product, problem and actions to be taken. The customers were instructed to cease distribution and remove the products from sale; to complete and return the enclosed Response Form via e-mail to DeNette@quasarbiotech.com or fax to: Quasar Bio-Tech, Attn: Quality/Recall Coordinator: at 941-306-5816, and return all remaining stock to Quasar Bio-Tech at Quasar Bio-Tech Inc., Attn: Quality/Recall Coordinator, 1431 Tallevast Road, Sarasota FL 34243.If you have any questions, contact the Quality/Recall Coordinator at 941-306-5812, Monday-Friday from 9:00 am to 5:00 pm Eastern Standard time. |  | 7734 | 7,734 devices total (included in batches identified for the Baby Quasar ) | Worldwide distribution: USA (nationwide) including states of: AZ, CA, CO, FL, GA, HI, ID, IL, KS, MD, MI, MN, MO, NC, NJ, NY, NV, OH, OR, TN, TX, UT, VA, and WA; and countries of: Aruba, Canada, Cayman Islands, China, Estonia, Ireland, Malaysia, New Zealand, Singapore, Sweden, Taiwan, and United Kingdom. | Terminated | May 21, 2013 | 208 | 7734 | Not_Computer
Z-0522-2013 | 63680 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 123 POC system |  |  | November 14, 2012 | 2012 | Roche Diagnostics Operations, Inc. | It has been determined that cobas b 123 Fluid Pack COOX cuvettes for lot numbers 21426121 and 21426171, exhibit a significant increase in the out-of-box failure rate caused by a change in the thickness in the cuvette layer. Internal investigations found the altered thickness of the cuvette was due to production equipment settings. Cuvettes were produced within specifications, but within the upper | Software design (manufacturing process) | Roche sent an Urgent Medical Device Removal (UMDR12-268) notification and faxback form on November 14, 2012, to all affected consignees. Customers were requested to discontinue the use of cobas b 123 Fluid Pack COOX, 200 lots 21426121 and21426171 and discard the product per local regulations. Customers were instructed to use the enclosed faxback form to order replacement product at 1-800-722-7222. For questions customers were instructed to contact Roche Diagnostics Technical Support at 1-800-526-2272.For questions regarding this recall call 317-521-3966. |  | 4600 | 23 units of Fluid Pack COOX, 200 were distributed | Nationwide Distribution including NE and AZ. | Terminated | August 07, 2013 | 266 | 4600 | Not_Computer
Z-0523-2013 | 63723 | No | GRO | Haemophilus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD BBL Taxo XV Factor Strips, Catalog number 231104, in shelf pack vials, fifty strips/vial labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152 USA 800-638-8663, www.bd.com/ds*** |  |  | October 22, 2012 | 2012 | Becton Dickinson &amp; Co. | In vitro diagnostic test kit was not manufactured according to specifications and may cause false negatives or misidentification of Hemophilus species present in patient specimens. | Employee error | BD Diagnostic Systems sent an Urgent Product Recall letter dated October 22, 2012, to all affected customers.. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to discontinue distribution and discard affected inventory for credit. . Customers were requested to return a response form by fax to 410-316-4258. For questions customers were instructed to contact BD Customer Service Department at 1-800-638-8663. For assistance with replacements, customers were instructed to call 1-800-675-0908.For questions regarding this recall call 410-316-4258. |  | 650 | 650 vials | Worldwide Distribution - USA (nationwide) to international customers in Belgium, India, Japan, South Korea, Thailand and Mexico | Terminated | January 24, 2013 | 94 | 650 | Not_Computer
Z-0497-2013 | 63712 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMedÂż II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The implantable components of the SynchroMed II Infusion System include the pump, catheter, and catheter accessories. |  |  | November 09, 2012 | 2012 | Medtronic Neuromodulation | Medtronic notified Healthcare Professionals of the impact of unapproved drugs on the performance of the SynchroMed infusion pump system. Use of unapproved drugs with SynchroMed pumps can result in an increased risk of permanent motor stall and cessation of drug infusion. | Nonconforming Material/Component | The firm Medtronic, sent an "Urgent medical Device Safety Notification" dated November 2012, to its customers beginning November 09, 2012. The letter described the product, problem and actions to be taken. The customers were instructed to follow the recommendations provided in the letter and complete and return the PHYSICIAN REPLY FORM via mail at Medtronic Neuromodulation Quality Assurance, 7000 Central Ave NE, Mailstop RCC150, Minneapolis, MN 55432; fax to: 800-897-3899 or email to: neuro.quality@medtronic.com.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 174700 | 174,700 devices total | Worldwide distribution: USA (nationwide) and countries including: Canada and Mexico. | Completed |  |  | 174700 | Not_Computer
Z-0521-2013 | 63694 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Ready-to-Use Mouse Monoclonal Antibody Terminal Deoxynuceotidyl Transferase Staining Reagent; an in vitro staining reagent for the qualitative identification by light microscopy of Terminal Deoxynucleotidyl Transferase molecules in paraffin sections; Catalog No. RTU-TdT-339; packaged in 7 mL vials; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom. |  |  | November 14, 2012 | 2012 | Leica Microsystems, Inc. | RTU-TdT-339 is not stable up to the expiry date on the product labeling, affecting the staining intensity. | Nonconforming Material/Component | Leica Microsystems, Inc. sent a Product Recall Notification letter dated November 14, 2012, via first class mail on the same date, to all affected customers. The letter identified the product ,the problem, and the actions to be taken by the customers. Customers were instructed to discontinue use of the product and to return any unused/part used reagents to the Richmond, IL, location for replacement. The accounts were asked to complete the attached Acknowledgement Form, acknowledging receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-236-3747. Additional questions should be addressed to the account's Leica Representative. |  | 322 | 322 vials | Worldwide Distribution -- USA including Arkansas, California, Florida, Illinois, Maryland, Michigan, Minnesota, Missouri, North Carolina, Oregon, Rhode Island, Tennessee, Texas and Wisconsin, and the country of Canada. | Open, Classified |  |  | 322 | Not_Computer
Z-0514-2013 | 63710 | No | LSZ |  |  | Anesthesiology | PMA | Patient Circuit used with the Life Pulse High Frequency Ventilator. The Patient Circuit is a required disposable component of the Life Pulse High Frequency Ventilator. |  |  | November 19, 2012 | 2012 | Bunnell, Inc. | Customer complaints received indicate the heater wire insulation can melt, causing a short which creates sparking and smoke, in the circuit close to the humidifier cartridge. Bunnell is recalling all lots of circuits distributed between March 12, 2012 and November 30, 2012. | Component design/selection | On 12/5/12, the firm, Brunnell, Inc., issued a public press notification to their consignees. A. On or before 12-12-2012 a certified letter will be sent to all customer listed in the customer database. Specifically each customer will receive two certified letters; one directed to Director of Respiratory Therapy and a second to the Director of Biomedical Engineering. The certified letter will contain a copy of the Recall Notice and a copy of the Certificate of Medical Necessity. The Certified Letter receipts will be maintained. On the mailing date a copy of the Recall Notice will be placed on the Bunnell website. B. An Excel database will be maintained for each customer/lot #. The database will contain the current status for each. Included will be: 1. Customer Information 2. Date Certified Letter Sent 3. Date receipt received and filed 4. Date product returned or that no stock remains 5. Lot number &amp; Quantity C. On or before 02-01-2013, a second letter will be sent to those customers who have not responded to the original certified letter. D. Data will be maintained in the database E. On or before 04-01-2013, a phone call will be placed to those customers who have not responded to either of the letters. F. Data will be maintained in the database G. On or before 04-15-2013, a second phone call will be placed to those customers who have not responded to the letters or previous phone call. H. Data will be maintained in the database I. On or before 06-07-2013, a Final Recall Notification certified letter will be sent to those customers who have not responded to either of the previous letters or phone calls. J. The Certified Letter receipt will be maintained K. Data will be maintained in the database L. When PMA for alternate heater wire is approved, steps A-K will be repeated. Dates will be determined. M. Following completion of Step L, a final review of the database will occur. Final tabulation of the status for each custo |  | 5743 | 5743 patient circuits distributed in commerce | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, Malaysia, and Uruguay (animal study only). | Terminated | February 21, 2014 | 459 | 5743 | Not_Computer
Z-0524-2013 | 63655 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module with v2.0 software used with Solar 8000M/l, Transport Pro or Carescape B650/B850 host monitor. The Patient Data Module (PDM) is intended to provide physiologic parameter data on adult, pediatric and neonatal patients during bedside and transport patient care episodes. |  |  | February 17, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of two issues with the Patient Data Module using Version 2.0 software that may impact patient care. | Software design | Consignees were sent on 2/17/2012 three GE Healthcare "Important Product Information" letters dated 10/17/ 2011. The letters were addressed to Healthcare Administrator/Risk Manager, Director of Biomedical Engineering, and Chief of Nursing. The letters described the Issue, Instructions, Affected Product Details, Product Correction and Contact Information. |  | 10250 | 10,250 | Worldwide distribution, including the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, LA, KS, KY, LA, ME,MD, MA, MI, MN, MS,MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK OR, PA, SC, TN,TX, UT,VA, WA, WT, WY and the countries of the UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAK IA, SINGAPORE, SAUDI ARABIA, QATAR, POLAND, PERU, PANAMA, NORWAY, NEW ZEALAND, NETHERLAND, MEXICO, MALAYSIA, LEBANON, KUWAIT, KOREA, JAPAN, ITALY, INDONESIA, INDIA, IRELAND, GERMANY FRANCE, FINLAND, EGYPT, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHILE CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, ARGENTINA, and ALGERIA. | Terminated | December 03, 2013 | 655 | 10250 | Computer
Z-0494-2013 | 63650 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000/4000/5000. |  |  | June 26, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to an unexpected loss of CustomAutomatic NBP Measurement cycling associated with the Dash 3000/4000/5000 patient monitors. | Software design | GE Healthcare sent a Important Product Information Letter dated June 21, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If your facility is using the Custom Auto NBP Measurement feature in conjunction with Automatic View On Alarm and/or View Other Beds options, GE Healthcare recommends that you stop using the Custom Automatic NBP Measurement feature and use the Auto NBP function instead.All Dash3000/4000/5000 patient monitors with software version 6.0 and later with the networking option enabled may be affected. You can determine the software version of the Dash monitors by selecting MORE MENUS &gt; MONITOR SETUP &gt; SOFTWARE CONFIGURATION. Product CorrectionGE Healthcare is working on a software correction that will be provided to you at no charge once it is available. We will contact you to schedule the correction.Contact InformationIf you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044 or your local Service Representative.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above.. |  | 119260 | 119,260 devices | Worldwide Distribution -- USA (nationwide) including the countries of AUSTRALIA, ANGOLA,ARGENTINA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNG ARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHSTAN, KENYA, KOREA, KUWAIT, LATIVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MALTA, MOROCCO, NEPAL, NETHERLAND, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PALESTINIAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN THAILAND, TRINIDAD AND TOBAGO,K TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARB EMIRATES , UNITED KINGDOM and URUGUAY | Terminated | June 23, 2017 | 1823 | 119260 | Computer
Z-0535-2013 | 63549 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Titan Spine Endoskeleton TA Implant products. |  |  | October 19, 2012 | 2012 | TITAN SPINE, LLC | Titan Spine LLC, is conducting a recall on the Endoskeleton TA VBR products. The labeling accompanying this device contains incorrect indications for use. | Labeling False and Misleading | Titan sent an Urgent Medical Device Recall letter dated October 19, 2012, to all affected customers. The letter described the problem , the product affected and the corrective measures to be taken. Customers were instructed to review any surgeries that were performed at the L5-51 level with the affected device. Customers were asked to communicate any complications associated with those cases to Titan Spine. Customers with questions should call 1-262-242-7801.For questions regarding this recall call 262-242-7801. |  | 92 | 92 pcs. | Nationwide Distribution including NV | Terminated | January 14, 2013 | 87 | 92 | Not_Computer
Z-0536-2013 | 63653 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Polaris 5.5 Button Lock Screw Inserter. Catalog number 14-500178 |  |  | September 29, 2012 | 2012 | Ebi, Llc | Broken screw insertor hardness was measured and found to be below the acceptable limits. | Process control | Biomet mailed Urgent Medical Device Recall Notice/Fax Back Response Form dated 9/25/2012 to the distributors on 9/29/2012. |  | 15 | 15 units | Nationwide Distribution, including the states of MO, TX, and IN. | Terminated | December 07, 2015 | 1164 | 15 | Not_Computer
Z-0527-2013 | 63725 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | EBI XFIX DFS Rail System (Adult Rail Gradual Swivel Clamp). The Biomet Rail System consists of a unilateral external fixator rail, modular bone screw clamps, and compression and distraction units. The device is used in the treatment of bone fractures. |  |  | November 09, 2012 | 2012 | Biomet, Inc. | It was reported that the number scale/graduation markings were missing on 14 pieces of the swivel clamp. | Process design | On 11/9/2012 an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter was sent to all consignees (distributors and Doctor notices) informing them of the issue and providing recommended actions. Customer contact number (574) 372-1570 Monday through Friday from 8 am to 5 pm. |  | 49 | 49 | Worldwide Distribution, including the states of TX, TN, PA, CA, NY, FL, and MI and the countries of Japan and Columbia. | Terminated | August 28, 2013 | 292 | 49 | Not_Computer
Z-0526-2013 | 63824 | No | HCC | Biofeedback device. | Neurology | Neurology | 510(K) Exempt | EEG NeuroAmp, Model Number: CS 10090; CS 10137. |  |  | November 15, 2012 | 2012 | EEG Info | EEG Info, Inc. is recalling the EEG Neuroamp device because they have identified a potential risk associated with the mislabeling of certain devices. | Labeling False and Misleading | EEGInfo sent a recall notification letterdated November 29, 2012, to all affected customers. The letter informed the customers they have not received approval from the FDA to market and distribute a device for therapeutic applications other than that of its FDA intended use-biofeedback relaxation. Customers are informed that the safety and effectiveness for other indications has not been established. The risk or likelihood of serious adverse health consequences is remote. Customers were instructed to forward this notice if affected product was further distributed. Customers with questions about the notification letter are instructed to contact EEG Info's customer support line during normal business hours, Monday-Friday, from 8am-5pm PST. For questions regarding this recall call 818-456=5965. |  | 1308 | 1308 units | Nationwide Distribution | Terminated | June 11, 2013 | 208 | 1308 | Not_Computer
Z-0532-2013 | 63520 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Medline Total Knee CDS; a single patient prescription procedure pack and disposables, including a back table cover; |  |  | September 20, 2012 | 2012 | Medline Industries Inc | The Back Table Cover component in the custom pack was placed in the non-sterile portion of the pack when it should have been in the sterile portion. | Packaging | Medline Industries Inc notified the sole direct account on September 20, 2012 via telephone that the affected custom pack were recalled. Instruction was provided to the staff at the facility not to use the pack until it had been reworked with a sterile back table cover. |  | 18 | 18 packs | US (nationwide) in the state of West Virginia | Terminated | December 17, 2012 | 88 | 18 | Not_Computer
Z-0531-2013 | 63713 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand ARIA Radiation Oncology, ARIA Oncology Information System Radiation Oncology, Import Export Application, Model Number: HIT, Affected Versions: ARIA version 11.0 below 11.0.55, Build 11.0.28 and 11.0.34, Reference/FSCA Identifier: CP-09798; |  |  | November 05, 2012 | 2012 | Varian Medical Systems, Inc. Oncology Systems | Varian Medical Systems has identified an anomaly with the ARIA Oncology Information System Radiation Oncology [ARIA RO] DICOM import/export functionality where a 360 degree arc field may convert to static field during DICOM export or import. | Software design | Varian sent an Urgent Medical Device Correction Urgent Field Safety Notice dated November 6, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that the plan was not converted unintentionally during import/export. Customers would be contacted by a Customer Service Representative when the correction was available to schedule installation on their system. Customers with questions were instructed to contact their local Varian Medical Systems Customer Support District or Regional Manager.US and Canada 1-888-827-4265.For questions regarding this recall call 650-424-5731. |  | 264 | 264 units | Worldwide Distribution | Terminated | March 03, 2014 | 483 | 264 | Not_Computer
Z-0529-2013 | 63732 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Mitek 11 MM Fully Fluted Reamer, Sterile |  |  | November 16, 2012 | 2012 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Specific units of the DePuy Mitek Sterile Reamers are not labeled correctly, 7 units were labeled as 11MM while the box actually contained a 10 MM reamer | Packaging process control | DePuy Mitek sent an Urgent Voluntary Product Recall letter dated November 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to refer to the attached instructions to report their inventory status and for returning the affected product. Customers with questions should call 1-508-880-8477. For questions regarding this recall call 508-880-8100. |  | 20 | 20 | Nationwide Distribution including IL and NJ | Terminated | December 21, 2012 | 35 | 20 | Not_Computer
Z-0533-2013 | 63781 | No | MBV | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Offset Tibial Tray 2.5 mm Adaptor, REF 141490 |  |  | November 15, 2012 | 2012 | Biomet, Inc. | Biomet investigation identified that the square end of the offset tibial tray adaptor locking insert is oversized and may not engage insert tool screwdriver. This could cause a possible delay in surgery greater than 30 minutes, potentially exposing the patient to the increased risks from being under anesthesia for a longer duration. | Employee error | Biomet Recall sent an Urgent Medical Device Recall Notices letter dated November 15, 2012 to distributors, hospital operating room(OR)managers and Surgeons receiving the affected product. The letter identified the affected product, the reason for the recall and actions to be take, including directions to immediately locate and discontinue use of the affected product. Customers were asked to confirm receipt of the notice by calling 1-888-348-9500 extension 1570. Customers with affected product to be returned are asked to obtain RGA # by calling 574-372-6677 or email: rgarequest@biomet.com. A FAX Back Response Form is to be filled out and sent within 3 days of receipt to 574-372-1683. Return product to the Attention of Audrey Daenzer, Field Action Coordinator, Biomet, Attention Return Goods, Building B Recalls, 56 East Bell Drive, Warsaw, IN 46580. Questions related to the notice should be directed to 474-372-3983 or 574-371-3009, Monday- Friday 8 a.m. to 5 p.m. |  | 200 | 200 | Worldwide distribution - USA (nationwide) and the countries of Australia, Canada, Chile, Europe, Japan, Korea, Panama, Puerto Rico and Thailand. | Terminated | November 08, 2013 | 358 | 200 | Not_Computer
Z-0534-2013 | 63808 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Arobella Medical 6mm Qurette used with the Arobella Medical AR1000 Wound Therapy System. |  |  | November 05, 2012 | 2012 | Arobella Medical, LLC | This recall has been initiated because the Arobella Medical 6mm Qurette Q07 Lot Numbers 0001 and 0002 can fracture during use and may release a fragment or fragments. The 6 mm Qurette is a component of the AR1000 Quostic Wound Therapy System. | Nonconforming Material/Component | Arobella Medical sent an " Urgent: Medical Device Recall" letter dated November 2, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine their inventory, quarantine and return any product subject to the recall. If the affected product have been further distributed, customers were advised to notify their customers of the product recall. For questions call 952-288-2810. |  | 99 | 99 | Worldwide distribution - US (nationwide) and the states of AL, CA, ID, IN, IA, MI, MN, MS, NJ, NM, OH, PA. SC, TX, UT, VA, WI. including the country of SAUDI ARABIA. | Terminated | March 04, 2015 | 849 | 99 | Not_Computer
Z-0528-2013 | 63652 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray V12 Size: 320 mm X 320 mm X 450 mm N.W.: 3 kg G.W.: 6 kg Qty:1 Manufactured in China |  |  | September 24, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the V Series Monitor where the monitors touch screen may stop responding to touch. | Nonconforming Material/Component | Mindray issued an "Urgent: V Series Monitoring System Corrective Action" letter dated 9/24/2012 via certified mail with return receipt verification to all hospitals.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 to arrange for relplacement of the affected product. For questions regarding this recall call 201-995-8407. |  | 206 | 206 units (180 units domestic, 26 units foreign) | Worldwide Distribution - USA (nationwide) and internationally to Great Britian, Colombia, and Australia. | Terminated | August 15, 2013 | 325 | 206 | Not_Computer
Z-0538-2013 | 63641 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Coonrad/Morrey New Ulnar Revision Kit for Use with Non-Interchangeable Regular Humerals STERILE Rx, Tivanium QTY/PKG:1 |  |  | November 19, 2012 | 2012 | Zimmer, Inc. | Zimmer is initiating a lot specific recall because affected lots could include an extra Small Inner Pin instead of the required Regular Inner Pin in the Coonrad/Morrey Elbow Replacement Set and the New Ulnar Revision Kit . | Process change control | Zimmer sent "URGENT MEDICAL DEVICE RECALL" letter dated November 19, 2012 to Distributors, Risk Managers/ Implanting Surgeons, and direct accounts. The notice describes the product, problem, risks and actions to be taken. Consignees were instructed to immediately locate, quarantine and return affected product with the completed Inventory Return Certification form to Zimmer Product Service Department, 1777 West center St. Warsaw, IN 46580 (US accounts) or Zimmer International Logistics GmbH, ATTN: Fao Tanja Herold (Recall Warsaw), Max-Immelmann-Allee 12 79427 Eschbach Germany for international accounts. If an implant is removed from customer inventory, they were directed to provide a copy of the Risk Manager letter upon retrieval of that implant(s) to ensure all facilities are aware of the removal. Surgeons are asked to consider risk information in the letter should a patient present with unexplained symtoms. Hospitals that received direct shipments from the Warsaw Distribution Center will be sent a copy of the Risk Manager letter directly. Questions and concerns along with information on any implanted product is to be addressed to CorporateQuality.PostMarket@zimmer.com using the provided spreadsheet or contact Zimmer at 1-800-447-5633. |  | 17 | 17 | Worldwide distribution - USA nationwide and the countries of: Canada, Germany, Finland, Japan, Italy, San Juan, Switzerland, and UK. | Terminated | April 17, 2014 | 514 | 40 | Not_Computer
Z-0546-2013 | 63690 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona, The Personalized Knee System, Tibial Drill Guide, Cemented, 15.7 mm Diameter, Nonsterile, Nonsterile |  |  | November 14, 2012 | 2012 | Zimmer, Inc. | All lots of the Persona Cemented Tibial Drill, 5 Degree Cemented Stemmed Natural Tibia implants, and associated tools including the Tibial Drill Guide, Broach, and Sizing Plate are being removed and replaced with enhanced products and instruments to reduce risks of tibial perforation which could result in tibial bone fracture due to the stress riser created by a perforation; tibial plate loosening | Device Design | Zimmer sent distributors beginning the week of November 14, 2012, E-mail information requesting all equipment be collected in distributors office before shipment back to Zimmer and replacement. A telephone conference discussing the recall and amended surgical technique was held the week of November 14, 2012, and Urgent Recall notices were sent to Distributors describing the devices; risks; updated surgical technique with directions to hand deliver surgeon/risk manager letters to all affected surgeons and risk managers in hospitals receiving the product. The notices included directions to remove all affected product and ensure risk manager letter is provided. The letter stated a Zimmer Inc. employee will conduct onsite reconciliation and asked return of distribution Spreadsheet to CorporateQuality.Postmarket@zimmer.com within 24 hours of receiving notification. Product is to be returned with certificate of sterilization to Zimmer Product Service Department 1777 West Center St. Warsaw, IN 46580 or Zimmer International Logistics, Eschbach Germany for International Accounts.For questions regarding this recall call 574-372-4807. |  | 94 | 94 | Worldwide Distribution - USA (nationwide) and internationally to Germany and Europe. | Terminated | April 16, 2014 | 518 | 4035 | Not_Computer
Z-0556-2013 | 63775 | cane, safety walk | KHY | Cane. | Physical Medicine | Physical Medicine | 510(K) Exempt | Invacare Height Adjustable (walking) Canes, Model Numbers: 8916, 8917, and 40918-6. The device is packaged six (6) canes per shipping carton. |  |  | November 06, 2012 | 2012 | Invacare Corporation | Following their receipt of several customer complaints, Invacare recalled their height adjustable walking canes. The device was recalled from distribution based on reports indicating that there is a possibility that the button which holds the cane at the desired height, may suddenly and unexpectedly compress, causing the cane to collapse and cause possible injury to the user. | Device Design | The firm, Invacare, sent an "URGENT MEDICAL DEVICE RECALL (Removal)" letter dated November 6, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to examine their existing stock using the enclosed detailed 'Product Tracking Sheet' and quarantine any affected canes; Acknowledge receipt and understanding of the urgent recall notice and indicate the number of canes remaining in inventory on the enclosed, 'Provider Response Card;'. Fax the completed Provider Response Card to Invacare Regulatory Affairs at: 1- 440-326-3544; disassemble each cane and discard both pieces in the trash; (Disassembly instructions are also provided); and lastly, contact Invacare Customer Service at 1-800-668-2337, Monday - Friday, 8 a.m. - 8 p.m. EST, to order replacement merchandise at no charge.If you have any questions concerning these instructions, contact Invacare Customer Service at 1-800-668-2337, Monday - Friday, 8 a.m. - 6 p.m. ET. |  | 38280 | 38,280 canes | Worldwide distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN KS, KY, LA,MA, MD ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH,OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Australia, Puerto Rico, and Taiwan Province of China. | Terminated | August 20, 2013 | 287 | 38280 | Not_Computer
Z-0557-2013 | 62826 | No | ILK | Powered patient transport. | Physical Medicine | Physical Medicine | 510(k) | ThyssenKrupp Accessibility LEVANT Stairlift. |  |  | August 06, 2012 | 2012 | ThyssenKrupp Access Corp | The seat could break as a result of damage resulting from a seat belt being trapped underneath the foldable part of the seat when the seat belt is not buckled. | Device Design | The firm, ThyssenKrupp Access sent an "IMPORTANT SAFETY AND PRODUCT USE INFORMATION" letter dated August 6, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Contact the firm at 1-800-925-3100 ext. 5484 for assistance regarding this recall. |  | 5413 | 5,413 stairlifts | Worldwide Distribution-USA (nationwide) including the states of IL, IN, MI, OH, WI, IA, KS, MN, MO, NE, SD, AR, LA,OK, TX. AZ, CO, ID, MT, NM, UT, MT, AK, CA, HI, OR, WA, AL, FL, GA, KY, NC, SC, TN, MD, PA, VA, WV, CT, MA, ME, NH, NJ, NY, RI and VT and the countries of Canada, Nova Scotia and British Columbia | Terminated | January 29, 2014 | 541 | 5413 | Not_Computer
Z-0568-2013 | 63739 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Medline Angio Drape Pack Radiology-LF; a sterile custom single patient prescription procedure pack, including a waste bag, for use by the Interventional Radiology Department; Packaged for Medline Industries, Inc.; Reorder #DYNJ35993D |  |  | November 08, 2012 | 2012 | Medline Industries Inc | The non-sterile Waste Bag component in the custom pack was attached to the outside of the pack in its own packaging. The product was not labeled or identified as sterile, but was assumed to be sterile by the customer. | Packaging | The firm, Medline Industries, Inc., telephoned the consignee/customer on November 8, 2012. The phone call informed the customer of the product, problem and actions to be taken. The customer was instructed to inform the staff at the facility not to use the pack until it had been reworked to replace the non-sterile waste bag with the sterile equivalent. It was also noted that the product does not need to be returned as it is being replaced directly at the facility.If you have any question call 847-837-2759 or email lsimmons@medline.com. |  | 105 | 105 packs | Nationwide distribution: New York only. | Terminated | December 20, 2012 | 42 | 105 | Not_Computer
Z-0561-2013 | 63744 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | AutoTac Delivery Handle, REF 400-200, Rx Only, non-sterile, BIOHORIZONS, BIRMINGHAM, AL 35244. |  |  | November 19, 2012 | 2012 | BioHorizons Implant Systems Inc | The cleaning instructions provided may not be sufficient to remove residual contaminants after surgical use. | Component design/selection | BioHorizons sent an "URGENT DEVICE CORRECTION" letter on November 19, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact the Customer Care Department at (888) 246-8338 for questions regarding this recall. |  | 174 | 174 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Brazil, Bulgaria, Canada, Cayman Islands, Chile, Colombia, Costa Rica, Ecuador, Egypt, Fiji, France, Germany, Greece, Hong Kong, India, Indonesia, Iran, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Panama, Peru, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | September 24, 2013 | 309 | 1420 | Not_Computer
Z-0566-2013 | 63604 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | The Stryker Model 6390, Power-LOAD, is a power-loading cot fastener system designed to lift, lower, or steer ambulance cots into and out of the ambulance. |  |  | October 11, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Model 6390 Power-LOAD systems built between September 12, 2012 and September 27, 2012 may have been built with a load switch slide that does not meet specifications causing the lifting ability of the arms to be negatively impacted. | Component design/selection | All in process and Load Switch Slides within Stryker control were quarantined on 09/27/2012. Direct accounts were telephoned beginning 10/26/2012 with "Urgent Medical Device Correction" notices sent beginning on 11/30/2012. The notices explained the reason for the recall, hazards involved, and recommended actions. Stryker field service was to contact consignees to schedule replacement of the affected load switch slide. Consignees were asked to forward a copy of his notice to the new users and advise Stryker of their new location. Questions or concerns should be directed to (269) 389-6927 or (269) 389-6954. |  | 24 | 24 | USA, Nationwide including the states of PA, NM, MI, IA, CA, CO, WA, VA, IL, WI, NY, WY, MO, TX, and CT. | Terminated | August 30, 2013 | 323 | 24 | Not_Computer
Z-0569-2013 | 62764 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Cancellous Screw |  |  | November 28, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary recall of ten lots of the PinnacleÂż Cancellous Screw due to an incorrect description on the label. The label description reads Pinnacle Cannulated Screw and should read Pinnacle Cancellous Screw. | Labeling Change Control | The firm, DePuy Orthopedics, Inc., sent an "URGENT INFORMATION - MEDICAL DEVICE RECALL NOTICE" dated July 2012 to its customers via email and mail. The notice described the product, problem and actions to be taken. The customers were instructed to determine whether they would like to retain and use the product or return affected product to DePuy in Warsaw, IN, and to complete and return the Reconciliation forms to your DePuy Sales Representative or by fax to: 574-371-4939. DePuy sales representatives are expected to aid customers in returning of recalled products.For questions about recall information provided, please contact the Manager, Post-Market Surveillance, 574-372-7333 (M-F; 8 am - 5 pm EDT.) |  | 118 | 118 units | Nationwide distribution: USA states of: GA, MA, MN, SC and VA. | Terminated | March 20, 2013 | 112 | 118 | Not_Computer
Z-0560-2013 | 61593 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Advantage Workstation Volume Share 2 (versionAW4.4_04) with Volume Viewer 3 and 3.1 (version7.0 to version7.5 and version 8.1) and only when or more of these applications is installed: |  |  | May 29, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has became aware of two issues with the Advantage Workstation.1) A possible mismatch between the label of the tracked vessel and the underlying image associated with the cardio- vascular applications of the Advantage workstation that may impact patient safety.2) A refresh problem of the stenosis / aneurysmmeasurement tools associated with the cardio-vascular applications of | Software design | GE Healthcare sent two Urgent Medical Device Correction letters dated May 23, 2012, to all affected customers. Each letter addressed a separate issue. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. Customers were instructed to use the following precautions until corrected software was installed:1. Please avoid the use of leading or trailing spaces in Vessel Names2. Be sure to press "Show Tracking" after depositing new bifurcation points3. When eating a multi-point vessel tracking, remove and replace the entire branch, not only the final end-point.4. When analyzing multi-phase cardiac images, do not delete vessel branches from a phase other than the one in which it was created.Customers were instructed that the new application version that addresses this issue will soon be installed on all systems. Customers with questions or concerns were instructed to call the following:United States 1-800-437-1171.Japan 0120-055-919For other countries they were instructed to contact their GE Healthcare Service Representative.For questions regarding this recalla call 262-513-4122 |  | 548 | 548 (118 USA; 430 OUS) | Worldwide Distribution - USA including AL, CA, CO, FL, GA, IL, IN, IA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX , UT, VA, WI and Internationally to the following countries:VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SOUTH AFRICA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIA, ROMANIA, REUNION, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, NORWAY, NIGERIA, NEW CALEDONIA, NEPAL, MEXICO, MALAYSIA, KOREA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRIA, AUSTRALIA, ARGENTINA ALGERIA. | Terminated | January 17, 2013 | 233 | 548 | Computer
Z-0563-2013 | 63807 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | CoverStainer Slide Rack |  |  | November 26, 2012 | 2012 | Dako Denmark A/S | The expected life time of the CoverStainer slide racks (CS 10330) does not meet established performance requirements, originating from variations in the manufacturing process. | Device Design | Dako sent a Voluntary Product Recall Notification dated November 13, 2012, to all affected customers. This voluntary recall notification was initiated because the expected life time of the CoverStainer slide racks (CS 10330) does not meet established performance requirements, originating from variations in the manufacturing process. The result can be that slides are falling out of the rack during the staining process. While this has been observed at only a few customer sites, the corrective action includes all CoverStainer installations. Customers were notified that Dako would correct the rack situation by using a new manufacturer and has imroved the rack design. Customers with questions were instructed to contact their local Dako representative.For questions regarding this recall call 805-566-3036. |  | 2650 | 25 Units US (50-100 racks per unit) 106 total | Worldwide Distribution - USA (nationwide) and the following countries:AustraliaFrance Rep. of KoreaBelgium Germany RussiaBrazil Israel Saudi ArabiaCanada Italy SpainChile Japan SwedenChina Mexico TaiwanDenmark Netherlands TurkeyEcuador Norway UKFinland Poland | Terminated | May 14, 2013 | 169 | 2650 | Not_Computer
Z-0567-2013 | 63777 | No | MHY |  |  | Neurology | PMA | Medtronic, DBS Extension Kit for Deep Brain Stimulation, Catalog # 7482, 37085, 37086. Rx Only, Manufactured in: Medtronic, Inc. Villalba, Puerto Rico. |  |  | December 03, 2012 | 2012 | Medtronic Neuromodulation | Medtronic has identified 140 Restore DBS Extension Kits (model numbers 7482, 37085, and 37086), which may contain a damaged winged connector boot. The connector boot may not have been molded properly and may be cracked. Approximately 14% of the 140 affected extension kits may contain a damaged radiopaque winged connector boot. | Nonconforming Material/Component | Medtronic beginning November 15, 2012 , initiated a retrieval of unused inventory in the field. In the US, three customers were identified. Medtronic field representatives visited these US accounts to facilitate product identification and removal. If product was removed a letter was provided to the facility describing the recall. The letter dated November 2012, "Medtronic Medical Device recall" was addressed to the Hospital. Beginning on December 3, 2012, Medtronic mailed healthcare professional letters to physician who have used the affected product. The Medtronic " Medical Device Recall" letter was dated November 2012. The letter identified the problem, the product involved in the recall, and provided advise if the product was used. A reply card was provided to each physician to confirm they have received and understood the communication. For further questions please call (763) 514-2000. |  | 140 | 140 units | Worldwide Distribution -- USA (nationwide) including the states of CT, MI, NH, NY, OH, PA and VT., and the countries of AUSTRALIA, JAPAN, BELGIUM, FRANCE, GERMANY, ITALY, NETHERLANDS, POLAND, SPAIN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | November 14, 2013 | 346 | 140 | Not_Computer
Z-0558-2013 | 63776 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | MAYFIELD Composite Series Base Unit, Model A3101, a Swivel Adaptor (A3018) is an integral component of the Base Unit, Model A3101. |  |  | November 05, 2012 | 2012 | Integra LifeSciences Corporation | An investigation of an adverse trend of complaints conducted by Integra LifeSciences identified various product performance issues which could impact the functionality of the firm's Mayfield Composite Series Base Units and Mayfield Skull Clamps. In order to prevent additional complaint reports, the firm decided to recall and replace certain lot codes of the Base Units and Skull Clamps that were d | Under Investigation by firm | Integra issued an 'Urgent: Voluntary Medical Device Recall Notification with 'Recall Acknowledgment and Return Form' to all Integra Sales Representatives and Customers via visit and E-mail on November 5, 2012. The recall notice informs the customers of the firm's receipt of complaints pin-pointing quality issues with the MAYFIELD Base Units and MAYFIELD Skull Clamps which may impair their functionality. The customers are asked to sign and return the enclosed Recall Acknowledgement and Return Form so as to document their receipt of the recall notification and replacement products. The replacement products are to be delivered by the Integra NeuroSpecialist who is also responsible for returning the signed forms to Integra as documentation that the recalled devices have been replaced. Should the customer(s) have questions or concerns, they are directed to contact Integra Customer Service at: 1-877-444-1114,Option 3 for Technical Support. |  | 5 | 5 base units were distributed | Nationwide Distribution including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NY, OH, OK, OR, PA TN, TX, UT, VA, WA and WI. | Terminated | February 21, 2014 | 473 | 13 | Not_Computer
Z-0570-2013 | 63737 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix PMIC-108 Panels, Catalog number 448418, labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, USA. |  |  | November 06, 2012 | 2012 | Becton Dickinson &amp; Co. | An antibiotic susceptibility test kit may produce false intermediate or false resistant results with clinically significant bacteria. | Nonconforming Material/Component | The firm, BD, sent an "URGENT PRODUCT RECALL" letter dated November 2012 via the telephone and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and discard any remaining packages for replacement and to complete and return the Acknowledgement Receipt Form, whether or not you have any inventory remaining, via fax to: BD Regulatory Compliance at 410-316-4258. BD will issue replacements for the discarded materials.If further assistance is needed regarding replacements please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 12250 | 12,250 tests | Nationwide distribution: USA including states of: AR, CA, KY, OK, RI, WI, and WY. | Terminated | January 09, 2013 | 64 | 12250 | Not_Computer
Z-0565-2013 | 63714 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators are anesthesia gas machine. |  |  | November 02, 2012 | 2012 | Del Mar Reynolds Medical, Ltd. | The BleaseFocus Anesthesia Workstation and 700/900 Series Ventilators were shipped with manuals that describe a secondary O2 Supply Failure alarm which has not been incorporated into the system. The manuals were prematurely updated for a feature that is not yet available. | Labeling design | Spacelabs has sent the Urgent -Medical Device Correction-Spacelabs Healthcare BleaseFocus Anesthesia Workstation Blease 700/900 Series Ventilators letter, dated 16 November 2012, to US customers.On November 26, 2012, a customer letter (translated as necessary) wasemailed to all international subsidiaries and distributors of record.Customers are advised that the primary alarm will sound only in the event of an O2 supply failure and they will be provided with the corrected user and service manuals that do not reference a secondary O2 supply failure alarm.Customers with questions can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 3124 | 3 in US and 3121 Outside US | Worldwide Distribution--USA (nationwide) including states of North Carolina and Oklahoma and the countries of Algeria, Argentina, Bahrain, Bangladesh, Bolivia, Chile, China, Columbia, Dominican Republic, Ecuador, Egypt, Finland, Ghana, Great Britian, Guatemala, Honduras, Hong Kong, India, Indonesia, Iraq, Jamaica, Jordan, Lebanon, Libya, Malaysia, Mauritius, Mexico, Moldova, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Pakistan, Palestine, Paraguay, Peru, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sudan, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates , Uruguay, Venezuela, and Vietnam. | Terminated | July 11, 2013 | 251 | 3124 | Not_Computer
Z-0564-2013 | 62394 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Case Carton: |  |  | October 02, 2012 | 2012 | Becton Dickinson &amp; Company | BD received a complaint report for mixed syringe tips (oral and non-oral syringe tips) within the same lot. | Mixed-up of materials/components | Becton Dickinson sent a Urgent Product Recall letter/Customer Packing list dated October 2, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to review inventory and determine if they have any of the affected product. If so, it should be removed from inventory. The enclosed customer list should be completed as well. Any questions please call 1-201-847-4267. |  | 2786000 | 2,786,000 units | (USA) nationwide | Open, Classified |  |  | 2786000 | Not_Computer
Z-0593-2013 | 63057 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Bausch + Lomb Stellaris PC Vision Enhancement System, Posterior Fluidics Module (PFM). The system is intended for the emulsification and removal of cataracts, anterior, and posterior segment vitrectomy. |  |  | August 31, 2012 | 2012 | Bausch &amp; Lomb Inc | Posterior fluidics modules installed in the system may need to be recalibrated. | Process control | "Urgent Product Recall" notification letters dated August 31, 2012 were issued via first class mail, e-mail, and/or phone calls to affected customers. The product issue and the field correction was provided. Customers in the US may contact 1-800 338-2020 for questions. |  | 124 | 124 units | Worldwide Distribution was made in the U.S., Canada, Australia, Singapore, Korea, Switzerland, France, Germany, United Kingdom, Portugal, Italy, Spain, Ireland, Africa, UAE, Jordan, Turkey, Ukraine, and Lebanon.There was no government or military distribution. | Terminated | July 11, 2013 | 314 | 124 | Not_Computer
Z-0572-2013 | 63784 | No | CEP | Luteinizing hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Access hLH assay is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of luteinizing hormone (LH) levels in human serum and plasma using the Access Immunoassay Systems. |  |  | November 05, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter Inc. (BEC) has demonstrated that real time stability testing of hLH Calibrator of certain lots have failed the 13-month and 14-month time points. Per BEC's real time stability procedure, in order to verify expiration dating, one additional month is needed beyond the claimed expiration period. Therefore, these lots did not meet BEC's procedural requirements for a 12-month shelf-lif | Nonconforming Material/Component | Beckman Coulter Inc. (BEC) sent an "URGENT FIELD SAFETY NOTICE" dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Per BEC's real time stability procedure, in order to verify expiration dating, one additional month is needed beyond the claimed expiration period. Therefore, these the affected lots did not meet BEC's procedural requirements for a 12-month shelf-life claim.Contact the firm at 952-448-4848 for assistance with this recall. |  | 14409 | 14409 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Djibouti, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Georgia, Germany, Ghana, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Macao, Malaysia, Mexico, Moldova Republic, Morocco, Myanmar, Namibia, Netherlands, Netherlands Antilles, New Zealand, Nigeria, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, South Africa, Spain, Swaziland, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, Yemen, Zambia, and Zimbabwe | Terminated | May 14, 2014 | 555 | 14409 | Not_Computer
Z-0607-2013 | 63667 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cervix Applicator Set, part number GM11004310, a component used with the tube for segments and cervical sleeves |  |  | September 28, 2012 | 2012 | Varian Medical Systems, Inc. | The guide tube which is part of the Segmented Cervix Applicator Set is difficult to disconnect from cervical sleeve. | Device Design | The firm, Varian Medical Systems, sent an "URGENT MEDICAL DEVICE CORRECTION - URGENT FIELD SAFETY NOTICE" letter dated September 28, 2012 to all affected consignees/Customers. The letter identified the affected product, description of problem and recommended user actions to be taken. The consignees/customers were instructed to immediately remove all affected guiding tubes from clinical use, and sequester the components for pending return to Varian Brachy Therapy. Consignees/Customers were requested to complete the attached proof of notification form and return it to Varian Medical Systems.For questions contact-AT&T; international direct dialing network (http://www.business.att.com/bt/dial_guide.isp. Then after hearing the AT&T; signature noise, and. "Please enter the number you are calling", dial 800-360-7909; UK:+44 1293-601-327 or US:+1-434-979-1540. To submit a support request or general inquiry, compose an email with a blank subject line to: brachyhelp@varian.com. |  | 15 | 15 units | Worldwide Distribution - U.S. state of Georgia and the countries of Canada, Germany, India, Poland, and Thailand. | Terminated | April 02, 2013 | 186 | 15 | Not_Computer
Z-0577-2013 | 63545 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SOL SYS 8IN CALC 1.5/18.0 |  |  | December 03, 2012 | 2012 | DePuy Orthopaedics, Inc. | An Engineering Analysis has determined the Solution Calcar Stem has the potential to result in impingement with a variety of heads. This impingement may not allow the stem and head to lock as intended. The amount of possible impingement varies depending on the femoral head. Worst case condition: Calcar stem with -2mm offset articuleze maximum of 0.937mm radial impingement. It is evident that a | Device Design | Deputh Orthopaedics sent an Urgent Information - Medical Device Recall Notice dated December , 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The notification letter instructed consignees to please cease further distribution or use of products immediately. The purpose of the communication was to inform them of this recall and request acknowledgement of receipt of this letter by their facility. Reconciliation forms should be returned to your DePuy Sales Representative or faxed to the following Fax Number: 574-371-4939.For questions regarding this recall call 574-267-8143. |  | 1732 | 1732 total product codes | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR | Terminated | November 13, 2013 | 345 | 1732 | Not_Computer
Z-0571-2013 | 63788 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | SpineSix consists of a motorized treatment table that is controlled by an integral computer. The system is operated by means of 12-inch Color LCD touch screen that displays treatment time, positioning and magnitude and direction of motion. |  |  | November 12, 2012 | 2012 | BioMotion Medical Systems, LLC | BioMotion is issuing a field correction for all models of the SpineSix (SpineSix01 and SpineSix02; serial numbers 1001 through 1054) to rectify all regulatory deficiencies involving the former owners of the SpineSix product and concurrently preparing the submission of a 510(k) application in order to receive premarket clearance for SpineSix. | No Marketing Application | BioMotion issued a field correction letter dated October 18, 2012 for all models of the SpineSix (SpineSix01 and SpineSix02; serial numbers 1001 through 1054) to all affected customers. The letter identified the product, problem, and actions to be taken. BioMotion is preparing the submission of a 510(k) application in order to receive premarket clearance for SpineSix. Contact the firm at 772-405-7899 for questions concerning this letter. |  | 28 | 28 units | Distributed in Florida, Georgia, Tennessee, New Jersey, New York, Pennsylvania, and Wisconsin. | Terminated | November 05, 2013 | 358 | 28 | Not_Computer
Z-0596-2013 | 61734 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medrad(R) Continuum MR Infusion System Standard Administration Kit |  |  | April 18, 2012 | 2012 | Medrad Inc | Bayer HealthCare is recalling all MEDRAD(R) Continuum MF Infusion System tubing sets (macrobore and microbore), catalog numbers MIK200A, MIL 200B, 3015153, 3015155, 3015156, 3015157, 3015158, 3015159, and 3015161. The firm is initiating this recall because this tubing exhibits a variation that, when used in conjunction with any Continuum pump, may result in flow rate accuracy performance issues. | Nonconforming Material/Component | The firm, Bayer HealthCare, sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 17, 2012 via the USPS certified mail to its consignees/customers. The firm expanded the recall on May 11, 2012 and sent a second letter entitled "EXPANDED URGENT MEDICAL DEVICE RECALL" to its affected consignees/customers notifying them of the additional issue. The letters described the product, problem and actions to be taken. The customers were instructed to return pump now; contact customer support at 1-877-229-3767 for a return goods authorization (RGA); complete and return Response Form, include RGA numbers, via Fax to Bayer Compliance at 412-4060942, and/or wait to return the pump when they begin contacting customers with this subset of pumps in August to make arrangements for a replacement pump.If you have questions, please contact our customer support team at 1-877-229-3767 or visit www.medrad.com for more information. |  | 1229 | 1,229 Pumps total | Worldwide distribution: USA (nationwide) and countries of: Australia, Bermuda, Brazil, Canada, China, Chile, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Great Britain, Guam, Hong Kong, India, Israel, Italy, Japan, Korea, Lebanon, Mexico, Malaysia, Netherland, Poland, Portugal, Puerto Rico, Sweden, Singapore, Taiwan, Thailand, and Venezula. | Terminated | August 17, 2015 | 1216 | 1229 | Not_Computer
Z-0608-2013 | 37530 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Qualicheck 1+, Aqueous Quality Control, Level 2, S7140, REF 944-042. |  |  | June 06, 2006 | 2006 | Radiometer America Inc | The barcode for the ABL77 on the insert incorrectly enters a control range for pH of 7.90-7.95 into the analyzer. The insert correctly states that the control range is 7.40-7.45. | Process control | Radiometer sent Field Action Notes regarding the affected product. The note identified the product, the problem and the action to be taken by the customer. Instead of scanning the barcorde on the insert to enter the QC data into the anyalyzer, customers were instructed to manually key in the QUALICHECK type, lot number, and correct control range.For questions regarding this recall call 440-871-8900. |  | 840 | 840 Ampules | Distributed to KY, NC, and NY. | Terminated | February 05, 2013 | 2436 | 840 | Not_Computer
Z-0525-2013 | 63729 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Praxair Vantage Grab n Go - Gas Cylinder pressure regulator |  |  | November 16, 2012 | 2012 | Praxair Inc. | Isolated incidents of ignition inside Grab n' Go cylinders that had been knocked over or otherwise subject to significant physical impact. | Device Design | The firm, PRAXAIR, sent a "MEDICAL DEVICE CORRECTION" letter dated November 26, 2012 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Grab'n Go operating instructions provided in the letter which include: storing the Grab'n Go systems indoors in a safe area; never store a system in an elevated location; handle systems safely and avoid jarring or dropping the system and locate at the following web link: www.praxair.com/GNGinstructions. The customers were also instructed to ensure that all of the Grab'n Go units that were sent by Praxair are sent to Praxair for servicing and refilling after use, in accordance with their normal process.If you have any questions, contact your sales representative at 630-320-4452 or email: Thelma_Brantley@Praxair.com. |  | 320451 | 320,451 units (255,690 US ; 55,884 Canada; 8877 OUS) | Worldwide distribution: USA (nationwide) and countries of: Brazil, Canada, China, Columbia, Costa Rica, India, Paraguay and Uruguay. | Terminated | March 24, 2014 | 493 | 320451 | Not_Computer
Z-0479-2013 | 63394 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System, Extended Brilliance Workspace (EBW), Model #728260 Philips Healthcare, Highland Heights, OH |  |  | September 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips is issued a Field Change Order to update Extended Brilliance Workspace (EBW) systems to software version 4.5.5. The update will provide a number of improvements to address nonhazardous reliability issues which will improve the overall quality and functionality of the EBW. | Software design | Philips sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Philips is issuing a system update to all of the Extended Brilliance Workspace (EBW) systems currently running software version 4.5.4. The update will provide a number of improvements to the overall quality and functionality of the EBW systems.If you need further information or support concerning this issue, please contact your local Philips representative. In the United States and Canada, please call the Philips Healthcare Customer Care Solutions Center at 1-800-722-9377, and follow the recorded menu options to reach a Customer Solutions Engineer; in all other countries please dial your local Philips Healthcare office. |  | 265 | 265 units | Worldwide Distribution--US (nationwide) including the states of AR, OH, OR, PA, TX VA &amp; WI., and the countries of Algeria, Australia, Brazil, China, Ecuador, Finland, France, Germany, Hong Kong, India, Italy, Lebanon, Libyan Arab Jamahiriya, Russia, South Korea &amp; United Kingdom | Terminated | December 24, 2013 | 452 | 265 | Computer
Z-0611-2013 | 63676 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster MobiCath Bi-Directional Guiding Sheath. Small Curve catalog no. D140010, Large Curve catalog no. D140011. The MobiCath Bi-Directional Guiding Sheath is a flexible tipped percutaneous catheter introducer (or guiding sheath) designed for gaining access to the coronary systems. The device features adjustable tip geometry through use of a rotating handle to deflect the catheter. This catheter is assembled with a hemostasis valve and a sideport infusion line. |  |  | November 12, 2012 | 2012 | Greatbatch Medical | Greatbatch Medical observed some anomalies of the inner lumen of the MobiCath Bi-Directional Guiding Sheaths upon implementation of a new visual inspection tool. The anomalies may include loose, string-like liner material, liner abrasion (fraying), and/or attached string-like liner material (from scratches or damage on the inner wall of sheath lumen). | Nonconforming Material/Component | Biosense Webster sent an "Urgent Field Safety Notice" dated November 12, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Voluntary Field Removal Certification Form was attached for customers to complete and return. Contact the firm at 800-729-9010 for questions related to this issue. |  | 10201 | 10,201 units | Worldwide Distribution-USA including the states of AL, AK, AZ, ALR, CA, CO, CT, DC, FL, GA, HI, IL IN, IA, KY, MD, MA, MI, MN, MO, NE, NV,NM, NJ, NH, NY, NC, OH, OK OR, PA, RI,SC,SD, TN, TX, UT,VT, VA, WA, and WI, and the countries of AUSTRIA, BELGIUM, SWITZERLAND, GERMANY, DENMARK, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HUNGARY, IRELAND, ISRAEL, ITALY, THE NETHERLAND, NORWAY, SWEDEN, AUSTRALIA, and CANADA | Terminated | May 15, 2014 | 549 | 10201 | Not_Computer
Z-0610-2013 | 63457 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics; Stryker T2 Humerus System |  |  | August 06, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | During the manufacturing of the affected lot a manufacturing error has led to embrittlement of the tube material of the T2 Humerus Teflon Tube. | Nonconforming Material/Component | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter &amp; Acknowledgement form dated August 7, 2012 via FedEx Return Receipt to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to inform their users of this Medical Device Removal and pass this notice to all those individuals who need to be aware within their organization; return all affected products available at their location to: Stryker Osteosynthesis, 325 Corporate Drive, Mahwah, NJ 07430 REF: PFA#2012-124 or contact Stryker customer service; complete and sign the enclosed Product Recall Acknowledgement Form and Fax to 1-855-251-3635 or email to Recall Coordinator at janice.pevide@stryker.com and keep a copy of their records.If you have any questions, contact the Manager, Regulatory Compliance, at 1-201-972-2100 or email: colleen.omeara@stryker.com. |  | 100 | 100 units (73 units US; 27 units OUS) | Worldwide distribution: USA (nationwide) including states of: AL, CO, CT, IL, IN, MO, MS, NC, NJ, NY, OR,and TX; and countries of: Australia, Canada, France, Spain and United Kingdom. | Terminated | September 11, 2013 | 401 | 100 | Not_Computer
Z-0609-2013 | 63803 | [6 empty] | Diamigo i-Phone App |  |  | September 21, 2012 | 2012 | sanofi-aventis US, Inc. | Diamigo, a software application, was released through the i-Phone Global Store which inadvertently allowed global access to the application as opposed to the original intended restriction of access to Brazil. | No Marketing Application | Apple sent a message ("push notification") on September 21, 2012 to its customers. The message describes the product, problem and actions to be taken. The customers were instructed not to use the app to calculate insulin does and should delete if from their device(s). If you have additional questions about the app or this message, contact Sanofi's offices in your country. If you are in the U.S. contact Sanofi US Medical Information at 1-800-633-1610, for Brazil please contact 0800 703 0080. |  | 1600 | 1600 downloads (140 US; 1460 Foreign) | Worldwide distribution. | Terminated | July 22, 2015 | 1034 | 1600 | Computer
Z-0613-2013 | 61212 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Total T4 Calibrators, Part Number: 33805. |  |  | July 27, 2011 | 2011 | Beckman Coulter Inc. | A recall communication was initiated because Beckman has identified that Access Total T4 calibrator lots 019818, 021654, and 024072 do not meet their 12 month expiration date claim. | Nonconforming Material/Component | A recall communication was initiated on July 27, 2011 with Beckman forwarding a Product Corrective Action (PCA) letter with attached PCA response form to all their customers who purchased the Access Immunoassay Systems Total T4 Calibrators. The letter provides the customers with an explanation of the problems identified and actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers can call for product replacement in the US, by contacting Client Services at (800) 526-3821, option 1. Customers in Canada can contact Customer service at (800) 463-7828. Customers with questions regarding this notice can contact Customer Support Center at http:www.beckmancoulter.com or call (800) 854-3633 in the US and Canada. |  | 6634 | 6,634 units total (2,742 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Egypt, El Salvador, Ethiopia, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Kazakhstan, Lebanon, Libyan Arab Jamahiriya, Liechtenstein, Macao, Malaysia, Mexico, Netherlands, Nigeria, Pakistan, Panama, Philippines, Puerto Rico, Romania, Russian Federation, Serbia, Singapore, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Viet Nam, and Zambia. | Terminated | December 26, 2012 | 518 | 6634 | Not_Computer
Z-0594-2013 | 63683 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Zimmer Spine Ardis Inserter, Rx, non-sterile |  |  | November 12, 2012 | 2012 | Zimmer, Inc. | Zimmer Spine has received reports of the PEEK Ardis Interbody Spacer breaking into fragments when too much lateral force is applied to the Ardis Inserter. If the PEEK Ardis Interbody Spacer breaks during implantation, surgical delays of up to 60 minutes could occur. Sharp edges of the broken implant may contact and damage adjacent structures, requiring surgical intervention. Observed health risks | Device Design | Zimmer Spine, sent a "MEDICAL DEVICE CORRECTION" memo dated November 12, 2012 to all consignees / customers describing product, problem and actions to be taken including revised surgical technique. Beginning the week of December 17, 2012, Zimmer began the immediate retrieval of the affected units as a Class I recall. December 20, 2012 nation wide press outlining the recall was issued. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer Spine by calling 866-774-6368 between the hours of 8AM and 5PM CST or by email at qualitycompliance@zimmer.com. |  | 315 | 315 (172 in USA) | Worldwide distribution: USA (nationwide) including states of: AR, AZ, CA, CO, CT, DC, FL, GA, ID, LA, MD, MI, MN, MO, NC, NV, NY, OH, OK, PA, TN, TX, WI, and WV; and countries of: Australia, El Salvador, France and Singapore. | Terminated | February 25, 2014 | 470 | 315 | Not_Computer
Z-0612-2013 | 63632 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay System, |  |  | November 01, 2012 | 2012 | Beckman Coulter Inc. | It has been determined that, in some cases,pulley drive components have broken parts or have become loose, mispositioned, ordisconnected from their drive shafts. When these defects occur, the pulley may not turn at all, ormay turn at an incorrect speed, resulting in inadequate washing or mixing of the contents of thereaction vessel. | Device Design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 1, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to complete and return an enclosed Response Form. Contact Technical Support for questions regarding this notice. |  | 6858 | 6,858 units | Worldwide Distribution-USA (nationwide) and the country of Canada. | Terminated | December 23, 2013 | 417 | 6858 | Not_Computer
Z-0614-2013 | 63814 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS AmpliPrep Instrument (CAP) with AMPLILINK (AL) Software v.3.3.5 |  |  | October 11, 2012 | 2012 | Roche Molecular Systems, Inc. | Certain COBAS AmpliPrep instruments using AMPILINK software v.3.3.4 had the sample clot detection flagging feature disabled during quality control release resulting in samples that may not be flagged as having clots. | Device Design | Roche sent an Urgent Medical Device Correction notification/Fax Back forms dated 10/11/12 to potentially affected customers. The letter identified the affected product, the issue, clinical significance, workaround, and required actions, along with contact information. Customers are to use the workaround until their Roche FSE has performed a Master Initialization on their affected instruments. Also, the letter should be filed for future reference. Questions should be directed to Roche Molecular Diagnostics Technical Support at 1-800-526-1247. |  | 7 | 7 units | Distributed in the states of GA, CA, NC, and MN. | Terminated | September 11, 2013 | 335 | 7 | Computer
Z-0606-2013 | 63123 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Giraffe and Panda T-Piece Resuscitation System labeled ***T-Piece Neonatal Patient Circuit Kit Disposable with Size 1 Mask***M1091335 QTY: 10 T-Piece Disposable Circuit, QTY: 10, M1091316 T-Piece Disposable Circuit Kit with Mask Size O; M1091365 QTY: 10 T-Piece Disposable Circuit Kit with Mask, Size 1, Lot 1981713 *** Ohmeda Medical, A Division of Datex-Ohmeda, Inc. A General Electric Company Laurel, MD 20723 USA MADE IN USA*** |  |  | June 26, 2012 | 2012 | GE Healthcare, LLC | Disposable T-Piece circuits used with resuscitation systems are unable to achieve maximum Peak Inspiratory Pressure (PIP) during the pre-use check out procedure. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter, dated June 26, 2012, with a return receipt to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer along with the product correction. Required Action for Users: Inspect existing stock of T-piece circuits to look for the Red inspection stamp, the T-pieces with the blue stamp are acceptable for clinical use. Separate affected circuits (with Red inspection stamp) from unaffected stock (with Blue inspection stamp). Destroy affected circuits that have the Red inspection stamp. Fill out the customer acknowledgement form and fax or forward to the GE contact address on the form. Further, it is recommended that healthcare professionals continue to emphasize the recommended pre-use checkout practices, and do not use a T-piece patient circuit that does not reach the maximum PIP level of 45 +/- 5 cmH20. Clinicians should always use the built-in Airway Pressure Manometer to verify the PIP value. For questions or concerns regarding this notification, customers should call the following phone number: 1-800-345-2700. |  | 209804 | 209,804 circuits | Worldwide Distribution - USA, including DC and the states of AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IN, IL, KY, KS, LA, MA, MD, ME, MI, MN, MO,MS, MT, NC, NE, NJ, NM, NY, NV, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Albania, Australia, Austria, Bahrain, Bangladesh, Belgium, Canada, China, Czech Rep., Egypt, Finland, France, Germany, Great Britain, Greece, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, S. Korea, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, and Arab Emirates, | Terminated | April 14, 2014 | 657 | 209804 | Not_Computer
Z-0616-2013 | 63497 | [6 empty] | Brookstone Heated Body Bean, electrically heated |  |  | November 02, 2012 | 2012 | Brookstone Company, Inc. | Additional Warning Statement to prevent burns to the skin | Labeling design | Brookstone sent a "MEDICAL DEVICE CORRECTION" letter dated December 10, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Recall Information is on the website link from the homepage, www. Brookstone.com. Recall posted: http://www.brookstone.com/recall-information. Questions call Brookstone Customer Care at 800.846.3000 or visit any Brookstone Store. |  | 109956 | 109,956 units | Nationwide Distribution. | Terminated | June 13, 2013 | 223 | 109956 | Not_Computer
Z-0615-2013 | 63857 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | Flowtron Trio DVT Pump; an Rx intermittent pneumatic compression deep vein thrombosis pump; ArjoHuntleigh AB, Eslov, Sweden; Model 512003; a non-invasive prophylaxis system for reducing the incidence of deep vein thrombosis. |  |  | December 10, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | The Flowtron Trio DVT Pump has been marketed without proper FDA Approval. | No Marketing Application | ArjoHuntleigh sent an Urgent Field Correction Recall letters dated December 10, 2012, to all affected customers, informing the accounts that the pumps were marketed without proper FDA approval, and that ArjoHuntleigh is removing the pumps from the market and replacing them with a suitable alternative. The accounts were requested to disseminate the information to all organizations where the devices may have been transferred, and to complete the Customer Response Form and return it to ArjoHuntleigh by mail, fax or e-mail. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 2859 | 2,859 unts | USA (nationwide) including the states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Dakota, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin and Wyoming. | Open, Classified |  |  | 2859 | Not_Computer
Z-0619-2013 | 63825 | No | JIQ | Urinary protein or albumin (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Analyzing system consists of a small and portable analyzer and plastic microcuvettes. The microcuvette contains reagents deposited on its inner walls. The urine sample is drawn into the cavity by capillary action. The filled cuvette is inserted into the device where the contents of the cuvette are mixed through vibration. Within 90 seconds, the immunochemical reaction is completed and the turbidity is measured photometrically at 610 nm. The albumin concentration is proportional to the turbidity. When the end point is reached, the result is displayed in mg/L. |  |  | November 07, 2012 | 2012 | HemoCue AB | Incorrect version of Quick Reference Guide (QRG) was delivered with products in shipment made from April1 st 2011 and onwards. In addition, a warranty statement was missing in the analyzer package. | Labeling mix-ups | HemoCue sent an "IMPORTANT MEDICAL DEVICE CORRECTION NOTIFICATION" letter dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Correction Notification Verification Form was included for customers to complete and return. Contact the firm at 562-668-5741 for questions regarding this notification. |  | 48 | 48 affected analyzers | Distributed in the states of FL, MA, CA, OR, NY, NC, MO, ID, IN, KY, PA, TX, OK, IL, GA, KS, and UT. | Terminated | March 14, 2013 | 127 | 48 | Computer
Z-0618-2013 | 63769 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack with or without coating ( Carmeda BioActive Surfact, Carmeda BioActive Surfact and Trillium Biosurface,or Balance Biosurface) with the following Model numbers: |  |  | November 19, 2012 | 2012 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is initiating an Urgent Medical Device Customer Notification. We have confirmed that certain Medtronic Intersept Tubing Packs contain packaging trays that are susceptible to damage that can compromise product sterility. | Package design/selection | Medtronic sent a "Urgent Medical Device Customer Notification" letter dated November 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately identify and quarantine all affected product and to return affected product, but if they have concerns about replacement than they recommend to inspect and return any damaged product..For further questions please contact your local Medtronic sales representative or Lifeline Technical Services at 1-877-526-7890. Outside the US, please call1+763-526-7890. We apologize for any inconveniences this issue may have caused you or your institution. |  | 1552 | 1552 | Worldwide Distribution--USA (nationwide) including the states of AZ, CA, CT, FL, GA, IL IA, MD, MA, MI, MS, NJ, NY, OH, OK, SC, TN, TX and WI. and the countries of CANADA and MALAYSIA. | Terminated | April 04, 2013 | 136 | 1552 | Not_Computer
Z-0620-2013 | 63915 | No | DSZ | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Companion External Battery, Part Number: 293001-001 |  |  | December 05, 2012 | 2012 | SynCardia Systems Inc. | Syncardia Systems is recalling the Companion External Batteries used with the Companion 2 Driver System because they may potentially be at risk of thermal events and/or other physical damage such as melting, charring and/or burning. | Process control | Syncardia Systems sent a" Urgent Field Safety Notice" dated December 10, 2012, to all affected customers. The letter identified the product the problem and what action the customer needs to take.. The letter informed the customers that there were 3 reported thermal battery events (2 of which occurred on-site at either SynCardia or at the battery manufacturer, Dow Kokam, and one incident that occurred in the field) which caused physical damage to the batteries including burning and/or charring. The batteries were not in use with a patient at time of the reported issue, and there were no reports of any adverse impact to any patient or user. Customers are instructed to follow the directions on the attached Return page, and return their customer response form to SynCardia as soon as possible.Customers with questions are instructed to call (520) 545-1234 or (520) 440-7593. |  | 81 | 81 units | Worldwide Distribution -- US (nationwide) | Terminated | June 07, 2013 | 184 | 81 | Not_Computer
Z-0617-2013 | 63855 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Fast-Cath Transseptal Guiding Introducer, 90, 8F, REF 406586, Length 60 cm, Sterile EO, Rx only, St. Jude Medical, 14901 DeVeau Place, Minnetonka, MN 55343 |  |  | March 28, 2012 | 2012 | St Jude Medical Inc | St. Jude Medical is recalling a single batch of Fast Cath Transseptal Guiding Introducers which was packaged as a 180 degree curve style introducer instead of a 90 degree curve style introducer as indicated on the labeling. | Process control | St Jude Medical hand delivered the" Product Recall" letter dated March 27, 2012, to the affected customers on April 5, 2012, and March 28, 2012. The letter identified the product, the problem and the action needed to be taken by the customer.The returned devices and those not distributed were quarantined at the SJM Minnetonka distribution center until their destruction.Customers were requested to to complete and return the reconciliation form.Further questions please call (651) 756-6526 |  | 26 | 26 | US Distribution only including the states of FL, KY and OK. | Terminated | July 12, 2013 | 471 | 26 | Not_Computer
Z-0623-2013 | 63862 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Space Pump Pole Clamps for the Infusomat Space Pump, model number 8713130. |  |  | March 23, 2012 | 2012 | B. Braun Medical, Inc. | Firm has become aware of the potential for breakage of the pole clamp locking grids, or rasters, in the Space Pole Clamps. | Device Design | B. Braun Medical, Inc sent an Urgent Medical Device letter dated March 23, 2012 , to all affected customers. The letter identified the product the problem and the action needed by the customer to take.B. Braun has implemented an improvement to the pump locking mechanism of t he older style pole clamps and will be upgrading any older style pole clamps that may remain within your facility.Please return the completed form from attachment 1 in the self-addressed envelope provided or fax the form to (610) 849-1197 within two (2) weeks of receipt\.If you require any additional information or have any questions, comments or concerns please contact your B. Braun Medical Customer Service at 800-627-7867. |  | 13078 | 13,078 | Worldwide Distribution--US (nationwide) including the states of AL, AZ, NC and TN. | Terminated | February 02, 2015 | 1046 | 13078 | Not_Computer
Z-0626-2013 | 63772 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Hyperbaric Single Channel Infusion Pumps; Hospira, Inc., Lake Forest, IL 60045; list number 11005 Product Usage: The Plum A+ Hyperbaric Infusion System is exclusively designed for hyperbaric therapy in monospace and multiplace chambers. |  |  | December 07, 2012 | 2012 | Hospira Inc. | The door roller assembly on the Plum A+/A+3 infuser family of infusers has the potential to break which can lead to possible unrestricted flow and/or over-delivery during the removal of the IV administration set's cassette from the pump. | Device Design | Hospira Inc. sent "URGENT DEVICE FIELD CORRECTION" letters, dated December 7, 2012, to all affected customers. The letters identified the products, problem and action(s) to be taken until doors have been replaced. The customers were instructed to read the information provided carefully, including instructions to inspect the door roller assembly prior to loading a cassette for signs of damage. Customers are to complete the attached Reply Form and return it to the firm via fax at 1-877-650-8359. Customers were instructed to notify their consignees if they have further distributed the affected product. If customers require additional assistance, they are to contact Stericycle at 1-877-650-0362.. |  | 180 | 180 units | Worldwide Distribution -- USA, including Puerto Rico, and the countries of Austria, Australia, Argentina, Bahrain, Belgium, Brazil, Canada, Chile, Colombia, Denmark, Finland, France, Greece, Hong Kong, Indonesia, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Malaysia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, United Arab Emirates, and Vietnam. | Open, Classified |  |  | 326612 | Not_Computer
Z-0622-2013 | 62413 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Standard Slingbar 450. |  |  | December 06, 2012 | 2012 | Hill-Rom, Inc. | Liko/Hill-Rom has received reports from facilities which indicate that serious injuries have been received related to the Standard Sling bar 450. There have been 6 complaints of injury related to the product over the past 32 years (6 out of 838,687,856 estimated patient uses). One of the injuries was associated with a patient during use, four injuries occurred when the slingbar was not in use wher | Device Design | Hill-rom sent an URGENT MEDICAL DEVICE RECALL notification letter dated December 6, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.. Customers were requested to please read, fill in completely, and return the Response Form to Hill-Rom immediately. This information will be compiled and reported to the applicable regulatory governing bodies. If we do not receive your response we will be required to contact you asking for your response.Continue to follow safe lifting practices, and do not allow unauthorized persons, especially children, to play around or operate the lift.If you have any questions concerning this Customer Safety Notification, please contact Hill-Rom Technical Support at 800-445-3720. |  | 261 | 261 units | Worldwide Distribution -- US (nationwide) including the states of AR, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, MA, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OK, OR, PA, TN, TX, UT, WA, WV, and WY., and the country Canada,We have not yet defined our European or International facilities at this time. | Terminated | November 06, 2014 | 700 | 261 | Not_Computer
Z-0630-2013 | 63616 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM Central Station |  |  | May 12, 2011 | 2011 | Mindray DS USA, Inc. d.b.a. Mindray North America | There is a software anomaly with the DPM Central Monitoring System where the trend data of one patient maybe replaced with the trend data of a second patient. | Software design | Mindray sent a Product Corrective Action letter dated May 12, 2012, to all affected customers via certified mail return receipt requested to two medical facilities. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the issue would be corrected through a softweare upgrade of the affected product. A Mindray Service Representative would contact their facility to arrange for the software upgrade at no cost to them.For questions regarding this recall call 1-800-288-2121. |  | 2 | two units | Nationwide Distribution including OH and GA | Terminated | February 19, 2013 | 649 | 2 | Computer
Z-0629-2013 | 63602 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Medical Surgical Bed, Model 3005S3EX |  |  | December 18, 2012 | 2012 | Stryker Medical Division of Stryker Corporation | Stryker Medical Model 3002S3EX and Model 3005S3EX beds built between March 15, 2012 and May 23, 2012. may have been built with brake rings that do not meet specifications, negatively impacting the holding force of the brakes. | Process control | Stryker sent an Urgent Medical Device Correction letterdated December 18, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate the affected product and ensure that the brakes are holding properly. Consignees were directed to contact Stryker (1-800-327-0770, option 2) to make arrangements to replace all brake rings and continue to routinely inspect brake functionality. Customers were asked to Fax (269-488-8691) the enclosed acknowledgment and Stryker field service will contact customers to schedule service. Consignees were directed to forward a copy of this notice to the new users and advise us of their new location.For questions regarding this recall call 269-488-8691. |  | 1004 | 1004 | Worldwide Distribution - Nationwide and Canada | Terminated | August 30, 2013 | 255 | 1013 | Not_Computer
Z-0634-2013 | 63746 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp Instrument; an automated fluid handling system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; Abbott m2000sp Instrument; an automated system for performing sample preparation for nucleic acid testing; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 09K14 |  |  | December 04, 2012 | 2012 | Abbott Molecular | An issue with the m2000sp Instrument software has been identified whereby, after a series of steps, if changes are made to the control or calibrator information on the Sample Extraction Assay Details screen, the changes are not saved. | Software design | Abbott Molecular Inc. sent a Field Correction Recall/Urgent Field Safety Notice letter dated December 4, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer was instructed to follow the instructions identified until an upgrade to their m2000sp software can be installed. Abbott Molecular estimates that this software version will be available the first quarter of 2013. They were requested to complete the enclosed reply sheet and fax it to (847) 775-6728 at prior to December 18, 2012 to acknowledge receipt and understanding of the letter.For questions customers were instructed to contact their local Abbott Molecular representative.For questions regarding this recall call 224-361-7274. |  | 215 | 215 units - U.S., 693 units - ROW | Worldwide Distribution - USA (nationwide) and internationally to Algeria, Argentina, Australia, Austria, Azerbaijan, Belgium, Bolivia, Brazil, Burkina Faso, Burundi, Cameroon, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ethiopia, Estonia, Finland, France, Gaza &amp; Jericho, Germany, Greece, Guinea, Haiti, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Mali, Mexico, Mozambique, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam and Zambia | Terminated | March 08, 2016 | 1190 | 215 | Computer
Z-0633-2013 | 63838 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PENTRA C200, Part Number: A11A01629. |  |  | October 10, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA AST CP because during the AST application validation on the Pentra C200 using an Audit Systems linearity kit, it was determined that two higher values for AST (expected at 800 U/L and 1000 U/L) were recorded at 280 and 199 U/L without alarm. | Software design | Horiba Medical sent a Field Safety Notification and Acknowledgement Letter dated November 20, 2012, to all affected customers. The letter notified the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 calendar days via mail, email or fax. Customers with questions were instructed to contact their local representative.For questions regarding this recall call 714-273-9254. |  | 384 | 384 units | Nationwide Distribution including CA, OK, NY, FL, AZ, NY, VA, and TX. | Terminated | September 12, 2013 | 337 | 67270 | Not_Computer
Z-0621-2013 | 63794 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Fisher &amp; Paykel Healthcare reusable breathing circuit. |  |  | November 19, 2012 | 2012 | Fisher &amp; Paykel Healthcare, Ltd. | FPH has become aware that one batch of tubes supplied to FPH may contain hole damage. These tubes were incorporated by FPH into certain 900MR068 reusable circuits with lot numbers 110810 and 111020. If this hole damage is not detected during the standard leak test before patient use, it could potentially result in a gas leak in the breathing system, which may lead to a loss of pressure for the int | Nonconforming Material/Component | Fisher &amp; Paykel sent an Urgent Medical Device Recall letter dated November 19, 2012, with a Product Recall Response Form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to destroy and discard any affected circuits in their possession. Customers were instructed to complete and return the Product Recall Response form to the address indicated on the form, even if they have no affected product. Customers were asked to forward the recall notice to all those who need to be aware within their organization. Customers with questions should call 1-800-792-3912.For questions regarding this recall call 949-443-4095. |  | 200 | 200 units total, 26 units US | Worldwide Distribution - USA including IL, IN, AR and GA. | Terminated | July 11, 2013 | 234 | 200 | Not_Computer
Z-0631-2013 | 63719 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | AVAIRA (enfilcon A) Soft (Hydrophilic) Daily Wear Toric Contact Lenses |  |  | November 09, 2012 | 2012 | Cooper Vision Caribbean Corp. | The lenses exceeded the acceptance limit for silicone oil residual. | Material/Component Contamination | The recall was conducted to the wholesale/distributor and Eye Care Practitioner level. The company contacted the Authorized Distributors by telephone, followed by written notification dated November 9, 2012. All written notifications to first consignees, for both revenue and Fit Sets, were sent by overnight mail with signature required as proof of receipt and the envelopes were marked "Urgent". The notifications described the issue and provided recommended actions. Customers may contact 800-341-2020 if questions. |  | 246 | 246 | US Nationwide, including Puerto Rico, and Canada. | Terminated | May 09, 2014 | 546 | 246 | Not_Computer
Z-0627-2013 | 63878 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Yuno Extension Table Device |  |  | November 02, 2012 | 2012 | Maquet Inc. | There is a potential issue with the screw tension of the extension device 1433.62A1 used with the YUNO table. | Device Design | Maquet sent Urgent Product Field Correction letters/Field Correction Confirmation Form dated November 2, 2012 on November 15, 2012, via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a MAQUET authorized Service Technicians would contact them to arrange an appointment to assess their device and to replace the release lever at no cost to them at their facility. Customers were asked to completed and return the enclosed confirmation form, acknowledging their receipt and understanding of the communication, in the enclosed postage prepaid envelope, by e-mailing a scanned copy to FieldActions@maquet.com or via fax at 973-709-6508.For questions customers were instructed to call MAQUET Technical Support at 800-777-4222, Press option2, followed by option 2.For questions regarding this recall call 973-709-7660. |  | 63 | 63 devices | Worldwide Distribution - USA including AR, GA, SC, TX, TN and internationally to France, United Arab Emirates, Australia, Brazil, Canada, Switzerland, China, Germany, Spain, United Kingdom, Italy, Japan, Sweden, Thailand, Turkey, Serbia, and South Africa. | Terminated | August 21, 2013 | 292 | 63 | Not_Computer
Z-0642-2013 | 63030 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FRx 861304 Automated External Defibrillator Manufactured by Philips Medical Systems, Seattle, WA 98121 USA; and Laerdal branded FRx 861305 |  |  | July 26, 2012 | 2012 | Philips Medical Systems North America Co. Phillips | Philips determined that it is important to clarify information in the Owner's Manual and keep customers informed about the maintenance of their FRx (models 861304 and 861305), OnSite (models M5066A and M5067A), or HeartStart Home (model M5068A) automated external defibrillators (AEDs) shipped between December 2005 and July 2012. Philips had determined there is a need to emphasize the importance o | Nonconforming Material/Component | Philips Healthcare sent an Emergency Care and Resuscitation letter Customer Information Letter dated July 26, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their Philip's representative to obtain the tools and information they need to ensure that their device is functioning properly.For questions regarding this recall, please call AED Recall Support at 1-800-263-3342 (Option 6). |  | 201598 | 201,598 | Worldwide Distribution - USA (nationwide) and the countries of Japan and Malaysia. | Open, Classified |  |  | 814598 | Not_Computer
Z-0635-2013 | 63809 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Chek-Stix Positive and Negative Control Strips for Urine Chemistry |  |  | November 21, 2012 | 2012 | Siemens Healthcare Diagnostics Inc | When these lots of control material are reconstituted, the positive control solution leukocyte result may become negative if the solution is held for more than two hours above 25ÂżC (77ÂżF). | Nonconforming Material/Component | Siemens issued An Urgent Field Safety Notice via e-mail to all affected SiemensHealthcare Diagnostics Regional offices on November 21, 2012 both in the United States and foreign accounts. This notice informed customers of the issue with the CLINITEK AtlasÂż and Chek-StixÂż Positive Urinalysis Control Solutions When these lots of control material are reconstituted, the positive control solution leukocyte result may become negative if the solution is held for more than two hours above 25ÂżC (77ÂżF). You can use the reconstituted positive control solution for up to eight hours if the solution is held at 25ÂżC (77ÂżF) or lower. If the operating temperature is greater than 25ÂżC and lower than 30ÂżC (77ÂżF and 86ÂżF), the solution can be used for up to two hours.For questions regarding this recall call 781-269-3158. |  | 7865 | 7865 | Worldwide Distribution - USA (nationwide) and internationally to Canada, Aruba, Brazil, Denmark, India, Mexico, New Zealand, Singapore,Thailand, Bahamas, South Wales, and Hong Kong, | Terminated | November 06, 2014 | 715 | 37839 | Not_Computer
Z-0649-2013 | 63896 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | Sterile Rigid Tip Suction Wand |  |  | November 02, 2012 | 2012 | California Medical Laboratories Inc | California Medical Laboratories, Inc. has identified an issue with the Rigid Tip Suction Wand. Due to insufficient adhesive application, the distal tip has the potential to come loose from the wand and fall into the patient's chest cavity. | Nonconforming Material/Component | CalMed sent an Urgent Field Safety Notice dated November 2, 2012, to all affected customers by certified mail and additionally followed up via electronic mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Sorin Group USA's Customers Service Department at +1-800-650-2623 to return any unused product and order replacements. Customers were also asked to communicate the important safety information to all potential users within their organization or any other organization where the affected product may have been transferred and to maintain awareness of this communication for an appropriate period of time to ensure effectiveness of the corrective action. Customers were asked to return all unused product to CalMed Labs. Customers with questions were instructed to contact California Medical Laboratories at +1-714-556-7365.For questions regarding this recall call 714-556-7365. |  | 1430 | S112462=995 units &amp; S111480 1430 units | Worldwide Distribution - USA including AZ, FL, CA, LA, CO, and MS.Internationally to Canada, Israel, Greece, Hong Kong, India, Ireland, Jordan, Lebanon, Malaysia, Netherlands, Russia, Saudi Arabia, South Africa and Spain. | Terminated | March 04, 2013 | 122 | 357891 | Not_Computer
Z-0645-2013 | 63829 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA ML, Software Versions: V8.0.x and Below; V.9.0.1 and V9.0.2 |  |  | September 20, 2012 | 2012 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX PENTRA ML because they have confirmed an issue with the ABX PENTRA ML data management and validation system. An incorrect patient result is transmitted to the Laboratory Information System. This result is from a previous run of the same patient that is stored in the ML database. This issue occurs in a very rare sequence of specific events. The affected software | Software design | .A recall letter dated 11/09/2012 was sent to all their customers who received the ABX PENTRA ML products. The letter informed the customers of the problem identified and the action to be taken. Customers are instructed to complete and return the enclose response form within 10 days. Customers with questions about the letter are instructed to contact their local HORIBA Medical representative.Recall Letter, as produced per FCD-0377, Field Safety Notice &amp; Acknowledgement will be mailed via FedEx to the ten (10) customers noted on the list the week of November 12, 2012. Effectivity will be evaluated by receipt of the Field Safety Notice Acknowledgment Form. Follow-up communications with customers will be conducted, as needed. |  | 10 | 10 units | Nationwide in the US, including the states of CA, MI, IN, FL, and GA. Center Recall Depth is Retail. | Terminated | March 12, 2013 | 173 | 10 | Computer
Z-0644-2013 | 63842 | No | CGJ | Follicle-stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Laparoscope, general; plastic surgery |  |  | November 05, 2012 | 2012 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall on a specific lot number of its 12x150mm Kii shielded bladed access systems. This is due to a potential for the blade to remain exposed after shield deployment. Further use of any affected product should cease immediately.The affected model number is CTB71, and the affected lot is 1179107. | Nonconforming Material/Component | Applied Medical sent an Urgent Medical Device Recall to all affected customers in November 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product, complete the attached Recall Notification Confirmation Form to acknowledge the recall and indcated any product their facility is returning. The Recall confirmation form should be returned to Applied Medical by emailing to recall60494844#applied medical.com or fax to 949-713-8832. Return affecte4d product and a copy of the confirmation form to Applied Medical. Customers with questions were instructed to call 949-713-8662. For questions regarding this recall 949-713-8662. |  | 72 | 72 units | Worldwide Distribution - USA (nationwide) and France. | Terminated | January 16, 2013 | 72 | 72 | Not_Computer
Z-0648-2013 | 63852 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Immunoradiometric assay. |  |  | November 28, 2012 | 2012 | Siemens Healthcare Diagnostics | The recall was initiated because Siemens Healthcare Diagnostics confirmed an under-recovery of dose values with some vials of Coat-A-Consta (CAC) IRMA TSH calibrators TSI4 (level B), TSI5 (level C), TSI6 (level D), TSI8 (level F), and TSI9 (level G) lot 035. The affected calibrator vials can be identified by a darker yellow color than normal. The TSI3 (level A), TSI7 (level E), and TSIX (level H) | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated November 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product, complete and return the Urgent Medical Device Recall Effectiveness Check form included with the letter by fax to 302-631-7597. Customers were also instructed to determine their replacements needed, and to forward the recall letter to whomever they may have distributed the affected product. For questions customers were instructed to contact their Siemens Technical Solutions Ceter or their local technical support representative. For questions regarding this recalla call 914-524-2955. |  | 8 | Lot 742= 8 units all foreign consignees) Lot 743=7 units (all foreign consignees) | Worldwide Distribution - USA including CA, GA, IL, IN, MI, MN, NY, and Puerto Rico.Internationally to Agentina, Bolivia, Brazil, Canada, Chile, Guatemala, South Korea, Philippines, Paraguay, Taiwan, and Uraguay, | Terminated | October 28, 2013 | 334 | 1987 | Not_Computer
Z-0651-2013 | 63226 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Ultrafill DBM labeled in part: TISSUENET Sterilized by T10^6 Description: UltraFill DBM Size: 1cc Product Code: RT53001 Donated Human Tissue Allograft SINGLE PATIENT USE ONLY FOR INTERNATIONAL USE ONLY See Package Insert for Additional Information |  |  | March 14, 2012 | 2012 | Surgical Tissue Network, Inc. | Surgical Tissue Network Inc., DBA TissueNet Inc. recalled their UltraFill DBM Putty-Porcine (1 cc; 5cc; 10cc) because products may contain trace amounts of 316L Stainless Steel particulates. | Employee error | Surgical Tissue Network sent a Notification of Voluntary Tissue Recall dated March 22, 2012, to all affected customers. The firm issued an additional notification letter on August 30, 2012. The letter identified the product, the problem, and action to be taken by the customer. Consignees were asked to return all listed products that remain in inventory. If the product was further distributed, they were asked to forward the recall information to their customers.Customers were instructed to contact TissueNet's Customer Service Department at 800-465-8800 x283 to coordinate return/replacement of the affected product. Customers with questions were instructed to call 800-465-8800 x230.For questions regarding this recall call 407-380-2424. |  | 25 | 25 units | Worldwide Distribution including Turkey, Greece, &amp; Mexico. | Terminated | December 02, 2014 | 993 | 393 | Not_Computer
Z-0656-2013 | 61718 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #2.5 and Accolade TMZF Plus 127 degree Neck Angle V40 Hip Stem #3.5, Howmedica Osteonics Corp. The intended use is for reconstruction of the head and neck of the femoral joint. |  |  | October 23, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Stryker has received a report that single size 3.5 stem was packaged as size 2.5 stem resulting in a possible product mix. | Labeling mix-ups | Stryker sent "Urgent Medical Device Recall Notification" letters and product accountability forms via Fed Ex with return receipt on 10/23/12 to branches/agencies and to hospitals and physicians on 10/24/12. The letter described the issue and provided recommended actions to mitigate the risk. Customers can call (201) 972-2100 if questions. |  | 18107 | 6021-2530- four units; 6021-3535- | Nationwide distribution, including the states of NY, NJ, CA and FL. | Terminated | September 11, 2013 | 323 | 18107 | Not_Computer
Z-0655-2013 | 63938 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FlowGate Balloon Guide Catheter; |  |  | December 20, 2012 | 2012 | Concentric Medical Inc | Firm received complaints of resistance and possible collapse of the distal tip of the FlowGate Balloon Guide Catheters during use for aspiration. | Device Design | Concentric Medical sent a Urgent Medical Device Recall Notification letter dated December 19, 2012, to all affected customers. The letter identified the product the problme and the action needed to be taken by the customer.Customers were instructed to follow these instructions.1. Immediately locate the subject devices and quarantine them.2. Distribute this notice to all affected Departments in your facility.3. Inform Concentric Medical (Stryker Neurovascular) if any of the subjectdevices have been distributed to other organizations.a. Please provide contact details so that Concentric Medical (StrykerNeurovascular) can inform those recipients appropriately.4. Complete and sign the attached Customer Response Form and: Fax it to: 650-237-5230 Please complete this form even if you do not have any subject devices to return to prevent the need of follow up notices.5. Keep a copy of the completed, signed Customer Response Form for yourrecords. Upon receipt of the returned customer response form, a Stryker Representative will contact you to arrange for credit of returned devices.On behalf of Stryker Neurovascular, we sincerely thank you for your help and support and regret any inconvenience and appreciate your assistance in completing this recall.Should you have any questions concerning this matter please do not hesitate to contact the undersigned. |  | 13 | 13 units | US (nationwide) including the states of CA, FL, GA, OR and TN. | Terminated | February 05, 2013 | 47 | 13 | Not_Computer
Z-0660-2013 | 63850 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PT113950 PT HYBRID Glenoid Post |  |  | December 06, 0012 | 12 | Biomet, Inc. | Biomet is recalling Part Number PT-113950 PT Hybrid Glenoid Post, following an investigation which identified that the male thread of the post may be oversized. This oversized condition can vary in degree and may lead to the following three events:1) If the PT Hybrid Glenoid Post is not fully seated into the Hybrid Base and the implant construct isimplanted, then a gap of 1-3 mm will be prese | Nonconforming Material/Component | Biomet Inc. sent a Urgent Medical Device Recall Notices dated December 6, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The notice identified the reason for the recall including directions to immediately locate, discontinue use of the device. Surgeons were asked to continue monitoring of patients implanted. Customers were asked to confirm receipt of the notice by calling 1-888-348-9500 extension 1570. Customers with product to be returned are asked to obtain RGA # by calling 574-372-6677 . A FAX Back Response Form is to be filled out and sent within 3 days of receipt to 574-372-1683. Biomet, Attention Return Goods, Building B Recalls, 56 East Bell Drive, Warsaw, IN 46580. Questions related to the notice should be directed to 574-371-1570, Monday- Friday 8 a.m. to 5 p.m. |  | 1050 | 1050 | Worldwide Distribution -- US (nationwide) including the countries of , Canada, Japan, Argentina, Australia ,Netherlands, Europe and Chile. | Terminated | November 08, 2013 | 730822 | 1050 | Not_Computer
Z-0661-2013 | 63877 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Cassettes, P/N 10133. |  |  | September 20, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products is recalling three lots of STERRAD NX System Cassettes because it does not have adequate data to support the entire duration of the product's labeled shelf-life. | Labeling design | Advanced Sterilization Product sent an Urgent Medical Device Recall (Removal) letter dated December 4, 2012, to all affected customers. The letter informed the customers that the Sterrad NX Cassettes do not have adequate data to support the entire duration of the product's labeled shelf-life. Customers are instructed of the problems identifed and the actions to be taken. Customers are instructed to notify their customers if they have further distributed the recalled products and to have them contact Stericycle at (877) 650-7686 to arrange for return of the product. Customers with questions are instructed to call Stericycle at (877) 650-7686 Mon-Fri from 8am to 5pm, ET.For questions regarding this recall call 949-453-6400. |  | 12280 | 12,280 units | Worldwide Distribution - USA (nationwide) and internationally to the following countries:AR, CO, TW, SI, AU, MX, MY, HU, PH, BE, FR, DE, IE, IT, ES, NO, CH, and GB. | Terminated | October 17, 2013 | 392 | 12280 | Not_Computer
Z-0659-2013 | 63559 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The TM Ardis Interbody System Inserter Stainless Steel, Rubber, Silicone Non Sterile Rx only. |  |  | September 20, 2012 | 2012 | Zimmer, Inc. | Zimmer Spine is has received complaints of difficulty in turning the implant release knob of the TM Ardis Inserter after final positioning of the TM Ardis interbody spacer implant. There has been two reports where the inserter instrument was inadvertently removed from patient with the implant still engaged. The correction includes revised directions explaining the technique to remove the implant | Device Design | Zimmer sent a "MEDICAL DEVICE CORRECTION" letter dated September 26, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Risk Manager Confirmation Form was attached for customers to complete and return. Questions and concerns should be addressed to Zimmer Spine Customer Service at 866-774-6368.When re-design is approved, Zimmer will remove and replace all affected devices. |  | 170 | 170 | USA nationwide; Europe, Sweden | Terminated | April 23, 2014 | 580 | 170 | Not_Computer
Z-0658-2013 | 62950 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | AutoPulse Resuscitation System Model 100, |  |  | August 08, 2012 | 2012 | Zoll Circulation, Inc. | Complaints associated with an unexpected stop in AutoPulse compressions. | Employee error | Zoll sent Customer Notification letters dated August 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Zoll Technical Support Line at 1 (800) 348-9011 for further assistance. |  | 5727 | 3881 in US, 1846 ROW. | Worldwide distribution. | Terminated | May 12, 2014 | 642 | 5727 | Not_Computer
Z-0657-2013 | 63355 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | (l) Visum Halogen Surgical Light, (2) Visum LED Surgical Light, (3) Stryker Flat Panel and Navigation Arm |  |  | September 25, 2012 | 2012 | Stryker Communications, Inc. | Stryker has become aware that there is a low likelihood of incomplete engagement of the Spring Arm Circlip during service or replacementof the Spring Arm. If the Spring Arm Circlip is not fully seated, the Spring Arm could separate and fall along with the surgical lights or monitors mounted to the end of the Spring Arm. To date, there have been no serious injuries reported from the incomplete e | Equipment maintenance | Stryker sent an Urgent Medical Device Correction letter dated September 25, 2012, to all affected customers. The letter identified the product, the prolem, and the action to be taken by the customer. Customers were instructed to immediately inspect their Spring Arm. If customers experienced any concerns with their Spring Arm they were instructed to contact Stryker immediately at 1-866-841-5663 to schedule a service technician to inspect and correct the Spring Arm circlips.For questions regarding this recall call 972-410-7100. |  | 1281 | United States: 1,082 units and Outside United States: 199 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Japan, Taiwan, Canada, United Kingdom, France, Netherlands, Spain, Mexico, India, Korea, Singapore, New Zealand, Hong Kong, Philippines and Brazil. | Terminated | April 22, 2013 | 209 | 1281 | Not_Computer
Z-0662-2013 | 62121 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon Pump |  |  | March 27, 2012 | 2012 | Maquet Cardiovascular, LLC | In the Cardiosave Intra-aortic Balloon Pump, there is a possibility that the coiled cord assembly, a prominent connector cord between the Cardiosave IABP housing and the dual display head which secured atop the IABP may fail as a result of mishandling. | Device Design | Maquet Cardiovascular sent an Urgent Product Recall Letter dated March 27, 2012, and Urgent Product Recall Acknowledgement to all affected customers . The letter identified the product, the problem, and the action to be taken by the customer. Datascope Corp determined that the best course of action was to replace the affected product. Customers were instructed that they would be contacted by a service representative to make arrangements for the removal and replacement of their affected product. Customers with questions were instructed to contact the US Call Center at 1-800-777-4222.For questions regarding this recall call 973-709-7498. |  | 29 | 29 units (total) | Worldwide Distribution - USA including ID and Fl and Internationally to United Arab Emirates, France, Luxembourg, Poland, Spain, Sweden, Germany, Italy, Netherlands, Republic of China, Spain, Sweden, and Thailand | Terminated | September 30, 2015 | 1282 | 29 | Not_Computer
Z-0665-2013 | 63941 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Siemens Emit(R) tox(TM) Serum Calibrators, catalog number 7B019UL |  |  | December 06, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Incorrect expiration date was assigned to the D1 lot of Emit(R) tox(TM) Serum Calibrators. The correct expiration date is 2013-12-31. The incorrect assigned expiration date is 2014-12-31. | Error in labeling | Siemens sent a Customer Notification letter dated December 2012 to all affected customers of Emit(R) tox(TM) Serum Calibrators lot D1, informing them that Siemens Healthcare Diagnostics, Inc. has confirmed that the expiration of 2014-12-31 is incorrect. Siemens requested customers to discontinue use and discard any remaining inventory of the D1 lots. Customers were instructed to complete the Effectiveness Check form included with the letter even if no replacement reagents were needed and fax to 302-631-8467. Siemens would replace any unused kits of the affected lots at no charge.Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 302-631-6311. |  | 1998 | 1,998 cartons | Nationwide Distribution to the following states: AL, AR, CA, CO, FL, GA, IL, IN, KY, MA, MD, MI, MO, NC, NY, OH, PA, TN, TX and WI. | Terminated | December 24, 2013 | 383 | 1998 | Not_Computer
Z-0663-2013 | 63231 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Merit Medical's Custom Procedural Tray or Kits K12T-03162B &amp;amp; K12T-02641F.Custom Procedural Tray or Kits is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. |  |  | September 14, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural Trays/Kits contain 0.9% Sodium Chloride Injection which were recalled because one confirmed customer report where four separate particulate issues were identified in four individual flexible containers. The particular matter may result acutely in local inflammation, phlebitis and/or generalized low-level allergic response to the particulate and/or embolize to other organs in the | Under Investigation by firm | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" dated September 14, 2012 to all affected customers. The letter identified the affected product, description of problem and recommended user actions to be taken. Consignees were requested to ensure that all appropriate personnel are notified of the affected Merit trays so that they can take immediate action. They were instructed to quarantine all affected kits, remove the bag of 0.9% Sodium Chloride and return the affected product to Merit for credit or replacement. Consignees were requested to complete and sign the Customer Response Form and immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031 and mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. Consignees were given product return instructions. |  | 422 | 422 total kits/trays | Distributed in Mississippi and New Jersey. | Terminated | January 14, 2013 | 122 | 422 | Not_Computer
Z-0664-2013 | 63892 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit(R) 2000 Phenobarbital Assay (OSR4D229) and Syva(R) Emit(R) Phenobarbital Assay (4D019UL). |  |  | December 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed higher imprecision with the recalled lots on Beckman Coulter AU(R) Systems and the Syva(R) Emit(R) 2000 Phenobarbital (4D019UL) lots E1, E2, and E3 as an application if the AU(R) System is used, the issue may be experienced. Firm has found that within-run imprecision of the recalled lots frequency of outliers that differ from the expected value by 20%. This is an expansion of | Nonconforming Material/Component | Siemens sent an "URGENT MEDICAL DEVICE RECALL" notice letter dated November 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and discard any remaining inventory of the affected product. Customers were asked to share the content of the letter with their laboratory director regarding the need to review previous Phenobarbital results, conduct patient follow up and/or repeat testing. For a no charge replacement of remaining inventory customers were instructed to contact Siemens Customer Service Center at 888-588-3916. For questions contact Siemens Technical Solutions Center at 800-227-8994. |  | 1979 | 1979 cartons | Distributed in the states of CA, FL, IL, KS, MA, MI, MN, MS, NM, NY, NC, OH, OR, PA, TN, TX, UT, and WA. | Terminated | August 27, 2014 | 625 | 1979 | Not_Computer
Z-0666-2013 | 64007 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns flexible arterial cannula, 6.7 mm (20Fr) OD with 3/8 in. connector, suture ring, 9.5 in (24 cm) long |  |  | January 07, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Flexible Arterial Cannula and High-Flow Aortic Arch Cannula connectors were identified to contain flash on the exterior of the cannula connector that exceeded established specifications. Because the flash has the potential to detach in the operative field, TCVS has decided to remove all affected lots that were manufactured with the affected connector part. | Process control | TERUMO sent an URGENT MEDICAL DEVICE RECALL notification to all consignees on January 7, 2013. Terumo CVS is advising users to discontinue use and return all affected cannulae in inventory. Terumo CVS will issue credit for returned cannulae. Terumo CVS providied a list of marketed devices that can be used as alternatives. Contact Terumo CVS Customer Service with questions or concerns: 800- 521- 2818 Monday  Friday, 8 AM - 6 PM ESTFor questions regarding this recall call 734-741-6173. |  | 6830 | 6830 total units for all products | Nationwide and worldwide: AL , AR , AZ , CA , CO , CT , FL , GA , HI , IA , IL , IN , KS , KY , LA , MA , ME , MI, MO , NC , ND , NV , NY , OH , OK , PA , SD , TN , TX , UT , VA , WA , and WI. Bahrain, BRAZIL, CANADA, CHILE, Honduras, Singapore, Taiwan, Trinidad and Tobago, and UNITED ARAB EMIRATES (UAE). | Terminated | August 26, 2013 | 231 | 6830 | Not_Computer
Z-0674-2013 | 63817 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity CT Computed Tomography X-Ray Scanning Systems: Brilliance 64 - Model #728231; and Ingenuity CT - Model # 728326, Distributed By: Philips Healthcare, Cleveland, OH. The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | December 03, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software anomalies. Philips Medical Systems recalled their Brilliance 64 running software v4.0.0xx516 and their Ingenuity CT running software v4.0.0xx518 due to software issues detected in these CT scanning systems. If these issues were to re-occur they could pose the risk of injury to patients undergoing scans; device users; or device service personnel. | Software design | The firm, Philips Healthcare, sent an 'URGENT- Medical Device Correction" letters' dated November 15, 2012, and December 3, 2012, to their customers. Both letters describe the product, problem and actions to be taken. The letters state that Philips is notifying all affected customers of this information via a 'Field Safety Notice and will be installing a software update on the affected Brilliance 64 and Ingenuity CT system in order to address these issues. A Philips' Field Service Engineer will be contacting each customer site and make arrangements to install the necessary software update on these device systems. The customers should contact their local Philips representative or local Philips Healthcare office. In North America and Canada needing further information or support concerning this update, should contact our Customer Care Solutions Center at 1-800-722-9377. Selection option 5, and enter their site ID or follow the prompts. |  | 6 | A total of nine (9) units were distributed (3 units in the U.S &amp; 6 units overseas) | Worldwide distribution: USA including IN, OH and OR; and the countries: Australia, Belgium, China, Israel, and Sweden. | Terminated | June 18, 2013 | 197 | 6 | Computer
Z-0675-2013 | 63393 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System Brilliance Workspace Portal, Model #728269 Philips Medical Systems, Highland Heights, OH. |  |  | September 11, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips was notified that when performing post processing with the Brilliance Workspace Portal platform with software versions 2.6.1.4, the coronary arteries in the rendered image are segmented automatically. However, the center line is not extracted or highlighted upon hovering and making it impossible to activate the vessel and choose the appropriate name (label) from the drop-down menu. | Software design | Philips Healthcare sent a "CUSTOMER INFORMATION -URGENT-MEDICAL DEVICE CORRECTION" letter dated September 13, 2012. The letter identified the product, problem, and actions to be taken by the customers. Philips Service Engineers will contact customers for implementation of the software update. Contact your local Philips representative or local Philips office for further information concerning this notice. |  | 108 | 108 units | Worldwide Distribution-USA including the states of AZ, CA, CO, IL, IL, IN, MA, ME, MI, MO, NC, NJ, NY, OR, TN, TX, VA and WI and the countries of Austria, Canada, China, France, India, Ireland, Italy, Japan, Netherlands, Norway, Spain, Switzerland, Thailand and United Kingdom. | Terminated | February 14, 2014 | 521 | 108 | Computer
Z-0685-2013 | 62719 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Spine EXPEDIUM Offset Torque Wrench |  |  | July 24, 2012 | 2012 | DePuy Spine, Inc. | A component within the handle of the Expedium Offset Torque Wrench is prematurely wearing, potentially resulting in the instrument recording torque values outside of the specification. | Device Design | DePuy Spine issued an Urgent - Voluntary Product Recall Notification dated July 24, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconstinue the use of the affected product and to remove the EXPEDIUMÂż Offset Torque Wrenches from service and return to DePuy Synthes Spine until direction can be provided by DePuy Synthes Spine regarding a replacement torque limiting device.For questions regarding this recall call 508-828-3769. |  | 140 | 140 units | Worldwide Distribution - USA (nationwide) and internationally to Belgium, Egypt, Ireland, Italy, Portugal, and the UK. | Terminated | December 30, 2013 | 524 | 140 | Not_Computer
Z-0688-2013 | 63771 | No | JLA | Creatine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NEONATAL TOTAL GALACTOSE KIT |  |  | November 26, 2012 | 2012 | Perkinelmer | Truncated V-Bottomed Microplate Wells used in assays contain defective wells. | Process control | Perkin Elmer initiated the recall on November 23, 2012. The Recall Letter will be delivered to the end customer by email facilitated through our PerkinElmer Technical Support group located in Akron Ohio. Technical Support will be asked to make sure the customers receive the Recall Letter and implement the recommended precautions:- Inspect the wells of row A and row H of the affected microplate lots for possible holes and cracks prior to use.- If a hole or crack is visually detected in a well of the microplate, the well should not be used for testing.- If a hole or crack is not visually detected in a well of the plate, it is required during use of the well, the volume of solution in the well should be visually confirmed to be consistent before and after incubation.Please contact PerkinEimer Technical Support for further information (800) 321 -9632, select prompt 2. |  | 30290010 | 3029-0010 =1562 kits (10 plats/kit) and 3029-001B=13 kits (50plates/kit) | Nationwide Distribution. | Terminated | March 07, 2014 | 466 | 30291065 | Not_Computer
Z-0679-2013 | 63768 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius GranuFlo Naturalyte Dry Acid Concentrate For Bicarbonate Dialysis 2.0K 2.25 Ca 1Mg .16.5GAL |  |  | November 26, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | GranuFlo Dry Acid Concentrate may show discoloration | Nonconforming Material/Component | Fresenius Medical Care North America notified consignees by telephone on November 26, 2012 and follow-up letter on November 30, 2012 of the product recall. On January 10, 2013, the firm expanded there recall to an additional 32 lots (34 lots total). Customers were instructed to examine their inventory immediately, discontinue use, place all units in quarantine and return to Fresenius Medical Care North America. Customers were advised to contact their Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. For questions contact the GranuFio Recall Line at 1-888-803-0509 |  | 13409 | 13,409 units | US Nationwide Distribution | Open, Classified |  |  | 158107 | Not_Computer
Z-0692-2013 | 64039 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Eterna Giovinezza System, Powered Laser Surgical Instrument. |  |  | November 07, 2012 | 2012 | Quanta System, S.p.A. | Quanta System S.p.A, Italy, is notifying U.S. consignees that their website contained promotional materials for products distributed in the U.S. which did not specify that some of the claims made therein were not meant to apply to the U.S. distributed devices. | Labeling False and Misleading | Quanta System sent a letter dated November 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to perform the following: quarantine and destroy all promotional literature printed from the manufacturer, website or obtained from the manufacturer. Provide any promotional literature independently produced by the initial importer or distributor containing information about Quanta System products. Provide URL of initial importers or distributors website if information about Quanta System products is contained therein. Provide a written certification that all actions have been taken. Quanta System has or will: review all promotional materials and websiteinformation provided by initial importers and distributors; advise initial importers and distributors of any actions that need to be taken to change promotional materials and/or website information.For questions regarding this recall call 303-223-4336. |  | 12 | 12 units | Class II Recall - Worldwide Distribution - USA including TX, AZ, OR, and CO.Only US product is affected by this recall. | Terminated | January 23, 2013 | 77 | 109 | Not_Computer
Z-0694-2013 | 63931 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Smart Disclosure System, Model 92810, is an integrated component of the Intesys Clinical Suite (ICS) G2, Version 4.03. |  |  | December 05, 2012 | 2012 | Spacelabs Healthcare, Llc | Displaying or printing of an incorrect waveform for the time indicated in the printout or display in the Spacelabs Healthcare Intesys Clinical Suite (ICS) G2, Version 4.03, Clinical Access portion of Smart Disclosure, Model 92810. | Software design | Spacelabs has sent the Urgent -Medical Device Correction Spacelabs Healthcare ICS G2 Version 4.03 Model 92810 Smart Closure letter, dated 19 December 2012, to all U.S. customers. The firm plans to send a customer letter (translated as necessary) via e-mail to all international subsidiaries and distributors of record on December 27, 2012.The letter is addressed to Hospital Administrator/Biomedical Manager/Office Manager.Users are advised on the following actions:1) Advise your staff of this potential failure and advise them to always click the"Refresh" button when scrolling to new waveform data prior to printing or viewing waveform information or storing saved events.2) We will contact you at the earliest possible date to schedule a convenient time for your hospital to upgrade your affected software at no cost.Customers with questions can contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 44 | 27 in US and 17 International | Worldwide Distribution-USA including Puerto Rico and the states of California, Florida, Iowa, Illinois, Louisiana, Michigan, Minnesota, North Carolina, South Carolina, Texas, and Washington, and the countries of Argentina, Mexico, France, and India. | Terminated | July 08, 2013 | 215 | 44 | Computer
Z-0672-2013 | 63698 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Bausch and Lomb 27G sterile cannula packed in the Bausch and Lomb Amvisc 1.2% Sodium Hyaluronate ( Model 59051, 59081, 59051L, 59081L) and Amvisc Plus 1.6% Sodium Hyaluronate (Model 60081, 60051, 60051L, 60081L) ophthalmic viscosurgical device. |  |  | November 15, 2012 | 2012 | Bausch and Lomb, Incorporated | Some disposable cannulas (with 4-digit lot numbers) provided with the Amvisc or Amvisc Plus OVD are not properly engaging to the Luer-Lock on the Amvisc and Amvisc Plus sterile glass syringes. The cannulas may leak viscoelastic material or detach from the syringe during injection. In rare incidents, detachment has resulted in serious patient injury. | Nonconforming Material/Component | The firm, Bausch &amp; Lomb, initially notified their customers via telephone November 14, 2012 and sent a follow-up letter entitled "URGENT VOLUNTARY PRODUCT RECALL AND FIELD CORRECTION" on November 15, 2012 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine the product; identify the affected cannulas; and complete and return the attached Recall and Field Correction Acknowledgment Form via Fax to (800) 822-8918 or email at: Luis.Leyva@bausch.com. Note: Based on your form selections, replacement cannulas will either be expedited directly to your facility or your Bausch &amp; Lomb Sales Representative will contact you shortly to furnish replacements and properly dispose of the cannulas with 4-digit lot numbers in your product inventory.Please feel free to contact Customer Service with any questions or concerns at (800) 338-2020 (option 2, followed again by Option 2). |  | 2515760 | 2,515,760 | Worldwide distribution: USA including DC and PR; and countries of: Canada and Mexico accounts not available at this time. | Terminated | December 04, 2013 | 384 | 2515760 | Not_Computer
Z-0695-2013 | 63863 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Hoffman LRF Wire Bolt, Catalog number 4933-1-001 through -003. |  |  | October 25, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Two complaints have been filed where Hoffman LRF Wires have broken during load bearing application by patients. | Device Design | Stryker sent "Urgent Medical Device Removal" notification letters and acknowledgement forms dated 10/25/2012 via Fed Ex to Branches/Agencies with return receipt. Stryker sent notification letters/acknowledgement forms dated 10/26/202 via Fed Ex to Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letters described the device issue and provided recommended actions. Questions were directed to Stryker Customer Service. |  | 476 | 476 units | Nationwide Distribution, including the states of PA, OH, NY, CA and IL. | Terminated | December 01, 2015 | 1132 | 476 | Not_Computer
Z-0673-2013 | 63319 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Vycor Medical ViewSite Brain Access System 17mm/ 11mm / 5cm Device***STERILE*** |  |  | August 21, 2012 | 2012 | Vycor Medical, Inc. | Vycor Medical recalled their VBAS (Vycor Viewsite Brain Access System) because an unidentified black fiber was found on the device. | Equipment maintenance | Vycor Medical, Inc. notified all customers of the product recall by telephone and follow-up letter. Customers were instructed to quarantine affected products and hold the shipment(s) in quarantine unti further notice. The units will be replaced with inspected and passed units. All distributors were instructed to notify their customers of the recall and ask to have them return their stock. For question call 561-558-2020. |  | 3000 | 3000 devices | Worldwide Distribution - USA (Nationwide) in the states of TX, IL, NJ, FL, OH, PA and (International) in the country of Japan. | Terminated | May 13, 2013 | 265 | 3000 | Not_Computer
Z-0697-2013 | 63840 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 25-Hydroxy Vitamin D EIA, Catalog #AC-57F1. |  |  | November 07, 2012 | 2012 | IDS (Immunodiagnostic Systems Ltd.) | Immunodiagnostic Systems is recalling the 25-Hydroxy Vitamin D EIA, AC-57F1 because the QC report was found to have an erroneous calibrator value assigned. The value of 24.4 ng/mL for calibrator 4 was stated on the QC report, however the correct value should have read 28.4 ng/mL. | Nonconforming Material/Component | IDS sent a recall notification letter on November 7, 2012 to all customers that received the the 25-Hydroxy Vitamin D EIA (AC-57F1). The letter informed the customers of the problem identified and the action to be taken.Customers are instructed to complete the attached Recall and Acknowledgement Form and fax back to Immunodiagnostic Systems Inc. Customers with questions are instructed to contact their local IDS office. |  | 675 | 675 units | US Nationwide Distribution | Terminated | May 31, 2013 | 205 | 675 | Not_Computer
Z-0696-2013 | 63617 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS INTEGRA 400 and 400 plus Analyzer, |  |  | November 19, 2012 | 2012 | Roche Diagnostics Operations, Inc. | A software security issue with Oracles TNS-Listener component has been identified. TNS-Listener is supplied with the Roche COBAS INTEGRA 400/400 plus Analyzers. The Oracle TNS-Listener can be considered as an authentication and redirection component between the client and the database inside the Roche product. This means an attacker could register their own non-Roche database via remote access | Software design | Roche sent an Urgent Medical Device Correction letter on November 19, 2012, to all affected customers via UPS Ground (receipt required). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete the attached form and return by fax and file this Medical Device Correction for future reference. For technical support customers were instructed to contact Diagnostics Technical Support at 1-900-428-2336.For questions regarding this recall call 317-521-3911. |  | 845 | 845 Units | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY and GUAYNABO, PR. | Terminated | January 13, 2014 | 420 | 845 | Computer
Z-0699-2013 | 63958 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BBL Crystal Enteric/Nonfermenter ID Kit, Catalog # 24500, cartons of 20 kits, In-vitro diagnostic reagents for microbiological testing. |  |  | December 13, 2012 | 2012 | Becton Dickinson &amp; Co. | In-vitro diagnostic reagents for microbiological testing may be reversed in well positions in the test kit and when utilized may lead to bacteria misidentification. | Mixed-up of materials/components | Becton Dickinson initiated the recall by letter to direct accounts on 12/13/2012. Ex-US BD sites were notified via email, distributors were notified by fax and customer letters were delivered by UPS ground. The notification advised that the recalled kit contained citrate (CIT} and malonic acid (MLO} reagents in the reversed prong/well positions. Customers were advised to discontinue use of identified lot numbers and discard any remaining packages for replacements, and also to return the enclosed reply form. The recalling firm requested distributors to submit their customer list for direct BD notification to all end users. Contact the BD Customer Service Department at 1-800-638-8663 if further assistance is required. |  | 67820 | 67,820 units | Worldwide Distribution-USA including the states of MA, TN, KY, IN, GA, NC, TX, NY, FL, AZ, WA, NJ, MI, and CA, and the countries of Uruguay, Taiwan, Singapore, Australia, South Korea, Mexico, Hong Kong, Guatemala, Colombia, Chile, Canada, India, Japan, Brazil, China and Belgium. | Terminated | June 03, 2013 | 172 | 67820 | Not_Computer
Z-0700-2013 | 63972 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight HPS (High Performance System) |  |  | December 11, 2012 | 2012 | AMS Innovative Center - San Jose | During electrical recertification testing it was found that the Voltage Scaling Board was non-conforming to the IEC 60601-1-2 standard. | Nonconforming Material/Component | AMS sent "NOTICE OF DEVICE FIELD CORRECTIVE ACTION" letters dated December 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-408-456-3320 for questions about this notification. |  | 389 | 389 devices, all model numbers. | Worldwide Distribution. | Terminated | February 13, 2014 | 429 | 389 | Not_Computer
Z-0701-2013 | 63949 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Becton, Dickinson and Company, BD MAX (tm) (6 channel) Instruments, Catalog number 441916, clinical diagnostic instrument. |  |  | September 18, 2012 | 2012 | Becton Dickinson &amp; Co. | Clinical diagnostic instrument may exhibit minor defects, which could lead to erroneous test results. | Device Design | Becton Dickinson sent an Urgent Field Corrective Action letter September 18, 2012,by fax. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that BD Field Service Associates will contact them within 7-10 days of the notification to determine if their instrument is impacted. Customers were requested to return a response form indicating the serial number on their instrument. If customers needed further assistance they were instrycted ti contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4054. |  | 86 | 86 Instruments | USA including CT, FL, IL, KS, MD, MI, MO, NY, PA, RI, TX and WI. Internationally to Canada, Europe and Japan. | Terminated | February 04, 2013 | 139 | 86 | Not_Computer
Z-0698-2013 | 63912 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Laboratories, ARCHITECT 25-OH Vitamin D Reagent Kit, 03L52-25 (ROW) and 03L52-27 (U.S.) (1 x 100 tests), 03L52-35 (ROW) and 03L52-37 (U.S.) (5 x 100 tests); in vitro diagnostic. |  |  | December 12, 2012 | 2012 | Abbott Laboratories | The Architect 25-OH Vitamin D assay has the potential to generate falsely elevated or falsely depressed results when used with certain lots of Architect Reaction Vessels. | Mixed-up of materials/components | Abbott Laboratories sent a "Product Correction Immediate Action Required" letter dated December 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to immediately discontinue use and destroy any remaining inventory of the reaction vessel lots listed; to clear the ARCHITECT RV hopper of impacted RV lots following the instructions provided, prior to continued use of the ARCHITECT 25-OH Vitamin D assay; to review the patient results obtained with the affected lots with the treating physician or heath care provider to ensure they are aware of the potential analytical issue and can determine whether re-evaluation of the patient is necessary; to forward a copy of the letter to any laboratory to whom they may have forwarded the reagent; and to retain a copy of the letter for their records and to complete and return the Customer Reply Form via fax to: 1-800-777-0051 For further questions please call your local Customer Service representative at 877-422-2688. |  | 28600 | 28,600 total | Worldwide Distribution -- USA (nationwide) and the countries of Austria, Barbados, Belgium, Canada, China, Colombia, Cyprus, Czech Republic, Denmark, France, Germany, Greece, Hungary, India, Ireland, Italy, Japan, Latvia, Libya, Malaysia, Netherlands, New Zealand, Poland, Portugal, Russia, Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open, Classified |  |  | 28600 | Not_Computer
Z-0705-2013 | 63665 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | DePuy Spine VIPER POLY SCREW 5.5 8 X70MM Tl |  |  | November 07, 2012 | 2012 | DePuy Spine, Inc. | Polyaxial screws 8x 80 and larger may disassociate when inserting in the Ilium., VIPER large diameter long screws in high-torque circumstances, the screw shank has the potential to separate from the polyaxial head | Component change control | DePuy Synthes Spine sent an Urgent Voluntary Product Recall Notification letter dated November 7, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, quarantine the material and return remaining inventory. Customers were advised to complete the enclosed Business Reply Form and fax back to 508-828-3762 or email to lregan3@its.jnj.com. For questions call Linda Regan @508-977-6606. |  | 34 | 34 units | Worldwide Distribution - US (nationwide) including the states of: AZ, CA, FL, HI, IL, MD,MI, MO,NC, TN, TX and the country of Australia. | Terminated | December 24, 2013 | 412 | 169 | Not_Computer
Z-0717-2013 | 63873 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Lifestyles Pleasure Collection |  |  | November 05, 2012 | 2012 | Ansell Healthcare Products LLC | One lot of Lifestyles Pleasure Collection Lubricated Condoms is misbranded. Ultra Thin was included rather than Skyn. | Packaging | Ansell Healthcare Products LLC sent an "URGENT: VOLUNTARY CONDOM RECALL" letter/return response form dated November 5, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm's Customer Service Department at 1-800-305-1054 for questions or if additional information is needed. |  | 445 | 445 cases (10680 consumer units) | Nationwide Distribution. | Terminated | July 31, 2013 | 268 | 445 | Not_Computer
Z-0716-2013 | 64043 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Point of Care i-Stat cTnI cartridge. |  |  | November 01, 2012 | 2012 | Abbott Point of Care Canada Limited | APOC has determined that 10 boxes of i-Stat cTnI cartridges from lot P12179 have an incorrect barcode applied to the cartridge portion pack. | Error in labeling | APOC sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Abbott Point Of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point Of Care representative for questions regarding this information. |  | 560 | 560 boxes (25 cartridges in each box) | Worldwide Distribution. | Terminated | June 14, 2013 | 225 | 560 | Not_Computer
Z-0718-2013 | 63981 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ELECTROLASE(R) HYFRECATOR(R) ELECTRODES - SHARP, Disposable Hyfrecator(R) Tips, REF/Catalog No. 7-100-8CS, Sterile, Rx ONLY -- CONMED CORPORATION Electrosurgical, Cutting and Coagulation and Accessories. |  |  | December 27, 2012 | 2012 | ConMed Corporation | ConMed received complaints of some devices which had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. In no instance has it been reported to ConMed that a compromise in the sterile barrier has resulted in illness or injury. The compromised seals were discovered prior to use. | Equipment maintenance | ConMed sent an Urgent Field Safety Notice dated January 7, 2013, and Business Reply Form to all affected customers via USPS Priority Mail. International notices were sent via FedEx International Priority on December 27, 2012. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II and return it with unused affected product to:ConMed Corporation525 French Road, Utica, NY 13502Customer were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. |  | 313190 | 313,190 (total for all devices) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Belgium, Colombia, Czech Republic, Dubai, Guatemala, Hong Kong, India, Israel, Italy, Japan, Kenya, Kuwait, Malaysia, Mexico, New Zealand, Norway, Oman, Pakistan, Portugal, Saudi Arabia, Thailand, and Uruguay. | Open, Classified |  |  | 313190 | Not_Computer
Z-0714-2013 | 63959 | No | FIE | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medisystems 15 Gauge x 3/5 inch Buttonhole Needle Set |  |  | December 20, 2012 | 2012 | Medisystems a NX Stage Company | A total of 5 (ea) single needle pouches are mislabeled as 15 Ga x1" instead of 15 Ga x 3/5". | Labeling mix-ups | Medisystems sent a notification letter dated December 20, 2012 to affected customers. The letter identified the affected products, problem and actions to be taken. Customers were instructed to immediately open and inspect inventory for the affected products. Results of the inspection are to be provided to Medisystems and the product is to be returned to Medisystems. For questions call (978) 332-5973. |  | 26 | 26 casesx 250 needles (6500needles) | US Nationwide Distribution in the states of NY, PA and TN. | Terminated | October 10, 2013 | 294 | 42 | Not_Computer
Z-0738-2013 | 64054 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT SA Blood Culture Bottle. |  |  | December 25, 2012 | 2012 | bioMerieux, Inc. | BacT/ALERT SA Culture Bottle lot 1031866 may contain bottles with uncured sensors. | Nonconforming Material/Component | bioMerieux sent an "URGENT PRODUCT REMOVAL NOTICE" and Acknowledgement Form on January 4, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Contact your local bioMerieux Customer Service Representative for additional assistance. |  | 1690 | 1690 cases | Worldwide Distribution.USA including the states of AR, AZ, CA, CO, DE, FL, GA, IL, IN, KY, LA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WI, and WV, and the countries of China, Curacao and Japan. | Terminated | June 03, 2014 | 525 | 1690 | Not_Computer
Z-0740-2013 | 63827 | No | MGB |  |  | Cardiovascular | PMA | Cordis, a Johnson and Johnson Company, ExoSeal(TM) 6F Vascular Closure Device, cat. No EX500, EX600, and EX700. Used for femoral artery puncture site closure, reducing times to hemostasis and ambulation in patients who have undergone diagnostic or interventional procedures. |  |  | October 29, 2012 | 2012 | Cordis Corporation | Affected lots of Cordis ExoSeal (TM) VCD have been sterilized using a process that did not achieve the sterility assurance level that is required by the standards of this class of product. Model Numbers EX500, EX600, and EX700. | Material/Component Contamination | Cordis Corporation sent an "Urgent Medical Device Recall (Removal ) letter dated October 29, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately identify and set aside all product listed in a manner that ensures the affected product will not be used. A product recall response form was attached for customers to review, complete, sign and return the enclosed acknowledgement form directly to Cordis.For further questions please call (786) 313-2000. |  | 98675 | 98,675 units distributed in US. | US (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Hawaii. | Terminated | November 06, 2013 | 373 | 98675 | Not_Computer
Z-0739-2013 | 64008 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL(TM) Fildes Enrichment, catalog number 211866, 10 tubes/pack, labeled in part ***Becton Dickinson and Company, Sparks, MD 21152 USA*** |  |  | December 10, 2012 | 2012 | Becton Dickinson &amp; Co. | Enrichment media intended to enhance the cultivation of various microorganisms may fail to support the growth of Haemophilus species. | Unknown/Undetermined by firm | Becton Dickinson sent an Urgent Product Recall letter dated December 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to discontinue use and discard any remaining product for replacement. Customers were also advised that due to an anticipated extended delay for the availability of Fildes Enrichment, BD will substitute DIFCO" Supplement B with Reconstituting Fluid as the replacement product. Customers were advised to ensure that this replacement product is appropriately validated in accordance with their laboratory's procedures. Customers with questions should contact BD Technical Services Department at 1-800-638-8663.For questions regarding this recall call 410-316-4000. |  | 11710 | 11,710 tubes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Canada, South Korea, Taiwan, Hong Kong, Japan, Mexico, and Singapore. | Terminated | July 01, 2013 | 203 | 11710 | Not_Computer
Z-0742-2013 | 63800 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DPM5 Monitor, a vital signs monitor used on human patients in healthcare facilities. |  |  | September 05, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified an issue with the DPM5 monitors where ST and Arrhythmia features may not be available. | Device Design | Mindray DS USA Inc. sent an Urgent Product Corrective Action letter on September 5, 2012 via Fed Ex with delivery confirmation to all users. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday 8:30-5:30 pm, to schedule a software upgrade. |  | 13 | 13 units | US Nationwide Distribution including the states of: VA, TN, FL, &amp; TX | Terminated | October 22, 2013 | 412 | 13 | Computer
Z-0702-2013 | 63868 | No | RCA |  |  |  |  | Juno DRF , X-ray system, diagnostic, Fluoroscopic, general-purpose |  |  | December 13, 2012 | 2012 | Philips Healthcare Inc. | Values of fluoroscopy time, Air Kerma Rate andCumulated Air Kerma may not be displayed "atthe fluoroscopist's working position" as per 21 CFR 1003.21 | Radiation Control for Health and Safety Act | Philips Medical issued a notification by letter dated December 13, 2012. via Certified Mail to affected customer. The letter identified the affected product, problem and actions to be taken. Philips will contact affected customers to schedule an upgrade of the equipment . For questions or support concerning this issue, please contact Dominic Siewko at 978-659-7936. |  | 16 | 16 units | US Nationwide Distribution including the states of AL, CA, CO, CT, IL, MD, MI, MO, NY, OH, PA | Open, Classified |  |  | 16 | Not_Computer
Z-0744-2013 | 64135 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | KODAK INSIGHT DENTAL FILM in SureSoft Packets, Size 2 Single Packet Film, Model Number IP21S, REF/Catalog Number 8118226 --- COMMON/USUAL NAME: Dental Intraoral Radiographic Film -- 20 cartons of 80 packets. |  |  | January 17, 2013 | 2013 | Carestream Health, Inc. | Missing symbols and manufacturing information on the clear carton label. The expiration dating symbol is missing but the expiration date is not missing. The Lot symbol is missing but the lot number is not missing. The Manufacturer symbol, name and address including country of origin are missing. All of the remaining symbols and information are on the case label. Missing information is only a | Process change control | Carestream sent an Urgent Customer Notification and Field Corrective Action Letters dated January 17, 2013 and response form to the domestic consignees via UPS Overnight Mail. The letter identified the affected product, problem and actions to be taken. In addition, the consignees were instructed to complete the response form and fax it back to Carestream to fax # 585-323-7643. For questions or concerns customers were instructed to contact the Carestream Customer Care Center in the U.S. at 1-800-933-8031. |  | 202 | 101 cartons: Domestic - 95 cartons; Foreign - 6 cartons | US Nationwide Distribution including the states of California, Florida, Texas, Illinois, Indiana, Maryland, Pennsylvania and Canada. | Open, Classified |  |  | 202 | Not_Computer
Z-0743-2013 | 63866 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alfa Wassermann System Diluent 750 mL |  |  | October 10, 2012 | 2012 | Alfa Wassermann, Inc. | Inorganic phosphorus assays using System Diluent Lot 155599 may not product acceptable calibration results. | Device Design | Alfa Wassermann sent an important advisory notice/product response form to their distributors on October 10, 2012. The letter identified the product, the problem, and the action to be taken by the customer. The letter informs customers about the affected lot and provide instruction regarding the receipt of replacement product. Any recalled product should be destroyed. If there is further distribution, those customers should provide Alfa with a distribution list so that they may be notified. Any questions can be directed to (866) 419 - 2532. |  | 4104 | 4,104 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada and Cayman Islands | Terminated | October 25, 2013 | 380 | 4104 | Not_Computer
Z-0745-2013 | 64081 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) IRON Calibrator (DC85), Siemens material number 10445010. |  |  | December 20, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints of IRON calibrations outside of acceptable calibration slope (m) values (0.97 - 1.03) when using IRON calibrator lots 2DN060 and 2FN063. Lots 2GN068 and 2JN074 may also be impacted but no complaints have been received. | Nonconforming Material/Component | Siemens sent a "CUSTOMER NOTIFICATION" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Siemens Technical Solutions Center at 1-800-441-9250 or your local Siemens Healthcare Diagnostics representative for assistance with this issue. |  | 5999 | 5999 | Nationwide Distribution. | Terminated | January 08, 2014 | 384 | 5999 | Not_Computer
Z-0747-2013 | 63918 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Calibration Verification, 4 vial set, Catalog Number: 88770 |  |  | October 22, 2012 | 2012 | Alere San Diego, Inc. | Out of range (increased) HDL Cholesterol results. | Component design/selection | A recall letter was sent on 10/22/12 to all customers who purchased the Alere Cholestech LDX Mutianalyte Control, 2 vial set, Catalog Number 88769, Lot #C2510 &amp; Lot #C2510A. The letter informed the customers that these lots of controls may yield results which are out of range high for the Level 2 control when tested with the Alere Cholestech LDX HDL test. Customers are instructed of the actions to be taken.On 11/16/12, the recall expanded to include the Alere Cholestech LDX Calibration Verification, 4 vial set, Catalog 88770, Lot#271097. The problem identified is that these lots of calibration verification materials may yield results which are out of range high when tested with the Alere Cholestech LDX HDL test. Customers are instruted of the actions to be taken.On 3/6/13, the recall expanded to include the Alere Cholestech LDX Calibration Verification, 4 Vial set, Catalog # 88770, Lot#298148 and 301886. The recall expanded to include the Alere Cholestech LDX Multianalyte Control, 2 vial Set, Catalog # 88769, Lots#C2572, C2572A, C2593 and C2593A.Customers are instructed to complete and fax the enclosed verification form within 10 days.Customers with questions are instructed to contact Alere San Diego, 9975 Summers Ridge Rd, San Diego, CA 92121, (877) 308-8289, fax (858) 805-8457, and responses.ts@alere.com. |  | 614 | 614 kits | Worldwide Distribution - USA (nationwide) Puerto Rico and South Korea. | Terminated | September 26, 2013 | 339 | 11780 | Not_Computer
Z-0754-2013 | 61611 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092-2 Version 7.4, 5160092-3 Version 7.5, 5160092-4 version 7.6. |  |  | May 24, 2012 | 2012 | GE Healthcare, LLC | It was reported by a customer site that when using GE AdvantageSim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 11,2012, to all consignees on May 24, 2012. The letter identified the product, the problem, and the action to be taken by the consignee. The letter addressed the Safety Issue, Safety Instructions, Affected Product,, Product Correction and Contact Information. GE indicated that a GE Healthcare service representative would contact they would correct all affected systems by providing a software upgrade at no cost to the consignee. For questions customers were instructed to call one of the following numers:USA 800-437-1171Japan 0120 - 055 - 919For questions regarding this recall call 262-513-4122. |  | 65 | 65 | Worldwide Distribution - USA including AL, CT, FL, IL, MA, MO, NJ, NY, PA, TX VA, WA, WI and Internationally to TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SOUTH AFRICA, NIGERIA, MALAYSIA, KOREA, JAPAN, INDIA, GERMANY, FRANCE, EGYPT, DENMARK, CHINA, CHILE, CANDA, BRAZIL, AUSTRIA, and GREAT BRITAIN. | Terminated | January 16, 2014 | 602 | 65 | Computer
Z-0516-2013 | 61870 | No | RFZ |  |  |  |  | Motorola Solutions PCS3000/CS3070 Laser |  |  | April 27, 2012 | 2012 | Motorola Inc | It was discovered the CS3000/CS3070 was manufactured using the laser power settings for a Class 2 laser device instead of the Class 1 power levels as specified on the product label. | Radiation Control for Health and Safety Act | Motorola Solutions, Inc. agrees to provide instructions for downloading and installing a software update from the internet, which will reduce the laser power to IEC Class 1 AELs. Purchasers will be provided an internet site and instructions for notifying Motorola Solutions, Inc. that the software installation and power reduction were successful, thereby completing the corrective action.corrections to all units is anticipated to be completed by May 2013.I. Notifications to all dealers and purchasers must be made, and the FDA NY District Office Recall Coordinator is to be included in the notifications. 2.Correction of all units to be completed by May 2013.3.All products will be brought into compliance with the laser performance standard at no cost to the purchaser. For questions regarding this recall call 631-738-3474. |  | 16961 | 16,961 units domestically | Nationwide distribution | Open, Classified |  |  | 16961 | Computer
Z-0755-2013 | 63934 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Ortho Clinical Diagnostics VITROSÂż Chemistry Products Vapor Adsorption Cartridge REF/CAT No. 680 0100. VITROSÂż Vapor Adsorption Cartridges are used to remove vapor contaminants that may potentially interfere with Signal Reagent used with MicroWell Assays on the VITROSÂż 3600, VITROSÂż 5600, and VITROSÂż ECi/ECiQ Systems. |  |  | December 13, 2012 | 2012 | Ortho-Clinical Diagnostics | Some VITROSÂż Chemistry Products Vapor Adsorption Cartridge CAT No. 6800100 cartons incorrectly contain a Particulate Cartridge Part No. J19612. In some instances, the Particulate Cartridge has been inadvertently installed on the VITROSÂż System instead of the intended VITROSÂż Vapor Adsorption Cartridge which may bias results. | Packaging process control | Ortho-Clinical Diagnostics sent URGENT PRODUCT CORRECTION Letters (dated 12/13/12) along with Q&A; and Confirmation of Receipt Forms to distributors and consignees on 12/13/12. They were advised of the issue and instructed to discard their current inventory of VITROSÂż Vapor Adsorption Cartridges. On 12/13/12, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 5237 | Domestic: 15,891; Foreign: 5237 | Worldwide Distribution, including Nationwide (USA) and the countries of AU, GM, FR, IT, SP, BR, BM, AR, CA, JA, MX, UK, CH, CL, CO, IN, EC, PE, PY, PA, PT, NZ, SG, VE and UY. | Open, Classified |  |  | 5237 | Not_Computer
Z-0751-2013 | 64129 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 45Âż Angle Driver / Z-H, Greatbatch Medical 10 000 Wahrie Drive, Clarance, NY 14031. |  |  | January 16, 2013 | 2013 | Greatbatch Medical | Greatbatch Medical has initiated a global voluntary recall and market withdrawal of its entire Small Angled Driver 45Âż product line. Greatbatch has received reports of noisy and / or rough operation when using these Small Angled Drivers. | Nonconforming Material/Component | Greatbatch sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 15, 2012 to all affected customers. The letter described the Reason For Recall, Risk to Health, and Instructions to Customers. Advised consignees to examine their inventory and quarantine the product in question. If they have further distributed the product to notify their customers and to complete and return the "Recall Response Form". Contact the firm at 1-763-951-8235 for questions relating to this notice. |  | 913 | 913 | Worldwide Distribution-USA including the states of CA, IN, KS, MA, NJ, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CANADA, DENMARK, FRANCE, GERMANY, ITALY, JAPAN, PORTUGAL, and SWITZERLAND. | Terminated | September 16, 2014 | 608 | 913 | Not_Computer
Z-0752-2013 | 63805 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Extremity Medical, IO FiX System Accessory. Catalog Number 101-00010. |  |  | March 22, 2012 | 2012 | Extremity Medical LLC | Lag screw measurement is short (approximately 5 mm) when using current screw measuring device (p/n 101-00010) per the IO FIX Surgical Technique Rev C. which can lead to possible inappropriate implant size selection (e.g. shorter screw size selected). | Device Design | Extremity Medical LLC sent e-mail notification dated March 22, 2012, to all affected customers. The email identified the product the problem and the action to be taken by the customers.The email includes a buleeting which informes the customer that an upgrade to the existing instrumentation is being released. The upgrade will replace the existing gauge. Updated pdf versions of the technique guide are also included. Customers are instructed to return the recalled device after swapping out the recalled device with the replacement unit. Customers will be contacted by Extremity and Sales Reps should contact their principal.For further questions please call (973) 588-8980. |  | 93 | 93 units | Worldwide Distribution -- Nationwide Distribution including the states of AR, AZ, CA, FL, GA, IL, IN, MD, MT, NC, NE, NH, NJ, NY, OH, PA, TN, TX, WI and WA. | Terminated | June 13, 2013 | 448 | 93 | Not_Computer
Z-0753-2013 | 64004 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | TUNNELOC ROUND HEAD INTERFERENCE SCREW |  |  | January 03, 2013 | 2013 | Biomet, Inc. | Thru hole at the tip of the screw is undersized. | Reprocessing Controls | Biomet sent an URGENT MEDICAL DEVICE RECALL letter dated January 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed locate and discontinue use of the affected product and return it to Biomet. recall notification was sent to all consignees on 1/4/2013. The action requires the immediate location and discontinued use of the product and its return to Biomet. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8a.m. to 5 p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 44 | 44 units distributed | Worldwide Distribution - USA including CA, FL, IL, AL, OH, KY, TX , and Internationally to Australia, Columbia, Mexico, Netherlands, and Chile. | Terminated | August 02, 2013 | 211 | 44 | Not_Computer
Z-0750-2013 | 63875 | No | OAF |  |  | Ear Nose &amp; Throat | PMA | Esteem Programmer Model 6001, US and ISA, Model 3001 Dell E54XX, part of the Esteem Totally Implantable Hearing System. |  |  | January 09, 2013 | 2013 | Envoy Medical Corporation | Envoy Medical is conducting a voluntary correction of a limited number of Esteem Programmers, part of the Esteem Totally Implantable Hearing System, to reduce their susceptibility to noise interference. | Software design | Envoy Medical Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated December 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-233-1204 for questions relating to this notice. |  | 44 | 44 | Worldwide Distribution-USA including the states of AL, AZ, CA, CT, FL, GA, IL, MA, MI, MN, NC, OK, RI, and TX, and the country of GERMANY. | Terminated | April 09, 2013 | 90 | 44 | Not_Computer
Z-0762-2013 | 63997 | No | MWA | Nucleic acid-based in vitro diagnostic devices for the detection of Mycobacterium tuberculosis complex in respiratory specimens. | Microbiology | Microbiology | 510(k) | COBAS TaqMan Mycobacteria tuberculosis (MTB) Test FOR IN VITRO DIAGNOSTIC USE. |  |  | November 28, 2012 | 2012 | Roche Molecular Systems, Inc. | The COBAS TaqMan MTB Test Definition Files (TDF) v3.0, for both CE-IVD and JPN-IVD tests, being used with AMPLILINK software version 3.3 series and COBAS TaqMan MTB Test Kits, both CE-IVD and JPN-IVD, have an incorrect Target Elbow Max cutoff value of 48 creating the potential for false negative results to be generated for patient samples with Target Elbows in the range of 48 to 50. | Component change control | Roche issued Product Advisory Notice (PAN) 2012-11 and Urgent Field Safety Notice PAN 2012-11 by email to Local Safety Officers (LSO) and General/Country Managers from all countries on 28-Nov-2012. The notice identifed the affected product, problem and actions to be taken. LSOs are to provide the information in the Field Safety Notice (FSN) to each customer and return a Feedback Form to Roche (email address: rmdsafeb@roche.com) once actions in FSN are executed. |  | 383 | 335 CE-IVD TDFs; 48 JPN-IVD TDFs | Worldwide Distribution - Foreign (22 EU countries - Austria, Belgium, Bulgaria, Croatia, Cypress, Czech Republic, Denmark, Germany, Finland, France, Greenland, Hungary, Iceland, Italy, Luxembourg, Netherland, Norway, Sweden, Switzerland, Spain, Turkey, UK); (41 Rest of World countries - See Attachment for full country distribution + Japan, Portugal and Mexico) | Terminated | August 06, 2015 | 981 | 383 | Computer
Z-0764-2013 | 63485 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Infinity Enteralite Ambulatory Feeding Pumps, List number: INFKIT2 - Addendum to Operator's Manual |  |  | October 12, 2012 | 2012 | MOOG Medical Devices Group | The firm is adding an addendum to the operator's manual for Infinity Enteralite Ambulatory Feeding Pumps pertaining to the Infinite Dose mode and the inability of the Pump to detect air in line. | Nonconforming Material/Component | An Urgent Field Safety Notice was sent to all original consignees by certified mail beginning on Monday, October 15, 2012. Consignees were provided with an Addendum to the Operator's Manual as part of the Notice. Consignees were instructed to assure pump users are aware of the information in the addendum and to store the addendum with the Operator's Manual. Consignees were also requested to complete and return a Medical Device Recall reply form to the firm. For additional information or clarification, consignees are to contact Holly Imes, Manager, Regulatory Affairs at 801-264-1001, ext. 230 or via email at: himesstark@moog.com. |  | 92488 | 92,488 units | Worldwide Distribution and US Nationwide in the countries of (Germany, Sweeden, Croatia, Lithuania, and Dubai). | Terminated | April 10, 2014 | 545 | 92488 | Not_Computer
Z-0757-2013 | 60906 | No | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Brand Name: RX Accunet Embolic Protection System |  |  | November 30, 2011 | 2011 | Abbott Vascular | The recall was initiated because Abbot Vascular has discovered that the identified lots of the RX Accunet Embolic Protection System may exhibit difficult removal of the peel away sheath due to higher than normal wall thickness. | Nonconforming Material/Component | Abbott Vascular sent an "URGENT DEVICE RECALL" letter dated November 30, 2011 to all affected customers. The letter provides the customers with an explanation of the problem identified and an action to be taken. Customers were instructed to work with their local account representative to review their inventory, complete the attached Field Action Reconciliation/ Effectiveness Check Form and return any unused identified products to Abbott Vascular. |  | 1253 | 1253 units (962 units in the US &amp; Puerto Rico) | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, China, Germany, Spain, France, United Kingdom, Italy, Pakistan, Poland, Russia, Slovakia, Taiwan, and South Africa. | Terminated | January 31, 2013 | 428 | 1253 | Not_Computer
Z-0758-2013 | 64055 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Endo-Surgery Inc. The ENSEALÂż G2 Tissue Sealers are indicated for bipolar coagulation and mechanical transection of tissue during laparoscopic and open procedures. |  |  | November 06, 2012 | 2012 | Ethicon Endo-Surgery Inc | Potential for self activation while the device is in use, which may result in unintended thermal damage. | Process control | Ethicon Endo-Surgery Inc. sent a "URGENT MEDICAL DEVICE RECALL" letter dated November 28, 2012, to all affected customers. The letter identified the the product the problem and the action needed to be taken by the customer.The customers were instructed to:1. Examine your inventory immediately to determine if you have affected product on hand and remove the affected product(s).2. Complete the Business Reply Form and fax it to Stericycle at 1-877-787-0367 within 3 business days, even if you do not have affected product. Please reference Event 3436. If you have product to be returned, keep a copy of this form for your records. 3.To return affected product, photocopy your completed Business Reply Form, place it in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Note: Returned product will be credited.Should you have any questions, please contact your local representative or call our Customer Support Center, 1-800-873-3636, Option 6. |  | 58338 | 58,338 Instruments | Worldwide Distribution - USA including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, &amp; WV. and the countries Arab Emirates, Australia, Belgium, Canada, Chile, China, Colombia, Cyprus, Hong Kong, Indonesia, Israel, India, South Korea, Kuwait, Malaysia, Panama, Philippines, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, Thailand, Turkey &amp; Vietnam. | Terminated | April 11, 2014 | 521 | 58338 | Not_Computer
Z-0741-2013 | 64117 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-T1 Ventilator with software versions 1.1.2 or lower. |  |  | December 19, 2012 | 2012 | Hamilton Medical Inc | Analysis of a customer complaint has shown that during ventilation of small pediatric patients with high airway resistance and low lung compliance, the oxygen consumption of a Hamilton T1 ventilator with software versions 1.1.2 and lower must be calculated using a larger margin than originally expected. | Software design | Hamilton sent a Field Safety alert letter dated January 11, 2013 to all customers. The letter identified the affected product, explained the potential problem and a included Medical Device Safety Alert and Corrective Action document. Hamilton is working on a new software version for this issue. Customers are asked to update their operator manual with the update oxygen consumption formula page which provided revised information for calculating oxygen capacity. When the new software is available, Hamilton will provide loaner devices to each location while their unit is being upgraded.Questions should be addressed to Hamilton Medical at 800-426-6331 ext 215 |  | 6693 | 97 total ventilators; 69 of part number 161005 and 28 of part number 161006 | US Nationwide Distribution including the country of Switzerland. | Terminated | April 22, 2014 | 489 | 6693 | Computer
Z-0759-2013 | 61879 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Rectal Retractors, stainless steel, item # AL07289000, a component of Varian Medical Ring and Tandem Applicator Sets; Intended for use in Brachytherapy. |  |  | April 12, 2012 | 2012 | Varian Medical Systems, Inc. | A medical device used in brachytherapy may become contaminated and can cause infections, toxicity or allergic reactions. | Component design/selection | Varian Medical Systems an an Urgent Medical Device Correction letters dated April 11, 2012 and May 8, 2012 to all affected customers. The April notification advised of an anomaly with titanium rectal retractors, and the correction was expanded in May to include stainless steel retractors. In both notifications, users were instructed to discontinue use and quarantine their cleaned and sterilized devices for collection by Varian field service engineers. Users were requested to return an enclosed reply form to respond with number of units in inventory. |  | 344 | 344 pieces | US Nationwide Distribution | Terminated | April 10, 2014 | 728 | 496 | Not_Computer
Z-0767-2013 | 64127 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Phillips Easy Diagnost Eleva DRF. Model Number: 706034 |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | Software design (manufacturing process) | Phillips issued an Urgent Field Safety Notice to all customers on January 14, 2013, to notify users that the Phillips EasyDiagnost Eleva systems, release 3.0 has a mirror icon that could be misinterpereted and could potentially lead to misdiagnosis. Philips will upgrade the system to version 3.1 which involves software and hardware upgrades of the system. This software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Additionally, Philips will provide updated Instructions for Use (IFU) for the Eleva workspot and the EasyDiagnost 3.1 system and an addendum to the EasyDiagnost system IFU. Users will be contacted by Phillips to schedule the installation. For questions contact your local Philips representative: 1-800-722-9377. |  | 14 | 14 units | US Nationwide Distribution including the states of FL, CA, WA, AZ, NC, KY, DE, AZ and MO. | Open, Classified |  |  | 14 | Computer
Z-0765-2013 | 64163 | No | EKS | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Sybron Endo Lightspeed: LSX Files, 25MM, Part Numbers: 835-2525; LSX Files, 25MM Assorted Kit, Part Numbers: 835-2455; LSX Files, 30MM, Part Numbers: 835-2031. |  |  | December 12, 2012 | 2012 | Ormco/Sybronendo | The length markings on the Lightspeed LSX files were wrong. | Process control | Sybron Endo sent a Urgent: Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It has come to our attention that the length rings laser-etched on several lots of the Lightspeed LSX Files are inaccurate. We are asking for the product to be returned since use of the file could be a safety risk due to misadjustment of the file length. Please contact SybronEndo Customer Care at 1-800-346-3636 directly to handle the arrangements of a quick return and replacement. |  | 740 | 740 units | Worldwide Distribution - USA and the country of Canada. | Terminated | November 07, 2013 | 330 | 740 | Not_Computer
Z-0766-2013 | 64099 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | P10 Tobii battery box (product number 510410) |  |  | January 14, 2013 | 2013 | Tobii Assistive Technology, Inc. | Tobii ATI customers P10 external battery pack, which contains sealed lead-acid batteries, malfunctioned while recharging and produced smoke with a chemical smell. | Nonconforming Material/Component | The recalling firm, Tobii Assistive Technology, Inc. sent a notification letter to all customers who received the recalled product via certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediatley discontinue the use of the P10 battery pack and provded with instructions to receive a complimentary replacement solution. The letter instructed customers to dispose of their P10 External Battery Pack in a manner stipulated with local laws. An acknowledgment form is included to be completed by customers. For questions contact Tobii ATI at 1-800-793-9227 or P10batterypack@tobiiati.com. |  | 483 | 483 units | US Nationwide Distribution | Terminated |  |  | 483 | Not_Computer
Z-0756-2013 | 64010 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Buffered Sodium Citrate (9NC) Plus Blood Collection Tubes; Sterile; IVD; REF 363083; 13 X 75 mm, 2.7 mL |  |  | November 12, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer 2.7mL Plus Citrate Blood Collection Tube due to extended Partial Thromboplastin Time (aPTT) test results reported in two customer complaints. | Material/Component Contamination | BD sent "ATTENTION: PRODUCT RECALL" letters dated November 12, 2012 via UPS to distributors and customers instructing them to examine their inventory immediately and discontinue the shipment of the affected product. All customers were instructed to return all products from the lot affected, by following the instructions on the enclosed packing slip,and BD will send them replacement products at no charge. Contact the firm at 10201-847-4267 for assistance. |  | 2944700 | 2,944,700 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, India, Japan, Nepal, Malaysia, Singapore, Brazil, Costa Rica, and Virgin Islands. | Terminated | September 06, 2013 | 298 | 2944700 | Not_Computer
Z-0768-2013 | 64057 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Li-Ion battery used in the Oxus Portable Oxygen Concentrator (POC), RS-00400. |  |  | January 14, 2013 | 2013 | Oxus America, Inc. | Due to past battery inventory storage conditions and deficiencies in the maintenance procedure, it was possible that some batteries were over-discharged to point of damaging the cells. | Nonconforming Material/Component | Oxus, Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-888-475-1568 for information related to this recall. |  | 1724 | 1724 total units | Worldwide Distribution-USA including the states of New York, Michigan, Colorado, Texas, Iowa, Florida, Georgia, Ohio, Wisconsin, Oregon, Alabama, Illinois, Utah, Oklahoma, West Virginia and California and the countries of Canada, European Union, South Africa, Korea and Brazil. | Terminated | July 17, 2013 | 184 | 1724 | Not_Computer
Z-0777-2013 | 63651 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | PREMIER LIFECARE RECLINER 5400 |  |  | September 24, 2012 | 2012 | Winco Mfg., LLC | The performance of the chemicals in the foam and vinyl are not consistent enough to meet the TB 133 fire protection mandate required by California and 4 other states. | Component design/selection | Winco sent a Chair Field Correction Notification dated September 24, 2012, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The consignees were instructed to forward the "Field Correction Form - REPAIR" form back to Winco and the firm would send a Winco Representative to the consignees' facility to correct the chairs in the field.Customers with questions were instructed to contact Winco at 800-237-3377 or email customerservice@wincomfg.com.For questions regarding this recall call 352-854-2929, ext 110. |  | 95 | 95 units (all models) | Nationwide Distribution including CA, FL, IL, MA, MD, MN, NM, NY, OH, OK, and TX. | Terminated | May 23, 2013 | 241 | 95 | Not_Computer
Z-0785-2013 | 64044 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Solara 3G Custom Manual Wheelchair |  |  | September 10, 2012 | 2012 | Invacare Corporation | The potential exists for the wheel to rotate freely despite engagement of hub brake. | Process control | Invacare sent an URGENT: RECALL INFORMATION Letter dated September 10, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1) Complete the Provider Response Card indicating your receipt and understanding of these instructions. 2) Contact your customers to inform them of the recall. It is recommended that you provide them with the enclosed Customer Letter. Set up an appointment to correct the customer's wheelchair as soon as possible. 3) Please examine the wheelchair to confirm it has one of the serial numbers listed above. The serial number can be found on the front lower crossbar located on the right while sitting in the chair. 4) Using the enclosed retrofit kit, make the necessary correction. Installation instructions can be found on page 109 of the Solara service manual, which is enclosed.5) Document the correction of the wheelchair on the enclosed tracking sheet.6) Using the enclosed pre-paid UPS mailing label and original box, return the replaced parts to the address listed on the label. 7) Fax or email the Provider Response Form and tracking sheet to 330-329-3458 or recall@invacare.com. For questions please contact Invacare Customer Service at 800-333-6900. |  | 3 | 3 units | US Nationwide Distribution including the states of CA &amp; MT. | Terminated | May 13, 2014 | 610 | 3 | Not_Computer
Z-0774-2013 | 63531 | No | MRA |  |  | Orthopedic | PMA | Trilogy Acetabular AB Shells with Cluster Holes. |  |  | September 14, 2012 | 2012 | Zimmer, Inc. | Some units of Trilogy shells manufactured since March 2009 may be below the lower range of the pore size specification (as low as 28% vs. 30% porosity specification). On 10/31/2012, the firm expanded the recall action to include lots of Uni-Hole Trilogy Acetabular System shells. | Process control | Zimmer sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated September 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.The firm expanded the scope of the recall in January 2013 to include devices that use a secondary mean of fixation. New recall letters will be submitted to distributors and surgeons. |  | 275952 | 275,952 units | Worldwide Distribution-USA (nationwide) and the countries of Denmark, Germany, Greece, Hungary, Ireland, Italy, Netherlands, Spain, Sweden, UK, Poland, France, Australia, Chile, Canada, India, Japan, New Zealand, Peru, South Africa, Colombia, Ukraine, and UAE. | Terminated | June 20, 2014 | 644 | 275952 | Not_Computer
Z-0781-2013 | 62943 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes(R) Spine Vectra, Vectra-T and Vectra-One Anterior Cervical Plating for Spinal Fusion - 2.5mm Drill Bit, graduated 12mm - 26mm adjustable depth, Part Number 03.613.011. |  |  | August 06, 2012 | 2012 | Synthes USA HQ, Inc. | Although the reported incidence is low, there is the potential for the drill stop to slip on the drill bit resulting in an adverse event. | Component design/selection | Synthes sent a "NOTICE: MEDICAL DEVICE RECALL" letter dated August 6, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 610-719-5450 for questions related to this notice. |  | 63 | 63 | Nationwide Distribution-including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY. | Terminated | August 31, 2015 | 1120 | 111 | Not_Computer
Z-0783-2013 | 63793 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anspach Effort, Inc., 98-0061 MA-15C Bearing Sleeve with Removable Bur Guard. A reusable device used with dissection tools; designed for Transphenoidal and Skull base procedures. |  |  | September 05, 2012 | 2012 | The Anspach Effort, Inc. | Contactbetween the bur and bur guard could generate metal fragments that may or may not be visible to thesurgeon and can potentially remain in the surgical site. The materials used to fabricate the bur guard are not traceable, design validation did not effectively evaluate adequate protection of adjacent tissue and inspection results for each of these bur guards were not documented. | Nonconforming Material/Component | The firm, Anspach Effort, Inc. sent an "URGENT: Medical Device Removal" letter dated November 7, 2012, to its customer. The letter identified the product, problem and actions to be taken. The customer was instructed to do the following: 1) Screen their inventory and remove and return all products immediately. 2) Complete and return the attached Customer Reply Form via fax or email to the address provided on the form. Should the customer have any questions, please contact Anspach Product Support at (800) 327-6887. |  | 8 | 8 devices | Distributed in the state of Massachusetts | Terminated | August 13, 2014 | 707 | 8 | Not_Computer
Z-0784-2013 | 63390 | pack, hot or cold, reuseable | IME | Cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | THERA PEARL 3-in-1 BREAST THERAPY, 2 hot and cold packs in plastic box. |  |  | August 14, 2012 | 2012 | Lansinoh Laboratories Inc | The Spanish translation on the Instructions for Use (IFU) insert and product package for hot and cold packs has incorrect heating instructions which would cause product to over heat. | Error in labeling | Lansinoh Laboratories, Inc. sent an Urgent Product Recall letter dated August 14, 2012, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop distributing and quarantine the affected product. Customers were informed that a credit would be issued for the affected product and they were asked to mail the postage paid Business Reply Card even if they do not have the affected product. If necessary call 1-877-366-1182 to arrange pickup. For questions customers were instructed to call the Customer Service Center at 1-800-292-4794.For questions regarding this recall call 703-299-6477. |  | 51472 | 51,472 units | Nationwide Distribution, including the states of TX, CA, NJ, NY, SD, RI, WV, IL PA, NC and MN. | Terminated | March 08, 2013 | 206 | 51472 | Not_Computer
Z-0787-2013 | 64028 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech NOVATION CROWN CUP, NEUTRAL LINER, Use with Crown Cup Only, GXL UHMWPE, SN 0123456, REF 130-32-53. |  |  | November 05, 2012 | 2012 | Exactech, Inc. | Recall of Novation Crown Cup Connection GXL Neutral Liners Catalog No. (130-32-53) and Novation Crown Cup Connection GXL + 5mm Lateralized Liners Catalog No. (136-32-53). Exactech discovered the device labeling was inadvertently transposed. | Labeling mix-ups | Exactech, Inc. notified consignees by telephone on 11/05/12 of the affected devices and sent an "IMPORTANT PRODUCT RECALL NOTICE" letter dated 11/06/12. The letter describes the purpose of the recall, the product affected, the problem and actions to be taken by the consignees. A Product Recall Response Form was attached for the consignees to contact the Exactech inventory representative within 5 business days to confirm quantities of product in their inventory and return them to Exactech. Consignees with any questions relating to these devices should call 1-800-392-2832. |  | 26 | 26 Units | Distributed in the states of TN, FL, VA, WA and NV. | Terminated | December 24, 2013 | 414 | 26 | Not_Computer
Z-0789-2013 | 63823 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ConMed Linvatec Biomaterials, PopLok bone Punch PKL-35M, 3.5mm |  |  | November 15, 2012 | 2012 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatech recalled one lot of their PopLok bone Punch 3.5mm because the metal shaft diameter is 4.5mm when it should be 3.5mm. | Employee error | ConMed Linvatech issued a "URGENT MEDICAL DEVICE RECALL NOTIFICATION" to its customers on November 15, 2012 for a recall that was conducted at the end user/distributor level. All recall notifications were sent FedEx next day delivery or equivalent method to notify all direct consignees. The notifications identified the affected product along with reason for the recall health risk. All accounts were asked to immediately check their facility for the affected product, stop using any of the product with specified lot number, and to segregate the product. The letter also asked customers to return all affected product to ConMed Linvatec and respond using a Reply Form. If customers have any additional questions, they should contact the Customer Service department at 800-237-0169.For questions regarding this recall call 727-399-5205. |  | 8 | 8 | Worldwide Distribution - USA including California and Internationally to Spain and Korea | Terminated | May 22, 2013 | 188 | 8 | Not_Computer
Z-0790-2013 | 63555 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, Aztreonam 30 ug, IVD, REF CT0264B, Oxoid Ltd., UK, oxoid@thrmofisher.com (Both Remel &amp;amp;amp;amp;amp; Oxoid, Ltd are a part of Thermo Fisher Scientific) |  |  | August 21, 2012 | 2012 | Remel Inc | The firm is recalling the product due to some cartridges may contain individual discs that are not impregnated with antibiotic. The firm notified their customers of the recall with a notification letter and requested a confirmation fax back of receipt. | Nonconforming Material/Component | ThermoFisher sent an Important Medical Device Product Recall Notice dated August 17, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were instructed to review results reported using this lot and consider restesting and/or seek approproiate expert advice. Customers were requested to destroy any remaining inventory of the affected lot and contact Customer Services or their local Oxoid supplier who will issue a replacement. Customer were asked to return the attached "Acknowledgement &amp; Receipt Form by FAX to 1-877-428-1922 or 1-913-895-4190. Customers with questions were instructed to call the Technical Support Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 30 | 30 | Nationwide Distribution including WI, MO, MA, CT, AL, IL, WA, CA, RI, MD, NJ, NC and TX. | Terminated | June 04, 2013 | 287 | 30 | Not_Computer
Z-0792-2013 | 64115 | No | LKA | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Diagnostic Hybrids, Inc., RhMK, Model #:49-0600, 49-0600A, 49-0102A, and 49-T075A. Culture media, used in the cultivation and amplification of viruses, such as respiratory viruses, enteroviruses, measles, and poliovirus from patient specimens. |  |  | December 31, 2012 | 2012 | Diagnostic Hybrids Inc | RhMK product fungal contamination. | Process control | Diagnostic Hybrids sent a Urgent: Medical Device Recall letter dated December 31, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were instructed to immediately review your inventory and determine if you have any of the affected product.If so, please remove it from inventory. Complete the enclosed Product Recall Fax-Back form. Upon receipt of Fax-Back form, Diagnostic Hybrids will send you replacement product or issue credit.Representatives are available to assist you in this process and answer any questions you may have about this recall, and how to obtain replacement product. Please contact Technical Support at technicalsupport@quidel.com or by calling 800.874.1517, Monday through Friday, 8:00 a.m. to 5:00 p.m. EST. |  | 13005 | 13,005 units | Nationwide Distribution including the states of AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA &amp; WI. | Terminated | December 05, 2013 | 339 | 13005 | Not_Computer
Z-0791-2013 | 64160 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Medacta Quadra Trial Broach size 0 part # 01.10.10.045, for use with the Medacta Total Hip Prosthesis System. |  |  | December 17, 2012 | 2012 | Medacta Usa Inc | The tip of the Quadra Trial Broach size 0, broke during a surgery. | Device Design | Medacta sent a Urgent Safety Notification letter dated December 13, 2012 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to check inventory for affected product and contact QA@medacta.us.com for an RMA# to return items to Medacta USA. A new version of the recalled product will be provided to all affected customers. For questions call 805-437-7085 x26. |  | 26 | 26 pieces U.S.A. (226 pieces Worldwide) | US Nationwide Distribution including the states of Arizona, California, Idaho, Illinois, North Carolina, Ohio, Utah, Washington and US Virgin Islands | Terminated | February 07, 2013 | 52 | 26 | Not_Computer
Z-0796-2013 | 64084 | No | FCH | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Custom Breast Biopsy Tray, Catalog # B 9-16261, packaged in bags, 1 tray/bag, 20 bags/case. |  |  | November 26, 2012 | 2012 | Custom Healthcare Systems, Inc. | Surgical procedure trays are packed with pharmaceuticals for injection that may contain particulate matter. | Material/Component Contamination | Custom Healthcare Systems, Inc. sent an Urgent Recall Letter dated November 19, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to check inventory, quarantine affected product and return affected product to Stericycle using the label provided with the letter. Customer were also requested to complete and return the enclosed recall response form. For questions contact Hospira Customer Care at 1-877-946-7747. |  | 80 | 80 trays | USA Nationwide Distribution including the states GA, OH, GA, NC and ME. | Terminated | February 08, 2013 | 74 | 80 | Not_Computer
Z-0793-2013 | 64132 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | OPTIGEN Food 30; Catalog number: 85035. For allergen testing. |  |  | December 21, 2012 | 2012 | Hitachi Chemical Diagnostics Inc | There is no evidence that a visual inspection was performed on one tray of coverslips. Possibility of a low fill volume which could lead to a false negative results. | Process control | Hitachi sent an Urgent Product Recall Notification letter dated December 21, 2012, to the affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was instructed to discontinue distribution and destroy any remaining inventory of the affected product at their location per their local waste ordinances, determine if any of their customers are using or have inventory of the affected product, rquest that their customers review all patient results reported with the use of the affectted product. The customer was instructed to tell their customers to discontinue use and destroy any remaining inventory per their local waste ordinances. The customer was instructed to retain a copy of the notification letter for their records. The customer was also instructed to send confirmation of the destruction within 5 days Customers were instructed to complete and return the attached Certificate of Destruction. For questions customers were instructed to contact their local area Customer Service.For questions regarding this recall 650-961-5501. |  | 20 | 20 kits | Worldwide Distribution - France only | Terminated | May 16, 2013 | 146 | 20 | Not_Computer
Z-0794-2013 | 64069 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SeriScaffold Surgical Scaffold |  |  | January 08, 2013 | 2013 | Allergen Medical | Product is dual packaged in an inner and outer pouch, the outer pouch seal may be compromised and sterility may not be assured | Packaging process control | Allergan sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers are requested to quarantine and return any unused devices in inventory . In addition, customers are asked to complete and return the acknowledgement form. Customers with questions should contact Allergan at 781-658-2800, Mon- Fri, 9 a.m. to 5 p.m. EST. |  | 225 | 225 units (US) and 81 (Outside US) | Worldwide Distribution-USA including the states of CA, FL, GA, MI, MN, NJ, NY, OH, OR, RI, TX, VA, and WA, and the countries of Germany, Spain, and UK. | Terminated | July 29, 2013 | 202 | 225 | Not_Computer
Z-0749-2013 | 63761 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Giraffe and Panda Bag and Mask Resuscitation System |  |  | October 15, 2012 | 2012 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly. | Employee error | GE Healthcare sent an Urgent Medical Device Correction letter dated October 15, 2012, with a return receipt to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer along with the product correction. Required Action for Users: It is recommended that a qualified biomedical technician or respiratory therapist identify the affected resuscitation system and locate the air and oxygen fittings and labels at the back of the unit. With the table provided, identify the appropriate fitting for their country and location and verify the air fitting is located on the left side and the oxygen fitting is located on the right. Verify the air label is located on the left side and the oxygen label is on the right and verify that the color of the air and oxygen labels match the picture on table provided. If the unit does not meet all of the criteria, the unit will require repair. If the unit in use with patient, disconnect the wall supply gas and switch to tank supply. Once the patient no longer requires respiratory support, transfer the patient to another bed and remove the unit from service. A Field Engineer will be deployed to your site to complete the required repair. If the unit NOT in use with Patient, take the unit ouf of service. A Field Engineer will be deployed to your site to complete the required repair. For questions or concerns regarding this notification, please call the phone number 1-800-345-2700. For questions regarding this recall call 262-513-4122. |  | 30084 | 276 units total, 109 of those nationwide. | Worldwide Distribution - Worldwide Distribution - USA including AL, AZ, CA, CO, FL,GA, IA, ID, IL, KY, MA, MD, MI, MN, MO, MS, MT, NC, ND ,NE, NM, NV, NY, OH, OR, PA ,TX, VT, WA, WI and WV; and the countries of Australia, Belgium, Canada, China, France, New Zealand, Poland, Saudi Arabia, Switzerland and United Kingdom. | Terminated | January 18, 2017 | 1556 | 30084 | Not_Computer
Z-0801-2013 | 63584 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | The product is a lyophilized growth supplement for use with dehydrated culture media to prepare BCYE Medium used for isolation of Legionellae from clinical and environmental samples. |  |  | October 10, 2012 | 2012 | Remel Inc | The product may contain low level microbial contamination which could result in incorrect results. | Material/Component Contamination | ThermoFisher Scientific notified customers on October 10, 2012, with a Field Safety Notice letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised use of this lot could result in incorrect results reporting. This notice needs to be passed on to all who need to be aware within your organization.The firm request the information form be faxed back confirming receipt of notice.For further information please contact Technical Support via our main switchboard on +44 (0) 1256 841144. |  | 10 | 10 units | Distribution included the states of IL, IN and MI. | Terminated | June 04, 2013 | 237 | 10 | Not_Computer
Z-0799-2013 | 64175 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Set Single Conn. Ext. DL. Part Number: 050-87216. |  |  | October 12, 2012 | 2012 | Fresenius Medical Care Holdings, Inc. | There is a small risk of a leak in Cycler Sets and if the Cycler Set is punctured, the dialysate may become contaminated potentially resulting in peritonitis. | Process control | Fresenius Medical Care began contacting the affected customers by telephone and mailing an "URGENT: FMCNA LIBERTY CYCLER SET RECALL" letter and Return Receipt form dated October 31, 2012. The letters identified the specific part number and product lot and request that customers discontinue use of the lots immediately and to return unused product to FMCNA. Instructions are provided for returning the product, asking customers to isolate and discontinue use of the lot number 12BR08009 and place the product in a secure, segregated area. Consignees were instructed to contact their Fresenius Medical Care Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product. |  | 560 | 560 cases, totaling 5590 units | Nationwide Distribution. | Completed |  |  | 560 | Not_Computer
Z-0763-2013 | 64155 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, AGA Medical corporation, AMPLATZERÂż TorqVueÂż FX Delivery System. Model #'s: 9-ITVFX06F45/60, 9-ITVFX07F45/60, 9-ITVFX007F45/80, 9-ITVFX08F45/60, 9-ITVFX08F45/80, 9-ITVFX09F45/80, 9-ITVFX10F45/80 9-ITVFX12F45/80, 9-ITVFX13F45/80. Cardiac surgical instrument, used to close atrial septal defects between the right and left atriums. |  |  | January 18, 2013 | 2013 | St Jude Medical Inc | The distal end of the core wire of the TorqVue FX Delivery System could potentially fracture when exposed to a combination of certain cardiac anatomies and usage conditions. | Device Design | St. Jude Medical sent an "URGENT MEDICAL DEVICE RECALL NOTICE" dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were advised to discontinue use of the device and remove it from inventory. Customers should contact their SJM sales representative for questions relating to this recall. |  | 635 | 635 | Nationwide Distribution-including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TX, VA, WA, and WI. | Terminated | May 21, 2013 | 123 | 635 | Not_Computer
Z-0798-2013 | 64114 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DS USA, Inc., DPM Central Monitoring System with ambulatory Telepack. Part numbers: 300BF-PA00002, and 300BF-PA00003 (DPM Central Monitoring System) 115-006908-01, -03, -04 (DPM software upgrade kit). The DPM Central Monitoring System transfers information between DPM Central station and other networked devices. |  |  | November 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | Mindray has identified a software anomaly that may where, in the presence of an asystole, the Telepack ECG waveform may produce a unique step pattern (instead of a flat line) if Lead 1 or V is being utilized for detecting cardiac arrhythmia. | Software design | Mindray North America/ Mindray DS USA Inc. sent an "URGENT PRODUCT CORRECTION ACTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your Mindray Service Representative at 1-800-288-2121 for questions relating to this issue. |  | 51 | 51 units | Nationwide Distribution. | Terminated | August 15, 2013 | 260 | 51 | Computer
Z-0800-2013 | 63433 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Bacterin International, Inc., Osteoselect Demineralized Bone Matrix DMB Putty; Graft ID# B090056-664 Donor# B090056 2.5cc Expiration Date: 12/25/2011; LBL 210-0010-1 REV: 040109; and Graft ID# B090056-675. Bone void filler for voids or gaps in bone. Indicated for treatment of surgically created osseous defects. |  |  | April 18, 2012 | 2012 | Bacterin International, Inc. | Expired product was shipped to two customers. | Incorrect or no expiration date | Bacterin International phoned the two customers and they were instructed to return any of the unused product.For questions regarding this recall call 406-388-0480.The 3rd expired product was discovered by Bacterin Sales Rep 7/12/2012 at the consignee. |  | 3 | 3 packages each containing 2.5cc of putty | Nationwide Distribution in TX , GA and CA. | Terminated | May 13, 2013 | 390 | 3 | Not_Computer
Z-0802-2013 | 60695 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Terumo CDI 101 Hematocrit/Oxygen Saturation Monitoring System. Monitors blood/oxygen saturation, hematocrit and hemoglobin levels. |  |  | November 01, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Inaccurate readings. Software contains unused (turned off) accuracy correction factors for slope (gain) and offset (bias) for Hematocrit/Saturation calculations. The accuracy on any cuvette size did not meet the specifications. | Software design | Terumo sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 1, 2010 to all affected customers. The letter identified the product, problem, and the actions to be taken by the customers. Contact Terumo CVS Customer Service at 1-800-521-2818 for questions or concerns. |  | 622 | 622 units | Worldwide Distribution. | Terminated | November 08, 2013 | 1103 | 622 | Computer
Z-0808-2013 | 63556 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Discs, Cefoxitin 30ug REF CT0119B, packaged with 50 discs per cartridge, 5 cartridges per box. The responsible firm name on the label is Oxoid Ltd., Basingstoke, Hants, UK. |  |  | August 21, 2012 | 2012 | Remel Inc | The firm is recalling the product because the cartridges may contain discs that have an insufficient dose of antibiotic. | Nonconforming Material/Component | The recalling firm sent a recall notification letter dated 8/17/ 2012 and 8/21/2012. |  | 23 | 23 units | US Nationwide Distribution including the states of: CA, CT, GA, IA, KS, OK, RI, TX, WA, WI, | Terminated | June 04, 2013 | 287 | 23 | Not_Computer
Z-0805-2013 | 63244 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap(R) - Miethke Shunt System, FT040T - Control Reservoir Set w/Distal Catheter and FT075P 0 Ventric Cathet, w/Stylet and RT-ANG Guide |  |  | June 04, 2012 | 2012 | Aesculap, Inc. | The deflector may not move freely on the ventricular catheter. The deflector comes pre-positioned at the 5 cm mark on the catheter. In some cases the surgeon will want to adjust the deflector to a different location on the catheter before placing it in the ventricle. There is a small percentage of cases where aggressively attempting to adjust the deflector has a potential to cause the catheter | Process control | Aesculap sent an Important Recall Notification letter dated June 4, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the contact their local sales representative or customer service department at 1-800-282-9000 to return the affected product. Customers were also instructed to complete the Inventory Sheet attached to the letter with the quantity being returned. If customers could not locate the affected product they should provide an explanation as to why the inventory will not be returned. Customers with questions should call 610-984-9251 or 610-984-9291.For questions regarding this recall call 1-800-234-9179, ext 5067. |  | 11 | 11 | Nationwide Distribution including MI, PA and TX. | Terminated | January 08, 2014 | 583 | 11 | Not_Computer
Z-0803-2013 | 64177 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | BREG Cold Therapy Combination Units containing Sterile Polar Pads. 02356 Rev M, PAD M/U XL STER, Mfg.Date: 112012. Model #s: 10903, 10703, 09101, 09111, 09131, 09731, 09611, 09621, and 09631. Local anesthetic effect. |  |  | January 14, 2013 | 2013 | Breg Inc | The recall was initiated because Breg has determined that some Sterile Polar Pads products manufactured from January 2010 to October 2012 may have sustained damage to the product packaging which may compromise product sterility assurance. | Packaging | The firm, BREG, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and quarantine the product; complete the Return Response Form by email, Fax or mail; and contact Breg for a Return Authorization for the affected product and request replacement by contacting Breg Customer Care at 800-321-0607. Note: response is required even if you have no affected inventory.Should you have any questions regarding this communication or need to report an adverse event, please contact Breg Customer Care at 800-321-0607. |  | 44883 | 44,883 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Singapore, Chile, Latvia. | Terminated | July 01, 2013 | 168 | 44883 | Not_Computer
Z-0806-2013 | 63558 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel Xpect Clostridium difficile Toxin A/B, IVD, 20 clostridium difficile toxin A/B tests, REF 24650. |  |  | August 20, 2012 | 2012 | Remel Inc | The firm is recalling the products due to a potential for false negative test results. | Material/Component Contamination | Thermo Fisher sent a recall notification letter with details of the recall on August 20, 2012 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers were requested to follow the instructions provided and fax back the attached "Acknowledgement &amp; Receipt Form" to 1-877-428-1924 (toll free in the US) or 1-913-895-4190 (International). For questions contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 260797 | 493 of lot 082505; 529 of lot 131988; and 448 of lot 131994 | Worldwide Distribution - USA Nationwide and the countries Canada and the United Kingdom | Terminated | June 04, 2013 | 288 | 260847 | Not_Computer
Z-0810-2013 | 63991 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mizuho Orthopedic Table, model 5855, sub assembly 5855-901 Orthopedic Table Spar. |  |  | December 17, 2012 | 2012 | Mizuho OSI | Possible instability of Orthopedic Table Top. If the epoxy bond fails, the aluminum plug may migrate out of the carbon fiber tube. If this occurs, the Orthopedic Table Spar may need to be replaced to ensure stability of the Orthopedic Table Top. | Process control | Mizuho sent a Recall Notification letter dated December 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Orthopedic Table Spar will be replaced at no cost by Mizuho technicians. Customers were informed the potential failure and were advised inspection of tables prior to use. If tables are experiencing the described failure, contact Mizuho Resource Group at 1-800-777-4674, option 2. . |  | 869 | 869 devices | USA Distribution -- Worldwide Distribution including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV and WY., and the countries of Australia, Canada, Egypt, Finland, France, Germany, China, Israel, Italy, Japan, Netherlands, New Zealand, P.R. China, Saudi Arabia, Singapore, Slovenia, South Korea, Spain, Taiwan R.O.C., UAE and UK. | Terminated | April 15, 2016 | 1215 | 869 | Not_Computer
Z-0809-2013 | 64110 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Chenica Gunnell Rehab Series of Wheelchairs (GRSW); Mechanical Wheelchair, Sold as Rehab Tough and Tilt (TNT), Rehab Recline and Mobility (RAM), Rehab Kidster (Kidster). Provide mobility to individuals limited to a seating position. |  |  | January 18, 2013 | 2013 | CHENICA INC. | The wheelchairs may not meet the specifications of WC-19 (crash testing), therefore the WC-19 verbiage and corresponding logo are removed from the flyers on the Gunnell Rehab Series of Wheelchairs (GRSW). | Labeling False and Misleading | Gunnell Inc. sent a "RE: URGENT MEDICAL DEVICE CORRECTION/RECALL" letter dated January 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm 1-800-551-0055 for questions relating to this letter. |  | 53 | 53 | Nationwide Distribution-including the states of Florida, Michigan, South Carolina, California, Texas, Maryland, New York, Tennessee, Virginia, Illinois, Ohio, Massachusetts, Arizona and Washington. | Terminated | April 15, 2014 | 452 | 53 | Not_Computer
Z-0813-2013 | 63654 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000 Physiological Patient Monitor and accessories. |  |  | June 13, 2012 | 2012 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 Patient Monitor. These issues may occur when the Ethernet controller in the XPC823 CPU on the Dash and DashPort 2motherboard functions intermittently. GE Healthcare has identified that these issues occur most often at sites with large Unity Network installations. It should be noted that the Dash 3000/4000/5000 patient m | Nonconforming Material/Component | GE Healthcare sent an Important Product Information letter dated June 13, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were told the otherboard would be replaced, on request, should the customer continue to expeience the issue with the affected product. A second letter dated August 30, 2012, was sent on September 5, 2012, informing consignees that the motherboard replacement did not fully correct the issues. Therefore, the motherboard replacement field action had been suspended. It further explained that the correction will be provided when it is available. Customers with questions were instructed to call 1-800-558-7044.For questions regarding this recall call 262-513-4122. |  | 119262 | 119,262 | Nationwide Distribution - ALL STATES IN CONTINENTAL USA INCLUDING PR and DC excluding AK, DE, ND, and VT.ALGERIA, AUSTRALIA, ANGOLA, ARGENTINA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNG ARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHSTAN, KENYA, KOREA, KUWAIT, LATIVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MALTA, MOROCCO, NEPAL, NETHERLAND, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PALESTINIAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN THAILAND, TRINIDAD AND TOBAGO,K TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARB EMIRATES , UNITED KINGDOM, URUGUAY.PUERTO RICO | Open, Classified |  |  | 119262 | Not_Computer
Z-0811-2013 | 64272 | No | ELW | Impression material. | Dental | Dental | 510(k) | Ultradent Products, Chromaclone PVS, Regular Set heavy body, Bubble Gum Flavor, 50 ml, Impression Material, Part No. 5012, Lot No. B7PPC. Used in prosthodontics. |  |  | February 04, 2013 | 2013 | Ultradent Products, Inc. | Incorrect packaging of regular set and fast set cartridges. | Packaging process control | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected products. They may continue to use product or dispose unwanted cartridges and reply back to firm indicating acknowledgement of letter and if replacement is needed. Consumers with questions can call the firm at 801-553-4420.For questions regarding this recall call 801-553-4420. |  | 32 | 32 cartridges | Nationwide Distribution | Terminated | May 30, 2013 | 115 | 132 | Not_Computer
Z-0818-2013 | 63927 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | Foam electrodes with Silver-Carbon construction: |  |  | November 30, 2009 | 2009 | NAImco Inc dba Richmar Inc | Electrode Socket Terminal pulling away from the electrode lead wire when the electrical stim lead wire was removed. | Unknown/Undetermined by firm | Rich-Mar sent "URGENT: MEDICAL DEVICE VOLUNTARY RECALL" letters to affected customers beginning on 11/30/2009. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 423-648-7730 ext. 214 for questions relating to t his recall. |  | 56980 | 56980 electrodes | Distributed in the states of GA and TN. | Terminated | September 03, 2013 | 1373 | 422100 | Not_Computer
Z-0822-2013 | 64111 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Viterion TeleHealth Monitor Model Number V500; Product Codes 08832984 (US) and 02080468 (UK) - Physiological patient monitor (without arrhythmia detection or alarms). |  |  | November 15, 2012 | 2012 | Viterion TeleHealthcare Llc | Viterion telehealth monitors do not transmit "HI" (above the meter's measurable range) and "LO" (below the meter's measurable range) blood glucose readings from Bayer's Contour Elite XL blood glucose meters, or from Roche's Accu-Chek Advantage 03 and 04 meters to healthcare professionals. | Software design | Bayer Healthcare sent "IMPORTANT RECALL INFORMATION" to all affected customers.The letter identified the product, problem, and actions to be taken by the customers. Contact Viterion Support at 1-800-866-0133 option 1 for questions relating to this issue. |  | 369 | Domestic - 296 units, Foreign - 73 units | Worldwide Distribution-USA (nationwide) and the United Kingdom. | Open, Classified |  |  | 27440 | Not_Computer
Z-0824-2013 | 64070 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Philips Healthcare SureSigns VS2+ |  |  | January 14, 2013 | 2013 | Philips Healthcare Inc. | VS2+ screen may lock up and stop monitoring during patient use | Software design | Philips issued an Urgent Medical Device Correction letter on January 14, 2013. The field action will take place in two phases: Phase 1 Philips sent customer an Urgent Medical Device Correction Notification/Field Safety Notice. This notification explains the issues and gives the customers instructions on actions to take while they await the hardware correction.Phase 2 Once the hardware correction Philips will contact customers and perform the correction free of charge. Until the correction is available.Affected products may continue to be used while waiting for the correction from Philips if the unit is not used in NBP interval mode.A Philips Healthcare representative will contact you to arrange for the correction of all affected devices. If you need any further information or support concerning this issue, please contact the Philips Customer Care Center at 1-800-722-9377.On 3/14/13, firm stated Initially the fix would be a hardware upgrade but have since changed that to a software update. Customers within the US will be sent a new Field Safety Notice as well as a USB software update package and will be required to notify us upon the upgrade of their affected units. They will be performing the updates themselves |  | 3248 | 3248 units | Worldwide Distribution - US Nationwide and the countries of: AUSTRALIA, AUSTRIA, BELGIUM,CANADA, CHILE, DENMARK, FRANCEGERMANY, ICELAND, INDIA, IRELAND,ISRAEL,ITALY, JORDAN, KENYAKUWAIT, LEBANON, MALAYSIA, NAMIBIA, NETHERLANDS, NEW ZEALANDNORWAY,PALESTINE, POLAND,PORTUGAL, SAUDI ARABIA, SINGAPORESLOVENIA, SOUTH AFRICA, SOUTH KOREA,SPAIN,SWEDEN, SWITZERLANDTHAILAND, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESUNITED KINGDOM | Open, Classified |  |  | 3248 | Computer
Z-0795-2013 | 63858 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The Lumenis VersaCut Tissue Morcellator |  |  | January 02, 2013 | 2013 | Lumenis Limited | Lumenis Ltd is recalling certain models of VersaCut Morcellator devices to perform corrective labeling as a result of three reported events of air embolism secondary to incorrect reversal tubing system hook-up. | Labeling design | Lumenis sent an Urgent Safety Advisory Notice letter dated January 2, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to re-label the devices with the enclosed labels provided, replace the Operation Manual with the current revised Manual, complete the information requested on the Confirmation Card and return the Confirmation Card using the enclosed pre-paid shipping label by UPS. For question contact VersaCut Morcellator field correction administrator by either Cell: +972-(54)-9599196, Tel: +972-(4)-9599296 or E-mail: VersaCutMorcellator@lumenis.com. |  | 931 | 931 units | Worldwide Distribution - US Nationwide including: CA, WI, NC, FL, PA, IN, NM, NJ, IA, WA, OH, FL, MD, MA, WY, NY, AZ, NH, MN, TN, OK, KY, TX, AK, MI, VT. Internationally: Japan, Nepal, India, UK, South Korea, Netherlands, Germany, Australia, China, Canada, Turkey, Spain, Italy, Egypt, Mexico, Bulgaria, Utd. Arab, Emir., Poland, Kuwait, Hong Kong, Saudi Arabia, Yemen, Qatar, Czech Republic, New Zealand, Malaysia, Portugal, France, Argentina, Singapore, Romania, Chile, Colombia, Switzerland, Bahrain, Brazil, Serbia/Montenero, Taiwan, Israel, Indonesia, Peru, Pakistan, Nepal, Belarus, Latvia, Sri Lanka, Russian Fed, Thailand, Azerbaijan, Belgium, and Denmark. | Terminated | May 20, 2013 | 138 | 931 | Not_Computer
Z-0819-2013 | 64017 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Quidel, Sofia Influenza A + B FIA, Kit #20218 (25-Test Kit), in vitro diagnostic. |  |  | December 03, 2012 | 2012 | Quidel Corporation | Quidel Corporation has initiated a voluntary recall of specific lots of the Sofia Influenza A+B FIA kits because of false positive results which may lead to misdiagnosis. | Nonconforming Material/Component | Quidel Corporation sent a Medical Device Recall dated December 3, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and remove the affected product from their inventory. Customers should contact Quide Corporation for instructions on how to obtain replacement product. Customers were instructed to fax the enclosed inventory Form to Quidel at 858-552-7905.For technical support customers should call 800-874-1517 or 858-552-1100.For questions regarding this recall call 858-552-1100, ext 7922. |  | 13090 | 13,090 kits (327,250 tests) | Worldwide Distribution - USA (nationwide) and internationally to Korea, Norway, Germany, Chile, Hong Kong, Switzerland, Netherlands, Thailand, Japan, China, and the UK | Terminated | January 31, 2014 | 424 | 13090 | Not_Computer
Z-0814-2013 | 63662 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash Port 2 Docking Station. |  |  | June 13, 2012 | 2012 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 and DashPort 2monitor. This report of correction encompasses the DashPort 2 device, in which the DashPort 2 stopscommunicating with the Dash, resulting in Loss of Monitoring at the CIC. The DashPort 2 status LEDsin the back of the unit indicate Unity Network and DashPort 2 to Dash communication has stopped. | Nonconforming Material/Component | GE Healthcare sent an Important Product Information Letter dated June 13, 2012 via FedEx to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter indicated that the motherboard would be replaced, on request, should the customer continue to experience the issue. A second letter dated August 30th 2012 was issued informing consignees that the correction that was initially proposed did not fully correct the issues. Therefore, the motherboard replacement field action had been suspended. It further explained that the correction will be provided at no charge when it is available. For questions contact Technical Support at 1-800-558-7044. |  | 10497 | 10,497 | Worldwide Distribution and US (nationwide) including DC and the countries of ALGERIA, ARAB EMIRATES, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FRANCE, GERMANY, INDIA, ISRAEL, ITA LY, JAPAN, JORDAN, KUWAIT, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE SLOVENIA, SOUTH KOREA, SPAIN SWEDEN SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM. | Open, Classified |  |  | 10497 | Not_Computer
Z-0823-2013 | 64242 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. |  |  | September 30, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of potential issues associated with the CARESCAPE Monitor B650. There have been no reported illnesses or injuries that have occurred with the use of the device. | Software design | GE Healthcare sent an Important Product Notification letteron September 30, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to keep the patient under close observation or use an alternate monitoring device. Customers should not disconnect the affected USB devices from the patient monitor during an active patient case. If the license is enabled, the alarm limit adjustments from the iCentral Client has an L-NETM-00 license should be disabled. GE Healthcare would contact customers to schedule the update. Customers with questions should contact Technical Support or their local representative.For questions regarding this recall call 262-513-4122. |  | 2813 | 2813 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, FL, GA,IL,IN, IA, KS, KY,LA, MD, MA, MI, M, MO, NE, NE,NJ, NM, NY, NC, OH, OK PA, SC, TN, TX, UT, VA, WI. Internationally to UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, ROMANIA, PORTUGAL, POLAND, PERU, NORWAY, NEW ZEALAND, NETHERLAND, MALAYSIA, LITHUANIA, KUWAIT, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, ICELAND, HONG KONG, GERMANY, FRANCE, FINLAND, EGYPT, DENMARK, CZECH REPUBLIC, COLOMBIA, CHILE, CANADA, BELGIUM BAHRAIN, AUSTRIA, AUSTRALIA. | Terminated | January 27, 2014 | 850 | 2813 | Not_Computer
Z-0654-2013 | 64029 | No | REX |  |  |  |  | Lockheed Martin Gyrocam Systems ( LMGS) LMGS 15" camera systems. Laser pointer and illuminator used for law enforcement surveillance from aircraft. Models DNV, TS and DS |  |  | January 21, 2010 | 2010 | Lockheed Martin Gyrocam Systems, Inc. | These camera systems were manufactured by LMGS and found to fail to comply with the FDA laser performance standard during an FDA inspection. These camera systems incorporate Class IIIb surveying, leveling, and alignment lasers and were entered into US commerce without a variance. Additionally, the manufacturer did not submit a product report to FDA for the camera system. | Radiation Control for Health and Safety Act | Please note that LMGS is responsible for the correction of all noncompliant products.1. Notifications to all purchasers must be made, and the FDA district office Recall Coordinator is to be included in the notifications. 2. Correction of all units to be completed by November 8, 2013.3. All products will be brought into compliance with the laser performance standard at no cost to the purchaser.4. LMGS will submit a product report for the new camera system prior to installation of the replacement lasers as per 21 CFR 1002.10 and certify that the new camera systems comply with the general performance standard in 21 CFR 1010 and the laser performance standard found in 21 CFR 1040.10 and 1040.115. Updated operator manuals will be provided to purchasers.Further questions please call (941) 355-3206. |  | 33 | 33 units | Nationwide Distribution | Terminated | February 24, 2015 | 1860 | 33 | Not_Computer
Z-0832-2013 | 64204 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CareFusion SmartSite Gravity Blood Set Model 42082E. Used to administer blood and blood products through a needle or catheter inserted into the patient's artery or vein. |  |  | December 20, 2012 | 2012 | Carefusion 303 Inc | CareFusion is recalling the SmartSiteÂż Gravity Blood Set as a result of a potential for separation at the connection of the drip chamber and tubing . The separation may be observed at priming or observed during use. If a separation between the drip chamber and tubing occurs, this could cause a delay in treatment. | Nonconforming Material/Component | Carefusion sent an "URGENT: Medical Device Recall Notification" letter dated January 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to return all recalled product directly to the distributor from whom it was purchased immediately, complete and return the enclosed pre-addressed and postage paid mandatory Customer Recall Response Card and to contact your distributor directly, or call the Carefusion Support Center, if you have any questions.For questions Recall Related Questions contact CareFusion Support Center at 888-562-6018; 7am to 4pm (Pacific). For technical questions contact Technical Support at 888-812-3229, 7am to 5pm (Pacific). |  | 5050 | 5,050 units | Nationwide Distribution including the states of AK, CA, CO, IL, MO, NV, NY, TX and UT. | Terminated | July 25, 2013 | 217 | 5050 | Not_Computer
Z-0825-2013 | 63933 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HealthTronics Endocare 3.8mm RenalCryo Cryoprobe, R3.8L. The Endocare CryoProbe is an accessory probe intended for use with the CryoCare Surgical System, for freezing and ablation of tissue. |  |  | December 20, 2012 | 2012 | Healthtronics, Inc. | The Endocare 3.8mm RenalCryo Cryoprobe is being recalled because the label on the outer box has the incorrect expiration date. | Incorrect or no expiration date | Healthtronics sent an Urgent Device Recall letter dated December 20, 2012, to all affected customers that allows tracking. The letter included a suggested format/content for a correction status/acknowledgement letter. The customer was requested to complete the verification section of the recall letter and return all affected products to Healthtronics. A return material authorization (RMA) and pre-paid FedEx return labels was provided for all customers. All affected Healthtronics field personnell were notified of the recall via email.For questions regarding this recall call 512-439-8355. |  | 50 | 50 units | Nationwide Distribution including CA, CO, FL, GA, NC, NJ, OH, TX, and WI | Terminated | April 22, 2013 | 123 | 50 | Not_Computer
Z-0830-2013 | 63783 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Diaphyseal Sleeves |  |  | January 04, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics is initiating a voluntary recall of the LPS Diaphyseal sleeve because the product has the potential for fracture at the taper junction. The LPS Diaphyseal Sleeve to Diaphyseal Sleeve Base taper connection may not be sufficient to accommodate potential physiologic loads that may be transferred to the junction during normal gait activities by some patients. | Device Design | DePuy Othopaedics sent an Urgent Information Medical Device Recall Notice letter dated January 4, 2013 to all affected customers. The letter identified the affected products, reason for recall, clinical implications and actions to be taken. The letter instructed customers to check inventory, discontinue use, cease further distribution and return all inventory on hand for credit. Customers were requested to complete and return the Reconciliation Form. For product-related questions contact your local DePuy Othopaedics Sales Representative, clinical questions contact DePuy's Scientific Information Office at 1-888-554-2482 and recall notice information call 574-372-7333. |  | 162 | 162 total units (US) and 102 units total (OUS) | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, FL, GA, ID, IL, MA, MD, MI, MN, NC, NM, NV, NY, OH, PA, SC, SD, TN, TX, VA, WA, WI, and WV and the countries of Holland, Australia, Austria, South Africa and Canada including Veteran Administration hospitals. | Terminated | May 20, 2014 | 501 | 162 | Not_Computer
Z-0835-2013 | 64321 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | VACUETTE Safety Blood Collection Set + Luer Adapter, 23x12" Safety Blood Collect LA Product, Item number 450096. Usage: The Safety Blood Collection Set is used in routine venipuncture procedures. |  |  | February 01, 2013 | 2013 | Greiner Bio-One North America, Inc. | The VACUETTE Safety Blood Collection Set + Luer Adapter may display leaking (small droplets) during use due to a crack on the luer connector piece. | Other | Greiner Bio-One sent a notification letter dated January 31, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to stop distributing the affected product, complete the attached Product Disposition form and fax to GBO at 800-726-0052 or email to patech@gbo.com. GBO will contact customers for the return of the affected product after receiving the attached Product Disposition form. Customers were requested to provide the names of customers who may have received the product. For questions call 800-515-8112. |  | 192000 | 192,000 pieces (192 cases) | US Nationwide Distribution in the states of FL, IL, KY, NY, NC, PA, OH, TX, VA | Terminated | October 29, 2013 | 270 | 192000 | Not_Computer
Z-0827-2013 | 64248 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ELECTROLASE(R) HYFRECATOR(R) ELECTRODES - BLUNT, REF/Catalog Number 7-101-8BX, Disposable Hyfrecator(R) Tips, Sterile, Rx Only --- CONMED CORPORATION USA. |  |  | January 28, 2013 | 2013 | ConMed Corporation | Some devices had broken through the seal of the sterile pouch. ConMed confirmed instances where the pouch seal was compromised on the affected products. | Equipment maintenance | ConMed sent Urgent Device Recall Letters (dated 1/28/13) and Business Reply Forms to all domestic customers via USPS Priority Mail. International notices were sent via FedEx International Priority to the foreign consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and complete Attachment II response form and return it with unused affected product to: ConMed Corporation 525 French Road, Utica, NY 13502. Customers were asked to return Attachment II even if they don't have any affected product. For questions regarding this recall call 315-624-3533. |  | 104250 | Domestic Distribution: 97,450 units; Foreign Distribution: 6,800 units | Worldwide Distribution-USA (nationwide) and the countries of China, Dubai and Indonesia. | Terminated | April 11, 2017 | 1534 | 104250 | Not_Computer
Z-0836-2013 | 64307 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020, Animas(R) IR 1250 (1200 plus), Animas(R) IR 1200 Insulin Infusion Pumps |  |  | December 17, 2012 | 2012 | Animas Corporation | Due to an internal calendar date limit (December 31, 2015) these pumps will cease to operate on January 1, 2016, and will result in a call service alarm that is displayed on the pump display. | Device Design | The firm initiated their recall of this product on December 17, 2012 by sending a letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Animas Customer Support Department at (866) 796-6373 for questions or concerns. |  | 37998 | 37,998 | Worldwide Distribution-USA (nationwide) including Washington, DC, Puerto Rico, US Virgin Island and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Austria, Australia, Bermuda, Canada, New Zealand, Czech Republic, Italy, France, Finland, Hungary, Israel, United Kingdom, Spain, Sweden, Norway, Denmark, Germany, Mexico and Ireland. | Terminated | December 24, 2013 | 372 | 37998 | Not_Computer
Z-0837-2013 | 64076 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pro Time Microcoagulation System/ Pro Time 3 Cuvette, Pro3-4, Pro3-25, and Pro3-50. |  |  | September 28, 2012 | 2012 | International Technidyne Corp. | Some Pro Time3 Test Cuvettes within a specified lot range may recover lower than expected Prothrombin Time/International Normalized Ratio (PT/INR) results. | Process change control | ITC sent an 'Urgent Medical Device Recall' letter/Health Professional Account Tracking Form dated October 2, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete the form whether or not they have remaining inventory of these lots and return the form promptly. All unused returned product will be replaced. Any questions can be directed to ITC Technical Support at 800-631-5945 Extension 4707 or e-mail ProTime@itcmed.com.For questions regarding this recall call 800-631-5945, ext. 4707. |  | 269690256 | 11,742 boxes US (25 cuvettes per box); 22,968 boxes OUS (25 cuvettes per box) | Worldwide Distribution - USA (nationwide) and Internationally to Bermuda, Canada, Chile, Denmark, Germany, Italy, Latvia, Lithuania, Malaysia, Philippines, South Africa, Spain, Switzerland, Taiwan, and United Arab Emirates. | Terminated | January 09, 2014 | 468 | 269690256 | Not_Computer
Z-0842-2013 | 64158 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | NEUROPATH 6Fx90 GUIDING CATH MPC |  |  | January 11, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Sterility of device may be compromised due to lack of packaging integrity | Component design/selection | Codman sent an Urgent Voluntary Product Recall Notification letter dated January 11, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to check all inventory to determine if they have any affected product, use the included instructions to report inventory status, return and obtain replacement product. Customers were instructed to contact your local Codman Representative or Stericycle at 1-866-201-9077.For questions regarding this recall call 508-880-8100. |  | 867 | 867 units | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Belgium, Canada, Chile, Costa, Rica, Czech Republic, Denmark, France, Germany, Great Britian/UK, Ireland, Italy, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Poland, Portugal, Singapore, Slovakia, Spain, and Switzerland. | Terminated | August 10, 2015 | 941 | 5911 | Not_Computer
Z-0838-2013 | 64246 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | LifeStyles ZERO 10 uber-thin lubricated latex condoms. Mfd. for Ansell Healthcare Products LLC, Dothan, AL 96303. |  |  | January 14, 2013 | 2013 | Ansell Healthcare Products LLC | Out of an abundance of caution, Ansell is implementing a voluntary recall of the Zero Latex Condom (regular) product. The condom has met or exceeded test and quality standards required by the Food and Drug Administration. However, we are experiencing a number of complaints concerning the difficulty in unrolling and subsequent breakage, and so we are asking to have all products returned. | Device Design | Ansell HealthCare Product LLC sent an Recall letter dated January 14, 2013, via registered mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recall notification informs customers about the problem with the affected product and provides instructions to the customers on how to handle affected product. Customers are asked to complete the Response Form and return via fax (1-800) 722-8155.If you have further questions, please contact our Customer Service Department at 1-800-952-9916. |  | 469134 | 469,134 units | .Nationwide Distribution including the states of AL, AZ, CA, CO, DE, FL, GA, IA, IN, IL, KS, KY, LA, MI, MN, MO.MS, NC, NH, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WI and UT. | Terminated | August 06, 2015 | 934 | 469134 | Not_Computer
Z-0841-2013 | 64328 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System. |  |  | February 11, 2013 | 2013 | Elekta, Inc. | The beam is displayed at the "gantry" angle, but the dose is calculated at the "rotation_start" angle. | Nonconforming Material/Component | Elekta provided customers with an "Important Safety Notice - Dose Display Not Matching Beam Geometry - Monaco RFC 35748". The notice identified the product, problem, and actions to be taken by the customers. Contact the form at 800-878-4267 for questions relating to this notice. |  | 13 | 13 | Nationwide Distribution-including the states of AK, CA, FL, NJ, NC, OH, OR, TN, TX, WA, and WI. | Terminated | January 21, 2014 | 344 | 13 | Not_Computer
Z-0839-2013 | 64288 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet ELS Cannula 15 FR. Used for extracorporeal membrane oxygenation (ECMO) for the treatment of respiratory insufficiency in infants. |  |  | December 26, 2012 | 2012 | Maquet Cardiovascular Us Sales, Llc | ELS 15 Fr Cannula could potentially contain an occluded arterial lumen. | Nonconforming Material/Component | Maquet sent Urgent Device Recall Letters/Recall Acknowledgement Forms to their customers on 12/26/2012. |  | 257 | 41 units - US Distirbution; 216 units - OUS Distribution | Worldwide Distribution-USA including the states of AZ, CO, FL, TN, and TX and the countries of Australia, Canada, Czech Republic, Dubai/Bahrain/Kuwait/Lebanon/Saudi Arabia/Yemen, France, Germany, Hong Kong, Italy, Japan, Netherlands, Singapore, South Africa, Sweden, Switzerland, Thailand, Turkey, and UK. | Terminated | May 08, 2014 | 498 | 257 | Not_Computer
Z-0868-2013 | 64215 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions, SuperCross Microcatheter, Sterilized using ethylene oxide, Rx Only, Model numbers 5300 - 5307, 5324, 5325, 5329, 5332, 5336, 5340, and 5341. |  |  | January 29, 2013 | 2013 | Vascular Solutions, Inc. | Vascular Solutions, Inc has discovered a potential problem with a number of lots of Supercross microcatheters. A potential for surface contamination on aged catheters was observed during in-house testing. | Material/Component Contamination | Vascular Solutions sent an "Urgent Medical device Recall" letter dated January 28, 29013 to all affected customers. The letter identified the affected product, described the problem and actions to be taken. The letter advised customers to remove the product from their inventory and to complete and return the VSI Account Inventory Form. For questions contact your local Sales Representative. |  | 4731 | 4731 | Worldwide Distribution - USA Nationwide including the states of AL, AZ, AR, CA, CO, CT, DE, CD, FL, GA, HI, ID, IL ,IN, IA, KY, LA, ME, MD, MA, MI,MN, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT VA, WA, WV, WI, WY and AUSTRIA, CANDA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, ISRAEL,. ITALY, MEXICO, NETHERLANDS, NORWAY, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM. | Terminated | June 25, 2013 | 147 | 4731 | Not_Computer
Z-0866-2013 | 64202 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Merit Medical Systems Inc., Resolve Biliary Locking Drainage Catheter, Models RBC and RBDC, used for drainage of bile within the biliary system. |  |  | January 16, 2013 | 2013 | Merit Medical Systems, Inc. | A consumer complaint that a catheter was observed to be in several pieces under patient fluoroscopic visualization. | Nonconforming Material/Component | Merit initiated the recall on January 16, 2013. A voicemail script notification to users and Q&A; document was provided to those sales representative that had received and or distributed product. On January 16, 2013, Merit sales representatives began contacting each physician and/or user facility that had been provided product.Written communication distribution began on 22 January 2013. Affected sales representatives, physicians, and hospitals were provided written identification information, instruction to not distribute or continue placing these devices, and on how to return any unused product if applicable.All product on hand at Merit was inventoried and remains under quarantined status.Further questions please call (801) 208-4301, |  | 387 | 387 units | Worldwide Distribution - USA including the states of NJ, FL, GA, MO, AZ, UT and the country of Denmark | Terminated | September 16, 2015 | 973 | 387 | Not_Computer
Z-0840-2013 | 63692 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | STRYKER Sustainability Solutions, AC Power Adaptor for Restep Deep Vein Thrombosis (DVT) System, an external compression system. PUMPKIT (Pump Kit), and RSP159-01 (AC Power Adaptor). |  |  | August 30, 2012 | 2012 | Stryker Sustainability Solutions | Reports that the AC Power Adaptor may not operate due to wires coming loose in the unit. | Process control | Stryker Sustainability Solutions sent a Urgent Medical Device Removal Notification letter dated September 5, 2012, to all affected customers. The letter identified the product, problem and the action to be taken by the customer. The firm notified that affected AC Power Adaptors would be replaced with improved Adaptors. Any affected lots would be removed and shipped back to Stryker Sustainability Solutions for final disposition. Customers were instructed to fill out the attached Recall Effectiveness Form even if no affected product was found and fax to Stryker Sustainability Solutions at (480) 809-4957 |  | 1311 | 1,311 | Distributed in Arizona and New Jersey. | Terminated | September 30, 2014 | 761 | 1311 | Not_Computer
Z-0867-2013 | 64319 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Eclipse Treatment Planning System for Radiotherapy; Model Number: H48. In addition, the Eclipse Proton Eye algorithm is specifically indicated for planning proton treatment of neoplasms of the eye. |  |  | January 28, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | The calculation of dose dynamic plans containing different energy fields is incorrect when the Pencil Beam Convolution 11.0 algorithm (PBC 11.0) is used. | Software design | Varian Medical Systems, Inc. Oncology Systems sent an Urgent Medical Device Correction Letter dated January 30, 2013 to all affected customers. The Letter identified the affected product, problem and actions to be taken. Varian Medical Systems is developing a technical correction for this issue, Its customers will be contacted by a Varian Customer Service representatives when this correction is available to schedule its installation. The letter provides customers with help desk contact information. |  | 1168 | 1168 units | Worldwide Distribution - USA Nationwide | Terminated | April 04, 2014 | 431 | 1168 | Not_Computer
Z-0870-2013 | 64329 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Hand Controller Sheath, Catalog number AVA 500 HCS/3010903, for cardiac catheterization procedures. The Sterile Sheath is specifically intended for use in the x-ray angiography environment to cover the Medrad Avanta Sterile Hand Controller to prevent cross contamination between patients. |  |  | November 05, 2012 | 2012 | Medrad Inc | The seal between the white polyethylene and the clear plastic may be missing on the Hand Controller Sheath package, resulting in a potential breach of the sterility of this accessory. | Packaging | Medrad sent a Recall letter dated November 5, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were instructed to review their inventory and quarantine any packages with an expiration date prior to April, 2017. Customers were also given the option of conducting a visual inspection on site to determine if the product should be returned. Alternatively, the customer has the option of returning the affected product. If the customer has affected product, they are instructed to contact Medrad Customer Support to make arrangements for the return and replacement of the product.For further questions please call our Customer Support Team at 1-877-229-3767. |  | 832175 | 832,175 | Worldwide Distribution including USA nationwide. | Terminated | March 10, 2015 | 855 | 832175 | Not_Computer
Z-0872-2013 | 62334 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko Overhead Rail System, Ceiling Mounted Rail Systems for all Liko Stationary Lifts including Likorall, Multirall and Masterlift Systems |  |  | October 31, 2012 | 2012 | Hill-Rom, Inc. | Liko/Hill-Rom to date has received two reports from facilities alleging that the overhead rail system has failed to the extent that the overhead rails have fallen down while in use in high humidity (swimming pool) areas. Analysis of the incidents suggest that the components (hardware) used to secure the rail to the ceiling had corroded and may be unsuitable for environments with an atmosphere con | Device Design | Liko, a Hill-Rom company, sent an URGENT MEDICAL DEVICE RECALL notice dated October 31, 2012 to all affected customers. Letter identified the affected product, problem and actions to be taken until a new design is finalized and the field corrective action implemented. The letter instructed customers to complete and return the Response Form immediately even if you do not have any rail systems in a chlorinated environment. For questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 3816 | 3816 in total | Worldwide Distribution: USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and CANADA LOCATIONS: AB, BC, MB, NB, NS, ON, and SK. | Terminated | March 09, 2016 | 1225 | 3816 | Not_Computer
Z-0871-2013 | 64201 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | January 28, 2013 | 2013 | Philips Healthcare Inc. | Philips HeartStart XL+ Defibrillator/Monitor may become locked out of clinical use | Software design | Philips Healthcare sent an Urgent Medical Device Recall letter dated January 2013 on January 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised if the Philips HeartStart XL+ Defibrillator/Monitor is first powered on and then there is disconnection or an interruption in AC power, the device may be locked out of clinical mode until the user enters the service mode password and performs an Operational Check. Users are provided with the Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Customers were advised to follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 3785 | 3785 units | Worldwide Distribution - USA (nationwide) and Internationally to ALGERIA, ANGOLA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BULGARIA, COLOMBIA, CROATIA (LOCAL NAME HYVATSKA), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GABON, GERMANY, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MALAYSIA, MALTA, MAURITIUS, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PALESTINE, POLAND, PORTUGAL, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA (SLOVAK REPUBLIC), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, AND VIET NAM. | Open, Classified |  |  | 3785 | Computer
Z-0869-2013 | 64185 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | MAYFIELDÂż Composite Series Base Unit Standard, Model (A3101), Rx ONLY, Integra LifeScience Corporation, 4900 Charlemar Drive, Building A, Cincinnati, OH 45227 |  |  | December 21, 2012 | 2012 | Integra LifeSciences Corp. | Integra determined there is a possibility that the Base Units may break during use. | Device Design | Integra LifeSciences Corporation sent an Urgent Medical Device Recall letter dated December 21, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using the affected product immediately and also remove them from service to prevent any further use of the affected product. Integra would provide an alternative base unit to replace those customers could no longer use. Customers were asked to complete the Recall Acknowledgement and Return Form to ensure they received this information. For assistance customers were instructed to call 1-513-533-7944 or 1-513-533-7946 |  | 96 | 96 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA &amp; WI. Internationally to Belgium. | Terminated | July 10, 2013 | 201 | 96 | Not_Computer
Z-0875-2013 | 64322 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | KY TINGLING JELLY Personal Lubricant. The product is packaged into 5 oz tubes with the UPC printed on it. Distributed By: McNeil-PPC Inc. |  |  | January 25, 2013 | 2013 | Mc-NEIL-PPC, Inc. | This product is not cleared for marketing in the US. | No Marketing Application | Johnson and Johnson Healthcare Products sent recall notification letters dated 1/25/2013 to its customers (retailers and wholesalers). The letter identified the product, problem, and actions to be taken by the customers. Customers were asked to return all cases of affected product in their possession. The letter included delivery confirmation and a return packing slip, return authorization label and business reply card. Contact J &amp; J Sales &amp; Logistics at 866-912-8414 for questions regarding the return of affected product. Contact your J &amp; J Sales &amp; Logistics Company Sales Representative for questions about the product. |  | 425304 | 425,304 units | Nationwide Distribution. | Terminated | April 11, 2014 | 441 | 2361232 | Not_Computer
Z-0876-2013 | 64303 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Avex CX Suite, Avex CX 0.22 UL4,5 HK REF #224-017, |  |  | February 08, 2013 | 2013 | Ultradent Products, Inc. | Orthodontic brackets distributed was incorrectly labeled. | Packaging | Ultradent Products, Inc. sent an Urgent Medical Device Correction letter dated February 8 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to return affected products for replacement. Customers with questions can call firm representative, Corey Jaseph at (801) 553-4220. |  | 124 | 124 | USA Nationwide Distribution including the states of CA, AZ, CO and WA. | Terminated | June 07, 2013 | 119 | 124 | Not_Computer
Z-0877-2013 | 63957 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | QuietCare-Networked product, Facility Server/model: QC101601; |  |  | December 16, 2012 | 2012 | Intel-GE Care Innovations LLC | Care Innovations detected a software anomaly: During the engineering investigation into the cause of why a Wander Alert was not generated, a defect in the implementation of the Wander and Night Motion ADLs was discovered. | Software design | Care Innovations sent a "DEVICE CORRECTION NOTIFICATION" letter dated 12/20/12 to all affected customers. The letter identified the product and problem. Correction was performed remotely by the firm. Contact the firm at 917-232-4775 for questions regarding this notice. |  | 19 | 19 units | Nationwide Distribution. | Terminated | February 26, 2013 | 72 | 19 | Computer
Z-0878-2013 | 64294 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Versafitcup Impacting Ring |  |  | January 15, 2010 | 2010 | Medacta Usa Inc | Medacta USA is recalling Versafitcup Impacting Ring because the impaction ring may not fit properly in the implant shell in all cases. This could result in surgical delay potentially increasing possible complications under anesthesia, having to ream up to a larger shell size leaving less bone than planned and repositiong/removal of the shell after initial impaction. | Process control | Medacta USA sent notification letters dated 1/15/10 to Sales Representatives for customers that purchased the Versafitcup Impacting Ring. The letter informed the customers of the problem identified and the actions to be taken. Customers are informed that Medacta USA personnel will contact them by telephone to provide them with a RMA number. Customers are informed that they will receive a new set of redesigned/reworked impacting rings and a second near the end of the January 2010. Customers with questions are instructed to contact the firm at (805) 437-7085. |  | 293 | 293 units | Nationwide Distribution including the states of ID, CO, TX, WI, NY, CT, WA, CA, UT, and AZ. | Terminated | February 23, 2013 | 1135 | 293 | Not_Computer
Z-0761-2013 | 64226 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Brivo XR385 DR-F/ X-Ray System |  |  | August 03, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue on the interface of the Brivo XR385 DR-F/ Model 5271997 radiographic product related to Visual indication as mandated by the US Code of Federal Regulations (21CFR) and IEC 60601-1-3. | Radiation Control for Health and Safety Act | GE Healthcare will, without charge, remedy the issue or bring the product intocompliance with each applicable Federal and IEC standard in accordance with a planto be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you or through a GE Healthcare field engineer site visit. A GE Healthcare Service Representative will update the software on the system to address the issue.Contact Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification |  | 9 | 9 units installed in US | Nationwide Distribution including NM, WY, FL, MO, MA, and MI. | Terminated | December 05, 2013 | 489 | 9 | Computer
Z-0890-2013 | 64316 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Atrapro Antipruritic Hydrogel, 1.5 oz gel tube, 4 Net wt/ 1.5 oz.. |  |  | February 12, 2013 | 2013 | Oculus Innovative Sciences Inc | Out of specification result - stability test failure of free available chlorine (FAC) which functions as a preservative for the hydrogel. The 13 month test result obtained was 4 ppm FAC and the stability specification is &gt;10 ppm FAC. | Incorrect or no expiration date | Oculus sent an Urgent Medical Device Recall letter dated February 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue distribution and segregate the affected product. Immediately remove the affected lots of Recall product from the inventory of sales representative, sub contracted distribution and or fulfillment centers. Segregate the affected product in a secure location for return or destruction. If the product was further distributed customers were instructed to notify healthcare professionals of this recall. Customers were also instructed to completed and return the Business Reply Card and fax to 1-415-462-5163. Once the firm receives the affected product the customer would be contacted by an Oculus Representative and provided a Returns Goods Authorization (RGA) number.For questions customers were instructed to call 1-707-559-2445.For questions regarding this recall 707-559-7191, ext 14. |  | 56567.5 | 56,566 1.5 oz tubes | Nationwide Distribution - One distributor in PA | Terminated | November 25, 2013 | 286 | 56567.5 | Not_Computer
Z-0884-2013 | 63674 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost X-Ray System with Eleva Software 2.1.3 |  |  | October 25, 2012 | 2012 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | Software design | Philips Medical Systems issued on October 25, 2012 an Urgent Field Safety Notice for the Digital Diagnost System informing consignees of the software upgrade. The UFSN described the problem, the hazard involved, the action to be taken by the consumer/user and the action planned by Philips. For questions contact your local Philips representative: Technical Support Line 1-800-722-9377. |  | 551 | 551 | Worldwide Distribution - USA Nationwide and the Foreign countries of: Austria, Canada, Belgium, China,Denmark, France, germany, Italy, Netherlands,New Zealand, Norway, Philippines, Portugual, SOuth Korea, Spain, Switzerland, Thailand, and UK. | Terminated | March 31, 2014 | 522 | 551 | Computer
Z-0879-2013 | 64250 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | BMC RF Cannula Curved Sharp RadiOpaque, Non-Pyrogenic, Active Tip: 10mm, Gauge 16 Ga, Length: 100mm. The BMC RF Cannula is sealed in a Tyvek pouch (primary packaging) and then placed into a white shelf-box (secondary packaging) with the instructions for use. A predetermined number of white shelf-boxes are then placed into shipping cartons (tertiary packaging). The BMC RF Cannula is intended for use in radiofrequency heat lesion procedures for relief of pain. |  |  | October 04, 2012 | 2012 | Baylis Medical Company Inc. | A report of an event was received indicating that upon opening a box containing 10 sterile cannula by a nurse, the pouch package of 5 of the 10 individually packaged cannula had not been sealed along one edge. Thus the pouch was open and the device could fall out of the pouch and on to the floor when retrieved from the shelf box. The cannulas were not used on any patients, as the unsealed packages | Packaging | A recall was initiated on 10/04/2012 and a letter was sent out to the distributor with a Fax-back Acknowledgment Form. The letter described the issue and provided recommendations and a contact number. |  | 1780 | 1780 | International distribution to the country of Australia. | Terminated | May 23, 2013 | 231 | 1780 | Not_Computer
Z-0786-2013 | 63392 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Mx8000 IDT Whole Body Computed Tomography X-ray system Distributed by Philips Medical Systems, Highland Heights, OH 44143 |  |  | September 18, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The low contrast pins in the Image Quality (IQ) phantom deteriorates over time. This prevents the customers from doing the Daily Image Quality Checks as defined in the current Instructions for Use (IfU). | Radiation Control for Health and Safety Act | Philips sent a Urgent Medical Device Recall Notice letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken.This issue does not effect the performance of the Scanner. The Customer Information is intended to inform you about:1. What the problem is and under what circumstances it can occur.2. The action you as a customer can take to minimize the effect of this problem.3 The action planned by Philips to correct this problem.If you need any further information or support concerning this issue please contact your local Philips Representative. In the US and Canada please contact the Philip's Healthcare customer care solutions at 1-800-722-9377, and follow the recorded menu options to reach a Customer Solutions Engineer. In all other countries please contact your local Philips Healthcare office. |  | 1474 | 1,474 Units | Worldwide Distribution - USA ( nationwide) including the states of AK, AL, AR, AZ, CA, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, VA, WA, WV and WY, and the countries of Albania, Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Canada, China, Colombia, Denmark, Dominican Republic, Ecuador, Egypt, France, Germany, Ghana, Gibraltar, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Turkmenistan, United Arab Emirates, United Kingdom, Venezuela, Viet Nam &amp; Yemen. | Terminated | September 02, 2015 | 1079 | 1474 | Not_Computer
Z-0886-2013 | 64357 | No | KNF | Unipolar endoscopic coagulator-cutter and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Integra(R) Jarit(R) Monopolar Cable (Catalogue No. 600-290). |  |  | February 12, 2013 | 2013 | Integra LifeSciences Corp. | Addendum to the Instructions for Use.The addendum emphasizes the warnings and precautions already contained in the device's IFU and also provides additional warnings and precautions that are applicable to the use of electrosurgical equipment in health care facility environments. | Other | Integra sent an Urgent Medical Device Correction Field Safety Notice letter dated February 12, 2013 and a copy of the Instructions for Use addendum to all affected customers. The letter identified the product, nature of the issue and actions to be taken. Customers were requested to examine their inventory for the affected product and include a copy of the addendum to that inventory. Customers were instructed to complete and return the Acknowledgement Return Form and provide a copy of the IFU addendum to any of their customers that had purchased the Jarit(R0 Monopolar Cable. For questions call Integra Surgical Customer Service at 800-431-1123. |  | 22500 | 22,500 | Worldwide Distribution - USA Nationwide including Canada. | Terminated | December 24, 2013 | 315 | 22500 | Not_Computer
Z-0880-2013 | 63822 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems, Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore, Model #728243; and Brilliance Computed Tomography (CT) X-Ray Scanners, Big Bore Oncology, Model #728244. |  |  | October 24, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Software anomaly. Philips determined that the TumorLOC software program for the referenced PET/CT X-Ray scanning systems has the capability to generate Intensity Projection data sets from respiratory gated data. A circumstance was identified by the firm whereby these generated datasets are 'flipped' and incorrectly labeled left-to-right when the are saved to disk. If the flipped Intensity Project | Software design | Philips Medical issued an URGENT - Medical Device Correction letter to their customers concerning the Tumor LOC IPD issue on October 25, 2012. The letter informs the customer of the nature of this software problem and names the Brilliance and GEMINI scanning systems affected by this issue. The letter goes on to describe the hazard inherent to this software issue and provides 'Workaround procedures which the customers can use in order to help mitigate the effects of this problem. Customers with questions should contact their Customers Care Solutions Center at 1-800-722-9377, option 5.For questions regarding this recall call 1-800-722-9377. |  | 765 | 765 units | Worldwide Distribution - USA including AK, AK, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, OA MA, MD MI, MN, MO, MS, NC, D, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV; Internationally to Canada and other countries.. | Terminated | July 23, 2015 | 1002 | 871 | Computer
Z-0885-2013 | 64213 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | Elekta, Apex add-on Micro MLC, collimator for radiation therapy. |  |  | January 18, 2013 | 2013 | Elekta, Inc. | If the Apex collimator is incorrectly installed on the radiation head, a touchguard inhibit activates the view restrictions dialog box only shows the top 12 restrictions. | Employee error | Elekta sent an Important Field Safety Notice Field Safety Notice FCO Ref: 20001407029 "Important information for the safe and satisfactory operation of Apex" to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obey all the applicable instructions, warning, and cautions in the user documentation for each piece of equipment.For questions regarding this recall call 770-300-9725. |  | 1 | 1 | Nationwide Distribution including TX | Terminated | May 08, 2014 | 475 | 1 | Not_Computer
Z-0887-2013 | 64407 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aesculap Bipolar Generator Foot Pedal, Catalog No. GK226, for electrosurgical cutting and coagulation. |  |  | January 18, 2013 | 2013 | Aesculap, Inc. | The bipolar energy did not stop after release of the foot pedal. | Nonconforming Material/Component | Aesculap AG sent an Important Recall Notification letter dated January 21, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory and return for a replacement product or credit. The letter included an inventory sheet to be completed by the customer and returned to the recalling firm. For question call 610-984-9265 or 610-984-9291. |  | 59 | 59 | USA Nationwide Distribution including the states of: AR, CA, FL, IN, IA, KS, MD, MN, MO, OR, TN, TX, VA and WA. | Terminated | January 08, 2014 | 355 | 59 | Not_Computer
Z-0889-2013 | 64228 | No | NCD |  |  | Microbiology | PMA | Vacuette Cellestis 1 ml QuantiFERONÂż TB ANTIGEN (HA) High Altitude: Shake 10x; REF 454413. HE0396. greiner bio-one; 13x75 50pcs.,1200 pcs in total. |  |  | September 27, 2012 | 2012 | Cellestis Inc | Cellestis is recalling QuantiFERONÂżÂż - TB Gold TB Blood Collection Tubes and HA TB Blood Collection Tubes due to potential contamination. | Process design | Cellestis, inc., sent a "Notification of market withdrawal of QuantiFERONÂż" letter dated September 27, 2012 to its domestic and international customers. The letter described the product, problem and actions to be taken. The customers were informed to discontinue use of these products and any product sequestered such that it cannot be used or distributed. Cellestis will provide replacement materials, if and when a retest is required.Please do not hesitate to contact your local QuantiFERON sales representative or customer service representative at +613 8527 3500 or 661-775-7480 should you have any questions or concerns. |  | 29690 | 29,690 units | Worldwide distribution: USA (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Bosnia and Herzegovina, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Ireland, Israel, Jordan, Luxembourg, Mexico, Netherlands, New Zealand, Peru, Poland, Saudi Arabia, Slovakia, Slovenia, South Africa, Switzerland, and Turkey. | Terminated | April 25, 2013 | 210 | 339090 | Not_Computer
Z-0704-2013 | 64134 | No | JAC | Radiographic film marking system. | Radiology | Radiology | 510(K) Exempt | GE Healthcare Definium 5000/5220493-2 X-Ray System |  |  | September 26, 2012 | 2012 | GE Healthcare, LLC | When quickly switching the protocols from AEC to Fixed under the same view, if the user selection process time is less than 0.5 seconds, the User Interface (UI) may display different exposure techniques than selected, and may not display AEC setting. The exposure can still be made as commanded, however it may not be properly indicated. | Software design | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. A GE Healthcare Service Representative will update the software on the system to address the issue.Call Center phone number: 800-437-1171 (United States). For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 312 | 312 installed in US | Nationwide Distribution | Terminated | December 05, 2013 | 435 | 312 | Computer
Z-0891-2013 | 64434 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Technocut Plus, Sterile Single Use Safety Scalpel, Distributed by: Myco Medical. |  |  | February 07, 2013 | 2013 | Myco Medical Supplies Inc | Pouch labeling/printing errors and defective packaging resulting in possible compromised sterility. | Packaging | MYCO Medical sent an Advisory Notice dated February 7, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers wsere instructed to review their inventory for the affected product and immediately quarantine the affected product per their internal procedures. Customers were also instructed to complete the attached Material for Return form and return to MYCO via email QA@mycomedical.com or fax to 919-800-3919. Customers were have the returned material packaged and ready for UPS to pick up at their location. If product was shipped to other customers, customers were asked to contact a MYCO Customer Service Representative with the additional customer's contact and ship-to information. Customers with questions were instructed to call 919-460-2535.For questions regarding this recall call 919-460-2535, ext 105. |  | 110768 | 110,768 scalpels in individual pouches | Worldwide Distribution - USA including CA, FL, IL, MI, MO, NY, OH, PA, SC, TN, WI and Internationally to Canada | Terminated | July 09, 2013 | 152 | 110768 | Not_Computer
Z-0893-2013 | 64313 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker XIA 3 Iliac Screwdriver, 3-PCS. |  |  | December 27, 2012 | 2012 | Stryker Spine | Surgeons have experienced unthreading of the screwdrivers outer shafts during use. | Device Design | Stryker Spine sent a "Urgent Medical Device Removal" letter dated January 14 2013, via FedEx overnight to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to examine your inventory and hospital inventory locations to identify your product. Please fax a copy of the Customer Response form to (855) 632-9049 to Regulatory Compliance. For further questions please call (201) 760-8298. |  | 46 | 46 units US | Nationwide Distribution including the states of AL, AZ, CA, CO, FL, GA, IA, IL,IN, KS, LA, MA, MD, MI, MO, MN, NC, NY. OH, OK, OR, PA, TN, TX, UT and WI. | Terminated | September 30, 2015 | 1007 | 46 | Not_Computer
Z-0895-2013 | 62804 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Orthofix ISKD Limb Lengthener. ISKD - Intramedullary Skeletal Kinetic Distractor. |  |  | August 06, 2012 | 2012 | Orthofix, Inc | There is a possibility that the ISKD limb lengthener may stop distracting post-operatively during treatment, which may result in premature bone consolidation (limb not achieving the desired length) leading to revision surgery to remove and/or replace the device. | Nonconforming Material/Component | Orthofix sent an Urgent Product Recall Notification letter dated August 6, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue distribution of devices, identify the affected and remove from their inventory, and returned to the firm. Credit will be issued upon receipt of the returned product. Customers were asked to take the actions listed in the enclosed Product Return Instructions immediately. Customers with questions should contact Orthofix Customers Service by telephone at 800-266-3349 or 214-937-2047.For questions regarding this recall call 214-937-2509. |  | 1224 | 1224 | Worldwide Distribution - USA (nationwide) Internationally to Australia, Hong Kong, Saudi Arabia, Korea, Turkey, Greece, Finland, South Africa, Spain, Singapore, Switzerland, Brazil, Germany, France, Japan, Netherlands, New Zealand, Israel, Canada and Mexico | Terminated | May 17, 2013 | 284 | 1224 | Not_Computer
Z-0894-2013 | 61554 | No | MHY |  |  | Neurology | PMA | Advanced Neuromodulation Systems, Brio DBS System implantable neurostimulator, IPG or R-IPG, Model 6788 Implantable Pulse Generator. |  |  | March 28, 2012 | 2012 | Advanced Neuromodulation Systems Inc. | Unintended Fluid ingress into the device header has been shown. | Device Design | St. Jude Medical sent an Important Product Field Corrective Action Device Information letter dated March 28, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for any unused product, isolate the affected product for return to St. Jude Medical. Their local representatiave would collect the unused inventory. Any unused product would be replaced with product once inventory becomes available. For patients with implanted product and no degradation of therapy, it is recommended dthat the devices not be unnecesssarily explanted. For patient with implanted product who experience diminished or loss of therapy customers were advised to follow the steps provided in the letter. Customers were asked to forward the notification to any institution in the event one or more patients or products could be potentially affected by this action. Customers wsere asked to maintain a record of thenotice along with the recommendations to ensure the effectiveness of the communication.For questions regarding this recall call 972-309-2154. |  | 388 | 388 units | Worldwide Distribution - 53 consignees out of the US. | Terminated | August 15, 2013 | 505 | 388 | Not_Computer
Z-0899-2013 | 64389 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips MX 16-Slice CT Scanner System, a whole-body computed tomography X-ray system. |  |  | December 19, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | The four screws that secure the patient table top to the carrier became loose, which caused the table top to detach. | Component design/selection | Philips Healthcare sent an Urgent - Medical Device Correction letter dated December 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to have BioMedical Engineering personnel inspect the patient table top screws that secure it to the carrier. Customers were instructed that If it is found that the table top screws are improperly seated, they should stop using the system and contact their local Philips Field Service Engineer for repair. A Philips Field Service Engineer will schedule a time to inspect and repair all the affected systems. For further information customers were instructed to contact their Philips Healthcare Customer Care Solutions Center at 1-800-722-9377 and follow the recorded menu option to reach a Customer Solutions. For questions regarding this recall call 800-722-9377, ext 5. |  | 412 | 412 Units | Worldwide Distribution - USA including AL, AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, NJ, NV, NY, OH, PA, PR and VA. | Terminated | December 03, 2013 | 349 | 412 | Not_Computer
Z-0896-2013 | 64265 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Firebird Spinal Fixation System, Parallel Rod Connector (52-6800) and Axial Rod Connector (52-6700). |  |  | February 01, 2013 | 2013 | Orthofix, Inc | There is a possibility that the Set Screw (p/n 90-2923) sub-component of the Parallel Rod Connector Assembly (p/n 52-6800) &amp; Axial Rod Connector (p/n 52-6700) may crack during final tightening. Orthofix Inc received six (6) customer complaints related to set screw cracking between November 15, 2010 and December 2012. | Device Design | Orthofix sent an Urgent Medical Device Recall Notification dated February 1, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to identify and remove the affected products from inventory immediately and follow the actions listed in the enclosed Product Return Instructions. For questions or concerns contact Regulatory Affairs at 214-937-2100. |  | 565 | 565 total (287 PN 52-6800 and 278 PN 52- 6700) | Worldwidewide Distribution - USA Nationwide including Puerto Rico and the countries of: Australia, Lebanon, Germany, Spain, Sweden and Israel. | Terminated | July 01, 2013 | 150 | 565 | Not_Computer
Z-0897-2013 | 64403 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. Diagnostic imaging system |  |  | January 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The system may not default to the same gating trigger with which the data were acquired. This could cause a risk for patients. | Software design | Philips Medical Systems (Cleveland), Inc. sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated January 21, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact your local Philips representative or Philips Healthcare office for further information concerning this issue. |  | 1 | 1 Unit | Distributed only in Ohio. | Terminated | July 19, 2013 | 179 | 1 | Not_Computer
Z-0898-2013 | 64235 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | REF CMS-CT6T225, 6F Triple Lumen CT PICC with M.S.T. Components |  |  | December 13, 2012 | 2012 | Medical Components, Inc dba MedComp | The affected product was packaged with the incorrect introducer needle. The kit label indicates a "safety" needle is included; the kits were packaged with a non-safety needle. | Nonconforming Material/Component | Medcomp sent a notification letter dated December 13, 2012 via email to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns call 215-256-4201 |  | 785 | 785 | USA Nationwide Distribution including the state of PA. | Terminated | July 31, 2013 | 230 | 785 | Not_Computer
Z-0902-2013 | 64383 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Century Console Base, Model #043625051b. |  |  | January 31, 2013 | 2013 | Salyer Prn Biomedical | Concerns that the addition of the UPS Battery Backup to the Century Console Base Shelf did not comply with the FDA premarket notification regulations. | No Marketing Application | An Urgent: Medical Device Recall letter, dated 1/31/13, was sent to all the customers that received the Century Console Base Shelf option with the Linear/1.5k 115v UPS Battery Backup for the Century Heart Lung Machine. The letter informed the customers that Salyer PRN Biomedical will have a Salyer Biomedical technician return the customers' Century Console Bases to their original configuration without the Shelf option and restore the Linear/1.5k UPS Battery Backup to its original condition as a stand-alone unit with its own wheel base for ease of use. A refund or credit will be issued for the return of the product. Customers are instructed to contact Jeff Poland at (314) 497-5232. |  | 9 | 9 units | Distributed in the states of MA,CA, and WI. | Terminated | April 15, 2013 | 74 | 9 | Not_Computer
Z-0900-2013 | 63671 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Digital Diagnost |  |  | October 30, 2012 | 2012 | Philips Healthcare Inc. | When the operator for a wall stand view selects an "image rotation" different from default, or such image rotation is preset in the examination database, then the resulting image may be rotated in the wrong direction. When the operator manually rotates the image back, electronic side markers (if programmed) are rotated also and as a result may be placed in a wrong position inside the image. | Software design | Philips Healthcare issued an Urgent Field Safety Notice dated October 25, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review the information contained in the Notice with all member of their staff who need to be aware of the contents. Also, customers were asked to retain a copy with the equipment Instruction for Use. For further information or questions customers were instructed to contact their local Phillips representative Technical Support Line at 1-800-722-9377.For questions regarding this recall call 978-659-4519. |  | 360 | 360 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Canada, Chile, Czech republic, Estonia,France, Germany, Hungary, Inida, Ireland, Italy, Japan, Libyan Arab Jamahari, Mexico, Netherlands, New Zealand, Norway , Oman, Panama, Peru, Poland, Qatar, Reunion, Russia, Saudi Arabia, South Africa, South Korea,Spain, Sweden, Swizerland, Thailand, UAE, and the UK. | Terminated | July 29, 2013 | 272 | 360 | Not_Computer
Z-0901-2013 | 64071 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Volume Viewer Plus 5.x (Advantage Workstation 4.2) Volume Viewer 3.x (Advantage Workstation 4.1) Volume Analysis (Advantage Workstation 4.0) |  |  | May 25, 2012 | 2012 | GE Healthcare, LLC | Incorrect views and measurements possible. GE Healthcare became aware of multiple potential safety issues involving the Advantage workstation, Volume Viewer Plus and Volume Analysis. | Software design | GE Healthcare sent a "Urgent Medical Device Correction " letter dated May 23, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to discontinue use of the Edit Contour feature until this software functionality is corrected or replaced.GE Healthcare will correct all affected systems by providing a software upgrade at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 - 055 - 919For other countries, please contact your local GE Healthcare Service Representative. Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 363 | 363 | Worldwide Distribution - USA including the states of AR, CA, CO, FL, GA, IL IN, IA, LA, MD, MA, MI, MN, MO, NJ, NM, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX and VA. and the countries of ALGERIA, AUSTRALIA, AUSTRIA, ARGENTINA, BELARUS, CANADA, CHINA, CHILE, EGYPT CZECH REPUBLIC, FRANCE, FINLAND, GERMANY, GABON, HONG KONG, INDIA, IRELAND, ITALY KENYA, KOREA, JAPAN, MEXICO, NIGERIA, PHILIPPINES, PORTUGAL REPUBLIC OF MALTA, RUSSIAN FEDERATION SAUDI ARABIA, SPAIN SWITZERLAND, TURKEY, TAIWAN, TUNISIA, UNTIED KINGDOM and VENEZUELA. | Terminated | January 28, 2014 | 613 | 363 | Computer
Z-0905-2013 | 64195 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Synthes Multi-Vector Distractor - Multi-axial Mandible Distractor (Left), Part Number 487.967. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that the sheer pin on the Multi-Vector Distractor may not be made according to product specifications. | Device Design | The firm, Synthes (USA) sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firmat 610-719-5450 for questions regarding this notice. |  | 18 | 18 | Nationwide Distribution-including the states of CO, LA, MD, MN, NY, PA, RI, TN, TX and UT. | Terminated | August 31, 2015 | 962 | 88 | Not_Computer
Z-0892-2013 | 64136 | No | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Ad-Tech Medical Instrument Corporation Macro Micro Subdural Electrode. For temporary (&lt;30 days) use with recording, monitoring and stimulation equipment, for the recording, monitoring and stimulation of electrical signals on the surface level of the brain. The recording of electrical activity supports definition of the location of epileptogenic foci and brain mapping. |  |  | December 18, 2012 | 2012 | Ad-Tech Medical Instrument Corporation | The microelectrodes are not flush with the silastic substrate, may contain sharp edges capable of abrading the cortical surface, and may have remaining pieces which may fracture and remain in the cortical tissue with the electrode is explanted. Recall expanded on 5/2/2013 to include additional Macro Micro Subdural Electrodes. | Device Design | AD-TECH Medical Instrument Corporation sent an "Urgent Medical Device Recall" letter dated December 18, 2012 to all affected consignees. The letter described the Reason for the Voluntary Recall, Risk to Health, Actions to be taken by the Customer, Product and Distribution Information, Type of Action by Ad-Tech Medical Instrument Corporation and Other Information. Consignees were instructed to quarantine the affected product and contact the firm for a Return Material Authorization number. If the product have been further distributed consignees were advised to notify their customers of the recall. The letter requested consignees to return the Acknowledgement and Receipt Form.For questions contact your Ad-Tech Clinical Specialist.Expanded recall letters dated May 2, 2013 were sent out April 29, 2013. The letters were formatted as the initial ones sent on 12/18/2013. |  | 144 | 144 (103 US; 41 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, GA, LA, MA, MD, MI, MN, NH, PA, RI, TX, UT, WI, NY IA, WI and the countries of: Australia, CANADA, ENGLAND, JAPAN, KOREA, SPAIN, FRANCE, and SWITZERLAND. | Terminated | October 15, 2014 | 666 | 144 | Not_Computer
Z-0907-2013 | 64214 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE syngo RT Therapist (RTT) version 4.3, ARTISTE Model, Part No. 08162815, Digital linear accelerators, MOSAIQ OIS system, Distributed by SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Manufactured by SIEMENS AG, MEDICAL SOLUTIONS, GERMANY. |  |  | January 17, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential safety issue when using the "In-session resumption" feature of the syngo RT Therapist (RTT) 4.3 software. The RTT will send again the previously completed beam or segment for delivery to the control console, thus repeating the already delivered Monitor Units of the same beam or segment. | Software design | The firm issued "Urgent Medical Device Correction" letters via mail to all its direct consignees, informing them of the affected product issue and providing instructions on the recall. The firm stated that the release of this UI TH005/13/S to distribute the Customer Safety Advisory Notice began on January 17, 2013. |  | 5 | 5 units | 5 devices, all located outside of the United States in Germany. | Terminated | March 05, 2013 | 47 | 5 | Computer
Z-0906-2013 | 64259 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System Extended Brilliance Workspace (EBW), Computed Tomography X-Ray System. |  |  | September 28, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Incorrect values may be displayed. Philips is issuing a software update (version 4.5.5) to the Extended Brilliance Workspace (EBW) systems currently running software version(s) 4.0.2, V4.5.2 or V4.5.3, addressing a number of non-conformances. | Software design | Philips Healthcare sent a Urgent Medical Device Correction letter dated September 28, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.This Field Safety notification is to inform the actions that should be taken by the customer/user to prevent risks for patient or users, and the actions planned by Philips to correct the problem.If you need further information or supplies please contact your local Philips Healthcare office. For North American and Canada contact the Customer Care Solutions Center ( 1-800 722-9377) option 5 enter site or follow the prompts. |  | 2580 | 2,580 units | Worldwide Distribution - USA including the states of AL, AT, AZ, BAH, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NV, NY, OH, OR, PA, PR, PR, SC, TN, TX, VA, VT, WA, WI &amp; WV. and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bolivia, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican, Republic, Ecuador, Egypt, Ethiopia, Finland, France, French Polynesia, Gabon, Georgia, Germany, Greece, Guadeloupe, Guinea, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, LAO People's Dem, Rep., Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Malta, Martinique, Mexico, Netherlands, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela &amp; Viet Nam. | Terminated | December 26, 2013 | 454 | 2580 | Computer
Z-0908-2013 | 64417 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DPL Single Stage Venous Cannula with Metal Tip, used in cardiopulmonary bypass procedures. |  |  | February 19, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Potential for uncharacteristically rough metal edges to occur along the inner diameter of the metal tip side ports of the cannula. | Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent Medical Device Recall" letter dated February 19, 2013. The letter described the problem and the product involved in the recall. Advised consignees to quarantine affected product, and return the product to Medtronic. Requested consignees the complete and return the attached Customer confirmation certificate. If product has been forwarded to another facility, requested consignees to notify the facility of the issue. |  | 13691 | 13,691 (6711 USA; 6980 OUS) | Worldwide distribution, including USA nationwide, Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Kazakhstan, Korea, Libya, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Viet Nam. | Terminated | November 21, 2013 | 275 | 13691 | Not_Computer
Z-0913-2013 | 64000 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer. An automated chemistry analyzer that measures analytes such as Glucose, Magnesium, and Potassium. |  |  | November 28, 2012 | 2012 | Beckman Coulter Inc. | Due to a software data storage limitation which only occurs under certain conditions, measurements of the Reagent Blank (RB) or Calibration (CAL) sample results are not saved and the software reverts to a previously stored, acceptable RB/CAL result. | Software design | Beckman Coulter sent an "URGENT PRODUCT CORRECTION" letter dated November 28, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Call Center Hotline at (800) 854-3633 for questions regarding the notice. |  | 350 | 28 in U.S.; 4 Canada; 318 to other foreign countries. | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Australia, Austria, Brazil, China, Czech Republic, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea Republic, Malaysia, Netherlands, Russian Federation, Spain, Switzerland, Taiwan, Turkey, United Kingdom, and Viet Nam. | Terminated | February 24, 2014 | 453 | 350 | Computer
Z-0912-2013 | 64404 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Atricure Inc., Max1 Isolator Transpolar Pen. Model A000163. Used to ablate cardiac tissues during cardiac surgery using radiofrequency energy when connected to the AtriCure Ablation and Sensing Unit or for temporary cardiac pacing, sensing, recording, and stimulation during the evaluation of cardiac arrhythmias during surgery when connected to a temporary external cardiac pacemaker or recording device. |  |  | January 14, 2013 | 2013 | Atricure Inc | Products from lot no. 41865 may have been sent out with Incorrect Instructions for Use (IFU) labeling. | Employee error | AtriCure sent a Urgent Medical Device Recall/Correction letter dated January 14, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.AtriCure has identified that products from lot no. 41865 may have been sent out with Incorrect Instructions for Use (IFU) labeling. If the MAX1 is used as intended there is no risk to the patient as a result of this issue.The correct instructions for Use for the MAX1 Pen is included with this notification. AtriCure will replace any affected product if requested by the facility.Please review and fill out the attached " Acknowledgement Form" and either fax to AtriCure at (513) 644-4753. If you need to contact customer service, their number is 1-866-349-2342. |  | 5 | 5 devices | Nationwide Distribution including the states of MD, PA and TX. | Terminated | June 11, 2013 | 148 | 5 | Not_Computer
Z-0914-2013 | 64464 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Lifeguard Infusion Set with Smartsite, 20Ga, 1, CLGY-2010 |  |  | January 21, 2013 | 2013 | Churchill Medical Systems, Inc. | The label on the device states that the item number is CLGY-2210 when the item number on the device label should read CLGY-2010. This is a 20Ga Huber needle set; however, the product label states the pouch contains a 22Ga set. | Under Investigation by firm | Vycon sent an Urgent Medical Device Recall letter dated January 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed check their stock for the affected product and immediately cease use and distribution and quarantine all affected product immediately. Count and document their affected inventory on the Recall Acknowledgement and Inventory Return Form. Customers were instructed to fax or email the form to Vygon including customer contact information. All affected product would be credited or replaced at no cost to the customer. Customers with questions were instructed to call 1-800-473-5414, or by e-mail at customerservice@vygonus.com.For questions regarding this recall call 603-743-5988. |  | 2340 | 2340 | Nationwide Distribution including MO, IL, FL, PA, VA, MO, NY, NE, NH, and OH. | Terminated | September 10, 2013 | 232 | 2340 | Not_Computer
Z-0917-2013 | 64367 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | D3000I Advantage Drive System, 230v, surgical instrument motor. |  |  | January 24, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The Membrane Switch Panel (referred to as "front panel" of the component part number A34-003-207, used on both the D3000 and D3000I Advantage Drive System, contained the incorrect symbology per third edition IEC60606-01 Medical Electrical Equipment: General requirements for basic safety and essential performance. | Error in labeling | Linvatec Corp sent an "URGENT: Medical Device Recall Notification" letter dated January 24, 2013 to its customers after initial phone notifications. The letter describes the affected product, problem and actions to be taken. Customers were instructed to immediate check facility's inventory for the affected product, use the enclosed FedEx air-bill when returning the product, and complete and return to the Recall Reply Form. For questions contact your Customer Service department at 800-535-8536 or fax at 727-319-5701 or email Custserv1@linvatec.com. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) including the states of: CA, FL, GA, IL, KS, MA, MD, MO, NE, NY, and TX and the countries of: Canada, Germany, Romania, Spain, Palestine, South Africa, France, Sweden and Singapore | Terminated | July 05, 2013 | 162 | 54 | Not_Computer
Z-0915-2013 | 64440 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers designed for sterilization of heat and moisture-stable materials used in healthcare facilities. |  |  | February 12, 2013 | 2013 | Steris Corporation | The control boards alarm set points and temperature tolerances allowed for erroneous activation of the AD (analog/digital) alarm as temperature fluctuated during operation. When this alarm occurs, the sterilizer locks up and cannot be used until serviced, causing Customer nuisance. | Software design | Steris Corporation initiated their recall on February 12, 2013. They will notify customers via onsite visits by STERIS Field Service Technicians to each affected Customer site.For further questions, please call (440) 392-7601. |  | 395 | 395 Units | Worldwide Distribution - USA including Canada and the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, SK, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | December 05, 2013 | 296 | 395 | Not_Computer
Z-0909-2013 | 64425 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Hickman 9.0 French Dual Lumen Catheter with SureCuff Tissue Ingrowth Cuff with Peel-Apart Percutaneous Introducer system (pediatric product code 0600580; adult product code 0600600). |  |  | February 21, 2013 | 2013 | Bard Access Systems | Bard Access Systems is recalling two lots of Hickman 9.0 French and one lot of Leonard 10.0 French Dual Lumen Catheters with SureCuff Tissue Ingrowth Cuff with Peel-Apart Percutaneous Introducer Systems due to the wrong size introducer being included with the kit. | Mixed-up of materials/components | Bard sent an Urgent Recall Notification letters dated February 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to examine their inventory for the affected product, return any remaining in stock and complete the enclosed Reply Form and Inventory Reconciliation Form. For questions contact Bard Access Systems Customer Service at 1-800-290-1689. |  | 195 | 195 units. | USA Nationwide Distribution and the countries Europe, Belgium, Australia, and Latin America. | Terminated | January 03, 2014 | 316 | 268 | Not_Computer
Z-0911-2013 | 64013 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer PST Gel and Lithium Heparin (LH); 83 Units; Plus Blood Collection Tubes; Sterile; IVD; REF 367962; 4.5 mL, 13 X 100 mm; BD, Franklin Lakes, NJ USA; Made in USA. |  |  | March 05, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer Plus Plastic Plasma Separator Tube (PST) with Light Green/BD Homogard Closure due to a deformed top of the tube reservoir reported in two customer complaints. | Equipment maintenance | BD sent advisory letters via UPS dated 3/5/2012 to customers. The letter identified the affected product and the reason for recall. BD has taken steps to modify the existing manufacturing and assembly processes to help reduce the potential for this condition. Questions regarding this issue should be directed to BD Global Technical Services at 1-800-631-0174 or email at www.bd.com/Vascutainer/contact. |  | 225000000 | 225 million units | Nationwide Distribution | Terminated | July 20, 2015 | 1232 | 225000000 | Not_Computer
Z-0918-2013 | 64343 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PSX PROXIMATE Skin Staple Extractor, |  |  | October 04, 2012 | 2012 | Ethicon Endo-Surgery Inc | Damage to the packaging may have compromised the sterility of the device. | Package design/selection | On 10/28/2012, the firm sent Urgent: Device Recall, Event 5928 notification letters to their customers. The letters identified the affected product, the reason for the recall, the action to be taken by customers, as well as the actions taken by the firm. Customers are to examine their inventory immediately and remove the affected product. The Business Reply Form should be completed and faxed to Stericycle at 1-877-650-0360 within 3 business days, even if there is no affected product on hand. Customers must complete the "Replacement Purchase Order Number Field" on the Business Reply Form to ensure adequate identification of the return product in their facility. Customers should follow the instructions provided in order to return the product. The information should be shared with all appropriate staff at the customer's facility. EES planned to replace any unopened an unexpired recall product that was returned by December 31, 2012. Questions regarding the issue should be directed to Stericycle at 1-877-650-5412. |  | 5894975 | 5,894,975 pieces | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY; and, the countries of ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, NICOSIA, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PROVINCE OF CHINA, PUERTO RICO, REPUBLIC OF KOREA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA &amp;, YEMEN. | Terminated | June 11, 2013 | 250 | 5894975 | Not_Computer
Z-0922-2013 | 63551 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Endotrig Trigger Release Sterile Micro Hook Blade |  |  | September 18, 2012 | 2012 | Instratek, Incorporated | Instratek, Inc.has received multiple complaints of the 1052 Trigger Finger Release Blades which could not be advanced into the cannula through which it is used during trigger finger release procedures. | Nonconforming Material/Component | The firm decided to recall the product and sent notification letters to consignees dated 09/18/2012. |  | 235 | 235 | USA Nationwide Distribution including the states of: AZ, MI, NE, NV, and TX. | Terminated | April 10, 2013 | 204 | 235 | Not_Computer
Z-0920-2013 | 64318 | No | LBK |  |  | Neurology | 510(k) | Integra CUSA excel. CEM Nosecone for CUSA EXcel System. |  |  | January 07, 2013 | 2013 | Integra LifeSciences Corp. | There is a potential for erosion of the CUSA EXcel tip used with the Nosecone. Specifically, the potential for tip erosion occurs when using a CUSA CEM Nosecone if the side of the tip, rather than the annulus of the tip, is in contact with the tissue. This erosion can potentially lead to tip breakage. There were no reports of a patient injury associated with these complaints. | Nonconforming Material/Component | Integra sent a Urgent Medical Device Recall Field Safety Notices dated January 7, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the nature of the issue and how if affects CUSA top erosion/breakage; advises customers and provided a copy of the CEM Nosecone IFU additional Warning; and requess customers to complete and return the Acknowledgement Return Form.Please feel free to contact our service hotline at ( 888 ) 772-7378. |  | 29450 | USA 6,190, EU 22,972 and ROW 288 units | Worldwide Distribution - USA and the countries of Austria, Australia, Bangladesh, Belgium, Bulgaria, Brazil, Switzerland, Canada, China, Costa Rica, Czech Republic, Germany, Dominican Republic, Egypt, Emirates, Spain, Finland, France, Great Britian, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Italy, Japan, Korea (South), Kuwait, Lebanon, Sri Lanka, Latvia, Macedonia, Mexico, Malaysia, Netherlands, Philippines, Pakistan, Poland, Portugal, Qatar, Serbia, Saudi Arabia, Singapore, Slovenia, Thailand, Tunisia, Taiwan, United Arab, Viet Nam and South Africa. | Terminated | August 06, 2013 | 211 | 29450 | Not_Computer
Z-0921-2013 | 64421 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | 3.75mm x 12mm Fixed Cervical Screws. Components of the Genesys Spine Anterior Cervical Plate System |  |  | February 20, 2013 | 2013 | Genesys Orthopedic Systems, LLC | Manufacturing residue may be present on the cervicalscrews. Use of this product may result in patient infection, delayed fusion or non-fusion (pseudarthrosis), inflammatory response, or revision surgery. | Employee error | Genesys Orthopedic Systems, LLC sent a Recall Notification letter dated February 20, 2013 to all affected customers. The letter identified the affected product, problem and necessary actions to be taken. Customers were instructed to immediately return the affected product to the Genesys Spine office and complete and return the enclosed Response Form as soon as possible. For questions call 512-381-7070. |  | 9 | 9 cervical screws | USA Nationwide Distribution including the states of: CA and TX. | Terminated | September 09, 2013 | 201 | 9 | Not_Computer
Z-0926-2013 | 64256 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-arm Imaging System |  |  | January 31, 2013 | 2013 | Medtronic Navigation, Inc. | Potential failure of the braking system that controls the O-arm Imaging System gantry movement in the vertical direction | Component design/selection | Medtronic Navigation issued a Field Safety Notice letter dated January 18, 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. The letter instructed customers not to discontinue use and that a Medtronic Navigation Representative will be contacting them shortly to schedule a time to conduct the motion control box upgrade. For questions contact your local Medtronic Navigation representative or call Technical Services Group at 800-595-9709 or (720) 890-3200, or email at rs.navtechsupport@ medtronic.com. |  | 463 | 463 units | Worldwide Distribution - USA Nationwide and the countries of: Canada, AUSTRIA, BELGIUM, BRAZIL, BULGARIA, CHILE, CHINA, CROATIA, CZECH REPUBLIC, DENMARK,FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN ,KOREA, REPUBLIC OF KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MEXICO, NETHERLANDS, New Zealand, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, Taiwan(R.O.C.) , THAILAND, TURKEY, UNITED ARAB EMIRATES. and UNITED KINGDOM. | Completed |  |  | 463 | Not_Computer
Z-0919-2013 | 64306 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Command Instrument Broach Handle. Used in Total Hip Arthroplasty. Catalog Number: 6266-0-100 |  |  | December 20, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | Reports of handle fracture and subsequent disassociation of cam assembly subcomponents during use. | Process control | Stryker Orthopaedics sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" dated December 20, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (201) 972-2100 for questions regarding this issue. |  | 48 | 48 units | Nationwide Distribution. | Terminated | September 30, 2015 | 1014 | 48 | Not_Computer
Z-0923-2013 | 64477 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Genesis Straight Esthetic 4.0 mm 1mm Cuff Abutment |  |  | July 05, 2011 | 2011 | Keystone Dental Inc | Mislabeled: Genesis Straight Esthetic 5.0mm Contour ,3mm cuff labeled as Genesis Straight Esthetic 4.0 mm, 1 mm cuff abutment, | Mixed-up of materials/components | Keystone issued Market Withdrawal notice via fax on July 5, 2011, to affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were asked to complete and return the fax-back form to facilitate whether they hae any affected product and indicate whether they would like to receive a replacement of a full refund. Customers ewre asked to contact Keyston'es Customer Relations Department to coordinate shipment of the returned product. For questions regarding this recall call 781-328-3390. |  | 35 | 35 units | Nationwide Distribution including CA, IN, MA, MD, MN, MI, , NH, NY, NC, PA, and WI. | Terminated | March 11, 2013 | 615 | 57 | Not_Computer
Z-0925-2013 | 64354 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF S313152 Straight Acetabular Inserter Handle Ringloc |  |  | February 05, 2013 | 2013 | Biomet, Inc. | The Straight Ringloc Button Latch Inserter Handle may have been manufactured at the higher end of the tolerance causing interference with the Ringloc cup. The inserter handle may lock onto the cup during impaction. If alternative inserters are not available this could cause a surgical delay greater than 30 minutes. | Process control | BIOMET sent an Urgent Recall Notice beginning February 13, 2013, to all affected customers. The notice included possible adverse events and directions to locate and remove identified devices by returning a Fax Back Response Form to 574-372-1683 and obtaining an RGA # prior to return of product to Biomet, Inc. 56 East Bell Drive Warsaw, IN 46580. Questions should be directed to 574 372-1570 M-F, 8 a.m.-5.p.m.For questions regarding this recall call 574-267-6639, ext 1676. |  | 58 | 58 | Worldwide Distribution - No USA distribution -Internationally to Canada, Japan, and Chile | Terminated | September 30, 2013 | 237 | 58 | Not_Computer
Z-0927-2013 | 64348 | No | LSR | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | ANTI-BORRELIA (LYME) MICROPLATE EIA Kit, Catalog # 32507 Kit Lot Number 120960. Package Insert Supplement: Calibrator-1206 U/mL, Positive Control-228 - 532 U/mL, and Negative Control-&lt;120 U/mL; Manufactured by Innominata dba GenBio for Bio-Rad. The kit is used as an aid in diagnosis of Lyme disease. |  |  | February 11, 2013 | 2013 | Innominata Dba Genbio | Multiple incidents of low-reacting assay plates. | Nonconforming Material/Component | The firm, Innominata dba GenBio, sent an "Urgent Product Correction" letter dated February 11, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review the listed questions to determine sites risk and take the appropriate actions. For questions regarding this issue, customers may contact Bio-Rad Technical Support at : 1-800-224-6723, option #2, and then option #3 for assistance. |  | 195 | 195 | Nationwide distribution: CA, IL, MA, MD, ME, NC, ND, NJ, NY, OR, VA and WI. | Terminated | August 27, 2013 | 197 | 195 | Not_Computer
Z-0928-2013 | 64344 | Yes | MRM |  |  | Cardiovascular | PMA | Isoline Implantable Defibrillation Lead, Model Isoline 2CR5, 2CR6 and 2CT6. |  |  | January 28, 2013 | 2013 | Sorin Group USA, Inc. | Isoline defibrillation leads distributed may have internal insulation breach. | Device Design | Sorin C.R.M., srl sent an Urgent Medical Device Field Safety Notice dated January28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return any unused products by certified mail with signed acknowledgment form, patients should continue to be monitored by their physician field reps during notification visit, and customers should make sure all personnel involved in the management of patients implanted with affected product are aware of the information described in the Field Safety Notice. Customers with questions should contact their local Sorin representative at 1-610-304-9728.For questions regarding this recall call 303-467-6222. |  | 1376 | 1376 Isoline leads | Worldlwide Distribution - USA (Nationwide) and worldwide to Argentina, Australia, Austria, Belgium, Brazil, Canada, Cuba, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Iran, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Morocco, The Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Ukraine, Uruguay and the United Kingdom. | Terminated | December 09, 2015 | 1045 | 1376 | Not_Computer
Z-0930-2013 | 64231 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 Automated Breast Volume Scanner (ABVS); Model number 10042541 |  |  | January 18, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | On some ACUSON S2000 ABVS systems, the inside of the transducer pod may contain sharp edges. | Process design | Siemens sent a Customer Safety Advisory notification letter to Siemens Service Personnel on January 18, 2013 and which would be provided to all affected sites during their visit to correct the problem. The notification identified the product, the problem and the action to be taken. The local Customer Service Engineer will perform the test at their site. If the system passes the test, no further action is required. If the ABVS pod fails the inspection, the Service Engineer will correct the sharp edge.For questions regarding this recall call 650-694-5398. |  | 214 | 214 affected devices | Worldwide Distribution - USA including WV and OH; Internationally to Australia, Austria, Bahrain, Belgium, Brazil, China (P.R.) Estonia, France, Germany, Greece, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Macedonia, Malaysia, Mexico, Netherlands, Peru, Philippines, Poland, Portugal, Qatar, Rep of Korea, Romania, Sweden, Taiwan, Thailand, Turkey, Unite Arab Emirates, United Kingdom and Venezuela. | Terminated | July 16, 2013 | 179 | 214 | Not_Computer
Z-0929-2013 | 64015 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO Sensar Intraocular Lenses (IOLs), Model AR40e, |  |  | December 20, 2012 | 2012 | Abbott Medical Optics Inc (AMO) | Two production orders of these IOLs that contain mislabeled IOLs. | Process control | AMO sent an Urgent Field Safety Notice letter to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to stop using and remove affected product from inventory. The letter advice customers to complete and return the enclosed Facsimile Form. For questions contact AMO Customer Service Representative. |  | 155 | 155 units | Worldwide Distribution - USA Nationwide including the state of OH and the countries of: Germany, Switzerland, Sweden, Italy, China, France, Singapore, UK, Croatia, Hungary, Turkey and Brazil. | Terminated | January 08, 2014 | 384 | 155 | Not_Computer
Z-0931-2013 | 64165 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI. Radiation Therapy Digital Imager. Product XVI R3.5.1, R4.2.1, and R4.5.1 |  |  | December 19, 2012 | 2012 | Elekta, Inc. | The DICOM tags do not contain sufficient information for XVI to identify the different volumes that have the same name. | Software design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated December 2012 to all affected customers. The letter identifies the product, problem, and the actions to be taken by the customers. The IFSN informs the user in the TPS, to give the different volumes different structure names. After importing the patient data set from the TPS, make sure that the structure contours are correct. Users are requested to return the signed acknowledgement. Contact your local Elekta office for questions regarding this notice. |  | 164 | 164 | Nationwide Distribution-DC, Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI. | Open, Classified |  |  | 164 | Not_Computer
Z-0932-2013 | 62993 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Cynch Lordotic T-PLlF Implant. Intervertebral fusion device with bone graft, for use with supplemental fixation. |  |  | August 24, 2012 | 2012 | Spine Smith Partners LP | Implants were incorrectly laser marked as 23mm instead of 28mm. | Labeling mix-ups | Spine Smith sent a Field Action Notice dated August 24, 2012 to its affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 512-637-2097 for questions concerning this notice. |  | 3 | 3 units | Distributed in Texas. | Terminated | May 17, 2013 | 266 | 3 | Not_Computer
Z-0941-2013 | 64159 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Large Combination Clamp/Large External Fixation System, Part Number 390.005 |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | The L1 hex nuts, used to tighten the Large Combination Clamp, were oversized and the 11 mm sockets and wrenches will not fit on the hex nuts. | Nonconforming Material/Component | Synthes sent an Urgent Notice: Medical Device Recall letter dated January 11, 2013 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to examine their inventory for the affected product(s), remove them from inventory and return them to Synthes. For questions call 610-719-5450 or email FieldAction@synthes.com |  | 136 | 136 | USA Nationwide Distribution including the states of AL, AR, AK, CA, FL, ID, KS, LA, MD, MA, MO, NJ, NY, OH, PA, TN, TX, and VA. | Terminated | August 31, 2015 | 962 | 136 | Not_Computer
Z-0939-2013 | 64447 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System (IACS) Monitoring Solution. Catalog Number MS20783 (Infinity MCable-Dual Hemo) |  |  | December 13, 2012 | 2012 | Draeger Medical, Inc. | The displayed value for the invasive blood pressure measurement of the IACS was different than the actual patient values. Liquid ingress in the cable connection may cause inaccurate measurements. This can happen if the Dual Hemo MCable is used in a non-vertical position, a vertical position but upside down, or laying it on the floor or on the patient bed. | Labeling design | Drager sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 1-800-543-5047 for questions relating to this notice. |  | 260 | 260 | Distributed in the states of CA, IL, and MA. | Terminated | May 20, 2013 | 158 | 260 | Not_Computer
Z-0935-2013 | 64291 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR450w, Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, Signa Vibrant, and GE 1.5T Signa HDe MR System. |  |  | May 10, 2011 | 2011 | GE Healthcare, LLC | An error on the SR Viewer Reporting Tool in which edits can be made on a report and can be saved into incorrect reports. | Software design | GE Healthcare sent an Urgent Field Safety Notice dated May 5, 2011, to all affected customers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to open only one patient report at a time when editing a report in SR Viewer.US Customers did not receive a customer letter, however all US systems have been updated.For questions customers were instructed to contact their local Sales/Service Representative.For questions regarding this recall call 262-513-4122. |  | 4026 | 4026 (1672 US; 2354 OUS) | Worldwide Distribution - USA (All States in continental USA including PR except RI.). Internationall to:ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BELARUS, BELGIUM, BULGARIA, BOSNIA &amp; HERZEGOVINA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY, GREECE, GUADELOUPE, HONDURAS, HUNGARY, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF IRAN, IRAQ, ITALY, IRELAND, ISRAEL, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, KAZAKHSTAN, LATVIA, LEBANON, LIBYA, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, MACEDONIA, MARTINIQUE, MOROCCO, NETHERLANDS, NORWAY, PAKISTAN, PERU, PHILIPPINES, REUNION, NEW ZEALAND, OMAN, POLAND, PORTUGAL, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SINGAPORE , SLOVAKIA, SOUTH KOREA, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, URUGUAY, VIET NAAM, and YUGOSLAVIA. | Open, Classified |  |  | 4026 | Not_Computer
Z-0934-2013 | 64274 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ. Cardiovascular X-ray imaging |  |  | February 04, 2011 | 2011 | GE Healthcare, LLC | Exposed high voltage wires inside the cabinet . On a cable located in the system cabinet, two of the shrink tubings covering the 400V live cable leave an area potentially unprotected. | Nonconforming Material/Component | No consignee letter sent out. GE Healthcare's recall strategy was to replace the impacted cable on all affected Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ systems that were distributed. At this time, deployment of this recall is considered complete as of 9/6/2011. |  | 48 | 48 | Worldwide Distribution-USA including the states of FL, IL, IN, NJ, NC, OK, TX, and WI, and the countries of UNITED KINGDOM, TUNISIA, SERBIA, RUSSIAN FEDERATION, NORWAY, NIGERIA, LITHUANIA, LEBANON, KOREA, REPUBLIC OF KOREA, ITALY, INDONESIA, INDIA, FRANCE, FINLAND, COLOMBIA, CHINA, and BRAZIL. | Terminated | January 27, 2014 | 1088 | 48 | Not_Computer
Z-0936-2013 | 64326 | No | OEI | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bacterin Elutia, Hubless Coated Silicone Surgical Wound Drain 24 Fr Round Hubless, Full-Fluted Coated with Silver Sulfadiazine STERILE, Box of 10 Part #7118. Indicated for use in surgical wound drainage. |  |  | January 28, 2013 | 2013 | Bacterin International, Inc. | A portion of a silicone wound drain was found to be cracked. | Nonconforming Material/Component | Consignees were informed of the recall via e-mail beginning 1/28/13 with a request for a return response. This was followed by a "Voluntary Product Recall Information" letter dated 2/13/13 that described the product issue and provided recommended actions. The firm asked for all product to be returned to them. |  | 60 | 15 boxes (10/box) to 4 consignees in US and 81 boxes (10/box) to 1 foreign consignee | 13 US consignees located in CA, CO, MT, TX, UT, and VA. This represents 4 customers and 9 other individual contacts in USA that received marketing samples. Plus 4 Bacterin representives nationwide received samples.Foreign Consingee | Terminated | July 08, 2013 | 161 | 60 | Not_Computer
Z-0937-2013 | 64458 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems, Inc. A Vygon Company - All sizes of CMS-CT6T225 6F Triple Lumen CT PICC with M.S.T. Components, Peripherally Inserted Central Catheter. |  |  | December 17, 2012 | 2012 | Vygon Corporation | Product was packaged with the incorrect introducer needle. The kit label indicates a "safety" needle is included; the kits were packaged with a non-safety needle. | Component change control | Churchill Medical Systems sent a letter to their consignees on December 19, 2012. |  | 210 | 210 | Product was distributed to 4 hospitals (under one hospital network) within the State of Pennsylvania | Terminated | January 14, 2014 | 393 | 210 | Not_Computer
Z-0933-2013 | 64063 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Total Iron Binding Capacity (IBCT) Flex Reagent Cartridge. |  |  | December 20, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | The analytical sensitivity claim of 6 ug/dL (1.1 umol/L) in the Instructions for Use (IFU) is no longer valid. Siemens is conducting a voluntary corrective action for Dimension Total Iron Binding Capacity (IBCT) to modify the analytical sensitivity claim to 36 ug/dL (6.44 umol/L). | Process design | Siemens Healthcare sent a "Urgent Medical Device Correction" letter dated Dec 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were notified that theanalytical sensitivity of 6 ug/dL (1.1 umol/L) in the Dimension(R) IBCT FLex(R) Reagent Cartridge instructions for us (IFU) is no longer valid. Please make adjustments to middle wave products or Laboratory Information Systems ( LIS ) as appropriate for your Laboratory.Please complete the attached form and fax it to (302) 631-8467, to indicate you have received this information.If you have any questions please contact your Siemens Technical Solutions Center or your local Siemens technicial support representative at (302) 631-6311. |  | 19503 | 19, 503 cartons | Nationwide Distribution | Terminated | August 26, 2014 | 614 | 19503 | Not_Computer
Z-0938-2013 | 64173 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes Spine 5.2mm Titanium Click ' X Pedicle Screw Dual Core 40 mm Thread Length, Part Number 487.072. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | Lot number 7542279, of the 45 mm Click 'X Pedicle Screw (Part number 487.072), was incorrectly etched as "Click 'X Pedicle Screw 05.2 L40 TAN violet" although the actual length of the screw is 45 mm. | Employee error | The firm initiated a recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter, dated January 11, 2013, to all consignees. The letter identified the affected product and the issue, as well as requested that consignees examine their inventory for the product(s) with the identified Part and Lot Numbers, remove them from use, and return them to Synthes. Customers were to follow the steps provided in the letter regarding if they had or did not have any of the identified affected product. Questions were to be directed to 610-719-5450 or Fieldaction@synthes.com. |  | 18 | 18 | Distributed in the states of AL, CA, MS, MO, NJ, OR, and TN. | Terminated | September 10, 2015 | 972 | 18 | Not_Computer
Z-0942-2013 | 64287 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Synthes Titanium Midface Distractor, MR Conditional, Part number 487.982. For temporary stabilization and gradual lengthening of the cranial or midfacial bones. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | There is the potential for the anterior footplates to not fit onto the Titanium Midface Distractor Assembly as the tab of some distractors may be oversized. | Employee error | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to all consignees. The letter identified the affected product and the problem. Also, the firm requested consignees to examine their inventory for products(s) with the specified part number, remove them from use, and return them to Synthes. Customers were to follow the instructions provided on what to do if they have or do not have any identified affected product. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 596 | 596 | Worldwide Distribution -- USA, including the states of CA, LA, MA, NY, and PA, and the countries of Canada and Switzerland. | Terminated | August 31, 2015 | 962 | 596 | Not_Computer
Z-0946-2013 | 64479 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Nicolet EEG Wireless Amplifier 32/64 Channel. Catalog/Part Numbers : 515-015400 for the 32 channel amplifier and 515-015500 for the 64 channel amplifier. |  |  | February 27, 2013 | 2013 | Natus Neurology Incorporated | Natus Neurology Incorporated is recalling the Nicolet EEG Wireless Amplifier which may overheat and become uncomfortable to touch around the area of the on/off button. The overheating may result in discomfort ranging from too hot up to the potential for a burn. | Nonconforming Material/Component | Natus sent an "Urgent Medical Device Correction and Removal Notification' letter dated February 27, 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter advised customers that a Natus Field Service technician will remove and replace the affected product at no charge to the users. The letter also advised if any of the affected products have been transferred to another location to please forward a copy of the letter and to notify Natus. For questions call 800-356-0007 ext. 5129. |  | 114 | 114 (31 USA; 83 OUS) | Worldwide Distribution - USA Nationwide including the states of CA, FL, LA, MD, MS, MO, NJ, NY, PA, TX and the country of AUSTRALIA.BAHRAINBRAZILCANADACountry DENMARKEGYPTGERMANYHONG KONGINDIAINDONESIAIRELANDITALYJAPANKOREAKUWAITMALAYSIAMEXICONETHERLANDSROMANIARUSSIASINGAPORESPAINSWITZERLANDTAIWANTHAILANDTURKEYUNITED KINGDOMURUGUAYVIET NAM | Terminated | April 25, 2014 | 422 | 114 | Not_Computer
Z-0940-2013 | 64423 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | GOPump Elastometric infusion PumpKit with/ and without accesssories all containing Dual GOPump, part number 510076: One infusion pump comprising two, 150 mL elastomeric reservoirs, each of which connects to its own catheter. Liquid flow rate is labeled at 2 mL/hr per elastomer. |  |  | February 25, 2013 | 2013 | Symbios Medical Products, LLC | Received 2 complaints out of 458 distributed pumps in this lot, where flow restrictor bead became displaced, which were reported to permit fast flow of contents. | Pending | Symbios sent an Urgent Medical Device Recall letter dated February 25, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to follow the enclosed 'Steps for Voluntary Field Action' in their entirety. These steps contain information about: I) identifying the product to be returned, 2) obtaining a Returned Goods Authorization number, and 3) returning the recalled product. If product(s) is (are) located, please call 317-450-0145 from 7:00 a.In '0 5:00 p.m. EDT to receive a Returned Goods Authorization umber ("RGA"). Please indicate the RGA number on the outside of the package. Attached is a Verification Form. Customers were instructed to complete and return the form even if they do not have any product to return. Their local sales representative can assist them in completing this form. Distributors of the affected product were instructed to contact Symbios at the number provided and forward the recall notice to all their affected accounts.For questions regarding this recall call 317-225-4447. |  | 458 | 458 | Nationwide Distribution including PA, CA, OH, AL, FL, WV, TN, NJ, NC, TX, CO, NY, TX, MO, MS, SC, and GA. | Terminated | February 13, 2014 | 353 | 458 | Not_Computer
Z-0943-2013 | 64342 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | INTERNAL HEX RECON SCREW, TI-6AL-4V, REF 7162380, TRIGEN (R) INTERNAL HEX CAPTURED RECON SCREW, QTY: (1), 6.4 MM, 80 MM, TITANIUM, Smith &amp; Nephew, Inc. Orthopedic. Bone screw for intramedullary rod fixation. |  |  | February 07, 2013 | 2013 | Smith &amp; Nephew Inc | 75 mm screws were packaged and labeled using an 80 mm screw outer box and label, while 80 mm screws were packaged and labeled using a 75 mm screw outer box and label. | Labeling mix-ups | Smith &amp; Nephew sales representatives and Smith &amp; Nephew international distributors were notified of problem and the recall via email on February 7, 2013. They were instructed to cease use and distribution of the product and quarantine for return. Non responding consignees were notified again on February 28, 2013 by email. Hospitals were notified by letter beginning on February 7, 2013. Smith &amp; Nephew also contacted the surgeons who implanted the screws by letter the week of February 11th.For questions regarding this recall call 978-749-1330. |  | 69 | 69 units | Worldwide Distribution - USA including NM, TN, TX, NC, CA, MI, KY, MA, NJ, OH, and HI. Internationally to Canada, United Arab Emirates, Mexico, Argentina, Japan, and Chile. | Terminated | January 11, 2016 | 1068 | 159 | Not_Computer
Z-0947-2013 | 64178 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Advantx, Innova2000, Innova3100, Innova4100, Innova2100IQ, |  |  | January 21, 2013 | 2013 | GE Healthcare, LLC | A locking nut might loosen out of the overhead Video Monitor Suspension yoke affecting Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, and Innova 4100IQ Cardiovascular X-ray imaging systems. | Nonconforming Material/Component | GE Healthcare sent "URGENT MEDICAL DEVICE CORRECTION" letters dated January 18, 2013 and February 11, 2013 to all affected customers. The letters were addressed to Hospital Administrators / Risk Managers, Mangers of Radiology / Cardiology and Radiologists / Cardiologists. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information.The second letter dated February 11, 2013 was sent out to consignees which incorporated an additional model number (2270677-2) which was initially missed. Contact the firm 800-437-1171 for questions regarding these letters. |  | 297 | 297 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, ME, MD, MA, MN, MS,MO, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT , VA, WA, WV, WI and the countries of VENEZUELA, UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, SPAIN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, POLAND, NORWAY, NIGERIA, NICARAGUA, MOROCCO, MALTA, LEBANON, KAZAKHSTAN, JAPAN, ITALY, INDONESIA, INDIA, HUNGARY, GREECE, FRANCE, EGYPT, CHINA, CANADA, BRAZIL, BELGIUM, and NEW ZEALAND. | Terminated | January 16, 2014 | 360 | 297 | Not_Computer
Z-0945-2013 | 64401 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 6, 10, 16, 40, 64, Big Bore and the iCT, iCT SP, Ingenuity CT and MX8000 Dual v. EXP are Computed Tomography X-Ray Systems. |  |  | December 20, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field stating the patient table on a system had an unexpected downward movement while a patient was on the table. Philips has determined the cause was a malfunction of the vertical brake. | Component design/selection | Philips Healthcare sent a "Urgent Medical Device Correction" letter dated December 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed Philips received reports from the field stating that the screws holding the vertical brake together failed.A Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. If you need any further information please contact your local Philips representative or local Philips Healthcare office. For North America and Canada Contact the Customer Care Solutions Center at ( 1-800-722-9377, option 5: Enter Site ID or follow the prompts. |  | 4797 | 4797 Units | Worldwide Distribution - Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY. | Terminated | April 17, 2014 | 483 | 4797 | Not_Computer
Z-0948-2013 | 64376 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens Somatom CTs equipped with software syngo CT2010A, syngo CT2010B, syngo CT2011A and CT2012B. |  |  | November 27, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Sporadic hang-up malfunctions concerning 3D reconstructions during activated "Preview Image" functionality. | Software design | Siemens sent a "CUSTOMER SAFETY ADVISORY NOTICE" dated November 27, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The correction for this problem will be available with Service Pack 1 for syngo CT2011A And Service Pack 4 for syngo CT2012B. Both Service pack are planned to be released starting February 2013. Systems still running on syngo CT2010A or syngo CT2010B are already scheduled to be upgraded to syngo CT2011A or syngo CT2012B respectively. Contact your local Siemens Application Specialist for additional information. |  | 896 | 896 | Nationwide Distribution. | Terminated | January 31, 2014 | 430 | 896 | Not_Computer
Z-0949-2013 | 64580 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare SpO2RT2 bed. |  |  | February 26, 2013 | 2013 | Hill-Rom, Inc. | In certain situations a software problem with the TotalCare SpO2RTÂż 2 bed may occur that can lead to inadvertent loss of sufficient surface cushionpressure to adequately support the patient. This may increase the likelihoodof skin irritation for patients at higher risk of developing pressure-related skin conditions. A correction to the software program is currently being developed andtested, | Software design | Hill-Rom sent an Urgent Field Safey Notice letter dated February 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter advised customers to follow steps provided until the updated software is available. Customers were requested to forward a copy of this letter to any other facility personnel you deem appropriate. Questions should be directed to Hill-Rom Technical Support at 800-445-3720 |  | 309 | 309 | USA Nationwide Distribution including the state of Canada | Terminated | April 16, 2014 | 414 | 309 | Computer
Z-0950-2013 | 64436 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 and Ingenuity Computed Tomography X-Ray Systems |  |  | February 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | This field change order is being released to update software and customer release notes to the affected installed base. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated February 13, 2013, via certified mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review the information contained in the letter with all members of their staff who need to be aware of the contents of the letter. Customers were instructed to retain a copy with the equipment Instruction for Use. For questions customers were instructed to contact their local Phillips representative or local Philips Healthcare office. For North Americal and Canada contact Customer Care Solutions Center at 1-800-722-9377, option 5: Enter Site ID or follow the promptsFor questions regarding this recall call 1-800-722-9377. |  | 42 | 42 Units | Worldwide Distribution - USA including AR, AZ, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, WY, and worldwide including the contries of Australia, Belgium, China, France, Germany, Israel, Japan, Norway, Sweeden, and Switzerland. | Terminated | April 11, 2014 | 420 | 42 | Computer
Z-0951-2013 | 64320 | No | GJT | Factor deficiency test. | Hematology | Hematology | 510(k) | Vial Label: |  |  | January 28, 2013 | 2013 | Diagnostica Stago, Inc. | Diagnostica Stago has confirmed that sometimes the delta OD measured can be changed leading to a lowered percentage result. | Software design (manufacturing process) | Diagnostica Stago Inc. sent e-mail notifications on 1/28/2013 to all consignees (laboratory supervisor). The written field safety notice, entitled Important Field Safety Notice, with its response form was attached. The letter identified the affected product and gave a description of the defect. Customers were asked to confirm all the results obtained in the abnormal range by either of the two methods provided. Users are to complete and return the enclosed response form as soon as possible to the Stago Hotline via fax, 973-644-9348. Questions should be directed to the Stago Hotline at 800-725-0607. |  | 2901 | 2901 kits | Worldwide Distribution - USA and Canada. | Terminated | July 19, 2013 | 172 | 2901 | Not_Computer
Z-0958-2013 | 64162 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Creatinine, Part Numbers OSR6178, OSR6678 |  |  | December 28, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. | Labeling design | Beckman Coulter sent an Urgent Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed the customers that Beckman Coulter is including additional information and instructions in the Creatinine (OSR6x78) IFU and Setting Sheet, relating to calibration stability and quality control frequency to assure product performance. Customers are instructed to complete and return the enclosed response form. Customers with questions about the notification letter are instructed to contact the call center (hotline) at (800) 854-3633 in the US and Canada. |  | 57962 | 57,962 units (50,452 units in the US) | Worldwide Distribution - USA Nationwide and the countries of: Canada, Brazil, Colombia, Mexico, Puerto Rico, Venezuela, Guam, Honduras, Peru, Ecuador, Costa Rica. | Terminated | October 17, 2013 | 293 | 57962 | Not_Computer
Z-0956-2013 | 64194 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Compression Plate/VA-LCP Dorsal Distal Radius Plate, Part Number 02.115.431. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | It was discovered that one of the VA-LCP Dorsal Distal Radius Plates was incorrectly etched as Part Number 02.115.431 (110 degree) instead of the correct Part Number 02.115.231 (90 degree). | Labeling mix-ups | The firm initiated their recall of this product on January 11, 2013 by sending an Urgent Notice: Medical Device Recall letter to consignees. The letter identified the affected product and the issue. Synthes requested customers to examine their inventory for the affected product, remove them from use, and return them to Synthes. Additionally, customers were to follow the steps provided in the event that they did or did not have any affected product on hand. Questions should be directed to 610-719-5450 or FieldAction@synthes.com. |  | 4 | 4 | Distributed in the states of IN, OH, PA, and TX. | Terminated | August 31, 2015 | 962 | 4 | Not_Computer
Z-0954-2013 | 64452 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Spine Polaris 5.5 Button Lock Screw Inserter. Catalog Number 14-500178. The Polaris Spinal System is a non-cervical, pedicle screw spinal fixation device. |  |  | February 07, 2013 | 2013 | Ebi, Llc | Increased risk of failure at the driver tip during screw insertion or removal. | Process control | Biomet Spine sent a "Urgent Medical Device Recall" notices dated February 6, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately locate and remove affected items from circulation and return affected product to EBI, LLC.Please follow the instructions on the enclosed "FAX Back Response Form, "and fax a copy of the Response Form to (973) 257-0232, prior to return of the items.Questions related to this notice should be directed to (973) 299-9300 ex. 2322.Monday through Friday, 8am to 5pm. |  | 59 | 59 units. | Nationwide Distribution including the states of AL, IN, MO, NE TN and TX. | Terminated | December 07, 2015 | 1033 | 59 | Not_Computer
Z-0957-2013 | 63767 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Siemens Dimension(R) Tacrolimus Flex(R) reagent cartridge (DF107). The TACR method is used as an in vitro diagnostic test. |  |  | November 27, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed customer complaints of low patient sample recovery with Dimension(R) TACR Flex(R) lot FA3085. Quality Control (QC) materials have not exhibited low recovery of the same magnitude.Internal method regression testing versus LC/MS has shown an average 25% low bias compared to LC/MS across the assay range. Individual patients, particularly samples with lower TACR values | Process control | The firm initiated their recall of this product on November 27, 2012 by sending a letter to all consignees. The letter included a recall effectiveness check form to be completed by the consignee and returned to the firm. The firm expanded their recall on January 29, 2013 by sending another letter to consignees. This letter also included a recall effectiveness check form to be completed by the consignee and returned to the firm. |  | 2976 | 2,976 | Product was distributed within the United States to AL, AZ, CA, DC, FL, LA, MD, MI, MO, NJ, NM, NC, OH, OR, PA, TN, TX, UT, VA, and WA. | Terminated | August 04, 2014 | 615 | 2976 | Not_Computer
Z-0953-2013 | 64465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Logic Comprehensive Knee Prosthesis System, Proximal Tibial Spacer and screws, size 3.5, 8mm, Catalog # 02-012-42-3508. |  |  | February 01, 2013 | 2013 | Exactech, Inc. | A dimensional mismatch was identified which can potentially allow for interference between devices. | Device Design | Exactech Inc. sent a *** Important *** Product Market Recall Notice dated January 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Customers were instructed to immediately cease distribution or use of these products. 2. Extend this information to your accounts that may have this product in their possession.3. Verify if you have any of the subject Optetrak Logic Proxima Tibial Spacer (PTS) in the specified lots.Please complete the attached fax back form and return it to Exactech within the next 5 working days. Thank you for your prompt attention to this matter. For further questions please call (352) 377-1140. |  | 123 | 123 devices | Nationwide Distribution including the states of FL, AL, LA, PA, VA, IL, OH, NY, NC, KS, TN, NJ, CO, GA, CA, MD, MN, IN, &amp; Puerto Rico. | Terminated | September 22, 2014 | 598 | 123 | Not_Computer
Z-0952-2013 | 64330 | No | HXI | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Extremity Medical - Disposable Orthopedic Surgical Instrument - 1.6 mm Guidewire - Reference Number 101-00006 - Non-Sterile. Contained in Extremity Medical Surgical Kits including CompressX, Hallu.X, IO Fix and XMCP. Used in the selection of screw size when used with appropriate depth gauge. |  |  | January 10, 2013 | 2013 | Extremity Medical LLC | The 1.6 mm Guidewire is undersized up to 2 mm in length. When used with the depth gauge, the measurement reading will not correspond with the desired screw length. The selected screw size based on the depth gauge reading will be about 2 mm longer than intended. | Component change control | Extremity Medical sent an "URGENT: FIELD SAFETY NOTICE" dated January 10, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Response Form was included for customers to complete and return. Contact Customer Service at 1 (973) 588-8980 for questions regarding this notice.On April 10, 2013, Extremity sent notifications to their customers at the surgeon level. They were asked to complete the attached product field action response form. |  | 238 | 238 units | Worldwide Distribution-USA including the states of Florida, Pennsylvania, New Hampshire, Oklahoma, California, Maryland, Arizona, Ohio, North Carolina, Texas, Virginia, Illinois, New York and Tennessee, and the countries of Australia, Belgium, Israel, Netherlands and United Kingdom. | Terminated | July 18, 2013 | 189 | 238 | Not_Computer
Z-0967-2013 | 55155 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | ILLICO Straight CP TI Rod, -5.5MM x 110M; REF#73601-110; Qty 1; LBL-005; REV.B; RX ONLY; Alphatec Spine, Inc. The ILLICOÂż MIS Posterior Spinal Fixation System is intended to facilitate the surgical correction of non-cervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. |  |  | July 06, 2011 | 2011 | Alphatec Spine, Inc. | The rods were incorrectly laser etched as ILLICOÂż Straight CP Tl Rods, 110mm, when they were in fact manufactured to 100mm in length. | Process control | Alphatec Spine sent an Urgent Medical Device Recall letter dated July 6, 2011, to all affected customers. This notification described the product, problem and actions to be taken. The consignees were instructed to review their inventory to verify the affected device; abstain from use; immediately contact Alphatec Spine Sales Administration (salesadministration@alphatecspine.com) for instructions on how to return this affected product; and complete, sign and return confirmation form to Alphatec Spine, Inc., 5818 El Camino Real, Carlsbad, CA 92008.If you have any questions, call 760-431-9286. |  | 7 | 7 | Worldwide Distribution - USA including CO and NC; Internationally to Belgium. | Terminated | March 20, 2013 | 623 | 7 | Not_Computer
Z-0959-2013 | 64509 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Allen Traction Boot, Modular, OT 1000 Traction Device Boot Assembly, Fracture Table Orthopedic Leg Holders, non-invasive medical devices designed to attach to orthopedic fracture tables to aid clinical staff in applying traction to the leg of a patient during surgical procedures. |  |  | February 28, 2013 | 2013 | Allen Medical Systems, Inc. | Some units may have been assembled incorrectly which could result in the boot breaking free from the mounting base and allowing the patient's leg to fall. | Process control | Allen Medical Systems sent a "Urgent Medical Device Recall" letter dated February 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken. Customers were instructed to: 1. Immediately remove from service your Allen Traction Boot with the serial number(s) XXXXX 2. Contact Allen Medical Customer Service at (800) 433-5774 for an RMA number and instructions to return your Allen Traction Boot for inspection and repair if necessary. There will be no cost to you. 3. Return your Allen Traction Boot to Allen Medical as soon as possible. In order to minimize the inconvenience, we will repair your Allen Traction Boot and return the device to you within 48 hours of receipt. |  | 24 | 24 devices | Worldwide Distribution- USA including the states of TX, NV, KY, GA, OH, PA, NY, AL, WA, IL and the countries of New Zealand and Japan. | Terminated | August 07, 2013 | 160 | 24 | Not_Computer
Z-0969-2013 | 64471 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sterile Access &amp; Return Line Set, 36" Length, Part # UM-872-36. |  |  | June 19, 2012 | 2012 | Molded Products Inc | The recalled product leaks at the connection site. | Process control | The firm's one customer was notified of the recall by phone call on 5/8/2012. The customer returned the affected product once the replacement product was received. |  | 1200 | 1200 sets | Distributed to one customer in Michigan. | Terminated | August 14, 2013 | 421 | 1200 | Not_Computer
Z-0965-2013 | 64546 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific/Cameron Health SQ-RX Model 1010 subcutaneous pulse generators. The S-ICD System is intended to provide therapy for the treatment of life-threatening ventricular tachyarrhythmias. |  |  | March 01, 2013 | 2013 | Boston Scientific CRM Corp | An internal protective fuse can be unintentionally activated while the device is charging its capacitors for shock delivery or induction. Should this occur, the defibrillator would not be able to deliver therapy or communicate with the Q-TECH Model 2020 programmer, and would be unable to emit tones or otherwise respond to magnet application. | Software design | Consignees were sent a Boston Scientific "Important Medical Device Information" letter dated March 1, 2012. The letter was addressed to "Dear Doctor". The letter described the product involved in the recall, the problem and provided Recommendations/Actions. US customers may contact Technical Services for more information at 800-227-3422. |  | 2090 | 2090 (570 USA; 1520 OUS) | Worldwide Distribution, USA including the states of AZ, CA, CO, DC, GA, IL, MD, MA, MN, NJ, NY, NC, OH, PA, SC, VA, WA and the countries of AUSTRIA, BELGIUM, CANADA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IRELAND, ITALY, NETHERLANDS, NEW ZEALAND, PORTUGAL, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, and the UNITED KINGDOM. | Terminated | May 23, 2013 | 83 | 2090 | Not_Computer
Z-0964-2013 | 64456 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp. Stryker Orthopaedics Specialty Triathlon Left and Right A//P Sizer Adjustment Housing PER File K2401. For use in knee prosthesis implantation. NON-STERILE |  |  | January 24, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | The standard Triathlon Femoral Stylus does not fit into the Specialty Sizers and the "R" and "L" markings on the Right Sizer are reversed. | Process design | A Notification Letter/Product Accountability Form dated 01/28/2013 was sent via FedEx on 01/28/2013 to the sole US branch receiving the affected devices. The branch was instructed to also contact the hospitals in its territory that have the affected product to arrange return of the product. Affected products to be returned to Stryker Orthopaedics, 325 Corporate Drive, Mahway, NJ 07430. |  | 8 | 4 units (total of 2 sets) | 1 Branch (NJ) with one consignee (DE). | Terminated | May 06, 2015 | 832 | 8 | Not_Computer
Z-0962-2013 | 64072 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | DePuy CONFIDENCE SPINAL CEMENT SYSTEM 11C Plus KIT |  |  | January 15, 2013 | 2013 | DePuy Spine, Inc. | Cement mixer used to prepare cement for surgery may not turn properly and may jam or stuck | Process control | DePuy Synthes Spine issued notification letter dated 1/11/13 issued via UPS Next Day Air Saver on 17 January 2013. Accounts requested to review inventory, quarantine any products listed in this recall notice and to complete and return the enclosed Business Reply Form. Contact the firm at 508-977-6606 for questions relating to this notice. |  | 7773 | 7773 | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Armenia, Kazakhstan, Russia, Ukraine, Australia, Austria Belgium, Luxembourg , Brazil, Canada, China, Colombia, Cyprus, Jordan Lebanon, Saudi Arabia, UAE, Czech Republic, Finland ,France, Germany, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Latvia, Malaysia Mexico, Netherlands, New Zealand, Norway, Sweden, Poland, Portugal Singapore, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey, and UK. | Terminated | June 30, 2014 | 531 | 16013 | Not_Computer
Z-0966-2013 | 64410 | lens, condensing, diagnostic | HJL | Diagnostic condensing lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | 20 Diopter Binocular Indirect Ophthalmoscope (BIO) Lens |  |  | October 25, 2012 | 2012 | Volk Optical Inc | The defect was with mis-engraving of the lot # directly on the product and its packaging. Mixed up device engraving of V20LC (lot # BB09213) with that from V60C (lot # BB09205). | Employee error | Volk sent a Voluntary Medical Device Field Correction letter dated October 25, 2012, to all affected customers. The letter identifed the product, the problem, and the action to be taken by the customer. Customers were asked to correct their records with the correct lot number and to notify their customers of the voluntary field correction and request their compliance in updating their records. For questions customers were instructed to call 440-510-0800 or 440-942-6161.For questions regarding this recall call 440-942-6161. |  | 40 | 40 devices | Worldwide Distribution - USA including the states of AZ, CA, GA, MN, NJ, and NM. Internationally to Japan. | Terminated | June 11, 2013 | 229 | 40 | Not_Computer
Z-0970-2013 | 64529 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CR MD1.0 General Set, x-ray imager image plate. CR 10-X Image Plate 35CM X 43CM. An image plate, contained in an X-Ray cassette, used to capture images from X-Ray exposures. The cassette with the image plate, once exposed, is inserted in the Agfa Digitizer CR 10-X where the image plate is taken out of the cassette and then scanned so that the scanned image can be displayed for viewing. |  |  | February 07, 2013 | 2013 | AGFA Corp. | The cassettes were shipped with the wrong IP (image plate) size bar code. | Nonconforming Material/Component | On February 7, 2013, AGFA called dealers involved to alert them of the safety issue. Also, an 'URGENT FIELD SAFETY NOTICE" letter dated February 26, 2013 was sent to the customers. The letter described the affected product, problem and actions to be taken. Customers were requested to check their inventories for the affected product and block any identified product within their inventory. Customers were instructed to complete the attached Feedback form and return via FAX or email. For questions call 877-777-2432 Prompt 5. |  | 17 | 17 | USA Nationwide Distribution including the states of: AL, CA, FL, IN, MN, NE, NV, OH, SC and UT. | Terminated | June 06, 2013 | 119 | 17 | Not_Computer
Z-0971-2013 | 63973 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 IM Rod, Non-Sterile and associated Surgical Techniques: |  |  | February 05, 2013 | 2013 | DePuy Orthopaedics, Inc. | Potential for the IM rod to break, leaving fragments in the patient. Surgical Techniques guidance has been updated until design changes can be made to reduce likelihood of potential failure. | Device Design | DePuy Orthopaedics sent an Urgent Medical Device Correction Notice letter dated February 5, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to complete and return the FAX back form to FAX 574-371-4939 or email kseppa@its.jnj.com Questions about device correction information call Post Market Surveillance, 574-372-7333, Sales force questions call Product Specialist, 574-372-7303 and Surgeon questions DePuy's Scientific Information Office at 1-888-554-2482. |  | 2600 | 4000; 2600 in USA | Worldwide Distribution - USA Nationwide. | Terminated | January 06, 2014 | 335 | 2600 | Not_Computer
Z-0955-2013 | 64511 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Vascular Solutions, Inc., Guardian II and Guardian II NC Hemostasis Valve, used in cardiopulmonary bypass. |  |  | February 27, 2013 | 2013 | Vascular Solutions, Inc. | Risk that air may be introduced into the device which may lead to an air embolism. | Vendor change control | Consignees were sent a Vascular Solutions "Urgent Medical Device Recall" letter dated February 28, 2013. The letter described the problem and the product involved in the recall. In addition, consignees were advised to identify the location of and immediately remove the recalled product from their inventory and secure it. Consignees were requested to complete and return the VSI Account Inventory Form. The letter also stated that VSI Customer Service would contact customers within 1-2 business days after receiving the completed VSI Account Inventory Form. All returned affected product will be replaced upon receipt. Customers are to contact their local Sales Representative if they have any questions or concerns regarding this recall. |  | 79869 | 79,869 | Worldwide Distribution -- USA, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT , VT, VA, WA, WV, and WI, and countries of GREECE, DENMARK, AUSTRIA, UNITED KINGDOM, SWITZERLAND, IRELAND, INDIA, TURKEY, POLAND, HUNGARY, ITALY, EGYPT, KOREA, PORTUGAL HONG KONG, GERMANY, SPAIN, MEXICO, SWEDEN, NORWAY SLOVENIA, NETHERLANDS, and BELGIUM. | Terminated | August 05, 2013 | 159 | 79869 | Not_Computer
Z-0974-2013 | 64435 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | INTIO Inc., ClearStart-SVM Segmentation and Volumetric Measurement System, Users Manual. Part #: UM-06-20-0003. ClearStart-SVM is used for patient assessment using CT (computed tomography) images, in order to plan potential therapies or monitor tumor response to treatment. |  |  | February 26, 2013 | 2013 | Intio Inc | INTIO Inc. distributed a SVM User's Manual and it has an editing mistake that could be interpreted as promoting an unapproved use. | Labeling False and Misleading | INTIO sent a Recall Notice to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. INTIO has corrected this version of the SVM User's Manual. An INTIO Inc. employee will contact the customers to replace the current manual with the correct one. Consignees were notified by letter that an INTIO Inc. employee will contact you within the next 30 days to replace the current manual with the correct one. The new SVM User's Manual will be provided to you once INTIO Inc. has retrieved the incorrect version of the SVM User's Manual. Customers with questions were instructed to email support@intio.us or call 303-396-1777.For questions regarding this recall call 303-396-9291. |  | 7 | 7 | Nationwide Distribution including MN, FL, PA, WI, NY, and RI. | Terminated | April 17, 2013 | 50 | 7 | Not_Computer
Z-0972-2013 | 64285 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ReClaim Reamer Extension; Product code 2975-00-500, used with Reclaim Modular Revision Hip Prosthesis System. The ReClaim Reamer Extension is as an attachment to the distal reamers to allow the surgeon to clear soft tissue while reaming the distal femoral canal. |  |  | February 25, 2013 | 2013 | DePuy Orthopaedics, Inc. | The tabs on the reamer extension device have broken. DePuy is notifying surgeons of the issue to increase the awareness and to ask surgeons and users to look for tabs before and after surgery to reduce any potential that the tabs are left in the patient. Once the root cause is determined and a design change is implemented, DePuy will conduct a formal swap-out of the affected devices. | Device Design | DePuy Synthes sent a Urgent Information Medical Device Correction Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed upon visual inspection, if the reamer extension is found to have one or more tabs broken, please do not use the reamer extension. The reamer extension will still engage with one tab intact and no clinical implications have been identified with the use of the reamer extension with one tab, but DePuy recommends it be returned and the alternate reamer extension device be used.If both tabs are broken, the reamer extension will not engage and DePuy recommends it be returned and the alternate reamer extension device be used. A reamer extractor is available in the instrument kit for removal of the reamer if necessary.In order to reduce the risk of breaking the tabs during surgery, the Reamer Extension should only be used to drive the Distal Reamer axially within the patients femoral canal. The instrument is NOT intended to be used for side-cutting or lateralizing the femoral canal, therefore, no off-axis force should be applied during distal reaming.The purpose of this communication is to inform you of this device correction and request acknowledgement of receipt of this letter by signing and returning the Surgeon Response Card within 14 days of receiving this notice. Response cards should be faxed to Fax Number: 574-371-4939.For product-related questions, please contact your local DePuy Orthopaedics Sales Representative. For clinical-related questions from surgeons, please have them contact DePuys Scientific Information Office at 1-888-554-2482 (M-F; 8 am to 5 pm EST.) For questions about device correction information provided, 574-372-7333 (M-F: 8 am to 5 pm EST.) |  | 750 | 750 | Worldwide Distribution - USA including the states of ME, MA, PA, MD, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, TN, NY, AR, NC and VA and the countries of Canada, Chile, Germany, Czech Republic, Israel , Norway, Poland, Slovenia, Australia, New Zealand, South Africa, Belgium, Austria, Switzerland, France, England and County Cork, Ireland. | Terminated | December 30, 2013 | 308 | 750 | Not_Computer
Z-0973-2013 | 64633 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | PrimaConnex Tapered Implant TC, SD Internal Connection 3.5 x 10mm |  |  | March 11, 2013 | 2013 | Keystone Dental Inc | Mislabeled: Cover screw that was in the package was a WD cover screw and not a SD cover screw as labeled | Packaging process control | Keystone Dental initially called customers on 3/11/13 and forms were completed to document these calls. In addition, a letter sent via Federal Express/or hand delivered to all distributors and customers. Product requested to be returned. |  | 120 | 120 units | Worldwide Distribution-USA including DC and the states of CA, FL, IN, MA, MN, NJ, NY, NH, PA, WI, and NY, and the countries of Republic of China, Germany, and France. | Terminated | September 26, 2013 | 199 | 120 | Not_Computer
Z-0968-2013 | 64535 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs BleaseSirius Anesthesia Workstation, BleaseFocus Anesthesia Workstation, and CAS I/II Absorber contained in Service Kits Part Number 12200902, 050-0659-00 and 050-0901-00 and any workstation installed with the CAS I/II Absorber. |  |  | January 15, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | A defect in CAS I/II Absorbers in the Spacelabs Anesthesia Workstations and Service Kits may cause a condition leading to an increase in the carbon dioxide concentration within the inspired gas being delivered to the patient. | Process control | Spacelabs sent an URGENT-MEDICAL DEVICE CORRECTION letter, dated 16 January 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers until a Spacelabs representative replaces the affected product: 1) Immediately cease the use of BleaseSirius and BleaseFocus Anesthesia Workstations that shipped after 31 October 2012. 2) Immediately cease use of Service Kits Part Number 12200902, 050-0659-00 and 050-0901-00 and any workstation containing a CAS I/II Absorber from such Service Kit. 3) Advise your staff of this situation Customers can call Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical support. |  | 108 | 108 total (17 units in the US and 91 units outside the US) | Worldwide Distribution - USA Nationwide including the states of New Mexico, North Carolina and the countries of Argentina, China, Mexico, and Taiwan. | Terminated | May 14, 2013 | 119 | 108 | Not_Computer
Z-0975-2013 | 64485 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Intraosseous Fixation System (IFS) 1.7 mm, 2.0 mm &amp; 2.5 mm Cannulated Drill Bit. |  |  | October 30, 2012 | 2012 | Orthohelix Surgical Designs Inc | The firm was notified by a customer that there was a burr on the inside end of the drill bit resulting in the drill bit not fitting over the associated guide wire. | Nonconforming Material/Component | OrthoHelix sent a "FIELD REMOVAL NOTIFICATION" letter dated November 7, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 330-247-1444 for questions concerning this notice. |  | 5040415514 | Lot # 2399041, 97 pcs; Lot # 2399051, 143 pcs; Lot # 2399051R, 4 pcs &amp; Lot # 2399061, 119 pcs | Nationwide Distribution including the states of CA, CO, IL, MI, MD, NC, NM, OH, OK, PA, SD, TX, UT, VA, WA and WI. | Terminated | December 05, 2013 | 401 | 5040415514 | Not_Computer
Z-0977-2013 | 64112 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | BS-200 Chemistry Analyzer |  |  | November 28, 2012 | 2012 | Mindray DS USA, Inc. d.b.a. Mindray North America | An issue has been identified with the BS-200 where clinical chemistry results may be inaccurate. | Device Design | Mindray North America /Mindray DS USA Inc. sent Urgent Product Corrective Action letters via certified mail, return receipt requested, signature required on 11/28/2012. The letters identified the affected product as well as the issue, effect on patients, and the corrective action. Customers are to remove the analyzer from service until the affected pump has been replaced. Also, customers are to contact Mindray Customer Service Representative Tony Li at 1-800-288-2121 x2112 for technical assistance, and to arrange for replacement pumps at no charge. |  | 3 | 3 units | Distributed in the states of NC and TX. | Terminated | June 12, 2013 | 196 | 3 | Not_Computer
Z-0976-2013 | 64310 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD MAX PCR Cartridges Catalog #437519, box of 24 cartridges labeled in part:***BD MAX PCR Cartridges, GeneOhm Sciences Canada, Inc. 2555 boul du Parc-Technologique, Quebec, Canada G1P 4S5***BD Max Group B Streptococcus (GBS) Assay is an in vitro diagnostic (IVD) test for detection of GBS DNA in Lim Broth cultures. |  |  | December 11, 2012 | 2012 | Becton Dickinson &amp; Co. | Cartridges are not sealing properly during PCR testing and may cause false results, indeterminates and/or invalid runs. | Process control | BD Diagnostic Systems issued an Urgent Product Recall letter via UPS delivery on December 11, 2012. The letter identified the product, the problems, and the action to be taken by the customer. Customers were instructed to take the following actions: Discontinue and discard any affected product. BD would send replacements for the discarded material. Complete the attached form whether or not they have any inventory remaining so that the firm may acknowledge their receipt of the notification. Fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258 or Email to RegulatoryComplianceFax@BD.com.. If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, contact BD Technical Services Department at 1-800-638-8663. The BD site in Belgium was notified via email.For questions regarding this recall call 410-316-4054. |  | 17496 | 27 boxes/648 cartridges | Worldwide Distribution - USA including CA, FL, IN, IL, KS, NY, OH, OR, PA, TX, and WI and the country of Belgium | Terminated | July 22, 2013 | 223 | 17496 | Not_Computer
Z-0988-2013 | 63494 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension Vista(R) Alkaline Phosphatase (ALP) Calibrator |  |  | October 10, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Correlation slope bias and accuracy shift | Nonconforming Material/Component | Siemens sent a Urgent Medical Device Correction letter dated October 16, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to calibrate with newer lots of Dimension Vista(R) ALP Calibrator as they become available, beginning with lot 2GD011. Customers were also asked to discontinue use and discard any remaining inventory of lot 2AD052. Customers further instructed to complete the attached response form and fax to (302) 631-8467, to indicate they had received the information and to facilitate their no charge replacement.Further questions please call (302) 631-6311. |  | 1985 | 1985 cartons | Nationwide distribution: USA including Puerto Rico and the states of AL, OH, MO, GA, IL, CT, IA, SC, VA, TX, FL, NY, NC, CA, NM, AZ, CA, MI, NJ, PR, MS, NE, WV, PA, MT, MA, MD, KS, NH, TN, VT, LA, CO, ID, WA, NH, DE and OR. | Terminated | January 08, 2014 | 455 | 1985 | Not_Computer
Z-0994-2013 | 64674 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix-i X-Ray System. A radiographic and fluoroscopic x-ray system for use in diagnostic and angiographic procedures. |  |  | November 08, 2012 | 2012 | Toshiba American Medical Systems Inc | Lost and incorrect images. When using the "F-Rec" fluoroscopy image acquisition recording option in the "manual mode" an infrequent software error may cause images from another study to be displayed on the monitor. The acquired images may be lost requiring additional image acquisitions to be performed. | Software design | In November, 2012, Toshiba America Medical Systems, Inc. sent an Urgent: Medical Device Correction letter to customers. The letter contains a return reply form that is to be faxed to TAMS for retention. |  | 95 | 95 units | USA nationwide. No international. | Terminated | March 26, 2013 | 138 | 95 | Computer
Z-0992-2013 | 64684 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Axiom Artis zeego systems with software version VC14, VC20 and VC21 in conjunction with specific technical configuration |  |  | November 26, 2012 | 2012 | Siemens Medical Solutions USA, Inc | In the course of the firm's product monitoring activities, a potential risk for patient or operators was determined during the operation of an Artis zeego system with software revision VC14, VC20 and VC21 in conjunction with a specific technical configuration which cannot be completely executed. | Software design | Siemens sent an Urgent Customer Information letter dated November 8, 2012 to affected consignees. The letter notified consignees of the issue and informed them the update AX039/12/S would be started and affected systems would be checked and corrected if required. Customers were instructed that they would be contacted to arrange a date for checking their system.For questions regarding this recall call 610-219-6300. |  | 52 | 52 | Nationwide Distribution including AL, AR, AZ, CA, CT, DC, FL, GA, IA, IL, LA, MD, MI, MO, NC, NH, OR, PA, SD, TN, TX, WA and WI and Puerto Rico. | Terminated | December 24, 2013 | 393 | 52 | Computer
Z-0991-2013 | 64394 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | VITEK 2 Compact computer system when equipped with the following Hewlett-Packard PC models: |  |  | February 18, 2013 | 2013 | Biomerieux Inc | There is a potential to link an isolate result to the wrong patient and then upload the results to the LIS. | Software design | bioMerieux sent an "URGENT PRODUCT CORRECTION NOTICE" dated February 20, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Emails were issued by the recalling firm on 2/18/2013 and 2/19/2013 to their subsidiaries and to distributors. Letters began issuing to customers via e-mail on 2/20/2013. Contact bioMerieux Clinical Customer Service at (800) 682-2666 for questions regarding this notice. |  | 17293 | For all 3 computers types: 16,736 originally distributed; 554 refurbished and redistributed | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaidjian, Bahrain, Belarus, Belgium, Benin, Bolivia, Bosnia, Brazil, Bulgaria, Burkina Faso, Camaroun, Canada, Chad, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Democratic Republic of Congo, Demark, Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gabon, Georgia, Germany, Greece, Guatemala, Guinea, Herzegovina, Honduras, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Mexico, Moldavia, Mongolia, Morocco, Myanmar, Netherlands, Netherlands, Nicaragua, Niger Niamey, Nigeria, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Yemen. | Terminated | August 14, 2014 | 542 | 17293 | Not_Computer
Z-0987-2013 | 63946 | No | LOM |  |  | Microbiology | PMA | ARCHITECT HBsAg Qualitative Assay, List Number 4P53 and ARCHITECT HBsAg Qualitative Confirmatory Assay, List Number 4P54 |  |  | December 18, 2012 | 2012 | Abbott Laboratories, Inc | Abbott has identified that assay files for the ARCHITECT HBsAg Qualitative and HBsAg Qualitative Confirmatory assays (LNs 4P53 and 4P54) on the ARCHITECT i1000SR System were mistakenly made available on abbottdiagnostics.com in October 2012. Additionally, some i1000SR customers received notification through AbbottLink in November 2012 that assay files for LNs 4P53 and 4P54 were available. The AR | No Marketing Application | The firm decided to correct the issue and and sent out Urgent Field Safety Notices to their customers dated 10/18/2012. |  | 12 | 12 | Nationwide Distribution including AZ, CA, FL, GA, HI, LA, MA, MD, MO, NJ, NY, PA, RI, TN, TX, UT, and WA. | Terminated | July 31, 2013 | 225 | 12 | Not_Computer
Z-0997-2013 | 64662 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Systems with Software Artis VC20x/VC21A/VD10x that have a DSA license installed. |  |  | March 08, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm has become aware of a potential issue on Artis systems with the software Artis VC20x/VC21A/VD10x that have a DSA license installed. Under certain preconditions, executing roadmap OGP (organ program) in the DSA overlay mode may lead to imprecise registration of the DSA mask image with the roadmap image displayed on the live monitor. | Software design | Siemens sent a Customer Safety Advisory Notice on March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to pass the information on to all those who need to be aware of it and to instruct their personnel accordingly. Customers were asked to forward the safety information to other organizations that couald be affected by this action.For questions regarding this recall call 610-219-6300. |  | 81 | 81 | Nationwide Distribution including AL, CA, CO, FL, IA, IL, IN, KY, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, PA, SC, TX, UT, VA, WA, and WI and to Puerto Rico. | Terminated | November 06, 2014 | 608 | 81 | Computer
Z-1001-2013 | 64585 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Sepp, Compound Benzoin Tincture Packaged and Tincture 3,000 Bulk. |  |  | March 04, 2013 | 2013 | CareFusion 213, LLC | An internal review of CareFusion's labeling for preoperative skin prep product codes was recently conducted. The labeling of the product codes identified above state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterilized. | Labeling False and Misleading | CareFusion sent an Urgent Recall Notice dated March 4, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, quarantine and return the affected product for replacement. For questions call (913) 345-3504. |  | 13749000 | 13,749,000 | USA Nationwide Distribution | Terminated | May 12, 2014 | 434 | 13749000 | Not_Computer
Z-0995-2013 | 64689 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | KIMGUARD Container Filters, 7.5in (19cm), Round and 9in. X 9in. (23cm x 23cm), Square. |  |  | March 06, 2013 | 2013 | Kimberly-Clark Corporation | Some individual container filter units might contain thin areas that may potentially compromise the ability of the filter to maintain a sterile barrier. | Component design/selection | All US distributors and clinics were sent an Urgent: Voluntary Product Recall letter via Federal Express, and the Canadian importer/distributor was also notified with the same letter via Federal Express and e-mail. Distributors were instructed to contact their customers by sending a copy of the letter. The letter identified the affected product and stated the reason for the recall. The recall notification includes a Response Sheet that allows the customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. The customer is instructed to immediately cease further distribution of the product, physically quarantine and destroy the product, and fax the completed Response Sheet to the Kimberly-Clark Recall Coordinator. |  | 144717 | Code 68507 - 6420 cases; Code 68509 - 1281 cases | Worldwide Distribution -- USA, including the states of AL, AK, AZ, AR, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NV, NJ, NM, NY, NC, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI, and the countries of Canada, Australia, Belgium, Hong Kong, Japan, Portugal, and Singapore. | Terminated | September 16, 2013 | 194 | 144717 | Not_Computer
Z-0998-2013 | 64609 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 System Carriage, P/N 10208. |  |  | March 13, 2012 | 2012 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 System Carriage because it can potentially become dislodged when an excessive load is applied. | Device Design | Advanced Sterilization Products (ASP) sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 13, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter informed the customers that ASP recently determined that the push bar located on the STERRAD 200 System Carriage can potentially become dislodged when an excessive load is applied. Customers with questions are instructed to contact Stericycle at (877) 931-0642. |  | 2457 | 2,457 units | Worldwide Distribution-USA (nationwide) and the countries of Egypt, Germany, Italy, France, Saudi Arabia, UAE, Japan, Singapore, China, Thailand, Hong Kong, Hungary, Israel, Slovania, Poland, Czech Republic, Taiwan, Sweden, Greece, Korea, Ireland, and United Kingdom. | Terminated | May 28, 2014 | 806 | 2457 | Not_Computer
Z-0996-2013 | 63621 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion Model 3500 Syringe Pump, Model #3500-415 and 3500-500. |  |  | May 14, 2010 | 2010 | Smiths Medical ASD, Inc. | An issue with the plunger rod seal was identified where the seal had the potential to migrate inside the pump case top during slow rate/ high back pressure infusions. | Device Design | On May 14th, 2010, Smiths Medical contacted the facility that received the pumps via e-mail and requested that they immediately return the pumps. The pumps were returned per Returned Goods Authorization (RGA) and exchanged for replacement pumps manufactured with plunger rod seal G6000077. |  | 6 | 6 | Distributed in the state of NH. | Terminated | April 11, 2013 | 1063 | 6 | Not_Computer
Z-1013-2013 | 64327 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux VITEK 2 AST-ST01 Streptococcus Susceptibility Card, REF 410 028. |  |  | February 06, 2013 | 2013 | Biomerieux Inc | The product package did not contain the correct version of the package insert listing the limitations of use and performance requirements. | Error in labeling | bioMerieux sent an "PRODUCT SAFETY CORRECTION NOTICE" dated February 13, 2013 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at (800) 682-2666 for questions regarding this notice. |  | 4 | 4/20-card cartons | Distributed only in HI. | Terminated | October 30, 2013 | 266 | 4 | Computer
Z-1002-2013 | 64461 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3 |  |  | February 05, 2013 | 2013 | Vital Images, Inc. | Vital Images has found a potential error in measurement generated through the software which affects all versions of Vitrea Enterprise Suite, Vitrea, VitreaAdvanced, VitreaCore (VitalConnect) and Vitrea fX prior to version 6.3.3, except 6.1.6, 6.2.3 and 6.3.This error occurs when images are rotated at the time of reconstruction by the scanner in non-90 degree increments of the transverse/axia | Software design | Vital Images sent an "Urgent Field Safety Notice' letter dated February 5, 2013, to all affected customers. The letter identified the probelm the product and the action needed to be taken by the customer. Please immediately inform all potential users of Vitrea of the information contained in this letter. We recommend that all measurements associated with the error described in this letter be verified against results from other technologies.You will be contacted by Vital Images Customer Support or by the customer support department of your distributor to arrange for installation of this patch on your system. We recommend that you implement the patch. If you have an earlier version, Vital Images Customer Support or your distributor will contact you to discuss your options for replacement.Please sign this Notice in the space below to verify that you have received and understood it, and return your signed Notice by email to fieldnotices@vitalimages.com.If you have questions, direct customers of Vital Images should contact Vital Images Customer Support at 1-800-208-3005 or by e-mail to support@vitalimages.com; customers through a distributor of Vital Images products should contact that distributor.Updated 5/28/2013:A second consignee Vital Images, Inc "Urgent Field safety Notice" letter dated March 26, 2013 was sent to consignees on 3/25/13. The follow-up letter was sent via e-mail and mail to 4,616 already-notified consginees, who had received affected versions 6.3.0, 6.2.3, or 6.1.6. The letter described the problem and the affected product being recalled and provided a number of steps to follow which included to retun the Effectiveness Check form. |  | 38245 | 7649 (Expnaded 5/28/13 to 7654) | Worldwide Distribution - All states in continental USA including DC, PR and the countries of ALGERIA, ARGENTINA, ANGOLA, ARMENIA, AUSTRALIA, AUSTRIA, BENGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL BULAGARIA, CAMBODIA, CANADA, CENTRAL AFRICAN REPUBLIC, CHAD, CHILE, CHINA , COLOMBIA, CONGO, CROATIA, CYRPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GREAT BRITAIN, GEORGIA, GERMANY, GHANA, GREECE, GUAM, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAPAN, JORDAN, KAZAKHASTAN, KOREA, LEBANON, LITHUANIA, MACEDONIA, MALYASIA, MALI, MAURITIUS, MEXICO, MALDOVA, MOSCOW, NYNMAR, NAMBIA, NAMIRGIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGER, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILLIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, RUSSIAN FEDERATION, SAUDI ARABIA, SENEGAL, SERBIA, SEBIA AND MONROBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOMALIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, THE NETHERLANDS, TIBET, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBERKISTAN, VENEZUELA, VIETNAM and YEMEN. | Terminated | May 08, 2015 | 822 | 38245 | Computer
Z-1004-2013 | 64531 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical Incontinence Ring Flexible Pessary #4, 2 3/4", 70mm. Milex |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | Pessary mislabeled outer package may not reflect the size of the product contained in the package | Packaging process control | Cooper Surgical initiateda recall by letter on February 26, 2013, advising accounts to examine inventory and report quantity back to the firm for replacement. For questions customers should call Customer Service at 1-800-243-2974.For questions regarding this recall call 203-601-9825. |  | 80 | 80 units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Canada, Colombia, Demarkt, France, Holland, New ZeaJand, andThailand. | Terminated | March 04, 2014 | 371 | 750 | Not_Computer
Z-1015-2013 | 53991 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System Model 2000 (PAS 2000) |  |  | November 20, 2009 | 2009 | CareFusion Corporation | The recall was initiated because Carefusion has confirmed that when the PAS 2000 drawers due to inaccessability of drawers caused by system down time or lock-up. | Device Design | Carefusion sent a Safety Alert Notification letter dated January 13, 2010 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to ensure back-up procedures were in place in case of system down-time or lock-up, including review of the back-up procedures with staff. The letter included an update to the PAS 2000 User Guide. The letterstates that a CareFusion field representative will be contacting customers to make arrangements to apply warning labels to the PAS 2000 electronics doors. Customers were instructed to complete and return the Response Card included. For questions contact your local field support representative ot the Advocacy department at customer.advocacy@carefusion.com. |  | 2637 | 2637 | Worldwide Distribution - USA (Nationwide) including CANADA. | Terminated | March 28, 2013 | 1224 | 2637 | Not_Computer
Z-1016-2013 | 64271 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Dens Instrument Set (Dens Graphic Case), Part Number 687.030 (also known as the Anterior Column Retractor Set) |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | The Dens Instrument Set has been designed to facilitate the anterior internal fixation of the Dens (odontoid process). There are currently no screws specifically indicated to treat Dens fractures via anterior screw fixation therefore the affected product will be removed. | No Marketing Application | Synthes sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory, remove the recalled product from use and return it to Synthes. Customers were asked to completed the attached verification Section at the end of the letter by checking the appropriate box indicating that no affected product had been located and fax to 610-251-9005.For questions customers were instructed to call 610-719-5450 or email at FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 148 | 148 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, KS, KY, LA, MD, ME, MI, MN, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, TN, TX, VA, VT, WI, WV, and WY. | Terminated | September 10, 2015 | 972 | 148 | Not_Computer
Z-1022-2013 | 64501 | No | LCO | Platelet factor 4 radioimmunoassay. | Hematology | Hematology | 510(k) | Gen-Probe LIFECODES PF4 Enhanced assay, Catalog No. X-HAT45 |  |  | February 04, 2013 | 2013 | Genetic Testing Institute,inc | Lifecodes PF4 Enhanced assay, lot number 3000389, is being recalled because of observed inconsistent replicate values when testing the kit Positive Control and patient samples. The variation in reactivity can lead to invalid assays. If an invalid assay is obtained, those results should not be reported and the test should be re-run. An invalid assay would not result in significant harm to a pati | Process control | Hologic/Gen-Probe sent an Urgent Medical Device Recall letter dated February 4, 2013, via FedEx overnight deliver yto all affected customers. The letter identified affected product, stated the issue, and asked that product from the affected lot stop being used. Reported results for all samples tested using the affected lot should be reviewed. The referring physician should be notified that any negative patient result may be erroneous. Remaining inventory should be returned to Gen-Probe. A response form was also asked to be returned. For questions customers were instructed to contact Technical Support at 262-754-1000 or waukeshatechsupport@gen-probe.com.For questions regarding this recall call 262-754-1026. |  | 386 | 386 kits (370 US, 16 OUS) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, China, Slovakia, and Spain. | Terminated | March 21, 2016 | 1141 | 386 | Not_Computer
Z-1018-2013 | 64510 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Omniflex Diaphragm - Size 65 (Milex Wide Seal Silicone Vaginal Diaphragm) Model MXWF65 |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | Incorrect size was printed on the outer product packaging. | Packaging | CSI notified consignees affected by this recall by a Recall Notification letter, dated February 26, 2013, sent via Federal Express with confirmed delivery receipt. The letter identified the affected product, and asked customers to contact the firm's Customer Service Department at 1-800-243-2974, if they had any of the affected lot in their inventory. Customers also have the option of completing the form attached to the letter and faxing it back to 800-262-0105. At CooperSurgical's expense, arrangements will made to replace the affected products that customers have on hand. The firm will provide further instruction and disposition of the product at the time of replacement. |  | 80 | 80 individual diaphragms | Worldwide Distribution -- USA, including the states of NE, MA, FL, TN, CA, ME, SC, CA, NH, OK, TX, NM, MO, NY, PA, and AZ, and countries of Germany, Italy, and Canada. | Terminated | March 04, 2014 | 371 | 80 | Not_Computer
Z-0983-2013 | 63951 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Needle Counters; 12.2 x 10.5 x 3.5 cm; 40 Count/70 Capacity; Sterile; Removable Numbered Foam Block; Adhesive Area (Cat. No. 3FA70SSA) |  |  | December 18, 2012 | 2012 | Cardinal Health | Small holes (approximately 1 mm or less) found in device unit level packaging potentially compromising product sterility. | Package design/selection | Cardinal Health sent an Urgent Product Recall letter dated December 18, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immedaitely stop using affected product and discard any product that may be on hand, return the enclosed acknowledgement form, notify any customers to whom you may have distributed product affected by this recall, and contact the appropriate Customer Service group to arrange for credit of any affected product discarded. For questions contact Cardinal Health Professional Services at 800-292-9332. |  | 54912 | Cat. No. 3FA70SSA: 23,808; Cat. No. 3FM70SSA: 54,912 | Worldwide Distribution - USA (Nationwide) and the countried of Australia, Canada and Japan. | Open, Classified |  |  | 316752 | Not_Computer
Z-1019-2013 | 64623 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewire, Models ARCSJ200W, ARCSJ260W, and ARCDJ260W. Sterile using ethylene oxide. |  |  | March 13, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified an issue involving specific lot numbers of the Medtronic Archer 0.035 inch (0.89mm) Super Stiff Guidewire, Models ARCSJ200W, ARCSJ260W, and ARCDJ260W, where a change in the manufacturing process has been identified as the cause for potential kinking and breaking at the proximal weld end of the outer spring coil during in-vivo use. | Component change control | Medtronic sent an Urgent Medical Device Recall letter dated March 2013 to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers were instructed to immediately quarantine all affected products and returned to Medtronic. Customers were requested to complete the attached Customer Confirmation Certificate and fax it to Medtronic to the attention of Customer Focused Quality. For question contact your local Medtronic representative or Medtronic Customer Service at 1-888-283-7868. |  | 27682 | 27,682 devices | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Argentina, Armenia, Australia, Austria, Belgium, Bolivia, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Israel, Italy, Jordan, Korea, Kuwait, Latvia, Lebanon, Libya, Malaysia, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Viet Nam. | Terminated | February 06, 2014 | 330 | 27682 | Not_Computer
Z-1017-2013 | 64620 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Archer Super Stiff Guidewire 0.035 inch (0.89mm), Model ARCSJ200W. |  |  | March 13, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified an issue with a specific lot of the Archer 0.035 inch (0.89mm) Super Stiff Guidewire, model ARCSJ200W, where the PTFE (Teflon) coating has the potential to separate from the wire during in-vivo use. | Mixed-up of materials/components | Medtronic sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Medtronic requested any unused guidewires from the lot number to be quarantined, and returned by contacting the customer service phone number at 1-888-283-7868. A Customer Confirmation Certificate was requested to be faxed back to Medtronic at 651-367-0612. |  | 225 | 225 units | Worldwide Distribution-USA including the states of FL, OH, and MN and the countries of Belgium, France Germany, Netherlands, Turkey, and United Kingdom. | Terminated | November 21, 2013 | 253 | 225 | Not_Computer
Z-1023-2013 | 64579 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | 20-Pole Eco Cable, Catalog #EM-5050-060 |  |  | March 05, 2013 | 2013 | Biosense Webster, Inc. | Biosense is recalling the 20-Pole Eco Cable because it was incorrectly calibrated. When a catheter is connected to an affected to an affected 20-Pole Eco Cable, the catheter will not be recognized by the impacted 54 cable serial numbers within the field. | Employee error | Biosense Webster sent an Urgent Field Notice Medical Device Letter dated March 5, 2013 to all affected customers. The letter identified the affected product, problem and the actions to be taken. Customers with questions are instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Mon-Friday from 7amto 8pm EST. |  | 54 | 54 units | Worldwide Distribution - USA (Nationwide) including the states of: DE, NL, CH, FR, GB, DK. | Terminated | October 17, 2013 | 226 | 54 | Not_Computer
Z-1021-2013 | 64469 | No | LJG | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Microbiologics KWIK-QC SLIDES ( KWIK-QC Acid Fast Stain Slides and KWIK-QC Mycobacterium Slides). |  |  | February 11, 2013 | 2013 | Microbiologics Inc | Microbiologics Inc is initiating a voluntary recall on products KWIK-QC Acid Fast Stain Slide, SL42-10 and KWIK-QC Mycobacterium, SL43-10. These products are being recalled because the control organisms are not properly fixated on the slide, which may cause the user to unintentionally contaminate the patient area of the slide by flooding the control well with stain. This contamination could pote | Device Design | Microbiologists an "Urgent Medical Device Recall" letter dated February 11, 2013. The Letter described the product and the reason for the recall. Advised consignees to "Do not Distribute" and "Return" the product. Urged consignees to "Notify" end users and to complete the attached "Recall Response Form". Contact the form at 320-229-7057 for questions regarding this notice. |  | 3960 | 396 boxes of 10 slides each. | Worldwide Distribution-USA including the states of CA, IN, MS, NY, OH, PA, and SD, and the countries of AUSTRALIA, AUSTRIA, UNITED ARAB EMIRATES, MALAYSIA, TAIWAN, JAPAN, TURKEY, BELGIUM, COLOMBIA, CANADA, SOUTH AFRICA, and VENEZUELA. | Terminated | March 10, 2016 | 1123 | 3960 | Not_Computer
Z-1025-2013 | 64075 | No | FZZ | Skin marker. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Dual Tip Marker, Sterile |  |  | August 22, 2012 | 2012 | Viscot Medical LLC | Dual Tip Marker has inconsistent ink flow on the black permanent ink side. | Device Design | Viscot Medical LLC sent a Customer Notification Letter dated August 22, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to return the affected product for a replacement. To return the product call 1-800-221-0658 or e-mail jeffl@viscot .com for the RGA. |  | 458589 | 458,589 markers | USA Nationwide Distribution | Terminated | July 21, 2014 | 698 | 458589 | Not_Computer
Z-1024-2013 | 64448 | No | OPP | Infusion pump. | General Hospital | General Hospital | 510(k) | Z-800 Volumetric Infusion System. |  |  | June 13, 2011 | 2011 | Zyno Medical LLC | Update Instruction For Use for Preventative Maintenance (PM) Schedule and addition of Odometer software to monitor accumulative fluid volume infused by the pump | Process design | Zyno Medical initiated recall 6/13/11 to accounts advising of a Preventative Maintenance (PM) Schedule and a software upgrade Odometer to monitor counts of accumulative fluid volume infused by the unit. On 3/1/13, firm issued correction notice to users that have not had the software implemented on their unit. Units will be serviced at their scheduled PM. The notice identified the product, problem, and actions to be taken by the customers. |  | 2589 | 2589 units | Nationwide Distribution. | Completed |  |  | 2589 | Computer
Z-1014-2013 | 64543 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical, BCIÂż Remote Alarm Cable (Normally Open &amp; Normally Closed), Reorder Numbers: WW3398NO &amp; WW3398NC. For use with respiration carbon dioxide measurement equipment. Product Usage: The BCIÂż Remote Alarm Cable is designed for use with the 9004 CapnocheckÂż Capnography (Plus or Sleep) system to transfer alarm signals from the CapnocheckÂż Capnography system to a remote alarm system. The BCIÂż Remote Alarm Cable and the 9004 CapnocheckÂż Capnography systems are intended for use in all critical environments including ventilatory applications, patient transport, and anesthesia. |  |  | February 06, 2013 | 2013 | Smiths Medical ASD, Inc. | In rare instances, the BCIÂż Remote Alarm Cables (BCIÂż Cable) are not transferring alarms when used with some remote nurse alarm systems. | Nonconforming Material/Component | Smiths Medical sent an "Urgent Medical Device Field Safety Notice-Recall" letter dated February 96, 2013 via UPS to all affected customers. The letter described the affected product, problem and actions to be taken. Customers were instructed to inspect their inventory and remove all affected devices from use, complete and return the attached "Confirmation Form", by fax to 1-800-237-8033 or email to BCIRemote.AlarmCable@smiths-medical.com. The letter stated that a Customer representative will contact customers with instructions for returning affected products for replacement or credit. For questions contact Smiths Medical's Customer Service Department at 10-800-258-5361. |  | 39 | 39 | USA Nationwide Distribution - including the states of: CO, IL, IN, MA, SC and WI. | Terminated | July 12, 2013 | 156 | 39 | Not_Computer
Z-1026-2013 | 64692 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Sterilization System, P/N 10104 |  |  | March 22, 2013 | 2013 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products (ASP) has determined that the STERRAD 100NX and STERRAD 200 Sterilization Systems may emit an odor or smell into the surrounding environment that is not typical to normal operating conditions. | Nonconforming Material/Component | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notification letter dated March 22, 2013 to all customers. The letter identified the affected products, problem and actions to be taken. Customers are instructed to report any complaints or suspected problems with their STERRAD 100NX and/or STERRAD 200 System to ASP Professional Services directly at (888) 783-7723 and select option #2. |  | 1500 | 1500 units | US Nationwide Distribution | Terminated | September 30, 2013 | 192 | 2538 | Not_Computer
Z-1028-2013 | 64453 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Molecular Systems cobas x 480 System 50 mL reagent reservoir, Catalog number 05232732001. Used with the cobasÂż CT/NG 4800 System, for in vitro diagnostics. |  |  | March 01, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Certain lots of Reagent reservoirs, may not have a separation in the lower part of the reservoir. As a consequence, pipetting errors can occur due to uneven distribution of reagent. These pipetting errors may lead to delayed results and wasted reagent. | Process control | Roche sent a Urgent Medical Device Correction letter via UPS to all affected customers ( receipt required) The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to visually inspect affected product for the molding defect and to verify equal distribution of reagent in both chambers of the reservoir after adding reagent. All customers are requested to fill out the FAX Back form to schedule shipment of defective part. Non-responding accounts will be monitored on an ongoing basis and follow-up attempts will be made. Questions and concerns are to be addressed to Roche Molecular Diagnostics Technical Support 24 hours/day 1-800-526-1247. |  | 189 | 189 units | Nationwide Distribution including Puerto Rico and the states of MO, NE, CO, MT, NJ, FL, WA, TX, GA, NY, KS, NH, VA, HI, PA, MI, IA, CA, NC, AZ, UT, LA, IL, TN, ME, MN, KY, CT, OR, NV, AL and NM. | Terminated | June 04, 2014 | 460 | 189 | Not_Computer
Z-1029-2013 | 64393 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT. |  |  | December 20, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Issues have been detected in the Ingenuity TF PET/CT currently running software version 4.0.60605.0 that if they were to re-occur, could cause a risk for patients, users or service personnel, including inaccurate standard uptake values (SUV). Philips Medical Systems (Cleveland), Inc. is issuing a software update to address the issues. | Software design | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated December 19, 2012 to the affected customer. The letter identified the product, problem, and actions to be taken by the customer. A Philips Service Engineer will contact the customer for implementation of the software update on the affected system. |  | 1 | 1 | Distributed to one customer in Ohio. | Terminated | June 18, 2013 | 180 | 1 | Computer
Z-1045-2013 | 64569 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Zimmer NexPosure, retractor portals of the NexPosure MIS Access System, Part Numbers or REF numbers: N1830R80-100, N1830R80-090Âż, N1830R80-080Âż, N1830R80-070Âż, N1830R80-060, N1830R50-100, N1830R50-090, N1830R50-080, Âż N1830R50-070, N1830R50-060, N1830L80-100, N1830L80-090, N1830L80-080, N1830L80-070, N1830L80-060, N1830L50-100, N1830L50-090, N1830L50-080, N1830L50-070, N1830L50-060, Âż N182450-085, N182450-075, N182450-065, N182450-055, N182450-045.Âż Retractor access system that allows placement of posterior spinal fixation implants. |  |  | February 08, 2013 | 2013 | Zimmer, Inc. | Possibility that the inner pouch may be compromised. The product is packaged in two pouches - a sterile inner pouch within an external pouch. NexPosure Retractors contained in a breached inner pouch have an increased risk of contamination (i.e.; loss of product sterility) if proper aseptic technique is not followed when transferring the inner package into the sterile field. Patient infection may o | Package design/selection | Consignees were sent on 2/8/13 a Zimmer Spine "Urgent Medical Device Recall" letter dated February 08, 2013. The letter was sent to Surgeons and Facilities using the NexPosure MIS Access System. The letter described the product involved in the recall and the problem. Advised consignees to discontinue the use of the product and to contact their sales representative for removal.A consignee letter dated February 8, 2013 was also sent to Distributors, Sales Representatives, and Distribution Operation Managers Distributing the NexPosureÂż MIS Access System. The letter contained instructions if they had the affected product or not and the steps to follow. |  | 58 | 58 | Nationwide Distribution including the states of AZ, CO, DE, FL, ID, LMA, MI, NH, NJ, NM, PA, RI, TX, and VA. | Terminated | December 09, 2014 | 669 | 58 | Not_Computer
Z-1048-2013 | 64570 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension(R) EXL(TM) 200 (Siemens Material Number 10636929). |  |  | February 27, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Through investigation of customer complaints, Siemens Healthcare Diagnostics has confirmed an increase in System Check failures due to high System Check LOCI results. Upon further investigation, Siemens has determined that many of these errors are due to the current limit set for System Check LOCI results. The errors have no impact on results, but they may cause customer inconvenience. Siemens h | Under Investigation by firm | Siemens initiated their recall of this product on February 27, 2013 by sending an Important Customer Notification letter to consignees by FedEx. The letter identified the product, problem, and actions to be taken by the customers. |  | 2207 | 2207 total | Worldwide Distribution. | Terminated | January 14, 2014 | 321 | 2207 | Not_Computer
Z-1033-2013 | 64332 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System; P/N 107405S 2/12 |  |  | February 12, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall due to a defective newly released aluminum incubator belt pulley that may rub against the light shield, which causes the formation of debris. Debris that is caused by the incubator belt pulley rubbing against the light shield can fall into the reaction vessels (RVs) and present a risk of erroneous test results. | Device Design | Beckman Coulter sent an "Urgent Product Correction" letter dated February 12, 2013 via mail and email to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to place your Access system in the Ready or Not Ready mode. Lift the cover (as you do to access the probes) and look at the left side of the incubator unit under the peripump by the aspirate probe, if you see a silver pulley, immediately discontinue use and contact Technical Support also check for QC failures and system check failures; If you see a white pulley, this is the correct part and you can resume normal activity. Beckman will replace the metal pulley with the plastic pulley if you verified the your instrument has the silver pulley. The customers were also instructed to share this information with staff; and complete and return enclosed Response Form within 10 days.If you have any questions regarding this notice, please contact Technical support from the website: http://www.beckmancoulter.com or by phone at: 1-800-854-3633 in the United States and Canada. |  | 3500 | 3500 units | Worldwide distribution: USA (nationwide) and countries of: Aurstria, Belgium, Bulgaria, Canada, Croatia, Czech Republic, Denmark, Djibouti, Finland, France, Germany, Greece, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | December 23, 2013 | 314 | 10236 | Not_Computer
Z-1030-2013 | 64665 | No | GYC | Cortical electrode. | Neurology | Neurology | 510(k) | Integra Ojemann Cortical Stimulator; Product / Catalogue No: OCS2. |  |  | March 18, 2013 | 2013 | Integra Burlington MA, Inc. | If the headphone jack is in use during a surgical procedure with the OCS2 and a non-intended voltage (such as static electricity) comes in contact with the outer case of the unit, that voltage could be transmitted to the patient and could be a potential source of injury. | Device Design | Integra initiated a voluntary recall on March 18, 2013 by providing written notification to consignees either by traceable courier service or traceable emails regarding the correction that Integra records indicate have an OCS2 since it was introduced to the market in 2007. The consignee notification:Advised of the nature of the issue and to immediately stop using the headphone jack while using the OCS2. Requested identification by S/N any OCS2's they have, and complete / return the Acknowledgment Return Form.Advised that Integra will contact the consignee to schedule a time / date to correct the OCS2's they have identified. Questions, contact service hotline at 1-888-772-7378. |  | 371 | 371 units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Czech Republic, Belgium, Denmark, Spain, Great Britain, Hungary, Ireland, Italy, Netherlands, Poland, Morocco, Netherlands, Poland, Portugal, Saudia Arabia, Sweden, Turkey, and South Africa. | Terminated | February 25, 2014 | 344 | 371 | Not_Computer
Z-1038-2013 | 64675 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand COHERENCE RT Therapist versions RTT MR1.5, RTT 2.1/2.1A and RTT 2.2/ 2.2B; Software release for COHERENCE Therapist 2.3 for Primus and Oncor Linear Accelerators; |  |  | March 05, 0013 | 13 | Siemens Medical Solutions USA, Inc | Siemen's issued a Customer Information Letter under Update Instruction TH016/12S to alert all its affected customers of the software issues. The software release is a result of the final system testing addressing multiple performance defects and several safety issues that were not considered critical (for immediate correction) for the safe use of the device. | Software design (manufacturing process) | Siemens Medical Solutions sent a Customer Information letter via mail to all affected customers. The letter identified the affected product, problem and provide instructions on the correction. Customers were instructed to include the Customer Information letter in the Owner Manual chapter "Safety Advisory Letters".9-16-13 UPDATE: - Control Console SW 9.2.18 was released to support the rol-out10/18/2013; Distribution of update instructions for TH017/13S and TH018/13S - Distribtuion of the Control Console SW 9.2.18 released |  | 621 | 621 units (foreign account) in total | Worldwide Distribution | Terminated | August 28, 2014 | 731026 | 621 | Computer
Z-1031-2013 | 64497 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Prostiva RF Therapy Generator. |  |  | February 27, 2013 | 2013 | Medtronic Neuromodulation | Medtronic has found through device testing that if the optional footswitch is not connected to the Model 8930 RF generator during use, an electrostatic discharge (i.e. static electricity) to the footswitch connector port may toggle the RF output on or off. If this were to occur, it could result in unintentionally starting or stopping the lesion process, depending on the current status of the equi | Nonconforming Material/Component | Medtronic sent a "Urgent Medical Device Correction" letter dated January 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Medtronic has found through device testing that if the optional footswitch (Model 60883) is not connected to the Model 8930 RF generator during use, an electrostatic discharge (i.e. static electricity) to the footswitchconnector port may toggle the RF output on or off. If this were to occur, it could result in unintentionally starting or stopping the lesion process, depending on the current status of the equipment. The status messages on the display screen and associated audible alarms operate as intended, and would notify the clinician of any change in status as described in chapter 2 of the system manual.Enclosed you will find an insulating cap that should be placed over the footswitch connector port when the Model 8930 RF Generator is used without the footswitch. The footswitch connector port is located on the rearpanel of the Model 8930 RF Generator as shown below.For further questions In the United States, contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 or your local product representative. Outside of the United States, contact your local product representative. |  | 848 | 848 (696 US, 152 OUS) | Nationwide Distribution including Puerto Rico and all States in continental USA except ID, MT, and NH. | Terminated | November 21, 2013 | 267 | 848 | Not_Computer
Z-1049-2013 | 64463 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension(R) Phosphorus (PHOS) Flex(R) reagent cartridge (DF61) |  |  | February 22, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has initiated a voluntary field action for Dimension(R) PHOS Flex(R) Reagent due to reagent interaction causing falsely depressed Creatinine results when EZCR is processed from open wells that are in close proximity to open wells of PHOS reagent. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action to be taken by Customer:1. Customers with two or more Dimension systems should perform testing of PHOS and EZCR on separate Dimension systems.2. Customers with only one Dimension system should discontinue use of either PHOS or EZCR on the Dimension system. Please contact your local sales representative to discuss alternate solutions that will best meet the needs of your laboratory.Please complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at (302) 631-8467, to indicate you have received the information. If you have any questions, please contact the Siemens Technical Solutions enter or your local Siemens technical support representative. We apologize for the disruption this situation will cause your laboratory. We are working diligently to resolve this issue with a permanent solution.. |  | 60674 | 60,674 | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. | Terminated | January 29, 2015 | 706 | 60674 | Not_Computer
Z-1053-2013 | 64528 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | LEEP REDIKIT Product No. 6060 |  |  | February 26, 2013 | 2013 | Cooper Surgical, Inc. | LEEP RediKit Products (for use in Loop Electrosurgical Excision Procedures) have Xylocaine (Lidocaine) HCL (2%) listed on the tray lid label but this is not included in the contents of the kit, which contained Polocaine (Mepivacaine) HCL (3%) instead (which was not listed on the tray label). | Labeling mix-ups | CooperSurgical sent a Recall Notification dated February 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CSI intends to track and quarantine returned product in a segregated location at its manufacturing facility to ensure its separation from acceptable product and prevent its misuse. CSI intends to render used product useless by physical destruction and to rework unused product by correcting the labeling for redistribution.If you have any one of these affected lot numbers in your current inventory, please contact our Customer Service Department at 800-243-2974, or simply complete the attached form and fax to 800-262-0105. Arrangements will be made to replace product in your possession at CooperSurgical's expense. Further instruction and disposition of product will be communicated to you at that time.CooperSurgical deeply regrets the inconvenience that this problem has caused. |  | 4395 | 4395 kits | Nationwide Distribution including Puerto Rico | Terminated | June 12, 2014 | 471 | 4395 | Not_Computer
Z-1050-2013 | 64386 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX Micros IM2, Part# M60CSB110EN04 and D00A00202 |  |  | February 15, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling ABX Micros IM2 Data Management System because it may display incorrect patient test results under unanticipated operating conditions. | Device Design | Horiba Medical sent an Urgent Safety Notice dated February 15, 2013, to all affected customers. The letter informed the customers that Horiba Medical is informing all ABX Micros IM2 Data Management System customers of their intention to remove the ABX Micros IM2 product from the market. The letter informed the customers of the problem identified and the action to be taken. Customers were instructed to complete and return the enclosed response form within 10 days.For questions regarding this recall call 949-453-0500, ext 208. |  | 180 | 180 units | Nationwide Distribution | Terminated | August 11, 2014 | 542 | 180 | Not_Computer
Z-0993-2013 | 64370 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas(R) 2020 Insulin Infusion Pump |  |  | January 03, 2013 | 2013 | Animas Corporation | Animas has identified a component issue affecting a small supply of the Animas(R) 2020 insulin pumps. The component issue may trigger the pumps to sound a false alarm or warning related to one of the following: - Loss of prime; -Occlusion; or -No cartridge detected. | Nonconforming Material/Component | Animas sent an Urgent Recall letter dated December 17, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that Animas will send a replacement pump free of charge, along with instructions for returning their pump. For questions or concerns contact at 866-796-6373. |  | 7157 | 7,157 (494 US pumps affected) | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and Puerto Rico and the countries of: Australia, Austria, Bermuda, Canada, Czech Republic, Finland, France, Germany, Ireland, Israel, Italy, Mexico, New Zealand, Norway, Spain and the United Kingdom. | Terminated | May 14, 2015 | 861 | 7157 | Not_Computer
Z-1056-2013 | 64494 | No | FMH | Specimen transport and storage container. | Pathology | General Hospital | 510(K) Exempt | BD Vacutainer C&S; Transfer Straw Kit C&S; Preservative Plus Urine Tube. 4.0 ml, 13 X 75 mm STERILE REF 364953. Becton Dickinson &amp; Company, Franklin Lakes, NJ USA, |  |  | February 14, 2013 | 2013 | Becton Dickinson &amp; Company | The BD Vacutainer Urinalysis Transfer Straw Kit #364991, Lot 1335484 has preprinted polybags which are incorrectly labeled as Catalog #364953. BD advises any Customer who purchased this Kit that the Tube in this Kit is not a C&S; tube. If a Customer relies solely on the polybag labeling and uses the Tube inside the Kit for the purpose of collecting urine for culture and sensitivity, it is possible | Employee error | BD Diagnostics sent a "Product Advisory Notice" dated February 14, 2013 via UPS Second Day Air to all end user customers. The notice advises customers of the product issue, identifies affected product and provides instructions on how to handle affected product.For questions regarding this recall call 201-847-5033. |  | 236400 | 236,400 units | Nationwide Distribution including Arizona, California, Colorado, Georgia, Florida, Illinois, Iowa, Kentucky, Louisiana, Maryland, Massachussets, Michigan, Minnesota, Missouri, New Jersey, New Mexico, New York, North Carolina, Ohio, Oregon, Pennsylvania, Tennessee, Texas, Utah, Virginia, Washington DC, Washington, West Virginia, and Wisconsin. | Terminated | August 06, 2013 | 173 | 236400 | Not_Computer
Z-1054-2013 | 64704 | No | KFX | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | LeMaitre Stent Guide: 1102-00 (100 strips); 1102-50 (50 strips); 1102-20 (20 strips). |  |  | March 18, 2013 | 2013 | LeMaitre Vascular, Inc. | Devices were not sealed correctly during the manufacturing process, and the sterility of these products have been compromised. | Packaging process control | LeMaitre sent an "URGENT:VASCUTAPE RADIOPAQUE TAPE DEVICE FIELD SAFETY NOTICE" dated March 20, 2013 to affected customers. The recall letter contains a form that requests customers to return to LeMaitre Vascular, Inc. as a record of notification and reconciliation. The letter provides instructions on how the customer can inspect the defective tape seals and/or return the affected products for replacement/credit. |  | 6970 | 6970 | Worldwide Distribution-USA (nationwide) and the countries of Canada, EU, and Asia. | Terminated | December 03, 2014 | 625 | 20530 | Not_Computer
Z-1059-2013 | 64537 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | Test Strips packaged in the OC-Light Manual iFOBT Kit, Catalog No. FOB50. |  |  | February 12, 2013 | 2013 | Polymedco, Inc | The test strips in the affected product lots of the OC Light immunological fecal occult blood test are being recalled because they may not be stable through their labeled shelf-life. There may be a possibility of obtaining false negative results when using the affected lots. | Process control | Polymedco initiated phone calls to clients on 2/12/13. When the contact was made with the client, an URGENT DEVICE RECALL letter and Fax Form were sent by e-mail/fax. Mass mailing of the URGENT DEVICE RECALL letter dated February 12, 2013 and Fax Form was done by First Class USPS on 2/14/13-2/18/13. The recall affects the test strips only. For questions, please contact Polymedco Technical Support at 1-800-431-2123 ext 285. |  | 355350 | 355,350 test strips | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Canada and Guam. | Terminated | February 10, 2016 | 1093 | 584950 | Not_Computer
Z-1057-2013 | 64593 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow/Teleflex Medical Berman Angiographic Balloon Catheter and Reverse Berman Angiographic Balloon Catheter. |  |  | March 04, 2013 | 2013 | Arrow International Inc | The product specifications sheet, part of the Instructions for Use (IFU), for the specified product numbers, reference Renografin-76 contrast medium (viscosity of 8.4 centipoise), which is no longer available. Firm is in the process of updating the labeling for these products. In the meantime, when choosing an alternative contrast medium, health care providers are instructed to select a contrast | Labeling Change Control | Arrow sent an Urgent Medical Device Customer Notification letter dated March 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Arrow is in the process of updating the labeling; in the meantime they wanted to instruct all customers to select a contrast medium with a viscosity of 8.4 centipoise or lower for use with the pressure and flow settings specified in the current IFU. In addition, they are placing an insert card in kits for the specified product codes alerting them that Renografin-76 is no longer available and to select a contract medium with a viscosity of 8/4 centipose or lower for use with the pressure and flow setting specified in the current IFU.Arrow remains committed to providing high quality, safe and effective products. If you require additional information or clarification regarding this matter, please contact your Sales Representative or Arrow's Customer Service Department at (866) 396-2111. |  | 99627 | 99627ea | Worldwide Distribution - USA including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. and the countries of Australia, Dominican Republic, Uruguay, Ecuador, Chile, Taiwan, Pakistan, Colombia, Argentina, Korea Republic, New Zealand, Guatemala, Thailand, Peru, Bahamas, Philippines, Trinidad and Tobago, Japan, Ireland, India, Singapore, China, Canada and Mexico. | Terminated | December 21, 2013 | 292 | 99627 | Not_Computer
Z-1051-2013 | 64189 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBedÂż and Giraffe Incubator p/n 6600-0723-700 (rev A and B) labeled in part: ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 U.S.A. www.gehealthcare.com***. |  |  | January 08, 2013 | 2013 | GE Healthcare, LLC | Neonatal beds and incubator display boards may randomly overwrite stored information and radiant heater and oxygen set points upon power up or recovery from a power failure. | Process control | GE Healthcare sent an "Urgent Medical Device Correction" letter, dated January 8, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were given the safety issue and instructions for continued use of the Giraffe OmniBed and Incubators, as well as given the affected product details. GE Healthcare will correct all affected Giraffe Incubators and Giraffe OmniBeds at no cost, and will contact customers to arrange for this correction. If customers have any questions or concerns regarding this notification, please call: United States: 1-800-345-2700. For other countries, please contact your local GE Healthcare Service Representative. |  | 679 | 679 units | Worldwide Distribution -- USA in the states of AR ,CA, CO, DC, FL, GA, HI, ID, IL, IN, KS, MA, MD, MI, MO, MS, NC, NE, NJ, NY, NV, OH, PA, TN, TX, VA, WA, WI, and WV, and the countries of Arab Emirates, Albania, Australia, Belgium, Brazil, Canada, Chile, China, Croatia, Czech., Denmark, Ecuador, Estonia, Finland, France, Germany, Hungary, Italy ,Korea, Kurdistan, Lebanon, Macedonia, Malaysia, Mexico, Norway, Oman, Poland, Russia, San Salvador, Saudi Arabia, Singapore, South Africa, Sweden, Turkey, Uruguay, Venezuela, and United Kingdom. | Terminated | July 10, 2013 | 183 | 679 | Not_Computer
Z-1062-2013 | 64200 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Flexible Grip which is part of the Synthes Universal Nail System. |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary recall of the Flexible Grip (Part number 355.28) which is part of the Universal Nail System, due to the potential lack of the ability to thoroughly clean the instrument. | Device Design | Synthes sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" dated January 11, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter requested consignees examine inventory, remove the recalled products from use and return them to Synthes. If you have any questions, contact the firm at 610-719-5450. |  | 328 | 328 | Nationwide Distribution including the states of AZ, CA, CO, FL, IA, ID, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, TN, TX, UT, VT, WA, WI, and WV. | Terminated | August 31, 2015 | 962 | 328 | Not_Computer
Z-1063-2013 | 64420 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Covidien Endo GIA AutoSuture Universal Loading Unit, 30mm-2.0mm |  |  | February 11, 2013 | 2013 | Covidien LP | Single Use Loading Unit contained two staples loaded in each cartridge pocket and may result in difficulty firing and removing the device from the application site, which may require medical intervention | Process control | Covidien notified customers with an Urgent Medical Device Recall notification on February 11, 2013. The letter identified the affected product and the reason for the recall. Accounts were requested to quarantine and discontinue use of the affected devices, and notify Universal Loading Unit 30mm-2.0mm users of this recall. Customers should follow the instructions provided in order to return the affected product. Customers are to complete and return the Product Return Form regardless if they have any affected product or not. For questions regarding the RGA I return process, please contact Covidien Customer Service, M-F, 8 am - 6:30 pm ET at (800) 962-9888, option 1, and then option 2. |  | 6751 | 6,751 units | Worldwide Distribution -- USA, Asia Pacific, Canada, Europe, and Japan. | Terminated | March 20, 2014 | 402 | 6751 | Not_Computer
Z-1061-2013 | 64655 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Baseplate Impactor Extractor TRIATHLON Instruments. Howmedica Osteonics Corp. 325 Corporate Drive, Mahwah, NJ 07430 USA A subsidiary of Stryker Corporation. Made in USA. |  |  | February 23, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received reports from the field indicating that the handle detached from the cam of the Baseplate Impactor/Extractor Assembly. | Nonconforming Material/Component | Stryker sent Notification Letters/Product Accountablility Forms dated March 1, 2013 via FedEx to their Branches/Agencies on March 1, 2013.. Notification Letters/Product Recall Acknowledgement Forms dated March 4, 2013, were sent via FedEx to Hospital Risk Management on March 4, 2013. The Notification letters identified the product, the problem, and instructed customers on how to identify affected product; informs customers of the issues and potential hazards with the affected product and instructs them to contact their representative to arrange for return/replacement of any affected product. Customers are also asked to fax back the provided "Urgent Medical Device Removal Notification Acknowledgement Form" to Stryker within 5 days of receipt.Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5523. |  | 871 | 871 units | Worldwide Distribution - USA (nationwide) and internationally to Canada, Japan, Australia, China, New Zealand, France, Italy, Netherlands, Poland, Spain, Sweden, Switzerland and the United Kingdom. | Terminated | June 15, 2015 | 842 | 871 | Not_Computer
Z-1064-2013 | 64527 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow International FiberOptix Intra-Aortic Balloon Catheter Kit (IAB) 8 Fr 30 cc; |  |  | November 27, 2012 | 2012 | Arrow International Inc | Potential to have an internal leak at the fill port of the stainless steel tip, which may necessitate removal and replacement ofthe catheter. | Process control | Arrow International issued Urgent Medical Device Recall notification on 11/27/12. The letter identified the affected product and the reason for the recall. Accounts are requested to check their stock for the products included within the scope of this recall. Also, customers should cease use and distribution, and quarantine all affected product immediately. Contact Arrow's Customer Service Department at 800-523-8446 for a Return AuthorizationNumber. Customers are to complete and fax the Recall Acknowledgement &amp; Stock Status Form to the number provided, and return the affected product. If customers have questions or need clarification regarding this issue, they should contact Arrow's Customer Service Department at 1-800-523-8446. |  | 188 | 188 units | Worldwide Distribution -- USA, Belgium, Great Britain, Italy, Kazakhstan, Latvia, Russia, Slovenia, Turkey, and South Africa. | Terminated | June 02, 2015 | 917 | 825 | Not_Computer
Z-1066-2013 | 64505 | No | OQG | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Cross-over Acetabular Shell and Liner Hip System, Part Numbers SC3458-52, SC3474-52, SC3459-52, and SC3460-52. |  |  | January 24, 2013 | 2013 | Stelkast Co | Firm was made aware of incidents relating to the difficulty of properly engaging and securing the Cross-Over 36mm liner into the Cross-Over shell. | Device Design | The firm initiated their recall of this product on January 24, 2013 by notifying their consignees by telephone informing them of the recall. |  | 254 | 254 | Nationwide Distribution including the states of NV, OK and PA. | Terminated | October 08, 2014 | 622 | 254 | Not_Computer
Z-1070-2013 | 64656 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Reporting Cardiac Catheterization module and IMPAX CV Outbound Reporting. |  |  | March 15, 2013 | 2013 | AGFA Corp. | Baseline pulmonary capillary wedge (PCW) pressure values from IMPAX CV Reporting Cardiac Catheterization module were not correctly transferring into the hospital's electronic health record (EHR) system. | Software design | AGFA sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013 to the consignees who use the IMPAX CV Reporting tool described for this issue. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 877-777-2432, referencing HQ_10020006 for questions regarding this notice. |  | 26 | 26 | Nationwide Distribution including the states of AK, GA, IL, KY, MA, NY, NC, RI, SC, TX, and VA. | Terminated | January 09, 2014 | 300 | 26 | Not_Computer
Z-1067-2013 | 64667 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-IOI: AQUILION 32/64; System, X-Ray, Tomography, Computed. |  |  | November 30, 2011 | 2011 | Toshiba American Medical Systems Inc | The firm initiated this recall due to a potential software issue. The DLP (Dose Length Product) value may be displayed incorrectly when the vHP option is used. | Process design | Toshiba sent "URGENT: MEDICAL DEVICE CORRECTION" letters to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The letter contained a return reply form that is to be faxed to TAMS for retention. Contact the firm at (800) 421-1968 for questions relating to this letter. |  | 23 | 23 units | Nationwide Distribution including the states of WA, OH, CA, TX, WV, IL, AZ, CA, WI, ME, WI, NC, FL, MT, and HI. | Terminated | June 17, 2013 | 565 | 23 | Computer
Z-1069-2013 | 64654 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV 7.8 SU3 - OCR Service. |  |  | March 15, 2013 | 2013 | AGFA Corp. | Software design error in IMPAX CV 7.8 SU3 | Software design | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to distribute this information within your facility to all individuals who need to be aware of this. Please fax back to (864) 421-1664 that the information was received and understood was requested from the customers.If you have any questions about this matter please contact your local AGFA Healthcare organization, or call (864) 421-1754. |  | 2 | 2 | Distribute in the state of FL and KY | Terminated | April 01, 2014 | 382 | 2 | Computer
Z-1068-2013 | 64673 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba TSX-30IC AQUlLlON ONE System, Computed Tomography X-Ray. This system supports Whole body scanning. |  |  | November 14, 2012 | 2012 | Toshiba American Medical Systems Inc | Due to a problem with the control software of the x-ray high-voltage generator, if scanning is performed with a tube current less than 30mA, the control software may detect an error and perform error processing even though the x-ray output is normal. | Process design | Toshiba American Medical Systems Inc. sent a Urgent Medical Device Correction letter dated November 14, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer was informed a New software to prevent the occurrence of this issue will be installed on your system. Toshiba Service Representative will contact you for an appointment to install the New Software on your Aquilion One System.Please fax the return reply form to (877) 349-3054.We thank you for your urgent attention to this matter. If you have any further questions regarding this letter please call (800) 521-1968. |  | 1 | 1 unit | Distributed in the state of MD. | Terminated | April 08, 2013 | 145 | 1 | Computer
Z-1071-2013 | 63622 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Biothane Restraint Cuffs, Catalog Number 2900. |  |  | October 31, 2012 | 2012 | J T Posey Company | The firm voluntarily recalled Biothane Restraint Cuffs, Catalog No. 2900, that did not meet the required specifications for tensile strength. Failure of these rivets could allow the cuff to become disconnected from the connecting strap and allow patients to regain free use of the limb. | Process design | J. T. Posey sent an Medical Device Recall (Removal) letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed the customer that Posey Company is voluntarily recalling all units of Posey BiothaneQD Cuffs, Catalog Number 2900. All lots of this product are being recalled. Reason for recall was listed:"Posey has identified rivets on the Biothane~ Restraint Cuffs, Catalog No. 2900, that did not meet therequired specifications for tensile strength. Failure of these rivets could allow the cuff to become disconnected from the connecting strap and allow patients to regain free use of the limb. No patient injuries or deaths have reported."Please immediately discontinue further distribution or use of Posey Biothane~ Cuffs, Catalog Number 2900.Please Complete and return the Acknowledgement and Receipt Form to Posey within 7 days using the provided self-addressed, stamped envelope, by fax to 626-443-5418 or via e-mail to RegulatoryAffairs@Posey.com. Return this form even if you do not have any devices in your facility.Please return directly to Posey ill! units you have of Biothane~ Cuffs, Catalog Number 2900, using the enclosed pre-paid shipping label containing an RMA number. Posey will provide replacement product or credit upon receipt of the product.If you have any further questions please call (626) 443-3143. |  | 1673 | 1673 units | Nationwide Distribution | Terminated | July 17, 2014 | 624 | 1673 | Not_Computer
Z-1073-2013 | 63804 | No | GGW | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade(R) Actin(R) FSL Activated PTT Reagent. |  |  | March 11, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has observed that the listed lot numbers show an increase in Heparin sensitivity over the shelf life which is demonstrated by testing with Dade(R) Citrol(R) Heparin Control, low and high. | Unknown/Undetermined by firm | Siemens sent an "URGENT FIELD SAFETY NOTICE" dated February 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. The safety notice instructed consignees to discontinue use of the affected lots and discard any remaining inventory. Contact the firm at 1-800-242-3233, option 1 for questions about this notice. |  | 13083 | 13,083 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and Puerto Rico and the country of Canada. | Terminated | December 24, 2013 | 288 | 13083 | Not_Computer
Z-1072-2013 | 64664 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table. |  |  | March 14, 2013 | 2013 | Elekta, Inc. | The Column rotation can accidentally be moved to a different position which in the locked position. | Device Design | Elekta sent an "IMPORTANT FIELD SAFETY NOTICE" dated March 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Elekta office for questions about this notice. |  | 443 | 443 | Nationwide Distribution and Puerto Rico. | Terminated | March 17, 2015 | 733 | 443 | Not_Computer
Z-1082-2013 | 64427 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK CR Plus and LIFEPAK EXPRESS defibrillators. Product part's numbers: 3200731-xxx, 3202177-xxx, U3200731-xxx, and U3202177-006. |  |  | February 08, 2013 | 2013 | Physio Control, Inc. | LIFEPAK CR Plus or LIFEPAK Express Automated External Defibrillators (AEDs) are recalled due to the potential to experience an early internal battery depletion issue. | Component design/selection | The firm start sending the "URGENT MEDICAL DEVICE CORRECTION LIFEPAK CR Plus AND LIFEPAK EXPRESS AEDs" letter, dated February 2013, to their consignees.Consignees are advised to check the readiness of the AEDs and should contact Physio if the OK symbol is NOT visible and either the ATTENTION or WRENCH symbols are present.In the event a device is returned to Physio-Control for assessment, the firm will analyze it to determine if it is affected by the early internal battery depletion issue. If it is verified the device is affected, the consignee will be provided with a permanent replacement device.Customers with questions can visit the firm's website www.physio-control-notices.com/CHARGE-PAK or call Technical Support at 1-800-442-1142. |  | 54197 | 54,197 units in US and 78,627 units internationally (total 132, 824 units) | Worldwide Distribution. | Terminated | November 05, 2015 | 1000 | 54197 | Not_Computer
Z-1058-2013 | 64751 | No | BSJ | Anesthetic gas mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Respironics SealFlex Single Port Ribbed Mask (Infant). |  |  | September 12, 2011 | 2011 | Engineered Medical Systems,Inc | The mask inside the package has been identified as "Neonatal" in size and not "Infant" as the labeling indicates. | Process change control | Consignees were notified 9/12/2011 with a Medical Device Voluntary Recall letter identifying the product and the problem. Consignees were directed to quarantine, cease further distribution, and destroy the affected product, returning a proof of confirmation/disposal letter, to be submitted for credit. Correspondence should be sent to Tamara Lefevers, Quality Manager, at tlefevers@ engmedsys.com or fax 317-246-5501. |  | 1595 | 1500pcs, 75cs {20pcs per cases) | Distributed in the states of PA, IN, NY, CO, and CA. | Terminated | August 02, 2013 | 690 | 1595 | Not_Computer
Z-1079-2013 | 64418 | bath, incubators/water, all | JTQ | Microbiological incubator. | Microbiology | Microbiology | 510(K) Exempt | CO2 Incubators (General Purpose Incubator), model numbers 370, 371, 380, 381, 3307, 3308, 3310, 3311, 3540, 3541, 3542, 3543, 3584, 3585, 3598, and 3599. |  |  | November 29, 2012 | 2012 | Thermo Fisher Scientific (Asheville) LLC | Incubators manufactured prior to mid 2009 have coated metal hinges that have reportedly failed in the field. | Component design/selection | Thermo Fisher sent an Urgent Field Correction Notice letter dated November 19, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate all affected products and determine if it is within the recall. If so, they are to return the Acknowledgement Form in order to get replacement hinges and instructions. Upon receipt of the new hinges they are to replace them according to the instructions. |  | 15567 | 15,567 incubators | Worldwide Distribution - USA (nationwide) including the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SN, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY. and the countries: AFGHANISTAN, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, COSTA RICA, DENMARK, EGYPT, ENGLAND, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KOREA, LEBANON, MALAYSIA, MEXICO, NEPAL, NETHERLANDS, NEW ZEALAND, NORWAY, PALESTINE, PEOPLES REPUBLIC OF CHINA, PERU, PHILIPPINES, PORTUGAL, QATAR, REPUBLIC OF PANAMA, ROMANIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, VENEZUELA, VIETNAM &amp; WEST INDIES. | Terminated | April 14, 2015 | 866 | 15567 | Not_Computer
Z-1077-2013 | 63974 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ELEKTA MOSAIQ SW version below 2.41 |  |  | December 20, 2012 | 2012 | Siemens Medical Solutions USA, Inc | A potential safety risk with the MOSAIQ OncologyInformation System that has been identified by the manufacturer IMPAC Medical Systems, Inc. | Software design | Siemens sent an Urgent Medical Device Correction letter with update instructions to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your responsible Siemens Service Engineer or Application Specialist. |  | 101 | 101 | Worldwide Distribution | Terminated | July 19, 2013 | 211 | 101 | Not_Computer
Z-1093-2013 | 64776 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | LHB, Perishable Items Pack Type IV. Kit number: 2114016-100. |  |  | March 22, 2013 | 2013 | Lighthouse For The Blind | The product contains an iodine 10% Ampule (10% Povidone Iodine swab) which is being recalled by the firm's distributor. The labeling of the iodine 10% Ampule labels state that the contents of the package are sterile and does not clarify that this sterility claim is only applicable to the applicator. The antiseptic solution that is applied to the skin is a cutaneous disinfectant, and is not sterili | Labeling False and Misleading | The firm sent "URGENT DEVICE RECALL" letters dated March 22, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-542-3697 for questions regarding this notice. |  | 425 | 425 kits | Nationwide Distribution including DC and the states of AK, AL, AR, AZ, CA, CO, DE, FL, GA, ID, IL, IN, KS, KY, MI, MD, MN, MO, MT, NC, NJ, NM, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA and WY. | Terminated | November 25, 2013 | 248 | 108939 | Not_Computer
Z-1085-2013 | 64754 | No | OYC |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Paradigm Insulin Infusion Pumps |  |  | March 22, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic Paradigm Insulin Pump because the pump's drive support cap may become detached from the pump case and protrude from the lower right side of the pump. When the pump is expose to water it may result in damage to the pump's internal electronics. This moisture damage can prevent the pump's buttons from working properly or can cause the pump to alarm. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Safety Notification letters to all affected customers. The recall letter informs the customers of the problem identified and the actions to be taken. The letter informs the healthcare professionals that Medtronic will inform their patients about the potential safety issues. Customers are instructed to find online information at www.medtronicdiabetes.com/support/product-updates. Customers are instructed to call Medtronic at 1.888.204.7616 (option 1 for loose drive support cap related questions and option 2 for water damage related questions). Distributors are instructed to provide Medtronic with an Excel spreadsheet list of all end users to whom they have shipped any of the listed serialized products. |  | 428000 | 428,000 units | Worldwide Distribution - USA (nationwide) the countries of: AE, AR, AT, AU, BA, BH, BM, BR, BS, CA, CH, CL, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HR, HU, IE, IL, IN, IQ, IT, JO, JP, KR, KW, LB, LC, LV, MK, MX, MY, NL, NO, NZ, OM, PA, PE, PL, PR, PY, QA, RO, SA, SE, SG, SI, SK, SV, SY, TF, TH, TR, TT, TW, UA, UY, VE, ZA. | Terminated | November 04, 2013 | 227 | 428000 | Not_Computer
Z-1083-2013 | 64759 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | VMX, VMX Plus Mobile Radiographic Unit with the following models: |  |  | January 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware of an issue associated with the base casting for the articulated arm of your VMX andVMX Plus Mobile X-ray systems that may impact Operator or Patient safety. | Device Design | GE Healthcare sent and "Urgent Medical Device Correction" letter dated January 23, 2012. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details and Product Correction. For questions contact your local GE Healthcare Service Represenative. |  | 1296 | 1296 | Worldwide Distribution - USA (nationwide) including the states of: AR, CA, GA, IL, IN,IA,LA, MA, MN, MT, NY, OH, PA, TX, VA and the countries of BELGIUM, BURKINA FASO, BRAZIL, CHINA, CANADA, SWITZERLAND, CHILE, COLOMBIA, CYPRUS, CZECH REPUBLIC, GERMANY, DENMARK, DOMINICAN REPUBLIC, EGYPT, SPAIN, FRANCE, UNITED KINGDOM, CHANA, GUADELOUPE, GREECE, GUYANA, HONG KONG, CROATIA, HUNGARY, INDONESIA, ISRAEL, ITALY, KUWAIT, LEBANON LUXEMBURG, LIBYAN, MOROCCO, MARTINIQUE, MEXICO, NETHERLANDS, NORWAY, PANAMA, PERU, POLAND, PORTUGAL , REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SWEDEN, SINGAPORE, SLOVENIA, SLOVAKIA, SIERRA LEONE, THAILAND, TUNISIA, TURKEY, VENEZUELA, MACEDONIA, PALESTINE, SOUTH AFRICA. | Terminated | February 04, 2014 | 743 | 1296 | Not_Computer
Z-1076-2013 | 64559 | No | MZW | Dental cement. | Dental | Dental | 510(k) | Kerr Maxcem Elite, self-etch/self-adhesive resin cement; Intro Kit Part No. 34417; DMC 0446 Rev 0. |  |  | February 22, 2013 | 2013 | Kerr/Pentron, DBA Kerr Corporation and Pentron Clinical | Kerr Corporation is initiating this recall because raw material used to produce the Maxcem Elite is causing the material to prematurely polymerize in the syringe. The product could be difficult to extrude or demonstrate a shorter work-time as the material may prematurely set prior to placing restorations on the tooth or core build-up. As a result, the restoration may not be seated properly. | Nonconforming Material/Component | The firm, Kerr Corporation/Sybron Dental Specialties, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 22, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers are instructed to determine if you have any of the affected products in your inventory; If you have the affected lots, please contact Kerr Customer Care at 1-800-537-7123 to receive an RMA number for quick return; complete and return the Maxcem Elite Acknowledgement and Recall Return Form via fax to 714-288-4609 regardless of whether you have any product in your inventory. If you are an authorized Kerr Corporation distributor, identify and contact those customers that may have been shipped the affected product within 48 hours of receipt of this field correction notification.If you have any question, contact Kerr Customer Care at 1-800-537-7123. |  | 40819 | 40,819 kit and syringe units | Worldwide distribution: USA (nationwide) and countries of: Australia, Canada, Belgium, Czech Republic, France, Great Britain, Poland, Albania, Austria, Bulgaria, Switzerland, Cyprus, Germany, Denmark, Estonia, Spain, Finland, Croatia, Hungary, Ireland, ITtaly Luxembourg, Latvia, Netherlands, Norway, Portugal, Romania, Serbia (Republic of Serbia), Sweden, Slovenia, Turkey, Ukraine, United Arab Emirates, Congo, The Democratic Republic of the (formerly Zaire), Kazakhstan, Lebanon, Russian Federation, South Africa, Mexico, Thailand, Brazil, China, Hong Kong, Israel, India, Japan, Malaysia, Panama, Singapore, Viet Nam. | Terminated | January 22, 2014 | 334 | 40819 | Not_Computer
Z-1078-2013 | 64398 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Quasar Nuclear Medicine System, Hawkeye Option and Hawkeye 4 Option for Dual-Head Variable Angle Gamma Camera. Nuclear Medicine imaging systems. Recall includes all Infinia systems, all configurations. |  |  | March 06, 2012 | 2012 | GE Healthcare, LLC | Accelerated fatigue of the lateral axis motion subsystem of the Infinia Computed Emission Tomography System could result in mechanical failure causing the detector to slide; may impact patient and operators. | Process control | GE Healthcare sent an " IMPORTANT PRODUCT INFORMATION" letter dated March 6, 2012. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers and Managers of Radiology/Cardiology. The letter described the Issue, instructions, Affected Product Details, and Product Correction. For questions contact your local GE Healthcare Service Representative or call the GE Healthcare Call Center: United States or Canada: 800-437-1171 and Other Countries 262-896-2890. |  | 2180 | 2180 | Worldwide Distribution - USA (nationwide) including DC and PR and the countries of UKRAINE, TAIWAN, TURKEY, TUNISIA, THAILAND, SINGAPORE, SWEDEN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, POLAND, PAKISTAN, PHILIPPINES, PANAMA, NEW ZEALAND, NORWAY NETHERLANDS, MALAYSIA, MEXICO, MALTA, MARTINIQUE, MOROCCO, LATVIA, LITHUANIA, LEBANON, KUWAIT, KOREA, JAPAN, JORDAN, ITALY, IRAN, INDIA, ISRAEL, IRELAND, INDONESIA, HUNGARY, CORATIA, HONG KONG, GREECE, GEORGIA, UNITED KINGDOM, FRANCE, FINLAND, SPAIN, EGYPT, ECUADOR, ALGERIA, DOMINICAN REPUBLIC, DENMARK, GERMANY, CZECH REPUBLIC, CYPRUS, COLUMBIA, CHILE, SWITZERLAND, CANADA, CHINA, BRAZIL, BELGIUM, AUSTRALIA, AUSTRIA, ARGENTINA, SOUTH AFRICA, VENEZUELA, URUGUAY, ALGERIA. | Terminated | January 27, 2014 | 692 | 2180 | Not_Computer
Z-1081-2013 | 64855 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens brand ACUSON SC2000 Ultrasound System, Model 10433816, hardware that contains version 4 of the MPI board 10434200; |  |  | March 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | When using the V5M transducer, the display of the temperature on the ACUSON SC2000 system can be lower than the actual temperature of the V5M transducer. This occurs only if the system hardware contains revision 4 of the MPI Board. There is a potential to cause esophageal burns in patients. | Device Design | Siemens Medical Solutions USA Inc. sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A return response form is included for customers to complete and mail to the recalling establishment. Siemens stated additional follow up action will occur, once material is available, where Siemens field service engineers will visit each unit and replace the affected MPI 4 board.For further questions please call (650) 694-5747. |  | 120 | 120 units | Worldwide Distribution - USA and countries of Australia, Austria, Bosnia Herzegovina, Belgium, Finland, France, Germany, India, Italy, Japan, Kuwait, Malaysia, Netherlands, Norway, P.R. China, Portugal, Rep.of Korea, Romania, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan and United Kingdom. | Terminated | February 14, 2014 | 323 | 120 | Not_Computer
Z-1075-2013 | 56584 | No | JPD | Hemoglobin A[bdi2] assay. | Hematology | Hematology | 510(k) | G7 HPLC. |  |  | August 18, 2010 | 2010 | Tosoh Bioscience Inc | The Beta-thalessemia Mode of the G7 Analyzer-Beta Thalassemia Mode is an automated High Performance Liquid Chromatography (HPLC) system that separates and reports HbF and HbA2 quantitative percentages in whole blood. The firm received complaints that the A2 peak could not be detected with the G7 Beta-thalessemia columns. | Component change control | Please be aware that this is not a new recall. The firm has taken action. The firm stated that they have notified consignees via mail, with the list of notified customers maintained by the Regulatory Affairs/Quality Assurance department. The firm stated that they have notified distributors in the same way that they have notified other consignees, with a required confirmation to the firm of corrective actions. |  | 633 | 633 | Worldwide Distribution-USA (nationwide) including DC, Puerto Rico, Virgin Islands and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV, and the countries of Canada, Columbia, Chile, Brazil, and Jamaica. | Terminated | May 13, 2014 | 1364 | 633 | Not_Computer
Z-1084-2013 | 64749 | No | DZD | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Bracket Buccal Tube U7L DB/UL22 Roth MT HK 0/-14/0 |  |  | March 08, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one (1) lot of Bracket Buccal Tubes as a result of a packaging error. More specifically, the BracketBuccal Tubes product package correctly identifies the product as gingival (0Âż) offset brackets, butactually contains distal (10Âż) offset brackets. Use of the affected product may cause unintended rotationof the molars. Although reversible, a delay in | Packaging | Ormco sent an Urgent Medical Device Recall letter dated March 8, 2013 to all affected customers, via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. The consignees were instructed to complete the Return Form and return any affected product identified in stock. For questions contact Ormco Customer Care. |  | 8210 | 8210 | Worldwide Distribution - USA (nationwide) including the country of EU. | Terminated | September 04, 2013 | 180 | 8210 | Not_Computer
Z-1100-2013 | 64763 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750w 3.0T System. |  |  | April 23, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to image flip in frequency encoding direction associated with a specific workflow of your Discovery MR750w system.Image flip issue in frequency encoding direction (including S/I, A/P and R/L) can result in incorrect localization ofanatomy and pathological findings. | Software design | GE Healthcare sent an Important Product Information letter dated April 23, 2012. The letter described the Issue, Instructions, Affected Product Details and Product Correction. For questions contact your local service representative. |  | 46 | 46 | Worldwide Distribution - USA (nationwide) including the states of: AZ, CA, CO, FL, IL, MD, NJ, NY, OH, OK, TX. and the countries of Australia, Belarus, Canada,France, Germany, India, Japan, Lebanon, New Zealand, Saudi Arabia, Singapore, Spain, Sweden, Switzerland and United Kingdom. | Terminated | March 14, 2014 | 690 | 46 | Not_Computer
Z-1097-2013 | 64678 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fusion or Luminos Select. |  |  | January 15, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential malfunction and possible hazard to patients and staff when using the Iconos R200 C20, Iconos R200 T20, Iconos R100, Iconos MD, Iconos RHD, Luminos RF Classic, Luminos Fusion or Luminos Select. This malfunction is possible in cases where the table vertical lift movement is obstructed by a solid object which could potentially result in collision. This can re | Process design | Siemens sent a Customer Safety Advisory letter dated January 15th, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The SIEMENS service representative will investigate if parts need to be exchanged in your product, an if so, which parts need to be exchanged to restore the system safety and function. Please kindly confirm receipt of the Field Safety Notice by sending back the attached form.For further questions please call (610) 448-3237. |  | 51 | 51 | Nationwide Distribution including the states of CA, FL, GA, IA, IN, KS, LA, ME, MI, MO, NC, NH, NJ, NY, NV, OH, PA, TN, TX, VA and WA. | Terminated | April 23, 2013 | 98 | 51 | Not_Computer
Z-1099-2013 | 64853 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Torrent" irrigation tubing. |  |  | March 07, 2013 | 2013 | US Endoscopy Group Inc | The firm was notified by their customers that there was a water leakage from the white tubing connector. | Device Design | US Endoscopy Group sent an Urgent Recall Notification letters to their customers. The letter identified the affected product, problem and actions to be taken. They instructed their customers to remove any of the affected products from their inventory and quarantine them. They also requested each customer complete the attached Customer Response Card and return it to US Endoscopy via fax at 1-440-639-4495 or by email at recall@endoscopy.com as soon as possible. For questions contact Customer Service at 1-800-769-8229. |  | 21189 | 21,189 units | USA Nationwide Distribution - including the states of: AK, AL, AZ, CA, CO, CT, DE, FL, GA, ID, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI &amp; WV. | Terminated | March 20, 2014 | 378 | 21189 | Not_Computer